## Sheet1
公告(開)號 | KindCode | 公告(開)日 | 名稱 | 申請日 | 摘要 | 申請範圍 | 專利權人 | IPC
US09823259 | B1 | 2017-11-21 | Detection device for cannabinoid use | 2016-04-08 | A detection device for cannabinoid use including a base unit and a display screen disposed on the base unit. A central processing unit and a tetrahydrocannabinol testing device are disposed within the base unit. A testing slot is disposed on a bottom portion of the base unit, with a back end of the testing slot disposed within the tetrahydrocannabinol testing device. Each of a plurality of testing strips is slidably and removably disposed within the testing slot. A printer has a pair of T-shaped attachment extensions. Each of the pair of attachment extensions and a universal serial bus plug disposed on the printer simultaneously removably and slidably engages each of a pair of attachment slots and a universal serial bus port disposed on the base unit, respectively. The tetrahydrocannabinol testing device is configured to analyze and calculate the presence and amount of tetrahydrocannabinol in a person's bloodstream. | 1. A detection device for cannabinoid use comprising: a substantially trapezoidal handheld portable base unit having a convexly curved bottom portion, a linear top edge, a right side, a left side, a front side, and an exterior surface;a display screen disposed on the front side of the base unit proximal the top edge;an antimicrobial covering disposed on the exterior surface of the base unit;a first battery compartment disposed within the base unit;a pair of rechargeable batteries disposed within the first battery compartment;a central processing unit disposed within the base unit;a tetrahydrocannabinol testing device disposed within the base unit;a testing slot having a front end medially disposed on the bottom portion of the base unit and a back end disposed within the tetrahydrocannabinol testing device;a plurality of elongated rectangular sterile testing strips, each of the plurality of testing strips slidably and removably disposed within the testing slot;a testing activation control disposed on the front side of the base unit proximal the bottom portion;a pair of T-shaped attachment slots comprising a right attachment slot and a left attachment slot, each of the right attachment slot and the left attachment slot disposed within the top edge of the base unit proximal the right side and the left side, respectively;a universal serial bus port medially disposed within the top edge of the base unit between the right attachment slot and the left attachment slot;a handheld portable printer having a linear bottom surface, a back surface, a right surface, a left surface, a universal serial bus plug substantially medially disposed on the bottom surface, and a pair of T-shaped attachment extensions comprising a right attachment extension and a left attachment extension, each of the right attachment extension and the left attachment extension disposed on the bottom surface proximal the right surface and the left surface, respectively;a second battery compartment disposed within the portable printer; anda rechargeable battery disposed within the second battery compartment;wherein a distance between the right attachment slot and the left attachment slot is equal to a distance between the right attachment extension and the left attachment extension;wherein each of the right attachment extension, the left attachment extension, and the universal serial bus plug simultaneously removably and slidably engages the right attachment slot, the left attachment slot, and the universal serial bus port, respectively;wherein the tetrahydrocannabinol testing device is configured to analyze and calculate the presence and amount of tetrahydrocannabinol in a person's bloodstream when one of the plurality of testing strips containing a sample of one of the person's saliva and blood is disposed within the testing slot and the testing activation control is depressed by a user;wherein the display screen is configured to display the presence and amount of tetrahydrocannabinol in... | - | G01N033/94 G01N001/10 G01N033/48 A61B005/15 G01N021/75 A61B010/00 G01N033/487 G01N031/22 G01N021/01 G01N025/48 C12Q001/68 A61B005/157
US09775562 | B2 | 2017-10-03 | Method and apparatus for a smart vehicle gateway with connection context aware radio communication management and multi-radio technology | 2016-01-20 | An information handling system operating as a smart vehicle gateway and includes a wireless adapter for communicating with a wireless link and a storage device for storing a spatial-temporal user profile comprising wireless device usage trend data for a location in or near a predicted smart vehicle gateway path during a future time interval for a smart vehicle gateway. The smart vehicle gateway may operate to establish a wireless link on one of several WWAN link options as a home network via a programmable eSIM. The information handling system further includes positional detector and an application processor that determines a trajectory estimation during a future time interval. | 1. An information handling system functioning as a smart vehicle gateway comprising: a local wireless adapter to wirelessly communicate with a plurality of wireless devices proximate to a vehicle;a wireless adapter to communicate with a wireless link and provide connectivity for the plurality of wireless devices proximate to the vehicle;a storage device to store a spatial-temporal user profile comprising wireless device usage trend data for a plurality of locations where the smart vehicle gateway has operated;a positional detector to detect a location of the smart vehicle gateway;an application processor to correlate the wireless device usage trend data for a first location in or near a predicted smart vehicle gateway path during a future time interval;the application processor to determine a list of optimal wireless service carriers and available wireless protocols at the first location based on a mobile wireless traffic report with a spatial-temporal radio frequency profile indicating signal quality for the plurality of wireless links and the wireless device usage trend data for the first location along the predicted smart vehicle gateway path during the future time interval;the application processor to select an international mobile subscriber identity (IMSI) for one of the optimal wireless links from an electronic subscriber identity module (eSIM) programmable to switch between multiple available wireless service carriers; andthe wireless adapter to switch to a selected optimal wireless link on a wireless service carrier corresponding to the selected IMSI. 2. The smart vehicle gateway of claim 1 further comprising: the application processor selects a second IMSI for a second optimal wireless link determined for a second location;the application processor to reprogram the eSIM to the second IMSI upon the smart vehicle gateway reaching the second location; andthe wireless adapter to switch to the second selected optimal wireless link on the wireless service carrier corresponding to the second selected IMSI. 3. The smart vehicle gateway of claim 1 further comprising: the application processor determines a list of optimal wireless service carriers and available wireless protocols across a plurality of locations along the predicted smart vehicle gateway path;the application processor selects an IMSI for a selected optimal wireless service carrier for the plurality of locations along the predicted smart vehicle gateway path; andthe wireless adapter to switch to the selected optimal wireless link for the plurality of locations along the predicted smart vehicle gateway path. 4. The smart vehicle gateway of claim 1 further comprising: the application processor to select a second IMSI for a second optimal wireless link at the first location along the predicted smart vehicle gateway path via a second eSIM programmable to switch between multiple available wireless service carriers;a second wireless adapter to establish the second selected optimal wire... | Dell Products, LP | H04W004/00 A61B005/00 A61B005/145
US09808207 | B2 | 2017-11-07 | Detecting, assessing and managing a risk of death in epilepsy | 2016-10-25 | A method for determining and responding in real-time to an increased risk of death relating to a patient with epilepsy is provided. The method includes receiving cardiac data and determining a cardiac index based upon the cardiac data. The method includes determining an increased risk of death associated with epilepsy if the indices are extreme, issuing a warning of the increased risk of death and logging information related to the increased risk of death. Also presented is a second method for determining and responding in real-time to an increased risk of death relating to a patient with epilepsy comprising receiving at least one of arousal data, responsiveness data or awareness data and determining an arousal index, a responsiveness index or an awareness index, where the indices are based on arousal data, responsiveness data or awareness data respectively. The second method includes determining an increased risk of death related to epilepsy if indices are extreme values, issuing a warning of the increased risk of death and logging information related to the increased risk of death. A computer readable program storage device is also provided. Also provided is a method for receiving body data, determining a cardiac, an arousal, a responsiveness, or a kinetic index, determining an increased or increasing risk of death over a first time window relating to a patient with epilepsy and issuing a warning and logging relevant information. | 1. A method for determining and responding in real-time to an increased risk of death relating to a patient with epilepsy, comprising: receiving cardiac data into a medical device;determining, at least one cardiac index based upon said cardiac data;determining whether said cardiac index is an extreme cardiac index by determining whether said cardiac index is below a low cardiac threshold value or above a high cardiac threshold value;determining, using the medical device, the increased risk of death associated with epilepsy in response to a determination that said cardiac index is the extreme cardiac index; andperforming a responsive action based upon an occurrence of the increased risk of death, said responsive action comprises at least one of issuing automatically a warning; providing a therapy, or logging automatically information related to the increased risk of death. 2. The method of claim 1, wherein determining the increased risk of death relating to the patient with epilepsy comprises determining at least one of a time at which at least one cardiac index becomes the extreme cardiac index or if there is a likelihood of said at least one cardiac index becoming the extreme cardiac index. 3. The method of claim 1, wherein providing said therapy comprises providing at least one of an electrical therapy, a chemical therapy, a thermal therapy, a cardiac defibrillation, a cardiac pacing or a supportive treatment. 4. The method of claim 1, further comprising at least one of: wherein the cardiac index is selected from a group consisting of a heart rate, a heart rhythm, an EKG complex morphology, and a blood pressure;wherein the cardiac index becomes an extreme value based at least upon one of:the cardiac index is above an eightieth percentile of a plurality of reference autonomic index values;the cardiac index is below a twentieth percentile of the plurality of reference autonomic index values; orthe cardiac index is beyond two standard deviations to a right or a left of a mean for a normal or a normalized distribution of autonomic index values for the patient or a population of patients; ordetermining an additional autonomic index based upon a respiratory index, wherein the respiratory index is at least one of a respiratory rate, a respiratory rhythm, a respiratory pattern, a breath morphology, an oxygen saturation of the patient or a carbon dioxide concentration in the patient. 5. The method of claim 1, further comprising determining at least one additional index from at least one of a metabolic index, an endocrine index, a tissue stress index, a neurologic index, or a physical fitness/integrity index, said at least one additional index being based upon body data; wherein the metabolic index comprises a blood index selected from an arterial pH, a lactic acid concentration, a serum potassium concentration, a serum glucose concentration, a lactate-pyruvate ratio, at least one electrolyte concentration, and a bicarbonate concentration;wherein the... | Flint Hills Scientific, L.L.C. | A61B005/00 A61N001/365 A61N001/36 A61N001/39 A61B005/16 A61B005/0452 A61B005/0205 A61B005/08 A61B005/021 A61B005/0245 A61B005/11 G06F019/00 A61F007/00 A61M005/172 A61N001/372
US09818212 | B2 | 2017-11-14 | Magnetic resonance imaging (MRI) apparatus and method of processing MR image | 2015-10-19 | A magnetic resonance imaging (MRI) apparatus and a method of processing an MR image are provided. The MRI apparatus includes a scanner configured to acquire a first image that is a T1-weighted image and a second image that is a fluid attenuated inversion recovery (FLAIR) image by performing an MRI scan on a brain. The MRI apparatus further includes an image processor configured to determine a white matter region in the second image based on the first image and the second image, and detect a white matter hyperintensity (WMH) region in the determined white matter region. The MRI apparatus further includes an output interface configured to display the detected WMH region and a change in the WMH region over time. | 1. A magnetic resonance imaging (MRI) apparatus comprising: a scanner configured to: acquire, at a first time point, a first image that is a T1-weighted image and a second image that is a fluid attenuated inversion recovery (FLAIR) image by performing a first MRI scan on a brain; andacquire, at a second time point, the first image and the second image by performing a second MRI scan on the brain;an image processor configured to: determine a white matter region in the second image, based on the first image and the second image; anddetect a white matter hyperintensity (WMH) region in the determined white matter region; andan output interface configured to display the detected WMH region and a generation and a loss of the detected WMH region over time between the first time point and the second time point. 2. The MRI apparatus of claim 1, wherein the output interface is further configured to indicate portions where the WMH region is generated and lost at corresponding locations in the second image. 3. The MRI apparatus of claim 1, wherein the output interface is further configured to: display the second image in a first region of a screen; anddisplay, as numerical values, a volume of the WMH region and an amount of change in the WMH region over time, in a second region of the screen. 4. The MRI apparatus of claim 1, wherein the generation and the loss of the WMH region over time comprises a change over time in the WMH region in at least one among frontal, parietal, temporal, and occipital lobes. 5. The MRI apparatus of claim 1, further comprising an input interface configured to receive an input selecting at least one among frontal, parietal, temporal, and occipital lobes. 6. The MRI apparatus of claim 5, wherein the output interface is further configured to display a change over time in the WMH region in the selected at least one among frontal, parietal, temporal, and occipital lobes. 7. The MRI apparatus of claim 1, wherein the image processor is further configured to: generate a blended image by performing a weighted sum on the first image and the second image; anddetermine the white matter region based on the blended image. 8. The MRI apparatus of claim 1, wherein the image processor is further configured to normalize an intensity of an image signal in each of the second image acquired at the first time point and the second image acquired at the second time point. 9. The MRI apparatus of claim 8, wherein the image processor is further configured to normalize the intensity of the image signal, based on an intensity of an image signal in a gray matter region in each of the second image acquired at the first time point and the second image acquired at the second time point. 10. A method of processing a magnetic resonance (MR) image, the method comprising: acquiring, at a first time point, a first image that is a T1-weighted image and a second image that is a fluid attenuated inversion recovery (FLAIR) image by performing a first magnetic r... | SAMSUNG ELECTRONICS CO., LTD. | G06K009/00 G06T011/60 A61B005/055 G06T007/11 G06T007/174 A61B005/00
US09795304 | B2 | 2017-10-24 | Low profile temperature transducer | 2012-12-04 | A low profile temperature transducer has a working surface for placement against a body surface and a first output. The transducer is a flat laminate composed of alternating conductive and dielectric layers. The laminate defines at least one slotline antenna for exposure to the body surface to pick up thermal emissions from the underlying tissue at depth. A feed network having a characteristic impedance is connected to the first output and a slotline-to-stripline transition is connected between the at least one antenna and the feed network, the transition providing a match between the impedance at the at least one antenna and the characteristic impedance. Also, a temperature sensor may be present at the working surface to detect the body surface temperature under the transducer, that surface temperature being used to calculate actual temperature at depth. | 1. A low profile temperature transducer having a working surface for placement against a body surface to pick up thermal emissions at a depth below the body surface and produce a corresponding transducer output, said transducer comprising a flat laminate composed of alternating conductive and dielectric layers which define at least one slotline antenna adapted to receive thermal radiation emanating from a depth below the body surface and produce a corresponding antenna output, a feed network having a characteristic impedance connected to said transducer output and a slotline-to-stripline transition connected between the at least one antenna and the feed network, said transition providing a match between impedance at the at least one antenna and said characteristic impedance, wherein the flat laminate includes:a first metal layer at the working surface that defines an antenna aperture:a second metal layer beyond the first metal layer is slotted to form the at least one antenna, the first metal layer and the second metal layer being grounded:a third metal layer beyond the second metal layer that includes said feed network and saidtransition: and a fourth metal layer beyond the third metal layer is slotted in the same configuration as the second metal layer and the corresponding slots of the second metal layer and the fourth metallayer are connected electrically at slot perimeters thereof to form the at least one antenna. 2. The transducer defined in claim 1 wherein said transition comprises a pair of stripline traces which bridge each slot antenna at selected spaced-apart locations along the length thereof, one of each stripline trace being connected to the second metal layer, and a stub trace connected to the other end of each stripline trace to provide said impedance match, andsaid feed network comprises electrical connections between the stub traces associated with each pair of stripline traces and said transducer output. 3. The transducer defined in claim 2 wherein said at least one antenna comprises two or more similar antennas, andthe feed network combines the outputs of all of the two or more similar antennas at said transducer output. 4. The transducer defined in claim 1 wherein plated through holes are provided in the dielectric layers between the second metal layer and the fourth metal layer to electrically connect the slot perimeters of the second metal layer and the fourth metal layer. 5. The transducer defined in claim 1 wherein said first metal layer is part of a grounded conductive shield that surrounds all the other layers to exclude outsideelectrical effects. 6. The transducer defined in claim 1 and further including a temperature sensor affixed to said working surface for sensing the temperature of the body surface under the transducer, an output of the sensor constituting a second transducer output. 7. The transducer defined in claim 6 wherein said sensor is selected from the group consisting of an infrared sensor, a thermi... | Coral Sand Beach LLC | G01K001/16 G01J005/00 G01J005/12 G01J005/20 A61B005/01 G01K011/00 G01K013/00 A61B005/1455 A61B005/00 H01Q001/27 H01Q013/10 A61F007/00 A61F007/02
US09775535 | B2 | 2017-10-03 | Non-invasive prediction of risk for sudden cardiac death | 2014-11-07 | A method and apparatus for the quantitative determination of an individual's risk for sudden cardiac death (SCD) is described. Risk stratification is accomplished (and may have a sensitivity and specificity of greater than about 90%) by determining the presence in any individual being tested for SCD risk of sequences identified herein to correlate quantitatively with SCD risk. Both the number of such sequences present and their alignment scores (similarity) with the SCD risk sequence ensemble are used to calculate quantitative SCD risk. | 1. A method for determining an individual's risk for sudden cardiac death (SCD), the method comprising: receiving a digital electrocardiogram (ECG) measurement taken from the individual; andquantitatively determining. in, a processor, in the received ECG measurement presences of predetermined digital electrocardiogram (ECG) sequences indicative of a risk for SCD, wherein the processor assesses quality of alignment between (a) the predetermined digital ECG sequences, first derivatives of the predetermined ECG sequences and second derivatives of the predetermined ECG sequences and (b) EGG sequences measured from the individual, first derivatives of the ECG sequences measured from the individual and second derivatives of the ECG sequences measured from the individual. 2. The method of claim 1, wherein the processor detects presence of a first predetermined ECG sequence of the predetermined ECG sequences by computing an alignment error score related to the first predetermined ECG sequence. 3. The method of claim 2, wherein the ECG measurement is obtained from a standard resting 12-lead ECG machine. 4. The method of claim 3, wherein the processor detects, in the received ECG measurement, the first predetermined digital ECG sequence indicative of the risk for SCD by: receiving the first predetermined digital ECG sequence indicative of the risk for SCD;aligning the first predetermined digital ECG sequence with respect to an ECG sequence obtained from the received ECG measurement, to obtain an alignment;determining an error score for the alignment; anddetermining, in the received ECG measurement, presence of the first predetermined digital ECG sequence indicative of the risk for SCD based on the error score. 5. The method of claim 3, further comprising preprocessing, with a preprocessor, the ECG measurement taken from the individual to obtain first and second derivatives of digital ECG data from the ECG measurement. 6. The method of claim 5, wherein the preprocessor preprocesses the digital ECG measurement taken from the individual by: denoising the digital ECG data, andderiving first and second derivatives from the denoised digital ECG data. 7. The method of claim 6, wherein the ECG measurement is denoised by at least one of a Finite Impulse Response (FIR) filter and a wavelet packet operation using entropy calculations to optimize threshold settings. 8. The method of claim 7, wherein the noise includes at least one of electrical noise, mechanical noise, respiratory artifacts, white noise, movement artifact, and baseline drift. 9. The method of claim 7, wherein the processor detects, in the received digital ECG measurement, the predetermined digital ECG sequences indicative of the risk for SCD by: receiving the predetermined digital ECG sequences indicative of the risk for SCD;aligning ECG sequences from the digital ECG measurement with the predetermined digital ECG sequences associated with SCD risk, to obtain an alignment; anddetermining an er... | Spangler Scientific LLC | A61B005/04 A61B005/00 A61B005/0402 G06K009/00
US09830005 | B2 | 2017-11-28 | Devices, systems, and methods for empathetic computing | 2015-11-11 | Devices, systems, and methods for empathetic computing are described herein. An example empathetic computing device includes an enclosure configured to fit into the palm of a user. The enclosure may have an upper portion and a lower portion and comprise an interface board, a processor, and a touch sensor. The interface board may have a plurality of light emitting devices configured to provide a light pattern of a visual response. The processor may be coupled to a memory device. The touch sensor may be configured to detect the touch of the user. The touch sensor may comprise a touch belt disposed along the enclosure circumferentially and a touch pad. | 1. An empathetic computing device, comprising: an enclosure having a first portion and a second portion;an interface board arranged within the enclosure and comprising a light emitting device configured to provide a visual response, a sound generator configured to provide an auditory response, a vibrator configured to provide a vibrational response, or combinations thereof; a processor coupled to the interface board and a memory device; anda sensor configured to detect a proximity of the user prior to any physical contact between the empathetic computing device and the user and configured to transmit a proximity signal to the processor for providing a first response, based in part on a determination that the user is in proximity, and further configured to detect a natural action of the user and transmit a second signal to the processor for providing a second response to the user, wherein the second response is configured to emulate an action, expression, or emotion of the user. 2. The empathetic computing device of claim 1, wherein at least one of the first portion or the second portion of the enclosure is partially translucent. 3. The empathetic computing device of claim 1, wherein the sensor includes one or more sensors selected from a touch sensor, a proximity sensor, a motion sensor, an accelerometer, a gyroscope, a light sensor, a pressure sensor, a heat sensor, or combinations thereof. 4. The empathetic computing device of claim 1, wherein the sensor comprises a plurality of infrared sensors disposed around a circumference of the empathetic computing device for detecting the proximity of the user to the empathetic computing device. 5. The empathetic computing device of claim 1, wherein the empathetic computing device is configured to enter a one of a plurality of different modes responsive to detection of one of a plurality of different of natural actions of the user, and wherein the empathetic computing device is configured to provide a first visual, auditory, or vibrational responses responsive to entering a first mode of the plurality of different modes and a second different visual, auditory, or vibrational response responsive to entering a second mode of the plurality of different modes. 6. The empathetic computing device of claim 1, wherein the interface board comprises one or more light emitting diodes configured to provide a light expression. 7. The empathetic computing device of claim 1 wherein empathetic computing device is configured to detect a facial expression of the user and generate a light pattern emulating the facial expression of the user. 8. The empathetic computing device of claim 1, wherein the empathetic computing device is configured to generate a relatively long vibration in response to the determination that the user is in proximity and further generate a relatively short vibration in response to the user touching the empathetic computing device. 9. An empathetic computing system, comprising: an empathetic ... | SomniQ, Inc. | G06F003/041 G06F003/01 G06F001/16 A61B005/16
US09775516 | B2 | 2017-10-03 | Outpatient monitoring systems and methods | 2012-12-05 | Described are computer-based methods and apparatuses, including computer program products, for outpatient monitoring. In some examples, the outpatient monitoring technology includes a system for remote monitoring and consultations of patients' state associated with automated interpretation of vital signs signals. In other examples, the outpatient monitoring technology includes a system for remote monitoring of patients' state associated with analysis of full-disclosure vital signs digital signals with reduced bit data or channel data. In some examples, the outpatient monitoring technology includes a system for remote monitoring of patients' state associated with ongoing automated analysis and transmission of full-disclosure vital signs digital signals and automatically generated signals' interpretation results and manually marked patient symptoms. | 1. A method for remote monitoring of patients' state associated with interpretation results data representative of the patients' vital signs, comprising the steps of: remotely initializing a portable patient monitor to begin a monitoring session having an associated unique identification number, the remote initialization causing the portable patient monitor to acquire, via a vital signs sensor, continuous full-disclosure vital signs;analyzing, in real time via the portable patient monitor, the continuous full-disclosure vital signs to identify at least one of: a pathological event and emergency event;selecting, via the portable patient monitor, portions of the continuous full-disclosure vital signs to be stored, at the portable patient monitor, as uncompressed/unreduced data fragments based on generated interpretation results and/or a signal condition representing at least one significant signal event;transmitting data signals representative of the continuous full-disclosure vital signs with interpretation results to a remote location, wherein the full-disclosure vital signs are real-time patient vital signs, by creating, at the portable patient monitor, data signals that includes a combination of bit-reduced/channel-reduced data signals and the uncompressed/unreduced data signals, wherein the uncompressed/unreduced data signals represent at least one of: the pathological event and emergency event; andsynchronizing verified interpretation results data with a data interpretation center and repository of a remote monitoring station. 2. The method of claim 1, wherein the sessions' identification number is a combination of the session initialization time, date, and a portable monitor's identification number. 3. The method of claim 1, further comprising creating an event list based on interpretation results data, each event having an unconfirmed status indicator by the one or more processors. 4. The method of claim 3, further comprising changing the unconfirmed status indicator to a confirmed status indicator when the event has been detected properly, deleted, determined to be a false event, determined to be unclear, or when the consultant is not sure whether it is a true event or an incorrectly detected event by the one or more processors. 5. A non-transitory computer program product, tangibly embodied in a non-transitory information carrier, the computer program product including instructions being operable to cause a data processing apparatus to: remotely initialize a portable patient monitor to begin a monitoring session having an associated unique identification number, the remote initialization causing the portable patient monitor to acquire, via a vital signs sensor, continuous full-disclosure vital signs;analyze, in real time via the portable patient monitor, the continuous full-disclosure vital signs to identify at least one of: a pathological event and emergency event;select, via the portable patient monitor, portions of the continuous f... | MEDICALGORITHMICS S.A. | A61B005/00 G06F019/00 G06Q050/22 H04L029/08
US09826902 | B2 | 2017-11-28 | Sensor guide wire device and system including a sensor guide wire device | 2011-01-17 | A sensor guide wire device (200) for intravascular measurements of a physiological variable in a living body is disclosed. The sensor guide wire (200) may comprise a sensor element (213) and a connector unit (220). The connector unit (220) is attachable to a signal converting device (310). The sensor element (213) provides the signal converting device (310) with a signal indicative of a physiological variable sensed by the sensor element (213). Furthermore, the connector unit (220) comprises a battery (222). A system (300) comprising the sensor guide wire device (200) and the signal converting device (310) is also disclosed. | 1. A sensor guide wire device comprising: a sensor guide wire configured to perform an intravascular measurement of a physiological variable in a living body, the sensor guide wire comprising at least one sensor element configured to sense the physiological variable in the living body; anda connector unit configured to be directly or indirectly connected to a proximal end of the sensor guide wire, the connector unit being physically attachable to a signal converter via electrical contacts, the connector unit being configured to provide the signal converter with a signal indicative of the physiological variable sensed by the sensor element, the signal converter being configured to convert the signal received from the connector unit into a processed signal that is receivable by a physiology monitor,wherein the connector unit comprises a housing and a power supply disposed in the housing, the housing of the connector unit being separate from a housing of the signal converter,wherein the housing of the connector unit is reversibly attachable to and detachable from the housing of the signal converter,wherein the connector unit is configured such that, when the housing of the connector unit is attached to the housing of the signal converter, a surface of the housing of the connector unit abuts a surface of the housing of the signal converter, and the power supply of the connector unit powers the signal converter via the electrical contacts, andwherein the sensor guide wire and the connector unit are disposable. 2. The sensor guide wire device according to claim 1, wherein said sensor element is energized by said signal converter when the connector unit is attached to said signal converter. 3. The sensor guide wire device according to claim 2, wherein said connector unit comprises a calibration memory that includes calibration data related to said sensor element. 4. The sensor guide wire device according to claim 1, wherein said connector unit comprises a calibration memory that includes calibration data related to said sensor element. 5. The sensor guide wire device according to claim 1, wherein the power supply comprises at least one battery. 6. The sensor guide wire device according to claim 1, wherein the connector unit is positioned at a proximal portion of the sensor guide wire. 7. The sensor guide wire device according to claim 1, wherein the connector unit comprises a male connector configured to connect with a female connector of the signal converter. 8. The sensor guide wire device according to claim 1, wherein the connector unit comprises a female connector configured to connect with a male connector of the signal converter. 9. The sensor guide wire device according to claim 1, wherein the signal converter is configured to process said signal such that the processed signal is receivable by any of multiple different physiology monitors having differing signal requirements. 10. The sensor guide wire device according to claim 1, wherein... | ST. JUDE MEDICAL COORDINATION CENTER BVBA | A61B005/00 A61B005/0215 A61M025/09 A61B005/01
US09820515 | B2 | 2017-11-21 | Apparel with thermochromic dyes for monitoring and optimizing exercise performance | 2013-03-15 | Apparel for monitoring optimal body shell temperature during physical activity. The apparel uses thermochromic dyes and thermochromic inks to indicate changes in body shell temperature as physical activity commences and progresses. The thermochromic dyes and inks reveal when the body shell temperature has reached a temperature indicating optimal muscle performance, and inform the wearer of the apparel whether to increase or decrease the intensity of the physical activity to establish or maintain the optimal body shell temperature. | 1. Exercise apparel for monitoring optimal body shell temperature in a wearer of the apparel, comprising a first set of fibers having a first thermochromic dye infused therein that changes color at 32 degrees C., a second set of fibers having a second thermochromic dye infused therein that changes color at 34 degrees C., and a third set of fibers having a third thermochromic dye infused therein that changes color at 36 degrees C.; andwherein the apparel is capable of showing the temperatures of portions of the wearer's body based on the color changes of the thermochromatic dyes. 2. The exercise apparel of claim 1, further comprising a logo or a design comprising a thermochromic ink that changes color at 34 degrees C. 3. The exercise apparel of claim 1, wherein the fibers comprise cotton and nylon fibers. 4. The exercise apparel of claim 1, further comprising a fourth set of fibers having a fourth thermochromic dye infused therein that changes color at 38 degrees C. 5. The exercise apparel of claim 1, wherein the apparel is a shirt. 6. The exercise apparel of claim 5, wherein the first thermochromic dye, the second thermochromic dye, and the third thermochromic dye are present only in a sleeve of the exercise apparel. 7. The exercise apparel of claim 5, wherein the first thermochromic dye, the second thermochromic dye, and the third thermochromic dye are present only on a shoulder portion of the exercise apparel. 8. The exercise apparel of claim 5, wherein the first thermochromic dye, the second thermochromic dye, and the third thermochromic dye are present only on a chest portion of the exercise apparel. 9. The exercise apparel as claimed in claim 1 wherein a portion of the fibers are moisture wicking fibers. 10. Exercise apparel for monitoring optimal body shell temperature in a wearer of the apparel, comprising a first set of fibers having a first thermochromic dye infused therein that changes color at 33 degrees C., a second set of fibers having a second thermochromic dye infused therein that changes color at 35 degrees C., and a third set of fibers having a third thermochromic dye infused therein that changes color at 37 degrees C.; andwherein the apparel is capable of showing the temperatures of portions of the wearer's body based on the color changes of the thermochromatic dyes. 11. The exercise apparel of claim 10, further comprising a logo or a design comprising a thermochromic ink that changes color at 35 degrees C. 12. The exercise apparel of claim 10, wherein the fibers comprise cotton and nylon fibers. 13. The exercise apparel of claim 10, wherein the apparel is a shirt. 14. The exercise apparel of claim 10, wherein the first thermochromic dye, the second thermochromic dye, and the third thermochromic dye are present only in a sleeve of the exercise apparel. 15. The exercise apparel of claim 10, wherein the first thermochromic dye, the second thermochromic dye, and the third thermochromic dye are present only on a should... | - | A41D013/00 A61B005/01 A63B071/06 A61B005/00
US09808175 | B1 | 2017-11-07 | Method and system for analyzing images to quantify brain atrophy | 2010-02-02 | A technique for reliably measuring brain atrophy based on image data, such as data collected as part of an MRI scan of a subject's brain. Rather than measure anatomical features of the brain in one of the scan planes as output by an MRI system, the image data is reformatted to generate a slice that may be transverse to those scan planes. The transverse slice is generated by determining a location of anatomical landmarks and reformatting the image data to represent a slice through these anatomical landmarks. The Anterior and Posterior Commissures are useful landmarks and ventricle area is a useful characteristic to measure from the slice. The rate of change of this measured characteristic can be tracked to determine disease progression and may be useful for diagnosis or evaluation of treatments, and the technique may be applied during clinical trials. | 1. A method of analyzing image data comprising a 3-dimensional representation of a brain collected with an imaging device, the method comprising: with at least one processor:obtaining an indication of a position of a first anatomical landmark in a 3-dimensional representation of a brain collected during a first imaging session and an indication of a position of a second anatomical landmark in the 3-dimensional representation of the brain collected during the first imaging session; andcomputing, from the 3-dimensional representation of the brain collected during the first imaging session, a first slice with an orientation in the 3-dimensions, wherein the orientation of the first slice is selected based on the obtained position of the first anatomical landmark and the obtained position of the second anatomical landmark such that the orientation of the first slice is not constrained by an orientation of the imaging device and the first slice encompasses:a position selected based on at least the position of the first anatomical landmark anda position selected based on at least the position of the second anatomical landmark;obtaining an indication of a position of the first anatomical landmark in a 3-dimensional representation of a brain collected during a second imaging session and an indication of a position of a second anatomical landmark in the 3-dimensional representation of the brain collected during the second imaging session;computing, from the 3-dimensional representation of the brain collected during the second imaging session, a second slice with an orientation in the 3-dimensions,wherein the orientation of the second slice is selected based on the obtained position of the first anatomical landmark and the obtained position of the second anatomical landmark such that the orientation of the second slice is not constrained by an orientation of the imaging device, wherein the orientation of the second slice with respect to the orientation of the imaging device is different than the orientation of the first slice with respect to the orientation of the imaging device, and wherein the second slice encompasses:a position selected based on at least the position of the first anatomical landmark anda position selected based on at least the position of the second anatomical landmark; andmeasuring a change in a characteristic of an anatomical feature based on a comparison of information from the first slice to information from the second slice to track the change in the characteristic of the anatomical feature over time. 2. The method of claim 1, wherein obtaining an indication of a position of a first anatomical landmark and an indication of a position of a second anatomical landmark comprises receiving user input. 3. The method of claim 1, wherein obtaining an indication of the position of the first anatomical landmark and an indication of the position of the second anatomical landmark comprises performing an automated analysis of the 3-dimensional... | PAREXEL International Corporation | A61B005/055
US09788738 | B2 | 2017-10-17 | System and method for deep vein thrombosis prevention and diagnosis | 2013-02-05 | A system and method prevents and diagnoses deep vein thrombosis in a body limb by providing a pressure sleeve having a plurality of individually fillable cells, the pressure sleeve being configurable to be placed around a body limb. A source fills each fillable cell individually, and a pressure sensor measures a pressure in a fillable cell. A controller establishes a fill sequence of each individually fillable cell and a fill time for each individually fillable cell. The controller causes a first individually fillable cell of the pressure sleeve to be filled to a predetermined pressure and causes the pressure of first individually fillable cell of the pressure sleeve to be measured while a second individually fillable cell of the pressure sleeve is filled. The controller determines a presence of deep vein thrombosis in a body limb having the pressure sleeve therearound based upon a measured pressure change in the first individually fillable cell of the pressure sleeve. | 1. A system for diagnosing deep vein thrombosis in a body limb, comprising: a compression system for applying external pressure to a body limb; anda venous phasic flow monitoring system to detect a thrombosis in the body limb;said compression system including a pressure sleeve;said pressure sleeve including a first individually fillable cell, a second individually fillable cell, and a third individually fillable cell;said first individually fillable cell being configured to be located on the body limb proximal to a heart;said second individually fillable cell being configured to be located on the body limb distal to the heart;said third individually fillable cell being configured to be located on the body limb between said first individually fillable cell and said second individually fillable cell;said compression system including a first pneumatic pressure sensor to measure pneumatic pressure;said compression system including a second pneumatic pressure sensor to measure pneumatic pressure;said compression system performing a first operation and a second operation;said compression system, during said first operation, filling said first individually fillable cell to a predetermined pressure, said pressure of said first individually fillable cell being measured by said first pneumatic pressure sensor;said compression system, during said first operation, filling said second individually fillable cell while said compression system measures, using said second pneumatic pressure sensor, a pneumatic pressure change in said filled first individually fillable cell caused by a change in a volume of the body limb within said pressure sleeve and generating a first signal corresponding to changes in the pneumatic pressure in said filled first individually fillable cell, measured by said second pneumatic pressure sensor, caused by the change in the volume of the body limb within said pressure sleeve;said compression system, during said second operation, filling said first individually fillable cell and said second individually fillable cell to said predetermined pressure, said pressure of said first and second individually fillable cells being measured by said first pneumatic pressure sensor;said compression system, during said second operation, filling said third individually fillable cell while said compression system measures, using said second pneumatic pressure sensor, pneumatic pressure changes in said filled first individually fillable cell and said filled second individually fillable cell, measured by said second pneumatic pressure sensor, caused by a change in a volume of the body limb within said pressure sleeve and generating a second signal corresponding to changes in the pneumatic pressure in said filled first individually fillable cell and said filled second individually fillable cell caused by the change in the volume of the body limb within said pressure sleeve;said venous phasic flow monitoring system monitoring, over a period of time, said f... | MEDICAL COMPRESSION SYSTEMS (DBN) LTD. | A61B005/02 A61B005/0285 A61B005/021 A61B005/00 A61H009/00 A61B005/0205 A61B005/022 A61B005/0225
US09775542 | B2 | 2017-10-03 | Apparatus for evaluating the integrity of a uterine cavity | 2013-07-18 | Methods, systems and devices for evaluating the integrity of a uterine cavity. A method comprises introducing transcervically a probe into a patient's uterine cavity, providing a flow of a fluid (e.g., CO2) through the probe into the uterine cavity and monitoring the rate of the flow to characterize the uterine cavity as perforated or non-perforated based on a change in the flow rate. If the flow rate drops to zero or close to zero, this indicates that the uterine cavity is intact and not perforated. If the flow rate does not drop to zero or close to zero, this indicates that a fluid flow is leaking through a perforation in the uterine cavity into the uterine cavity or escaping around an occlusion balloon that occludes the cervical canal. | 1. A system for performing a procedure in a patient's gynecologic tract, comprising: an ablation device;first and second expandable members carried by the device, the expandable members having fluid-tight walls;a source of a pressurized flow of a fluid;a passageway in the device, connected to the source, and configured to deliver the fluid from the source to a region surrounding an exterior of device when the expandable members are expanded within or proximate to at least a portion of the cervical canal and the uterine cavity of a patient;a flow sensor for measuring a flow rate of the fluid in the passageway; anda controller operatively coupled to the flow sensor and to the source, the controller being configured to (1) initiate a flow of fluid from the source through the passageway, (2) measure the flow rate using the flow sensor, and (3) determine if the measured flow rate drops below a predetermined minimum level within a predetermined time period ranging from 1 second to 60 seconds. 2. The system of claim 1 wherein the controller is configured to generate a signal indicating that the uterine cavity has a perforation if the flow rate fails to fall below said predetermined minimum level within said predetermined time period. 3. The system of claim 1, wherein the controller is configured to generate a signal indicating that the uterine cavity is not perforated if the flow rate drops below said predetermined minimum level within said predetermined period of time. 4. The system of claim 1, wherein the ablation device carries at least one electrode coupled to an RF generator. 5. The system of claim 4, wherein the controller is configured to disable activation of the at least one electrode if the flow rate fails to fall below said predetermined minimum level within said predetermined time period. 6. The system of claim 4, wherein the controller is configured to automatically activate the at least one electrode if the flow rate falls below said predetermined minimum level within said predetermined time period. 7. A system for performing a procedure in a patient's gynecologic tract, comprising: an ablation device having an elongated member extending about an axis to a distal portion configured for positioning in a cervical canal and uterine cavity of a patient;a first expandable member having a fluid-tight wall carried by the elongated member;a second expandable member having a fluid-tight wall carried by the elongated member wherein the first expandable member is axially moveable relative to the second expandable member;a source of a pressurized flow of a fluid;a passageway in the elongated member, connected to the source, and configured to deliver the fluid from the source to an outlet in said distal of the elongated member;a flow sensor for measuring a flow rate of the fluid in the passageway; anda controller operatively coupled to the flow sensor and to the source, the controller being configured to (1) initiate a flow of fluid from the sour... | Minerva Surgical, Inc. | A61B005/00 A61B005/107 A61B005/03 A61B017/42 A61B018/14 A61B018/18 A61B018/04 A61B017/00 A61B017/02 A61B017/22 A61B018/00
US09808594 | B2 | 2017-11-07 | Flow regulation vent | 2016-10-20 | A flow regulation vent for use in controlling a vent flow of washout gas in a system for supplying breathable gas pressurized to within a therapeutic pressure range above atmospheric pressure to a patient in the treatment of a sleep disordered breathing condition is disclosed. The flow regulation vent may include a housing having at least one vent orifice and at least one fixed bleed orifice; and a flap configured to regulate the vent flow of washout gas through the at least one vent orifice, wherein the at least one fixed bleed orifice is not covered by the flap in any position and the flow regulation vent provides a minimum vent flow independent of the position of the flap. | 1. A flow regulation vent for use in controlling a vent flow of washout gas to atmosphere in a system for supplying breathable gas pressurized to within a therapeutic pressure range above atmospheric pressure to a patient in the treatment of a sleep disordered breathing condition, the flow regulation vent comprising: a housing having at least one vent orifice and at least one fixed bleed orifice; anda flap configured to regulate the vent flow of washout gas to atmosphere through the at least one vent orifice,wherein the at least one fixed bleed orifice is not covered by the flap in any position and the flow regulation vent provides a minimum vent flow to atmosphere independent of the position of the flap, andwherein the at least one vent orifice and the at least one fixed bleed orifice are structured such that the vent flow through the at least one vent orifice and the at least one fixed bleed orifice has a substantially constant flow rate when the breathable gas is pressurized to any pressure within the therapeutic pressure range. 2. The flow regulation vent of claim 1, wherein the flap is configured such that increasing the pressure of the breathable gas within the therapeutic pressure range causes the flap to move and reduce a cross-sectional area of the at least one vent orifice to reduce the vent flow through the at least one vent orifice. 3. The flow regulation vent of claim 2, wherein the at least one vent orifice and the at least one fixed bleed orifice are structured such that the vent flow through the at least one fixed bleed orifice increases in response to a reduction of the vent flow through the at least one vent orifice. 4. The flow regulation vent of claim 1, wherein the therapeutic pressure range is 4 cmH2O to 30 cmH2O. 5. The flow regulation vent of claim 1, wherein the housing further comprises positive operating stops to limit travel of the flap. 6. The flow regulation vent of claim 1, wherein the at least one vent orifice is positioned downstream of the flap relative to the vent flow. 7. The flow regulation vent of claim 1, wherein the flap extends from the housing such that the flap is cantilevered over the at least one vent orifice. 8. The flow regulation vent of claim 1, further comprising a diffuser to diffuse the vent flow. 9. The flow regulation vent of claim 1, wherein the flap has an outer diameter that is greater than a thickness of the flap. 10. The flow regulation vent of claim 1, wherein the at least one vent orifice comprises a plurality of vent orifices, and the at least one fixed bleed orifice comprises a plurality of fixed bleed orifices. 11. The flow regulation vent of claim 1, wherein the flap comprises an elastically deformable material. 12. The flow regulation vent of claim 11, wherein when the flap is not elastically deformed, the flap is configured to not restrict the vent flow through the at least one vent orifice. 13. The flow regulation vent of claim 12, wherein when the breathable gas is p... | RESMED LIMITED | A61M016/20 A61B005/08 A61M016/06 A61B005/087 G01F001/28 A61M016/08 A61M016/00
US09797859 | B2 | 2017-10-24 | Methods of using information from recovery pulses in electrochemical analyte measurements as well as devices, apparatuses and systems incorporating the same | 2015-09-11 | Methods are disclosed for measuring an analyte concentration in a fluidic sample. Such methods allow one to correct and/or compensate for confounding variables such as temperature before providing an analyte concentration. The measurement methods use response information from a test sequence having at least one DC block, where the DC block includes at least one excitation pulse and at least one recovery pulse, and where a closed circuit condition of an electrode system is maintained during the at least one recovery pulse. Information encoded in the at least one recovery pulse is used to correct/compensate for temperature effects on the analyte concentration. Also disclosed are devices, apparatuses and systems incorporating the various measurement methods. | 1. A method of compensating for temperature while electrochemically measuring an analyte in a fluid sample, the method comprising the steps of: applying an electrical test sequence to an electrochemical biosensor, the biosensor comprising:an electrode system,a reagent in electrical communication with the electrode system, anda receptacle configured to contact the fluid sample provided to the biosensor with the reagent,with the fluid sample in fluidic contact with the reagent, wherein the test sequence comprising at least one direct current (DC) block having a sequence of at least one excitation potential pulse and at least one recovery potential pulse, wherein a closed circuit condition of the electrode system is maintained during the DC block; anddetermining the analyte concentration based upon current response information from the DC block, wherein information from the at least one recovery potential pulse is used to compensate for temperature effects on the analyte concentration based upon a partial least squares (PLS) regression model. 2. The method of claim 1, wherein the at least one excitation potential pulse is about +450 mV and the at least one recovery potential pulse is about 0 mV, and wherein each pulse is from about 50 msec to about 500 msec. 3. The method of claim 1 further comprising: measuring current response information to the at least one excitation potential pulse and to the at least one recovery potential pulse; anddetermining the analyte concentration from the excitation current response and the recovery current response. 4. The method of claim 1, wherein the test sequence further comprises an alternating current (AC) block of low-amplitude signals of at least two different frequencies. 5. The method of claim 4, wherein the AC block is applied before the at least one DC block, after the at least one DC block or interspersed within the at least one DC block. 6. The method of claim 1, wherein the PLS regression model is based upon a covariate dataset comprising hematocrit, temperature and analyte concentration. 7. The method of claim 1, wherein the PLS regression model is based upon a covariate dataset comprising temperature and analyte concentration. 8. The method of claim 4, wherein the wherein the frequencies are about 10 kHz, about 20 kHz, about 10 kHz, about 2 kHz and about 1 kHz, and wherein each is applied for about 0.5 seconds to about 1.5 seconds. 9. The method of claim 1, wherein the information of an electrical response to a relaxation pulse includes unique information not found in information of an electrical response to an excitation pulse. 10. The method of claim 1, wherein the analyte concentration is a glucose concentration. | Roche Diabetes Care, Inc. | G01N027/327 A61B005/145 A61B005/1473 A61B005/1486 A61B005/1495
US09775758 | B2 | 2017-10-03 | Person support apparatus having physiological sensor | 2016-12-15 | A person support apparatus includes a frame and a support surface cooperating with the frame to support a person. The person support apparatus also has a sensor coupled to one of the frame and the support surface. The sensor detects at least one characteristic associated with the person. A controller is coupled to the sensor. In response to at least one of a condition of the frame, a condition of the support surface, a position of the person, or a condition of the person, the controller operates to control the sensor by at least one of changing a gain of the sensor and changing a manner in which a signal from the sensor is filtered. In some instances, the controller turns the sensor off. | 1. A person support apparatus comprising: a frame including a plurality of deck sections including a head section to support a person's head and a portion of the person's torso, the head section being pivotable between raised and lowered positions,a mattress supported by the plurality of deck sections;a physiological sensor to detect at least one physiological characteristic of the person, the physiological sensor being located above the head section and being pivotable with the head section as the head section pivots, the physiological sensor detecting the at least one physiological characteristic through at least a portion of the mattress;a force sensor coupled to the frame and operable to sense at least a portion of a weight of the person;a pressure sensor coupled to the frame and operable to sense pressure associated with an inflatable bladder of the mattress;at least one controller coupled to the frame; anda network bus coupling the physiological sensor, the force sensor, and the pressure sensor to the at least one controller, wherein the at least one controller generates an alert signal in response to data from the physiological sensor meeting a threshold condition. 2. The person support apparatus of claim 1, wherein the at least one controller is configured to adjust a sensitivity of the physiological sensor. 3. The person support apparatus of claim 2, wherein the at least one controller adjusts the sensitivity of the physiological sensor in response to information from the force sensor or the pressure sensor. 4. The person support apparatus of claim 2, wherein the at least one controller adjusts the sensitivity of the physiological sensor by changing a gain of the physiological sensor or by changing a manner in which a signal from the physiological sensor is filtered. 5. The person support apparatus of claim 2, wherein the at least one controller adjusts the sensitivity of the physiological sensor by changing a gain of the physiological sensor and by changing a manner in which a signal from the physiological sensor is filtered. 6. The person support apparatus of claim 1, wherein the at least one controller is configured to turn power to the physiological sensor on and off. 7. The person support apparatus of claim 6, wherein the at least one controller is configured to turn power to the physiological sensor on and off in response to information from the force sensor. 8. The person support apparatus of claim 7, wherein power to the physiological sensor is turned on by the at least one controller in response to the force sensor sensing that the person is present on the patient support apparatus. 9. The person support apparatus of claim 7, wherein power to the physiological sensor is turned off by the at least one controller in response to the force sensor sensing that the person is absent from the patient support apparatus. 10. The person support apparatus of claim 1, wherein a signal from the physiological sensor is used by the at ... | Hill-Rom Services, Inc. | G08B021/02 A61G007/057 A61B005/00 A61B005/024 A61B005/08 A61B005/11 A61G007/012 A61G007/005
US09795337 | B2 | 2017-10-24 | System, apparatus, and method for promoting usage of core muscles and other applications | 2013-12-19 | Embodiments disclosed include a system and method for development of core muscles' support, comprising a means for identifying a user qualifying movement, a means for detecting a core muscle contraction in the identified qualifying movement, a means for discriminating between a core muscle contraction and no core muscle contraction in the identified qualifying movement; and a means to provide feedback to the user. | 1. A system for development of core muscles' support, comprising: a first sensor coupled to a processor for detecting movements of a body;a second sensor coupled to the processor for detecting a core muscle contraction;a qualifying movement identifier running on the processor for identifying a qualifying movement of the body in an upright position for which core contraction is beneficial; anda movement discriminator running on the processor for discriminating between a protected qualifying movement where the core muscle contraction is detected by the second sensor during the qualifying movement and an unprotected qualifying movement where the core muscle contraction is not detected by the second sensor during the qualifying movement;wherein the processor provides feedback to a user. 2. The system of claim 1 further comprises: identifying a number of protected qualifying movements, and a number of unprotected qualifying movements. 3. The system of claim 1 wherein identifying the protected qualifying movement comprises identifying the core muscle contraction of the user starts before a start and ends after an end of the qualifying movement when stress is placed on the user's lumbo sacral junction. 4. The system of claim 1 wherein identifying the unprotected qualifying movement comprises identifying inadequate or no core muscle contraction of the user during the qualifying movement when stress is placed on the user's lumbo sacral junction. 5. The system of claim 1 wherein the identifying the qualifying movement comprises identifying at least one of: standing from a seated position, sitting from a standing position, rotating to the left or right while standing, or a sequence of movements associated with: removing a carry-on load from an overhead bin in an airplane, lifting groceries out of a car trunk, serving or spiking a volleyball, and stabilizing a golf putting stroke. 6. The system of claim 1 wherein the identifying the qualifying movement, comprises identifying a movement of the user's body wherein the core muscle contraction would be beneficial. 7. The system of claim 1 wherein the feedback comprises at least one of a buzzer, an audible sound, a tens unit zap, a wearable computing device, a personal computer, a smart phone, and a portable device. 8. The system of claim 1 further comprising an exercise application. 9. The system of claim 1 wherein the processor identifies when the core should be contracted during the protected qualifying movement or the unprotected qualifying movement via image on a display, audible instruction, or video. 10. The system of claim 1 wherein the processor stores information about each of the identified qualifying movements in a core score keeper, the information including a status as protected qualifying movement or unprotected qualifying movement and a time stamp and the processor uses a quantity of the protected qualifying movements, a quantity of the unprotected qualifying movements and the time stamps... | Alert Core, Inc. | A61B005/22 A63B024/00 A63B023/02 A61B005/11 G06F019/00 G06K009/00 G09B019/00 A61B005/00
US09795747 | B2 | 2017-10-24 | Methods and apparatus for lancet actuation | 2011-06-02 | A lancet driver is provided wherein the driver exerts a driving force on a lancet during a lancing cycle and is used on a tissue site. The driver comprises of a drive force generator for advancing the lancet along a path into the tissue site, and a manual switch for a user interface input. | 1. A lancet device configured to exert a driving force on a lancet during a lancing cycle and used on a tissue site, the lancet device including electrical components, comprising: a housing, the housing including a first housing and a second housing, where the first housing being larger than the second housing and the first and second housings having dissimilar shapes, the first housing configured to house a battery, and the second housing configured to house the electrical components, a lancet analyte sensor associated with the lancet; a processor coupled to a drive force generator, the processor and drive force generator positioned in the second housing, the drive force generator configured to be coupled to the lancet to advance the lancet along a path into the tissue site, create a penetration and a subsequent flow of a blood sample into a receiving channel of the second housing and then to the analyte sensor for analysis;a latching device that holds the a lancet from exiting the lancet device when the lancet device is not in use;a bearings guide associated with the lancet for accurately defining a lancing trajectory;a human interface input control configured to provide depth setting range for lancet penetration into the tissue site; anda human interface coupled to the drive force generator, the human interface position at the second housing, the human interface being a voice system which uses words to describe status or alarms related to the lancet device usage, the voice system including a sound generator configured to provide information to a user relative to at least one of a reminder when a next measurement of a blood sample is needed and an analytic measurement result of a blood sample. 2. The device of claim 1, further comprising: an LED or light source configured to indicate different user interface outputs selected from at least one of, low battery, charging, lancet present, device error condition, ready to launch, and an indication that the battery requires replacement. 3. The device claim 2, further comprising: a user interface indicator. 4. The device of claim 1, further comprising: a multi-position mechanical switch at the second housing. 5. The device of claim 1, wherein the device has a total moving mass that is less than 0.40 g. 6. The device of claim 1, further comprising: a sterility barrier that seals the lancet prior to use. 7. The device of claim 1, wherein the drive force generator is a controllable drive driver. 8. The device of claim 1, further comprising: an actuator arm coupled to the drive force generator. 9. The device of claim 1, further comprising: a coupler configured to be coupled to the drive force generator. | SANOFI-AVENTIS DEUTSCHLAND GMBH | A61B005/00 B65D081/00 A61B017/14 A61B017/32 A61M005/46 A61B005/15 A61B005/157 A61M005/145 A61M005/172 A61M005/20 A61M005/32 A61B005/151
US09808174 | B2 | 2017-11-07 | Magnetic-resonance imaging diagnosis apparatus and magnetic-resonance imaging method | 2009-12-02 | Upon detecting a body-motion before starting a main-imaging, a sequence-switching control-unit controls operation so as to switch from a usual-imaging sequence to a body-motion adaptive-sequence corresponding to an imaging-portion of a subject P by referring to a body-motion adaptive-sequence storage-unit. Moreover, upon detecting a body-motion during the main-imaging according to the usual-imaging sequence, the sequence-switching control-unit refers to a collected-data storage-unit, and controls operation so as to perform a retake by switching to the body-motion adaptive-sequence if an already-collected data-volume is less than a predetermined volume. By contrast, if the already-collected data-volume is equal to or more than the predetermined volume at a time of detecting a body-motion during the main-imaging, the sequence-switching control-unit stops the main-imaging, and controls a data-processing unit so as to reconstruct a magnetic resonance image only with collected data. | 1. A magnetic-resonance imaging diagnosis apparatus comprising: a computer;a data-collection controller configured to automatically detect a motion of a subject and to control a magnetic resonance image via main imaging by collecting magnetic resonance signals from an inside of the subject, a data collection method being initialized to a first image data-collecting method before collecting the magnetic resonance signals, said data-collection controller controlling to switch the data collection method from the first image data-collecting method to a second image data-collecting method when the motion of the subject is detected and to collect the magnetic resonance signals by the second image data-collecting method after switching the data collection method, the second image data-collecting method being one of a data collecting method having a shorter imaging time than the first image data-collecting method and a data collecting method correcting some of the magnetic resonance signals corresponding to the motion of the subject; andsaid computer connected to said data-collection controller and that reconstructs the magnetic resonance image from the magnetic resonance signals that are collected during the main imaging in accordance with the data collection method,wherein a pulse sequence of the first data-collecting method is created based on first imaging parameters of the first image data-collecting method, and a pulse sequence of the second image data-collecting method is created based on second imaging parameters of the second image data-collecting method that are different from those of the first imaging parameters. 2. The magnetic-resonance imaging diagnosis apparatus according to claim 1, wherein the data-collection controller controls to switch from the first image data-collecting method to the second image data-collecting method in a first case where the motion of the subject is detected before the main imaging, and in a second case where a data volume of already collected magnetic resonance signals is less than a predetermined volume at a time point when the motion of the subject is detected during the main imaging according to the first image data-collecting method, and the data-collection controller stops the main imaging in a third case where the data volume of already collected magnetic resonance signals is equal to or more than the predetermined volume at a time point when the motion of the subject is detected during the main imaging according to the first image data-collecting method, and controls said computer so as to reconstruct the magnetic resonance image only with collected magnetic resonance signals. 3. The magnetic-resonance imaging diagnosis apparatus according to claim 2, wherein the second image data-collecting method is a method of collecting the magnetic resonance signals by rotating every repetition time a belt-shaped region formed of a plurality of parallel data collection traces in a frequency space, and said compute... | Toshiba Medical Systems Corporation | A61B005/05 A61B005/055 A61B005/00 G01R033/48 G01R033/565 A61B005/11 G01R033/56 G01R033/561 G01R033/567
US09814408 | B2 | 2017-11-14 | Display system for displaying conditions of esophageal mucosa and indications of gastroesophageal reflux disease | 2011-07-13 | A mucosal impedance measuring system that is capable of direct and accurate impedance measurements of esophageal mucosa includes example multi-channel intraluminal impedance catheters with inflatable and deflatable components for positioning the impedance sensor electrodes in direct contact with the esophageal mucosa on the inside surface of the esophagus. Impedance measurements from multiple locations in the esophagus are processed for indications of mucosal damage that are indicative of damage due to gastrointestinal reflux disease (GERD), non-erosive reflux disease (NERD), and Barrett's esophagus. | 1. A catheter system, comprising: a plurality of electrically conductive impedance sensor electrodes dispersed longitudinally along a length of an exterior surface of a catheter that is sized and shaped for insertion into an esophagus in a manner that positions the impedance sensor electrodes at a plurality of locations along a length in the esophagus, said catheter including an inflatable and debatable balloon positioned in a manner such that inflation of the balloon in the esophagus forces the impedance sensor electrodes into intimate contact with esophageal mucosa at the locations along the length in the esophagus, and said catheter including a tube connected in fluid flow relation to the balloon for conducting an inflating fluid into and out of the balloon;inflation apparatus connected to the tube for directing and controlling a flow of the inflation fluid into and out of the balloon, the inflation apparatus being configured to receive and respond to inflation control signals to vary the fluid pressure in the balloon;an impedance measuring circuit connected to the impedance sensor electrodes, including circuitry that applies voltages across selected pairs of the impedance sensor electrodes and that produces signals indicative of impedance of the esophageal mucosa between the selected pairs of impedance sensor electrodes at the plurality of locations along the length in the esophagus;computer apparatus connected to the impedance measuring circuit and programmed to: (i) receive the signals from the impedance measuring circuit that are indicative of an impedance of the esophageal mucosa at the plurality of locations along the length in the esophagus, to compare and determine differences between the impedances of the esophageal mucosa at the plurality of locations along the length in the esophagus, and to output display signals that are indicative of the impedances of the esophageal mucosa at the plurality of locations along the length in the esophagus and of differences between the impedances of the esophageal mucosa at the plurality of locations along the length in the esophagus; (ii) output the inflation control signals for varying the fluid pressure in the balloon; (iii) correlate the signals from the impedance measuring circuit that are indicative of the impedance of the esophageal mucosa at the plurality of levels in the esophagus with particular pressure levels in the balloon; (iv) determine changes of impedance of the esophageal mucosa at particular levels in the esophagus as a function of the fluid pressure in the balloon; and (v) compare the changes in impedance of the esophageal mucosa at particular levels in the esophagus as a function of fluid pressure in the balloon to impedance change as a function of pressure change criteria that is indicative of abnormal esophageal mucosa as distinct from normal esophageal mucosa and to output display signals that are indicative of abnormal esophageal mucosa in the esophagus for which the impedan... | Diversatek Healthcare, Inc. | A61B005/05 A61B005/103 A61B005/117 A61M029/00 A61B005/053 A61B005/00
US09797973 | B2 | 2017-10-24 | MR imaging with extended field of view | 2009-06-18 | Magnetic resonance examination system comprises displaceable carrier for supporting an object to be examined. The carrier can be moved over a two dimensional area. The magnetic resonance examination system is configured to acquire sets of magnetic resonance signals from the object for various positions of the carrier in the two dimensional area. | 1. A magnetic resonance examination system, comprising: an acquisition system configured to acquire magnetic resonance signals;a displaceable carrier configured to support an object to be examined; anda controller configured to control the acquisition system and the carrier, and further configured to: move the carrier along a continuous motion direction in a two dimensional area,shift the carrier along a direction transverse to the continuous motion direction to respective lateral positions, andfor respective lateral positions, acquire a set of magnetic resonance signals during continuous motion of the carrier along the continuous motion direction. 2. The magnetic resonance examination system of claim 1 and further having two magnetic poles between which a main magnetic field extends. 3. The magnetic resonance examination system of claim 1, wherein the controller is configured to apply oversampling of the acquired magnetic resonance signals in the direction transverse to the continuous motion direction of the carrier. 4. The magnetic resonance examination system of claim 1, wherein the controller is further configured to apply oversampling along at least one phase encoding direction. 5. The system of claim 1, wherein each set of magnetic resonance signals corresponds to a field of view, and wherein the controller is configured to acquire magnetic resonance signals corresponding to a plurality of fields of view including at least a first field of view and a second field of view which are displaced with respect to each other in the continuous motion direction, and at least a third field of view and a fourth field of view which are displaced with respect to each other in the direction transverse to the continuous motion direction. 6. The magnetic resonance examination system of claim 1, wherein the continuous motion direction is a first linear direction and the direction transverse to the continuous motion direction is a second linear direction orthogonal to the first linear direction, wherein the object to be examined has a width which extends in the second linear direction, and wherein the controller is configured to move the carrier, for each of the respective lateral positions, linearly along the first linear direction to acquire the set of magnetic resonance signals which span less than the entire width of the object to be examined. 7. The magnetic resonance examination system of claim 6, wherein the sets of magnetic resonance signals acquired for all of the respective lateral positions together span the entire width along the second linear direction of the examination area be imaged. 8. A method of operating a magnetic resonance examination system which employs a carrier configured to support an object to be examined, the method comprising: acquiring a plurality of sets of magnetic resonance signals from the object for various positions of the carrier in two dimensions, wherein at least a first one of the positions and a second one of t... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/00 G01R033/563
US09775557 | B2 | 2017-10-03 | Quantifying breast tissue changes with spectrally selective MRI and MRS | 2014-04-03 | Systems and methods for magnetic resonance analysis and imaging are provided. IN particular, pulse sequences for DWI, APT, and MRS analysis and imaging are provided which rely on an RF excitation pulse for the signal of interest, followed by one or more refocusing pulses and acquisition steps, based on the type of imaging. | 1. A method for a magnetic resonance (MR) analysis of a biological structure, comprising: generating a spin echo imaging pulse sequence for acquisition comprising a spectrally selective excitation pulse, a plurality of refocusing pulses subsequent to the excitation pulse, and a pair of diffusion gradient pulses, the excitation pulse selected to excite water signals in the biological structure without exciting lipid signals in the biological structure associated with at least one dominant resonance for lipids;subsequent to the imaging pulse sequence, collecting echo signals from the biological structure; andproducing at least one diffusion weighted MR image based on the echo signals,wherein the plurality of refocusing pulses comprises at least one pair of adiabatic inversion pulses, and wherein the pair of diffusion gradient pulses are applied prior to and subsequent to one of the plurality of refocusing pulses. 2. The method of claim 1, wherein the excitation pulse comprises a pulse with an amplitude between 0.1 μT and 30μ, a center frequency between less than 12.0 ppm, a bandwidth between 20 Hz and 525 Hz, and a duration between 4 ms and 100 ms. 3. The method of claim 1, wherein the excitation pulse is one of a Gaussian pulse, a sinc Gaussian pulse, a minimal phase pulse, or an adiabatic half passage (AHP) pulse. 4. The method of claim 1, wherein the biological structure comprises breast tissue. 5. The method of claim 1, wherein a duration of each of the pair of diffusion pulses is longer than a duration of an associated one of the plurality of refocusing pulses. 6. A method for a magnetic resonance (MR) analysis of a biological structure, comprising: generating a first spin echo pulse sequence for acquisition comprising at least one saturation pulse followed by a spectrally selective excitation pulse, the at least one saturation pulse selected for saturating protons for protein species in the biological structure, and the excitation pulse selected to excite water signals in the biological structure without exciting lipid signals in the biological structure associated with at least one dominant resonance for lipids;applying a plurality of refocusing pulses subsequent to the first spin each pulse sequence;applying pair of diffusion gradient pulses;subsequent to the applying, collecting echo signals from the biological structure; andproducing at least one saturation transfer MR image based on the echo signals,wherein the plurality of refocusing pulses comprises at least one pair of adiabatic inversion pulses, and wherein the pair of diffusion gradient pulses are applied prior to and subsequent to one of the plurality of refocusing pulses. 7. The method of claim 6, wherein the excitation pulse comprises a pulse with an amplitude between 0.1 μT and 30 μT, a center frequency between 0 ppm and 12.0 ppm, a bandwidth between 20 Hz and 525 Hz, and a duration between 4 ms and 100 ms. 8. The method of claim 6, wherein the excitation pulse is one of ... | VANDERBILT UNIVERSITY | G01V003/00 A61B005/00 A61B005/055 G01R033/485 G01R033/56 G01R033/563
US09820690 | B1 | 2017-11-21 | Analyte detection system | 2014-07-16 | A wearable device includes a mount to mount the wearable device on a living body and a detector to detect an analyte response signal transmitted from tissue in the living body. The tissue contains a biologically active agent in an inactive state and functionalized particles. The biologically active agent can be converted to an active state that can affect a biological state of the living body. The functionalized particles are configured to bind with a target analyte, the presence or absence or concentration or abundance of which is correlated with the biological state. The analyte response signal is related to interaction of the target analyte with the functionalized particles. A source can apply directed energy into the tissue that is sufficient to convert the biologically active agent from the inactive state to the active state. A processor can determine a presence or absence or concentration or abundance of the analyte based on the analyte response signal. | 1. A system comprising: a wearable device for detecting an analyte response signal and an unbound functionalized particle signal from tissue through an external surface of a living body, the wearable device comprising a mount adapted for mounting the wearable device on the external surface of a living body and a detector, wherein the tissue contains a biologically active agent in an inactive state and functionalized particles, wherein the functionalized particles can interact with one or more target analytes, wherein the biologically active agent can be converted from the inactive state to an active state, wherein the concentration or abundance of the one or more target analytes is correlated with a biological state of the living body that is affected by the biologically active agent when in the active state, wherein the analyte response signal is related to interaction of the one or more target analytes with the functionalized particles, and wherein the unbound functionalized particle signal is related to functional particles that are not interacting with the one or more target analytes;a directed energy source;a modulation source;a processor; anda computer readable medium storing instructions executable by the processor to perform operations comprising: controlling the directed energy source to apply directed energy into the tissue through the external surface, wherein the directed energy is sufficient to convert the biologically active agent from the inactive state to the active state;controlling the detector to detect an analyte response signal transmitted from tissue through the external surface;controlling the modulation source to modulate the analyte response signal and the unbound functionalized particle signal, such that the analyte response signal is affected differently than the unbound functionalized particle signal; andcontrolling the processor to determine the concentration or abundance of the one or more target analytes by differentiating the analyte response signal from the unbound functionalized particle signal, based, at least in part, on the modulation. 2. The system of claim 1, wherein the biologically active agent comprises a pro-drug when in the inactive state and comprises a drug when in the active state. 3. The system of claim 1, wherein the biologically active agent is associated with a carrier when in the inactive state and is released from the carrier when in the active state. 4. The system of claim 3, wherein the carrier comprises a liposome, micelle, nanoparticle, nanocage or reservoir. 5. The system of claim 1, wherein the directed energy comprises electromagnetic radiation, acoustic energy, or heat. 6. The system of claim 5, wherein the directed energy comprises a series of one or more pulses for controlling a rate of conversion of the biologically active agent from the inactive state to the active state. 7. The system of claim 6, wherein the operations further comprise controlling the rate of conversion based... | Verily Life Sciences LLC | A61B005/1455 A61B005/00 A61B005/1468 A61B005/055
US09795362 | B2 | 2017-10-24 | Method, system, and apparatus for cranial anatomy evaluation | 2012-07-23 | Embodiments of portable cranial anatomy injury evaluation systems, apparatus, and methods are described generally herein where the system and apparatus may include multiple signal generation devices including photonic, acoustic, and electrical signals. Other embodiments may be described and claimed. | 1. A system for evaluating cranial anatomy, including: a flexible module configured to engage at least a portion of a patient's cranium;a photonic generation module (PGM) at a known location in the flexible module;a photonic detection module (PDM) at a known location in the flexible module;an acoustic generation module (AGM) at a known location in the flexible module, the AGM configured to generate acoustic energy and receive acoustic energy; anda controller, the controller directing the operation of the PGM to generate photonic energy having a first frequency, directing the operation of the AGM to generate acoustic energy having a second frequency, the first frequency different than the second frequency, receiving electrical signals from the PDM representing photonic energy and from the AGM representing acoustic energy; and evaluating a combination of the received signals to determine the status of a section of the at least portion of the patient's cranium. 2. The system for evaluating cranial anatomy of claim 1, further including a memory module, the memory module including signals representing received data for a PDM having various cranium anatomy status and the controller correlating received PDM data with the stored signals to determine a cranium anatomy status. 3. The system for evaluating cranial anatomy of claim 2, the controller applying signals to the plurality of PGM and receiving signals from the corresponding, nearby plurality of PDM and correlating stored PDM data with the received data to determine the physical location of one or more of the plurality of PGM relative to a patient cranial anatomy and correlating the position of the other of the plurality of PGM and the AGM relative to the patient cranial anatomy based on the determined physical location of one or more of the plurality of PGM. 4. The system for evaluating cranial anatomy of claim 3, further including a plurality of antennas at different, known locations in the flexible module to generate low-energy electrical signals and to receive electrical signals, the controller receiving electrical signals from plurality of antennas and evaluating a combination of the received PDM, AGM, and antenna signals to determine the status of a section of the at least portion of the patient's cranium. 5. The system for evaluating cranial anatomy of claim 4, the memory module further including signals representing received data for an antenna at a known location having various cranium anatomy status and the controller correlating received antenna data from a corresponding known location with the stored signals to determine a cranium anatomy status. 6. The system for evaluating cranial anatomy of claim 4, further including a memory module, the memory module including signals representing combined received data for a PDM at a known location, an antenna at a known location, and an AGM at a known location having various cranium anatomy status and the controller correlating received PDM da... | - | A61B005/1455 A61B008/00 A61B005/00 A61B005/026 A61B005/05 A61B008/08 A61B005/107 A61B005/03 A61B005/0476 A61B008/06
US09817009 | B2 | 2017-11-14 | Biological information measurement device | 2013-09-27 | It is an object of certain embodiments of the present invention to suppress variance in measurement results from a biological information measurement device and enhance measurement accuracy. A biological information measurement device comprises a main case having a sensor mounting component for mounting a sensor that senses biological information, a measurement component that connects to the sensor mounting component, a controller that connects to the measurement component, a first timer that connects to the controller, and a second timer that connects to the controller. The controller has a measurement mode in which a measurement operation is performed through the measurement component. In the measurement mode, the controller executes the measurement operation based on a preset measurement timing using the first timer, and determines the suitability of the measurement operation based on the preset measurement timing, using the second timer. | 1. A biological information measurement device, comprising: a main case including a sensor mounting component configured to mount a sensor that senses biological information;a measurement component that connects to the sensor mounting component;a controller that connects to the measurement component;a first timer that connects to the controller; anda second timer that connects to the controller,wherein the controller has a measurement mode in which a measurement operation is performed through the measurement component, andin the measurement mode, the controller executes the measurement operation based on a preset measurement timing using the first timer, and uses the second timer to determine a suitability of the measurement operation based on the preset measurement timing. 2. The biological information measurement device according to claim 1, wherein in the measurement mode the controller determines whether or not there is a discrepancy between the preset measurement timing and an actual measurement timing of the measurement component. 3. The biological information measurement device according to claim 2, further comprising a display component that connects to the controller,wherein the controller causes the display component to display information indicating whether or not there is a discrepancy between the preset measurement timing and the actual measurement timing. 4. The biological information measurement device according to claim 1, wherein the second timer is actuated by the controller before the measurement operation is performed through the measurement component. 5. The biological information measurement device according to claim 1, wherein the second timer is constituted by a ceramic oscillator. 6. The biological information measurement device according to claim 5, wherein the second timer is provided integrally inside the controller. 7. The biological information measurement device according to claim 1, wherein the controller:executes the measurement operation at a plurality of preset measurement timings in the measurement mode, anduses the second timer to determine the suitability of each measurement operation based on the plurality of preset measurement timings. 8. The biological information measurement device according to claim 7, wherein in the measurement mode the controller determines whether or not there is a discrepancy between one or more preset measurement timings out of the plurality of preset measurement timings and a corresponding actual measurement timing when the measurement operation is executed on the basis of said one or more preset measurement timings. 9. A biological information measurement device, comprising: a main case;a sensor mounting component configured to mount a sensor that senses biological information;a measurement component that connects to the sensor mounting component and measures the biological information;a controller that connects to the measurement component and causes the measurement compo... | Panasonic Healthcare Holdings Co., Ltd. | G01N033/49 G01N015/06 G01N033/00 G01N033/48 G01N035/00 G01N033/543 G01N033/66 A61B005/145
US09782279 | B2 | 2017-10-10 | Systems, devices, and methods for autoretroperfusion | 2014-11-18 | Devices, systems, and methods for autoretroperfusion. In at least one embodiment of a perfusion system of the present disclosure, the system comprises a first catheter having a proximal end, a distal end, and a first lumen therethrough, the distal end configured for insertion into a luminal organ of a patient, a coupler defining an outlet port and a first port, the coupler configured to engage the proximal end of the first catheter at the outlet port and to receive blood from a blood supply through the first port, a first tube having a proximal end, a distal end, and a lumen therethrough, the distal end of the first tube configured to engage the first port of the coupler and to receive the blood from the blood supply, and a first flow regulator in communication with one or more of the coupler and the first tube, the first flow regulator operable to regulate a flow and/or a pressure of the blood through at least part of the system, wherein the system is configured to permit the blood from the blood supply to flow through the first tube, the coupler, and the first catheter into the patient to treat a patient condition. | 1. A method of using a perfusion system, the method comprising the steps of: positioning at least a portion of a first catheter of a perfusion system into a luminal organ of a patient at a first time;positioning at least a portion of an arterial tube of the perfusion system into an arterial vessel of the patient at a second time and connecting the arterial tube and the first catheter to a coupler;selectively operating one or both of a first flow regulator and a second flow regulator of the perfusion system, while a heart of the patient is pumping blood, to regulate blood flow from the arterial vessel to the luminal organ and/or to regulate blood pressure within at least part of the first catheter, each without the use of an external pump;wherein the second flow regulator is configured for intracorporeal placement within the patient. 2. The method of claim 1, wherein the step of positioning at least a portion of a first catheter further comprises the step of inflating an expandable balloon positioned along the portion of the first catheter positioned into the luminal organ to secure the portion of the first catheter within the luminal organ. 3. The method of claim 1, wherein the first time is earlier than the second time. 4. The method of claim 1, wherein the first time is later than the second time. 5. The method of claim 4, wherein the step of positioning at least a portion of an arterial tube further comprises the step of operating one or more of the first flow regulator and the second flow regulator to regulate blood flow from the arterial vessel to the luminal organ so to substantially eliminate an introduction of a gas within at least a portion of the perfusion system to the luminal organ. 6. The method of claim 1, wherein the method is performed to treat a patient selected from the group consisting of a patient indicating a S-T segment elevated myocardial infarction, a cardiogenic shock patient, a percutaneous coronary intervention patient, an ischemic cardiac patient, a diabetic patient, an atherosclerotic patient, a hypertensive patient, a patient with cerebral and peripheral vascular disease, a patient with renal dysfunction, a patient with chronic pulmonary disease, and a no-option patient. 7. The method of claim 1, further comprising the step of: removing the at least a portion of a first catheter from the luminal organ within about 24 hours after positioning the at least a portion of a first catheter into the luminal organ. 8. The method of claim 1, further comprising the step of: removing the at least a portion of a first catheter from the luminal organ between about 24 hours and about 48 hours after positioning of the at least a portion of a first catheter into the luminal organ. 9. The method of claim 1, further comprising the step of: removing the at least a portion of a first catheter from the luminal organ after about 48 hours after positioning of the at least a portion of a first catheter into the luminal organ. 10. Th... | CVDevices, LLC | A61M001/36 A61F002/856 A61B005/0215 A61M027/00 A61M025/04 A61B017/00 A61F002/958 A61M001/10 A61M025/00 A61M025/10
US09788726 | B2 | 2017-10-17 | Device and method for measuring vibration transmittance of sternum | 2011-06-09 | A method for measuring vibration transmittance of a sternum noninvasively includes producing vibration, transmitting the vibration to a first side the sternum through skin and soft tissues, obtaining response data of the sternum from a second side of the sternum through skin and soft tissues, which first and second sides of the sternum are on different sides of the midline of the sternum, and processing said response data for determining transmittance of vibration of the sternum. The response data may be used for defining state of a patient such as sternal instability of the breastbone or abnormal healing in the sternum after open chest surgery by surgeons. | 1. A method for measuring vibration transmittance of a sternum noninvasively with an apparatus, comprising: producing vibration with spectral content,transmitting the vibration with spectral content with varying frequency to a first side of the sternum through skin and soft tissues by a transmitter,obtaining response data of the sternum from a second side of the sternum through skin and soft tissues by an accelerometer, which first and second sides of the sternum are on different sides of the midline of the sternum, and wherein there is a handle between the transmitter and the accelerometer, so that the distance of the transmitter and the accelerometer remains at a constant distance within 2 to 5 centimeters from the midline of the sternum during measuring, andprocessing said response data for determining integrated transmittance of vibration of the sternum. 2. A method according to claim 1, wherein spectral content of the produced vibration is in the 50-1500 Hz range. 3. A method according to claim 1, wherein spectral content of the produced vibration is in the 50-2000 Hz range. 4. A method according to claim 1, wherein the processing comprises producing a transfer function from response data and extracting parameters from the transfer function. 5. A method according to claim 1, wherein the measuring of vibration transmittance of the sternum is performed at least twice, wherein the first measuring is executed before a sternal incision of the sternum and the second measuring is executed after the sternal incision of the sternum, and wherein the response data of the first and second measurements are compared. 6. A method according to claim 1, wherein the measuring of vibration transmittance of the sternum is executed to the sternum at least at two different positions. 7. A system for measuring vibration transmittance of a sternum noninvasively, wherein the system comprises a device comprising a processor and a memory including computer program code for producing vibration, a transmitter for transmitting the vibration with spectral content with varying frequency to a first side of the sternum through skin and soft tissues, by an accelerometer for obtaining response data of the sternum from a second side of the sternum through skin and soft tissues, which first and second sides of the sternum are on different sides of the midline of the sternum, and wherein there is a handle between the transmitter and the accelerometer, so that the distance of the transmitter and the accelerometer remains at a constant distance within 2 to 5 centimeters from the midline of the sternum during measuring, and wherein the device is further arranged to process said response data for determining integrated transmittance of vibration of the sternum. 8. A system according to claim 7, wherein the device is further arranged to receive said obtained data for further data processing. 9. A system according to claim 7, wherein the device is connected to a computer that i... | TAYS SYDÄNKESKUS OY | A61B005/00
US09792409 | B2 | 2017-10-17 | Communicative water bottle and system thereof | 2015-03-13 | A communicative water bottle includes communication logic and wireless transmission logic technology electronically connected with a variety of sensors either on the water bottle or located remote from the water bottle. The sensors on the bottle create digital data associated with amount of fluid in the bottle and change thereof. The sensors remote from the bottle, which can be on an activity tracker, create digital data associated with an activity being performed by a user, such as running, or the absence of activity, such as remaining sedentary. A display on the bottle can indicate to the user the amount of fluid consumed or a reminder that fluid should be consumed. The fluid consumption data syncs with other remote devices such as mobile applications executable on smartphones. | 1. A communicative sports water bottle comprising: a bottle body having a base and an upwardly extending sidewall therein defining a fluid chamber and a first display coupled to the bottle body;a sensor determining one or more of the following an amount of fluid in the fluid chamber, an amount of fluid being dispensed from the fluid chamber, and a physical movement of the bottle body;fluid information logic electronically coupled with the sensor creating digital data associated with fluid in the chamber;data transmission and reception logic configured to send the digital data to a remote device and receive digital data from the remote device, wherein the data transmission and reception logic monitors wireless input sources for incoming wireless packets and analyzes any received packets in order to detect a source of the packet transmission and decide whether to sync the digital data associated with fluid in the chamber with the source of the packet;wherein the remote device is a wrist-worn health activity tracker including a second display that displays information associated with one or more of the following: the amount of fluid in the chamber, the amount of fluid dispensed from the chamber, an amount of time since fluid was last dispensed from the chamber, an amount of fluid dispensed over a given time period, a reminder notification if fluid has not been dispensed over a given time period, a reminder to consume an adjusted amount of fluid based, at least in part, on activity information obtained from the remote device;network connectivity electronically connecting the data transmission and reception logic and the remote device, wherein in the network connectivity is selected from a group comprising of: Bluetooth connection, wireless internet connection, wired internet connection, internet, and 3G/4G connection; anda record of information for a previous period at a fixed time interval;artificial intelligence configured to learn the record of information to detect behavior signatures from the previous period, which is implemented to predict a likelihood of a subject has certain activity levels at a present time. 2. The communicative sports water bottle of claim 1, further comprising: a cap repeatably attachable and detachable from the bottle body; and wherein the sensor is on the cap;cap sensor placement-specific algorithms executed by the fluid information logic to determine the one or more of the following the amount of fluid in the fluid chamber, the amount of fluid dispensed from the fluid chamber, the physical movement of the bottle body, the amount of time since fluid was last dispensed from the chamber, the amount of fluid dispensed over a given time period, and the reminder notification if fluid has not been dispensed over a given time period; andadaptive filtering logic to cancel out excessive movements of the cap. 3. The communicative sports water bottle of claim 1, wherein the sensor is on the sports water bottle body; further compr... | - | A63B069/00 G06F019/00 A61B005/00 A61B005/11
US09797780 | B2 | 2017-10-24 | Contact type internal thermometer | 2013-02-21 | Provided is a contact type internal thermometer including: a measurement surface that is brought into contact with a surface to be measured of a measurement object; a first temperature sensor stack including a first measurement surface side temperature sensor arranged on a measurement surface side, and a first back surface side temperature sensor arranged on a back surface side; a second temperature sensor stack including a second measurement surface side temperature sensor arranged on a measurement surface side, and a second back surface side temperature sensor arranged on a back surface side; and a controller configured to calculate an internal temperature of the measurement object based on measurement results of the respective temperature sensors, in which the first back surface side temperature sensor and the second back surface side temperature sensor have different surface areas. | 1. A contact type internal thermometer, comprising: a measurement surface that is brought into contact with a surface to be measured of a measurement object so as to calculate an internal temperature of the measurement object;a first temperature sensor stack comprising a first measurement surface side temperature sensor arranged on a measurement surface side of a first flexible printed board, and a first back surface side temperature sensor arranged on a back surface side of the first flexible printed board;a second temperature sensor stack comprising a second measurement surface side temperature sensor arranged on a measurement surface side of a second flexible printed board, and a second back surface side temperature sensor arranged on a back surface side of the second flexible printed board; anda controller configured to calculate the internal temperature of the measurement object based on measurement results of the first measurement surface side temperature sensor, the first back surface side temperature sensor, the second measurement surface side temperature sensor, and the second back surface side temperature sensor,wherein the first back surface side temperature sensor and the second back surface side temperature sensor have different surface areas such that one of the surface areas is intentionally larger than the other surface area. 2. The contact type internal thermometer according to claim 1, wherein the first back surface side temperature sensor and the second back surface side temperature sensor have different volumes such that one of the volumes is intentionally different from the other volume. 3. The contact type internal thermometer according to claim 1, wherein the first back surface side temperature sensor has an intentionally larger volume than a volume of the first measurement surface side temperature sensor. 4. The contact type internal thermometer according to claim 1, further comprising; a controller activates a cooling mechanism to cool the first temperature sensor stack and the second temperature sensor stack after the controller calculates an internal temperature. 5. The contact type internal temperature thermometer according to claim 4, wherein the cooling mechanism comprises: a blower;an intake port; andan exhaust port;wherein the blower is configured to draw external air through the intake port and cause air to be exhausted through the exhaust port. | CITIZEN WATCH CO., LTD. | G01K007/22 A61B005/00 G01K007/42 G01K013/00 A61B005/01
US09775568 | B2 | 2017-10-03 | Medical device alarm management | 2015-02-24 | It is determined, during monitoring of a patient, that a plurality of alarm conditions has been initiated. The alarm conditions characterize a measurement by and/or operation of at least one medical device and have an associated grade among a plurality of prioritized grades. Thereafter, alarms can be generated that corresponding to some or all of the plurality of alarm conditions such that among concurrent active alarms: (i) an alarm corresponding to an alarm condition having a highest priority grade is conveyed at a first level upon initiation and, if such alarm condition remains active for a pre-defined time period, at the first level upon expiration of the pre-defined time period, and (ii) alarms corresponding to alarm conditions having equal or lower priority grades than the priority grade of the current highest priority grade alarm condition are conveyed at a second level, the second level being different from the first level. | 1. A method for implementation by one or more data processors forming part of at least one computing system, the method comprising: determining, by at least one data processor during monitoring of a patient, that a plurality of alarm conditions have been triggered, each alarm condition characterizing a measurement by and/or operation of at least one medical device and having an associated grade among a plurality of prioritized grades; andgenerating and conveying, by at least one data processor, audio and/or vibrational alarms corresponding to each of the plurality of alarm conditions such that among concurrent active alarms: an alarm corresponding to an alarm condition having a highest priority grade is conveyed at a first level and, when such alarm condition remains active for a pre-defined time period, is conveyed again at the first level upon expiration of the pre-defined time period, andalarms corresponding to alarm conditions having equal or lower priority grades than the priority grade of the current highest priority grade alarm condition are conveyed at a second level while the alarm condition having the highest priority grade is active, the second level being different and distinguishable from the first level. 2. The method as in claim 1, wherein when the alarm condition having the highest priority grade is removed prior to the expiration of the pre-defined time period, an alarm corresponding to an alarm condition having a next highest priority grade is conveyed at the first level and, when such alarm condition remains active for a second pre-defined time period, is conveyed at the first level upon expiration of the second pre-defined time period. 3. The method as in claim 1, wherein the alarms at the first level are different at each priority grade and the alarms at the second level are all the same. 4. The method as in claim 1, wherein the at least one medical device is selected from a group consisting of: patient monitors, anesthesia delivery platforms, ventilators, and physiological sensors. 5. The method as in claim 1, wherein the at least one medical device comprises an alarm management system. 6. The method as in claim 1, wherein, for each triggered alarm condition, a new timer is initiated which is used to determine whether such alarm condition remains active at the expiration of the pre-defined time period. 7. The method as in claim 1, wherein the at least one computing system comprises the at least one medical device. 8. The method as in claim 1, wherein the at least one computing system forms part of an alarm management system. 9. The method as in claim 1 further comprising: initiating, by at least one data processor, an alarm pause for a third pre-defined time period in response to user-generated input via the at least one medical device;wherein at least a portion of alarms initiated during the third pre-defined time period result in an attention alarm being conveyed to a caregiver. 10. The method as in claim 9, where... | Draeger Medical Systems, Inc. | A61B005/00 G06F019/00 A61M016/00 G08B003/10 G08B006/00 G08B021/02 G08B025/00
US09808155 | B2 | 2017-11-07 | Apparatus to detect amyloid in a patient user's retina in a diagnosis, advancement and prognosis of alzheimer's disease, traumatic brain injury, macular degeneration and a plurality of neurodegenerative disorders and ocular diseases | 2014-02-26 | An eye imaging apparatus to produce an image of an eye of a patient user is disclosed. The eye imaging apparatus may include a video camera, a slit lamp chinrest, a camera housing, a plurality of illumination source optics and a processing system. The eye imaging apparatus may also include a video camera, a plurality of video camera optics, a camera housing, a microscope assembly, a plurality of illumination source optics and a floor base. The eye imaging apparatus may be utilized in combination with one or more contrasting agents. The eye imaging apparatus may detect amyloid-beta plaque, amyloid or amyloid-beta peptide by a spectral signature. | 1. An eye imaging apparatus, comprising: a digital video camera having a plurality of video camera optics, wherein the digital video camera records an image of an amyloid-beta plaque, an amyloid or an amyloid-beta peptide in a patient user's eye or in a patient user's brain;a slit lamp chinrest adapted to hold a patient user's chin and forehead in a fixed position, the slit lamp chinrest includes a head support, a movable base, a joystick and a housing support;a camera assembly having a camera housing, the digital video camera, a plurality of observation optics and a lamp;a plurality of illumination source optics controlling focus and resolution of the camera housing; anda processing system having a memory system and an observation video monitor, the observation video monitor having high-resolution color graphics capability to display the patient user's eye image or the patient user's brain image to be analyzed, the processing system also having a digitizing board to collect video and digital data and images from the digital video camera instantly and store the collected video and digital data and images into digital data in the memory system and the collected video and digital data and images are available to the processing system to analyze. 2. The eye imaging apparatus according to claim 1, wherein the digital video camera incorporates a color or monochrome charge coupled device or a complementary metal oxide semiconductor. 3. The eye imaging apparatus according to claim 1, wherein the head support is adapted to provide a plurality of elevation adjustments to provide a comfortable resting place to dispose a patient user's head. 4. The eye imaging apparatus according to claim 1, wherein the movable base includes one or more wheels to move the eye imaging apparatus along a ground surface. 5. The eye imaging apparatus according to claim 1, wherein the camera housing is adapted to adjust the head support in gross and fine increments utilizing the joystick. 6. The eye imaging apparatus according to claim 1, further comprising a rubber cup provided between the camera housing and a patient user's eyeball. 7. The eye imaging apparatus according to claim 1, wherein the eye imaging apparatus is utilized in combination with one or more contrasting agents that are selected from the group consisting of diarylheptanoid, one or more diarylheptanoid derivatives, methylated dehydrothiotoluidine with sulfonic acid, one or more methylated dehydrothiotoluidine with sulfonic acid derivatives, methylation of dehydrothiotoluidine with methanol in presence of hydrochloric acid, one or more methylation of dehydrothiotoluidine with methanol in presence of hydrochloric acid derivatives, sodium salt of 3,3′-([1,1′-biphenyl]-4,4′-diyl)bis(4-aminonaphthalene-1-sulfonic acid), one or more sodium salts of 3,3′-([1,1′-biphenyl]-4,4′-diyl)bis(4-aminonaphthalene-1-sulfonic acid) derivatives, 2,5-Bis(2-(4-hydroxyphenyl)vinyl)anisole or methoxy-X04, a radioactive analog of ... | NeuroVision Imaging LLC | A61B003/10 A61B003/13 A61B005/00
US09788709 | B2 | 2017-10-17 | Endoscope system and image generation method to generate images associated with irregularities of a subject | 2014-04-11 | An endoscope system comprising an image signal acquisition unit acquiring an image signal by imaging a subject, and an irregularity image generation unit generating an irregularity image by extracting information of irregularities on the subject, the irregularity image generation unit includes a microstructure image generation section generating a microstructure image as the irregularity image by extracting a microstructure of a body tissue surface layer from the image signal, the image signal including image signals of a plurality of colors having information of different wavelength components, the microstructure image generation section includes a first brightness signal generation section generating a brightness signal showing an average brightness of the subject based on the image signals of the plurality of colors, and a first image generation section generating the microstructure image by extracting a pixel region, in which a pixel value exceeds a fixed threshold value, of the brightness signal as a microstructure. | 1. An endoscope system, comprising: an electric endoscope that acquires an image signal by imaging a subject; anda processor device configured togenerate an irregularity image by extracting information of irregularities on the subject from the image signal; wherein the processor device generates a microstructure image as the irregularity image by extracting a microstructure of a body tissue surface layer from the image signal;wherein the image signal is configured to include image signals of a plurality of colors having information of different wavelength components,the processor device further configured to:generate a brightness signal showing an average brightness of the subject based on the image signals of the plurality of colors; andgenerate the microstructure image by extracting a pixel region, in which a pixel value exceeds a fixed threshold value, of the brightness signal as a microstructure. 2. The endoscope system according to claim 1, wherein the processor device further performs high-frequency filtering on the microstructure image. 3. The endoscope system according to claim 2, wherein imaging of the subject is performed by a color imaging device having pixels of a plurality of colors in which respective color separation filters are provided. 4. The endoscope system according to claim 2, further comprising: a light source device that sequentially eradiates the subject with light beams of a plurality of colors,wherein imaging of the subject is performed by a monochrome imaging device whenever sequential irradiation is performed by the light source device. 5. The endoscope system according to claim 4, further comprising: a display unit for displaying the irregularity image. 6. The endoscope system according to claim 2, further comprising: a display unit for displaying the irregularity image. 7. The endoscope system according to claim 1, wherein imaging of the subject is performed by a color imaging device having pixels of a plurality of colors in which respective color separation filters are provided. 8. The endoscope system according to claim 7, further comprising: a display unit for displaying the irregularity image. 9. The endoscope system according to claim 1, further comprising: a light source device that sequentially eradiates the subject with light beams of a plurality of colors,wherein imaging of the subject is performed by a monochrome imaging device whenever sequential irradiation is performed by the light source device. 10. The endoscope system according to claim 1, further comprising: a display unit for displaying the irregularity image. 11. An endoscope system, comprising: an electric endoscope that acquires an image signal by imaging a subject;a processor device configured togenerate an irregularity image by extracting information of irregularities on the subject from the image signal;wherein the processor device generates a hypertrophy image as the irregularity image by extracting a hypertrophy, which has a thick... | FUJIFILM Corporation | A61B001/04 A61B001/06 G02B023/24 A61B001/00 A61B001/005 A61B001/05 A61B005/00 H04N005/232 H04N005/235 H04N009/04 G06T007/11 G06T007/136 H04N005/225
US09802042 | B2 | 2017-10-31 | Method and system for providing electrical stimulation to a user | 2016-10-17 | A method for providing electrical stimulation to a user as a user performs a set of tasks during a time window, the method comprising: providing an electrical stimulation treatment, characterized by a stimulation parameter and a set of portions, to a brain region of the user in association with the time window; for each task of the set of tasks: receiving a signal stream characterizing a neurological state of the user; from the signal stream, identifying a neurological signature characterizing the neurological state associated with the task; and modulating the electrical stimulation treatment provided to the brain region of the user based upon the neurological signature, wherein modulating comprises delivering a portion of the set of portions of the electrical stimulation treatment to the brain region of the user, while maintaining an aggregate amount of the stimulation parameter of the treatment provided during the time window below a maximum limit. | 1. A method for providing electrical stimulation to a user comprising: at a biosignal detection device, receiving a signal stream characterizing a neurological state of the user within a time window, wherein receiving the signal stream comprises receiving signals of the user in association with performance of a task;at a processing system coupled to the biosignal detection device, identifying a neurological signature indicative of the neurological state from the signal stream;at an electrical stimulation device in communication with the biosignal detection device, providing an electrical stimulation treatment to a brain region of the user in association with the time window;modulating the electrical stimulation treatment provided to the user based upon the neurological signature, wherein modulating comprises detecting at least one of an entrance and an exit of the neurological signature at the biosignal detection device, and modulating the electrical stimulation treatment upon detection of at least one of the entrance and the exit of the neurological signature; andguiding the user in performing the task at a mobile computing device associated with the user, wherein the task includes an athletic performance training activity, and wherein modulating the electrical stimulation treatment comprises modulating based upon the neurological signature associated with the athletic performance training activity. 2. The method of claim 1, wherein modulating the electrical stimulation treatment comprises maintaining an aggregate amount of the stimulation parameter of the electrical stimulation treatment provided to the user during the time window below a maximum limit. 3. The method of claim 1, wherein modulating the electrical stimulation treatment comprises detecting the exit of the neurological signature at the biosignal detection device, and terminating the electrical stimulation treatment upon detection of the exit of the neurological signature. 4. The method of claim 1, wherein providing the electrical stimulation treatment comprises providing stimulation toward least at one of a motor cortex region and a premotor cortex region of the user. 5. The method of claim 1, further comprising detecting desynchronization in neural oscillations of the user, and modulating the electrical stimulation treatment in association with the desynchonization. 6. The method of claim 1, wherein receiving the signal stream characterizing the neurological state comprises receiving signals indicative of a learning state of the user in association with performance of the task. 7. The method of claim 6, wherein receiving the signal stream characterizing the neurological state of the user comprises receiving signals indicative of neural plasticity, derived from more than one of: a stage of the task, galvanic skin response data, eye tracking data, and heart rate data. 8. The method of claim 1, wherein delivering the electrical stimulation treatment comprises delivering the el... | Halo Neuro, Inc. | A61N001/36 A61N001/04 A61B005/0205 A61B005/00 A61B005/16 A61B005/04 A61N001/08 A61N002/00 A61B005/0476 A61N001/05 A61B005/0488 A61B005/0496 A61B005/053
US09801571 | B2 | 2017-10-31 | Method and apparatus for providing data processing and control in medical communication system | 2013-09-16 | Methods and apparatus for providing data processing and control for use in a medical communication system are provided. | 1. A method of determining glucose information, comprising: receiving a signal associated with a monitored glucose level from an in vivo glucose sensor;retrieving stored signals associated with the monitored glucose level;determining a number of corrupt signals of the retrieved stored signals and a number of non-corrupt signals of the retrieved stored signals;determining that the retrieved stored signals do not exceed a preset tolerance level such that the number of corrupt signals is not more than a predetermined number;determining a rate of change of the monitored glucose level based on the received signal associated with the monitored glucose level and the non-corrupt signals of the retrieved stored signals; anddelivering medicament using a medical device based on the determined rate of change. 2. The method of claim 1, wherein the rate of change is determined based on glucose sensor sensitivity. 3. The method of claim 1, including displaying the determined rate of change. 4. The method of claim 1, including modifying a display of the determined rate of change. 5. The method of claim 4, wherein modifying the display of the determined rate of change includes modifying a display of a trend arrow indicator. 6. The method of claim 5, wherein modifying the display of the trend arrow indicator includes displaying a change in the trend arrow indicator position. 7. The method of claim 6, wherein the trend arrow indicator position includes one of five or less displayed positions. 8. The method of claim 1, wherein the retrieved stored signals and the received signal associated with the monitored glucose level are temporally separated. 9. An apparatus, comprising: a data storage unit; anda data processing unit operatively coupled to the data storage unit including programming to receive a signal associated with a monitored glucose level from an in vivo glucose sensor, retrieve from the data storage unit stored signals associated with the monitored glucose level, determine a number of corrupt signals of the retrieved stored signals and a number of non-corrupt signals of the retrieved stored signals, determine that the retrieved stored signals do not exceed a preset tolerance level such that the number of corrupt signals is not more than a predetermined number, determine a rate of change of the monitored glucose level based on the received signal associated with the monitored glucose level and the non-corrupt signals of the retrieved stored signals, and deliver medicament using a medical device based on the determined rate of change. 10. The apparatus of claim 9, further including a display operatively coupled to the data processing unit to display the determined rate of change. 11. The apparatus of claim 10, wherein the data processing unit includes programming to modify the display of the determined rate of change. 12. The apparatus of claim 11, wherein the data processing unit includes programming to modify the display of a trend arrow indica... | Abbott Diabetes Care Inc. | A61B005/145 A61B005/1473 A61B005/1495 A61B005/00 G06F019/00
US09814406 | B2 | 2017-11-14 | Method and system to identify motion data associated with consistent electrical and mechanical behavior for a region of interest | 2014-09-05 | Methods and system are provided that identify motion data associated with consistent electrical and mechanical behavior for a region of interest of the heart. The methods and systems acquire electrical cardiac signals indicative of physiologic behavior of at least a portion of the heart over a plurality of cardiac cycles. The methods and systems acquires motion data indicative of mechanical behavior of a motion sensor over the plurality of cardiac cycles to form a motion data collection, the motion data indicative of mechanical behavior of the region of interest when the motion sensor is in contact with the region of interest. The designating ectopic beats within the cardiac cycles may be based on the electrical cardiac signals, the ectopic beats producing electrically inconsistent (EI) data within the motion data collection. The methods and systems identify mechanically inconsistent (MI) data within the motion data collection based on irregular changes in the motion data. The methods and systems remove at least a portion of the EI and MI data from the motion data collection based on the designating and identifying operations to form an electrically/mechanically consistent (EMC) motion data collection. | 1. A method to identify motion data associated with consistent electrical and mechanical behavior for a region of interest of the heart, the method comprising: acquiring, via an electrical sensor, electrical cardiac signals indicative of physiologic behavior of at least a portion of the heart over a plurality of cardiac cycles;acquiring, via a motion sensor, motion data indicative of mechanical behavior of a motion sensor over the plurality of cardiac cycles to form a motion data collection, the motion data indicative of mechanical behavior of the region of interest when the motion sensor is in contact with the region of interest;designating, via a processor, ectopic beats or arrhythmias within the cardiac cycles based on the electrical cardiac signals, the ectopic beats producing electrically inconsistent (EI) data within the motion data collection;identifying, via the processor, mechanically inconsistent (MI) data within the motion data collection based on irregular changes in the motion data; andremoving, via the processor, at least a portion of the EI and MI data from the motion data collection based on the designating and identifying operations to form an electrically/mechanically consistent (EMC) motion data collection. 2. The method of claim 1, wherein the removing operation removes the EI data associated with ectopic beats and removes motion data associated with at least a portion of neighboring beats that occur at least one of before or after the ectopic beats. 3. The method of claim 1, wherein the removing operation removes the MI data associated with the irregular changes and at least a portion of adjacent motion data that occurs at least one of before or after the irregular changes. 4. The method of claim 1, wherein the identifying operation analyzes the motion data to identify changes in the motion data that exceed a predetermined threshold within a select period of time. 5. The method of claim 1, wherein the identifying operation analyzes the motion data to identify at least one of sudden or non-periodic changes in the motion data, the at least one of sudden or non-periodic changes being identified as MI data. 6. The method of claim 1, wherein the motion sensor is located on a medical tool in contact with the region of interest, the identifying operation analyzes a waveform representative of the motion data over time to identify waveform characteristics indicative of separation between the medical tool and the region of interest. 7. The method of claim 1, further comprising repeating the acquiring, designating, identifying and removing operations for each of multiple map points on the region of interest to form separate EMC motion data collections for each of the multiple map points. 8. The method of claim 1, wherein the acquiring, designating, identifying and removing operations are performed for a select map point, the method further comprising determining a motion characterization waveform for the map point based on the EM... | Pacesetter, Inc. | A61B005/04 A61B005/0468 A61B005/11 A61B005/00 A61B034/20
US09807519 | B2 | 2017-10-31 | Method and apparatus for analyzing and visualizing the performance of frequency lowering hearing aids | 2014-08-11 | A method of analyzing performance of frequency lowering hearing aids. The method includes generating a sequentially of noise signals and transmitting acoustical sounds from a sound output device in response to the sequence of noise signals. A sound input device records the acoustical sounds and saves as a first device data. The sound input device with a frequency lowering hearing aid records the acoustical sounds and save as a second device data. The second device data is compared to the first device data and, in response to the comparison, at least one function of the frequency lowering hearing aid is optionally adjusted. | 1. A method of analyzing performance of a frequency lowering hearing aid, the method comprising: generating a sequence of successive noise signals, wherein the sequence of successive noise signals comprises a plurality of one-third octave band noise signals, and wherein the sequence of successive noise signals covers a range of frequencies;transmitting acoustical sounds from a sound output device in response to the sequence of noise signals;recording the acoustical sounds with a sound input device and saving as a first device data;recording the acoustical sounds with the sound input device and a frequency lowering hearing aid and saving as a second device data; andcomparing the second device data to the first device data using an arrow-based visualization method comprising generating a single graphical display of changes in center frequencies and output levels of the transmitted noise signal of the first device data and the second device data comprising plotting values of the center frequencies and the output levels of the transmitted noise signal of the first device data, and extending an arrow from the plotted values of the center frequencies and the output levels of the transmitted noise signal of the first device data to values of the center frequencies and the output levels of the transmitted noise signal of the corresponding second device data. 2. The method of claim 1, further comprising modulating the plurality of one-third octave band noise signals with an envelope having random variations within a speech perception frequencies. 3. The method of claim 2 wherein said random variations within speech perception frequencies range from about 4 Hz to about 16 Hz. 4. The method of claim 1, further comprising: calibrating at least one of the sound output device, the sound input device, and the frequency lowering hearing aid. 5. The method of claim 1, wherein the center frequency for each of the first and second device data is identified by: windowing each of the first and second device data;transforming each of the windowed first and second device data to respective first and second frequency data;smoothing the first and second frequency data; andfinding a maximum within the smoothed first and second frequency data. 6. The method of claim 5, wherein the windowing includes applying a Hanning window. 7. The method of claim 5, wherein the smoothing of the frequency domain includes convolving a rectangular window. 8. The method of claim 1, further comprising: verifying a validity of each of said identified center frequency of the first and second device data. 9. The method of claim 8, wherein said verifying of the validity of each of said identified center frequencies comprises estimating the envelope of the recorded acoustical sounds of the first device data by squaring the waveform of said recorded acoustical sounds of the first device data and applying a low-pass filter having a frequency of 32 Hz;estimating the envelope of the recorded a... | The United States of America as Represented by the Secretary of Defense | H04R025/00 H04R029/00 A61B005/0484 A61B005/12 A61B005/00
US09782287 | B2 | 2017-10-10 | Method of securing a patient onto an operating table when the patient is in the Trendelenburg position and apparatus therefor including a kit | 2017-01-19 | A method of securing a patient onto an operating table when the patient is in the Trendelenburg position and apparatus therefor including a kit. A viscoelastic pad is used to support and hold a patient on a medical procedure table during a medical procedure performed while the table, and thus the patient lying thereon, is in an inclined position, such as the Trendelenburg position. The viscoelastic pad has characteristics which promote a minimization of pressure forces on the patient's body, as well as promote a secure cushioning and holding of the patient in a desired position on the table, in order to minimize injury to the patient. | 1. A single use memory pad patient support system configured to minimize injuries caused by pressure on portions of a body of a patient and minimize unwanted movement of a patient upon a patient being in a Trendelenburg position on a medical procedure table, such as a surgical operating table or a patient examination table, said single use memory pad patient support system comprising a pad;said pad comprising: a viscoelastic polyurethane foam comprising a coefficient of static friction of at least 0.2 and said coefficient of static friction being sufficient to assist in minimizing movement of said viscoelastic pad on said medical procedure table upon said patient being on said medical procedure table;sufficient thickness and viscosity to sufficiently cushion a body of a patient to: minimize bottoming out or prevent bottoming out, on said medical procedure table, of one or more of the portions of a body of a patient during positioning of a patient and during a medical procedure, and to minimize injuries from pressure during a medical procedure; andsufficient compliance to conform to a substantial portion of the body of a patient;sufficient thinness to stabilize a patient on a medical procedure table upon a patient being in an inclined, Trendelenburg position; andsufficient thickness and sufficient compliance to permit formation of a cavity in said pad of a depth sufficient to at least one of: assist in holding a patient on a medical procedure table, and assist in minimizing undesired movement of the body of a patient on a medical procedure table, during a medical procedure performed while said patient is in an inclined Trendelenburg position, such that: a) the head of the patient is disposed lower than the torso of the patient, or b) the head of the patient is disposed higher than the torso of the patient, or c) the right side of the patient is disposed higher than the left side of the patient, or d) the left side of the patient is disposed higher than the right side of the patient, or e) a combination of a), b), c), and d);wherein said viscoelastic memory foam pad is configured to generate a static frictional force on the skin of a patient on said medical procedure table pad, and further said pad is configured to generate a friction or friction-like force due to the formation of said cavity in said viscoelastic foam pad, and which force generated by said coefficient of static friction both together being sufficient to hold the body of the patient in the inclined Trendelenburg position;wherein a holding quality or ability between the Trendelenburg Pad and a patient is a combination of the coefficient of friction between the patient and the viscoelastic foam and the holding ability of the impression made by the patient in the viscoelastic foam, since the Trendelenburg pad comprises a viscoelastic foam in which an impression is formed and held in the foam at least for a time after the patient is moved from one position to another or from the operati... | - | A61F005/30 A61G013/04 A61G013/12 A61G007/10 A61B005/00 A61F005/37
US09831985 | B2 | 2017-11-28 | Close proximity communication device and methods | 2015-09-29 | Disclosed herein are methods and systems for receiving an encoded data packet, one or more activation commands, and a communication identifier, decoding the received data packet, validating the decoded received data packet, and executing one or more routines associated with the respective one or more activation commands. | 1. A method, comprising: wirelessly receiving a data packet from a data processing unit when the data processing unit is positioned within a close proximity transmission range of a sensor unit, wherein the received data packet includes an activation command;validating the received data packet; andexecuting one or more routines associated with the received activation command,wherein the executed one or more routines activate electronics of the sensor unit and control the communication or processing of analyte related data. 2. The method of claim 1, including executing the one or more routines when the received data packet is validated. 3. The method of claim 1, wherein validating the received data packet includes comparing a communication identifier in the data packet with a stored value. 4. The method of claim 1, wherein validating the received data packet includes performing error detection on the data packet. 5. The method of claim 4, wherein performing error detection on the data packet includes comparing one or more data patterns in the received data packet. 6. The method of claim 4, wherein performing error detection on the data packet includes performing error detection based on one or more error detection bits. 7. The method of claim 6, wherein the received data packet includes the one or more error detection bits. 8. The method of claim 6, wherein the one or more error detection bits includes a predetermined bit pattern. 9. The method of claim 1, wherein the one or more routines include a power on routine, a data transfer initiation routine, or a data transfer disable routine. 10. The method of claim 1, wherein the close proximity transmission range is a distance of one foot or less. 11. An apparatus, comprising: a wireless communication interface adapted to receive a data packet from a data processing unit when the data processing unit is positioned within a close proximity transmission range;one or more processing units coupled to the wireless communication interface; anda memory for storing instructions which, when executed by the one or more processing units, causes the one or more processing units to: validate the received data packet, wherein the received data packet includes an activation command; andexecute one or more routines associated with the received activation command to control the communication or processing of analyte related data. 12. The apparatus of claim 11, wherein the one or more processing units process a communication identifier received from the data processing unit. 13. The apparatus of claim 11, wherein the one or more routines include a power on routine, a data transfer initiation routine, or a data transfer disable routine. 14. The apparatus of claim 11, wherein the one or more processing units include an application specific integrated circuit (ASIC). 15. The apparatus of claim 11, wherein the close proximity transmission range is a distance of one foot or less. 16. An apparatus, comprising: ... | ABBOTT DIABETES CARE INC. | H04L012/28 H04L001/00 A61B005/00 A61B005/145
US09817481 | B2 | 2017-11-14 | Methods, systems, and apparatuses to display visibility changes responsive to user gestures | 2015-10-06 | In one embodiment, an electronic device to be worn on a user's forearm includes a display and a set of one or more sensors that provide sensor data. In one aspect, a device may detect, using sensor data obtained from a set of sensors, that a first activity state of a user is active. The device may determine, while the first activity state is active, that the sensor data matches a watch check rule associated with the first activity state. Responsive to the detected match, the device may cause a change in visibility of the display. | 1. An apparatus comprising: a display;a set of one or more sensors configured to generate sensor data during a first time period and a second time period;a set of one or more processors coupled to the display and the set of one or more sensors; anda non-transitory machine readable storage medium coupled to the set of one or more processors and having stored therein a watch check gesture detector for receiving the sensor data, the watch check gesture detector comprising a watch check rule for causing a change to the display, wherein the watch check rule includes data or logic that represents one or more patterns of movement indicative of a user of the apparatus checking the display,wherein the storage medium further comprises instructions that, when executed by the set of one or more processors, cause the set of one or more processors to: detect, using the sensor data, that the user is performing a particular type of activity;enable the watch check rule based on the detection;determine, based on the sensor data generated during the first time period and while a first activity metric is displayed on the display, that the user has made a bodily adjustment matching the one or more patterns of movement indicative of the user checking the display such that the watch check rule is satisfied;in response to determining that the watch check rule is satisfied, cause a first change to the display, the first change causing the display displaying the first activity metric to display a second activity metric different from the first activity metric and indicative of the user's performance with respect to the particular type of activity;log first accelerometer data extracted from the sensor data generated during the first time period, the first accelerometer data comprising a first x-axis value and a first y-axis value;extract second accelerometer data from the sensor data generated during the second time period, the second accelerometer data comprising a second x-axis value and a second y-axis value;determine whether the second x-axis value of the second accelerometer data exceeds the first x-axis value of the logged first accelerometer data by a first threshold amount;determine whether the second y-axis value of the second accelerometer data exceeds the first y-axis value of the logged first accelerometer data by a second threshold amount; andin response to determining that (i) the second x-axis value exceeds the first x-axis value by the first threshold amount and (ii) the second y-axis value exceeds the first y-axis value by the second threshold amount, cause a second change to the display, the second change causing the display displaying the second activity metric to display the first activity metric. 2. The apparatus of claim 1, wherein the set of one or more sensors is further configured to generate sensor data during a third time period, and the instructions, when executed by the set of one or more processors, further cause the set of one or more proces... | Fitbit, Inc. | G06F003/01 G09G005/10 A61B005/00 G01P015/00
US09798862 | B2 | 2017-10-24 | Apparatus for dispensing medicine, vitamins and/or samples | 2012-11-28 | An apparatus for distributing medicines, vitamins and/or samples, including an organizer and dispensing units for medicines, vitamins and/or samples. The organizer includes seats in which to place the dispensing units. The apparatus includes an identification device connected to a data processing device and containing dispensing data on medicines, vitamins and/or samples, including detecting members controlled to a detecting state by detecting signals generated on the basis of dispensing data in the data processing device. The identification device is configured to be brought in connection with the organizer so that at least some of the detecting members are aligned with at least some of the seats of the organizer, wherein at least those dispensing units, which are aligned with at least one detecting member, can be identified by means of detecting signals generated on the basis of dispensing data in the data processing device. | 1. An apparatus for dispensing medicines, vitamins and/or samples, the apparatus comprising: an organizer,dispensing units for the medicines, vitamins and/or samples to be placed on the organizer, the organizer comprising seats in which the dispensing units can be placed, andan identification device without a data processing device and that enables a data processing device outside of the apparatus to control the apparatus through the identification device when the identification device is connected with the data processing device such that the data processing device can provide dispensing data on the medicines, vitamins and/or samples, wherein: the identification device comprises detecting members to be controlled to a detecting state by means of detecting signals to be generated on the basis of dispensing data in the data processing device,the identification device is separable from the organizer and is configured to be brought in connection with the organizer in such a way that at least some of the detecting members are aligned with at least some of the seats of the organizer,at least those dispensing units, which are aligned with at least one detecting member, can be identified by means of detecting signals generated on the basis of dispensing data in the data processing device, andthe organizer comprises an organizer tag and the identification device comprises a reader device connected to the data processing device outside of the apparatus and configured to identify the organizer connected to the identification device on the basis of the organizer tag, when the organizer is placed in connection with the identification device and the identification device is connected to the data processing device. 2. The apparatus according to claim 1, wherein the detecting members are, in the detecting state, light emitting detecting members which are arranged to be set in the detecting state when they receive a detecting signal. 3. The apparatus according to claim 2, wherein each seat is provided with a maximum of one detecting member. 4. The apparatus according to claim 3, wherein the detecting members are configured to detect whether the seat is occupied by a separate dispensing unit or whether the seat is empty. 5. The apparatus according to claim 2, wherein the detecting members are light emitting diodes and that the detecting signal is an electric current that lights the light emitting diode. 6. The apparatus according to claim 5, wherein the seats in the organizer are openings formed in the organizer, and that the dispensing units are configured to be supported to the seats in such a way that they come above the detecting members of the identification device. 7. The apparatus according to claim 1, wherein the identification device is a plate-like or at least partly plate-like piece, from whose upper surface upwards protrude detecting members spaced from each other, and that the upper part of the identification device comprises an adapter part, i... | NEWICO OY | G06F019/00 G07F017/00 A61J001/03 A61J007/00 B01L009/06 B01L003/00 A61B005/15
US09788785 | B2 | 2017-10-17 | Apparatus and methods for estimating time-state physiological parameters | 2016-12-05 | A method of determining a value of a physiological parameter for a subject at a selected state includes obtaining, via a device located a distance from the subject, a value of the physiological parameter of the subject at a particular time-of-day, and applying a time-dependent relationship function to the obtained physiological parameter value via a processor to determine a value of the physiological parameter at the selected state. | 1. A method of estimating a value of a physiological parameter for a subject at a selected metabolic state, the method comprising: obtaining, via a non-wearable device located at a stand-off distance from the subject, a value of the physiological parameter of the subiect at a time-of-day, wherein the non-wearable device comprises a sensor configured to measure heart rate, breathing rate, and/or physical activity of the subject, wherein the sensor is an electromagnetic energy sensor, an electrical energy sensor, a magnetic energy sensor, an inductive energy sensor, a capacitive energy sensor, a thermal energy sensor, or an acoustic energy sensor; andestimating the value of the physiological parameter for the subject at the selected metabolic state by applying a time-dependent relationship function to the obtained physiological parameter value via at least one processor, wherein the selected metabolic state is different from a current metabolic state of the subject. 2. The method of claim 1, wherein the time-dependent relationship function is derived, from a circadian rhythm for the subject. 3. The method of claim 1, wherein the time-dependent relationship function is a lookup table. 4. The method of claim 1, wherein the physiological parameter includes one or more of the following: subject body temperature, subject heart rate, subject heart rate variability, subject blood gas levels, subject metabolic rate, subject respiration rate, subject blood analyte levels, subject blood pressure, and subject pulse pressure. 5. The method of claim 1, wherein estimating the value of the physiological parameter for the subject at the selected metabolic state comprises obtaining the value of the physiological parameter at the same time-of-day for multiple days and determining an average value for the multiple obtained values, and wherein applying the time-dependent relationship function to the obtained physiological parameter value comprises applying the time-dependent relationship function to the average value. 6. The method of claim 1, further comprising determining if the subject is in a condition of heightened activity at the time-of-day by determining via the non-wearable device if at least one obtained physiological parameter value is at a level associated with heightened activity, and wherein the time-dependent relationship function is adjusted for the heightened activity condition of the subject prior to estimating the value of the physiological parameter at the selected metabolic state. 7. The method of claim 1, wherein prior to obtaining a value of the physiological parameter of the subject at the time-of-day, values of the physiological parameter of the subject are obtained at multiple times during at least one previous day, and a personalized time-dependent relationship function between at least one of value of the physiological parameter and a value of the physiological parameter at the selected metabolic state is generated for the subject usi... | Valencell, Inc. | A61B005/04 A61B005/00 A61B005/0205 A61B005/08 A61B005/1455 A61B005/11 A61B005/145 A61B005/021 A61B005/024
US09795319 | B2 | 2017-10-24 | Method and device for navigation of a surgical tool | 2013-11-12 | The invention relates to a method of navigation of a tool to be used in a surgical operation for treating a target region within the body of a patient. | 1. A method of navigation of a linear instrument to be used in a surgical intervention for treating a target region within the body of a patient, comprising detecting by an imaging system an image position of a reference marker disposed on the body of the patient in a vicinity of an entry point corresponding to a point through which the linear instrument should be inserted for reaching the target region, and detecting a spatial position of said reference marker by a localizing system,calculating a transform matrix between an image coordinate system of the imaging system and a spatial coordinate system associated with the reference marker by using detected image and spatial positions of the reference marker, in order to register data of the localizing system and imaging system together,applying a correction matrix to said transform matrix so as to determine a corrected transform matrix,wherein said correction matrix is computed by the following steps:obtaining image data of a zone of the body of the patient with the imaging system, said zone including the target region to be treated and at least a part of the linear instrument;displaying a real position of said part of the linear instrument as a linear trace visible on the image data;determining a virtual position of the linear instrument with the localizing system using a magnetic tracker attached to the linear instrument or to a guide of the linear instrument; andcomputing the correction matrix so as to make a position of the linear instrument on a navigation image correspond to a position of a selected linear shape using said linear trace and said virtual position;detecting a spatial position of the linear instrument determined by the localizing system; andconstructing and displaying the navigation image, wherein a representation of the spatial position of the linear instrument is displayed on said image data from the imaging system using the corrected transform matrix. 2. The method of claim 1, wherein the step of obtaining image data of the zone of the body of the patient is regularly and automatically performed, and wherein displaying the real position, determining, computing, detecting the spatial position, constructing and displaying steps are performed using the last obtained image data of the imaged zone. 3. The method of claim 1, wherein the localizing system is a magnetic tracking system based on magnetic detection and the imaging system is based on X-ray Computed Tomography. 4. The method of claim 3, wherein the reference marker comprises: a magnetic element having at least two coils arranged in a manner to define at least two vectors in the coordinate system of the magnetic tracking system; anda geometric element comprising components visible on image data of the imaging system and having a unique geometry detectable by the imaging system. 5. The method of claim 4, wherein the at least two coils are coils arranged in a different manner with respect to one another, and the compone... | IMACTIS | A61B005/06 A61B006/03 A61B017/34 A61B017/00 A61B090/00 A61B034/20
US09778111 | B2 | 2017-10-03 | Method for estimating a cardiac frequency and associated device | 2017-02-28 | The invention is a method for estimating a cardiac frequency via the detection of radiation backscattered or transmitted by a bodily zone. The part is illuminated, simultaneously or successively, by light radiation extending over a first spectral band and a second spectral band. A photodetector detects radiation emitted by the bodily zone under the effect of its illumination, in each of the spectral bands. A first detection function and a second detection function are formed from the radiation detected in each spectral band, respectively. The method allows the cardiac frequency to be determined via the determination of characteristic instants that are identified from the first detection function and the second detection function simultaneously. | 1. A Method for estimating a cardiac frequency of a living being including the following steps: a) illuminating a bodily zone of the living being with an incident light beam in a first spectral band;b) detecting light radiation transmitted or backscattered, in the first spectral band, by the bodily zone under the effect of the illumination;c) determining a first detection function, representing a variation as a function of time of an intensity of the light radiation thus detected;d) identifying characteristic instants from the first detection function, and calculating an occurrence frequency of the characteristic instants;e) estimating a cardiac frequency from the occurrence frequency calculated in step d);wherein: steps a) to b) are also implemented in a second spectral band, such that step c) includes determining a second detection function representing a variation as a function of time of an intensity of the light radiation detected in the second spectral band;step d) includes identifying characteristic instants from the second detection function and selecting characteristic instants, identified from each detection function, and appearing in temporal coincidence, the occurrence frequency being calculated from the characteristic instants thus selected. 2. The method according to claim 1, wherein the second spectral band is different from the first spectral band. 3. The method according to claim 1, wherein step d) includes calculating a first derived function, derived from the first detection function, and a second derived function, derived from the second detection function, and identifying characteristic instants from each of the derived functions. 4. The method according to claim 3, wherein each derived function is obtained via a difference between the value of a detection function at two different instants. 5. The method according to claim 1 wherein: the first spectral band includes wavelengths comprised between 600 and 700 nm;the second spectral band includes wavelengths comprised between 750 nm and 1 μm. 6. The method according to claim 1, wherein the first spectral band extends between 600 nm and 700 nm, whereas the second spectral band extends between 750 nm and 1 μm. 7. The method according to claim 1, wherein the detected radiation is radiation backscattered by the bodily zone under the effect of its illumination. 8. A device for estimating a cardiac frequency of a living being, including: a light source configured to emit an incident light beam that propagates towards a bodily zone of the living being, in a first spectral band and in a second spectral band;a photodetector, configured to detect, in the first spectral band and in the second spectral band, radiation backscattered or transmitted by the bodily zone under the effect of its illumination by the incident light beam;a processor, configured to process the radiations detected by the photodetector and to implement the steps c) to e) of the method according to claim 1. 9. T... | Commissariat A L'Energie Atomique et aux Energies Alternatives | G01J003/00 G01J003/427 G01N021/359 G01N021/49 G01N021/59 A61B005/024 G01N021/31
US09827047 | B2 | 2017-11-28 | Control apparatus and control method of insertion apparatus, insertion apparatus having control apparatus, control program for insertion apparatus, and controlling integrated electronic circuit of insertion apparatus | 2014-06-02 | A control apparatus of an insertion apparatus in which a state specifying unit specifies a state of an insertion member to carry out vibration control depending on the state, thereby removing getting-stuck of a tip of the insertion member with a vibration having such a proper magnitude as not to cause an overload in a case where the tip of the insertion member is gotten stuck. | 1. A control apparatus of an insertion apparatus configured to insert an insertion member of a catheter into a body lumen, the control apparatus comprising: a tip movement information acquiring unit that acquires tip movement information indicative of presence of movement in the body lumen of a tip of the insertion member;a non-tip area movement information acquiring unit that acquires non-tip area movement information indicative of presence of a moving operation in the body lumen of a non-tip area other than the tip of the insertion member;a state specifying unit that, by using the non-tip area movement information acquired by the non-tip area movement information acquiring unit and the tip movement information acquired by the tip movement information acquiring unit,(a) specifies as a first tip stop state for vibrating the insertion member with a first magnitude of a vibration, a case where the non-tip area of the insertion member is moved and the tip of the insertion member is not moved, and(b) specifies as a second tip stop state for vibrating the insertion member with a second magnitude of a vibration greater than the first magnitude of the vibration in the first tip stop state of (a), a case where the non-tip area of the insertion member is not moved and the tip of the insertion member is not moved;an operation information generating unit that generates operation information about the insertion apparatus to execute the first tip stop state of (a) or the second tip stop state of (b) which is specified by the state specifying unit; anda control unit that controls an operation of the insertion member based on the operation information generated by the operation information generating unit,wherein a maximum value of the second magnitude of the vibration in the second tip stop state is set to be an amplitude of 0.3 mm, andthe control unit controls the operation of the insertion member by applying vibration control to execute the vibrating of the insertion member in the first tip stop state of (a) or the second tip stop state of (b). 2. The control apparatus of an insertion apparatus according to claim 1, wherein the state specifying unit specifies the first tip stop state of (a) and the second tip stop state of (b) by using the non-tip area movement information acquired by the non-tip area movement information acquiring unit and the tip movement information acquired by the tip movement information acquiring unit, and specifies as a state in which the insertion member is not vibrated, a case where the tip of the insertion member is moved regardless of the presence of the moving operation in the non-tip area of the insertion member, andthe operation information generating unit generates the operation information about the insertion apparatus for executing any of the states specified by the state specifying unit. 3. The control apparatus of an insertion apparatus according to claim 1, wherein the state specifying unit specifies the first tip stop ... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | A61B001/00 A61B001/04 A61B019/00 A61B005/06 G05B019/423 A61B034/32 A61M025/01 A61B017/22 A61B034/30
US09826937 | B2 | 2017-11-28 | Method and apparatus for reducing motion artifacts in ECG signals | 2013-12-13 | The present invention provides a method and apparatus for reducing motion artifacts in ECG signals. According to an aspect of the present invention, there is proposed a method of reducing motion artifacts in ECG signals, comprising: acquiring a current beat from a continuously measured ECG signal of a patient; calculating a correlation coefficient between a previous mean value beat and the current beat in the ECG signal; determining the weights to be assigned to the previous mean value beat and the current beat based on the correlation coefficient; and calculating a current mean value beat based on the previous mean value beat, the current beat, and the weights thereof. Accordingly, the novel method of deriving the current mean value beat may reduce ECG artifacts due to patient movement in such a manner that the SNR of the ECG signal can be improved substantially. | 1. A method of reducing motion artifacts in ECG signals, comprising: acquiring, via an acquiring unit, a current beat from a continuously measured ECG signal of a patient;calculating, via a first calculating unit, a correlation coefficient between a previous mean value beat and the current beat in the ECG signal;determining, via a determining unit, weights to be assigned to the previous mean value beat and the current beat based on the correlation coefficient; andcalculating, via a second calculating unit, a current mean value beat based on (i) the previous mean value beat, (ii) the current beat, and (iii) the determined weights respectively assigned to the previous mean value beat and the current beat, wherein the current mean value beat comprises a motion artifact reduced beat having reduced motion artifacts over the current beat in the ECG signal. 2. The method according to claim 1, further comprising a step of comparing, via a comparator, the correlation coefficient with a predetermined coefficient. 3. The method according to claim 1, further comprising: subtracting, via a subtracting unit, the current mean value beat from the current beat to obtain a residual signal;performing, via a filter unit, piecewise filtering on the residual signal based on signal characteristics of the ECG signal to obtain a filtered signal; andadding, via an adding unit, the filtered signal back to the current mean value beat to obtain a modified current beat. 4. The method according to claim 3, further comprising: extracting, via an extracting unit, from the ECG signal, the signal characteristics, wherein the signal characteristics include (i) a noise envelope, (ii) correlation information, (iii) peak positions of P wave, QRS complex wave, and T wave, and (iv) segment intervals, andselecting, via a selection unit, (i) one of P wave, QRS complex wave, and T wave, or (ii) any combination thereof, as the current beat to be processed. 5. The method according to claim 4, wherein, responsive to selecting the QRS complex wave as the current beat to be processed, the step of performing piecewise filtering further comprises the sub-steps of: identifying, via an identification unit, a QRS segment and T and P segments based on the signal characteristics; andperforming filtering on the QRS, T and P segments, respectively, via different filters. 6. The method according to claim 1, further comprising deriving, via a derivation unit, an initial mean value beat from a predetermined number of consecutive beats in the ECG signal. 7. The method according to claim 6, wherein the predetermined number of consecutive beats comprises (i) first predetermined number of consecutive beats and (ii) a second predetermined number of consecutive beats following the first predetermined number of consecutive beats, and wherein the step of deriving an initial mean value beat further comprises the sub-steps of: deriving an initial mean value beat from the first predetermined number of consecuti... | KONINKLIJKE PHILIPS N.V. | A61B005/04 A61B005/00 A61B005/0452
US09795336 | B2 | 2017-10-24 | Device for measuring urinary detrusor pressure | 2014-04-18 | The present invention relates to a device for measuring bladder pressure, comprising: a measurement rod inserted into the abdomen of a user to measure both abdomen pressure and bladder pressure; and a body unit provided to be worn on the waist of the user and electrically connected to the measurement rod so as to calculate and store urinary detrusor pressure measured by subtracting a measurement abdomen pressure from a measurement bladder pressure measured by the measurement rod. | 1. A device for measuring bladder pressure, comprising: a measurement rod configured to be inserted into an abdomen of a user, to measure both abdominal pressure and bladder pressure; anda body unit configured to be worn on a waist of the user, and be electrically connected with the measurement rod, to calculate and store detrusor pressure obtained by subtracting the measured abdominal pressure from the bladder pressure measured by the measurement rod,wherein the measurement rod comprises:a rod body having a predetermined length;a bladder pressure measurement unit configured to be provided at a distal end of the rod body, to measure the bladder pressure of the user; andan abdominal pressure measurement unit configured to be spaced apart from the bladder pressure measurement unit, to measure the abdominal pressure of the user,wherein the abdominal pressure measurement unit comprises a plurality of abdominal pressure measurement units spaced apart from each other. 2. The device of claim 1, further comprising a guide tube configured to penetrate and be inserted into the abdomen of the user, and receive the measurement rod within it, to guide the measurement rod. 3. The device of claim 2, wherein the body unit comprises: a body having a display unit and an input unit;an adapter provided on the backside of the body to come into contact with the abdomen of the user, to be electrically connected with the measurement rod; anda belt coupled to both sides of the body, and configured to be worn on the waist of the user. 4. The device of claim 3, further comprising a wireless control unit configured to be separate from the body unit, to enable an inputting person to input signals using buttons. 5. The device of claim 3, wherein the measurement rod further comprises: a connector configured to be provided at a proximal end of the rod unit, and to be connected to the adapter. 6. The device of claim 3, further comprising a length control bracket configured to be screw-coupled with the backside of the body, to control depth of insertion of the measurement rod; wherein the position within the abdomen, into which the abdominal pressure measurement unit is inserted, is controlled by controlling the screw coupling length between the length control bracket and the body. 7. The device of claim 2, wherein a measurement unit exposing hole having a predetermined size is formed penetrating through the outer circumferential surface of the guide tube at a position corresponding to the position of the abdominal pressure measurement unit. 8. The device of claim 2, wherein the guide tube is configured such that a sharp needle is formed at the distal end of the guide tube and a discharge slit from which the measurement rod can be separated is formed along the axial direction of the guide tube to have a diameter larger than that of the measurement rod. 9. The device of claim 2, wherein the guide tube is formed of a polymer material or a metal material. 10. The device of ... | GACHON UNIVERSITY OF INDUSTRY-ACADEMIC COOPERATION FOUNDATION GIL MEDICAL CENTER | A61B005/00 A61B005/20 A61B005/0488 A61B017/34 A61B017/00
US09826907 | B2 | 2017-11-28 | Wearable electronic device for determining user health status | 2013-12-28 | Particular embodiments described herein provide for a wearable electronic device, such as a bracelet, coupled to a plurality of electronic components (which may include any type of components, elements, circuitry, etc.). One particular implementation of a wearable electronic device may include a plurality of sensors configured to measure at least one health parameter of a first user associated with the wearable electronic device, and a control module in communication with the plurality of sensors. The control module includes a processor configured to receive a plurality of health parameter measurements from at least a subset of the plurality of sensors, and determine a general health state of the first user based upon the received health parameter measurements. | 1. A wearable electronic device, comprising: a plurality of sensors configured to measure at least one health parameter of a first user associated with the wearable electronic device;a control module in communication with the plurality of sensors, the control module including a processor configured to: receive a plurality of different health parameter measurements from at least a subset of the plurality of sensors;determine an overall value for each of a plurality of particular health parameters by calculating a mean value of each received different health parameter measurement for the particular health parameter;determine a general health state of the first user based upon a combination of the determined overall values of each of the plurality of particular health parameters;send a first message including information indicative of the general health state to an electronic device associated with a second user, wherein the electronic device associated with the second user includes another wearable electronic device; andreceive a second message from the electronic device associated with the second user, wherein the second message includes a request for additional information from the first user; anda display configured to provide feedback indicative of the determined general health state to the first user and the second received message. 2. The wearable electronic device of claim 1, wherein the processor is further configured to send a third message to the electronic device associated with the second user, wherein the third message includes the requested additional information from the first user. 3. The wearable electronic device of claim 1, wherein the processor is further configured to determine the general health state of the first user by determining whether the overall value of each health parameter measurement is at an acceptable level. 4. The wearable electronic device of claim 3, wherein the processor is further configured to provide feedback to the first user indicating that the first user has achieved an optimal health state if each of the overall health parameter measurements are determined to be at an acceptable level. 5. The wearable electronic device of claim 1, further comprising a strap portion, wherein the plurality of sensors are disposed upon a portion of a surface of the strap portion. 6. The wearable electronic device of claim 5, wherein the plurality of health measurements are associated with one or more health parameters of the user, wherein each health parameter is configured to be measured at a plurality of locations upon the surface of the strap portion. 7. The wearable electronic device of claim 1, wherein the plurality of sensors comprise a sensor array. 8. The wearable electronic device of claim 1, wherein at least one of the plurality of sensors includes at least one of a skin sensor, a temperature sensor, a pulse sensor and a galvanic skin response sensor. 9. The wearable electronic device of claim 1, wherein... | Intel Corporation | G08B001/08 A61B005/0205 A61B005/145 A61B005/00 A61B005/021 A61B005/024 A61B005/053
US09818285 | B2 | 2017-11-14 | Method and apparatus for moving while receiving information | 2014-03-18 | A method and apparatus to facilitate learning while moving, efficient multitasking involving movement while the user processes or responds to different stimuli. The stimuli may include but are not limited to information related to education or entertainment or feedback concerning the user's movement. More specifically this movement may be related to coordination, exercise or physiotherapy. The methods and apparatuses of the disclosure involve the user conducting movement, while simultaneously processing information (via: learning; creating through typing, moving, or talking; or being entertained) and receiving feedback or assistance related to that movement, processing, or any combination thereof. | 1. A method of receiving sensory feedback while moving involving: one or more sensor relays;any number of output devices;any number of movement devices; andone or more computer processors that:receive a signal from any of the one or more sensor relays, detecting the information related to the user;analyze the detected information against one or more default or user defined set points;and either:send a signal to any of the output devices, instructing the decision maker with feedback related to one or more control factors measured against one or more set points;orsend a signal to any of the movement devices, controlling said movement devices based on said analysis. 2. The method described in claim 1 wherein the user controls, is educated by, or is entertained by a media device. 3. The method described in claim 1 further comprising a moving walkway in a room for use by multiple users. 4. An apparatus comprising: one or more sensor relays;and either:one or more output devices;orone or more movement devices;wherein said sensor relay may detect information related to one or more user control factors;andwherein said sensor relay may send one or more signals to one or more other sensor relays, output devices or movement devices;said output devices translating and communicating said signals to a decision maker as measured against one or more default or user defined set points;said movement devices altering based on said signal measured against one or more default or user defined set points;orsaid other sensor relays detecting information related to the user and sending a signal to any of the one or more other sensor relays, output devices, or movement devices, or relaying the signal sent to it, to one or more other sensor relays, output devices, or movement devices measured against one or more default or user defined set points. 5. An apparatus as described in claim 4 further comprising one or more computer processors;wherein said sensor relays may detect information related to the user;wherein said sensor relay may send one or more signals to any of the one or more computer processors;said computer processors may then analyze the detected information against any number of default or user defined set points;and either:send a signal to any of the output devices, instructing the decision maker;orsend a signal to any of the movement devices, controlling said movement devices. 6. An apparatus as described in claim 5, further comprising one or more of output devices;andone or more movement devices. 7. The method described in claim 5 further comprising a moving walkway in a room for use by multiple users. 8. The method described in claim 5 wherein feedback from said output device allows the user to alter the media deviceorfeedback from said computer processor is sent to the media device, thereby automatically altering it. 9. A method described in claim 8 wherein one or more output devices is a headset comprising the steps of: having one or more users in... | Forward Entertainment & Technology, LLC | G08B021/18 B29D023/00 B66B003/00 F16L011/10 F16L011/12 A61B005/11 A61B005/00 A61B005/22
US09808160 | B2 | 2017-11-07 | Systems and methods for monitoring and analyzing cardiovascular states | 2014-08-12 | A system and method for hemodynamic dysfunction detection may include at least one sensor configured to received one or more signals from a patient, a computing device in data communication with the at least one sensor, a computer-readable storage medium in communication with the computing device, an input device, and an output device. The system may include computer readable instructions to cause the system to receive at least one signal in the time domain from the sensor, determine at least one metric in the frequency domain from the at least one signal in the time domain, and determine the cardiovascular state of the patient from a combination of the at least one metric in the frequency domain and information contained in at least one database of cardiovascular states. The system may also notify a user of a immanent patient cardiovascular event and recommend one or more interventions to mitigate it. | 1. A system for determining a cardiovascular state of a patient in a stress condition, the system comprising: at least one sensor configured to received one or more signals from the patient;a computing device in data communication with the at least one sensor; a non-transitory, computer-readable storage medium in operable communication with the computing device;an input device in operable communication with the computing device; andan output device in operable communication with the computing device;wherein the computer-readable storage medium contains one or more programming instructions that, when executed, cause the computing device to: receive at least one signal in the time domain from the at least one sensor, wherein the at least one signal in the time domain is a pulse wave measurement of the patient;determine at least one metric in the time domain from the at least one signal in the time domain;determine at least one metric in the frequency domain from the at least one signal in the time domain;derive one or more sets of parameters using the at least one metric in the time domain and the at least one metric in the frequency domain, a set of parameters being associated with a particular cardiovascular state;evaluate temporal variations of the parameters within the one or more sets of parameters; anddetermine the cardiovascular state of the patient from a combination of the temporal variations of parameters in a set of parameters and information from at least one database of cardiovascular states that are associated with the sets of parameters. 2. The system of claim 1, wherein the input device comprises one or more of a keyboard, a mouse, a touch screen, a joystick, a voice recognition system, a removable memory device, and a data communication link. 3. The system of claim 1, wherein the output device comprises one or more of a monitor, a flat-screen display, one or more LED devices, an audio device, and a data communication link. 4. The system of claim 1, wherein the computer-readable storage medium further contains the at least one database. 5. The system of claim 1, wherein the at least one database is contained in a memory storage device accessible to the computing device over a data communication link. 6. The system of claim 5, wherein the data communication link comprises one or more of an internet connection, a wireless connection, a telephonic connection, a LAN connection, a WAN connection, and a personal area network. 7. The system of claim 1, wherein the computer-readable storage medium contains one or more programming instructions that, when executed, further cause the computing device to: transform the at least one signal in the time domain into at least one signal in the frequency domain;select at least one frequency domain feature of the at least one signal in the frequency domain; andnormalize the at least one frequency domain feature to at least one frequency domain feature baseline. 8. The system of claim 1, wherein... | Intelomed, Inc. | A61B005/05 A61B005/0205 A61B003/16 A61B005/021 A61B005/024 G06F019/00 A61B005/00 A61B005/053 A61B005/0452 A61B005/08 A61B005/16
USRE046562 | E1 | 2017-10-03 | Vascular image co-registration | 2015-06-01 | A system and method for providing a vascular image are disclosed wherein a single composite display simultaneously provides a first view of a patient including an angiogram image and a second view including an intravascular image rendered from information provided by an imaging probe mounted on a distal end of a flexible elongate member. A cursor, having a position derived from image information provided by a radiopaque marker proximate the imaging probe, is displayed within the angiogram image to correlate the position of the imaging probe to a presently displayed intravascular image and thus provide an easily discernable identification of a position within a patient corresponding to a currently displayed intravascular image. The resulting composite display simultaneously provides: an intravascular image that includes information about a vessel that is not available from an angiogram and a current location within a vessel of a source of intravascular image data from which the intravascular image is rendered. | 1. A system for acquisition and co-registered display of intravascular information, comprising: an imaging flexible elongate member having a proximal end and a distal end;an imaging probe located near the distal end of the flexible elongate member, and configured to obtain information for generating an image of a vessel;a radiopaque marker located near the imaging probe;a first memory for storing angiogram image data;a second memory for storing intravascular image data derived from information obtained by the imaging probe;a third memory for storing radiopaque marker image data, distinct from the angiogram image data, the radiopaque marker image data being derived from information obtained from a fluoroscopic imaging device;a display processor configured to retrieve and combine data from the first memory, the second memory and the third memory, and further configured to render a composite image including: an enhanced radiological image derived from the angiogram image data comprising a superimposition of at least a portion of the angiogram data and the radiopaque marker image data and providing a location of the radiopaque marker based upon an actual location during active fluoroscopy and an estimated location during inactive fluoroscopy,an intravascular image element corresponding to the intravascular image data, wherein the enhanced radiological image and the intravascular image element are displayed proximate each other; anda cursor, displayed upon the enhanced radiological image, indicative of a location of the imaging probe while acquiring data for the intravascular image element presently displayed on the composite image, said cursor having a position that is based at least in part on third data derived from the radiopaque marker image data stored in the third memory;wherein the display processor is further configured to calculate an error function based on a difference between the estimated location of the radiopaque marker and the actual location of the radiopaque marker and wherein the display processor further configured to utilize the error function to correct the estimated location of the radiopaque marker for the preceding period of inactive fluoroscopy. 2. The system of claim 1 wherein the flexible elongate member is a catheter. 3. The system of claim 1 wherein the imaging probe comprises an ultrasound device. 4. The system of claim 3 wherein the ultrasound device is a side-firing intravascular ultrasound transducer assembly. 5. The system of claim 4 wherein the side-firing intravascular ultrasound transducer assembly comprises an array of transducer elements. 6. The system of claim 5 wherein the array of transducer elements are linearly arranged along a lengthwise axis of the flexible elongate member. 7. The system of claim 5 wherein the array of transducer elements are curvilinearly arranged about a lengthwise axis of the flexible elongate member. 8. The system of claim 5 wherein the array of transducer elements are circumf... | Volcano Corporation | A61B005/05 A61B005/06 A61B006/00 A61B008/12 A61B008/00 A61B034/20 A61B005/0215
US09789302 | B2 | 2017-10-17 | Symmetrical physiological signal sensing with a medical device | 2010-09-01 | A physiological signal of a patient is sensed with sense electrodes symmetrically arranged relative to a stimulation electrode. In some examples, a member includes a plurality of relatively small electrodes that are configured to function as both sense and stimulation electrodes. One or more of the electrodes may be selected as stimulation electrodes and two or more different electrodes of the member may be selected as sense electrodes that are symmetrically arranged relative to the one or more selected stimulation electrodes. In some examples, a member includes a plurality of levels of segmented sense electrodes and a plurality of levels of stimulation electrodes. The levels of sense electrodes are arranged such that each level of stimulation electrodes is adjacent at least two levels of sense electrodes symmetrically arranged relative to the level of stimulation electrodes. | 1. A method comprising: with a processor, selecting a first subset of electrodes of a plurality of electrodes of a member as stimulation electrodes, wherein each electrode of the plurality of electrodes is configured to function as a sense electrode or a stimulation electrode;with the processor, selecting a second subset of electrodes of the plurality of electrodes of the member as sense electrodes, wherein the sense electrodes are symmetrically arranged relative to the first subset of electrodes, and wherein a line or plane of symmetry substantially bisects the first subset of electrodes in a predetermined direction;with the processor, controlling a stimulation generator of a medical device to deliver stimulation to a patient via the first subset of electrodes; andwith the processor, controlling a sensing module of the medical device to sense a physiological signal of the patient via the second subset of electrodes, wherein selecting the first subset of electrodes of the plurality of electrodes of the member as stimulation electrodes comprises, with the processor, selecting the first subset of electrodes based on information stored in a memory, the information indicating the first subset of electrodes and sense electrodes associated with the first subset of electrodes, and wherein selecting the second subset of electrodes of the plurality of electrodes of the member as sense electrodes comprises determining the sense electrodes associated with the first subset of electrodes in the memory. 2. The method of claim 1, wherein the predetermined direction is substantially perpendicular to a longitudinal axis of the member. 3. The method of claim 1, wherein selecting the first subset of electrodes of the plurality of electrodes of the member as stimulation electrodes comprises selectively electrically coupling the first subset of electrodes to the stimulation generator of the medical device. 4. The method of claim 1, wherein selecting the second subset of electrodes of the plurality of electrodes of the member as sense electrodes comprises selectively electrically coupling the second subset of electrodes to the sensing module of the medical device. 5. The method of claim 1, wherein the first subset of electrodes has one electrode. 6. The method of claim 1, wherein the first subset of electrodes comprises a plurality of electrodes. 7. The method of claim 1, wherein each electrode of the plurality of electrodes has a length of about 0.5 millimeters to about 2.0 millimeters, the length being measured in a direction substantially parallel to a longitudinal axis of the member. 8. The method of claim 1, wherein each electrode of the plurality of electrodes is spaced from an adjacent electrode by about 0.2 millimeters to about 1.5 millimeters. 9. The method of claim 1, wherein controlling the sensing module to sense the physiological signal of the patient via the second subset of electrodes comprises controlling the sensing module to sense the physiol... | Medtronic, Inc. | A61N001/36 A61N001/02 A61B005/04 A61B005/0428 A61N001/05 A61B005/00
US09801609 | B2 | 2017-10-31 | Device and method for enhancing accuracy of recognizing fetus heart rate acceleration data | 2012-09-25 | The present invention relates to the field of biomedical signal processing, specifically, a device and method of implementation for enhancing the accuracy of fetal heart rate acceleration data recognition. The present invention comprises: collecting fetal heart rate data in a pre-configured period so as to obtain a fetal heart rate data sequence, H(n); performing baseline identification on the fetal heart rate sequence, H(n), to obtain a fetal heart rate baseline data sequence, B(n); pre-processing the fetal heart rate data sequence, H(n), to obtain a pre-processed fetal heart rate data sequence, C(n); performing acceleration recognition on the pre-processed fetal heart rate data sequence, C(n), according to pre-configured acceleration determination criteria and the fetal heart rate baseline data sequence, B(n), to obtain acceleration data segments; calculating an acceleration attribute value for each of the acceleration data segments, and outputting each acceleration data segment and the acceleration attribute value calculation result thereof. The technical solution provided by the present invention effectively distinguishes the variations between the accelerations and the baseline, accurately recognizes each acceleration in continuous accelerations, and avoids the situations of the detected number of fetal heart rate data accelerations being lower than the actual number or being mistaken due to current methods, thereby enhancing the accuracy in recognizing fetal heart rate curve accelerations. | 1. A device for enhancing the accuracy of recognizing fetal heart rate data acceleration, wherein the device comprises a fetal heart rate data collecting module, a fetal heart rate data baseline recognition module, a fetal heart rate data preprocessing module, a fetal heart rate data acceleration recognition module and an output module; wherein the fetal heart rate data collecting module is configured for performing the fetal heart rate data collection in a pre-configured period to obtain a fetal heart rate data sequence, H(n);wherein the fetal heart rate data baseline recognition module is connected with the fetal heart rate data collecting module and configured for receiving the fetal heart rate data sequence, H(n) transmitted by the fetal heart rate data collecting module and recognizing a fetal heart rate baseline sequence, B(n);wherein the fetal heart rate data preprocessing module is connected with the fetal heart rate data collecting module and configured for receiving the fetal heart rate data sequence, H(n), transmitted by the fetal heart rate data collecting module and performing preprocessing on the sequence H(n) to obtain a fetal heart rate data sequence, V(n);wherein the fetal heart rate data acceleration recognition module is connected with the fetal heart rate data baseline recognition module and the fetal heart rate data preprocessing module and configured for receiving the fetal heart rate baseline sequence, B(n), transmitted by the fetal heart rate data baseline recognition module and the preprocessed sequence, V(n), transmitted by the fetal heart rate data preprocessing module; the acceleration recognition on the preprocessed fetal heart rate data sequence V(n) is performed to obtain an acceleration data segment, according to a pre-configured acceleration determination criteria and the fetal heart rate data baseline sequence, B(n);wherein the fetal heart rate data acceleration recognition module comprises a criteria comparison unit configured for determining whether there is a sequence segment meeting a pre-configured acceleration determination criteria in the fetal heart rate data sequence, C(n), an acceleration and variation unit configured for inspecting and distinguishing the sequence segments of the variation between acceleration and baseline in the sequence C(n) meeting the criteria, according to the fetal heart rate data sequence C(n) meeting the determination criteria and the corresponding baseline sequence B(n) thereof, and an acceleration number analysis unit configured for analyzing the continuous acceleration number of an acceleration sequence in the sequence C(n) meeting pre-configured acceleration determination criteria; andwherein the output module is connected with the fetal heart rate data acceleration recognition module and configured for receiving the acceleration data segment transmitted by the fetal heart rate data acceleration recognition module and outputting the acceleration data segment. 2. The device ... | EDAN INSTRUMENTS INC. | A61B008/08 A61B008/02 A61B005/024 A61B005/00
US09788727 | B2 | 2017-10-17 | Physiological information measuring apparatus | 2016-12-22 | A physiological information measuring apparatus includes a first detection unit, having a first light emitting unit and a first light receiving unit, separated from one another by a first distance. A second detection unit of the physiological information measuring apparatus has a second light emitting unit and a second light receiving unit, separated from one another by a second, different distance. Alternatively, the second detection unit shares the first light emitting unit with the first detection unit and has a second light receiving unit. Alternatively, the second detection unit shares the first light receiving unit with the first detection unit and has a second light emitting unit. A measuring unit of the physiological information measuring apparatus measures the physiological information of a user based on light received by the light receiving unit or light receiving units. | 1. A physiological information measuring apparatus comprising: a first detection unit, comprising a first light emitting unit and a first light receiving unit, wherein the first light emitting unit and the first light receiving unit are separated from one another by a first distance;a second detection unit, comprising a second light emitting unit and a second light receiving unit, wherein the second light emitting unit and the second light receiving unit are separated from one another by a second distance, different than the first distance; anda measuring unit that measures the physiological information of a user based on light received by the first light receiving unit and light received by the second light receiving unit,wherein wavelengths of the first light emitting unit and the second light emitting unit are approximately 470 nm to approximately 610 nm, andwherein the first distance is from approximately 2 mm to approximately 4 mm. 2. The apparatus according to claim 1, wherein the second distance is from approximately 4 mm to approximately 10 mm. 3. The apparatus of claim 1, wherein the first light emitting unit, the first light receiving unit, the second light emitting unit, and the second light receiving unit are disposed on an exposed surface of the apparatus and configured to contact skin of the user during use. 4. The apparatus of claim 1, wherein the first light receiving unit and the second light receiving unit are configured to apply pressure to the user. 5. The apparatus according to claim 1, wherein a first region from the first light emitting unit to the first light receiving unit along the surface of the apparatus overlaps a second region from the second light emitting unit to the second light receiving unit along the surface of the apparatus. 6. The apparatus according to claim 1, wherein the measuring unit measures the heart rate by removing noise, wherein the noise is measured based on differences between the light received by the first light receiving unit and the light received by the second light receiving unit. 7. The apparatus according to claim 1, wherein the physiological information comprises at least one member of the group consisting of a heart rate of the user, a blood pressure of the user, a breathing rate of the user, and an autonomic balance of the user. 8. The apparatus according to claim 1, wherein a first region from the first light emitting unit to the first light receiving unit along the surface of the apparatus intersects a second region from the second light emitting unit to the second light receiving unit along the surface of the apparatus. 9. A physiological information measuring apparatus comprising: a first detection unit, comprising a first light emitting unit and a first light receiving unit, wherein the first light emitting unit and the first light receiving unit are separated from one another by a first distance; a second detection unit, comprising one member of the group consisting essent... | Seiko Epson Corporation | A61M005/00 A61B005/11 A61B005/0205 A61B005/00 A61B005/021 A61B005/024 A61B005/08
US09826939 | B2 | 2017-11-28 | Systems and methods for detecting medical treatment | 2015-06-30 | Systems and methods are provided for using stored physiologic information about a subject to detect a previous treatment event. Physiologic information can be sensed from a subject using one or more sensors. Using a detection circuit, a change in the sensed physiologic information, such as a change from reference physiologic information, can be used to identify a candidate previous treatment event. An alert or other information about the candidate treatment event can be provided to a patient or clinician. In an example, a candidate treatment event can include a heart failure or diuresis treatment that is identified using information about a change in one or more of a subject's circadian pattern, a subject's thoracic impedance, or a subject's respiration status. | 1. A system for detecting a previous or ongoing treatment of a subject, the system comprising: a first physiologic sensor configured to sense a physiologic signal from the subject; anda previous treatment detection circuit, the treatment detection circuit including: a memory circuit configured to store information about the physiologic signal from the first physiologic sensor; anda processor circuit configured to identify a candidate previous treatment event for the subject using the information about the physiologic signal stored in the memory circuit. 2. The system of claim 1, wherein the previous treatment detection circuit includes a heart failure (HF) hospitalization treatment detection circuit, and wherein the processor circuit is configured to identify a candidate previous HF hospitalization treatment event for the subject using the information about the physiologic signal stored in the memory circuit. 3. The system of claim 1, wherein the previous treatment detection circuit includes a diuresis treatment detection circuit, and wherein the processor circuit is configured to identify a candidate previous diuresis treatment event for the subject using the information about the physiologic signal stored in the memory circuit. 4. The system of claim 1, wherein the previous treatment detection circuit includes one of an extracellular volume reduction treatment detection circuit, a heart failure (HF) inotrope treatment detection circuit, and a fluid infusion detection circuit. 5. The system of claim 1, wherein the first physiologic sensor is configured to sense a physiologic signal that includes information about the subject's circadian rhythm, and wherein the processor circuit is configured to identify the candidate previous treatment event by identifying a change in the subject's circadian rhythm. 6. The system of claim 5, wherein the first physiologic sensor includes an accelerometer that is configured to sense the physiologic signal including being configured to sense the information about the subject's circadian rhythm using information about the subject's posture and using information about the subject's physical activity level. 7. The system of claim 1, wherein the first physiologic sensor is configured to sense a physiologic signal that includes information about the subject's thoracic impedance, and wherein the processor circuit is configured to identify the candidate previous treatment event using an inter-day trend in the information about the subject's thoracic impedance. 8. The system of claim 1, wherein the first physiologic sensor is configured to sense a physiologic signal that includes information about the subject's respiration status, and wherein the processor circuit is configured to identify the candidate previous treatment event using information about a change in the information about the subject's respiration status. 9. The system of claim 1, wherein the first physiologic sensor is configured to sense a physiologi... | Cardiac Pacemakers, Inc. | A61B005/00 A61B005/11 A61B005/053 A61B007/04 A61B005/091 A61N001/362 A61N001/39 A61B005/0205 A61N001/36 A61N001/365
US09827371 | B2 | 2017-11-28 | Method for controlling the delivery of insulin and the related system | 2012-03-23 | A method (400) controls the delivery of insulin in a diabetic patient (P) based on data (d) representative of at least a fraction of a meal (m(k+i)) that the patient (P) will consume. The method provides from a block (R) representative of conventional therapy or open loop rule that the patient (P) is subject to, based on the data (d) representative of at least a fraction of the meal (m(k+i)), a reference insulin value (u0). The method is also based on data representative of the difference between input data (ŷ), a reference glycemic level, and feedback data (yCGM) representative of the glycemic level detected in the patient (P). A control module (301; 401) provides a value of insulin (i) to be delivered to the patient (P) based on the various representative data. | 1. Method for controlling the delivery of insulin in a diabetic patient comprising steps of: providing data representative of at least a fraction of a meal that the patient will consume;providing, from a block representative of conventional therapy or open loop rule that the patient is subject to, based on the data representative of at least a fraction of the meal that the patient will consume, a reference insulin value;providing data representative of the difference between input data, representative of a reference glycemic level, and feedback data representative of the glycemic level detected in the patient;providing, from a control module, based on the data representative of the difference between the input data, representative of the reference glycemic level, and the feedback data, representative of the glycemic level detected in the patient and the data representative of at least a fraction of the meal that the patient will consume, data representative of a first insulin variation value;providing, based on the data, representative of the first insulin variation value, and data, representative of a basal insulin value, a value of insulin to be delivered to the patient;providing the data representative of the reference glycemic level based on data representative of the reference glycemic level in the absence of meals and data representative of a foreseen variation in the expected glycemic level due to data representative of a further fraction of the meal that the patient will consume;providing, based on the data representative of a further fraction of the meal that the patient will consume after a time interval, a second insulin variation value for compensating the data representative of a further fraction of the meal that the patient will consume;providing, from an empirical model block, based on the data representative of a further fraction of the meal that the patient will consume and the second insulin variation value, the data representative of the foreseen variation in expected glycemic level due to the data representative of a further fraction of the meal that the patient will consume. 2. Control method according to claim 1, also comprising the step of providing the insulin value to be delivered to the patient based on: the data representative of the first insulin variation value, the data representative of the basal insulin value, and the second insulin variation value. 3. Control method according to claim 1, wherein the control module is a predictive controller. 4. Control method according to claim 1, wherein said data representative of said at least a fraction of the meal that the patient will consume is provided by the patient or by an automatic prediction module. 5. Control method according to claim 1, wherein the step of the control module providing the data, representative of a first insulin variation value, is also based on previous data representative of the first insulin variation value provided by the control module. 6. ... | DIPARTIMENTO DI INGEGNERIA CIVILE E ARCHITETTURA DELL‘UNIVERSITA’ DEGLI STUDI DI PAVIA UNIVERSITA' DEGLI STUDI DI PADOVA | A61B005/145 A61M005/172 A61M005/142 G06F019/00
US09826911 | B2 | 2017-11-28 | Wearable device and determination method thereof | 2016-05-27 | A wearable device is provided. The wearable device includes a photon sensor, a processor, and an output unit. The photon sensor senses light reflected from a specific region and transforms the sensed light to a plurality of electric-signal components. The processor receives the electric-signal components sensed within a period to form a dimensional sensing signal. The processor extracts a feature of a waveform of the dimensional sensing signal and determines whether a predetermined heart condition of the object is present according to the feature of the waveform of the dimensional sensing signal to generate a determination signal. The output unit is coupled to the processor. The output unit receives the determination signal and generates an alarm signal according to the determination signal. | 1. A wearable device comprising: a photon sensor sensing light reflected from a specific region of an object and transforming the sensed light to a plurality of electric-signal components;a processor receiving the electric-signal components obtained within a period to form a dimensional sensing signal, detecting a plurality of feature points of the dimensional sensing signal, calculating a distance between two adjacent feature points among the plurality of feature points, extracting a feature of a waveform of the dimensional sensing signal according to the calculated distance, and determining whether a predetermined heart condition of the object is present according to the feature of the waveform of the dimensional sensing signal to generate a determination signal; andan output unit, coupled to the processor, receiving the determination signal and generating an alarm signal according to the determination signal. 2. The wearable device as claimed in claim 1, wherein when the processor determines that the predetermined heart condition is present, the output unit generates the alarm signal according to the determination signal to indicate that heartbeat of the object is irregular. 3. The wearable device as claimed in claim 1, wherein the output unit comprises a display to show information related to the predetermined heart condition according to the alarm signal, a vibration unit to enable a vibration alarm according to the alarm signal, a speaker to play an alarm sound according to the alarm signal, or a wireless transmission unit to communicate with a remote device according to the alarm signal. 4. The wearable device as claimed in claim 1, wherein the processor obtains statistical information of the calculated distance to serve as the feature of the waveform of the dimensional sensing signal. 5. The wearable device as claimed in claim 4, wherein the processor process the statistical information by applying an algorithm to determine whether the predetermined heart condition is present. 6. The wearable device as claimed in claim 1 further comprising: a band, wherein the photo sensor, the processor, and the output unit are disposed on the band, and the wearable device is tied onto the specific region through the band. 7. The wearable device as claimed in claim 1 further comprising: a light-emitting source providing light beams to the specific region, wherein the photon sensor senses the light reflected from the specific region being illuminated by the light beams. 8. The wearable device as claimed in claim 1, wherein intensity of the light beams provided by the light-emitting source is adjustable. 9. The wearable device as claimed in claim 1 further comprising: an electrocardiography recorder, coupled to the processor,wherein when the processor determines that the predetermined heart condition is present, the processor generates the determination signal to trigger the ECG recorder to perform an ECG measurement operation. 10. The wearable de... | MEDIATEK INC. | A61B005/11 A61B005/00 A61B005/024 A61B005/0402
US09801527 | B2 | 2017-10-31 | Lumen-traveling biological interface device | 2007-03-19 | Lumen-traveling biological interface devices and associated methods and systems are described. Lumen-traveling biological interface devices capable of traveling within a body lumen may include a propelling mechanism to produce movement of the lumen-traveling device within the lumen, electrodes or other electromagnetic transducers for detecting biological signals and electrodes, coils or other electromagnetic transducers for delivering electromagnetic stimuli to stimulus responsive tissues. Lumen-traveling biological interface devices may also include additional components such as sensors, an active portion, and/or control circuitry. | 1. A lumen-traveling device, comprising: a propelling mechanism capable of producing directional movement of the lumen-traveling device through a body lumen by producing relative extension and retraction of at least two lumen-wall-engaging structures with respect to each other in combination with alternate engagement and disengagement of at least one wall of the body lumen;a steering mechanism capable of modifying a direction of movement of the lumen-traveling device;a structural element configured to at least intermittently permit the movement of fluid through the lumen-traveling device;at least one electromagnetic transducer configured for at least one of producing an output signal representative of a bioelectromagnetic signal sensed from a target tissue, the bioelectromagnetic signal generated by a biological signal source in the target tissue, and delivering a non-ablative electromagnetic stimulus to the target tissue; andat least one of a signal processing portion capable of processing the output signal from the electromagnetic transducer and a stimulus source capable of producing the electromagnetic stimulus for delivery to the target tissue with the electromagnetic transducer. 2. The lumen-traveling device of claim 1, including at least one wall-engaging structure capable of engaging a wall of the body lumen to secure the lumen-traveling stimulation device with respect to the wall of the body lumen in the vicinity of the target tissue. 3. The lumen-traveling device of claim 1, including at least one electromagnetic transducer configured for producing an output signal representative of a bioelectromagnetic signal sensed from a target tissue, and at least one signal processing portion capable of processing the output signal from the electromagnetic transducer. 4. The lumen-traveling device of claim 1, including at least one electromagnetic transducer configured for delivering an electromagnetic stimulus to the target tissue, and at least one stimulus source capable of producing an electromagnetic stimulus for delivery to the target tissue with the electromagnetic transducer. 5. The lumen-traveling device of claim 1, including at least one electromagnetic transducer configured for producing an output signal representative of a bioelectromagnetic signal sensed from a target tissue, at least one electromagnetic transducer configured for delivering an electromagnetic stimulus to the target tissue; at least one signal processing portion capable of processing the output signal from the electromagnetic transducer, and at least one stimulus source capable of producing an electromagnetic stimulus for delivery to the target tissue with the electromagnetic transducer. 6. The lumen-traveling device of claim 5, including at least one electromagnetic transducer configured for producing an output signal representative of a bioelectromagnetic signal sensed from a target tissue and delivering an electromagnetic stimulus to the target tissue. 7. The lum... | Gearbox, LLC | A61B001/04 A61B001/00 A61B010/04 A61B017/29 A61B017/295 A61F002/04 A61B034/32 A61B034/20 A61B034/00 A61B005/00 A61B005/0205 A61B005/0215 A61B005/03 A61B005/04 A61B005/145 A61B007/00 A61B008/12 A61B008/00 A61B017/22 A61B017/3207 A61B017/00 A61F002/30 A61N001/05 A61N001/372 A61B090/00 A61B034/30
US09808164 | B2 | 2017-11-07 | Evaluation of cardiac dyssynchrony based on chest wall motion and electrical cardiac activity | 2012-08-28 | Cardiac dyssynchrony of a patient may be evaluated based on electrical activity of a heart of the patient and corresponding chest wall motion of the patient sensed via an external accelerometer. In one example, an acceleration signal indicative of the chest wall motion is generated by an external accelerometer positioned on the chest wall of the patient. A processor of a diagnostic device integrates the acceleration signal to generate a velocity signal and temporally correlates the velocity signal and an electrical cardiac signal. The processor determines a time delay between a deflection of the electrical cardiac signal indicating ventricular electrical activation and a subsequent greatest peak of the velocity signal. The time delay may indicate a degree of electromechanical delay of the left ventricle. In some examples, the processor generates an output indicative of a cardiac dyssynchrony status based on the time delay. | 1. A system comprising: an external accelerometer configured to generate an acceleration signal indicative of chest wall motion of a patient; anda processor configured to integrate the acceleration signal to generate a velocity signal, receive an electrical cardiac signal indicative of electrical cardiac activity of the patient, temporally correlate the velocity signal with the electrical cardiac signal, determine a time delay between a deflection of the electrical cardiac signal indicative of electrical activation of a ventricle of a heart of the patient and a subsequent greatest peak of the velocity signal, and generate an output indicative of a cardiac dyssynchrony status of the patient based on the time delay. 2. The system of claim 1, wherein the electrical cardiac signal comprises at least one of an electrocardiogram or an electrogram, and the deflection is a Q-wave. 3. The system of claim 1, wherein the electrical cardiac signal comprises at least one of an electrocardiogram or an electrogram, and the deflection is an R-wave. 4. The system of claim 1, wherein the accelerometer is configured to generate an acceleration signal indicative of movement of the chest wall of the patient in a single axis. 5. The system of claim 1, further comprising a user interface, wherein the processor is configured to generate the output indicative of the cardiac dyssynchrony status of the patient by at least controlling the user interface to display a visible indication of the time delay. 6. The system of claim 1, further comprising a user interface, wherein the processor is configured to generate the output indicative of the cardiac dyssynchrony status of the patient by at least controlling the user interface to generate an audible or somatosensory indication of the time delay. 7. The system of claim 1, further comprising a memory, wherein the processor stores the output indicative of the cardiac dyssynchrony status in the memory. 8. The system of claim 1, wherein the processor is configured to determine the cardiac dyssynchrony status by at least comparing the time delay to a threshold value, and wherein the processor is configured to determine a first cardiac dyssynchrony status and generate a first output indicative of the first cardiac dyssynchrony status in response to determining the time delay is less than the threshold value, and determine a second cardiac dyssynchrony status and generate a second output indicative of the second cardiac dyssynchrony status in response to determining the time delay is not less than the threshold value. 9. The system of claim 1, wherein the processor is configured to generate the output by at least automatically generating an output recommending a type of cardiac rhythm management therapy based on the time delay. 10. The system of claim 9, wherein the processor is configured to determine the type of cardiac rhythm management therapy by at least comparing the time delay to a plurality of time delay ranges assoc... | Medtronic, Inc. | A61B005/0245 A61B005/0452 A61B005/11
US09820699 | B2 | 2017-11-21 | Processing status information of a medical device | 2014-06-12 | A system (100) is provided for processing status information of a medical device (020-023). The medical device performs a medical function at a bedside of a patient and is arranged for generating an alarm signal (030) to alert a caregiver of the patient to an occurrence of an event which is associated with the performing of the medical function. A status interface (120) acquires a device signal (024) of the medical device, the device signal comprising status information which is indicative of a current status of the performing of the medical function. Moreover, an analysis subsystem (140) analyzes the status information to estimate an imminent occurrence of the event based on said current status, and a notification subsystem notifies the caregiver of the imminent occurrence away from the bedside of the patient of the event by a generating a notification signal (162) for a notification device (060-064). The priority of notifying the caregiver is based on an estimate of whether the patient is asleep. The system enables the caregiver to be notified about the event ahead of time, i.e., before the event occurs and thus before the medical device generates an alarm signal. Advantageously, the patient is less disturbed by alarm signals during sleep. | 1. A system for processing status information of a medical device performing a medical function at a bedside of a patient, the medical device being arranged for generating an alarm signal which is perceivable at the bedside of the patient to enable alerting a caregiver of the patient to an occurrence of an event which is associated with the performing of the medical function, the system comprising: a status interface for acquiring a device signal of the medical device, the device signal comprising status information which is indicative of a current status of the performing of the medical function;an analysis subsystem for analyzing the status information to estimate an imminent occurrence of the event based on said current status; anda notification subsystem for notifying the caregiver of the imminent occurrence of the event away from the bedside of the patient by a generating a notification signal for a notification device, the notification subsystem being arranged for i) determining a priority of the notifying of the caregiver based on estimating whether the patient is asleep, and ii) effecting the notifying of the caregiver based on the priority. 2. The system according to claim 1, wherein the notification subsystem is arranged for estimating whether the patient is asleep based on at least one of the group of: physiological data as obtained from a patient monitor, time data indicating a time of day, and light data being indicative of an amount of light in a room of the patient. 3. The system according to claim 1, wherein the notification subsystem is arranged for delaying or discarding said notifying if the priority does not exceed a threshold. 4. The system according to claim 1, wherein the analysis subsystem is arranged for i) tracking the status information over time to obtain a time series of status information, and ii) estimating the imminent occurrence of the event based on an extrapolation of the time series of status information. 5. The system according to claim 1, wherein the analysis subsystem is arranged for estimating a time of the imminent occurrence of the event, and wherein the notification subsystem is arranged for generating the notification signal to notify the caregiver of said time. 6. The system according to claim 1, wherein the status interface is further arranged for acquiring physiological data from a patient monitor of the patient, and wherein the analysis subsystem is arranged for estimating the imminent occurrence of the event further based on the physiological data. 7. The system according to claim 1, wherein the notification subsystem is arranged for generating the notification signal for a mobile notification device of the caregiver. 8. The system according to claim 1, wherein the notification subsystem is arranged for generating the notification signal for effecting a non-auditory notification of the caregiver. 9. The system according to claim 8, wherein the notification device comprises a light source, a... | Koninklijke Philips N.V. | G08B023/00 A61B005/00 G06F019/00 G08B021/04 A61B005/021 A61B005/024
US09782121 | B2 | 2017-10-10 | Distinguishing closed and open respiratory airway apneas by complex admittance values | 2017-01-31 | Methods and apparatus are disclosed for determining the occurrence of a closed or open apnea. Respiratory air flow from a patient is measured to give an air flow signal. The determination of an apnea is performed by applying an oscillatory pressure waveform of known frequency to a patient's airway, calculating a complex quantity representing a patient admittance (12) and comparing its value with ranges (14,16) indicative of open or closed apneas. The method distinguishes open from closed apneas even when the model used to calculate admittance is not based on details of the respiratory apparatus. In addition the patient admittance may be compared with admittance during normal breathing to avoid having to characterize the airway. | 1. An apparatus for providing a supply of continuous positive airway pressure to an airway of a patient at a desired treatment pressure, the apparatus comprising: a flow generator adapted to provide a supply of continuous positive airway pressure, the flow generator adapted to couple with an air delivery tube coupled with a patient interface configured to sealingly engaged with a face of the patient;a means for determining respiratory air flow and pressure in the patient interface;a means for measuring a patient impedance of the airway during a period of a breathing cycle when an apnea is not occurring; anda controller to control operation of the flow generator based on the determined respiratory air flow and pressure,the controller configured in use to detect the occurrence of an apnea and to determine patency of the airway of the patient during the apnea by performing: controlling application of an oscillatory pressure waveform of known frequency to the patient interface for delivery to the airway of the patient in use;measuring an apnea impedance of the airway during the apnea; andcomparing the measured apnea impedance with the patient impedance to determine whether the airway is patent. 2. The apparatus according to claim 1, wherein the means for determining respiratory air flow and pressure includes measuring the air flow and pressure using at least one sensor. 3. The apparatus according to claim 2, wherein the at least one sensor is a differential pressure sensor. 4. The apparatus according to claim 2, wherein the patient impedance is determined by the controller controlling application of an oscillatory pressure waveform of known frequency to the patient interface for delivery to the patient's airway during the period of the breathing cycle when an apnea is not occurring. 5. The apparatus according to claim 1, wherein the period of the breathing cycle is at an end of an expiratory portion of a breath. 6. The apparatus according to claim 1, wherein the patient impedance is a determined average of measured patient impedance over a predetermined time period. 7. The apparatus according to claim 1, wherein the controller is configured to perform a plausibility check on the measured patient impedance. 8. The apparatus according to claim 7, wherein if the measured patient impedance value is determined to be implausible, the controller indicates that a fault has occurred. 9. The apparatus according to claim 1, wherein the known frequency is between 1 and 16 Hz. 10. The apparatus according to claim 9, wherein the known frequency is between 2 and 8 Hz. 11. The apparatus according to claim 9, wherein the known frequency is 4 Hz. 12. The apparatus according to claim 1, wherein the controller is further configured to increase the treatment pressure in response to the determination of a patent airway. 13. The apparatus according to claim 1, wherein the controller is further configured to reduce the treatment pressure in response to the deter... | ResMed Limited | A61B005/08 A61B005/00 A61M016/00 A61B005/085 A61B005/087
US09814878 | B1 | 2017-11-14 | Self-sustaining piezoelectric assembly | 2016-11-18 | A self-sustaining piezoelectric assembly is disclosed herein for providing electrical stimulation to nerves near the surface of the skin. The assembly comprises a non-conductive mat having an aperture and a piezoelectric device positioned to fill at least a portion of the aperture. A flexible conductive material is position to be in electrical contact with a first and a second primary surface of the piezoelectric device, and with the surface of the skin, providing an electric current to the skin when the piezoelectric is mechanically compressed. | 1. A self-sustaining assembly for providing electrical stimulation to nerves near the surface of the skin of an individual, comprising: at least one non-conductive mat, said non-conductive mat comprising a heel position and a ball position and at least one aperture located at said heel position or said ball position;at least one piezoelectric device positioned to fill at least a portion of said aperture, said piezoelectric device having a first primary surface and a second primary surface generally opposed to each other and having opposing electrical charges when mechanical force is applied to said piezoelectric device; anda flexible conductive material comprising a first conductive fabric section and a second conductive fabric section, wherein said first conductive fabric section is in electrical contact with said first primary surface, said second conductive fabric section is in electrical contact with said second primary surface, and said first and second conductive fabric sections are separated by a gap that is positioned to enable the skin of an individual to fill at least a portion of the gap and create a closed electrical circuit upon use, said gap extending between an edge of said first conductive fabric section and an edge of said second conductive fabric section. 2. The assembly of claim 1, wherein said non-conductive mat comprises, a first mat and a second mat, wherein said second mat is adhered to said first mat on the side of said second primary surface of the piezoelectric device. 3. The assembly of claim 1, wherein said piezoelectric device is recessed within the aperture and said electrical contact between said first conductive fabric section and said first primary surface further comprises an electrical lead extending between said first primary surface of the piezoelectric device and said first flexible conductive fabric section. 4. A method of preparing a self-sustaining assembly for providing electrical stimulation to nerves near the surface of the skin of an individual, said method comprising: providing at least one non-conductive mat having an aperture as a holder for at least one piezoelectric device, said device having a first primary surface and a second primary surface;positioning said piezoelectric device so as to fill at least a portion of said aperture;providing a first flexible conductive fabric in electrical contact with the first primary surface of said piezoelectric device, and a second flexible conductive fabric in electrical contact with the second primary surface of said piezoelectric device, andproviding a gap between said first and second flexible conductive fabrics, said gap positioned to enable the skin of an individual to fill at least a portion of the gap and create a closed electrical circuit upon use, said gap extending between an edge of said first conductive fabric section and an edge of said second second conductive fabric section. 5. The method of claim 4, further comprising recessing the piezoel... | - | A61N001/36 A61B005/103 A61N001/04 H02N002/18
US09788773 | B2 | 2017-10-17 | Vein presentation enhancement device | 2013-03-15 | Preferably, an embodiment of a vein presentation enhancement device includes at least, an interior cover supporting a first fastening member, and an exterior cover communicating with the interior cover, wherein the exterior cover provides a first securement member that interacts with the first fastening member to secure the preferred device positioned about a limb of a subject. The preferred embodiment further includes a bladder formed between the interior and exterior covers, and an air transfer assembly connected to said bladder for transfer of air into and out of said bladder, wherein said interior and exterior covers collectively provide a plurality of projections defining a blood access window. | 1. A vein presentation enhancement apparatus comprising: an interior cover and an exterior cover, wherein the interior cover and the exterior cover are secured to one another to form a cuff, the interior cover including a first fastening member at a first end of the cuff, the exterior cover including a second fastening member at a second end of the cuff opposite to the first end of the cuff, and the first fastening member and the second fastening member are configured to removably attach to one another to wrap the cuff around a limb of a patient, wherein the first fastening member and the second fastening member are a respective hook portion and a respective loop portion of a hook and loop fastening system;an air transfer assembly;a blood access window configured to allow access to draw blood;a plurality of bladder walls;a plurality of bladders formed between the bladder walls, the plurality of bladders comprising a first bladder and a second bladder, wherein at least a portion of the first bladder is further from the blood access window than the second bladder;a first air passage connecting the first bladder to the second bladder, wherein the first air passage is configured to apply air pressure to the first bladder before the second bladder;a third bladder, the third bladder comprising at least one of the bladder walls; anda second air passage connecting the second bladder to the third bladder, wherein the third bladder is adjacent to at least three sides of the blood access window and configured to apply air pressure to the second bladder before the third bladder. 2. The vein presentation enhancement apparatus of claim 1 further comprising a filling nipple configured to receive air from the air transfer assembly. 3. The vein presentation enhancement apparatus of claim 1 further comprising a perimeter seal for retaining the air within the vein presentation enhancement apparatus. 4. The vein presentation enhancement apparatus of claim 1 further comprising a seal bag, wherein the seal bag is configured to rip away from a patient using a pull string. | - | A61M005/315 A61B005/15 A61B017/135
US09823189 | B2 | 2017-11-21 | Optical method for determining morphological parameters and physiological properties of tissue | 2009-03-18 | In an embodiment of the present disclosure, an optical method for determining morphological parameters and physiological properties of tissue is presented. The method includes using reflectance measurements from a tissue area for a plurality of wavelengths, using a bio-optical model, using radiative transfer modeling and using a non-linear inversion procedure. The method further includes systematically varying values of the morphological parameters and physiological properties of the tissue and simultaneously varying the inherent optical properties, which are linked to the morphological parameters and the physiological properties of the tissue, until the non-linear inversion procedure returns values for the morphological parameters and the physiological properties of the tissue such that an agreement between the reflectance measurements and reflectances computed by the radiative transfer model, based on the returned morphological parameters and the physiological properties of the tissue values and corresponding inherent optical properties values, reach a predetermined level of accuracy. | 1. An optical method for determining morphological parameters and physiological properties of a multilayered, vertically inhomogeneous skin tissue having inherent optical properties that are fixed in each layer, but vary between layers, the method comprising the steps of: generating a plurality of reflectance images using an optical transfer diagnosis (OTD) device, said plurality of reflectance images comprising reflectance measurements from above a skin surface for a plurality of different wavelengths from multiple angles of illumination and detection and using the reflectance measurements to determine the morphological parameters and the physiological properties for each layer in a plurality of layers of the multilayered, vertically inhomogeneous skin tissue,wherein said OTD device comprises a measurement head comprising a plurality of light-emitting diode (LED) lamps and a plurality of cameras where each of the plurality of LED lamps are placed at a different angle, relative to the skin surfacewherein determining the morphological parameters comprises using one of the reflectance images of the lesion from a visible channel for: determining a lesion border,measuring a size of the lesion,measuring a histogram-width that gives a measure of inhomogeneity of the reflectance image of the lesion,capturing a relative moment of inertia of the lesion,determining a center distance representing a physical distance between a geometric center of the lesion and a center of mass of absorptance,determining a fractal dimension of the lesion which describes the complexity of its border,determining an asphericity of the lesion, anddetermining a relative border length of the lesion, wherein the relative border length is defined as a ratio of a border length of a circle with a same area as the lesion to the border length of the lesion,and wherein determining the morphological parameters further comprises, and also wherein determining the physiological properties comprises, using a bio-optical model, radiative transfer modeling, and a non-linear inversion procedure;systematically varying input values of the morphological parameters and the physiological properties of the skin tissue while simultaneously varying inherent optical properties, wherein the inherent optical properties are linked to the morphological parameters and the physiological properties of the skin tissue by the bio-optical model,computing reflectances by the radiative transfer model for a plurality of different wavelengths and a plurality of different measurement configurations, each with specified directions of illumination and observation until an agreement between the reflectance measurements and reflectances computed by the radiative transfer model reaches a predetermined level of accuracy, andobtaining, for each layer in the plurality of layers of the multilayered, vertically inhomogeneous skin tissue, the morphological parameters and the physiological properties returned by the non-linear inv... | Balter, AS. | A61B006/00 G01N021/47 A61B005/00 A61B005/1455 G06K009/00 G06T007/62 G06T007/66
US09795318 | B2 | 2017-10-24 | Magnetic resonance imaging apparatus and magnetic resonance imaging method | 2013-08-15 | A magnetic resonance imaging apparatus includes a setting unit which sets a section position, a first image creating unit which creates a first image at the set section position by a multi planar reformat on the basis of a magnetic resonance signal collected from a subject by a first imaging sequence at a 3D region or multiple section positions different from the set section position, a determining unit which determines a section position of the first image, and a second image creating unit which creates a second image at the determined section position by the multi planar reformat on the basis of a magnetic resonance signal collected from the subject by a second imaging sequence different from the first imaging sequence at a 3D region or multiple section positions different from the determined section position. | 1. A magnetic resonance imaging apparatus comprising: a static magnetic field magnet configured to generate a uniform static magnetic field;a gradient magnetic field coil configured to generate a gradient magnetic field;a RF coil configured to receive magnetic resonance signals;a computer system, operatively coupled to at least the gradient magnetic field coil and the RF coil; the computer system including memory for storing instructions and a computer processor for executing the instructions so that the computer system is at least configured to: collect the magnetic resonance signals from a subject using a 2D imaging technique in a first protocol and collect magnetic resonance signals from the subject using a 3D imaging technique in a second protocol;generate a slice image at a first slice position on the basis of the magnetic resonance signals collected in the first protocol;determine a second slice position for obtaining a multi planar reformat image to be a same slice position as the first slice position; andgenerate the multi planar reformat image at the determined second slice position by a multi planar reformat on the basis of the magnetic resonance signals collected in the second protocol;set a region where the multi planar reformat image overlaps with the slice image as a cutout section so that the slice image outside the multi planar reformat image is not identical with the cutout section set based on the slice image; andinform an operator of the magnetic resonance imaging apparatus that the cutout section is not identical with the slice image at a time of displaying the multi planar reformat image created from the cutout section by at least one of a non-identical region displayed in the MPR image, a mark, a text message, an alarming sound, and/or a voice message. 2. A magnetic resonance imaging method comprising the steps of: generating a uniform static magnetic field using a static magnetic field magnet;generating a gradient magnetic field using a gradient magnetic field coil;receiving magnetic resonance signals using a RF coil;operating a computer system, operatively coupled to at least the gradient magnetic field coil and the RF coil; the computer system including memory for storing instructions and a computer processor for executing the instructions so that the computer system performs at least the following: collecting the magnetic resonance signals from a subject using a 2D imaging technique in a first protocol;collecting magnetic resonance signals from the subject using a 3D imaging technique in a second protocol;generating a slice image at a first slice position on the basis of the magnetic resonance signals collected in the first protocol;determining a second slice position for obtaining a multi planar reformat image to be a same slice position as the first slice position;generating, using a magnetic resonance imaging apparatus, the multi planar reformat image at the determined second slice position by a multi planar reformat... | TOSHIBA MEDICAL SYSTEMS CORPORATION | A61B005/055 A61B005/00
US09801589 | B2 | 2017-10-31 | Device and method for determining a comparison value of biodata and for recording biodata | 2007-04-04 | The method and the device are used for evaluating recorded measurement data. In particular, the use of the method and of the device makes it possible to detect the individual risk of a living being in respect of certain disease states and, if appropriate, to automatically evaluate correspondingly recognized risks in a device control. Measurement data in respect of heart-specific and/or circulation-specific parameters or respiration-specific parameters of a patient are preferably recorded and evaluated. The measurement data are determined over a defined period of time and the measurement data are evaluated taking into account the time information. A particularly effective implementation of the method can be based on pattern recognition in which individual patterns and/or sequence patterns are evaluated. | 1. A device for determining a reference value of biophysical data of an individual for determining an individual risk, said device comprising: at least one sensor for noninvasive measurement of at least three signals, which are CWF (continuous wave fluctuation), SpO2, heart rate; and an evaluation unit connected to the sensor, where the evaluation unit has at least one analyzer adapted to determine signal ranges that are definable by signal analysis, wherein the evaluation unit includes a comparator adapted to coordinate the signal ranges by taking additional parameters into account, and to produce a result that is output as an index value, wherein the sensor is a pulse oximeter adapted to record a plethysmogram and SpO2, wherein the evaluation unit extracts CWP's (continuous wave parameters) from the plethysmogram, the evaluation unit extracting the following as CWP's: pulse wave amplitude, a ratio of different integrals over different intervals of the plethysmogram, a quantity related to PTT (pulse transit time), the evaluation unit being adapted to compute a CWF (continuous wave fluctuation) signal from the CWP's as a function of time, wherein the CWF signal contains information about fluctuations of the plethysmogram or the CWP's derived from the plethysmogram, the signals being analyzed by pattern recognition and compared with stored values, the comparison providing a result that yields a patient-specific risk index that is suitable for predicting a risk of cardiovascular disease. 2. A device in accordance with claim 1, wherein the additional parameters include existing biophysical data of the individual and/or static data sets. 3. A device in accordance with claim 2, wherein the static data sets are compiled from groups of individuals, and include medication, age, sex, and diagnostic data. 4. A device in accordance with claim 1, wherein the comparator is adapted to be connected to peripheral devices. 5. A device in accordance with claim 1, wherein the evaluation unit is adapted for determining the CWF both from the at least one detected test signal and the CWP. 6. A method for determining an individual risk for an individual, comprising the following steps: determining measurement data of at least three heart-specific and/or circulation-specific and/or respiration-specific parameters of the user, selected from the group comprising: continuous wave fluctuation (CWF), SpO2, SaO2, pulse rate, heart rate, and blood pressure by at least one noninvasive sensor device;determining the measurement data over a period of at least one minute;determining measurement data with a synchronized time recording;bringing together time information and the measurement data;analyzing the measurement data with the time information being taken into account;carrying out the analysis as a pattern recognition, including, at least in intervals, detecting a course of the measurement with respect to time of at least one parameter can be and evaluating the course wit... | Loewenstein Medical Technology S.A. | A61B005/00 A61B005/024
US09801575 | B2 | 2017-10-31 | Advanced analyte sensor calibration and error detection | 2012-04-13 | Systems and methods for processing sensor data and self-calibration are provided. In some embodiments, systems and methods are provided which are capable of calibrating a continuous analyte sensor based on an initial sensitivity, and then continuously performing self-calibration without using, or with reduced use of, reference measurements. In certain embodiments, a sensitivity of the analyte sensor is determined by applying an estimative algorithm that is a function of certain parameters. Also described herein are systems and methods for determining a property of an analyte sensor using a stimulus signal. The sensor property can be used to compensate sensor data for sensitivity drift, or determine another property associated with the sensor, such as temperature, sensor membrane damage, moisture ingress in sensor electronics, and scaling factors. | 1. A system for measuring an analyte, the system comprising sensor electronics configured to be operably connected to a continuous analyte sensor, the sensor electronics configured to: apply a time-varying signal to the analyte sensor;measure a peak current from the sensor in response to the applied signal;determine a sensitivity of the analyte sensor, the determining comprising correlating the peak current to a predetermined sensor sensitivity curve that is a representation of a change in sensitivity over time, the determining further comprising selecting the predetermined sensitivity curve based on the peak current from a plurality of different predetermined sensitivity curves; andgenerate estimated analyte concentration values using the determined sensitivity and sensor data generated by the analyte sensor. 2. The system of claim 1, wherein the sensitivity curve comprises varying sensitivity values over time since implantation of the sensor. 3. The system of claim 1, wherein the predetermined sensitivity curve comprises a plurality of sensitivity values. 4. The system of claim 1, wherein the predetermined sensitivity curve is based on sensor sensitivity data generated from studying sensitivity changes of analyte sensors similar to the analyte sensor. 5. The system of claim 1, wherein the sensor electronics are further configured to apply a bias voltage to the sensor, wherein the time-varying signal comprises a step voltage above the bias voltage or a sine wave overlaying a voltage bias voltage. 6. The system of claim 1, wherein the determining further comprises calculating an impedance value based on the peak current and correlating the impedance value to a sensitivity value of the predetermined sensitivity curve. 7. The system of claim 1, wherein the sensor electronics are further configured to apply a DC bias voltage to the sensor to generate sensor data, wherein the generating the estimated analyte concentration values includes correcting the generated sensor data based on the measured response. 8. The system of claim 7, wherein the sensor electronics are further configured to apply a conversion function to the corrected sensor data to generate the estimated analyte concentration values. 9. The system of claim 1, wherein the sensor electronics are further configured to form a conversion function based at least in part of the determined sensitivity, and wherein the conversion function is applied to the sensor data to generate the estimated analyte concentration values. 10. The system of claim 1, wherein the selecting includes performing a data association analysis to determine a correlation between the peak current and each of the plurality of different predetermined sensitivity curves and wherein the selected predetermined sensitivity curve has the highest correlation. 11. The system of claim 1, wherein the sensor electronics are further configured to determine the sensitivity using the selected sensitivity curve. 12. The system ... | DexCom, Inc. | A61B005/1495 G01N037/00 G01N027/327 A61B005/145 A61B005/1486 A61B005/00 G01D018/00 G01N027/02 G01N033/49 A61B005/1473
US09795737 | B2 | 2017-10-24 | Method and system for closed-loop control of an artificial pancreas | 2013-03-15 | Methods and systems for controlling an insulin pump in response to glucose measurements are responsive to a base insulin delivery profile and a temporary insulin delivery profile. These can be used, e.g., to control blood glucose level of a subject using a continuous glucose monitor and an insulin infusion pump. During a selected time range, an insulin amount for the pump to supply is determined using the temporary insulin delivery profile. Outside that time range, the insulin amount is determined using the base insulin delivery profile. The temporary insulin delivery profile can specify an exact amount to be supplied (a “hard” profile), a nominal amount to be supplied if doing so does not drive glucose out of a desired zone (“soft”), or a soft profile with a minimum amount of insulin to be delivered (“semi-soft”). | 1. Apparatus for the delivery of insulin, the apparatus comprising: a) a glucose monitor adapted to measure respective glucose levels of a subject at discrete time intervals and provide respective glucose measurement data indicating each measured glucose level;b) an insulin infusion pump configured to deliver insulin in response to a delivery control signal;c) a memory configured to store a base insulin delivery profile;d) an interface adapted to selectively receive a temporary insulin delivery profile extending over a selected time range of the time intervals and to provide a first signal indicating whether the temporary insulin delivery profile was received; ande) a controller adapted to, for each of a plurality of the discrete time intervals: i) receive the glucose measurement data for that time interval from the glucose monitor;ii) determine an insulin delivery amount for that time interval using the received temporary insulin delivery profile when the time interval is in the selected time range and in response to the first signal and, when the time interval is outside the selected time range, using model predictive control based on a selected target glucose concentration range, the received glucose measurement data, and the stored base insulin delivery profile; andiii) provide to the insulin infusion pump a delivery control signal corresponding to the determined insulin delivery amount, whereby a corresponding amount of insulin is delivered. 2. The apparatus according to claim 1, wherein the interface is adapted to receive the temporary insulin delivery profile by receiving change information and modifying the stored base insulin delivery profile in the selected time range according to the change information to provide the temporary insulin delivery profile. 3. The apparatus according to claim 1, wherein the controller is adapted to determine the insulin delivery amount for a selected time interval using: a) if the temporary insulin delivery profile was not received or the selected time interval is outside the selected time range, the controller is adapted to use the selected target glucose concentration range, the received glucose measurement data, the stored base insulin delivery profile;b) otherwise, the controller is adapted to use the selected target glucose concentration range, the received glucose measurement data, and the received temporary insulin delivery profile. 4. The apparatus according to claim 1, wherein the controller is further adapted to, in response to the first signal and if a selected time interval is in the selected time range, constrain the determined insulin delivery amount for the selected time interval to be at least a difference between respective values of the temporary insulin delivery profile and the stored base insulin delivery profile for the selected time interval. 5. The apparatus according to claim 1, wherein the controller is further adapted to: a) predict an excursion of a glucose level of the subjec... | ANIMAS CORPORATION | A61M005/142 A61M005/172 A61B005/145 A61B005/00 G06F019/00
US09775561 | B2 | 2017-10-03 | System method and device for monitoring physiological parameters of a person | 2011-12-20 | Provided is a system, method and portable device for monitoring physiological parameters of a person in the field. | 1. An apparatus, comprising: an electronics portion;a housing enclosing at least a portion of said electronics portion and comprising a plurality of housing sensor pads, each of the plurality of housing sensor pads comprising a leaf spring;a garment, at least a portion of said garment having a plurality of stacked layers, one of said plurality of stacked layers being an electrically conductive sensor layer, said sensor layerconfigured to measure at least a first physiological signal of a user wearing said garment;an interface comprising a receptacle portion and a plurality of interface sensor pads, the interface configured to be fixedly attached to said garment in a chest area, each of the plurality of interface sensor pads comprising a protrusion;said housing configured to be removably attached to said interface by snapping into the receptacle portion of the interface, wherein the plurality of protrusions are configured to orient the housing in an attached state with respect to the interface to ensure that each leaf spring of the plurality of housing sensor pads aligns with and contacts the protrusion of the plurality interface sensor pads when the housing is snapped into the interface,the plurality of interface sensor pads including a first interface sensor pad coupled to a sensor ground, a second interface sensor pad coupled to an ECG sensor, and a third interface sensor pad coupled to a breathing rate sensor;at least one of said interface sensor pads being electrically coupled to said sensor layer. 2. The apparatus of claim 1, wherein said housing includes a user actuated member having a first state in which said user actuated member engages said interface and a second state in which said user actuated member does not engage said interface and wherein said user actuated member is configured to transition from said first state to said second state by actuation of the user. 3. The apparatus of claim 1, wherein said interface is fixedly attached to said garment by any of heat bonding or sewing, or a combination thereof. 4. The apparatus of claim 1, wherein said electronics portion comprises: a battery,at least one wireless transceiver,a processor,at least one light emitting LED, andan internal memory. 5. The apparatus of claim 4, wherein said at least one wireless transceiver comprises: at least two wireless transceivers, each operable to communicate simultaneously with at least one remote device. 6. The apparatus of claim 1, wherein said apparatus is configured to measure at least one of heart rate, temperature, activity, ECG data, breathing rate, or posture. | Covidien LP | A61B005/0408 A61B005/00 A61B005/024 A61B005/08 A61B005/11
US09816947 | B2 | 2017-11-14 | Specimen chamber for optical imaging of radiopharmaceuticals | 2013-08-05 | Apparatus for optical imaging of Cerenkov luminescence from an object subsequent to the object receiving a dose of a radiopharmaceutical, the apparatus comprising: a light tight enclosure within which the object can be received at a sample location; an imaging means; a means to mitigate direct particle impingement between the sample location and the imaging means; and one or more optical elements for transmitting Cerenkov photons from within the light tight enclosure to the imaging means. | 1. Apparatus for optical imaging of Cerenkov luminescence from a sample subsequent to the sample receiving a dose of a radiopharmaceutical, the apparatus comprising: a light tight enclosure within which the sample can be received at a sample location;an imaging means outside of the enclosure, the imaging means for optical imaging of Cerenkov photons;a radiation shield disposed between the sample location in the enclosure and the imaging means, wherein the radiation shield is configured to protect the imaging means from radiation emitted from the sample; andone or more optical elements for guiding the light path of the Cerenkov photons emitted from the sample along a hollow optical conduit, the one or more optical components being arranged in the hollow optical conduit to turn the Cerenkov photons through an angle around the radiation shield, to the imaging means outside the enclosure. 2. Apparatus according to claim 1, wherein the imaging means is positioned and/or oriented relative to the sample location to minimise radiation impingement on the imaging means. 3. Apparatus according to claim 1, wherein the enclosure is a light tight specimen chamber. 4. Apparatus according to claim 3, where in the chamber comprises a door in a wall of the chamber that, when open, allows access to the inside of the chamber and when closed ensures a light tight seal with surrounding parts of the chamber to maintain the light tightness of the chamber interior for imaging the sample. 5. Apparatus according to claim 4, comprising a seal around the perimeter of the door to provide the light tight seal between the door and surrounding parts of the chamber wall when the door is closed. 6. Apparatus according to claim 5, wherein the position of the sample platform can be adjusted within the enclosure relative to the optical element. 7. Apparatus according to claim 3, comprising means within the chamber that can be used to confirm whether or not the chamber is light tight once the door is closed. 8. Apparatus according to claim 1, wherein the sample location is provided by a sample platform mounted within the enclosure on which the sample can be placed. 9. Apparatus according to claim 1, comprising a frame within which the sample is held at the sample location to prevent the sample from deforming and/or to spatially position and orient the sample. 10. Apparatus according to claim 1, wherein the imaging means is a CCD camera. 11. Apparatus according to claim 10, wherein the camera is an emCCD camera. 12. Apparatus according to claim 1, wherein the imaging means includes an image detector with a surface on which the Cerenkov photons impinge and the imaging means is oriented so that the image detector surface is substantially normal to a wall of the enclosure it is mounted adjacent to. 13. Apparatus according to claim 1, wherein the radiation shield surrounds a majority of the sides of the imaging means. 14. Apparatus according to claim 1, wherein the one or more... | LIGHTPOINT MEDICAL LIMITED | G01N021/00 A61B005/00 G01T001/22 G01N023/04 G01N021/64
US09801590 | B2 | 2017-10-31 | Breathing disorder identification, characterization and diagnosis methods, devices and systems | 2012-05-17 | Disclosed herein are breathing disorder identification, characterization and diagnosis methods, devices and systems. In general the disclosed methods, devices and systems may rely on the characterization of breath sound amplitudes, periodic breath sounds and/or aperiodic breath sounds to characterize a breathing disorder as obstructive (e.g. obstructive sleep apnea—OSA) or non-obstructive (e.g. central sleep apnea—CSA). | 1. A computed-implemented method, automatically implemented by one or more processors of a computing device, for automatically characterizing recorded breath sounds as indicative of one of obstructive and central sleep apnea, the method comprising: identifying by the at least one process a segment of the recorded breath sounds encompassing at least one of an apnea and a hypopnea, said segment comprising a falling edge and a rising edge separated by a low-amplitude segment associated with said apnea or hypopnea;automatically evaluating by the at least one processor a breath sound amplitude profile pattern of said identified segment, wherein said pattern is at least partially defined by said falling edge and said rising edge, against one or more preset sound amplitude profile pattern criteria set to distinguish between segments indicative of obstructive sleep apnea and central sleep apnea;wherein said evaluating comprises evaluating a similarity between said rising edge and said falling edge of said identified segment, andwherein said one or more criteria comprise a similarity threshold set to distinguish obstructive from central apnea events;classifying by the at least one processor said identified segment as being indicative of one of obstructive and central sleep apnea based on a result of said evaluating step; andoutputting an indication as to one of obstructive and central sleep apnea based on said classifying. 2. The method of claim 1, said profile pattern comprising a fall/rise pattern characterized as either a crescendo-decrescendo pattern indicative of central sleep apnea or a gradual fall-abrupt rise pattern indicative of obstructive sleep apnea. 3. The method of claim 1, said criteria comprising a similarity index calculated by the at least one processor by comparing said falling edge and said rising edge using a Dynamic Time Warping (DTW) subroutine. 4. The method of claim 1 further comprising repeating the steps for multiple segments. 5. The method of claim 1, further comprising the steps of: automatically identifying by the at least one processor inspiration and expiration phases of the recorded breath sounds;isolating by the at least one processor expiration breath sounds from said recorded sounds; andextracting by the at least one processor said breath sound amplitude profile from said isolated expiration sounds, whereby said segment is representative of expiration sounds only. 6. The method of claim 1, said step of identifying said segment further comprising: scanning by the at least one processor said extracted breath sound amplitude profile to identify a prospect event segment;evaluating by the at least one processor characteristics of said prospect event segment for consistency with at least one of: one or more preset apnea-specific criteria, and one or more preset hypopnea-specific criteria distinct from said apnea-specific criteria; andidentifying by the at least one processor said prospect event as encompassing an apnea ... | University Health Network | A61B005/00 A61B007/00
US09792410 | B2 | 2017-10-17 | Multi-parameter diabetes risk evaluations | 2013-06-07 | Methods, systems and circuits evaluate a subject's risk of developing type 2 diabetes or having prediabetes using at least one defined mathematical model of risk of progression that can stratify risk for patients having the same glucose measurement. The model may include NMR derived measurements of GlycA and a plurality of selected lipoprotein components of at least one biosample of the subject. | 1. A method of evaluating a subject's risk of developing type 2 diabetes or all-cause death, comprising: placing an in vitro biosample of the subject in an NMR spectrometer,performing Nuclear Magnetic Resonance on the at least one in vitro biosample to measure at least one lipoprotein component and at least one of a measurement of (i) GlycA or GlycB or (ii) at least one branched chain amino acid, andprogrammatically calculating the risk index for diabetes or all-cause death of the subject using at least one defined mathematical model that includes a measurement of at least one lipoprotein component and at least one of the measurement of GlycA or GlycB,wherein GlycA is a composite Nuclear Magnetic Resonance signal from carbohydrate portions of acute phase reactant glycoproteins containing N-acetylglucosamine and/or N-acetylgalactosamine moieties that is detectable at about 2.00 ppm when Nuclear Magnetic Resonance measurements are performed at about 47 degrees C., andwherein GlycB is a composite NMR signal from the carbohydrate portions of acute phase reactant glycoproteins containing N-acetylneuraminic acid (sialic acid) moieties, more particularly from the protons of the 5-N acetyl methyl groups that is detectable at about 2.04 ppm when Nuclear Magnetic Resonance measurements are performed at about 47 degrees C. 2. The method of claim 1, wherein one or more of the at least one lipoprotein component forms a numerator, denominator or multiplication factor of at least one interaction parameter. 3. The method of claim 1, wherein the defined mathematical model of risk comprises GlycA, wherein the at least one lipoprotein component comprises an interaction parameter defined by the measurement of GlycA multiplied by a concentration of a defined subpopulation of high density lipoprotein (HDL) particles. 4. The method of claim 1, wherein the at least one lipoprotein component of the defined mathematical model of risk comprises a first interaction parameter of the measurement of GlycA multiplied by a concentration of a defined subpopulation of high density lipoprotein (HDL) particles and wherein the HDL subpopulation comprises only medium HDL particle subclasses with diameters between about 8.3 nm (average) to about 10.0 nm (average). 5. The method of claim 1, further comprising programmatically generating a report with a graph of risk of progression to type 2 diabetes in the future versus ranges of glucose levels with a visual indication of higher and lower risk values with respect to the calculated diabetes risk index for ease of identifying or understanding risk stratification for a particular glucose level. 6. The method of claim 1, wherein the at least one defined mathematical model of risk includes NMR derived measurements of GlycA and Valine and at least one interaction parameter comprising a high density lipoprotein (HDL) particle subpopulation. 7. The method of claim 1, wherein the defined mathematical risk model includes only NMR derived me... | LipoScience, Inc. | G01N033/48 G01N033/50 G06F019/00 G01R033/46 G06F017/10 A61B005/055 G01N024/08 G01R033/465 A61B005/00
US09795299 | B2 | 2017-10-24 | Modular physiological sensing patch | 2014-09-24 | A modular holder or patch is described that may be used with or as part of a wireless physiological sensing device. The wireless physiological sensing device may include a holder or patch, first and second electrodes, and an electronics package that may be removably coupled with the holder or patch and which may be in electrical contact with the first and second electrodes. The electronics package may include a housing, a wireless transceiver and electronic circuitry configured to process signals received via the first and second electrodes and the wireless transceiver. | 1. A remote physiological sensing device, comprising: a patch having first and second conductive traces for carrying signals from respective first and second signal input junctions of the patch to respective first and second signal output junctions of the patch, the patch also having a grounding conductor for receiving a ground signal without a corresponding conductive trace in the patch, the patch further including first and second electrodes coupled, respectively, with the first and second signal input junctions of the patch, wherein the patch is stretchable and is configured to determine one or more of a distance between the first and second electrodes or a relative position of the first and second electrodes based at least in part on the stretching; andan electronics package removably coupled with the patch and in electrical contact with the first and second signal output junctions of the patch and the grounding conductor, the electronics package including a housing, a wireless transceiver and electronic circuitry configured to process signals received via the first and second electrodes and the wireless transceiver. 2. The device of claim 1, further comprising a battery included within the patch. 3. The device of claim 2, wherein the battery is sealed within the patch. 4. The device of claim 1, further comprising: a battery located in the patch, wherein the battery is proximate to the first electrode and the electronics package is proximate to the second electrode. 5. The device of claim 4, the patch further comprising a connecting member between the electronics package and the battery through which the battery is electrically coupled to the electronics package. 6. The device of claim 1, further comprising a holder to which the electronics package is removably coupled, the holder being coupled to the patch. 7. The device of claim 6, wherein the holder is removably coupled to the patch. 8. The device of claim 1, further comprising an optical coupling between the patch and the electronics package. 9. The device of claim 1, the patch further comprises a unique identifier. 10. The device of claim 1, the patch further comprising at least a portion that is made of foam. 11. The device of claim 1, further comprising a battery included within the electronics package. 12. The device of claim 1, wherein the patch includes single-use electrodes. 13. The device of claim 1, the patch further comprising an adhesive layer for attachment to a person. 14. A method for remote monitoring of a person, the method comprising: receiving an electronics package at a patch having first and second conductive traces for carrying signals from respective first and second signal input junctions of the patch to respective first and second signal output junctions of the patch, the first and second signal input junctions being coupled with first and second electrodes, respectively, wherein the patch is stretchable and is configured to determine one or more of a ... | Covidien LP | A61B005/0408 A61B005/00 A61B005/0416 A61B005/08 A61B005/02
US09820636 | B2 | 2017-11-21 | Layered medium for three-dimensional imaging | 2017-05-31 | The attenuation and other optical properties of a medium are exploited to measure a thickness of the medium between a sensor and a target surface. Disclosed herein are various mediums, arrangements of hardware, and processing techniques that can be used to capture these thickness measurements and obtain three-dimensional images of the target surface in a variety of imaging contexts. This includes general techniques for imaging interior/concave surfaces as well as exterior/convex surfaces, as well as specific adaptations of these techniques to imaging ear canals, human dentition, and so forth. | 1. A system comprising: an inflatable membrane having a surface including a fluorescent substance that fluoresces over a range of wavelengths;a supply of a medium that absorbs a first wavelength of light within the range of wavelengths more than a second wavelength of light within the range of wavelengths, the supply being coupled to the inflatable membrane to permit selective delivery of the medium into the inflatable membrane;a pump coupled to the supply of the medium and configured to controllably deliver the medium to an interior of the inflatable membrane with a controlled pressure;a light source positioned within the inflatable membrane and capable of illuminating the surface of the inflatable membrane at an excitation wavelength for the fluorescent substance;a sensor positioned within the inflatable membrane to measure an intensity of the first wavelength and an intensity of the second wavelength in a direction of a location on the surface when illuminated by the light source; anda processor operatively coupled to the pump, the light source and the sensor, the processor programmed to determine a thickness of the medium in the direction based upon the intensity of the first wavelength and the intensity of the second wavelength, and to construct a three-dimensional image of a region of interest of the surface based upon a plurality of thickness measurements captured at the controlled pressure. 2. The system of claim 1, wherein the processor is further programmed to vary the controlled pressure between a first pressure and a second pressure between two or more consecutive calculations of the thickness. 3. The system of claim 2, wherein the processor is further programmed to calculate a compliance map based on the plurality of thickness measurements under different pressurizations. 4. The system of claim 1, wherein the inflatable membrane is an elastic membrane. 5. The system of claim 1, wherein the medium has a substantially greater coefficient of attenuation for the first wavelength than the second wavelength. 6. The system of claim 1, wherein the medium includes one or more of a gas, a liquid, and a gel. 7. The system of claim 1, wherein the sensor includes an optical coupling to an electronic imaging device outside the inflatable membrane. 8. The system of claim 1, wherein the sensor includes an electronic imaging device electronically coupled to the processor. 9. The system of claim 1, wherein the sensor includes a fiberscope with a lens for capturing images of the surface. 10. The system of claim 1, wherein the surface of the inflatable membrane has a known, uniform color. 11. The system of claim 1, wherein the inflatable membrane is shaped and sized for insertion into a human ear canal. 12. The system of claim 1, wherein the pump includes at least one of a rotary-type pump, a peristaltic pump, a reciprocating-type pump, a centrifugal pump, an educator-jet pump, and a hydraulic ram pump. 13. The system of claim 12, wherein th... | Massachusetts Institute of Technology | A61B001/04 G01B011/06 A61C005/62 A61B001/227 G01N021/64 G01B011/24 G01B011/02 A61B005/00 A61C019/04 A61B001/24 A61B001/015 A61B001/273 A61B001/018
US09814420 | B2 | 2017-11-14 | Burnout symptoms detection and prediction | 2016-03-09 | A method for detecting and presenting burnout symptoms associated with users is provided. The method may include receiving a plurality of predictor rules for monitoring and detecting a plurality of burnout indicators. The method may further include monitoring and receiving the plurality of burnout indicators. Additionally, the method may include storing the monitored and received plurality of burnout indicators. The method may also include detecting a plurality of burnout events associated with the stored monitored and received plurality of burnout indicators. The method may further include determining the plurality of burnout symptoms based on the detected plurality of burnout events. The method may also include determining at least one aggregate score based on the determined plurality of burnout symptoms. The method may further include presenting at least one alert and at least one recommendation based on the determined plurality of burnout symptoms and the determined at least one aggregate score. | 1. A method for detecting and presenting a plurality of burnout symptoms associated with at least one user, the method comprising: receiving a plurality of predictor rules for monitoring and detecting a plurality of burnout indicators;monitoring and receiving the plurality of burnout indicators;storing the monitored and received plurality of burnout indicators;detecting a plurality of burnout events associated with the stored monitored and received plurality of burnout indicators;determining the plurality of burnout symptoms based on the detected plurality of burnout events;determining at least one aggregate score based on the determined plurality of burnout symptoms; andpresenting at least one alert and at least one recommendation based on the determined plurality of burnout symptoms and the determined at least one aggregate score. 2. The method of claim 1, wherein the plurality of burnout indicators are selected from a group comprising at least one of a plurality of written text, a plurality of emails, a plurality of calendar entries, and a plurality of wearable device data. 3. The method of claim 1, wherein monitoring and receiving the plurality of burnout indicators further comprises: using a plurality of detectors selected from a group comprising at least one of at least one workstation, at least one wearable device, at least one email server, and at least one cloud repository, to monitor and receive the plurality of burnout indicators. 4. The method of claim 1, wherein monitoring and receiving the plurality of burnout indicators further comprises: using IBM Watson™ Personality Insights and IBM Watson™ Tone Analyzer to analyze the plurality of burnout indicators. 5. The method of claim 1, wherein detecting the plurality of burnout events further comprises: using a plurality of predictors to detect the plurality of burnout events based on the received plurality of predictor rules. 6. The method of claim 5, wherein detecting a plurality of burnout events further comprises: using a publish-subscribe messaging service associated with the plurality of predictors to detect the plurality of burnout events. 7. The method of claim 1, wherein presenting the at least one alert and the at least one recommendation further comprises: using at least one of a messaging communication, a phone communication, and an email communication to present the at least one alert and the at least one recommendation. 8. A computer system for detecting and presenting a plurality of burnout symptoms associated with at least one user, comprising: one or more processors, one or more computer-readable memories, one or more computer-readable tangible storage devices, and program instructions stored on at least one of the one or more storage devices for execution by at least one of the one or more processors via at least one of the one or more memories, wherein the computer system is capable of performing a method comprising:receiving a plurality of predictor rules for mo... | International Business Machines Corporation | G08B023/00 A61B005/16 G08B021/02 H04L029/08 H04L012/58
US09775992 | B2 | 2017-10-03 | Implantable electrode | 2016-02-04 | An electrode for use with an implantable medical device includes an alloy and a conductive oxide layer on a surface of the alloy. The alloy includes iridium and at least one of cobalt and iron. The conductive oxide layer includes iridium oxide. The conductive oxide layer has a thickness greater than about 5 nanometers. | 1. An electrode for use with an implantable medical device, the electrode comprising: an alloy consisting of: iridium;at least one of cobalt and iron; andoptionally, chromium; anda conductive oxide layer on a surface of the alloy, the conductive oxide layer having a thickness between 5 nanometers and 2,000 nanometers, the conductive oxide layer consisting of iridium oxide and, optionally, chromium oxide. 2. The electrode of claim 1, wherein the alloy includes the iridium in an amount from 10 wt. % to 60 wt. % and the balance is at least one of cobalt and iron. 3. The electrode of claim 1, wherein the alloy includes cobalt. 4. The electrode of claim 1, wherein the alloy includes chromium and the conductive oxide layer includes chromium oxide. 5. The electrode of claim 4, wherein the alloy includes the iridium in an amount from 10 wt. % to 50 wt. %, the chromium in an amount from 10 wt. % to 20 wt. %, and the balance is at least one of cobalt and iron. 6. The electrode of claim 1, wherein the conductive oxide layer has a thickness between 10 nanometers and 1000 nanometers. 7. The electrode of claim 6, wherein the conductive oxide layer has a thickness of 200 nanometers. 8. An implantable medical device comprising: a pulse generator including electronic circuitry for providing electro stimulation; andan elongate lead, the lead including: a proximal end connected to the pulse generator;a distal end opposite the proximal end; andat least one electrode disposed proximate to the distal end and electrically connected to the electronic circuitry, the at least one electrode including: an alloy consisting of: iridium;at least one of cobalt and iron; andoptionally, chromium; anda conductive oxide layer on a surface of the alloy, the conductive oxide layer having a thickness between 5 nanometers and 2,000 nanometers, the conductive oxide layer consisting of iridium oxide and, optionally, chromium oxide. 9. The device of claim 8, wherein the alloy is nickel-free and the alloy includes the iridium in an amount from 10 wt. % to 60 wt. % and the balance is at least one of cobalt and iron. 10. The device of claim 8, wherein the alloy includes cobalt. 11. The device of claim 8, wherein the alloy includes chromium and the conductive oxide layer includes chromium oxide. 12. The device of claim 11, wherein the alloy includes iridium in an amount from 10 wt. % to 50 wt. %, the chromium in an amount from 10 wt. % to 20 wt. %, and the balance is at least one of cobalt and iron. 13. The device of claim 8, wherein the implantable medical device is a cardiac function management system. 14. An electrode for use with an implantable medical device, the electrode comprising: an alloy including: iridium; andiron; anda conductive oxide layer on a surface of the alloy, the conductive oxide layer including iridium oxide and having a thickness between 5 nanometers and 2000 nanometers, wherein the alloy is cobalt-free. 15. The electrode of claim 14, wherein the alloy in... | Cardiac Pacemakers, Inc. | A61N001/05 A61B005/042 A61N001/36 H01J009/02 A61L031/02 C25D011/34 A61N001/375
US09810577 | B2 | 2017-11-07 | Apparatus and method for measuring a periodic signal | 2013-10-22 | A device for measuring a periodic signal includes: a first control unit for generating an electrical input signal (V1) of the period T; a light source for generating an optical input signal directed to an object being measured from the electrical input signal (V1); an optical receiver for detecting and converting the signal reflected from the object being measured, the signal corresponding to the optical input signal altered in terms of phase and amplitude, into an electrical measurement signal (V2); and a plurality of measurement channels connected in parallel between the optical receiver and a second control unit, each measurement channel being connected in series to a switching element, a filter element, and an analog-to-digital converter, wherein the second control unit is suitable for evaluating the measurement signals from the plurality of measurement channels, and in which the electrical measurement signal (V2) is applied to each of the plurality of measurement channels, the first control unit is connected to the plurality of switching elements and is suitable for actuating the switching elements for different time intervals in each case, and the analog-to-digital converters have a maximum sampling rate of less than 2×1/T. | 1. A device for measuring a periodic signal comprising: a first control unit for generating an electrical input signal of the period T;a light source for generating an optical input signal directed to an object being measured from the electrical input signal;an optical receiver for detecting and converting the signal reflected from the object being measured, said signal corresponding to the optical input signal altered in terms of phase and amplitude, into an electrical measurement signal; anda plurality of measurement channels connected in parallel between the optical receiver and a second control unit, each measurement channel being connected in series to a switching element, a filter element, and an analog-to-digital converter, the second control unit being suitable for evaluating the measurement signals from the plurality of measurement channels;wherein the electrical measurement signal is applied to each of the plurality of measurement channels, the same frequency of the electrical measurement signal being processed in each of the plurality of measurement channels;the first control unit is connected to the plurality of switching elements and is suitable for actuating the switching elements for different time intervals in each case; andthe analog-to-digital converters comprise a maximum sampling rate of less than 2×1/T. 2. The device of claim 1 wherein the electrical input signal is configured as a square-wave signal. 3. The device of claim 1 wherein the filter element is a low-pass filter. 4. The device of claim 1 wherein the switching elements in each of the plurality of measurement channels are configured as CMOS switches or CMOS changeover switches. 5. The device of claim 1 wherein the first control unit is connected to and synchronized with the second control unit. 6. The device of claim 1 wherein an integrated circuit comprises the first control unit, the second control unit, and the plurality of analog-to-digital converters. 7. The device of claim 1 wherein the first control unit comprises a pulse width modulation (PWM) generator. 8. A method for measuring a periodic measurement signal, the method comprising the following steps: sending to an object being measured an optical input signal from a light source based on an electrical input signal with the period T;receiving and converting an optical measurement signal corresponding to the optical input signal altered in terms of phase and amplitude into an electrical measurement signal in an optical receiver; andtapping the electrical measurement signal by each of a plurality of parallel-connected measurement channels, each of which is connected in series with a switching element, a low-pass filter, and an analog-to-digital converter; and whereinthe same frequency of the electrical measurement signal is processed in each of the plurality of measurement channels;the electrical measurement signal is tapped by each of the plurality of measurement channels at periodically recurring time... | Hamilton Bonaduz AG | G01J003/46 G01J003/28 A61B005/1455 G01J001/44 G01J003/44 G01J003/02
US09795322 | B1 | 2017-10-24 | Methods and systems for monitoring posture with alerts and analytics generated by a smart seat cover | 2017-04-17 | Disclosed are methods and systems for monitoring and correcting sitting posture of a user and for discouraging sedentary behavior. The system comprises a processor, a portable seat cover with embedded pressure sensors and an angle sensor, a notification device, and a non-transitory storage medium for storing program code. The program code, when executed by the processor, causes the processor to initiate an idleness timer to count how long the user has been sitting, monitor the user's sitting posture by receiving real-time sensor measurements; identify a sitting posture by applying a posture identification rule, generate a posture correction notification through the notification device if the identified sitting posture is not ergonomically correct, determine whether a sedentary threshold has been reached and generate a stand notification if the sedentary threshold has been reached. Embodiments of the present invention help improve sitting posture and reduce sedentary behavior to minimize associated adverse health effects. | 1. A system for monitoring and correcting sitting posture of a user, and for discouraging sedentary behavior of the user, comprising: a processor operatively connected to a clock;a portable seat cover comprising a backrest portion and a seat portion, adapted for attachment to a chair;a plurality of sensors embedded in the portable seat cover and each operatively connected to the processor, including at least three pressure sensors embedded in the backrest portion for monitoring shoulder and lower back positions of the user, at least one angle sensor embedded in the backrest portion for monitoring an inclination of the backrest portion relative to the seat portion, and at least two pressure sensors embedded in the seat portion for monitoring leg positions of the user;a notification device attached to the portable seat cover and operatively connected to the processor;a non-transitory physical storage medium for storing program code and operatively connected to the processor, the program code when executed by the processor causes the processor to: initiate an idleness timer to count how long the user has been sitting, wherein the idleness timer is activated when the user sits on the seat portion of the portable seat cover, and deactivated when the user stands up from sitting on the seat portion of the portable seat cover;monitor the user's sitting posture by receiving, at discrete intervals in real-time, a first plurality of sensor measurement values from the plurality of sensors;identify a sitting posture of the user by applying a posture identification rule to the first plurality of sensor measurement values, wherein the posture identification rule is based on the user's weight;determine whether the sitting posture is ergonomically correct by identifying a category that the sitting posture belongs to;in response to determining that the sitting posture is not ergonomically correct, generate a posture correction notification through the notification device to alert the user to correct the sitting posture;determine whether a sedentary threshold has been reached by comparing the idleness timer to the sedentary threshold;in response to determining that the sedentary threshold has been reached, generate a stand notification through the notification device to alert the user to stand up;initiate a stand timer to count how long the user has been standing, wherein the stand timer is activated when the user stands up from sitting on the seat portion of the portable seat cover, and deactivated when the user sits on the seat portion of the portable seat cover;determine whether the user has stood up in response to the stand notification;in response to determining that the user has not stood up in response to the stand notification, reduce the idleness timer by a predetermined amount of time;in response to determining that the user has stood up in response to the stand notification, pause the idleness timer until the user is seated again, increase the stand time... | Right Posture Pte. Ltd. | A61B005/00 A61B005/11 A47C007/42 A47C007/62 A47C007/00 A47C007/54
US09795313 | B2 | 2017-10-24 | Bioelectrode, and method and apparatus for processing biosignal using the same | 2015-05-14 | A bioelectrode including a plate, a first electrode disposed on a first side of the plate, and a second electrode disposed on the first side of the plate and separate from the first electrode. The bioelectrode further includes a first guard portion disposed on a second side of the plate, a second guard portion disposed on the second side of the plate and separate from the first guard portion, and a preamplifier configured to output a voltage signal based on a biosignal measured between the first electrode and the second electrode. | 1. A bioelectrode comprising: a plate having a first side and a second side, the first side being on an opposite side of the plate to the second side;electrodes in direct contact with the first side of the plate, the electrodes having no contact with each other and being configured to measure a biosignal; andguard portions in direct contact with the second side of the plate directly opposite to the electrodes, the guard portions having no contact with each other and being configured to reduce external noise affecting the biosignal measured by the electrodes. 2. The bioelectrode of claim 1, wherein the electrodes comprise a first electrode disposed on the first side of the plate and a second electrode disposed on the first side of the plate and separate from the first electrode, and wherein the guard portions comprise a first guard portion disposed on the second side of the plate and a second guard portion disposed on the second side of the plate and separate from the first guard portion. 3. The bioelectrode of claim 2, wherein: the first guard portion is disposed vertically separate from the first electrode; andthe second guard portion is disposed vertically separate from the second electrode. 4. The bioelectrode of claim 2, wherein the first guard portion and the second guard portion are configured to reduce external noise affecting the biosignal measured between the first electrode and the second electrode. 5. The bioelectrode of claim 2, wherein the first electrode and the second electrode are configured to measure an electromyogram of a user without making direct electrical contact with a body of the user. 6. The bioelectrode of claim 1, further comprising: a preamplifier configured to output a voltage signal based on the biosignal measured between the first electrode and the second electrode. 7. The bioelectrode of claim 6, wherein the first guard portion and the second guard portion comprise a conductive material, and are electrically connected to the preamplifier. 8. The bioelectrode of claim 1, further comprising: a shield disposed externally to the preamplifier, the first guard portion, and the second guard portion adjacent to the second side of the plate. 9. The bioelectrode of claim 8, wherein the shield is configured to form an air layer between the shield and the plate. 10. The bioelectrode of claim 8, wherein the shield comprises a conductive material, and is connected to a ground. 11. The bioelectrode of claim 1, wherein one of the electrodes and the guard portions are disposed between the plate and an insulation layer. 12. The bioelectrode of claim 1, wherein one of the electrodes and the guard portions are disposed internally into the plate. 13. A biosignal processing apparatus comprising: a bioelectrode configured to measure a biosignal without making direct electrical contact with a body of a user; anda biosignal processing portion configured to filter the biosignal, and amplify the filtered biosignal,wherein the bi... | Samsung Electronics Co., Ltd. SNU R&DB Foundation | G09G005/00 A61B005/04 A61B005/0492 A61B005/00 A61B005/0408 A61B005/0478 G06F003/01 A61B005/11
US09814935 | B2 | 2017-11-14 | Fitting system for sporting equipment | 2016-02-15 | Enables a fitting system for sporting equipment using an application that executes on a mobile phone for example to prompt and accept motion inputs from a given motion capture sensor to measure a user's size, range of motion, speed and then utilizes that same sensor to capture motion data from a piece of equipment, for example to further optimize the fit of, or suggest purchase of a particular piece of sporting equipment. Utilizes correlation or other data mining of motion data for size, range of motion, speed of other users to maximize the fit of a piece of equipment for the user based on other user's performance with particular equipment. For example, this enables a user of a similar size, range of motion and speed to data mine for the best performance equipment, e.g., longest drive, lowest putt scores, highest winning percentage, etc., associated with other users having similar characteristics. | 1. A fitting system for sporting equipment comprising: at least one motion capture element coupled to a piece of equipment and configured to capture motion capture data during a movement of said piece of equipment performed by a first user, said at least one motion capture element comprising a wireless communication interface configured to transmit said motion capture data from the at least one motion capture element;a database comprising measurements associated with a plurality of other users, and,an application configured to execute on a computer that is configured to wirelessly communicate with said at least one motion capture element, wherein the application is configured to accept said motion capture data from the motion capture element via the wireless communication interface;calculate at least one measurement of a characteristic of said first user orof said movement of said piece of equipment performed by said first userbased on the motion capture data;perform a comparison of said at least one measurement of said characteristic or of said movement to said database of said measurements associated with said plurality of other users;based on said comparison, identify a second user of said plurality of other users, wherein a corresponding measurement associated with said second user in said database of measurements is similar to said at least one measurement of a characteristic of said first user or of said movement of said piece of equipment performed by said first user; and,present information associated with a piece of equipment associated with said second user to the first user. 2. The fitting system for sporting equipment according to claim 1 wherein said identify said second user comprises: identify a user of said plurality of other users having a maximum value correlation with said at least one measurement associated with said first user. 3. The fitting system for sporting equipment according to claim 1 wherein said at least one measurement represents a height or an arm length, or a torso length, or a leg length of the first user or any combination thereof. 4. The fitting system for sporting equipment according to claim 1 wherein said at least one measurement represents a wrist to floor measurement of the first user. 5. The fitting system for sporting equipment according to claim 1 wherein said at least one measurement represents a hand size or longest finger size of the first user or both the hand size and longest finger size of the first user. 6. The fitting system for sporting equipment according to claim 1 wherein said application is further configured to: display one or more images to enable the first user to view a sporting event. 7. The fitting system for sporting equipment according to claim 1 wherein said application is further configured to: accept an input from the first user to purchase said piece of equipment associated with said second user. 8. The fitting system for sporting equipment according to claim 1 wherein s... | BLAST MOTION INC. | A63B024/00 G01P013/00 A61B005/11 A61B005/00
US09827333 | B2 | 2017-11-28 | Host cells with artificial endosymbionts | 2016-12-05 | The present invention is directed generally to eukaryotic cells comprising single-celled organisms that are introduced into the eukaryotic cell through human intervention and which transfer to daughter cells of the eukaryotic cell through at least five cell divisions, and methods of introducing such single-celled organisms into eukaryotic cells. The invention also provides methods of using such eukaryotic cells. The invention further provides single-celled organisms that introduce a phenotype to eukaryotic cells that is maintained in daughter cells. The invention additionally provides eukaryotic cells containing magnetotactic bacteria. | 1. A method of detecting viability of a mammalian cell, comprising the steps of: providing the mammalian cell, wherein the mammalian cell contains a magnetotactic bacterium, and imaging the mammalian cell wherein the imaging detects the magnetotactic bacterium whereby the viability of the mammalian cell is detected. 2. The method of claim 1, wherein the mammalian cell is a neural stem cell. 3. The method of claim 2, wherein the neural stem cell is imaged in a mammal. 4. The method of claim 3, wherein the mammal is a mouse. 5. The method of claim 4, wherein the neural stem cell is a murine neural stem cell. 6. The method of claim 4, wherein the neural stem cell is a human neural stem cell. 7. The method of claim 3, wherein the mammal is a human. 8. The method of claim 7, wherein the neural stem cell is a human neural stem cell. 9. The method of claim 1, wherein the mammalian cell is a T-cell. 10. A method of detecting viability of a stem cell, comprising the steps of: providing the stem cell, wherein the stem cell contains a magnetotactic bacterium; and imaging the mammalian cell wherein the imaging detects the magnetotactic bacterium whereby the viability of the stem cell is detected. 11. The method of claim 10, wherein the stem cell is imaged in a mammal. 12. The method of claim 11, wherein the mammal is a mouse. 13. The method of claim 12, wherein the stem cell is a murine stem cell. 14. The method of claim 12, wherein the stem cell is a human stem cell. 15. The method of claim 11, wherein the mammal is a human. 16. The method of claim 15, wherein the stem cell is a human stem cell. 17. The method of claim 1, wherein the imaging is a magnetic resonance imaging. 18. The method of claim 10, wherein the imaging is a magnetic resonance imaging. | Bell Biosystems, Inc. | C12Q001/02 A61K049/00 A61B005/05 A61N002/12 A61K049/18 B82Y005/00 B82Y015/00 C12N015/03 A01K067/027 G01N021/64 G01N033/50 G01N033/574 A61B005/055
US09782124 | B2 | 2017-10-10 | Apparel for physiological telemetry during athletics | 2014-11-07 | A user interface component and system for physiological parameters telemetry is provided. A user interface component may be reversibly or irreversibly attached to a garment that allows connection to a user's skin, such as via one or more holes in the garment. Signals transmitted through a conductive transfer layer in the user interface component are provided to a signal receiving unit which can collect and transmit physiological data. | 1. A system for physiological parameters telemetry, the system comprising: a garment having a skin-facing side and an outward-facing side, the garment configured to be worn by a user; andat least one user interface component located on the outward-facing side of the garment, the at least one user interface comprising at least: a conductive transfer layer having a first surface and a second surface, and a first terminal end and a second terminal end;a first insulating layer contacting the first surface of the conductive transfer layer, the first insulating layer having at least one aperture to allow at least a portion of the first surface of the conductive transfer layer to contact the user's skin;a second insulating layer having a first terminal end and a second terminal end, the second insulating layer contacting the second surface of the conductive transfer layer along the entire surface of the second surface of the conductive transfer layer except for at the first terminal end of the conductive transfer layer, wherein the first terminal end of the conductive transfer layer is oriented toward a head region of the user when the garment is being worn by the user, and wherein the first terminal end of the second insulating layer is generally aligned with the first terminal end of the conductive transfer layer; anda signal receiving unit adapted for connection to the conductive transfer layer second surface at the first terminal end of the conductive transfer layer. 2. The system of claim 1, wherein the garment is a shirt. 3. The system of claim 1, wherein the conductive transfer layer is an electrically conductive transfer layer. 4. The system of claim 1, wherein the first insulating layer is bonded to the outward-facing side of the garment. 5. The system of claim 1, wherein the first surface of the conductive transfer layer is adapted to contact the user's skin through an aperture in the garment when the garment is worn by the user. 6. The system of claim 5, wherein the at least one aperture of the first insulating layer is aligned with the aperture in the garment. 7. A physiological parameters telemetry system worn on a garment, the system comprising: a garment configured to be worn by a user;at least one user interface component comprising: a conductive transfer layer having a first surface and a second surface, and a first terminal end and a second terminal end, at least a first portion of the first surface adapted to contact the user's skin when the garment is worn by the user, wherein the first terminal end of the conductive transfer layer is oriented toward a head region of the user when the garment is worn by the user, and wherein the second terminal end of the conductive transfer layer is oriented toward a foot region of the user when the garment is worn by the user;a first insulating layer contacting at least a portion of the first surface of the conductive transfer layer, the first insulating layer having one or more apertures;a s... | NIKE, Inc. | A61B005/00 A61B005/024 A61B005/0245
US09775521 | B2 | 2017-10-03 | Remote patient management system | 2015-12-09 | A remote patient management system comprising a display and a processor. Execution of instructions cause the processor to: determine at least one first boundary value between a safe zone and a warning zone for each of at least one patient vital sign; determine at least one second boundary value between the warning zone and an alert zone for each of the at least one patient vital sign; repeatedly receive the at least one patient vital sign; determine a status value indicating the safe zone, warning zone, or alert zone; calculate a transition risk value; and display a patient status indicator which indicates the status level and the transition risk. | 1. A remote patient management system comprising: a display for displaying a graphical user interface;a plurality of sensors configured to measure vital signs of a patient, the sensors including a scale configured to measure a weight of the patient, a blood pressure cuff configured to measure a blood pressure of the patient, and a blood sugar sensor configured to measure blood sugar of a blood of the patient; anda computing device comprising a processor, wherein the computing device further comprises a memory containing machine readable instructions, wherein execution of the instructions cause the processor to:retrieve from the memory at least one first boundary value which defines a transition between a safe zone and a warning zone for individual ones of the vital signs of the patient;retrieve from the memory at least one second boundary value which defines a transition between the warning zone and an alert zone for individual ones of the vital signs of the patient;receive values of the vital signs of the patient as measured by the sensors;set a status value to alert if one or more of the values of the vital signs of the patient is in the alert zone;set the status value to warning if one or more of the values of the vital signs of the patient is in the warning zone and the alert is not displayed on the display;set the status value to safe if one or more of the values of the vital signs of the patient is within the safe zone; andcalculate a vital sign transition risk using a distance between individual ones of the vital signs of the patient and the first boundary value or the second boundary value. 2. The remote patient management system of claim 1, wherein the processor calculates the transition risk value using the values of the vital signs of the patient, the at least one first boundary value, and the at least one second boundary value, wherein the transition risk value is descriptive of a transition of the status value to warning if the status value is safe, wherein the transition risk value is descriptive of a transition of the status value to alert if the status value is warning. 3. The remote patient management system of claim 2, wherein the display displays a patient status indicator, wherein the patient status indicator indicates the status level and the transition risk. 4. The remote patient management system of claim 3, wherein the display displays a set of actions, wherein the actions are selected from a database using the status value and the transition risk. 5. The remote patient management system of claim 4, wherein the remote patient management system further comprises a home infrastructure device, wherein the home infrastructure device comprises at least one diagnostic medical device for measuring the vital signs of the patient; wherein the values of the vital signs of the patient is received from the home infrastructure device. 6. The remote patient management system of claim 5, wherein execution of the instructions causes ... | Koninklijke Philips N. V. | G08B019/00 G08B021/00 G08B023/00 A61B005/00 A61B005/0205 A61B005/021 A61B005/024 A61B005/145
US09808167 | B2 | 2017-11-07 | Systems, devices, and methods for determining severity of a stenosis within a luminal organ in the presence of a catheter | 2016-05-10 | Systems, devices, and methods for obtaining fractional flow reserve in the presence of a catheter. In a method of determining a fractional flow reserve in the presence of a catheter, the method comprises the steps of obtaining measurements of an inner luminal organ diameter proximal to, at, and distal to a stenosis and a length of the stenosis, obtaining a pressure drop measurement at the stenosis, calculating a volumetric flow of fluid through the inner luminal organ at the stenosis, and determining a stenotic pressure drop at the stenosis corresponding to dimensions of the guidewire as a function of the calculated volumetric flow of fluid through the inner luminal organ at the stenosis, wherein the stenotic pressure drop is indicative of a fractional flow reserve at or near the stenosis. | 1. A method of determining severity of a stenosis within a luminal organ in the presence of a catheter, comprising the steps of: obtaining a measurement of a pressure gradient change within the luminal organ at the stenosis using a catheter having a known size;obtaining luminal organ size measurements determined based upon conductance measurements obtained prior to, at, and distal to the stenosis using an impedance wire; anddetermining the severity of the stenosis based upon the measurement of the pressure gradient change, the luminal organ size measurements, and a size difference between the catheter and the impedance wire. 2. The method of claim 1, wherein the catheter has an outer diameter greater than 0.014″. 3. The method of claim 1, wherein the pressure gradient change is measured using the catheter selected from the group consisting of a pressure catheter and the catheter having a pressure sensor. 4. The method of claim 1, wherein the conductance measurements are obtained using the impedance wire at least partially positioned within a lumen of the catheter. 5. The method of claim 1, wherein the severity of the stenosis is based upon a flow reserve calculation. 6. The method of claim 1, wherein the severity of the stenosis is further determined using a mean aortic pressure within the luminal organ. 7. The method of claim 6, wherein the mean aortic pressure is measured using the catheter selected from the group consisting of a pressure catheter and the catheter having a pressure sensor. 8. The method of claim 6, wherein the mean aortic pressure is measured using the catheter. 9. The method of claim 1, wherein the severity of the stenosis is further determined using a calculated volumetric flow of fluid through the luminal organ at the stenosis. 10. A method of determining severity of a stenosis within a luminal organ in the presence of a catheter, comprising the steps of: obtaining a measurement of a pressure gradient change within the luminal organ at the stenosis using a catheter having a known size;obtaining luminal organ size measurements determined based upon conductance measurements obtained using an impedance wire positioned at least partially within a lumen of the catheter anddetermining the severity of the stenosis based upon the measurement of the pressure gradient change, the luminal organ size measurements, and a size difference between the catheter and the impedance wire. 11. The method of claim 10, wherein the luminal organ size measurements comprise a luminal organ size measurement proximal to the stenosis and a luminal organ size measurement distal to the stenosis. 12. The method of claim 10, wherein the severity of the stenosis is further determined using an identified length of the stenosis. 13. The method of claim 10, wherein the pressure gradient change is measured using the catheter selected from the group consisting of a pressure catheter and the catheter having a pressure sensor. 14. The method of claim 10... | 3DT Holdings, LLC | A61B005/027 A61B005/0215 A61B006/00 A61B008/12 A61B005/00 A61B005/053 A61B005/026 A61B008/06 A61B005/02
US09801586 | B2 | 2017-10-31 | Device for the, in particular, transperineal examination of patients | 2013-11-15 | Apparatus for the, in particular, transperineal examination of patients, consisting of two leg holders (1) arranged symmetrically in relation to the longitudinal axis of the patient, wherein the leg holders (1) have arranged between them a carriage (2) which runs parallel to the center plane of the patient and can be displaced, and arrested, in the direction of the longitudinal axis and, at its patient end, has a frame (3) which is oriented perpendicularly to the displacement direction and is suitable for accommodating a grid (4) and/or a targeting device or an NMR coil. | 1. A patient stabilization device, comprising: a carriage guidedly movable along a longitudinal axis of said patient stabilization device;a first leg holder;a second leg holder substantially mirror symmetric to said first leg holder relative to a plane comprising said longitudinal axis;an instrument guide frame mounted to said carriage; anda thigh spreading device comprising a plate-like first wing and a plate-like second wing, whereinsaid plate-like first wing and said plate-like second wing are guidedly moveable in a direction substantially perpendicular to said plane and form an acute angle with said longitudinal axis. 2. The device of claim 1, wherein said second wing is substantially mirror symmetric to said first wing relative to said plane. 3. The device of claim 1, comprising a mechanism that simultaneously moves said first wing in a first direction substantially perpendicular to said plane and said second wing in a second direction opposite to said first direction. 4. The device according to claim 1, wherein the instrument guide frame is capable of exchangeably receiving a grid and/or target device and/or a NMR tomographic imaging coil. 5. The device according to claim 1, wherein the instrument guide frame and/or the first and second leg holders are adjustable in their spatial orientation. 6. The device according to claim 1, wherein the carriage is movable along rails which extend parallel to the longitudinal direction of the device and can be fixed in any position along the rails. 7. The device according to claim 1, wherein the first and second leg holders and the carriage are integrated with one another as one piece. 8. The device according to claim 1, wherein the first and second wings are movable synchronously outwardly by means of a common drive. 9. The device according to claim 8, wherein the common drive is formed by coupling to the carriage. 10. The device according to claim 1, wherein the first and second wings are in connection with a scissor mechanism. 11. An instrument guiding method; comprising: positioning a patient on a patient stabilization device such that a spine of said patient extends substantially along a longitudinal axis of said patient stabilization device;positioning a first leg of said patient on a first leg holder of said patient stabilization device;positioning a second leg of said patient on a second leg holder of said patient stabilization device, which second leg holder is substantially mirror symmetric to said first leg holder relative to a plane comprising said longitudinal axis;spreading a thigh of said first leg from a thigh of said second leg using a thigh spreading device of said patient stabilization device, said thigh spreading device comprising a plate-like first wing and a plate-like second wing, each of said first wing and said second wing forming an acute angle with said longitudinal axis and being guidedly moveable in a direction substantially perpendicular to said plane;positioning a... | - | A61B005/00 A61B005/055 A61G013/12 A61G013/00 A61G013/10
US09788784 | B2 | 2017-10-17 | Accelerated evaluation of treatments to prevent clinical onset of neurodegenerative diseases | 2016-10-26 | Systems and methods for evaluating prospective treatments for progressive brain disorders like Alzheimer's disease and the progressive effects of aging involve the use of an imaging device communicatively coupled to a computing device. The imaging device may takes a plurality of brain imaging measurements from each of a plurality of human subjects who are divided into a treated group and an untreated group based on various subject data including whether the subjects carry certain alleles of genes known to increase the risk of developing the brain disorder at issue. The computing device receives the brain imaging measurements from the imaging device and applies an advanced processing method that permits enhanced computational efficiency when calculating rates of change for the two groups and determining whether any calculated difference between the rates is statistically significant. | 1. A method for determining the efficacy of a prospective prevention therapy for a neurodegenerative disease in persons having no clinical symptoms of the neurodegenerative disease, the method comprising: receiving at a computing device a plurality of brain imaging measurements from one or more imaging systems, the brain imaging measurements indicating a concentration of amyloid plaques or neurofibrillary tangles deposited in a brain of each of a plurality of subjects in a disease prevention therapy study, each subject having no clinical symptoms of the disease and at least one of Down syndrome or the concentration of amyloid plaques or neurofibrillary tangles deposited in the brain, wherein the concentration is at least equal to a threshold concentration of amyloid plaques or neurofibrillary tangles deposited in the brain, and wherein a first portion of the measurements were obtained by the one or more imaging systems from a first group of the subjects who first received an amount of a prospective prevention therapy for the disease, and a second portion of the measurements were obtained by the one or more imaging systems from a second group of the subjects who did not first receive the prospective prevention therapy; andexecuting instructions stored in memory of the computing device, wherein execution of the instructions by a processor of the computing device: extracts from the brain imaging measurements data corresponding to a region of interest identified within a brain atlas, the data indicating the concentration of amyloid plaques or neurofibrillary tangles deposited in the region of interest,compares a plurality of rates of change in the concentration of amyloid plaques or neurofibrillary tangles deposited in the region of interest calculated for the first group of the subjects to a plurality of rates of change in the concentration of amyloid plaques or neurofibrillary tangles deposited in the region of interest calculated for the second group of the subjects,determines that a difference between the rates of change in the concentration of amyloid plaques or neurofibrillary tangles deposited in the region of interest between the first group of the subjects and the second group of the subjects is statistically significant, andreports to a user of the computing device by way of a user interface that based on the statistically significant difference the prospective prevention therapy is efficacious in persons having no clinical symptoms of the disease and at least one of Down syndrome or at least the threshold concentration of amyloid plaques or neurofibrillary tangles deposited in the brain. 2. The method of claim 1, further comprising receiving at the computing device a second plurality of brain imaging measurements from a magnetic resonance imaging (MRI) machine, the MRI brain imaging measurements indicating a brain tissue volume associated with each subject. 3. The method of claim 1, further comprising receiving at the computing device a ... | Banner Health | A61B005/00 G06F019/00 A61B006/03 G01N033/68 A61B005/0476 A61B005/055 A61B005/05 G01N033/50
US09784730 | B2 | 2017-10-10 | Nanoparticle for targeting brain tumors and delivery of O6-benzylguanine | 2014-03-21 | Nanoparticle having a crosslinked chitosan-polyethylene oxide oligomer copolymer coating to which O6-benzylguanine is covalently coupled, compositions that include the nanoparticle, and methods for using the nanoparticle to treat brain cancers. | 1. A nanoparticle, comprising: (a) a core having a surface and comprising a core material;(b) a coating on the surface of the core, the coating comprising a crosslinked copolymer comprising a chitosan and a poly(ethylene oxide) oligomer, wherein the crosslinked polymer comprises crosslinks that are disulfide crosslinks; and(c) O6-benzylguanine covalently coupled to the coating. 2. The nanoparticle of claim 1, wherein the core material is a magnetic material. 3. The nanoparticle of claim 1, wherein the core material is selected from the group consisting of ferrous oxide, ferric oxide, silicon oxide, polycrystalline silicon oxide, silicon nitride, aluminum oxide, germanium oxide, zinc selenide, tin dioxide, titanium, titanium dioxide, nickel titanium, indium tin oxide, gadolinium oxide, stainless steel, gold, and mixtures thereof. 4. The nanoparticle of claim 1, wherein the copolymer is a graft copolymer having a chitosan backbone and poly(ethylene oxide) oligomer side chains. 5. The nanoparticle of claim 1, wherein the number of O6-benzylguanines/nanoparticle is from about 50 to about 2000. 6. The nanoparticle of claim 1 having a mean core size from about 2 to about 25 nm. 7. The nanoparticle of claim 1 having a hydrodynamic size from about 30 to about 250 nm. 8. The nanoparticle of claim 1 further comprising a targeting agent. 9. The nanoparticle of claim 8, wherein the targeting agent is chlorotoxin, or a variant or derivative thereof. 10. The nanoparticle of claim 1 further comprising a diagnostic agent. 11. A composition, comprising a nanoparticle of claim 1 and a carrier suitable for administration to a subject. 12. A method for introducing O6-benzyguanine into a cell comprising contacting a cell with a nanoparticle of claim 1. 13. A method for detecting cells or tissues by magnetic resonance imaging, comprising: (a) contacting cells or tissues of interest with a nanoparticle having affinity and specificity for the cells or tissues of interest, wherein the nanoparticle of claim 1; and(b) measuring the level of binding of the nanoparticle, wherein an elevated level of binding, relative to normal cells or tissues, is indicative of binding to the cells or tissues of interest. 14. A method for inhibiting O6-methylguanine-DNA methyltransferase (MGMT) in a subject, comprising administering a nanoparticle of claim 1 to the subject. 15. A method for treating a disease or condition treatable by inhibiting O6-methylguanine-DNA methyltransferase (MGMT), comprising administering a therapeutically effective amount of a nanoparticle of claim 1 to a subject in need thereof. 16. A method for treating a brain cancer, comprising administering a therapeutically effective amount of a nanoparticle of claim 1 to a subject in need thereof. 17. A method for treating a brain cancer, comprising administering a therapeutically effective amount of TMZ and a therapeutically effective amount of a nanoparticle of claim 1 to a subject in need thereof. 18. T... | University of Washington through its Center for Commercialization | A61B005/055 G01N033/50 A61K031/495 A61K047/48 A61K049/00 A61K009/00
US09826906 | B2 | 2017-11-28 | Apparatus and method for measuring biologic parameters | 2016-06-30 | Support structures for positioning sensors on a physiologic tunnel for measuring physical, chemical and biological parameters of the body and to produce an action according to the measured value of the parameters. The support structure includes a sensor fitted on the support structures using a special geometry for acquiring continuous and undisturbed data on the physiology of the body. Signals are transmitted to a remote station by wireless transmission such as by electromagnetic waves, radio waves, infrared, sound and the like or by being reported locally by audio or visual transmission. The physical and chemical parameters include brain function, metabolic function, hydrodynamic function, hydration status, levels of chemical compounds in the blood, and the like. The support structure includes patches, clips, eyeglasses, head mounted gear and the like, containing passive or active sensors positioned at the end of the tunnel with sensing systems positioned on and accessing a physiologic tunnel. | 1. A support structure housing for use in measuring biological parameters in a brain temperature tunnel, said support structure housing being used with a pair of eyeglasses, said support structure housing comprising a support structure mounted adjacent to nose pads of the pair of eyeglasses, said support structure extending from said nose pads,a temperature sensor mounted on the support structure being configured to measure temperature signals produced on the skin at an end of the brain temperature tunnel, andsaid nose pads locating said support structure so that said support structure is positioned adjacent to a medial corner of an eve above a medial canthal tendon and in a medial third of an upper eyelid when said eyeglasses is worn by an individual,said support structure including a reporting device. 2. The support structure as claimed in claim 1, wherein an attachment piece connects the temperature sensor to the support structure. 3. The support structure as claimed in claim 1, wherein a rotatable structure connects the temperature sensor to the support structure. 4. The support structure as claimed in claim 3, wherein the rotatable structure rotates through three axes. 5. The support structure as claimed in claim 2, wherein said attachment piece is a magnet. 6. The support structure as claimed in claim 1, wherein said reporting device is a visual display. 7. A support structure for use in measuring biological parameters in a brain temperature tunnel, said support structure comprising a pair of eyeglasses,a nose pad mounted on the pair of eyeglasses,a medial canthal cap mounted on the nose pad, said medial canthal cap extending from said nose pad,a sensor mounted on the medial canthal cap being configured to measure temperature signals produced on the skin at an end of the brain temperature tunnel,said nose pad locating said medial canthal cap so that said medial canthal cap is positioned adjacent to a medial corner of an eye above a medial canthal tendon and in a medial third of an upper eyelid when said eyeglasses is worn by an individual, anda reporting device mounted on the medial canthal cap, said reporting device producing a signal when the sensor reaches a predetermined threshold. 8. The support structure as claimed in claim 7, wherein an arm mounts said nose pad on the pair of eyeglasses. | GEELUX HOLDINGS, LTD. | A61B005/00 A61B005/01 A61B005/0205 G01K013/00 G02C011/00 A61B005/0476 A61B005/053 G08B005/36 A61B005/1455 A61B005/145 A61B005/1486
US09826923 | B2 | 2017-11-28 | Motion analysis method | 2014-10-30 | The present invention relates to an automated system of measuring and assessing an individual's improvement in motion (for example, a patient's recovery) through exercise performance. In an embodiment of the system of the present invention, two approaches for monitoring movement are introduced for quantifying an individual's performance. Both approaches consider the control population as their reference and consider the difference, or what is referred to in this application as the distance, between the individual's data and the control population as the measure of performance. | 1. A method for analysing an individual's motion through a computer programmed to process information, comprising the steps of: measuring control linear acceleration and angular velocity, using sensors on at least one control person performing a set of repetitions of an exercise;measuring individual linear acceleration and angular velocity, using sensors on said individual performing a set of repetitions of an exercise;inputting the measured control linear acceleration and angular velocity and individual linear acceleration and angular velocity into said computer to convert to joint angle positions, velocities and accelerations data for said individual and for said at least one control person;segmenting said data such that each segment begins with the start of an exercise repetition and ends when the exercise repetition is finished;extracting feature vectors from said data such that a control feature vector is V′H=VH(topfeatures) and an individual's feature vector is V′p=Vp(topfeatures) and said top features differentiate the individual from the at least one control person;calculating a mean of the at least one control person's feature vector such that μH=mean(V′H);calculating a diagonal matrix of standard deviations for the at least one control person's feature vector, such that ΣH=diaq(std(V′H));calculating a distance between one repetition of the exercise performed by said individual and the at least one control person's performance using δi=(V′Pi−μH)TΣH(V′Pi−μH); andmeasuring quality of a repetition set of said exercise, whereinQj=(ΔPj-μδ⁢⁢Hj)σδ⁢⁢Hja,and j is said exercise, ΔPj is the individual's said distance measure for the repetition set of said exercise, μδHj is the mean of the at least one control person's distance measure vector δHj , a is an index that penalizes Q based on exercise difficulty, and σδHj is the standard deviation of the at least one control person's distance measure vector δHj. 2. The method of claim 1, where a is 2. 3. A method for measuring the score for said individual in claim 1 for a set of more than one of said exercises j wherein the score isS=∑Ω∈Γ⁢⁢(nΩ∑Ω∈Γ⁢nΩ⁢μdHΩ⁢σdHΩ2)2-∑Ω∈Γ⁢⁢(nΩ∑Ω∈Γ⁢nΩ⁢QΩ)2where ┌ is said set of exercises, Ω is an exercise in the ┌, nΩ is the number of repetitions for exercise Ω, and QΩ is said quality measure of claim 1 calculated for exercise Ω. | - | G06F019/00 A61B005/11 A61B005/00
US09788796 | B2 | 2017-10-17 | System and method of adaptive interpretation of ECG waveforms | 2015-10-16 | A cluster database includes existing ECG datasets organized into clusters, wherein each existing ECG dataset includes an existing ECG waveform with at least one corresponding existing feature and existing interpretation. Each cluster is comprised of existing ECG datasets having a common existing feature. The cluster training module is executable by the processor to receive a new ECG waveform and a feature extracted from the new ECG waveform. The cluster training module then selects a cluster interpretation module based on the feature, wherein the cluster interpretation module is trained on one of the clusters from the cluster database. The cluster training module processes the new ECG waveform and/or the feature to provide a cluster interpretation output. The cluster interpretation output is then displayed on the user interface, and the cluster training module receives clinician input via the user interface accepting or rejecting the cluster interpretation output. | 1. A system for adaptive interpretation of ECG waveforms, the system comprising: a processor;a cluster database of existing ECG datasets organized into clusters, wherein each existing ECG dataset includes an existing ECG waveform with at least one corresponding existing feature and existing interpretation, and wherein the clusters are existing ECG datasets having a common existing feature;a user interface; anda cluster training module executable by the processor to: receive a new ECG waveform and a feature extracted from the new ECG waveform;select a cluster interpretation module based on the feature, wherein the cluster interpretation module is trained on one of the clusters from the cluster database;process the new ECG waveform and/or the feature to provide a cluster interpretation output;display the cluster interpretation output on the user interface;receive clinician input via the user interface accepting or rejecting the cluster interpretation output; andcreate a new cluster based on the extracted feature when the clinician input rejects the cluster interpretation output. 2. The system of claim 1, wherein the cluster training module is executable to further train the cluster interpretation module based on the clinician input. 3. The system of claim 1, further comprising a general interpretation module trained on a general database of existing ECG datasets, wherein the general interpretation module is executable by the processor to process the new ECG waveform and/or the extracted feature to provide a general interpretation output. 4. The system of claim 3, wherein the system is configured such that if the clinician input received by the cluster training module rejects the cluster interpretation output, then the general interpretation module displays the general interpretation output on the user interface. 5. The system of claim 4, wherein the general interpretation module is further executable by the processor to receive clinician input via the user interface accepting or rejecting the general interpretation output. 6. The system of claim 5, wherein the cluster training module is executable to create a new cluster based on the extracted feature if the clinician input rejects the cluster interpretation output and the general interpretation output. 7. The system of claim 3, wherein at least one of the general interpretation module and the cluster training module is a neural network or a decision tree. 8. A method for adaptive interpretation of ECG waveforms, the method comprising: receiving a new ECG waveform;extracting at least one feature from the new ECG waveform;selecting a cluster interpretation module based on the feature;processing the new ECG waveform and/or the extracted feature to provide a cluster interpretation output;displaying the cluster interpretation output on a user interface;receiving a clinician input via the user interface rejecting the cluster interpretation output; andusing the clinician input as negative feedback... | General Electric Company | A61B005/04 A61B005/00 A61B005/0402 A61B005/044
US09782086 | B2 | 2017-10-10 | Apparatus and method of measuring blood pressure | 2014-04-16 | Provided is a blood pressure measuring apparatus including: a pressurizing unit applying pressure to a blood vessel according to a first condition or a second condition; a pressure sensor sensing a sphygmus wave and a pressure of the blood vessel from the blood vessel under the first condition or the second condition; a standard blood pressure calculating unit for calculating a systolic standard blood pressure and a diastolic standard blood pressure; a continuous blood pressure calculating unit for calculating continuous blood pressure; and a repressurizing determining unit for determining whether pressure is applied to the blood vessel under the second condition during measuring of continuous blood pressure. | 1. An apparatus for measuring blood pressure comprising: a pressurizing unit for applying pressure to a blood vessel of an examinee under a first condition in which the pressure applied to the blood vessel changes or a second condition in which the pressure applied to the blood vessel is constant;a pressure sensor for sensing a sphygmus wave and a pressure of the blood vessel under the first condition or the second condition;a processor for calculating a systolic standard blood pressure and a diastolic standard blood pressure of the examinee, based on the sphygmus wave and the blood vessel pressure sensed by the pressure sensor under the first condition,for calculating continuous blood pressure of the examinee based on the sphygmus wave and the blood vessel pressure sensed by the pressure sensor under the second condition, the systolic standard blood pressure or the diastolic standard blood pressure under the first condition, and a regression equation for calculating tonometric blood pressure, andfor determining whether pressure applied to the blood vessel during measuring of the continuous blood pressure deviates from the second condition, by subtracting, from an amplitude of the sphygmus wave baseline sensed at a first time after starting the measurement of the continuous blood pressure, an amplitude of the sphygmus wave baseline sensed at a second time later than the first time, andcomparing a ratio of pulse amplitude of the sphygmus wave at the first time to pulse amplitude of the sphygmus wave at the second time to a second criterion,wherein when a result of the subtracting reaches at least a first criterion and the ratio is less than or equal to the second criterion, determining that said pressure applied to the blood vessel of the examinee deviates from the second condition during measuring of the continuous blood pressure and, based on the result, using the pressurizing unit to reapply pressure to the blood vessel of the examinee under the second condition during said measuring of the continuous blood pressure; andwhen the result of the subtracting reaches at least the first criterion, and the ratio is greater than the second criterion, determining that pressure applied to the blood vessel of the examinee satisfies the second condition during said measuring of the continuous blood pressure and that the actual blood pressure of the examinee is determined. 2. The apparatus of claim 1, wherein the sphygmus wave baseline is a line that is formed by connecting predetermined points of the sphygmus wave. 3. The apparatus of claim 2, wherein each of the predetermined points is a point between a high peak and a low peak of the sphygmus wave. 4. The apparatus of claim 1, wherein the sphygmus wave baseline is obtained by filtering the sphygmus wave using a low pass filter. 5. The apparatus of claim 1, wherein the first condition is an increased pressure condition under which the pressure is increased at a uniform pressure variation rate to apply... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/0225 A61B005/021 A61B005/022
US09820651 | B2 | 2017-11-21 | Methods and devices for labeling and/or matching | 2013-08-21 | Devices, such as computer readable media, and methods, such as automated methods, for labeling and/or matching. Some of the devices and methods are particularly useful for anatomical labeling of human airway trees. Some of the devices and methods are particularly useful for matching branch-points of human airway trees from represented in two or more graphs. | 1. A non-transitory computer readable medium comprising machine readable instructions for: accessing a representation of a target tree, the representation having edges and vertices;identifying a vertex in the representation that includes a first out-edge, a second out-edge and an in-edge, where the second out-edge is a terminal edge and where a first vertex defines one end of the first out-edge and a second vertex defines one end of the in-edge, and where the vertex, the first vertex, and the second vertex are different from each other; andintroducing an edge to the representation that extends between the first vertex and second vertex,where the target tree comprises a human airway tree. 2. The computer readable medium of claim 1, where the representation is based on volumetric image data taken using computed tomography. 3. The computer readable medium of claim 1, where the representation is based on volumetric image data taken using magnetic resonance. 4. The computer readable medium of claim 1, where the human airway tree is diseased. 5. The computer readable medium of claim 1, also including machine readable instructions for: computing an inheritance relationship for each possible pair of edges in the representation. 6. The computer readable medium of claim 5, also including machine readable instructions for: computing a topological distance between each possible pair of edges in the representation. 7. The computer readable medium of claim 1, also including machine readable instructions for: deleting from the representation any terminal branch having a length shorter than a pre-defined threshold length. 8. The computer readable medium of claim 1, also including machine readable instructions for: accessing data that includes (a) a reference tree having labels and reference tree edges, and (b) inheritance relationship data about some of the reference tree edges; andbuilding a tree association graph, where the building includes: adding a first edge to the tree association graph if (i) a corresponding edge of the representation has an inheritance relationship that is the same as the inheritance relationship of a corresponding reference tree edge, and (ii) the topological distance between a vertex of the first edge and a vertex of the corresponding edge of the representation tree is within a certain limit. 9. The computer readable medium of claim 8, where the tree association graph has a maximum weighted clique, and the computer readable medium also includes machine readable instructions for: identifying the maximum weighted clique. 10. The computer readable medium of claim 1, also including machine readable instructions for: building a tree association graph having segments and sub-trees; andlabeling some segments of the tree association graph in a series of steps, where one sub-tree is labeled in each step. 11. The computer readable medium of claim 1, also including machine readable instructions for: building a tree association graph ha... | University of Iowa Research Foundation | G06K009/00 A61B005/00 G06F019/00 A61B005/08
US09816923 | B2 | 2017-11-14 | System and method using OAM spectroscopy leveraging fractional orbital angular momentum as signature to detect materials | 2017-05-09 | An apparatus detects a material within a sample using signal generation circuitry that generates an orbital angular momentum (OAM) twisted light beam having at least one orbital angular momentum applied thereto. The signal generation circuitry applies the at least one OAM twisted light beam to the sample. The OAM twisted light beam includes at least one fractional OAM state, at least one intensity signature and at least one phase signature unique to the material within the sample. A detector receives the OAM twisted light beam after the OAM twisted light beam passes through the sample and detects the material responsive to a presence of a unique combination of the at least one fractional OAM state, the at least one intensity signature and the at least one phase signature. | 1. An apparatus that detects a material within a sample, comprising: signal generation circuitry that generates an orbital angular momentum (OAM) twisted light beam having at least one orbital angular momentum applied thereto, the signal generation circuitry applying the at least one OAM twisted light beam to the sample, the OAM twisted light beam including at least one fractional OAM state, at least one intensity signature and at least one phase signature; anda detector for receiving the OAM twisted light beam after the OAM twisted light beam passes through the sample and detecting the material responsive to a presence of a unique combination of the at least one fractional OAM state, the at least one intensity signature and the at least one phase signature. 2. The apparatus of claim 1, wherein the at least one intensity signature comprises at least one of a change of eccentricity, a shift of center of mass and a rotation of elliptical intensity. 3. The apparatus of claim 1, wherein the at least one phase signature comprises at least one change in a phase of the received OAM twisted light beam. 4. The apparatus of claim 1, wherein the signal generation circuitry further comprises: an emitting source that emits a first light beam comprising a plurality of plane waves; andorbital angular momentum generation circuitry that receives the first light beam and applies the at least one orbital angular momentum to the plurality of plane waves of the first light beam to generate the OAM twisted light beam, the OAM twisted light beam including the at least one fractional OAM state, the at least one intensity signature and the at least one phase signature unique to the material within the sample. 5. The apparatus of claim 4, wherein the orbital angular momentum generation circuitry comprises a spiral phase plate having a fraction step height to impart the at least one angular momentum to the first light beam. 6. The apparatus of claim 4, wherein the orbital angular momentum generation circuitry uses a spatial light modulator to impart the at least one orbital angular momentum to the first light beam. 7. The apparatus of claim 6, wherein the spatial light modulator limits a number of Gouy phases applied to the first light beam by the orbital angular momentum generation circuitry. 8. The apparatus of claim 6, wherein the spatial light modulator produces a superposition of a plurality of light modes, each mode having a different value m. 9. The apparatus of claim 6, wherein the spatial light modulator uses a spatial hologram to impart the at least one orbital angular momentum to the OAM twisted light beam. 10. The apparatus of claim 9, wherein the spatial hologram further includes a blazed grating to separate angularly a first fraction order of the at least one orbital angular momentum. 11. A method for detecting a material within a sample, comprising: generating an orbital angular momentum (OAM) twisted light beam having at least one orbital angular ... | NXGEN PARTNERS IP, LLC | G01N021/59 G01N033/483 G01N021/17 G01N033/487 G01N021/21 A61B005/145 A61B005/1455 G01N024/00
US09814909 | B2 | 2017-11-14 | Reference-based motion tracking during non-invasive therapy | 2013-02-06 | During a focused-ultrasound or other non-invasive treatment procedure, the motion of the treatment target or other object(s) of interest can be tracked in real time based on the comparison of treatment images against a reference library of images that have been acquired prior to treatment for the anticipated range of motion and have been processed to identify the location of the target or other object(s) therein. | 1. A method for tracking a moving anatomical object during a treatment sequence using a treatment device, the method comprising: (a) prior to the treatment sequence, (i) using an imaging device to acquire a series of reference images of an anatomical region comprising the anatomical object during motion thereof, each reference image corresponding to a different stage of the motion, and (ii) processing the images to determine, for each image separately, a location associated with the object in absolute coordinates in an imaging coordinate system; and(b) during the treatment sequence using the treatment device, (i) acquiring treatment images of the anatomical region using the imaging device, (ii) for a plurality of the treatment images, identifying best-matching reference images corresponding thereto based on a metric of image similarity therebetween, and (iii) tracking the object in the plurality of the treatment images based on the locations associated with the object in the absolute coordinates of the best-matching reference images, wherein the metric of image similarity is determined based on image data corresponding to unreconstructed images. 2. The method of claim 1, wherein the treatment sequence comprises treatment of the anatomical object. 3. The method of claim 2, wherein the treatment sequence comprises steering a focused ultrasound beam onto the object based on the tracking. 4. The method of claim 1, wherein the treatment sequence comprises treatment of a target other than the anatomical object. 5. The method of claim 4, further comprising, during the treatment, shaping a focused ultrasound beam onto the target so as to avoid the anatomical object based on the tracking. 6. The method of claim 1, wherein the treatment sequence is part of a treatment procedure comprising a plurality of time-separated treatment sequences each comprising at least one exposure of an anatomical target to therapeutic energy, wherein at least one of the acquired reference images used during a treatment sequence is a treatment image obtained during a previous treatment sequence. 7. The method of claim 6, wherein each exposure is subjection of the anatomical target to acoustic energy. 8. The method of claim 1, further comprising monitoring a temperature in the anatomical region by performing baseline subtraction between the treatment images and the corresponding reference images. 9. The method of claim 1, wherein processing the reference images comprises identifying at least one anatomical landmark in each of the reference images, the location associated with the object being a location of the at least one anatomical landmark, the location of the at least one anatomical landmark being fixed relative to a location of the object. 10. The method of claim 9, wherein tracking the target comprises inferring the location of the target from the location of the anatomical landmark in the corresponding reference image. 11. The method of claim 1 wherein the locati... | INSIGHTEC, LTD. | G06K009/00 A61N007/00 A61B005/055 A61N007/02 G06T007/73 G06T007/246 A61B005/01 A61B017/00 A61B090/00
US09804148 | B2 | 2017-10-31 | Analyte sensor with lag compensation | 2016-04-29 | In particular embodiments, methods, devices and systems including calibrating analyte data associated with a monitored analyte level received from an analyte sensor based on a reference measurement, determining a lag time constant associated with the calibrated analyte data, and performing lag correction of the calibrated analyte data based on the determined time lag constant are disclosed. | 1. A system for monitoring an analyte using an in vivo sensor, the system comprising: an analyte sensor configured for positioning in contact with bodily fluid under a skin surface;a display;a processor; anda memory operatively coupled to the processor, the memory storing instructions which, when executed by the processor, causes the processor to: obtain a plurality of analyte sensor data associated with a monitored analyte level from the analyte sensor;obtain one or more measured reference analyte data;generate a first subset of data pairs, each pair of the first subset comprising a reference analyte data and one or more of the plurality of analyte sensor data;determine a sensor sensitivity for the analyte sensor using the first subset of data pairs;obtain a plurality of calibrated sensor data by calibrating the plurality of analyte sensor data using the sensor sensitivity;generate a second subset of data pairs, each pair of the second subset comprising a reference analyte data and one or more of the plurality of calibrated sensor data;estimate a time lag constant based on the second subset of data pairs;estimate a lag corrected analyte level for the analyte sensor based on the estimated time lag constant; andoutput data for the lad corrected analyte level on the display as part of a graphical user interface. 2. The system of claim 1, further comprising instructions which, when executed by the processor, cause the processor to generate the first subset by including only pairs associated with analyte sensor data having an absolute rate of change that is below a predetermined threshold. 3. The system of claim 2, wherein the predetermined threshold is about 6% change in an analyte estimate per minute. 4. The system of claim 1, further comprising instructions which, when executed by the processor, cause the processor to obtain a plurality of time lag constants, to obtain a combined time lag constant based on the plurality of time lag constants and the estimated time lag constant, and to calibrate continuous glucose data from the analyte sensor using the combined time lag constant. 5. The system of claim 1, further comprising instructions which, when executed by the processor, cause the processor to determine the sensor sensitivity using the first subset of data pairs by, for each data pair of the first subset, forming a ratio for the reference data and the one or more of the plurality of analyte sensor data associated with the data pair. 6. The system of claim 1, further comprising instructions which, when executed by the processor, cause the processor to generate the second subset by at least including only pairs associated with calibrated analyte sensor data having an absolute rate of change exceeding a predetermined threshold. 7. The system of claim 1, further comprising instructions which, when executed by the processor, cause the processor to generate the first subset by at least excluding pairs associated with a reference below a predete... | Abbott Diabetes Care Inc. | A61B005/1486 A61B005/1495 A61B005/01 G01N033/487 A61B005/145 G06F019/00 A61B005/00
US09788752 | B2 | 2017-10-17 | Implantable dielectrometer | 2015-11-20 | Diagnostic apparatus (24) includes a sealed case (80), comprising a biocompatible material and configured for implantation within a body of a human subject (22). A dielectrometric probe (26, 50, 63, 66, 70, 102, 160) is connected to the case and includes first and second conductors (40, 42, 54, 56, 64, 67, 68, 72, 74, 162, 164), which are configured to be placed in proximity to a target tissue (34) in the body. A driving circuit (82), which is contained in the case, is coupled to apply a radio-frequency (RF) signal to the probe and to sense the signal returned from the probe. Processing circuitry (84) is configured to evaluate, responsively to the returned signal, a dielectric property of the target tissue. | 1. A diagnostic method, comprising: implanting a dielectrometer probe, comprising first and second conductors, in proximity to a target tissue in a body of a human subject;applying a radio-frequency (RF) signal to the probe; andevaluating, responsively to a signal from the probe, a dielectric property of the target tissue including finding a complex permittivity of the target tissue,wherein finding the complex permittivity of the target tissue includes one of: measuring a reflection of the signal from the probe, which is indicative of an impedance mismatch at the target tissue;measuring a delay of the signal transmitted through the probe; andmeasuring a resonant frequency of the probe. 2. The method according to claim 1, wherein applying the RF signal comprises applying RF signals and sensing returned signals at multiple different frequencies. 3. The method according to claim 1, wherein finding the complex permittivity further includes measuring an impedance between the first and second conductors. 4. A diagnostic method, comprising: implanting a dielectrometer probe, comprising first and second conductors, in proximity to a target tissue in a body of a human subject;applying a radio-frequency (RF) signal to the probe; andevaluating, responsively to a signal from the probe, a dielectric property of the target tissue;wherein the first and second conductors are arranged to: provide a coaxial contact to the target tissue at a distal end of the probe, ordefine a resonant ring in proximity to the target tissue. 5. The method according to claim 4, wherein the first and second conductors are arranged to provide a coaxial contact to the target tissue at a distal end of the probe, and the coaxial contact comprises a patch, which is placed against the target tissue. 6. The method according to claim 4, wherein the first and second conductors are arranged to provide a coaxial contact to the target tissue at a distal end of the probe, and the coaxial contact comprises a coaxial tip, in which the first conductor is concentrically contained within the second conductor, wherein the tip is inserted into the target tissue. 7. A diagnostic method, comprising: implanting a dielectrometer probe, comprising first and second conductors, in proximity to a target tissue in a body of a human subject; the first and second conductors arranged to define a transmission line in proximity to the target tissue;applying a radio-frequency (RF) signal to the probe; andevaluating, responsively to a signal from the probe, a dielectric property of the target tissue; wherein: the transmission line is defined by arranging the first conductor as a ground plane and the second conductor as a microstrip parallel to the ground plane; orthe transmission line comprises a coplanar waveguide. 8. The method according to claim 4, wherein the first and second conductors are arranged to define a transmission line in proximity to the target tissue and the transmission line is defined by arrang... | ZOLL MEDICAL ISRAEL LTD. | A61B005/00 A61B005/053 A61B005/05 A61N001/375 A61B005/0205 A61B001/313 A61B005/07 A61B005/08 A61B017/34 A61B005/01 A61B005/0215 A61B005/042 A61B005/145 A61N001/365 A61N001/37 A61N001/372
US09785745 | B2 | 2017-10-10 | System and method for providing multi-organ variability decision support for extubation management | 2015-04-07 | A decision support system is provided for the management of extubation in intensive care unit patients. Based on multi-organ variability analysis of physiological signals, the proposed system transforms acquired waveforms into clinical information such as the risk of failing extubation and the probability of passing extubation. Furthermore, a variety of mechanisms are provided for displaying the extracted information to support a clinician's decisions. | 1. A computer-implemented method for supporting extubation management, the method comprising: accessing a waveform dataset database to obtain waveform data previously collected from a plurality of patients, the waveform data comprising waveforms associated with patients that experienced one of a plurality of extubation outcomes;obtaining from one or more variability analyzers, waveform data and a plurality of variability metrics in a variability data set indicative of a degree and character to which at least one patient parameter changes over intervals of time for the plurality of patients;computing a plurality of different predictive models using the variability data and the waveform data obtained from the database;applying the waveform data and the variability data set for a corresponding patient to each of the plurality of predictive models to obtain a plurality of logistic regression predictive models indicative of an extubation outcome by selecting measures of variability from the variability data set and determining at least one threshold for determining the extubation outcome, each logistic regression providing a probability of the extubation outcome;averaging the plurality of logistic regression predictive models to generate an ensemble average of logistic regression predictive models indicating a final probability of the extubation outcome; anddisplaying via a user interface and/or outputting via a report, at least one decision support index determined from the ensemble of logistic regression predictive models that is indicative of the final probability of the extubation outcome, along with standard clinical factors for the corresponding patient, at least one of the standard clinical factors normally used to evaluate readiness for extubation, wherein the displaying is performed in a clinical location where the extubation is performed to assist in a clinical decision related to the extubation. 2. The method of claim 1, further comprising displaying with the at least one decision support index, additional information corresponding to a subject being analyzed. 3. The method of claim 1, wherein the at least one patient parameter comprises respiratory rate and heart rate. 4. The method of claim 1, wherein the at least one index comprises an index estimating a likelihood of failing extubation. 5. The method of claim 4, wherein the index estimating a likelihood of failing extubation is evaluated subsequent to a spontaneous breathing trial (SBT). 6. The method of claim 4, wherein the index estimating a likelihood of failing extubation is transformed into a fold increase in risk of failing extubation by computing for each of a plurality of ranges, a likelihood that the subject is going to fail extubation. 7. The method of claim 1, wherein the at least one index comprises an index estimating a likelihood of passing extubation. 8. The method of claim 7, wherein the index estimating a likelihood of passing extubation is evaluated subsequent t... | Ottawa Hospital Research Institute | G01N033/48 G06F019/00 A61B005/083 A61M016/04 G06G007/58 A61B005/01 A61B005/0402 A61B005/145
US09788801 | B2 | 2017-10-17 | Gas analyzer, respirator or anesthesia device with operating guide | 2014-01-23 | A process for the control of a gas analyzer, respirator or anesthesia device (10), a control module (19) and a gas analyzer, respirator or anesthesia device (10) are provided with a control module (19). The process is used for the output of online help and for guiding the user via a sequence of operating steps that are executed via defined operational controls (17). Each operational control (17) is designed with a selection signal generator and with an activation signal sensor in order to make possible a visual selection of the respective operational control. The selection may be coupled with the output of a visual indication on the graphical user interface (13). | 1. A control module for a gas analyzer, a respirator or an anesthesia device for the guided execution of a sequence of operating steps, which require at least partially actuation of operational controls, and wherein the gas analyzer, respirator or anesthesia device comprises a plurality of operational controls comprising: at least one selection signal generator receiving at least one request signal from the control module and consequently outputting at least one selection signal; andat least one activation signal generator generating at least one activation signal in the case of a correct operation of the operational control and for sending the activation signal to the control module and wherein the control module comprises:at least one I/O interface with a graphical user interface on which at least one visual indication of the respective position of the operational control to be operated is outputted, whereby the visual indication comprises a direction indicator, which indicates the respective position of the operational control to be operated relative to the position of the visual indication on the graphical user interface, said graphical user interface comprising a touch screen, said touch screen comprising an input element, said direction indicator being oriented in a direction of one of said input element and an element outside of said touch screen; anda device memory, in which a device operating file is stored, the device operating file comprising at least one sequence of operating steps and means for reading out the sequence of operating steps. 2. A control module in accordance with claim 1, further comprising a status detection unit continuously detecting a device status of the gas analyzer, respirator or anesthesia device during the operation of the device and which, in case of an absent and/or incorrect operation of an operational control, outputs a context-sensitive message and/or the respectively generated visual indication with the dynamically adapted direction indicator on the graphical user interface. 3. A gas analyzer, respirator or anesthesia device, comprising: a graphical user interface;a plurality of operational controls for operating the respirator or the gas analyzer, respirator or anesthesia device the plurality of operational controls comprising: at least one selection signal generator receiving at least one request signal from the control module consequently outputting at least one selection signal;at least one activation signal generator generating at least one activation signal in the case of a correct operation of the operational control; anda control module receiving the activation signal from the activation signal generator and controlling the respirator or gas analyzer, respirator or anesthesia device within the framework of executing a sequence of operating steps, wherein the control module comprises: at least one I/O interface with a graphical user interface on which at least one visual indication of the respect... | DRÄGERWERK AG & CO. KGAA | G08B005/00 A61B005/00 A61B005/08 A61M016/00 A61M016/01 G06F019/00
US09820659 | B1 | 2017-11-21 | Monitoring device with a pedometer | 2015-08-17 | A monitoring device with a pedometer function preferably comprises an article, an optical sensor, an accelerometer and processor. The optical sensor preferably comprises a photodetector and a plurality of light emitting diodes. A sensor signal from the optical sensor is processed with a filtered accelerometer output signal from the accelerometer to generate a pedometer function. | 1. An arm-based pedometer comprising: an article to be worn on an arm of the user, the article having a first surface and a second surface;a processor;a powering source;a short range wireless transceiver for transmitting data from the article;an accelerometer; andan optical sensor for generating a signal corresponding to the heart rate of the user,wherein the optical sensor is selected from the group consisting of an optical sensor comprising a light-to-voltage photodetector capable of transmitting a digital signal and at least one light emitting diode capable of radiating light ranging from 550 nanometers to 1100 nanometers, an optical sensor comprising a light-to-frequency photodetector capable of transmitting a digital signal and at least one light emitting diode capable of radiating light ranging from 600 nanometers to 1100 nanometers, an optical sensor comprising a plurality of light-to-voltage photodetectors capable of transmitting a digital signal and at least one light emitting diode capable of radiating light ranging from 550 nanometers to 1100 nanometers, an optical sensor comprising a plurality of light-to-frequency photodetectors capable of transmitting a digital signal and at least one light emitting diode capable of radiating light ranging from 550 nanometers to 1100 nanometers, an optical sensor comprising a light-to-voltage photodetector capable of transmitting a digital signal and at least one light emitting diode capable of radiating light at 567 nanometers, and an optical sensor comprising a light-to-voltage photodetector capable of transmitting a digital signal and a plurality of light emitting diodes capable of radiating green light ranging from 500 to 570 nanometers;wherein the processor is configured to generate a real-time heart rate from a signal from the optical sensor;wherein the processor is configured to process a signal from the accelerometer in a pedometer function to determine distance traveled by the user. 2. The arm-based pedometer according to claim 1 wherein the accelerometer is a 3-axis accelerometer. 3. The arm-based pedometer according to claim 1 wherein the short range wireless transceiver operates on a communication protocol using BLUETOOTH, PART 15, WiFi, Infrared, and ultrasound. 4. The arm-based pedometer according to claim 1 further comprising a display on the article to display pedometer data for the user. 5. A system for monitoring a plurality of real-time vital signs of a user and providing a pedometer function, the system comprising: a monitoring device integrated into clothing of the user, the monitoring device comprising an optical sensor for generating a real-time vital sign signal corresponding to the heart rate of the user, the optical sensor is selected from the group consisting of an optical sensor comprising a light-to-voltage photodetector capable of transmitting a digital signal and at least one light emitting diode capable of radiating light ranging from 550 nanometers to 1100 nanome... | Impact Sports Technologies, Inc. | A61B005/02 A61B005/0205 G01C022/00 A61B005/11 A61B005/00 A61B005/024 A61B005/021 A61B005/08
US09788802 | B2 | 2017-10-17 | Method for ascertaining an absolute scan region on a patient | 2015-07-21 | In a method and imaging device for ascertaining an absolute scan region on a patient, who is positioned on an examination table, for a subsequent medical imaging examination by operation of the medical imaging device, the height of the patient is ascertained, a relative examination region is ascertained, and the absolute scan region of the patient for the subsequent medical imaging examination is determined in a processor using the patient height and the relative examination region. | 1. A method for acquiring magnetic resonance (MR) data from a patient, comprising: providing an examination table of an MR data acquisition scanner with a marking on the examination table at a spacing A from an edge of the examination table, said examination table being movable in a longitudinal direction of said MR scanner;placing a patient on the examination table at a patient position Ref relative to said marker on said examination table;providing a control computer with an electronic input that designates a height H of the patient and with another electronic input that designates a section of the patient, said section of the patient being selected from the group consisting of a defined examination region of the patient and a defined organ of the patient;in said control computer, accessing a database, after placing the patient at said patient position on said examination table, that correlates said section with said patient of said height H and, dependent on said correlation, calculating, as a range or percentage of said height H of said patient, a relative examination region that encompasses said section within said patient, said relative examination region having a start SW_Rel and an end EW_Rel in said longitudinal direction and a relative examination region length between said start and said end that is said range or percentage of said height H of said patient;in said control computer, determining a final examination region in said scanner, from which MR image data will be acquired, said final examination region having a final examination region start SWAbs at which acquisition of said medical image data by said scanner begins, and a final examination region end EWAbs at which said acquisition of said medical image data by said scanner ceases, said final examination region and being at a length LAbs from said final examination region start SWAbs, according to:SWAbs=-1⁢(H-H⁢⁢EWRel100-(Ref-A)),andLAbs=-H⁢⁢EW_Rel-SW_Rel100; andwith said control computer, operating said MR scanner to acquire said MR image data from said final examination region, and making the acquired MR data available in electronic form as an output from said control computer. 2. A magnetic resonance (MR) imaging apparatus comprising: an MR data acquisition;an examination table movable in a longitudinal direction into and out of said MR scanner, said examination table having a marker therein at a spacing A from an edge of the examination table, and being adapted to receive a patient thereon at a patient position Ref relative to said marker on said examination table;a control computer configured to operate said scanner in order to acquire MR image data from said patient;said control computer being configured to receive an electronic input that designates a height H of the patient and another electronic input that designates a section of the patient, said section of the patient being selected from the group consisting of a defined examination region of the patient and a defin... | Siemens Aktiengesellschaft | A61B005/05 A61B005/00 A61B005/055 A61B005/107 G01R033/30 G01R033/54
US09788760 | B2 | 2017-10-17 | Ocular micro tremor (OMT) sensor, system and method | 2015-12-08 | An ocular micro tremor (OMT) sensor, system and method for measuring micromovements of an individual's eyeball having an amplitude of 40 micrometers or less to provide a variable voltage biosignal for measuring the individual's brain stem activity. The OMT sensor includes a flexible piezo-active sensing element mounted at the individual's closed or opened eyelid so as to be deflected by the micromovements of the eyeball. A shielded flexible ribbon assembly supplies the biosignal generated by the piezo-active sensing element to an OMT amplifier located on the individual's skin where the biosignal is amplified. The OMT amplifier is interconnected with a signal processor and a display by which graphical and numerical representations of the biosignal are made accessible to an anesthesiologist, intensivist or clinician. | 1. An apparatus comprising: a sensor comprising: a sensing element comprising: a piezoelectric element with a first surface and a second surface;a first electrically conductive surface in direct contact with the first surface;a second electrically conductive surface in direct contact with the second surface; anda flexible ribbon assembly comprising a conductive trace, wherein the conductive trace is in electrical communication with first electrically conducting surface;a conductive shield; andan electrical insulator wherein the electrical insulator is located between the first electrically conductive surface and the conductive shield;wherein the sensor is configured to conform to a shape of an eye; and the sensing element is configured to register micro-movements of an eyeball having an amplitude of 40 micrometers. 2. The apparatus of claim 1, wherein the piezoelectric element is 20 microns to 150 microns microns thick. 3. The apparatus of claim 1, wherein the piezoelectric element comprises Polyvinylidene fluoride. 4. The apparatus of claim 1, wherein the sensor further comprises an adhesive configured to attach to an eyelid. 5. The apparatus of claim 1, wherein the conductive shield comprises gold. 6. The apparatus of claim 1, wherein the sensing element is further configured to register micro-movements of an eyeball having amplitudes of 500 nanometers. 7. The apparatus of claim 1, wherein the sensing element is further configured to register micro-movements of an eyeball having a frequency of 84 Hz. 8. The apparatus of claim 1, wherein the conductive shield is in electrical communication with the first electrically conductive surface. 9. An apparatus comprising: a sensor comprising: a sensing element comprising: a piezoelectric element with a first surface and a second surface;a first electrically conductive surface in direct contact with the first surface;a second electrically conductive surface in direct contact with the second surface; anda flexible ribbon assembly comprising a conductive trace, wherein the conductive trace is in electrical communication with first electrically conducting surfacea conductive shield; andan electrical insulator wherein the electrical insulator is located between the first electrically conductive surface and the conductive shield;wherein the sensor is configured to conform to a shape of an eye; and a flexural rigidity of the flexible ribbon assembly is less than or equal to 10×10-4 lbs-in4. 10. The apparatus of claim 9, wherein the piezoelectric element is 20 microns to 150 microns microns thick. 11. The apparatus of claim 9, wherein the piezoelectric element comprises Polyvinylidene fluoride. 12. The apparatus of claim 9, wherein the sensor further comprises an adhesive configured to attach to an eyelid. 13. The apparatus of claim 9, wherein the conductive shield comprises gold. 14. The apparatus of claim 9, wherein the sensing element is further configured to register micro-movements of an eyeba... | - | A61B005/11 A61B003/113 A61B005/00
US09801582 | B2 | 2017-10-31 | Thigh adhesion quantitative measurement system | 2015-11-12 | The thigh adhesion quantitative measurement system is a sensor system for precisely locating fibers and/or bands of connective tissue formed between layers of body tissue, binding the layers together and not permitting them to glide over each other. This bonding is called an adhesion. The present system is designed to determine quantitatively the degree of adhesion and the location of high-pressure adhesion(s) such that medical removal procedures (e.g., liposuction) can be precisely performed. The system makes use of a scanning device having an array of pressure sensors, the scanning device being defined by a flexible plate wrapped by a belt to a patient's limb, such as a thigh. A flexible cable connects the plate to a body-worn transmitter that wirelessly transmits real-time pressure distribution images received from the flexible plate to a computer for use by doctors preparing for the liposuction procedure. | 1. A thigh adhesion quantitative measurement system, comprising: a processor;a scanning device, the scanning device consisting of a flexible plate and a flexible plate pressure sensor array disposed thereon and in operable communication with the processor, the flexible plate pressure sensor array having a central active sensor area and two inactive areas disposed on opposing ends of the central active sensor area, a (0,0) datum coordinate marker disposed on a corner of an active portion of the central active sensor area, the scanning device conforming to and taking pressure readings of the inner thigh region of a subject's body, sensor belts extending from opposite edges of the scanning device to secure the scanning device to the subject's inner thigh region;means for digitizing the pressure readings taken by the scanning device to obtain a pressure profile providing indicia of high pressure points within the region;a wireless data converter in operable communication with the processor; anda pressure scanning display computer in communication with the wireless data converter, the pressure scanning display computer having a screen for displaying the pressure profile as images, wherein the pressure profile images are presented as a color map of pressure over the subject's inner thigh region. 2. The thigh adhesion quantitative measurement system according to claim 1, further comprising a multiplexer connected between the scanning device and the processor, the multiplexer selecting inputs to scan elements of the pressure sensor continuously into the processor. 3. The thigh adhesion quantitative measurement system according to claim 2, wherein the processor further comprises a signal conditioner circuit for conditioning the pressure data before the data is digitized. 4. The thigh adhesion quantitative measurement system according to claim 3, wherein the signal conditioner circuit is selected from the group consisting of a programmable gain amplifier circuit and a low pass filter circuit. 5. The thigh adhesion quantitative measurement system according to claim 3, further comprising a DC offset compensator connected to the signal conditioner circuit. 6. The thigh adhesion quantitative measurement system according to claim 1, further comprising a connection cable extending from the flexible plate and attachable to the processor, the connection cable facilitating operable communication of the scanning device with the processor. 7. The thigh adhesion quantitative measurement system according to claim 1, further comprising: a transmitter unit belt; anda transmitter unit belt loop disposed on a housing of the wireless data converter, the transmitter unit belt being threaded through the transmitter unit belt loop to secure the wireless data converter to the subject's body. | KING SAUD UNIVERSITY | A61B005/00
US09821162 | B2 | 2017-11-21 | Deep brain stimulation for sleep disorders | 2015-04-14 | In some examples, a medical device delivers electrical stimulation to the external portion of the globus pallidus of a brain of a patient in order to treat a sleep impairment of the patient. In some examples, the electrical stimulation may be delivered via one or more electrodes implanted in the GPe of the brain. In some examples, an electrical stimulation device is configured to deliver electrical stimulation therapy to the GPe based on detection of a sleep state of a patient. The sleep state may include, for example, a state in which the patient is awake and intending on sleeping, is awake and attempting to sleep or has initiated sleep. In addition, in some examples, an electrical stimulation device is configured to deliver electrical stimulation therapy to the GPe based on detection of an awake state of a patient. | 1. A method comprising: determining, by a processor, a patient is in a sleep state; andcontrolling, by the processor, an electrical stimulation generator to deliver deep brain electrical stimulation to a globus pallidus externus of a brain of the patient according to a therapy program defining electrical stimulation configured to induce sleep of the patient, based on the determination that the patient is in the sleep state. 2. The method of claim 1, wherein controlling the electrical stimulation generator to deliver electrical stimulation to the globus pallidus externus comprises controlling the electrical stimulation generator to initiate delivery of electrical stimulation to the globus pallidus externus according to the therapy program in response to detecting the sleep state. 3. The method of claim 1, wherein determining the patient is in the sleep state comprises receiving user input indicating the sleep state. 4. The method of claim 1, wherein determining the patient is in the sleep state comprises: receiving, by the processor, a signal indicative of a patient parameter of the patient; anddetermining whether the patient is in the sleep state based on the signal indicative of the patient parameter. 5. The method of claim 4, wherein the patient parameter comprises at least one of an activity level, a patient posture, a heart rate, a respiration rate, a respiratory volume, a blood pressure, a blood oxygen saturation, a partial pressure of oxygen within blood, a partial pressure of oxygen within cerebrospinal fluid, muscular activity, a core temperature, an arterial blood flow or a galvanic skin response of the patient. 6. The method of claim 1, wherein determining the patient is in the sleep state comprises: determining an activity level of the patient; anddetermining whether the patient is in the sleep state based on the activity level of the patient. 7. The method of claim 1, further comprising: determining, by the processor, the patient is in an awake state; andcontrolling the electrical stimulation generator to deactivate delivery of electrical stimulation to the globus pallidus externus according to the therapy program in response to determining the patient is in the awake state. 8. The method of claim 7, wherein determining the patient is in the awake state comprises receiving user input indicating the awake state. 9. The method of claim 7, wherein determining the patient is in the awake state comprises: receiving, by the processor, a signal indicative of a physiological parameter of the patient; anddetermining whether the patient is in the awake state based on the signal indicative of the physiological parameter. 10. The method of claim 7, wherein determining the patient is in the awake state comprises: determining an activity level of the patient; anddetermining whether the patient is in the awake state based on the activity level of the patient. 11. The method of claim 1, wherein controlling the electrical stimulation generat... | Medtronic, Inc. | A61N001/36 A61B005/00 A61B005/048 A61N001/05
US09814396 | B2 | 2017-11-14 | Pulse wave signal processor | 2016-10-10 | In a pulse wave signal processor, a first filter attenuates a frequency component in an acquired pulse wave signal or in a differentiated pulse wave signal, and the attenuated frequency component is more than or equal to a first frequency. A second filter attenuates a frequency component in the acquired pulse wave signal, and the attenuated frequency component is less than the first frequency and more than or equal to a second frequency. A characteristic-point extractor extracts a characteristic point that exists in each single pulse of the pulse wave signal filtered by the second filter, and a signal separator partitioning the pulse wave signal or the differentiated signal filtered by the first filter into sections such that each section includes one of the extracted characteristic points. The partitioned sections are overlapped such that the characteristic points are coincident with each other, and arithmetically averaged. | 1. A pulse wave signal processor comprising: a signal acquirer configured to acquire a pulse wave signal;a first filter configured to attenuate a frequency component in the pulse wave signal acquired by the signal acquirer or in a differentiated signal that is obtained by differentiating the pulse wave signal, the frequency component attenuated by the first filter being more than or equal to a first frequency;a second filter configured to attenuate a frequency component in the pulse wave signal acquired by the signal acquirer, the frequency component attenuated by the second filter being less than the first frequency and more than or equal to a second frequency;a characteristic-point extractor configured to extract a characteristic point that exists in each single pulse of the pulse wave signal processed by the second filter;a signal separator configured to partition the pulse wave signal processed by the first filter or the differentiated signal processed by the first filter into sections corresponding to respective pulses of the pulse wave signal such that each section includes one of the extracted characteristic points; andan averaging calculator configured to overlap the partitioned sections such that the characteristic points are coincident with each other and arithmetically averaging the overlapped sections. 2. The pulse wave signal processor according to claim 1, wherein the first frequency is higher than or equal to 10 Hz. 3. The pulse wave signal processor according to claim 1, wherein the second frequency is higher than 1 Hz and lower than 10 Hz. 4. The pulse wave signal processor according to claim 1, wherein the signal acquirer is configured to acquire the pulse wave signal from a wrist of an examinee. 5. The pulse wave signal processor according to claim 1, wherein the differentiated signal is obtained by differentiating the pulse wave signal twice or four times. 6. The pulse wave signal processor according to claim 1, wherein the characteristic point is a largest point at which a value of the pulse wave signal is largest, a smallest point at which the value of the pulse wave signal is smallest, or a point which separates a distance between the largest point and the smallest point at a predetermined ratio. 7. The pulse wave signal processor according to claim 1, further comprising an estimator configured to estimate stiffness of a blood vessel or a blood pressure by using the pulse wave signal or the differentiated signal averaged by the averaging calculator. | DENSO CORPORATION | A61B005/024 A61B005/02 A61B005/0205 A61B005/00 A61B005/021
US09814486 | B2 | 2017-11-14 | Guide block for biopsy or surgical devices | 2015-05-13 | A method includes providing a guide block with a body having a proximal face, a distal face, a length between the proximal and distal faces, a central longitudinal axis extending between the proximal and distal faces, and at least one passageway disposed within the body that extends between the proximal and distal faces, which is configured to receive a tissue accessing device and which is not parallel to the central longitudinal axis; selecting a grid opening of the grid-like breast receiving pressure plate and selecting a rotational orientation of the guide block relative to the selected grid opening to provide access to a desired location in the patient's breast; and inserting the guide block into the selected grid opening of the grid-like breast receiving pressure plate of the magnetic resonance imaging device at the selected rotational orientation, and advancing the guide block in the selected grid opening. | 1. A guide block for use with a grid-like breast receiving pressure plate to guide a tissue accessing device to a desired location within a patient's breast, the guide block comprising: a body configured to be inserted into a grid opening of the grid-like breast receiving pressure plate, the body including:a proximal face with at least a first proximal opening;a distal face with at least a first distal opening, the distal face being spaced apart from the proximal face;a body length between the proximal face and the distal face, with a central longitudinal axis extending within the body between the proximal face and the distal face, with each of the first proximal opening of the proximal face and the first distal opening of the distal face being offset from the central longitudinal axis;an outer wall that extends from the distal face in a direction toward the proximal face, wherein an extent of the outer wall is less than an extent of the body length;a shoulder on a proximal portion of the body, the shoulder being defined at least in part by the proximal face having a larger surface area than a surface area of the distal face, the outer wall intersecting the shoulder, and the shoulder configured to project radially outwardly from the outer wall; anda first passageway having a single fixed orientation within the body and having a cylindrical shape, the first passageway configured to extend within the body in a distal direction from the first proximal opening in the proximal face to the first distal opening in the distal face, the first passageway being oriented at an angle that in the distal direction diverges from the central longitudinal axis within the body, andthe body configured to be inserted into the grid opening of the grid-like breast receiving pressure plate in a plurality of discrete orientations wherein each of the plurality of discrete orientations of the body represents a different direction of the distal direction divergence of the first passageway from the central longitudinal axis relative to the grid-like breast receiving pressure plate. 2. The guide block of claim 1, wherein each of the proximal face and the distal face is planar, the angle originating within the body. 3. The guide block of claim 1, wherein each of the proximal face and the distal face is planar, and the first passageway has a central axis that extends within the body to diagonally intersect the central longitudinal axis of the body at a location within the body. 4. The guide block of claim 1, wherein the outer wall is comprised of four equal sides that extend proximally from the distal face, and a distal portion of the first passageway extends away from the central longitudinal axis at an acute angle. 5. The guide block of claim 1, wherein the outer wall is comprised of four equal sides that extend proximally from the distal face, and the first passageway is oriented to intersect the central longitudinal axis within the body. 6. The guide block of claim 1, ... | SenoRx, Inc. | A61B017/34 A61B019/00 A61B005/055 A61B010/02 A61B005/00 A61B090/17 A61B090/11 A61B090/14
US09797768 | B2 | 2017-10-24 | Light sensing device for sensing ambient light intensity | 2013-11-26 | The invention relates to a light sensing device for sensing ambient light intensity, comprising at least one ambient light sensor and an occlusion detector for detecting an object occluding the ambient light sensor. The invention is further related to a corresponding method for sensing ambient light intensity. | 1. A light sensing device for sensing ambient light intensity, the light sensing device comprising: at least one ambient light sensor configured to measure ambient light intensity;an occlusion detector configured to detect an object occluding the ambient light sensor;a recording device configured to store ambient light intensity values measured by the ambient light sensor and corresponding occlusion status of the ambient light sensor as a function of time; anda processor configured to: determine an ambient light intensity threshold based on one or more parameters, the one or more parameters comprising date and/or time of day;activate the occlusion detector to measure a first occlusion measurement based on the determined ambient light intensity threshold; anddetermine timing for subsequent activations of the occlusion detector based on the first occlusion measurement and based on the one or more parameters. 2. The light sensing device of claim 1, wherein the occlusion detector comprises an infrared sensor and an infrared light source. 3. The light sensing device of claim 2, comprising a casing on which the infrared sensor, the infrared light source and the ambient light sensor are positioned on the same side. 4. The light sensing device of claim 1, wherein the light sensing device is a wrist-wearable device comprising a wristband. 5. The light sensing device of claim 1, further comprising an indicator device for generating an audible, visual and/or haptically perceivable notification signal in case of the detection of an object by the occlusion detector. 6. The light sensing device according to claim 1, wherein the processor is configured to activate the occlusion detector responsive to the ambient light intensity being below a predetermined threshold. 7. A method for sensing ambient light intensity by means of a light sensing device comprising at least one ambient light sensor an occlusion detector for detecting an object occluding the ambient light sensor a recording device and a processor, comprising the following steps: performing an ambient light intensity measurement with the ambient light sensor;determining, with the processor, an ambient light intensity threshold based on one or more parameters, the one or more parameters comprising date and/or time of day;activating, with the processor, the occlusion detector based on the determined ambient light intensity threshold;performing an occlusion measurement with the occlusion detector to determine if the ambient light sensor is occluded by the object;recording, with the recording device, ambient light intensity values measured by the ambient light sensor and corresponding occlusion status of the ambient light sensor; anddetermining, with the processor, timing for subsequent activations of the occlusion detector based on the first occlusion measurement and based on the one or more parameters. 8. The method of claim 7, wherein the ambient light measurement is performed within temporally li... | KONINKLIJKE PHILIPS N.V. | G01J001/44 G01J001/02 A61B005/00 G01J001/42 A61B005/16
US09789021 | B2 | 2017-10-17 | Medical imaging apparatus, bed for medical imaging apparatus, and pedal unit | 2012-11-14 | Realized is a pedal mechanism that does not lock and does not provide an operation feeling unless operated in a determined order. The pedal mechanism has first and second pedals 105 and 106, first and second plates 108 and 109, and a locking bar 115. The radius of the second plate 109 is smaller than that of the first plate 108, and the second plate does not come into contact with the locking bar in a state where a circular-arc region of the first plate supports the locking bar 115. Therefore, the second pedal does not give an operation feeling even when stepped on, and the second plate cannot be locked. In a locked state where the locking bar is located in a cutout region of the first plate, the locking bar contacts the second plate, the second pedal provides an operation feeling, and the second plate can be locked. | 1. A medical imaging apparatus comprising: a bed that mounts an object;a main body that images the object, anda pedal unit that is provided at the bed to operate the bed, whereinthe pedal unit has first and second pedals, a first plate that rotates through the operation of the first pedal, a second plate that rotates through the operation of the second pedal, and a locking bar that is arranged so as to be capable of being in contact with outer peripheries of the first and second plates,central axes of rotation of the first plate and second plate are coaxial,the first plate has a circular-arc region with a radius r1 from the central axis and a cutout region with a distance r2 from the central axis smaller than the radius r1, at an outer periphery thereof, and the second plate has a circular-arc region with a radius from the central axis equal to the distance r2 at an outer periphery thereof, andthe second plate does not come into contact with the locking bar in a state where the circular-arc region of the first plate supports the locking bar through the operation of the first pedal, and the second plate rotates while contacting the locking bar through the operation of the second pedal in a locked state where the locking bar is located in the cutout region of the first plate,wherein one end of a first wire and one end of a second wire are connected to the first and second plates, respectively, and the other ends of the first and second wires are respectively connected to members of the bed to be operated. 2. The medical imaging apparatus according to claim 1, wherein the pedal unit further includes a third pedal, a release plate that rotates through the operation of the third pedal, and a biasing member that applies a force that pushes the locking bar against the outer peripheries of the first and second plates, and the release plate that has rotated through the operation of the third pedal lifts the locking bar against the force of the biasing member to allow the locking bar to move from the cutout region of the first plate to the circular-arc region. 3. The medical imaging apparatus according to claim 1, wherein a protrusion is arranged between the circular-arc region and the cutout region at an outer periphery of the first plate, andthe locking bar rides over the protrusion from the circular-arc region and moves to the cutout region when the first plate rotates through the operation of the first pedal. 4. The medical imaging apparatus according to claim 3, wherein a locking region with a distance from the central axis equal to the radius r2 of the circular-arc region, and a protrusion arranged between the circular-arc region and the locking region are provided at an outer periphery of the second plate, andthe locking bar rides over the protrusion from the circular-arc region and moves to the locking region when the second plate rotates through the operation of the second pedal. 5. The medical imaging apparatus according to claim 4, wherein t... | HITACHI, LTD. | A61B005/05 A61G013/10 A61B005/055 A61B006/03 A61B006/04 A61B006/00
US09775527 | B2 | 2017-10-03 | Method and apparatus for simulation of pressure transducer for measurement of blood pressure | 2015-05-12 | Disclosed is an apparatus, system, and method for utilizing a pressure transducer simulator that is configured for use with a pressure sensor device and a patient monitoring device. The pressure transducer simulator may be configured to generate a simulation of the output of an analog pressure sensing device to be compatible with the patient monitoring device. The simulation may be based upon a pressure signal received from the pressure sensor device. In one embodiment, the pressure transducer simulator includes a digital potentiometer that is configured to generate an analog signal based upon the pressure signal received from the pressure sensor device and wherein the digital potentiometer is configured to transmit the analog signal to the patient monitoring device. | 1. A pressure transducer simulator configured for use with a pressure sensor device and a monitoring device, the pressure transducer simulator configured to generate a simulation of the output of an analog pressure sensing device compatible with the monitoring device, wherein the simulation is based upon a pressure signal received from a pressure sensor device, the pressure transducer simulator comprising: a digital potentiometer configured to generate an analog signal based upon the pressure signal, wherein the digital potentiometer is configured to transmit the analog signal to the monitoring device; anda processor to implement a scaling function applied to the digital potentiometer based upon the pressure signal from the pressure sensor device and a reference excitation signal from the monitoring device to ensure the analog signal transmitted to the monitoring device is approximately equivalent to the pressure signal from the pressure sensor device. 2. The pressure transducer simulator of claim 1, wherein the digital potentiometer includes an adjustable wiper that is adjustable to select the resistance of the digital potentiometer. 3. The pressure transducer simulator of claim 2, wherein the processor determines a digital wiper setting for the adjustable wiper to select the resistance of the digital potentiometer. 4. The pressure transducer simulator of claim 3, wherein the processor implements an auto-zeroing command. 5. The pressure transducer simulator of claim 4, wherein, when the auto-zeroing command is implemented, the adjustable wiper is adjusted to set the resistance of the digital potentiometer such that the analog signal produced by the digital potentiometer and transmitted to the monitoring device is approximately zero. 6. The pressure transducer simulator of claim 3, wherein the processor implements an operate command. 7. The pressure transducer simulator of claim 6, wherein, after implementation of the operate command, the digital potentiometer produces an analog signal based upon the pressure signal from the pressure sensor device and the digital wiper setting and transmits the analog signal to the monitoring device, wherein the analog signal is approximately equivalent to the pressure signal from the pressure sensor device. 8. The pressure transducer simulator of claim 3, further comprising a buffer coupled to the digital potentiometer. 9. A method for simulating a pressure sensor device as an analog pressure sensing device to generate a simulation of the output of an analog pressure sensing device compatible with a monitoring device, wherein the simulation is based upon a pressure signal received from the pressure sensor device, the method comprising: configuring a digital potentiometer to generate an analog signal based upon the pressure signal;transmitting the analog signal to the monitoring device; andimplementing a scaling function applied to the digital potentiometer based upon the pressure signal from the pressure... | Edwards Lifesciences Corporation | A61B005/00 A61B005/021
US09826924 | B2 | 2017-11-28 | Hearing assessment method and system | 2014-02-26 | A system and a method for a healthcare provider to easily, accurately and efficiently provide a hearing test and a pressure test, i.e., a hearing assessment, for their patients, in a way that accurately detects hearing loss and provides recommendations for seeking further assistance from qualified hearing health specialists is provided. | 1. A system to enable a healthcare provider to administer a portable interactive patient hearing assessment, the system comprising: a hand-held, mobile hearing test device for administration of a threshold hearing test by the healthcare provider, the hearing test device configured to transmit at least four distinct pure tone frequencies to a patient;a hand-held, mobile input device configured for activation by the patient upon hearing the at least four pure tone frequencies, and for transmitting one or more signals to a processor indicative of the patient's response to the hearing test in response to the patient's activation of the mobile input device, wherein the processor is configured to(i) process the one or more signals indicative of the patient's response to the hearing test,(ii) issue one or more hearing test result signals based on the one or more signals indicative of the patient's response to the hearing test which is based on the patient's ability to hear each of the at least four pure tone frequencies and(iii) process the hearing test result signals to determine the patient's hearing results; anda tympanometer configured to conduct a pressure test of each ear of the patient, the tympanometer configured to indicate pressure test results,wherein upon receipt of the one or more signals indicative of the patient's response to the hearing test and the pressure test results, the processor provides a report and recommendation based on the pressure test results and the patient's hearing results, wherein(i) when the patient's hearing results indicate no hearing loss, the report is provided to the healthcare provider and/or to the patient,(ii) when the patient's hearing results indicate a hearing loss, the pressure test results are normal bilaterally, and an input by the healthcare provider to the system that a physical exam of the patient's ears indicates that the ears are normal bilaterally, the report and the recommendation comprise a referral to an audiologist, and(iii) when the patient's hearing results indicate a hearing loss and when the pressure test results are not normal bilaterally and/or an input by the healthcare provider to the system that the physical exam of the patient's ears indicates that the ears are not normal bilaterally, the report and the recommendation comprise a referral to an otolaryngologist. 2. The system of claim 1, further comprising a calibration device configured to calibrate the hand-held, mobile hearing test device for uniformity of the at least four pure tone frequencies. 3. The system of claim 1, further comprising a simulator configured to demonstrate effects of the hearing loss, in an audible manner, measured by the hearing test to a person accompanying the patient so that the accompanying person understands a scope of the patient's hearing loss. 4. The system of claim 3, wherein the hand-held, mobile input device is a portable electronic device configured to host a plurality of modules, wherein the mod... | dB Diagnostic Systems, Inc. | A61B005/12
US09808190 | B2 | 2017-11-07 | Advanced analyte sensor calibration and error detection | 2013-12-30 | Systems and methods for processing sensor data and self-calibration are provided. In some embodiments, systems and methods are provided which are capable of calibrating a continuous analyte sensor based on an initial sensitivity, and then continuously performing self-calibration without using, or with reduced use of, reference measurements. In certain embodiments, a sensitivity of the analyte sensor is determined by applying an estimative algorithm that is a function of certain parameters. Also described herein are systems and methods for determining a property of an analyte sensor using a stimulus signal. The sensor property can be used to compensate sensor data for sensitivity drift, or determine another property associated with the sensor, such as temperature, sensor membrane damage, moisture ingress in sensor electronics, and scaling factors. | 1. A method for identifying failure of an analyte sensor using a sensor system electronically coupled to the analyte sensor, the method comprising: generating a data stream indicative of an analyte concentration of a host by applying a bias voltage to a first analyte sensor at least partially positioned within tissue of the host;applying a stimulus signal to the first analyte sensor, wherein the stimulus signal is a time-varying signal;measuring a signal response to the stimulus signal generated by the first analyte sensor;determining, by the sensor system, that the signal response satisfies one or more predetermined sensor fault criteria associated with membrane damage; andinitiating an error routine via the sensor system in response to determining that the signal response satisfies the one or more predetermined sensor fault criteria. 2. The method of claim 1, wherein the error routine includes triggering one or more of an audible alarm and a visual alarm on a display screen. 3. The method of claim 1, wherein the stimulus signal has a predetermined frequency. 4. The method of claim 1, wherein the stimulus signal comprises a spectrum of frequencies. 5. The method of claim 1, further comprising calculating an impedance based on the signal response. 6. The method of claim 5, wherein the calculated impedance comprises a magnitude value and a phase value, and wherein the one or more predetermined criteria comprises a predefined impedance magnitude threshold and a predefined phase threshold. 7. The method of claim 5, wherein the calculated impedance is a complex impedance value. 8. The method of claim 1, wherein the error routine includes suspending display of sensor data generated from the data stream. 9. The method of claim 1, further comprising measuring a signal generated by a second analyte sensor, the second analyte sensor transcutaneously positioned within the host at the same time as the first analyte sensor, wherein the analyzing further includes correlating the signal generated by the second analyte sensor to the signal response and determining if the one or more predetermined criteria are satisfied based on the correlation. 10. A system for monitoring an analyte concentration of a host using an analyte sensor, the system configured to also identify failure of the analyte sensor, the system comprising: an analyte sensor configured for transcutaneous implantation into a host; andsensor electronics coupled to the analyte sensor, the sensor electronics configured to: generate a data stream indicative of an analyte concentration of the host by applying a bias voltage to the analyte sensor,apply a stimulus signal to the analyte sensor, wherein the stimulus signal is a time-varying signal,measure a signal response to the stimulus signal generated by the first analyte sensor,determine that the signal response satisfies one or more predetermined sensor fault criteria associated with membrane damage, andinitiate an error routine in response ... | DexCom, Inc. | A61B005/05 A61B005/1495 G01N027/327 A61B005/145 A61B005/1486 A61B005/00 G01D018/00 G01N027/02 G01N033/49 A61B005/1473
US09792682 | B2 | 2017-10-17 | Training system for detection and classification of artificial objects in X-ray images | 2015-09-07 | The present invention involves a software based system and method which provides a series of training images with or without retained surgical items in scans that are used to train human physicians to detect such items at varying levels of difficulty. | 1. A server for training a user on detection of Retained Foreign Objects having radiographic markers adapted to change shape during surgery (RFOs) in images, said server comprising: an image database including a plurality of images of bodies, a first portion of the plurality of images having RFOs depicted, a second portion of the plurality of images lacking a depicted RFO;an interaction module for sending a sub-set of that plurality of images to a user device, the interaction module also adapted to receive user communication relating to the presence or absence of an RFO in a particular image; anda recording module for recording the correspondence between user communication received relating to each of the first portion and second portion of the images. 2. A method of using a computer to synthesize the presence of a Retained Foreign Object having radiographic markers adapted to change shape during surgery (RFO) in a body image, said method comprising the steps of: obtaining a first image of a body without any RFOs;obtaining a second image of a body having an RFO;extracting features of the RFO from the second image; andsuperimposing the extracted features with the first image to create said synthesized image. 3. A method of training human observers in visual detection of predetermined objects within areas of interest around a surgical field in a scan image, said method comprising the steps of: computing the contrast difference between areas of interest relating to the predetermined objects and their neighborhood using areas of interest in scan images;capturing scan images of foreign objects of interest with contrast with a known level of difficulty;segmenting the scan images of areas of interest to form a first binary mask; andtransforming the first mask and the areas of interest in the image according to the random rotation (R) and scale (S) parameters; andsmoothing the mask M1 using a Gaussian smoothing filter to produce a second binary mask so that the transition between areas of interest and other areas is smooth. 4. The method of claim 3 further including the step of detecting uniform background regions and exclude the uniform background regions from being candidates for implantation. 5. The method of claim 3 wherein the step of computing the contrast difference includes setting the mean and standard deviation of the normal distribution of the areas of interest intensity contrast. 6. The method of claim 3 wherein the step of segmenting includes having the first mask being composed of the set of pixels belonging to the area of interest. 7. The method of claim 6 wherein the segmentation involves marking a rectangular region where the areas of interest reside and applying an adaptive thresholding algorithm. 8. The method of claim 3 wherein the step of transforming includes randomly selecting a desired position (P), rotation (R), scale (S), thickness (T), and intensity contrast (C) of the area of interest. 9. The method of claim 8 wherein ... | RaPID Platforms LLC | G06K009/00 G06T007/00 G06F019/00 A61B008/00 A61B006/02 A61B006/03 A61B005/055
US09819830 | B2 | 2017-11-14 | Image forming apparatus, image forming system, and method for forming images including external devices | 2015-06-26 | An image forming apparatus (100) includes a biological information acquiring unit (402) that, upon accepting a print job transmitted by a user through a preset terminal device (103), identifies the terminal device (103) including a biological information detecting unit based on the terminal device's (103) identification information assigned to the print job, and acquires the user's biological information detected by the biological information detecting unit through a biological information receiving unit, a biological information determining unit (403) that determines whether or not the acquired biological information corresponds to reference biological information indicating that the user is in a hurry, and a calibration control unit (404) that, if the determination result shows that the acquired biological information corresponds to the reference biological information, disables regularly scheduled calibration at a time of, or during execution of the accepted print job. | 1. An image forming apparatus comprising: a central processing unit for controlling the image forming apparatus, whereinthe central processing unit executes a biological information acquiring unit that, upon accepting a print job transmitted by a user through a preset terminal device, identifies the terminal device including a biological information detecting unit based on the identification information for terminal device assigned to the print job, and acquires the user's biological information detected by the biological information detecting unit through a biological information receiving unit,wherein the biological information receiving unit acquire machine IDs of terminal devices within the communicable range of the biological information receiving unit for comparing the acquired machine IDs of terminal devices with the assigned print job;a biological information determining unit that determines whether or not the acquired biological information corresponds to reference biological information indicating that the user is in a hurry; anda calibration control unit that, if the determination result shows that the acquired biological information corresponds to the reference biological information, disables regularly scheduled calibration at a time of executing the accepted print job, or during execution of the accepted print job,wherein the biological information includes the user's heart rate, temperature, perspiration level. 2. An image forming system comprising an image forming apparatus, a first terminal device capable of communicating with the image forming apparatus, and a second terminal device capable of communicating with the first terminal device, wherein the second terminal device includes a biological information detecting unit that detects a user's biological information,the biological information includes the user's heart rate, temperature, perspiration level, and the biological information detecting unit monitors the biological information,the first terminal device includes a biological information receiving unit that receives the user's biological information detected by the biological information detecting unit, and a biological information acquiring unit that acquires the biological information received by the biological information receiving unit,the image forming apparatus or the first terminal device includes a biological information determining unit that determines whether or not the acquired biological information corresponds to reference biological information indicating that the user is in a hurry,upon accepting a print job from the second terminal device, the first terminal device identifies the second terminal device based on the second identification information for terminal device assigned to the print job, and acquires the biological information received by the biological information receiving unit of the second terminal device through the biological information acquiring unit,the biological information determining u... | KYOCERA DOCUMENT SOLUTIONS INC. | G06F003/12 H04N001/00 H04N001/44 A61B005/01 A61B005/0245 A61B005/16 B41J002/525 B41J029/00 B41J029/38 G03G021/00 G06F003/01 A61B005/0205 A61B005/117 A61B005/00
US09814788 | B2 | 2017-11-14 | Renal function analysis method and apparatus | 2013-10-24 | A method for measuring a glomerular filtration rate in a mammalian kidney comprises a source of reporter and marker fluorescent molecules. The fluorescent molecules are introduced into the blood stream of a mammalian subject. Over a period of time, a measurement of the intensities of the reporter and marker fluorescent molecules is taken. A ratio is calculated to determine the health of the subject's kidney. This method measures volume of plasma distribution based on a fluorescence of a marker molecule relative to a fluorescence of a reporter molecule. | 1. A composition for introduction into a mammalian subject's vascular system to analyze an organ function, the composition comprising: an amino fluorescein dextran reporter molecule or a fluorescein isothiocyanate-inulin reporter molecule; anda sulforhodamine 101 dextran marker molecule wherein said reporter molecule and said marker molecule share a molecular property, said reporter molecule molecular property having a first quality, and said marker molecule molecular property having a second quality distinguishable from said first quality; wherein said molecular property is molecular weight and said distinguishable quality is that said reporter molecule has a first molecular weight which is less than a second molecular weight of said marker molecule. 2. The composition of claim 1 wherein said reporter molecule has a first fluorescent characteristic, and said marker molecule has a second fluorescent characteristic, and wherein said first fluorescent characteristic is a first fluorescence excitation wavelength and a first fluorescence emission wavelength said second fluorescent characteristic is a second fluorescence excitation wavelength and a second fluorescence emission wavelength, said first and second fluorescence excitation wavelengths and said first and second fluorescence emission wavelengths being different. 3. The composition of claim 1 wherein said first molecular weight is chosen from a group of ranges consisting of 3 kD to 70 kD, 10 kD to 70 kD, and 20 kD to 70 kD. 4. The composition of claim 1 wherein said first molecular weight is between 3 kD and 70 kD. 5. The composition of claim 1 wherein said first molecular weight is between 3 kD and 20 kD. 6. The composition of claim 1 wherein said first molecular weight is about 5 kD. 7. The composition of claim 1 wherein said reporter molecule is amino fluorescein dextran. 8. The composition of claim 1 wherein said reporter molecule is a fluorescein isothiocyanate-inulin. 9. The composition of claim 1 wherein said marker molecule has a glomerular sieving coefficient of about 0. | PHARMACOPHOTONICS, INC. | A61B010/00 A61B005/00 A61B008/00 A61K049/00 A61B005/20
US09814404 | B2 | 2017-11-14 | Radiolucent ECG electrode system | 2016-07-25 | Disclosed is an ECG electrode lead system suitable for use during imaging procedures such as, without limitation, CT scans or MRI and methods of use. In embodiments, the system includes an ECG electrode connector formed from radiolucent materials to enhance performance during imaging procedures by reducing or eliminating shadows on imaging media. In some embodiments, the disclosed connector includes a housing having an opening configured to operably receive an electrode post of an ECG electrode pad, an electrode member having a generally semicircular contact face disposed along at least a part of the perimeter of the opening, and an engagement member having an actuation surface and an engaging face pivotable about a pivot to enable engaging face to move from a first position whereby the engaging face is closer to the contact face and a second position whereby the engaging face is further from the contact face. | 1. An electrocardiogram (ECG) lead system, comprising: a lead extension assembly that includes a first lead wire configured to be electrically coupled to an ECG monitor;a lead connector assembly that includes a second lead wire configured to be electrically coupled to the first lead wire; andan ECG connector assembly that comprises: a connector electrode that is disposed along at least part of a perimeter of an aperture of a housing and that is electrically coupled to the second lead wire; andan engagement member that comprises a resilient radiused member and that is coupled to a pivot within a cavity of the housing and that is arranged to rotate about the pivot under force imparted by the resilient radiused member, the engagement member being configured to electrically couple the connector electrode with a patient electrode when the patient electrode is inserted within the aperture. 2. The ECG lead system of claim 1, wherein the first lead wire comprises tinned copper wire. 3. The ECG lead system of claim 1, wherein the second lead wire comprises radiolucent conductor wire. 4. The ECG lead system of claim 1, wherein each one of the first lead wire and the second lead wire is formed from a conductive material selected from: conductive polymer; conductive carbon; conductive ink. 5. The ECG lead system of claim 1, wherein the connector electrode includes a contact face that extends into the aperture of the housing. 6. The ECG lead system of claim 1, wherein the second lead wire comprises at least one carbon fiber. 7. The ECG lead system of claim 1, wherein the connector electrode is formed from a radiolucent conductive material. 8. The ECG lead system of claim 1, further comprising a first lead set connector, a second lead set connector and a device connector; wherein a first end of the first lead wire terminates at the first lead set connector and a second end of the first lead wire terminates at the device connector, and a first end of the second lead wire terminates at the second lead set connector, and wherein the first lead set connector and the second lead set connector are configured to mate with each other and the device connector is configured to mate with the ECG monitor or an adaptor configured to mate with the ECG monitor. 9. The ECG lead system of claim 1, wherein the lead connector assembly is configured as a ribbon-cable comprising a plurality of second lead wires, each one of the second lead wires being separable from an adjacent second lead wire along a particular length of the ribbon cable. 10. An electrocardiogram (ECG) connector assembly, comprising: a housing defining an aperture that extends therethrough to a cavity of the housing;a conductor that terminates at an electrode disposed along at least part of a perimeter of the aperture and that is configured to electrically couple the electrode to an ECG monitor; andan engagement member that comprises a resilient radiused member and that is positioned within the cavity a... | Covidien LP | A61B005/0416
US09788778 | B2 | 2017-10-17 | Measuring psychological stress from cardiovascular and activity signals | 2016-04-11 | A method and system for measuring psychological stress disclosed. In a first aspect, the method comprises determining R-R intervals from an electrocardiogram (ECG) to calculate a standard deviation of the R-R intervals (SDNN) and determining a stress feature (SF) using the SDNN. In response to reaching a threshold, the method includes performing adaptation to update a probability mass function (PMF). The method includes determining a stress level (SL) using the SF and the updated PMF to continuously measure the psychological stress. In a second aspect, the system comprises a wireless sensor device coupled to a user via at least one electrode, wherein the wireless sensor device includes a processor and a memory device coupled to the processor, wherein the memory device stores an application which, when executed by the processor, causes the processor to carry out the steps of the method. | 1. A method for measuring psychological stress, the method comprising: determining R-R intervals from an electrocardiogram (ECG) to calculate a standard deviation of the R-R intervals (SDNN), wherein the determining the R-R intervals from the ECG to calculate the SDNN further comprises: detecting R peaks from a measured ECG within a predetermined time period, andcalculating the R-R intervals using the detected R peaks;determining a stress feature (SF) using the SDNN, wherein the determining the stress feature (SF) using the SDNN further comprises: calculating a mean heart rate (HR) from the ECG within the predetermined time period, andcalculating the SF utilizing an algorithm SF=HR+α*SDNN, wherein the α includes a predetermined negative variable that allows for combining the HR and the SDNN;in response to reaching a threshold, performing adaptation to update a probability mass function (PMF); anddetermining a stress level (SL) using the SF and the updated PMF to continuously measure the psychological stress. 2. The method of claim 1, further comprising: determining a posture state, wherein the psychological stress is not measured if the posture state is active; anddisplaying the determined SL to a user or another device. 3. The method of claim 2, wherein the posture state includes any of active, sitting, and standing. 4. The method of claim 1, wherein performing adaptation to update the PMF further comprises: grouping data into a predetermined distribution;calibrating the predetermined distribution according to a detected resting heart rate; andadjusting the predetermined distribution according to additional samples received. 5. The method of claim 4, wherein adjusting the predetermined distribution according to additional samples received further comprises: in response to data arriving, multiplying all bins of the predetermined distribution by 1−ε; andadding ε to a bin corresponding to the data, wherein the ε includes a forgetting parameter for how much the PMF changes with each adaptation run. 6. The method of claim 4, wherein determining the stress level (SL) using the SF and the PMF further comprises: adding all bins below a bin corresponding to the SF; andcomputing the SL utilizing an algorithm that includes a probability mass function for a given posture (PMFposture), the SF, and the added bins. 7. The method of claim 6, further comprising: adding a fraction of a current bin to improve granularity. 8. The method of claim 1, further comprising: tracking both a mean and a standard deviation of a probability distribution as the probability distribution adapts over time; andcombining the mean and the standard deviation to measure long-term stress. 9. The method of claim 1, wherein the application further causes the processor to: track both a mean and a standard deviation of a probability distribution as the probability distribution adapts over time; andcombine the mean and the standard deviation to measure long-term stress. 10. A syste... | Vital Connect, Inc. | A61B005/16 A61B005/00 A61B005/0468 A61B005/11 A61B005/024 A61B005/0245 A61B005/04 A61B005/0456 A61B005/044
US09790977 | B2 | 2017-10-17 | Pen-type injection device and electronic clip-on module therefor | 2013-02-08 | The present invention relates to a supplementary device for a manually operable injection device. The device has a body and a mating unit configured to releasably mount the body to the injection device in a specific position relative to an outside surface of the injection device. | 1. A supplementary device for attachment to a manually operable injection device, the supplementary device comprising a body having an injection device receiving channel; anda mating unit configured to releasably mount the body to the injection device in a specific position relative to an outside surface of the injection device,wherein the mating unit comprises a collar extending from the body,wherein the collar is configured to receive the injection device so that the injection device extends through the collar, andwherein the collar comprises a closure which is rotatable towards the body about an axis of rotation that is transverse to a longitudinal axis of the injection device receiving channel to secure the body to the injection device. 2. A supplementary device according to claim 1, wherein the mating unit further comprises an engaging unit configured to engage the supplementary device with the injection device. 3. A supplementary device according to claim 2, wherein the engaging unit comprises a support member having an engaging element, the engaging element being configured to engage in an indent on the injection device when the body is disposed in a specific position relative to an outside surface of the injection device. 4. A supplementary device according to claim 3, wherein the support member is a first support member and the engaging unit further comprises a second support member, engaging elements being disposed at the free end of each support member, and the free ends of the support members being biased towards each other. 5. A supplementary device according to claim 4, wherein each support member is pivotally mounted in the body about a mid-section so that the free ends of the support members may be biased away from each other to disengage the engaging elements from the indents. 6. A supplementary device according to claim 1, wherein the closure is configured to act on the support member to urge the engaging element into the indent when the closure is in its secured position so that the body is secured to the injection device. 7. A supplementary device according to claim 6, wherein the mating unit further comprises a catch arrangement configured to releasably retain the closure in its secured position. 8. A supplementary device according to claim 1, wherein the mating unit further comprises a locating recess in the body, the locating recess being configured to mate with a locating rib on the injection device. 9. A supplementary device according to claim 1, further comprising an optical reading arrangement and wherein the optical reading arrangement is directed at a display of the injection device when the body is mounted to the injection device in a specific position relative to an outside surface of the injection device. 10. A supplementary device according to claim 1, together with an injection device. 11. A supplementary device as claimed in claim 3, wherein the support member is disposed behind a wing which is formed ... | Sanofi-Aventis Deutschland GmbH | A61M005/24 F16B021/08 A61B005/145 A61M005/31 G06F019/00 A61B005/00
US09814448 | B2 | 2017-11-14 | Three-dimensional optical imaging and therapy of prostate cancer | 2013-05-20 | A system, method, and storage device storing computer executable instructions for use in tissue analysis and therapy. An optical probe array has at least two or more optical probes for illuminating tissue and generates light fluorescence and/or diffuse reflectance signals corresponding to the illuminated tissue. One or more ultrasound, CT and MRI imaging guidance systems identify the position of the optical probes relative to the tissue. The imaging guidance system can also be a fusion of an MRI or CT image provided by the MRI or CT imaging guidance system and an ultrasound image provided by the ultrasound imaging guidance system. An imaging system generates a three-dimensional image of the tissue based on the generated light signals and the identified position of the optical probes. | 1. A system evaluating tissue comprising: an optical probe array system having at least one or more optical probes configured to insert into an interior of the tissue, said optical probe comprising: at least one light source configured to illuminate a portion of the interior tissue adjacent to, contiguous to and/or in contact with the optical probe at a location within the tissue and configured to generate from the portion of the interior tissue light signals fluoresced or elastically scattered from the illuminated interior tissue, said fluoresced or scattered light signals comprising Auto-Fluorescence Spectra (AFS) or Elastic Scattering Spectra (ESS);an optical sensor configured to capture the AFS or ESS signals indicative of the AFS or ESS of the illuminated interior tissue; anda tissue characterization imaging system comprising a spectroscope configured to evaluate the AFS or ESS signals captured by the optical sensor based on a fluorescence or elastic scattering of the captured signals, said tissue characterization imaging system executing an algorithm analyzing selected data of the AFS or ESS signals to generate a diagnostic classification indicative of tissue condition of the illuminated interior tissue based only on the selected data, said tissue characterization imaging system configured to provide the classification by correlating the location of the classified interior tissue to a three-dimensional image of the tissue, and generating a three-dimensional optical spectroscopic image showing the location and classification of the illuminated interior tissue in relation to the three-dimensional image of the tissue. 2. The system of claim 1 further comprising an imaging guidance system for identifying a position of the one or more optical probes relative to the tissue, wherein the imaging guidance system comprises a three-dimensional imaging system for generating the three-dimensional image of the tissue. 3. The system of claim 2, wherein the imaging guidance system comprises an ultrasound imaging system. 4. The system of claim 2, wherein the imaging guidance system comprises an MRI or CT imaging system. 5. The system of claim 1 further comprising a treatment device for treating the illuminated interior tissue, wherein the treatment device for treating the illuminated interior tissue comprises at least one of the following: a Cryotherapy device;a Photodynamic Therapy device;a Brachytherapy device;a high-intensity focused ultrasound (HIFU) device;a tissue ablation device;a Laser ablation device;a RF ablation device;a Vapor ablation device;a Local drug delivery device. 6. The system of claim 1 further comprising a grid or template for aligning the one or more optical probes into a predetermined orientation relative to each other and relative to the tissue. 7. The system of claim 1 wherein the optical probe comprises: a shaft;at least one pair of optical devices supported by the shaft, each pair of optical devices comprising: the at least... | Precision Biopsy, Inc. | A61B005/00 A61B005/055 A61B006/03 A61B008/08 A61B010/04 A61B005/06 A61B001/00 A61B001/06 A61B001/07 A61B008/12 A61B008/00 A61B001/04 A61B006/12 A61B010/02 A61N005/06 A61N005/10 A61B018/04 A61B034/20 A61B017/34 A61B001/307 A61N007/02 A61B018/14 A61B018/22 A61B018/02 A61B090/00 A61B090/30
US09801550 | B2 | 2017-10-31 | Method and system for measuring anatomical dimensions from a digital photograph on a mobile device | 2013-09-26 | A mobile, hand-held communication device with a digital touch screen display and a camera for acquiring an image of the human body is programmed to function as a digital anthropometer. The user digitizes anatomical landmarks on the displayed image to quickly obtain linear measurements which are used with a known morphological relationship to make an anatomical prediction for clothing measurement, body composition and postural displacement with accuracy and without external equipment. | 1. A method of deriving an anatomical prediction using a known morphological relationship and a programmed apparatus including a digital touch screen display and a camera configured to acquire images of a person on the digital touch screen display, the method comprising: acquiring a plurality of images of a person on the digital touch screen display;digitizing points on a plurality of anatomical landmarks on the displayed images;determining linear anatomical dimensions of at least a portion of a body of a person in the displayed images using the digitized points and a scale factor for the displayed images; andmaking an anatomical circumferential measurement prediction of the person based on the determined linear anatomical dimensions and a known morphological relationship. 2. The method of claim 1, wherein the points on the plurality of anatomical landmarks are digitized by touching the digital touch screen display of the images at respective anatomical landmarks. 3. The method of claim 1, wherein acquiring the plurality of images of the person includes acquiring at least two different images corresponding to different views of the person on the digital display, and digitizing points on the plurality of anatomical landmarks includes digitizing points on anatomical landmarks on each displayed image and determining linear anatomical dimensions of the at least the portion of the body of the person using the digitized points and a scale factor for each displayed image for making the anatomical circumferential measurement prediction. 4. The method of claim 3, wherein the different views acquired include a front view and a side view of the person. 5. The method of claim 1, wherein the scale factor is a ratio of pixel to distance. 6. A method for measuring the dimensions of a body of a person, the method comprising: providing a digital anthropometer on a mobile device;digitizing anatomical landmarks provided on a plurality of images of at least a portion of a body of a person displayed on the device with established calibration methods;measuring dimensions of the at least the portion of the body of the person based on the digitized anatomical landmarks across the plurality of images; andmaking an anatomical circumferential measurement prediction of the person based on the measured dimensions and a known morphological relationship. 7. A system for measuring the dimensions of a body of a person, the system comprising: a programmed device comprising a digital touch screen display having an array of pixels and a camera configured to acquire a plurality of images of at least a portion of a body of a person on the digital touch screen display;a digital anthropometer configured to digitize anatomical landmarks on the images of the at least the portion of the body of the person displayed on the digital touch screen display in order to measure dimensions of the at least the portion of the body of the person; anda controller configured to make an anatomical... | - | A61B006/00 A61B005/00 A61B005/107 A61B005/11
US09815363 | B2 | 2017-11-14 | Operation of an electrically motorized vehicle | 2015-04-06 | A system, method, and device for operations of an electrically motorized vehicle. The vehicle can utilize an electrically motorized wheel to convert a non-motorized wheeled vehicle to an electrically motorized wheeled vehicle. A method of controlling the device of the electrically motorized wheel includes detecting a temperature of an electric motor of the electrically motorized wheel and controlling operation of the electric motor to maintain the detected temperature within a desired range. | 1. A method of controlling a motorized human-powered vehicle, the method comprising: detecting a temperature of a drive motor operable to propel the motorized human-powered vehicle, a battery system, or an electronic board within the motorized human-powered vehicle; andcontrolling, using a control unit of the motorized human-powered vehicle, a power output from the battery system to the drive motor to maintain the detected temperature below a respective predetermined maximum operational temperature limit for the drive motor, the battery system, or the electronic board within the motorized human-powered vehicle. 2. The method as recited in claim 1, wherein the motorized human-powered vehicle is a bicycle and controlling the power output from the battery system to the drive motor includes controlling assistance to a pedaling input transmitted to the motorized human-powered vehicle. 3. The method as recited in claim 1, wherein the motorized human-powered vehicle is a bicycle and controlling the power output from the battery system to the drive motor includes controlling resistance to a pedaling input transmitted to the motorized human-powered vehicle. 4. The method as recited in claim 1, wherein the motorized human-powered vehicle is a bicycle and controlling the power output from the battery system to the drive motor includes at least one of reducing and stopping assistance to a pedaling input transmitted to the motorized human-powered vehicle in response to the detected temperature of the drive motor, the battery system, or the electronic board being above the respective predetermined maximum operational temperature limit. 5. The method as recited in claim 1, wherein controlling the power output from the battery system to the drive motor includes at least one of reducing and stopping assistance or resistance to a user input transmitted to the motorized human-powered vehicle in response to the detected temperature of the drive motor, the battery system, or the electronic board being above the respective predetermined maximum operational temperature limit. 6. The method as recited in claim 1, wherein controlling the power output from the battery system to the drive motor includes controlling the drive motor to provide assistance to a user input transmitted to the motorized human-powered vehicle. 7. The method as recited in claim 1, wherein controlling the power output from the battery system to the drive motor includes controlling the drive motor to provide resistance to a user input transmitted to the motorized human-powered vehicle. 8. A system for controlling a motorized human-powered vehicle, the system comprising: a sensor system operable to detect a temperature of a drive motor operable to propel the motorized human-powered vehicle, a battery system, or an electronic board within the motorized human-powered vehicle; anda control system in communication with the sensor system, the control system operable to control a power output from th... | Superpedestrian, Inc. | B60L009/00 B60K007/00 B60L015/20 B60L003/00 B60L003/12 B60L007/12 B60L011/00 B60L011/18 A61G005/04 B62B003/00 B62B005/00 G06F009/445 G07C005/00 G07C005/08 H04L029/08 H04W004/00 B60Q005/00 B60Q009/00 A61B005/00 A61B005/11 B60W050/08 A63B024/00 B62M006/45 B62M025/08 H02P003/06 B60L007/00 H02P006/08 B62M006/80 A63B021/005 A63B021/22 A63B022/06 B60R016/02 G05D001/02 G07C005/02 G08G001/127 B60R025/04 B60R025/10 B60R025/20 E05B049/00 E05B081/54 G07C009/00 G08G001/01 G08G001/052 G08G001/123 G08G001/00 B61L025/02 G08G001/13 H04M001/725 H02P029/20 B60C005/00 B60C009/00 G01C021/36 B60W050/00 E05B047/00 G08G001/16
US09797967 | B2 | 2017-10-24 | Active coil to shift a volume of uniform magnetic field | 2015-07-15 | Gradient coils are operated to acquire magnetic resonance (MR) signals encoding a first MRI image over a first region inside a main magnet of the MRI system in which at least a portion of a subject is placed, the first region being located within a volume of uniform magnetic field with inhomogeneity below a defined threshold. An active coil is energized to shift the volume of uniform magnetic field such that a second region inside the main magnet of the MRI system is located within the shifted volume of uniform magnetic field, at least a portion of the second region being located outside of the volume of uniform magnetic field before the volume of uniform magnetic field has been shifted. The gradient coil is operated to acquire MR signals encoding a second MRI image over the second region. | 1. A method for operating a magnetic resonance imaging (MRI) system, the method comprising: operating gradient coils to acquire magnetic resonance (MR) signals encoding a first MRI image over a first region inside a main magnet of the MRI system in which at least a portion of a subject is placed, the first region being located within a volume of magnetic field where a field inhomogeneity is below a defined threshold;energizing an active coil to shift the volume of magnetic field such that a second region inside the main magnet of the MRI system is located within the shifted volume of magnetic field, at least a portion of the second region being located outside of the volume of magnetic field before the volume of magnetic field has been shifted and at least a portion of the first region being located outside of the shifted volume of magnetic field after the volume of magnetic field has been shifted; andoperating the gradient coils to acquire MR signals encoding a second MRI image over the second region. 2. The method of claim 1, wherein energizing the active coil comprises: energizing the active coil while operating the gradient coils to acquire the second MRI image. 3. The method of claim 1, wherein energizing the active coil comprises: energizing the active coil prior to operating the gradient coils to acquire the second MRI image. 4. The method of claim 1, wherein energizing the active coil comprises: driving the active coil with at least 2 amperes of electrical current. 5. The method of claim 1, wherein energizing the active coil comprises: cooling the active coil by running liquid coolant to the active coil. 6. The method of claim 1, further comprising: operating a shim coil to compensate susceptibility differences inside the portion of the subject located within the shifted volume of magnetic field with inhomogeneity below the defined threshold. 7. The method of claim 6, wherein operating the shim coil further comprises: operating the shim coil while operating the gradient coils to acquire MR signals encoding the second MRI image. 8. A magnetic resonance imaging (MRI) system, comprising: a housing having a bore in which at least a portion of a subject to be imaged is placed;a main magnet accommodated by the housing and configured to generate a volume of magnetic field with inhomogeneity below a defined threshold to form an MRI image over a region located within the volume of magnetic field;an active coil that when energized causes the volume of magnetic field to shift in location and transform in shape;pulse generating coils to generate and apply radio frequency (RF) pulses in sequence to scan the portion of the subject;gradient coils to provide perturbations to the volume of magnetic field such that MRI signals encoding an MRI image are acquired in response to the applied RF pulses; anda control unit coupled to the MRI system and configured to: operate the gradient coils to acquire MR signals encoding a first MRI image over a first r... | Synaptive Medical (Barbados) Inc. | G01V003/00 G01R033/3875 A61B005/055 G01R033/381
US09795306 | B2 | 2017-10-24 | Method of estimating blood pressure based on image | 2016-07-07 | A method of estimating a blood pressure based on an image is provided. The method includes obtaining, by camera, an image including a skin of a user, determining, by a computer device, a skin region, in which at least one portion of the skin is displayed, from the image, and storing a mean value of color data of a designated color model for each of two target regions which have different positions in the skin region from each other, estimating, by the computer device, a pulse-wave transit time (PTT) based on a pulse wave signal determined based on a change in the mean value of the color data for each of the two target regions, and estimating, by the computer device, the blood pressure using the PTT. | 1. A method of estimating a blood pressure based on a video of a user, comprising: obtaining, by camera, an image including a skin of the user;determining, by a computer device, a skin region, in which at least one portion of the skin is displayed, from the image, and storing a specific value generated from color data of a designated color model for each of two different portions of the skin;estimating, by the computer device, a pulse wave signal based on a change in specific value of color data for each of the two different portions of the skin;estimating, by the computer device, a pulse-wave transit time (PTT) based on a time difference between local extrema of the pulse wave signals from the two different portions of skin; andestimating, by the computer device, the blood pressure using the PTT. 2. The method of claim 1, wherein the storing of the specific value comprises: dividing the skin region in the image into two and storing the specific value of the color data for the two different portions of the skin; orextracting skin regions from images obtained by two cameras respectively and storing the specific value of the color data by determining the extracted two skin regions as the two different portions of the skin. 3. The method of claim 1, wherein the storing of the specific value comprises: setting, by the computer device, the two different portions of the skin by dividing the image; andgenerating a change in the specific value of the color data for each of the two different portions of the skin. 4. The method of claim 1, wherein: the obtaining of the image comprises obtaining, by the camera, an image including a facial region of the user; andthe storing of the specific value comprises extracting, by the computer device, the facial region from the image and classifying the facial region as two target different portions to be the two different portions of the skin. 5. The method of claim 1, wherein the storing of the specific value comprises storing, by the computer device, the specific value of the color data of the portions of the skin in units of frames or frames of a reference interval. 6. The method of claim 1, wherein the estimating of the PTT comprises generating, by the computer device, pulse wave signals for the portions of the skin from the change in the specific value of the color data using a band pass filter. 7. The method of claim 1, wherein the estimating of the PTT further comprises determining two associated peak points from the pulse wave signals for the respective target regions and estimates the PTT based on a time difference between the two peak points. 8. The method of claim 1, wherein the estimating the blood pressure comprises applying the PTT to an equation which is provided in advance through a multiple regression analysis and estimates the blood pressure. 9. The method of claim 1, further comprising before the estimating of the blood pressure, receiving age, height, and weight of the user, and applying th... | Research & Business Foundation Sungkyunkwan University | A61B005/02 A61B005/021 A61B005/00 A61B005/103
US09808192 | B2 | 2017-11-07 | Biological fluid sampling transfer device and biological fluid separation and testing system | 2016-12-14 | A blood separation and testing system for a blood sample is disclosed. The blood separation and testing system includes a blood sampling transfer device adapted to receive a blood sample, a blood separation device, and a blood testing device. The blood separation device is adapted to receive a portion of the blood sampling transfer device such that with the blood sampling transfer device received within the blood separation device and a rotational force applied to the blood sampling transfer device, a plasma portion of the blood sample is separated from a cellular portion of the blood sample. The blood testing device is adapted to receive a portion of the blood sampling transfer device to analyze the plasma portion of the blood sample and obtain test results. | 1. A biological fluid separation and testing system for a multi-component blood sample, the biological fluid separation and testing system comprising: a biological fluid sampling transfer device adapted to receive the blood sample, the biological fluid sampling transfer device comprising: a housing having a resiliently deformable upper portion, a lower portion removably connectable to the upper portion, the lower portion having an inlet port, an outlet port, and a flow channel, the inlet port and the outlet port in fluid communication via the flow channel, the flow channel adapted to receive the blood sample via the inlet port, the upper portion transitionable between an undeformed position and a deformed position; anda puncturing element, a portion of which is disposed within the housing and adapted for movement between a pre-actuated position wherein the puncturing element is retained within the housing and a puncturing position wherein the puncturing element extends through the inlet port of the lower portion of the housing,wherein actuation of the upper portion from the undeformed position to the deformed position moves the puncturing element from the pre-actuated position to the puncturing position, andwherein after the puncturing element moves to the puncturing position, the upper portion returns to the undeformed position and the puncturing element returns to the pre-actuated position; anda blood testing device having a receiving port adapted to receive the outlet port of the lower portion of the housing of the biological fluid sampling transfer device for closed transfer of a portion of the blood sample from the biological fluid sampling transfer device to the blood testing device via the outlet port. 2. The biological fluid separation and testing system of claim 1, wherein the blood testing device comprises a point-of-care testing device. 3. The biological fluid separation and testing system of claim 1, wherein with the biological fluid sampling transfer device received within the receiving port of the blood testing device, actuation of the upper portion from the undeformed position to the deformed position transfers the blood sample from the biological fluid sampling transfer device to the blood testing device via the outlet port. 4. The biological fluid separation and testing system of claim 1, wherein with the blood sample received within the flow channel, the lower portion is removable from the upper portion. 5. The biological fluid separation and testing system of claim 4, further comprising a blood separation device, wherein with the lower portion removed from the upper portion, the lower portion is receivable within the blood separation device. 6. The biological fluid separation and testing system of claim 5, wherein with the lower portion received within the blood separation device, the blood sample contained within the lower portion is separated. 7. The biological fluid separation and testing system of claim 1, wherein a p... | Becton, Dickinson and Company | A61B005/00 A61B005/15 A61B005/151 A61B005/157 A61M001/34 B01L003/00 B04B007/08 G01N001/28 G01N001/34 G01N001/40 G01N033/49
US09808225 | B2 | 2017-11-07 | Retrieval of biological materials from the human uterus, ovary and cervix by suction | 2014-08-28 | The invention provides devices and methods for self-administered noninvasive retrieval of biological materials of the uterus and/or cervix. The cervical device comprises a receptacle with a variable volume receptacle cavity, a controller configured to change the volume of the receptacle cavity, a surface for collection of biological materials, and a flexible pouch for generation of suction to facilitate efficient retrieval of biological materials. The biological materials retrieved include various biomarkers of diseases and disorders of reproduction, are directly isolated from a site of pathology, and are not metabolized or diluted. Information generated by analyzing these biological materials permits early diagnosis and prognosis assessments of disease and disorders of the female reproductive organs, irregularities of pregnancy, anomalies of the fetus in utero, and microbial infections. | 1. A device for retrieving biological materials of the uterus and/or cervix of a subject comprising: a receptacle having an open end configured to be placed over the cervix and a wall defining a variable volume receptacle cavity;wherein the device further comprises a flexible pouch and a controller, wherein expansion or compression of the pouch generates suction inside the cavity; and wherein the controller is configured to change the volume of the receptacle cavity by expanding or compressing the pouch;the controller having a proximal end and a distal end. 2. The device according to claim 1, further comprising: a removable surface for collection of the biological materials, the surface having an anterior wall facing the open end of the receptacle and a posterior wall facing opposite to the open end of the receptacle; andwherein the pouch is hollow inside and having a second anterior wall facing the open end of the receptacle and a second posterior wall facing opposite of the open end of the receptacle wall; wherein the surface is removably fitted within the receptacle cavity. 3. The device according to claim 1, wherein the receptacle further comprises a flexible rim extending from the open end and attaching the receptacle over the cervix. 4. The device according to claim 1, wherein the receptacle is made of silicone or a synthetic polymer. 5. The device according to claim 1, wherein the surface further comprises: a pull tab for detaching the surface from the receptacle cavity; anda platform holding a matrix or mesh of permeable substrate materials. 6. The device according to claim 5, wherein the substrate materials are selected from the group consisting of: cotton mesh, nylon mesh, collagen matrix, hydrogel matrix, synthetic sponge or tissue culture scaffold. 7. The device according to claim 2, wherein the pouch is an inflated pouch. 8. The device according to claim 7, the inflated pouch comprising one or more perforations on the second posterior wall. 9. The device according to claim 8, wherein the controller is configured to expel air contained in the pouch through the perforations of the posterior wall to generate suction. 10. The device according to claim 2, wherein the pouch is a compressed pouch. 11. The device according to claim 10, the compressed pouch comprising one or more pillars separating the second anterior wall and the second posterior wall. 12. The device according to claim 1, wherein the proximal end comprises a handle and the distal end comprises a bead. 13. The device according to claim 4, wherein the silicone is medical grade. 14. The device according to claim 1, further comprising: a removable surface; anda flexible pouch; wherein the surface is removably fitted within the receptacle cavity. 15. The device according to claim 1, wherein the receptacle further comprises a flexible rim extending from the open end. 16. The device according to claim 5, wherein the matrix or mesh contains a solution for irrigation a... | - | A61B010/00 A61B005/00 B65D081/00 B65D033/00 A61F005/00
US09795830 | B2 | 2017-10-24 | Associative object tracking systems and methods | 2016-05-24 | Systems and methods track a first object when continuous tracking information for the first object is not available. The systems and methods detect when the tracking information for the first object is not available. A last time of a last determined location of the first object is determined and a second object closest to the last determined location at the last time is determined. The location of the first object is associated with a location of the second object if tracking information for the first object is not available. | 1. An associative tracking apparatus for tracking a first object using tracking information for the first object and tracking information for a second object, comprising: a tracking reliability monitor for determining when the tracking information for the first object is not reliable;a proximity detector for identifying the second object as closest to the first object when the tracking reliability monitor determines that the tracking information for the first object is not reliable;an associative tracker for associating a location of the first object with a location of the second object when the tracking reliability monitor determines that the tracking information for the first object is not reliable, andan output generator for outputting, when the tracking reliability monitor determines that the tracking information for the first object is not reliable, location information for the second object as location information for the first object based upon the association of the first object with the second object. 2. An apparatus for tracking a first object using tracking information for the first object and tracking information for a second object, comprising: a tracking reliability monitor for determining when the tracking information for the first object is not reliable;a proximity detector for identifying the second object as closest to the first object when the tracking reliability monitor determines that the tracking information for the first object is not reliable, the proximity detector being configured to store a most recently determined location of the first object when the tracking reliability monitor determines that the tracking information for the first object is reliable for use by the proximity detector to identify the second object; andan associative tracker for associating a location of the first object with a location of the second object when the tracking reliability monitor determines that the tracking information for the first object is not reliable. 3. A method for tracking a first object using tracking information for a second object, comprising: sensing proximity of the first object to the second object through sensing, at the second object, presence of a short range radio signal transmitted from the first object;transmitting an indication of the sensed proximity with the tracking information for the second object;detecting, within a tracking apparatus receiving the tracking information, when tracking information for the first object is not available; andassociating the location of the first object with a location determined from the tracking information when the tracking information for the first object is not available. 4. A method for tracking a first object using tracking information for a second object, comprising: sensing proximity of the first object to the second object through detection of a magnetic field generated by the first object at the second object;transmitting an indication of the sensed proximity with the... | ISOLYNX, LLC | A63B024/00 A61B005/145 A61B005/08 A63B071/06 A61B005/024 G06Q010/08 G06Q050/28 G06T007/20 G06K009/00 G01S013/72 G01S013/75 A63B102/22
US09826917 | B1 | 2017-11-28 | Electrode harness and method of taking electrical measurements | 2011-08-08 | The present invention relates to an electrode harness and more particularly to an electrode harness with various features, which enhance the use and performance of the electrode harness. The present invention further relates to a method of taking biopotential measurements. The electrode harness and methods of the present invention allow for use with most applications where biopotential measurements are taken. The electrode harness can be used in ECG (or EKG), EEG, EMG, and other such biopotential measurement applications. Because of the versatility of various embodiments of the present invention, preferably the electrode harness can be adjusted for different applications or for application to various sized and shaped subjects. The electrode harness is further preferably part of a system, which includes either wireless or tethered bridges between the electrode harness and a monitor, and preferably includes various forms of processors for analyzing the biopotential signal. | 1. An electrode harness for physiological monitoring of a subject, the electrode harness comprising at least one dry electrode with adjustable connector;a material operable to mechanically and electrically interconnect at least two electrodes; anda track system attached or integral to the materialwherein the at least one electrode is movable along the track system with the adjustable connector. 2. The electrode harness in claim 1, wherein the material is a multi-layer laminate. 3. The electrode harness in claim 2, wherein the material comprises at least two expandable arms, each of the at least two expandable arms corresponding to the at least two electrodes. 4. The electrode harness in claim 3, wherein the material further comprises a slit or a mechanical weak point for separating one of the electrodes. 5. The electrode harness in claim 1, wherein the material comprises at least two expandable arms, each of the at least two expandable arms corresponding to the at least two electrodes. 6. The electrode harness in claim 1, further comprising at least one additional electrode wherein the material comprises an electrical pathway from each of the at least one and at least one additional electrodes, the electrical pathway being electrically shielded from large defibrillator voltages. 7. An electrode harness for physiological monitoring of a subject, the electrode harness comprising: at least two dry electrodes;a material operable to mechanically and electrically interconnect the at least two electrodes comprising separate electrical pathways to each of the at least two dry electrodes; anda lead connector to mechanically connect the material and electrically connect the separate electrical pathways to each of the at least two dry electrodes of the electrode harness to a tether or a wireless transmitter,wherein at least one of the at least two dry electrodes are each connected to the material with an adjustable connector. 8. The electrode harness in claim 7, wherein the material is a multi-layer laminate. 9. The electrode harness in claim 8, wherein the material further comprises a slit or a mechanical weak point for separating one of the electrodes. 10. The electrode harness in claim 8, wherein the material comprises at least two expandable arms, each of the at least two expandable arms corresponding to the at least two electrodes. 11. The electrode harness in claim 7, wherein the material comprises at least two expandable arms, each of the at least two expandable arms corresponding to the at least two electrodes. 12. The electrode harness in claim 7, wherein the material comprises an electrical pathway from the at least two dry electrodes, the electrical pathway being electrically shielded from large defibrillator voltages. 13. The electrode harness in claim 7, further comprising a track system attached or integral to the material wherein at least one electrode is movable along the track system with the adjustable connector. 14. An electr... | Orbital Research Inc. | A61B005/0408
US09820677 | B2 | 2017-11-21 | Cointegration filter for a catheter navigation system | 2013-01-03 | A system and method for determining a position of a medical device within a body are provided. The system includes an electronic control unit that receives position signals from position sensors of a first type and a second type disposed on the device and applies a filter to each of the position signals to obtain filtered estimated positions for each sensor. The unit computes a spline connecting the position sensors of the first type responsive to the filtered estimated positions for the sensors and estimates a spline position for the sensor of the second type along the spline. The unit generates maps between the spline position and filtered and unfiltered estimated positions for the sensor of the second type and determines actual positions for the sensors of the first type responsive to the filtered estimated position for the sensors and a composite map of the two maps. | 1. A system for determining a position of a medical device within a body, comprising: an electronic control unit configured to: receive a position signal from each of a plurality of position sensors of a first type disposed on the medical device and from at least one position sensor of a second type disposed on the medical device;apply a filter to each of said position signals to obtain filtered estimated positions for each of said plurality of position sensors of said first type and said at least one position sensor of said second type;compute a spline connecting said plurality of position sensors of said first type responsive to said filtered estimated positions for said plurality of positions sensors of said first type;estimate a spline position for said at least one position sensor of said second type along said spline;generate a first map between said spline position and said filtered estimated position of said at least one position sensor of said second type;generate a second map between said spline position and an unfiltered estimated position of said at least one position sensor of said second type;generate a composite map using said first and second maps;determine an actual position for one of said plurality of position sensors of said first type using said composite map and said filtered estimated position for said one of said plurality of position sensors of said first type; anda display communicatively coupled to the electronic control unit and configured and arranged to display the composite map overlaid with the actual position of said one of said plurality of position sensors of said first type. 2. The system of claim 1 wherein said first type of position sensor comprises an electrical position sensor. 3. The system of claim 2 wherein said second type of position sensor comprises a magnetic position sensor. 4. The system of claim 1 wherein said second type of position sensor comprises a magnetic position sensor. 5. The system of claim 1 wherein said first and second maps comprise vector offsets between said spline position and corresponding said filtered estimated position and said unfiltered estimated position of said at least one position sensor of said second type. 6. The system of claim 1 wherein said first and second maps comprise transformation matrices between said spline position and corresponding said filtered estimated position and said unfiltered estimated position of said at least one position sensor of said second type. 7. The system of claim 1 wherein said at least one position sensor of said second type comprises a plurality of position sensors of said second type and said first and second maps comprise transformation functions between said spline positions for said plurality of position sensors of said second type and corresponding ones of said filtered estimated positions and said unfiltered estimated positions of said plurality of position sensors of said second type. 8. The system of claim 7 wherein said t... | ST. JUDE MEDICAL, ATRIAL FIBRILLATION DIVISION, INC. | A61B005/05 A61B005/06 A61B017/00 A61B018/02 A61B018/18 A61B005/042 A61B005/053 A61B005/00 A61B008/08 A61B008/12 A61M025/01 A61M025/09 A61M005/142 A61M005/178 A61N007/02 A61B034/20
US09801559 | B2 | 2017-10-31 | Integrated optical neural probe | 2011-07-25 | In certain embodiments, a neural probe comprises a substrate comprising elongated shanks for penetrating neural tissue, each comprising a proximal end and a distal end; at least one optical source integral to the neural probe for illuminating the neural tissue; and microelectrodes located proximate the distal end of the elongated shanks for monitoring neuronal activity. In certain embodiments, a method of monitoring neuronal activity comprises implanting the neural probe into a test subject such that the elongated shanks protrude into neural tissue of the test subject; illuminating the neural tissue with the at least one optical source; and measuring neuronal activity in response to illuminating the neural tissue. In certain embodiments, a device comprises a semiconductor chip; at least one optical source integral to the semiconductor chip; and sensor elements integral to the semiconductor chip for collecting data responsive to light emitted from the at least one optical source. | 1. A neural probe, comprising: a semiconductor substrate comprising a back end and one or more elongated shanks for penetrating neural tissue, each elongated shank comprising a proximal end and a distal end and at least one optical source located proximate the distal end of one or more elongated shanks such that, when the probe is implanted, the neural tissue may be illuminated without the use of an optical waveguide; andone or more microelectrodes located proximate the distal end of one or more of the elongated shanks,wherein the at least one optical source comprises a buffer layer, an active semiconductor layer comprising a III-V semiconductor selected from the group consisting of GaN, GaAs, InP, and AlGaN, and a passivation layer. 2. The neural probe of claim 1, further comprising readout circuitry electrically connected to the one or more microelectrodes. 3. The neural probe of claim 1, the back end further comprising a power source electrically connected to the at least one optical source. 4. The neural probe of claim 1, wherein the semiconductor layer comprises GaN. 5. The neural probe of claim 1, wherein the at least one optical source is integrated with the neural probe by silicon microfabrication. 6. A method of monitoring neuronal activity comprising: implanting the neural probe of claim 1 into a test subject such that the one or more elongated shanks protrude into neural tissue of the test subject;illuminating the neural tissue with the at least one optical source; andmeasuring neuronal activity with the microelectrodes in response to illuminating the neural tissue. 7. The method of claim 6, further comprising transmitting the measured neuronal activity to an analysis unit remote from the test subject. 8. The method of claim 6, wherein the test subject is a selected one of an animal and a human. | DIAGNOSTIC BIOCHIPS, INC. | A61B005/04 A61B005/0478 A61B005/00 A61N005/06 A61B018/22
US09808196 | B2 | 2017-11-07 | Sensors | 2013-06-28 | The present invention relates to an assembly comprising an elastic semi-conductive or conductive track and a flexible conductive support base arranged on a fabric. The present invention also relates to the use of a silicone rubber and/or fluorosilicone rubber loaded with an electrically conductive material for the preparation of the semi-conductive or conductive track, as well as to the use of a conductive fabric, that comprises conductive fibers for the preparation of the conductive support base. It further relates to a sensor comprising the assembly, wherein one of the flexible conductive support bases there is arranged a rigid electrical component, and the non-contact area of the other flexible conductive support base is adapted to be used as an electrode, wherein the electrode is characterized in that the conductive layer comprises a plurality of orifices filled with a silicone rubber and/or fluorosilicone rubber throughout the conductive area. The present invention also relates to a device comprising the sensor, as well as a garment comprising the device. | 1. An assembly comprising: an elastic semi-conductive or conductive track and an at least one flexible conductive support base arranged on a textile fabric substrate of a garment;the at least one flexible conductive support base being a textile comprising conductive fibers and non-conductive fibers and having at least one of its ends round shaped;at least one end of the track in contact with said at least one round shaped end of at least one flexible conductive support base;an at least one further end of the track in contact with a further at least one flexible conductive support base;a rigid electrical component arranged on and in electrical contact with a non-contacted area of the at least one flexible conductive support base; anda non-contacted area of the further at least one flexible conductive support base adapted to be used as an electrode for obtaining physiological signals through its contact with the skin of a wearer of the garment, said further at least one flexible conductive support base comprising a plurality of orifices filled with a silicone rubber throughout the conductive area. 2. The assembly of claim 1, wherein the flexible conductive support base is attached to the fabric with an adhesive. 3. The assembly of claim 1, wherein the track comprises a layer of silicone rubber and/or fluorosilicone rubber loaded with an electrically conductive material. 4. The assembly of claim 1, wherein the track comprises a layer of a room temperature curing silicone rubber and/or fluorosilicone rubber loaded with an electrically conductive material selected from carbon fibres, carbon black, nickel coated graphite, copper fibres and mixtures thereof. 5. The assembly of claim 1, wherein a thickness of the track comprises a thickness of at least 25 μm, 50 μm, 75 μm, 100 μm, 120 μm, 130 μm, 140 μm, 150 μm, 160 μm, 170 μm, 180 μm, 190 μm, 200 μm, 210 μm, 220 μm, 230 μm, 240 μm, 250 μm, 260 μm, 270 μm, 280 μm, 290 μm, 300 μm, 325 μm, 350 μm, 375 μm, 400 μm, 425 μm, 450 μm, 475 μm, 500 μm, 525 μm, 550 μm, 575 μm, 600 μm, 625 μm, 650 μm, 675 μm, 700 μm, 725 μm, 750 μm, 775 μm, 800 μm, 825 μm, 850 μm, 875 μm, 900 μm, 925 μm, 950 μm, 975 μm, 1000 μm. 6. The assembly of claim 3, wherein the track is integrated into the textile fabric substrate and partially into the at least one round shaped end of the flexible conductive support base by anchoring the silicone with the structure of the fibres of the textile fabric substrate and the conductive support base when cured the silicone at room temperature after being screen-printed on them. 7. The assembly of claim 6, wherein the silicone rubber and/or fluorosilicone rubber is screen-printed on a fabric and on the at least one round shaped end of the flexible conductive support base applying a pressure comprising at least 0.1 Kg/m2, at least 0.2 Kg/m2, at least 0.3 Kg/m2, at least 0.4 Kg/m2, at least 0.5 Kg/m2, at least 0.6 Kg/m2, at least 0.7 Kg/m2, at least 0.8 Kg/m2, at least 0.9 Kg/m2, at least 1 Kg/m2... | SMART SOLUTIONS TECHNOLOGIES, S.L. | A61B005/0408 A61B005/00
US09820672 | B2 | 2017-11-21 | Colon screening by using magnetic particle imaging | 2011-11-07 | A method for colon screening and collecting data by using Magnetic Particle Imaging wherein an imaging magnetic field is generated with a spatial distribution of the magnetic field strength such that the area of examination in the colon consists of a first sub-area with lower magnetic field strength, where the magnetization of a magnetic particle which was pre-delivered to the colon is not saturated, and a second sub-area with a higher magnetic field strength, where the magnetization of said magnetic particle is saturated. The spatial location of both sub-areas in the area of examination is modified so that the magnetization of said particles changes locally. Signals are acquired and are evaluated to obtain information about the spatial distribution of the signals in the area of examination. The method may be carried out during an entire peristaltic cycle in a colon portion or segment. | 1. A method for collecting data by using Magnetic Particle Imaging comprising the steps of: (a) pre-delivering magnetic particles to a colon of a subject;(b) generating an imaging magnetic field with a spatial distribution of the magnetic field strength such that an area of examination in the colon consists of a first sub-area of the colon with lower magnetic field strength, where the magnetization of magnetic particles which were pre-delivered to the colon is not saturated, and a second sub-area of the colon with a higher magnetic field strength, where the magnetization of said magnetic particles is saturated;(c) changing the spatial location of both sub-areas in the area of examination so that the magnetization of said particles changes locally; and(d) acquiring signals that depend on the magnetization in the area of examination influenced by this change, wherein the area of examination in the colon comprises a portion or segment of the colon and wherein steps (b) to (d) are carried out during an entire peristaltic cycle in said colon portion or segment, followed by a reassignment of the area of examination to a different portion or segment of the colon. 2. The method of claim 1, wherein the step of changing the spatial location of the sub-areas comprises the generation of a spatially and temporarily varying magnetic field. 3. The method of claim 1, wherein an additional homogenous, temporarily variable electromagnetic focus field is employed for the reassignment of the area of examination. 4. The method of claim 1, wherein said method is combined with a method of data analysis which employs a visualization algorithm, and/or a statistical analysis method determining the quality or usability of the data. 5. The method of claim 1, wherein said magnetic particles are pre-delivered to the colon via oral uptake of a food stuff or liquid. 6. The method of claim 5, wherein said oral uptake is to be started 1 to 5 days before performing steps b-d. 7. The method of claim 1, wherein the step of pre-delivering the magnetic particles comprises pre-delivering iron oxide particles which include a gastrointestinal resistant coating or a shell structure. 8. The method of claim 7, wherein the step of pre-delivering the magnetic particles comprises pre-delivering iron oxide particles which include a gastrointestinal resistant coating that is formed from a polymer. 9. The method of claim 7, wherein the step of pre-delivering the magnetic particles comprises pre-delivering iron oxide particles which include a shell structure that is comprised of a polymer or latex. 10. The method of claim 9, wherein the step of pre-delivering the magnetic particles comprises pre-delivering iron oxide particles which include a shell structure that is comprised of a dispersion of polymer coated or latex coated iron oxide particles. 11. The method of claim 7, wherein the step of pre-delivering the magnetic particles comprises pre-delivering iron oxide particles which includ... | Koninklijke Philips N.V. | A61B005/05 A61K049/18
US09814403 | B2 | 2017-11-14 | Method and apparatus for measuring the levels of hormones, neuro transmitters, bio markers, or the like | 2014-01-10 | A method of predicting substance levels from EEG data is disclosed. The method includes analyzing EEG data to obtain the average power for each of a plurality of predetermined frequency bands and calculating a value from the average powers derived for each frequency band, said value being calculated by combining the average powers for each frequency band by dividing and/or multiplying according to a predetermined order. The method further includes obtaining an estimate of the hormone level from the equation Y=bX+C, where Y is the substance level to be predicted, X is the value and b and C are constants, wherein the substance is selected from hormones, neuro transmitters and bio markers. | 1. A computer implemented method of determining a substance level in a body of a human or non-human animal subject, the method comprising: obtaining EEG data comprising EEG signals collected from the subject and analyzing the EEG data to obtain the average power for each of a plurality of frequency bands;calculating a value from the average powers derived for each frequency band, said value being calculated by combining the average powers for each frequency band by dividing and/or multiplying according to a predetermined order, wherein the predetermined order is determined based upon the substance; andobtaining an estimate of the substance level from the equation Y=bX+C,where Y is the substance level to be predicted, X is the value and b and C are constants,wherein the substance is selected from hormones, neuro transmitters and bio markers. 2. The method of claim 1, wherein the value is a ratio of selected ones of the average powers. 3. The method of claim 1, wherein the frequency bands are selected from Delta, Theta, Alpha, Beta, SMR, High Beta and Gamma. 4. The method of claim 1, wherein the substance is cortisol. 5. The method of claim 4, wherein the frequency bands are the Theta, Alpha, Delta and SMR bands. 6. The method of claim 5, wherein the bands are divided in the order of Theta/Alpha/Delta/SMR. 7. The method of claim 1, wherein the EEG data is collected from the human or non-human animal subject for at least 30 seconds. 8. The method of claim 1, wherein the EEG data is collected from at least two electrode positions. 9. The method of claim 1, wherein the frequency bands are the frequency bands which have a correlation with increasing or decreasing levels of the substance. 10. The method of claim 1, further comprising selecting the frequency bands by measuring the average power for a plurality of frequency bands for a plurality of substance levels and selecting the bands which show a correlation with increasing or decreasing substance levels. 11. The method of claim 1, further comprising: determining, for each of at least three levels of the substance, a corresponding set of average power values to provide a plurality of sets of average power values, wherein each set of average power values comprises an average power value for each of the frequency bands;selecting a method of combining the average power values of the sets to provide a test value for each set, each test value corresponding to a level of the at least three levels of the substance;wherein the selecting comprises varying how the average power of the frequency bands are combined and selecting the method of combining which provides a relationship between the test values and the corresponding at least three levels of the substance that is closest to a straight line. 12. The method of claim 11, further comprising varying the boundaries of the frequency bands to obtain a better correlation to a straight line. 13. The method of claim 1, further comprising determining t... | - | A61B005/04 A61B005/048 A61B005/145 A61B005/0476 A61B005/00
US09782074 | B2 | 2017-10-10 | Hand-held medical-data capture-device having optical detection of a vital sign from multiple filters and interoperation with electronic medical record systems to transmit the vital sign and device information | 2014-12-31 | In one implementation, an apparatus estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point is described, estimates temperature from a digital infrared sensor and determines vital signs from a solid-state image transducer, or determines vital signs from a solid-state image transducer and estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point; after which the estimated and/or determined information is transmitted to an external database. | 1. An apparatus of variation amplification to communicate biological vital signs, the apparatus comprising: a microprocessor that is programmed to execute: a cropper that is operable to receive at least two images and that is operable to crop each of the images to exclude a border area of the images, generating at least two cropped images;a skin-pixel-identifier that is operably coupled to the cropper and that identifies pixel values that are representative of skin in the at least two cropped images, generating identified skin pixels;a spatial bandpass filter that is operably coupled to the skin-pixel-identifier and that processes the identified skin pixels, generating spatial bandpassed filtered identified pixel values of skin;a regional facial clusterial module that is operably coupled to the spatial bandpass filter and that applies spatial clustering to the spatial bandpassed filtered identified pixel values of skin, generating spatial clustered spatial bandpassed filtered identified pixel values of skin;a temporal bandpass filter that is operably coupled to the regional facial clusterial module and that is applied to the spatial clustered spatial bandpassed filtered identified pixel values of skin, generating temporal bandpass filtered spatial clustered spatial bandpass filtered identified pixel values of skin;a temporal-variation identifier that is operably coupled to the temporal bandpass filter and that identifies a temporal variation of the temporal bandpass filtered spatial clustered spatial bandpass filtered identified pixel values of skin, generating temporal-variation identified temporal bandpass filtered spatial clustered spatial bandpass filtered identified pixel values of skin;a vital-sign generator that is operably coupled to the temporal-variation identifier that generates at least one vital sign from the temporal-variation identified temporal bandpass filtered spatial clustered spatial bandpass filtered identified pixel values of skin; anda display device that is operably coupled to the vital-sign generator that is operable to display the at least one vital sign; anda wireless communication subsystem that is operably coupled to the microprocessor and that is operable to transmit a representation of the at least one vital sign,wherein the wireless communication subsystem further comprises a component that is operable to transmit a representation of date and time, operator identification, patient identification and manufacturer, model number and firmware revision of the apparatus. 2. The apparatus of claim 1, wherein the regional facial clusterial module further comprises: a fuzzy clusterer. 3. The apparatus of claim 1 further comprising: a storage device that is operably coupled to the vital-sign generator and that is operable to store the at least one vital sign in a nonvolatile memory. 4. The apparatus of claim 1 further comprising: a storage device that is operably coupled to the vital-sign generator and that is operable to... | Arc Devices, LTD | G08B023/00 A61B005/00 A61B005/01 G06F019/00 A61B005/021 A61B005/026 A61B005/08 A61B005/024 A61B005/0205 G06T007/00 H04W048/16 H04W076/02 A61B005/0402 A61B005/03 G01J005/00 H04W080/04 G01J005/02 G01J005/10 G08C017/02 H04L029/08 H04W012/06 H04B007/26 H04L012/46 H04N005/33 H04W048/08 G08B005/36 G08B021/18 G01K013/00 G06K009/62 G06T007/174 G06T007/269 G06T007/90 A61B005/0245 H04W012/08
US09814437 | B2 | 2017-11-14 | Anatomic range planning in positron emission tomography | 2014-02-18 | Anatomic range planning is provided in positron emission tomography (PET). The user indicates one or more ranges on an image of a patient based on anatomy. Rather than planning by bed position, the planning is based on the anatomy of the patient without reference to the length of the PET detector. The user interface for PET examination avoids overlapping boxes and other confusion based on bed position. Different anatomical ranges may be assigned different PET parameters, such as reconstruction parameters. A processor may automatically alter the examination (e.g., by extending the detection range beyond the region of interest or by increasing duration at an end position) to account for the sensitivity profile since the anatomical region of interest is known. Anatomical region specific directions may be included as part of planning, aiding in performing different protocols for different anatomical ranges. | 1. A method for anatomic range planning in positron emission tomography (PET), the method comprising: displaying an image of a patient as positioned on a bed of a PET scanner;receiving a selection on the image, the selection identifying locations specific to an anatomical starting point and an anatomical stopping point of anatomy of the patient represented in the image, the anatomical starting point and anatomical stopping point being different than a length of a detector of the PET scanner and being different than an integral multiplier of the length; and thenoperating the PET scanner as a function of the anatomical starting and stopping points based on the selection. 2. The method of claim 1 wherein displaying the image comprises displaying an x-ray based topogram of the patient, the topogram including the anatomy comprising internal anatomical structure of the patient. 3. The method of claim 2 wherein displaying comprises generating the image of the patient with an x-ray emitter and x-ray detector connected with the PET scanner. 4. The method of claim 1 wherein displaying comprises displaying on a display of the PET scanner, and wherein receiving comprises receiving on a user input of the PET scanner. 5. The method of claim 1 wherein receiving comprises receiving the anatomical starting and stopping points without an indication of bed position. 6. The method of claim 1 wherein receiving comprises receiving the anatomical starting and stopping points as a first range and receiving indications of a second range, the first range, second range, and any other ranges not overlapping and disconnected from the length of the detector. 7. The method of claim 1 wherein receiving comprises receiving the anatomical starting and stopping points as an entire range of interest of patient and further receiving at least one division of the entire detection range into at least two portions free of overlap, each portion being defined based on the anatomy of the patient. 8. The method of claim 7 further comprising receiving different PET reconstruction parameters for two of the portions; and wherein operating comprises reconstructing data for the different portions using the respective different reconstruction parameters. 9. The method of claim 7 further comprising receiving different speed of movement of the bed for two of the portions; and wherein operating comprises moving the bed at the different speeds in the respective portions. 10. The method of claim 1 wherein receiving comprises receiving user input of positions arbitrary relative to the PET scanner as the anatomical starting and stopping points relative to the image. 11. The method of claim 1 wherein receiving comprises receiving the anatomical starting and stopping points as an entire range of interest of patient; and further comprising:computing, by a processor, a speed, an additional region, or speed and additional region for detection by the PET scanner at ends of the range of interest; and... | Siemens Medical Solutions USA, Inc. | A61B005/05 A61B006/00 A61B006/03 A61B006/04
US09814452 | B2 | 2017-11-14 | Powered surgical instrument | 2014-03-31 | A surgical instrument including a housing, an endoscopic portion, a shaft portion, and an end effector is disclosed. The endoscopic portion extends distally from the housing and defines a longitudinal axis. The shaft portion is selectively connectable to a distal end of the endoscopic portion. The end effector is selectively connectable to a distal end of the shaft portion. | 1. A surgical instrument, comprising: a housing;an endoscopic portion extending distally from the housing and defining a longitudinal axis;a drive gear disposed at least partially within the housing, the drive gear rotatable about a drive gear axis extending therethrough;a drive motor configured to rotate the drive gear;a shift motor configured to move the drive motor along the drive gear axis; anda shift sensor arranged to determine the longitudinal position of the drive motor. 2. The surgical instrument according to claim 1, wherein the housing includes a screen for displaying information concerning the status of the surgical instrument. 3. The surgical instrument according to claim 1, further comprising a shift screw, the shift motor being configured to drive the shift screw. 4. The surgical instrument according to claim 3, further comprising a screw housing, the shift screw being in engagement with the screw housing for moving the screw housing longitudinally. 5. The surgical instrument according to claim 1, wherein information from the shift sensor is communicated to a screen associated with the instrument. 6. The surgical instrument according to claim 1, wherein the shift sensor communicates with at least one switch to start or stop the shift motor. 7. The surgical instrument according to claim 1, further comprising a position lock for engaging the drive motor in one of a plurality of longitudinally-spaced positions. 8. The surgical instrument according to claim 7, wherein the shift sensor interacts with the position lock to determine a status of the shift motor. 9. The surgical instrument according to claim 7, further comprising a drive motor casing at least partially surrounding the drive motor, the drive motor casing including at least three slots therein, wherein each slot is matable with the position lock to maintain the drive gear in a desired longitudinal position. 10. The surgical instrument according to claim 1, wherein the drive gear axis and the longitudinal axis are coaxial. 11. The surgical instrument according to claim 1, wherein the drive gear axis and the longitudinal axis are parallel. 12. The surgical instrument according to claim 1, wherein the drive motor includes a drive motor shaft extending distally therefrom, the drive motor shaft being mechanically engaged with the drive gear. 13. The surgical instrument according to claim 12, wherein the drive motor shaft includes a longitudinally flat portion which is configured to allow rotational float between the drive motor shaft and the drive gear. 14. The surgical instrument according to claim 1, wherein the drive gear is selectively movable along the drive gear axis between a proximal-most position and a distal-most position, and wherein the drive motor is disposed proximally of the proximal-most position of the drive gear. 15. The surgical instrument according to claim 1, wherein the shift motor is disposed in mechanical cooperation with the drive gear and is... | Covidien LP | A61B017/00 A61B005/08 A61B017/068 A61B017/072 A61B017/29 A61B090/00
US09801689 | B2 | 2017-10-31 | Method and system for patient-specific modeling of blood flow | 2015-09-25 | Embodiments include a system for determining cardiovascular information for a patient. The system may include at least one computer system configured to receive patient-specific data regarding a geometry of the patient's heart, and create a three-dimensional model representing at least a portion of the patient's heart based on the patient-specific data. The at least one computer system may be further configured to create a physics-based model relating to a blood flow characteristic of the patient's heart and determine a fractional flow reserve within the patient's heart based on the three-dimensional model and the physics-based model. | 1. A method for vascular assessment comprising: receiving a plurality of 2D angiographic images of a portion of a patient's vasculature having a stenotic lesion or plaque;generating an anatomical model of the patient's vasculature using the received plurality of 2D angiographic images;determining a flow characteristic of blood flow through the anatomical model;performing a plurality of iterations of: modifying the anatomical model by modifying measurements of the vasculature at one or more locations along one or more vessels of the patient's vasculature;determining a modified flow characteristic of blood flow through the modified anatomical model;calculating an index indicative of vascular function based, at least in part, on the first flow characteristic of blood flow through the anatomical model and the modified flow characteristic of blood flow through the modified anatomical model;generating and outputting, to an electronic storage medium or display, a report of a modified anatomical model from the one or more iterations, for which the modified flow characteristic of blood flow through the modified anatomical model maximizes the index indicative of vascular function; anddetermining and outputting, to an electronic storage medium or display, indicia of one or more treatment options based on the report. 2. The method according to claim 1, wherein said measurements of the vasculature are at one or more locations along a centerline of at least one branch of the vasculature. 3. The method according to claim 1, wherein said flow characteristic of the anatomical model comprises resistance to fluid flow. 4. The method according to claim 3, further comprising identifying in the anatomical model a stenosed vessel and a downstream portion of said stenosed vessel, and calculating said resistance to fluid flow in said downstream portion; wherein said index is calculated based on a volume of said downstream portion, and on a contribution of said stenosed vessel to said resistance to fluid flow. 5. The method according to claim 1, wherein said flow characteristic of the anatomical model comprises one or more of: blood flow rate, blood pressure, and fractional flow reserve. 6. The method according to claim 1, wherein the anatomical model is a three-dimensional vessel tree. 7. The method according to claim 6, wherein said vessel tree comprises data pertaining to location, orientation and size of vessels at a plurality of points within said portion of the vasculature. 8. The method according to claim 1, wherein the modified anatomical model comprises an inflated vessel replacing a stenotic vessel in the anatomical model. 9. The method according to claim 1, wherein the anatomical model is a three-dimensional vessel tree and the modified anatomical model is a modification of the three-dimensional vessel tree. 10. The method according to claim 1, wherein said flow characteristics of said anatomical model and said modified flow characteristics of said mod... | HeartFlow, Inc. | G06K009/00 A61B034/10 G06F017/10 A61B034/00 A61B005/02 A61B005/026 G06F019/00 G06G007/60 G06F019/26 A61B005/00 A61B005/029 A61B006/03 A61B006/00 A61B008/06 A61B008/08 G01R033/563 A61B005/107 G06T007/00 G06T017/00 G06K009/62 A61B005/021 G06F017/50 G06F019/12 A61B005/024 A61B005/22 G06K009/46 G06T007/20 G06T011/00 G06T011/20 A61M005/00 G06K009/52 A61B005/055 A61B005/11 G06T017/20 G06T015/10 G06T011/60 G06T007/60 A61B008/02 A61B008/04 G06T007/70 G06T007/73 G06T007/11 G06T007/12 G06T007/13 G06T007/149 G06T007/62 G01R033/56 A61B090/00
US09788745 | B2 | 2017-10-17 | Biological signal measuring equipment | 2010-06-28 | A biological signal measuring equipment used in a measurement of a biological signal in a head portion, has a jaw contact portion that comes into contact with a forehead; an occiput contact portion that comes into contact with an occiput; and a supporter that is interposed in a front and rear direction of the head portion and is supported by the head portion by using the jaw contact portion and the occiput contact portion as ends. The biological signal measuring equipment performs the measurement in the same manner as the case of measuring a wave motion of a short period, even in the case of measuring the wave motion of a long period generated in a head portion. | 1. A biological signal measuring equipment, comprising: a supporter including: an arm;a forehead contact portion configured to come into contact with a forehead of a head portion;an occiput contact portion configured to come into contact with an occiput of the head portion; anda fixing region, connected to the forehead contact portion,wherein the supporter is configured to be positioned between the forehead contact portion and the occiput contact portion when the supporter is arranged on the head portion;an attachment opening present in the fixing region and configured to attach the arm that extends from the supporter in a freely attachable and detachable manner;a first electrode attached to a portion of the supporter situated on a median sagittal plane of the supporter;a second electrode configured to be positioned at an earlobe and attached to the arm;a plurality of magnets, wherein the plurality of magnets and the second electrode are configured to affix to the earlobe so that the earlobe is interposed between the plurality of magnets and the second electrode;a case that protrudes from the supporter at the fixing region; andan electronic device configured to interface with the first electrode and the second electrode. 2. The biological signal measuring equipment of claim 1, wherein the electronic device includes an analog-based electronic circuit, which is present at the fixing region on the supporter, and a digital-based electronic circuit. 3. The biological signal measuring equipment of claim 1, further comprising: a plurality of amplification portions that are present at the fixing region on the supporter,wherein the plurality of amplification portions corresponds to the first electrode and the second electrode, andwherein a ground of the plurality of amplification portions is concentrated at the fixing region on the supporter. 4. The biological signal measuring equipment of claim 1, further comprising: a measurement portion configured to measure a potential difference between the first electrode and the second electrode as a biological signal of the head portion,wherein the measurement portion is further configured to measure a level in biological signals for each of a plurality of periods and compare an average of the level in the biological signals for each of the plurality of periods with a threshold value, and based on a determination that the average is lower than the threshold value, the measurement portion is further configured to notify, via a speaker, that the first electrode and the second electrode need to be remounted. 5. The biological signal measuring equipment of claim 1, wherein the supporter further comprises a supporter portion situated on the median sagittal plane of the supporter, and a pair of branching portions branched from the supporter portion, wherein the occiput contact portion comprises a first occiput contact portion which is configured to contact a first neck side slope portion of a right side external occ... | SONY CORPORATION | A61B005/0478 A61B005/0496 A61B005/00 A61B005/04 A61B005/0488
US09794705 | B2 | 2017-10-17 | Hearing aid tuning system and method | 2014-10-01 | Systems and methods are provided herein that provide for hearing aid tuning. | 1. A tuning-device-implemented method for tuning a hearing aid worn by a wearer, the hearing aid having a plurality of programmatically adjustable settings, the tuning device comprising at least one coupling body, the method comprising: displaying, on a display associated with the tuning device, a first graphical user interface (“GUI”) allowing the wearer to select one of a plurality of hearing aid issues;receiving first input, by the tuning device from the wearer via said first GUI, indicating a wearer-indicated hearing aid issue of said plurality of hearing aid issues;programmatically coupling, by the tuning device, with the hearing aid via said at least one coupling body;obtaining, by the tuning device, a plurality of current hearing-aid audio settings associated with the programmatically coupled hearing aid;selecting, by the tuning device, an audio stimulus of a plurality of audio stimuli based at least in part on said wearer-indicated hearing aid issue;presenting, by the tuning device to the wearer, said selected audio stimulus;selecting, by the tuning device according to at least one of said selected audio stimulus and said wearer-indicated issue, a second GUI soliciting wearer feedback related to presenting said selected audio stimulus;displaying said second GUI on said display;receiving second input, by the tuning device from the wearer via said second GUI, indicating said wearer feedback;automatically determining, by the tuning device, a change to at least one of said plurality of current hearing-aid audio settings according to said wearer feedback;programmatically adjusting, by the tuning device, said at least one of said plurality of current hearing-aid audio settings. 2. The method of claim 1, wherein obtaining said plurality of current hearing-aid audio settings comprises requesting said plurality of current hearing-aid audio settings from the programmatically coupled hearing aid. 3. The method of claim 1, wherein obtaining said plurality of current hearing-aid audio settings comprises retrieving said plurality of current hearing-aid audio settings from a user-data database associated with the tuning device. 4. The method of claim 3, further comprising storing said adjusted at least one of said plurality of current hearing-aid audio settings in said user-data database. 5. The method of claim 1, further comprising: programmatically coupling, by the tuning device, with a second hearing aid worn by the wearer via a second coupling body of said at least one coupling body; andprogrammatically isolating the hearing aid worn by the wearer from said second hearing aid prior to presenting said selected audio stimulus. 6. The method of claim 5, wherein isolating the hearing aid comprises a selected one of programmatically adjusting a gain setting of said second hearing aid to a minimum-gain setting and programmatically switching said second hearing aid off. 7. The method of claim 1, wherein presenting said selected audio stimulus compris... | Bowie-Wiggins LLC | H04R025/00 A61B005/12 G06F003/0484
US09782117 | B2 | 2017-10-10 | System with kinesthetic stimulation medical device for the non-invasive assessment of the sympathovagal balance of a patient | 2015-02-09 | A system for assessing a sympathovagal balance of a patient includes a generator configured to produce stimulation pulses, a stimulator that receives the stimulation pulses produced by the generator and outputs a stimulation energy, a sensor configured to measure a physiological signal, and a processor. The processor is configured to control the generator during a plurality of sequences to produce the stimulation pulses over a test period, determine a biological parameter of a current activity of the patient derived from the physiological signal, determine a variation of the biological parameter resulting from the stimulation pulses, and determine a sympathovagal balance index according to the variation in the biological parameter. | 1. A system for assessing a sympathovagal balance of a patient, comprising: a generator configured to produce stimulation pulses;a stimulator that receives the stimulation pulses produced by the generator and outputs a stimulation energy;a sensor configured to measure a physiological signal; anda processor configured to: control the generator during a plurality of sequences to produce the stimulation pulses over a test period;determine a biological parameter of a current activity of the patient derived from the physiological signal, the biological parameter comprising at least one of a PR interval as a dromotropic parameter or an endocardial acceleration as an inotropic parameter;determine a variation of the biological parameter resulting from the stimulation pulses from at least two values for each sequence of the test period, the at least two values for each sequence including: (i) a base value measured for a baseline period prior to the production of the stimulation pulses; (ii) a current value measured during an ANS period of production of the stimulation pulses; and (iii) a current value measured for a recovery period during which the stimulation pulses are stopped; wherein the generator produces a plurality of stimulation pulses during the ANS period; andwherein each of the plurality of stimulation pulses includes a rank i, such that the rank i ranges from 1 to n for each sequence; anddetermine a sympathovagal balance index according to the variation in the biological parameter. 2. The system of claim 1, wherein the physiological signal includes at least one of a cardiac activity, an ECG, a heart rate, breathing, an oxygen saturation, a pulse wave, and a phonocardiogram. 3. The system of claim 1, wherein the processor determines the current value of the biological parameter in synchronization with a cardiac rhythm of the patient during a plurality of successive cardiac cycles during the recovery period. 4. The system of claim 1, wherein each of the plurality of the stimulation pulses corresponds with one of a plurality of samples of the biological parameter, wherein the processor calculates an average of the plurality of samples of the biological parameter of the same rank i over all sequences of the test period, resulting in n averaged values of the biological parameter. 5. The system of claim 4, wherein the n averaged values of the biological parameter are the sympathovagal balance index of the test period, wherein the sympathovagal balance index is calculated during the test period. 6. The system of claim 1, wherein the processor inhibits the generator upon detection of an occurrence of at least one of a heart rate below a predetermined threshold, a presence of a patient's physical activity, a cardiac stimulation delivery, a presence of an apnea or hypopnea episode, and a presence of an arrhythmia episode. 7. A method for assessing a sympathovagal balance of a patient, comprising: acquiring, by a sensor, a physiological signal;cont... | SORIN CRM SAS UNIVERSITÉ DE RENNES 1 | A61B005/04 A61B005/00 A61B005/024 A61B005/0452 A61B005/0205 A61B005/0402 A61B005/08 A61B005/1455 A61B007/04
US09801780 | B2 | 2017-10-31 | Methods and devices for remote ischemic conditioning via partial limb occlusion | 2015-03-18 | Single- or dual-bladder devices for automated delivery of remote ischemic conditioning treatment via partial limb occlusion involve various methods of operating the cuff in which partial or full limb occlusion is achieved during the periods of cuff inflation. Achieving clinical benefits of remote ischemic conditioning without extended cessation of limb blood flow are advantageous due to lower required cuff pressure and reduced risk of clot formation in the limb vasculature. | 1. A device for delivery of remote ischemic conditioning treatment to a subject, said device comprising: a cuff configured for placement over a limb of the subject, said cuff is further configured to at least partially reduce blood flow in said limb upon inflation thereof;a controller operably connected to said cuff, said controller is configured upon activation to deliver said remote ischemic conditioning treatment by performing a plurality of remote ischemic conditioning treatment cycles, each treatment cycle comprising: a period of cuff inflation when said controller is causing said cuff to be inflated to and maintained at or above a pressure sufficient to cause at least a partial reduction of blood flow in said limb to produce an ischemic stress therein for at least about one minute, anda period of cuff deflation when said controller is causing said cuff to deflate so as to increase blood flow in said limb and relieve said ischemic stress,wherein said ischemic stress is sufficient to cause said remote ischemic conditioning treatment, said controller is further configured to avoid complete occlusion of said limb by said cuff lasting more than about one minute during at least one of said treatment cycles. 2. The device as in claim 1, wherein said controller is further configured to inflate said cuff during at least a portion of said period of cuff inflation to a pressure at or below a lesser pressure of either a systolic blood pressure of said subject or a limb occlusion pressure of said subject. 3. The device as in claim 2, wherein said controller is further configured during said entire period of cuff inflation to inflate said cuff at or above a pressure sufficient to cause about ninety percent reduction of unrestricted blood flow in said limb. 4. A device for delivery of remote ischemic conditioning treatment to a subject, said device comprising: a cuff configured for placement over a limb of the subject, said cuff is further configured to at least partially reduce blood flow in said limb upon inflation thereof;a controller operably connected to said cuff, said controller is configured upon activation to deliver said remote ischemic conditioning treatment by performing a plurality of remote ischemic conditioning treatment cycles, each treatment cycle comprising: a period of cuff inflation when said controller is causing said cuff to inflate to a pressure sufficient to at least partially reduce blood flow in said limb and produce an ischemic stress therein causing said remote ischemic conditioning, said period of cuff inflation lasting at least about one minute, anda period of cuff deflation when said controller is causing said cuff to deflate and increase blood flow in said limb to relieve said ischemic stress,wherein during at least one of said periods of cuff inflation said controller is further configured to vary pressure in said cuff so as to increase or decrease said ischemic stress in said limb. 5. The device as in claim 4, wherein... | LifeCuff Technologies Inc. | A61B005/022 A61H009/00 A61B017/135 A61B017/12 A61B017/00
US09788776 | B1 | 2017-10-17 | Protein M-based in vivo diagnostic system and detection method | 2014-09-22 | A wearable device mounted on a living body can detect an analyte response signal transmitted from tissue in the living body. The tissue contains at particles conjugated to at least one complex of (i) an antibody labeled with a fluorophore, the labeled antibody having a target analyte binding site; and (ii) protein M labeled with a quencher that is complimentary to the fluorophore of the labeled antibody, wherein the labeled protein M competes with the target analytes for the target analyte binding site on the labeled antibody; wherein the fluorophore and quencher are spectrally matched such that there is a detectable change in the fluorescent signal. The analyte response signal is related to interaction of the target analytes with the complexes. A processor can determine a presence or absence of the analytes based on the analyte response signal. | 1. A composition comprising: a particle conjugated to a complex comprising(i) antibodies labeled with a fluorophore, the labeled antibodies having target analyte binding sites; and(ii) protein M labeled with a quencher that is complimentary to the fluorophore of the labeled antibodies,wherein the labeled protein M competes with target analytes for the target analyte binding sites on the labeled antibodies;wherein the fluorophore and quencher are spectrally matched such that there is a detectable change in the fluorescent signal of the labeled antibodies when the fluorophore and the quencher are moved into or out of proximity to each other, andwherein fluorescent light levels change proportionately in response to the amount of target analytes that are able to bind with the labeled antibodies. 2. The composition of claim 1, wherein the particle is a nanoparticle. 3. The composition of claim 1, further comprising a linker to couple the labeled antibodies with labeled protein M. 4. A method comprising: (a) combining a solution with a nanoparticle conjugated to a complex, the complex comprising (i) antibodies labeled with a fluorophore, the labeled antibodies having target analyte binding sites; and(ii) protein M labeled with a quencher that is complimentary to the fluorophore of the labeled antibodies,wherein the labeled protein M competes with target analytes for the target analyte binding sites on the labeled antibodies;wherein the fluorophore and quencher are spectrally matched such that there is a detectable change in the fluorescent signal of the labeled antibodies when the fluorophore and the quencher are moved into or out of proximity to each other, andwherein fluorescent light levels change proportionately in response to the amount of target analytes that are able to bind with the labeled antibodies; and(b) measuring the fluorescence light level related to a presence of the target analytes in the solution. 5. The method of claim 4, wherein the particle is a nanoparticle. 6. The method of claim 4, further comprising a linker to couple the labeled antibodies with the labeled protein M. 7. A method comprising: introducing particles conjugated to complexes into the living body, the complexes comprising (i) antibodies labeled with a fluorophore, the labeled antibodies having target analyte binding sites; and (ii) protein M labeled with a quencher that is complimentary to the fluorophore of the labeled antibodies,wherein the labeled protein M competes with target analytes for the target analyte binding sites on the labeled antibodies; wherein the fluorophore and quencher are spectrally matched such that there is a detectable change in the fluorescent signal of the labeled antibodies when the fluorophore and the quencher are moved into or out of proximity to each other; wherein fluorescent light levels change proportionately in response to the amount of target analytes that are able to bind with the labeled antibodies; wherein the complexes ar... | Verily Life Sciences LLC | A61B005/1455 A61B005/154 G01N033/543 A61B005/00
US09788232 | B2 | 2017-10-10 | Opportunistic syncing methods for wearable devices | 2015-12-15 | A system and method for monitoring the health of an animal using multiple sensors is described. The wearable device may include one or more sensors whose resultant signal levels may be analyzed in the wearable device or uploaded to a data management server for additional analysis. One or more embodiments include variations of the UWB system to accommodate differences in animals, such as scheduling or attempting transmissions between the wearable device and the data management server in such a way as to increase the likelihood of a successful transmission. | 1. A method comprising: receiving, by a processor, first data indicative of one or more signals at a wearable device for an animal, wherein at least one of the one or more signals is a reflected Ultra-Wide Band (UWB) signal;determining, by the processor, a first syncing schedule for attempting to transmit said first data indicative of the one or more signals to a remote device;determining, by the processor, a second syncing schedule different from the first syncing schedule for attempting to transmit second data, different from the first data, indicative of the one or more signals to said remote device. 2. The method of claim 1, wherein the determining the second syncing schedule comprises modifying the first syncing schedule. 3. The method of claim 1, wherein the first syncing schedule and second syncing schedule comprise a plurality of attempted syncing times. 4. The method of claim 3, further comprising: determining, by the processor, a third syncing schedule different from the first syncing schedule and the second syncing schedule for attempting to transmit third data, indicative of the one or more signals to said remote device different from the first data and second data. 5. The method of claim 4, wherein the third syncing schedule comprises at least one attempted syncing time from the first syncing schedule and at least one attempted syncing time from the second syncing schedule. 6. The method of claim 1, wherein determining the second syncing schedule comprises mutating at least one element of the first syncing schedule. | i4C Innovations Inc. | G06F001/16 H04W028/02 H04W056/00 A01K027/00 A01K029/00 A61B005/11 A61D017/00 G06F019/00 H04W088/02 H04B001/7163 H04M001/725 G01S005/00 A61B005/00
US09795341 | B2 | 2017-10-24 | Physiological monitoring system featuring floormat and wired handheld sensor | 2016-01-05 | A physiological monitoring system features a Floormat and Handheld Sensor connected by a cable. A user stands on the Floormat and grips the Handheld Sensor. These components measure time-dependent physiological waveforms from a user over a conduction pathway extending from the user's hand or wrist to their feet. The Handheld Sensor and Floormat use a combination of electrodes that inject current into the user's body and collect bioelectric signals that, with processing, yield ECG, impedance, and bioreactance waveforms. Simultaneously, the Handheld Sensor measures photoplethysmogram waveforms with red and infrared radiation and pressure waveforms from the user's fingers and wrist, while the Floormat measures signals from load cells to determine ‘force’ waveforms to determine the user's weight, and ballistocardiogram waveforms to determine parameters related to cardiac contractility. Processing these waveforms with algorithms running on a microprocessor yield the vital sign, hemodynamic, and biometric parameters. | 1. A biometric sensor system configured to measure at least one physiological parameter, comprising: a first device comprising a generally flat Floormat configured to rest stably on a generally flat surface and to support the weight of a user standing thereon, the Floormat including: a weight-measuring system comprising at least one load cell and an amplifier system configured to measure a time-dependent load-cell voltage from at least one load cell and to process the time-dependent load-cell voltage to determine a current time-dependent load-cell waveform; anda first electrode disposed at an upper surface of the Floormat and in position to make contact with the sole of one of the user's feet when the user stands on the Floormat; anda second device comprising a Handheld Sensor configured to be supported at a region of one of the user's hands, the Handheld Sensor: 1) including a second electrode disposed in position to make contact with skin in the region of the user's hand when the Handheld Sensor is supported thereat; and 2) being electrically connected to the Floormat via a cable having one or more electrical conductors disposed therein;an analog system configured to receive biometric signals from the first electrode and from the second electrode and to process the biometric signals to generate a current ECG waveform, the analog system being located in one of the first and second devices and the electrode disposed in the other device being electrically connected to the analog system by means of said one or more electrical conductors; anda processing system configured to receive the current time-dependent load-cell waveform from the Floormat and the current ECG waveform from the analog system and to process the current load-cell waveform and the current ECG waveform, wherein the processing system comprises: 1) a computer memory configured to store therein historical ECG waveform data corresponding to a plurality of users who have stood on the Floormat; and 2) computer code configured to first identify a pool of potential users that could be standing on the Floormat based on a current weight value and then to compare the current ECG waveform against the historical ECG waveform data for the pool of potential users to identify a given user standing on the Floormat. 2. The biometric sensor system of claim 1, wherein the current ECG waveform comprises a plurality of successive QRS complexes and the processing system is configured to determine a numerical index value for the current ECG waveform from said plurality of successive QRS complexes, the processing system being further configured to identify the given user standing on the Floormat by comparing the numerical index value determined for the current ECG waveform to numerical index values associated with the historical ECG waveform data. 3. The biometric sensor system of claim 2, wherein the numerical index value is calculated based on the magnitude of the Q and R deflections within the success... | TOSENSE, INC. | A61B005/02 A61B005/00 A61B005/0205 A61B005/0402 A61B005/021 A61B005/04
US09786189 | B2 | 2017-10-10 | Aircraft control systems | 2016-01-25 | A method of controlling an aircraft in the event of pilot incapacity includes detecting, using at least one sensor, the occurrence of an event relating to a pilot's capacity to control the aircraft, determining whether the event justifies a controlled takeover of the aircraft from the pilot, asserting a controlled takeover of the aircraft from the pilot if it has been determined that a controlled takeover is justified, and executing a controlled landing of the aircraft without assistance of the pilot. | 1. A method of controlling an aircraft in the event of pilot incapacity, comprising: detecting, using at least one sensor, the occurrence of an event relating to a pilot's capacity to control the aircraft;determining, using a controller that is operatively connected to the at least one sensor, whether the event justifies a controlled takeover of the aircraft from the pilot;asserting a controlled takeover of the aircraft from the pilot if it has been determined that a controlled takeover is justified; andexecuting a controlled landing of the aircraft without assistance of the pilot. 2. The method of claim 1, wherein detecting the occurrence of an event includes detecting a biometric parameter using the sensor. 3. The method of claim 2, wherein determining whether the event justifies a controlled takeover includes determining whether the biometric parameter is outside of a predetermined range. 4. The method of claim 1, wherein detecting the occurrence of an event includes detecting aircraft course and/or location information. 5. The method of claim 4, wherein determining whether the event justifies a controlled takeover includes determining whether the course information represents a deviation from a predetermined flight course beyond a deviation threshold. 6. The method of claim 1, wherein detecting the occurrence of an event includes detecting an aircraft flight state. 7. The method of claim 6, wherein determining whether the event justifies a controlled takeover includes determining whether the aircraft flight state is within predetermined aircraft operational parameters. 8. The method of claim 1, wherein detecting the occurrence of an event includes detecting an action initiated by the pilot or other crew member. 9. The method of claim 1, wherein determining whether the event justifies a controlled takeover includes providing a signal to the pilot indicating that an event has occurred justifying a controlled takeover to confirm whether the pilot is incapacitated. 10. The method of claim 1, wherein determining whether the event justifies a controlled takeover includes confirming whether the pilot is incapacitated by remote visual observation. 11. The method of any of claim 1, wherein determining whether the event necessitates a controlled takeover includes confirming whether the pilot is incapacitated by illuminating the pilot with a laser and analyzing the reflected light. 12. The method of any of claim 1, wherein asserting a controlled takeover of the aircraft from the pilot includes activating an autonomous flight control of the aircraft. 13. The method of any of claim 1, wherein asserting a controlled takeover of the aircraft from the pilot involves asserting remote flight control of the aircraft. 14. The method of any of claim 1, wherein executing a controlled landing of the aircraft without the assistance of the pilot includes flying the aircraft to a predetermined landing location. 15. The method of any of claim 1, wherein t... | Sikorsky Aircraft Corp. | G06G007/70 G05D001/00 G08G005/02 G08G005/00 G05D001/10 A61B005/00 A61B005/18 A61B005/024 A61B005/021 A61B005/11
US09788723 | B2 | 2017-10-17 | Hand-held medical-data capture-device having determination of a temperature by a microprocessor from a signal from a digital infrared sensor and having interoperation with electronic medical record systems on a specific segment of a network to transmit the temperature and device information | 2014-12-31 | In one implementation, an apparatus estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point is described, estimates temperature from a digital infrared sensor and determines vital signs from a solid-state image transducer, or determines vital signs from a solid-state image transducer and estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point; after which the estimated and/or determined information is transmitted to an external database. | 1. A non-touch thermometer to measure a temperature, the non-touch thermometer comprising: a microprocessor;a wireless communication subsystem that is operably coupled to the microprocessor and that is operable to transmit a representation of the temperature;a battery operably coupled to the microprocessor;a single button operably coupled to the microprocessor; anda digital infrared sensor operably coupled to the microprocessor with no analog-to-digital converter operably coupled between the digital infrared sensor and the microprocessor, the digital infrared sensor having only digital readout ports, the digital infrared sensor comprising a low noise amplifier, an analog-to-digital converter, a digital signal processor, and a pulse width modulator (PWM) in a single industry standard package; wherein the digital infrared sensor is configured with no need for recalibration after an initial factory calibration with an emissivity range from 0.1 to 1.0,wherein the microprocessor is operable to receive from the digital readout ports a digital signal that is representative of an infrared signal detected by the digital infrared sensor and the microprocessor is operable to determine the temperature from the received digital signal by using the PWM. 2. The non-touch thermometer of claim 1, wherein a display device further comprises: a green traffic light operable to indicate that the temperature is good;an amber traffic light operable to indicate that the temperature is low; anda red traffic light operable to indicate that the temperature is high. 3. The non-touch thermometer of claim 1, the microprocessor further comprising: a camera operably coupled to the microprocessor and providing at least two images to the microprocessor; andthe microprocessor including a cropper that is operable to receive the at least two images and operable to crop the images to exclude a border area of the images, generating at least two cropped images,the microprocessor also including a temporal-variation-amplifier of at least two cropped images that is operable to generate a temporal variation, the microprocessor also including a vital-sign generator that is operably coupled to the temporal-variation-amplifier that is operable to generate at least one vital sign from the temporal variation and the microprocessor also operably coupled to a display device that is operable to display the at least one vital sign. 4. The non-touch thermometer of claim 3, wherein the temporal-variation-amplifier further comprises: skin-pixel-identifier. 5. The non-touch thermometer of claim 3, wherein the temporal-variation-amplifier further comprises: a regional facial clusterial module. 6. The non-touch thermometer of claim 3, wherein the temporal-variation-amplifier further comprises: a first frequency filter. 7. The non-touch thermometer of claim 1 further comprising: the digital infrared sensor having no analog sensor readout ports. 8. The non-touch thermometer of claim 1 wherein the wir... | Arc Devices, LTD | G08B023/00 A61B005/00 A61B005/01 G06F019/00 A61B005/021 A61B005/026 A61B005/08 A61B005/024 A61B005/0205 G06T003/40 G06T007/00 H04W048/16 H04W076/02 A61B005/0402 A61B005/03 G01J005/00 H04W080/04 G01J005/02 G01J005/10 G08C017/02 H04L029/08 H04W012/06 H04B007/26 H04L012/46 H04N005/33 H04W048/08 G08B005/36 G08B021/18 G01K013/00 G06K009/62 G06T007/174 G06T007/269 G06T007/90 A61B005/0245 H04W012/08
US09788799 | B2 | 2017-10-17 | Medical procedure monitoring system | 2013-07-16 | A system and method for monitoring a medical procedure performed in a clinical environment is provided. An audio recorder is configured to produce a verbal data signal that is representative of a verbal communication occurring in the clinical environment, and a data analyzer is configured to detect an adverse condition based upon the verbal data signal. An alert module is configured to alert an operator upon the detection of an adverse condition. | 1. A system for monitoring a medical procedure performed in a clinical environment comprising: an input module for receiving a checklist comprising a proper protocol for the medical procedure and adapted to be customized by an operator;an audio receiver constructed and arranged to produce a verbal data signal representative of verbal communication that occurs in the clinical environment;a data analyzer constructed and arranged to (i) receive the verbal data signal from the audio receiver, (ii) analyze the verbal data signal to determine real-time events, (iii) compare the real-time events to the checklist, and (iv) detect at least one adverse condition when a real-time event falls outside of the proper protocol; andan alert module constructed and arranged to alert at least one operator when the at least one adverse condition is detected by the data analyzer. 2. The system of claim 1 wherein the medical procedure comprises at least one of: a surgical procedure; a minimally invasive surgical procedure; a laparoscopic surgical procedure; an open surgical procedure; an interventional procedure; a reconstructive surgery; a robotic procedure; a robotically-enabled procedure; an outpatient procedure; a dental procedures; a fully anesthetized dental procedure; and combinations thereof. 3. The system of claim 1 wherein the clinical environment comprises at least one of: an operating room; a catheterization lab; an intensive care unit; a control room for an operating room; an outpatient surgery treatment room; a dentist's office; a surgeon's office; a maxillofacial surgeon's office; and combinations thereof. 4. The system of claim 1 wherein the audio receiver comprises at least one microphone. 5. The system of claim 1 wherein the audio receiver comprises multiple microphones. 6. The system of claim 5 wherein the multiple microphones are in proximity to the at least one operator. 7. The system of claim 5 wherein the multiple microphones comprise a first microphone positioned in proximity to a first operator and a second microphone in proximity to a second operator. 8. The system of claim 1 wherein the audio receiver comprises at least one of: a microphone; an operator-worn microphone; an operator-worn headset; a room microphone; an omnidirectional microphone; a Bluetooth device; a telephone; a mobile telephone; a wireless device; and combinations thereof. 9. The system of claim 1 wherein the audio receiver comprises an operator-worn headset. 10. The system of claim 9 wherein each operator of the at least one operator wears a headset. 11. The system of claim 10 wherein each headset comprises an identifier associated with a respective operator. 12. The system of claim 1 wherein the audio receiver comprises an intercom system. 13. The system of claim 1 wherein the audio receiver is further constructed and arranged to produce a non-verbal data signal. 14. The system of claim 13 wherein the non-verbal data signal comprises audio signals produced by ... | Surgical Safety Solutions, LLC | A61B005/00 G06F019/00 A61B017/00 A61C019/00 A61B090/00 A61B034/30 A61B090/90 A61B090/50
US09820676 | B2 | 2017-11-21 | Method and positionable patient-interface apparatus for an MRI system | 2014-11-19 | Apparatus and method for imaging a patient in an MRI system. This includes a frame, and at least one assembly that includes a patient-interface positioner connected to a reference position on the frame, a first lockable joint on the positioner; and a patient interface connected to a patient-proximal end of the positioner by a second joint, wherein the first patient-interface is moveably positioned to a selected pitch angle, a selected yaw angle, and a selected one of a plurality of distances relative to the reference position on the frame. The first lockable joint is configured to be tightened to yieldably hold the first patient-interface at the selected pitch and yaw angles, and at the selected one of the plurality of distances, relative to the reference position. Optionally a second substantially similar patient-interface and assembly are provided. The earpiece(s) optionally include audio transducer(s) and/or RF coil(s). | 1. An apparatus comprising: a frame;a first positionable patient-interface assembly, the first patient-interface assembly including: a positioning rod, wherein the positioning rod has a proximal end and a distal end;a housing;a first lockable joint assembly connected to the frame, wherein the first lockable joint assembly is configured to be loosened and tightened;a second joint, wherein the positioning rod is configured to be connected to the frame by the first lockable joint assembly, wherein the housing is connected to the proximal end of the rod by the second joint, wherein the rod passes through the first lockable joint assembly and is configured to be moveably positioned to a selected one of a plurality of possible pitch-and-yaw-angle combinations of the rod relative to the frame and to a selected one of a plurality of insertion depths relative to the frame, and wherein the first lockable joint assembly is configured to be tightened to yieldably hold the rod at the selected pitch angle, yaw angle and insertion depth. 2. The apparatus of claim 1, wherein the first positionable patient-interface assembly is MRI compatible, wherein the frame includes a coil frame with an RF coil; and wherein the apparatus further comprises: a second MRI-compatible positionable patient-interface assembly, wherein the second patient-interface assembly is substantially similar to the first patient-interface assembly but mirror-symmetric to the first patient-interface assembly about a central plane of the coil frame, and wherein the first and the second MRI-compatible patient-interface assemblies are attached to the coil frame by their respective first joints. 3. The apparatus of claim 1, wherein the first positionable patient-interface assembly includes a fluid circulation passageway configured to circulate a fluid within the first positionable patient-interface assembly to control a temperature of at least a portion of a patient. 4. The apparatus of claim 1, wherein the first positionable patient-interface assembly is MRI compatible, wherein the frame includes a coil frame with an RF coil; and wherein the RF coil is configured to both transmit and receive RF MRI signals. 5. The apparatus of claim 1, wherein the first patient-interface assembly includes a motor interfaced to the positioning rod to controllably extend and withdraw an end of the first patient-interface assembly. 6. The apparatus of claim 1, wherein the first patient-interface assembly includes one or more skin-contact electrodes configured to receive electrical signals from a tissue of a patient and conduct the signals through the positioning rod. 7. The apparatus of claim 1, wherein the first patient-interface assembly includes one or more liquid-delivery tubes configured to deliver a liquid to a patient through the positioning rod. 8. The apparatus of claim 1, wherein the first patient-interface assembly includes one or more gas-delivery tubes configured to deliver a gas to a patient throu... | Life Services, LLC | A61B005/055 G01R033/34 A61J015/00 A61B005/053 A61M016/00 G01R033/28 A61M021/02 A61B005/00 A61M021/00
US09824286 | B2 | 2017-11-21 | Method and apparatus for early detection of dynamic attentive states for providing an inattentive warning | 2016-02-24 | A method and apparatus for determining an inattentive state of an operator of a vehicle and for providing information to the operator of the vehicle by obtaining face images of the operator of the vehicle, obtaining images of an environment of the vehicle, determining one or more areas of interest in the environment of the vehicle based on the images of the environment, obtaining, from a relevance and priority database, relevance and priority values corresponding to the one or more areas of interest, determining a probability of attention of the operator of the vehicle to the one or more areas of interest based on the images of the environment and the relevance and priority values, determining an attention deficiency based on the determined probability of attention and the face images, and providing the information to the operator of the vehicle based on the determined attention deficiency. | 1. A method comprising: determining a probability of attention of an operator of a vehicle to one or more areas of interest in an environment of a vehicle based on relevance and priority values corresponding to the one or more areas of interest;determining an attention deficiency based on the determined probability of attention and eye movements of the operator of the vehicle; andproviding information to a warning or guidance device based on the determined attention deficiency. 2. The method of claim 1, wherein the one or more areas of interest includes at least one of a pedestrian, a cyclist, a traffic light, a traffic sign, an animal, an object, or another vehicle. 3. The method of claim 1, wherein the step of determining the probability of attention includes: determining linear distances of the one or more areas of interest from the vehicle;determining saliencies of the one or more areas of interest;detecting event frequencies within the one or more areas of interest; anddetermining the probability of attention of the operator of the vehicle to the one or more areas of interest based on the linear distances, the saliencies, the event frequencies, and the relevance and priority values. 4. The method of claim 3, wherein the step of determining the linear distances includes: setting a linear distance of an area of interest to zero when an angle between a line of sight of the operator of the vehicle and a line of sight corresponding to the area of interest is greater than or equal to 90 degrees; andsetting the linear distance of the area of interest to a distance between the area of interest and the vehicle divided by a cosine of the angle between the line of sight of the operator of the vehicle and the line of sight corresponding to the area of interest when said angle is smaller than 90 degrees. 5. The method of claim 1, wherein the step of determining the attention deficiency includes: detecting facial features of the operator of the vehicle by feature point extraction algorithms based on face images, the facial features including one or more of eye corners, upper eyelids, or lower eyelids;determining an iris location based on the detected facial features;determining gaze pointers based on the iris location and images of the environment of the vehicle;determining gaze fixation dwell times based on the gaze pointers;determining an observed attention level based on gaze fixation dwell times; anddetermining the attention deficiency based on the observed attention level and the probability of attention. 6. The method of claim 1, wherein the step of providing the information includes: providing, based on the determined attention deficiency, one or more of a failed-to-look warning, a look-but-fail-to-see warning, an audio/visual warning, and guidance information. 7. The method of claim 1, further comprising determining the one or more areas of interest in the environment of the vehicle based on environmental images of the environment of the veh... | TOYOTA JIDOSHA KABUSHIKI KAISHA CENTRAL MICHIGAN UNIVERSITY | G08B023/00 G06K009/00 A61B005/18 B60K028/06 G08B021/04 A61B005/00 A61B003/113 A61B005/11 G08G001/16
US09788777 | B1 | 2017-10-17 | Methods and apparatus to identify a mood of media | 2014-08-12 | Methods and apparatus to identify a mood of media are disclosed. An example method includes accessing a known audio known to evoke a first emotion. A first synthesized sample is synthesized based on the known audio sample. A first value of a first feature of the first synthesized sample is calculated. A mood model is created based on the first feature. The mood model is to establish a relationship between the first feature and the first emotion. A second value of the first feature of first media evoking an unknown emotion is identified. The unknown emotion is identified as the first emotion when the mood model indicates that the second value corresponds to the first value. | 1. A method to identify an emotion evoked by media, the method comprising: accessing a pre-verbal utterance known to evoke a first emotion;creating, using a musical instrument digital interface notator, a musical instrument digital interface representation of the pre-verbal utterance;synthesizing, using a digital musical instrument, a first synthesized sample based on the musical instrument digital interface representation of the pre-verbal utterance;calculating, with a processor, a first value of a first feature of the first synthesized sample;creating, with the processor, a mood model based on the first feature, the mood model to establish a relationship between the first value of the first feature and the first emotion;identifying a second value of the first feature of first media evoking an unknown emotion; andidentifying the unknown emotion as the first emotion when the mood model indicates that the second value corresponds to the first value. 2. The method as described in claim 1, wherein the first feature includes at least one of a number of zero crossings, a rolloff power, a brightness, a roughness, a presence of a minor third interval, a presence of a major third interval, an irregularity, a chroma, a main pitch, or a key. 3. The method as described in claim 1, further including: identifying a request for media evoking the first emotion; andrecommending second media as evoking the first emotion. 4. The method as described in claim 3, wherein the second media is an advertisement. 5. The method as described in claim 1, further including: selecting a first reference value of the first feature of the first emotion;calculating a first difference between the first reference value and the second value of the first feature of the first media; andapplying a first weighting value to the first difference to form a first feature score. 6. The method as described in claim 5, further including identifying the unknown emotion as the first emotion when the first feature score exceeds a threshold. 7. The method as described in claim 5, further including: identifying a second reference value of a second feature of the first emotion, the second feature different from the first feature;calculating a third value of the second feature of the first media;calculating a second difference between the second reference value and the third value; andapplying a second weighting value to the second difference to form a second feature score. 8. The method as described in claim 1, further including: selecting an advertisement that evokes the first emotion evoked by the first media; andpresenting the advertisement in connection with the first media. 9. The method as described in claim 8, wherein the advertisement includes audio and visual components, the method further including altering the audio component of the advertisement to more strongly evoke the first emotion. 10. The method as described in claim 1, further including: identifying a brand associated with... | The Neilsen Company (US), LLC | G06Q099/00 A61B005/16 G06Q030/02
US09795455 | B2 | 2017-10-24 | Tissue marker for multimodality radiographic imaging | 2014-08-22 | A radiographic marker containing a sensing medium that initially includes a gas and liquid mixture for producing increased signal intensity in a first imaging modality. The radiographic marker can have a permeable portion allowing the exchange of the gases or liquids of the sensing medium to be changed or exchanged for at least a second fluid in the surrounding tissue to reduce the bubbles or air pockets created by the gas portion of the sensing medium. The change or exchange of gas for the second liquid in the radiographic marker produces a detectable signal intensity of the radiographic marker in a second imaging modality different from the first imaging modality. | 1. A radiographic marker, comprising: a container wall enclosing an internal chamber and including a permeable portion; anda sensing medium received within the internal chamber, the sensing medium initially including at least one of a first liquid, a gas portion and a solid portion, the sensing medium configured to be detectable using a first imaging modality;wherein the permeable portion is semi-permeable to permit, when injected or surgically implanted into a patient's body, transfer of enough of a second liquid through the permeable portion of the chamber wall into the inner chamber while confining the sensing medium within the inner chamber to change a ratio of gas to liquid of the sensing medium within the internal chamber to produce a detectable change in signal intensity using a second imaging modality that is of a different type than the first imaging modality. 2. The radiographic marker of claim 1, wherein the gas portion of the sensing medium creates at least one bubble within the internal chamber to produce a detectable signal intensity in a first imaging modality. 3. The radiographic marker of claim 2, wherein the first imaging modality comprises ultrasound imaging. 4. The radiographic marker of claim 1, wherein transfer of the second liquid across the chamber wall produces a detectable signal intensity in a second imaging modality. 5. The radiographic marker of claim 4, wherein the second imaging modality comprises magnetic resonance imaging. 6. The radiographic marker of claim 1, wherein the sensing medium further includes an osmotic agent to create an osmotic gradient drawing the second fluid into the internal chamber through the permeable portion of the chamber walls. 7. The radiographic marker of claim 1, the sensing medium includes at least one contrast agent. 8. The radiographic marker of claim 7, wherein the permeable portion of the container wall retains the contrast agent within the internal chamber as the second fluid and the gas portion are exchanged through the container wall. 9. The radiographic marker of claim 1, wherein the gas portion is dispersed within the first liquid. 10. The radiographic marker of claim 1, wherein the sensing medium includes at least one of a solid and a gel. 11. The radiographic marker of claim 1, further comprising at least one object positioned within the internal chamber; wherein the object comprises an artificial shape identifiable under an imaging modality. 12. The radiographic marker of claim 11, wherein the at least one object positioned is configured to transmit an electromagnetic signal. 13. The radiographic marker of claim 1, wherein the permeable portion is a semi-permeable material. 14. The radiographic marker of claim 1, wherein the container wall comprises an impermeable portion. 15. The radiographic marker of claim 1, wherein the impermeable material is selected from at least one of glass, ceramics, polymers, metals and composites. 16. The radiographic marker of cla... | Breast-Med, Inc. | A61B005/05 A61B090/00 A61B019/00 A61B017/00
US09795301 | B2 | 2017-10-24 | Apparatus, systems, methods and computer-accessible medium for spectral analysis of optical coherence tomography images | 2011-05-25 | According to an exemplary embodiment of the present disclosure, apparatus and method can be provided for generating information for at least one structure. For example, using at least one first arrangement, it is possible to receive at least one first radiation from the at least one structure and at least one second radiation from a reference, and interfere the first and second radiations to generate at least one third radiation. Further, with at least one second arrangement, it is possible to generate spectroscopic data as a function of the at least one third radiation, and reduce at least one scattering effect in the spectroscopic data to generate the information. In addition or as an alternative, according to a further exemplary embodiment of the present disclosure, it is possible to classify a type of the structure based on the spectroscopic data to generate the information. | 1. An apparatus for generating information for at least one structure, comprising: at least one interferometer first arrangement which is configured to receive at least one single-spatial mode first radiation from the at least one structure and at least one second radiation from a reference, and combine only one of the first radiation with the at least one second radiation to generate a single third radiation;at least one computer second arrangement which is configured to (i) generate spectroscopic data as a function of the single third radiation; and (ii) reduce at least one scattering effect in the spectroscopic data using only the single third radiation to generate the information. 2. The apparatus according to claim 1, wherein the at least one second arrangement is further configured to reduce an intensity variation of the of the spectroscopic data. 3. The apparatus according to claim 1, wherein the at least one second arrangement is further configured to smooth the spectroscopic data based a polynomial fit to the spectroscopic data. 4. The apparatus according to claim 3, wherein the at least one structure is an anatomical structure. 5. The apparatus according to claim 4, wherein the anatomical structure is an artery wall. 6. The apparatus according to claim 1, wherein the at least one second arrangement is further configured to classify a type of the structure based on the spectroscopic data. 7. A method for generating information for at least one structure, comprising: receiving at least one single-spatial mode first radiation from the at least one structure and at least one second radiation from a reference;causing an interference of only one of the first radiation with the at least one second radiation to generate a single third radiation;generating spectroscopic data as a function of the single third radiation; andwith a computing arrangement, causing a reduction at least one scattering effect in the spectroscopic data using only the single third radiation to generate the information. 8. An apparatus for generating information for at least one structure, comprising: at least one interferometer first arrangement which is configured to receive at least one first radiation from the at least one structure and at least one second radiation from a reference, and interfere the first and second radiations to generate at least one third radiation; andat least one computer second arrangement which is configured to (i) generate spectroscopic data as a function of only the at least one third radiation, and (ii) classify a type of the structure based on the spectroscopic data using a statistical pattern recognition procedure to generate the information. 9. The apparatus according to claim 8, wherein the classification includes a generation of a probability that the structure is of a particular type. 10. The apparatus according to claim 9, wherein the probability is provided for a depth-resolved two-dimensional map. 11. The arrangement accor... | The General Hospital Corporation | G01B009/02 G01J003/45 A61B005/00 A61B005/02 G01N021/47 G01N021/17
US09814414 | B2 | 2017-11-14 | Transcutaneous analyte sensor | 2010-03-30 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A method for inserting a transcutaneous analyte sensor into a host using an applicator, the method comprising the steps of: adhering a base of a mounting unit to a skin of a host before insertion of a transcutaneous sensor into the host;pushing on a first component of an applicator for inserting the transcutaneous analyte sensor into the host, so as to insert an introducer and the sensor into the skin of the host by using a first springless force, wherein the pushing comprises pushing on a plunger; andpulling on a second component of the applicator so as to retract the introducer from the skin of the host by using a second springless force, while leaving the sensor in the host;wherein pushing is performed by a user, wherein the first component and the second component of the applicator are configured to interact so as to enable insertion of the sensor and retraction of the introducer. 2. The method of claim 1, wherein the pulling step comprises pulling on a body portion of the second component of the applicator. 3. The method of claim 1, wherein the first springless force or the second springless force is provided by a user's hand. 4. The method of claim 1, wherein the applicator is configured to insert the sensor and retract the introducer responsive to a squeezing motion between a user's thumb and fingers. 5. The method of claim 1, wherein the applicator is configured to sequentially insert the sensor and retract the introducer responsive to a squeezing motion between a user's thumb and fingers. 6. The method of claim 1, wherein the user is the host. 7. An applicator for inserting a transcutaneous analyte sensor into a host, comprising: a first component comprising a plunger; anda second component;wherein the applicator is configured and arranged such that pushing on the first component by a user inserts an introducer and a sensor into a host's skin after a base of a mounting unit has been adhered to the host's skin and pulling on the second component by a user retracts the introducer from the skin while leaving the sensor in the host, wherein pushing on the first component comprises pushing on the plunger, and wherein the first component and the second component are configured to interact so as to enable insertion of the sensor and retraction of the introducer; andwherein the applicator is configured to insert the sensor by a first springless force directly provided by a user and retract the introducer substantially by a second springless force directly provided by the user. 8. The applicator of claim 7, wherein the second component comprises a body portion, and wherein pulling on the second component comprises pulling on the body portion of the second component. 9. The applicator of claim 7, wherein the first springless force and the second springless force is provided by a user's hand. 10. The applicator of claim 7, wherein the applicator is configured to insert the sensor and retract the introducer responsive to a squeezing moti... | DexCom, Inc. | A61B005/00 A61B005/145 A61B005/1473 A61B005/15 A61B005/1495 A61B005/05 A61B017/34 A61M005/14 A61B005/1486 A61M005/172 A61M005/158 A61M005/142
US09820680 | B2 | 2017-11-21 | System and method for determining sleep and sleep stages of a person | 2012-05-15 | The present invention relates to a system and method for determining sleep, sleep stage and/or sleep stage transition of a person, including heart rate detecting means configured for detecting a heart rate of the person, movement detecting means configured for detecting a movement of a part of the body of the person, where the detected movement is caused by a skeletal muscle of the body, recording means configured for recording the detected heart rate and the detected movement of the part of the body, heart rate classifying means configured for classifying the recorded heart rate of the person into at least one heart rate class at least one heart rate variability class, movement classifying means configured for classifying the recorded movement into at least one movement class, and determining means configured for determining sleep, a sleep stage, a sleep stage transition and/or a sleep event of the person based at least partially on the at least one heart rate class and the at least one movement class. | 1. A system for determining sleep, a sleep stage, sleep stage transition, and/or a sleep event of a person, comprising: a wearable device configured for detecting and recording a heart rate of the person and configured for detecting and recording a movement of a part of the body of the person, wherein the movement is caused by a skeletal muscle of the body;an analysis device configured for classifying the recorded heart rate of the person into at least one heart rate class and at least one heart rate variability class, configured for classifying the recorded movement into a movement class comprising a large movement (LM), a small movement (SM) or a twitch (TM), based on a calculated intensity and/or duration of each movement, wherein the analysis device is capable of classifying the movement into any of the LM, SM, and TM, and configured for determining sleep, a sleep stage, a sleep stage transition and/or a sleep event of the person based at least partially on the at least one heart rate class, the at least one heart rate variability class, and the movement class; anda data connection configured for communicating data representing the recorded heart rate and the recorded movement from the wearable recording device to the analysis device, and wherein the analysis device is configured for classifying the recorded movement into the movement class based on the calculated intensity level of 1 to 2 counts per second, 3 to 5 counts per second, or 6 to 10 counts per second. 2. A system for determining sleep, a sleep stage, sleep stage transition, and/or a sleep event of a person, the system comprising: heart rate detecting means configured for detecting a heart rate of the person;movement detecting means configured for detecting a movement of a part of the body of the person caused by a skeletal muscle of the body and comprising movement sensing means configured for sensing an acceleration of the part of the body;recording means configured for recording the detected heart rate and the detected movement of the part of the body;movement calculating means configured for calculating, based on values of the sensed acceleration, at least an intensity and/or a duration of each movement of the part of the body;heart rate classifying means configured for classifying the recorded heart rate of the person into at least one heart rate class and at least one heart rate variability class;movement classifying means configured for classifying the recorded movement into a movement class comprising a large movement (LM), a small movement (SM), or a twitch movement (TM), based on the calculated intensity and/or the calculated duration of each movement, wherein the movement classifying means is capable of classifying the movement into any of the LM, SM, and TM; anddetermining means configured for determining sleep, a sleep stage, a sleep stage transition and/or a sleep event of the person based at least partially on the at least one heart rate class, the at least one hear... | V-WATCH SA | A61B005/0205 A61B005/00 A61B005/024 A61B005/11 A61B005/021 A61B005/145
US09826908 | B2 | 2017-11-28 | Devices, systems and methods for testing cardiac exercise functions | 2016-04-28 | The present disclosure discloses a device, system and method for testing cardiac exercise functions. Chronotropic Competence Indices (CCIs) are proposed to quantitatively describe the adaptation capability of cardiopulmonary system in response to exercise intensity variation in terms of heart rate changes, and thereby describes the dynamic process of the heart in the body metabolic process. The present disclosure discloses a device which measures the CCIs in real time by using the wearable technology, and referred to as Cardiac Exercise Test (CET). Compared with the Cardiopulmonary Exercise Testing (CPX) and parameters measured by the CPX such as a maximum oxygen uptake, the CCIs have clear clinical meanings and specific normal reference values; and the CET reduces the risk of the test. It is simple to use, and can be used anytime anywhere. It is of great importance in wide clinical applications, and is of great significance in prevention and rehabilitation of cardiopulmonary disease. | 1. A system for testing cardiac exercise functions, comprising: at least one wearable component worn by a testee, comprising an ECG acquisition module configured to acquire ECG signal of the testee, and an exercise acquisition module configured to acquire an exercise signal of the testee without or with a load;a handhold electronic device configured to acquire the ECG signal and the exercise signal of the testee from the at least one wearable component in a wired or wireless communication manner, process the ECG signal and the exercise signal, and control a process of cardiac exercise testing; anda workstation connected to the handhold electronic device in a wireless or wired communication manner, configured to receive data from the handhold electronic device, complete the process of cardiac exercise testing under the monitoring of a user, and generate a test report comprising at least ECG parameters, exercise parameters and Chronotropic Competence Indices (CCIs). 2. The system according to claim 1, wherein the handhold electronic device is configured to process the ECG signal and the exercise signal in real time in the following three periods of resting for at least 1 minute before an exercise test, exercising for 6 minutes or more in a certain mode after the exercise test is started, and resting for 1 minute or more after the exercise is stopped. 3. The system according to claim 1, wherein the CCIs comprise Resting Heart Rate (RHR),Chronotropic Limit (CL), Chronotropic Rate (CR), and Heartrate Recovery after Exercise (HRE). 4. The system according to claim 3, wherein the CR is measured according to the following equation: CR=(HRstage−HRrest)/(METstage−1)wherein HRstage and METstage represent a heart rate and an exercise metabolic equivalent of the testee at a certain time during the test respectively, and HRrest represents a heart rate of the testee during resting. 5. The system according to claim 3, wherein the chronotropic limits measured according to the following equation: CL=(HRmax−HRrest)/(HRPredM−HRrest)wherein HRmax and HRrest represent the maximum heart rate and Resting Heart Rate (RHR) respectively, and HRPredM represents a predicted value of the maximum heart rate. 6. The system according to claim 3, wherein the HRE is defined as the difference between the maximum exercise heart rate and the heart rate 1 minute after the exercise is stopped, which represents a speed at which the heart rate recovers to the RHR after the exercise is stopped. 7. The system according to claim 1, wherein the handhold electronic device comprises a man-machine interaction module configured to alert the testee of start or stop of the exercise or acceleration or deceleration of the exercise in a voice or image manner, and display information, an exercise mode, and an ECG signal of the testee in real time, and configured with an emergency button for use by the testee when feeling uncomfortable. 8. The system according to claim 1, wherein the ECG acquisi... | SmartHealth Electronics Ltd. | A61B005/04 A61B005/0205 A61B005/00 A61B005/0245 A61B005/0404 A61B005/044 A61B005/11 A61B005/22 G06F019/00
US09814384 | B2 | 2017-11-14 | High temporal resolution doppler OCT imaging of retinal blood flow | 2014-09-26 | Techniques are introduced to improve the ability of OCT to determine more accurately the nature of the flow of fluids in the eye, including faster measurements of the flow and a method to reduce geometric uncertainties due to eye movements. | 1. A method of determining blood flow in the vessels in a retina of an eye based on measurements obtained from an optical coherence tomography (OCT) system, said system including a scanner for scanning a measurement beam across the eye, said method comprising the steps of: generating an image of the retina;determining the location and angular orientation of blood vessels within the image and identifying sections of the vessels having an angular orientation with respect to the scanning beam where the velocity of moving scatterers inside the vessels is best measured;obtaining a set of Doppler OCT measurements by scanning the beam over the identified sections of the blood vessels;determining the velocity of moving scatterers in the vessels based on the Doppler OCT measurements; anddisplaying or storing information of the determined velocity. 2. A method as recited in claim 1, wherein the determined velocity information is used to calculate the blood flow in the vessels. 3. A method as recited in claim 1, wherein the image of the retina is generated with a fundus camera. 4. A method as recited in claim 1, wherein the image of the retina is generated by measurements made by the OCT system by scanning the beam over the retina. 5. A method as recited in claim 4, wherein the spatial density of the Doppler OCT measurements is higher than the spatial density of the measurements used to generate the image of the retina. 6. A method as recited in claim 4, wherein the spatial density of the Doppler OCT measurements is higher across the blood vessels than the spatial density of Doppler OCT measurements in between the blood vessels. 7. A method as recited in claim 1, wherein during the step of obtaining a set of Doppler OCT measurements by scanning the beam in the identified sections of the blood vessels, measurements are obtained to determine any changes to the orientation of the retina, said determination being used to update the previously determined location and orientation of the blood vessels. 8. A method as recited in claim 1, wherein the step of obtaining a set of Doppler OCT measurements includes obtaining temporally separated phase measurements at the same location in the retina. 9. A method as recited in claim 8, wherein the step of determining the velocity of moving scatterers in the vessels includes comparing the difference in the phase measurements obtained at the same location in the retina. 10. A method as recited in claim 1, wherein the Doppler OCT measurements are based on one of the following methods: phase difference Doppler OCT, phase-resolved Doppler variance, intensity-based Doppler variance, joint spectral and time domain OCT, optical microangiograpy, Doppler optical microangiography, bidirectional Doppler OCT, three-beam Doppler OCT. 11. A method as recited in claim 1, wherein the identified sections of the blood vessel have an angular orientation lower than a predetermined high limit and higher than a predetermined low limit.... | CARL ZEISS MEDITEC, INC. | A61B003/12 A61B003/00 A61B003/10 A61B003/14 A61B005/00 A61B005/026 G01B009/02
US09792801 | B2 | 2017-10-17 | Enhancing vehicle system control | 2016-08-04 | A method for alertness control of a vehicle operator includes obtaining, by an apparatus, physiological status data of a person acquired using at least one sensor including at least a biosignal sensor; determining, based on at least the physiological status data, at least one alertness value being indicative of alertness level of the person; determining, by comparing the at least one alertness value to at least one alertness reference value, whether the alertness level of the person is below a threshold alertness level for operating the vehicle; and as a response to the determining that the alertness level is below the threshold alertness level, causing an output of a control signal. | 1. An apparatus, configured to be attached to a wrist of a person in a vehicle system, comprising at least one processor and at least one memory including a computer program code that, when executed by the at least one processor and when it is determined that the person has initiated driving a vehicle, causes the apparatus to at least: obtain, using at least one sensor comprising at least a bio-signal sensor, physiological status data of the person driving the vehicle, wherein the at least one sensor comprises a heart activity sensor, and wherein the physiological status data comprises cardiac activity data of the person, wherein the cardiac activity data comprises cardiac activity history data of the person;determine, based on the cardiac activity history data, a cardiac activity of the person related to a first threshold alertness level for driving the vehicle;determine, based on real-time cardiac activity data, current cardiac activity of the person being indicative of a current alertness level of the person;determine, by comparing the current cardiac activity of the person to the cardiac activity of the person related to the first threshold alertness level, whether the current alertness level of the person is below the first threshold alertness level for driving the vehicle; andcause a wireless output of a control signal as a response to the determination that the current alertness level is below the first threshold alertness level, wherein the control signal is at least for the vehicle system to take control of driving operation of the vehicle. 2. The apparatus of claim 1, wherein the physiological status data further comprises sleep history data of the person. 3. The apparatus of claim 2, wherein the sleep history data comprises at least one of circadian rhythm data, a resting time during a time period, a sleep time during the time period, data being indicative of sleep quality during the time period. 4. The apparatus of claim 1, wherein the physiological status data further comprises exercise history data of the person. 5. The apparatus of claim 1, wherein the physiological status data further comprises personal characteristics data of the person. 6. The apparatus of claim 1, wherein the physiological status data further comprises respiration rate data of the person. 7. The apparatus of claim 1, wherein the at least one memory and the computer program code are configured, with the at least one processor, to cause the apparatus further to at least: acquire data related to the driving operation of the vehicle from the vehicle system, wherein determining the at least one alertness value being indicative of the alertness level of the person is further based on the acquired data related to the driving operation of the vehicle. 8. The apparatus of claim 1, wherein the at least one memory and the computer program code are configured, with the at least one processor, to cause the apparatus further to at least: transfer wirelessly data with ... | POLAR ELECTRO OY | G08B023/00 G08B021/04 A61B005/00 A61B005/18 B60L003/02 B60K028/02 A61B005/0205 A61B005/024 A61B005/0295 A61B005/04 A61B005/0402 A61B005/08 A61B005/11 G08B021/06
US09791341 | B2 | 2017-10-17 | Pressure sensor, microphone, blood pressure sensor, and touch panel | 2014-08-27 | According to one embodiment, a pressure sensor includes a support, a film unit supported by the support, having an upper surface, and capable of being deformed, and a first sensing element provided on the upper surface. The first sensing element includes a first magnetic layer, a second magnetic layer provided apart from the first magnetic layer and a first intermediate unit including a first intermediate layer including a portion provided between the first and second magnetic layers. The first magnetic layer extends in a first direction parallel to the upper surface, and a first major axis length of the first magnetic layer is longer than a first minor axis length. The second magnetic layer extends in a second direction parallel to the upper surface and crossing the first direction, and a second major axis length of the second magnetic layer is longer than a second minor axis length. | 1. A pressure sensor comprising: a support;a film unit supported by the support, the film unit having an upper surface, and being deformable; anda first sensing element provided on the upper surface,the first sensing element including: a first magnetic layer, a magnetization of the first magnetic layer being configured to change in accordance with a deformation of the film unit;a second magnetic layer provided apart from the first magnetic layer in a direction crossing the upper surface; anda first intermediate unit including a first intermediate layer including a portion provided between the first magnetic layer and the second magnetic layer,the first magnetic layer extending in a first direction parallel to the upper surface, a first major axis length of the first magnetic layer in the first direction being longer than a first minor axis length of the first magnetic layer in a second direction parallel to the upper surface and crossing the first direction, andthe second magnetic layer extending in the second direction, a second major axis length of the second magnetic layer in the second direction being longer than a second minor axis length of the second magnetic layer in the first direction. 2. The sensor according to claim 1, wherein the first intermediate unit has a first intermediate unit region,the first magnetic layer has an overlapping region where the first magnetic layer overlaps the second magnetic layer, andan outer edge of the first intermediate unit region in a plane parallel to the upper surface is located inside an outer edge of the overlapping region in the plane. 3. The sensor according to claim 1, wherein the first intermediate unit has a first intermediate unit region,the first magnetic layer has an overlapping region where the first magnetic layer overlaps the second magnetic layer, anda shape of the first intermediate unit in a plane parallel to the upper surface is same as a shape of the overlapping region in the plane. 4. The sensor according to claim 1, wherein a shape of the first intermediate unit in a plane parallel to the upper surface is same as a shape of the first magnetic layer in the plane. 5. The sensor according to claim 1, wherein the first intermediate unit further includes a first intermediate magnetic layer provided between the first intermediate layer and the first magnetic layer. 6. The sensor according to claim 1, wherein the first intermediate unit further includes a second intermediate magnetic layer provided between the first intermediate layer and the second magnetic layer. 7. The sensor according to claim 1, wherein an angle between the first direction and the second direction is not less than 60 degrees and not more than 120 degrees. 8. The sensor according to claim 1, wherein an angle between the first direction and the second direction is larger than 0 degrees and smaller than 90 degrees. 9. The sensor according to claim 1, wherein an angle between the first direction and the second dir... | Kabushiki Kaisha Toshiba | G01L009/16 A61B005/021 G01L009/00 G06F003/041 H04R019/00 H04R019/04 H04R031/00 A61B005/00
US09788759 | B2 | 2017-10-17 | Method and system for postural analysis and measuring anatomical dimensions from a digital three-dimensional image on a mobile device | 2015-01-16 | A mobile, hand-held communication device with a digital display, and a 3D camera for acquiring a three-dimensional image of the human body is programmed to function as a digital anthropometer. The user digitizes anatomical landmarks on the displayed image to quickly obtain anatomical measurements which are used with a known morphological relationship to make an anatomical prediction for clothing measurement, body composition, orthotic and insert manufacturing, and postural displacement with accuracy and without external equipment. | 1. A method of deriving an anatomical prediction using a known morphological relationship and a programmed apparatus including a digital touch screen display and a 3D camera configured to acquire an image of a person on the digital touch screen display, the method comprising: acquiring at least one digital three-dimensional image of a person on the digital touch screen display;digitizing points on a plurality of anatomical landmarks on the displayed three-dimensional image;calculating a circumferential measurement of at least a portion of a body of a person in the displayed three-dimensional image using at least the digitized points on the displayed three-dimensional image; andmaking an anatomical prediction based on the calculated circumferential measurement and a known morphological relationship. 2. The method of claim 1, wherein the anatomical prediction is a clothing measurement. 3. The method of claim 2, wherein the clothing measurement is selected from the group consisting of neck, overarm, chest, waist, hips, sleeve and outseam. 4. The method of claim 1, wherein the anatomical prediction is a body composition. 5. The method of claim 1, further comprising activating the 3D camera to acquire the three-dimensional image of the person, wherein the programmed apparatus is a mobile, hand-held communication device having the 3D camera. 6. The method of claim 1, wherein the anatomical prediction is a volume of a body part. 7. The method of claim 1, wherein the points on the plurality of anatomical landmarks are digitized by touching the digital touch screen display of the image at respective anatomical landmarks. 8. The method of claim 1, wherein the digital three-dimensional image is a foot of the person and the method further comprises using the anatomical prediction for orthotic or insert manufacturing. 9. The method of claim 1, wherein the circumferential measurement of the at least the portion of the body of the person in the displayed three-dimensional image is calculated by adding a linear distance between the digitized points on the displayed three-dimensional image along a predetermined two-dimensional plane of the image. 10. The method of claim 1, including displaying the anatomical prediction on the digital touch screen display. 11. The method of claim 1, wherein the anatomical prediction is a postural displacement. 12. A method for making an anatomical measurement of a body of a person, the method comprising: providing a digital anthropometer on a mobile device;acquiring at least one digital three-dimensional image of a body of a person and displaying the digital three-dimensional image on the device; anddigitizing anatomical landmarks provided on the digital three-dimensional image displayed on the device for calculating a circumferential measurement of the body of the person. 13. A system for determining an anatomical measurement of a body of a person, the system comprising: a programmed device comprising a digital touch ... | - | A61B006/00 A61B005/107 A61B005/00 G01B011/24 A61B005/11
US09810543 | B2 | 2017-11-07 | Method and device for carrying out travel route planning for a vehicle | 2013-10-01 | A method and device for carrying out travel route planning for a vehicle, in which schedule data are transmitted to a calculation unit, which is coupled to a data memory, in which data relating to a path network for the vehicle and data relating to the geographic positions of energy supply facilities are stored. The calculation unit checks if a route sequence which connects the geographic positions of the destinations associated with the schedule data is calculable such that the destinations are reached at the associated time data of the schedule data, wherein the prognosticated residual energy amount in the energy store of the vehicle for travelling the route sequence is determined and taken into account, if no such route sequence is calculable using the calculation unit, the latter determines matched schedule data for which such a route sequence is calculable, and the matched schedule data are output. | 1. A method for carrying out travel route planning for a vehicle, the method comprising: transmitting schedule data, which comprise geographic positions of destinations of the travel route to be planned and associated time data, to a computing unit that is coupled to a data memory, in which data based on a road network for the vehicle and data based on geographic positions of energy supply facilities are stored;checking, via the computing unit, whether a route sequence that connects the geographic positions of the destinations associated with the schedule data is calculable, so that destinations are reached at the associated time data of the schedule data, wherein a predicted remaining amount of energy in an energy storage device of the vehicle for traveling the route sequence is determined and taken into account;storing further data on geographic positions of parking areas comprising parking lots or energy supply facilities for the vehicle for the destinations of the route sequence in the data memory;determining, via the computing unit in calculating a route sequence also for the destinations of a destination sequence case, parking areas assigned to the destinations at the particular destination, determining an assigned parking area for each destination of the destination sequence, wherein, in determining the distance of the parking area from an assigned destination, the geographic position of the next destination or the geographic position of the parking areas of the next destination and the predicted remaining amount of energy in the energy storage device for traveling the route sequence are taken into account, and wherein the route sequence comprises routes between parking areas of successive destinations of the destination sequence;checking whether a route sequence connects the geographic positions of the destinations associated with the schedule data so that the destinations are reached at the associated time data of the schedule data, wherein a length of time is taken into account that the user requires to reach the destination from the parking area assigned to the destination;determining, if no such route sequence is calculable by the computing unit, adjusted schedule data for which such a route sequence is calculable; andoutputting the adjusted schedule data,wherein, during said checking, a probability of parking area availability at said destinations is taken into consideration,wherein the energy supply facilities are located in the parking area at said destinations,wherein, during said checking, time data for the destinations and a length of time for increasing energy reserves while the vehicle is at the destination is taken into consideration,wherein said increasing energy reserves includes a possibility of ending charging of the energy storage device before the energy storage device is filled so that destinations are reached at the associated time data of the schedule data, andwherein, during said checking, driving behavior based on ... | Volkswagen Aktiegesellschaft | G01C021/34 G01C021/36 G06Q010/10 A61B005/18 A61B005/0205 A61B005/021 A61B005/024 A61B005/0402 A61B005/08 A61B005/145
US09788946 | B2 | 2017-10-17 | Devices for remodeling a valve annulus and ventricle | 2015-08-14 | Devices and related methods of use are provided for improving heart function. In one embodiment of the present disclosure, a device includes a ring-like structure configured to be secured to a heart valve; at least one elongate member extending from the ring-like structure, wherein an end of the elongate member is configured to be secured to heart geometry other than a heart valve; and an adjustment mechanism for simultaneously altering a dimension of the ring-like structure and a length of the elongate member. | 1. A device for treating a heart, comprising: an elongated curvilinear body comprising an annuloplasty ring having structure that facilitates attachment to the atrioventricular heart valve annulus, the body having a shape that conforms around at least a portion of an atrioventricular heart valve annulus and is configured to be secured thereto, the body defining an atrial side opposite a ventricular side;a flexible first elongated member engaged by and extending from an adjustment mechanism through a channel that extends within and around a length of the body and out of an opening in the body on the ventricular side thereof, wherein an end of the first elongate member is configured to be secured to heart geometry in the ventricle, the adjustment mechanism being configured to alter a length or tension of the first elongate member and thus alter a distance between the body and the end of the first elongate member that is secured to the heart geometry; anda flexible second elongated member engaged by and extending from the adjustment mechanism through the channel in the body and connected to a portion of the body, the adjustment mechanism being configured to alter a length or tension of the second elongate member and thus alter the shape or dimension of the body, and whereinwherein the adjustment mechanism includes a component configured for rotation and accessible from the atrial side of the body, wherein the first and second elongated members pass around a pulley located within the body and rotated by the component. 2. The device of claim 1, wherein the structure that facilitates attachment to the atrioventricular heart valve annulus comprises a covering around the body configured to promote tissue ingrowth. 3. The device of claim 1, wherein the end of the first elongate member configured to be secured to heart geometry in the ventricle includes a tissue anchor. 4. The device of claim 3, wherein the tissue anchor includes a loop configured to be sewn through the heart geometry in the ventricle. 5. The device of claim 1, wherein the adjustment mechanism comprises a first pulley and a second pulley that are mechanically connected so that rotation of one pulley results in rotation of the other pulley, and vice versa, and the first pulley has a diameter different than that of the second pulley, and wherein the flexible first elongated member extends around the first pulley and the flexible second elongated member extends around the second pulley. 6. The device of claim 1, further including a stiffener provided in the channel in the body and wherein there are two flexible second elongated members each engaged by and extending in opposite directions from the adjustment mechanism through the channel and connected to opposite ends of the stiffening member. 7. The device of claim 6, wherein the stiffener comprises a plurality of spaced-apart segments. 8. The device of claim 1, wherein there are two of the first elongated members each engaged by and e... | Edwards Lifesciences Corporation | A61F002/24 A61B005/107 A61B006/00
US09797880 | B2 | 2017-10-24 | Method and apparatus for providing data processing and control in a medical communication system | 2013-10-11 | Methods and apparatus for providing data processing and control for use in a medical communication system are provided. | 1. A method for data processing using one or more processors, comprising: retrieving a predetermined number of stored signals associated with a monitored analyte level;processing a received signal associated with the monitored analyte level and the retrieved predetermined number of stored signals to determine a tolerance parameter, the tolerance parameter including a predetermined number of acceptable missing data points;comparing the tolerance parameter to a predetermined tolerance range corresponding to a dynamically varying uncertainty parameter determined based upon a number of data available for determining trend information;determining that the tolerance parameter is within the predetermined tolerance range by at least determining that no more than the predetermined number of acceptable missing data points from the retrieved predetermined number of stored signals is missing;in response to determining that the tolerance parameter is within the predetermined tolerance range, determining the trend information; andproviding an indicator on a display of a computing device indicating an analyte level trend based on the determined trend information. 2. The method of claim 1, wherein the tolerance parameter is related to a temporal spacing of the received signal and the retrieved predetermined number of stored signals. 3. The method of claim 1, wherein the trend information is determined based on analyte sensor sensitivity. 4. The method of claim 1, wherein the predetermined tolerance range includes a predetermined number of temporally spaced signals related to the monitored analyte level necessary to determine the trend information. 5. The method of claim 1, wherein each of the retrieved predetermined number of stored signals associated with the monitored analyte level and the received signal associated with the monitored analyte level are temporally separated. 6. The method of claim 1, wherein the trend information includes a projected alarm associated with one or more of a rate of change of the monitored analyte level, or a direction of a projected analyte level. 7. The method of claim 1, including outputting the trend information on an output device separate from the display. 8. The method of claim 1, further comprising administering a therapeutic treatment to a patient based at least on the determined trend information. 9. An analyte monitoring device, comprising: a data processing unit; anda memory operatively coupled to the data processing unit, the memory configured for storing instructions which, when executed by the data processing unit, causes the data processing unit to: retrieve a predetermined number of stored signals associated with a monitored analyte level, process a received signal associated with the monitored analyte level and the retrieved predetermined number of stored signals to determine a tolerance parameter, the tolerance parameter including a predetermined number of acceptable missing data points;compare the toler... | Abbott Diabetes Care Inc. | A61B005/1473 A61B005/1495 G01N033/49 A61B005/145 A61B005/00
US09804162 | B2 | 2017-10-31 | Pleural fluid markers for malignant pleural effusions | 2015-08-31 | Methods of diagnosis and/or prognosis of malignant disease from pleural effusion samples are disclosed. The methods typically include determining the normalized intensity of one or more biomarkers, such as acetone and 3-hydroxybutyate, in pleural effusion samples using nuclear magnetic resonance (NMR), such as proton NMR (1H-NMR) spectroscopy. Reference values for malignant disease diagnosis using the disclosed biomarkers are provided. Reference values for prognosis of median survival time based on the disclosed biomarkers are also provided. | 1. A method of diagnosing a malignant disease or survival term thereof, comprising obtaining samples of effusions from a subject in need of diagnosis,measuring the normalized intensity of a biomarker selected from the group consisting of acetone, 3-hydroxybutyrate, 1-methylhistamine, 1-methylnicotinamide, 2-methylglutarate, 2-oxoglutarate, 3-OH-3-methylglutarate, 3-methyladipate, 4-aminohippurate, acetone, adenine, alanine, creatine, dimethylamine, formate, fumarate, glucose, glycolate, imidazole, lactate, methylamine, O-acetylcarnitine, oxalacetate, phenylacetylglycine, phenylalanine, tryptophan, tyrosine, cis-aconitate, myo-inositol, trans-aconitate, leucine, valine, acetate, acetoacetate, creatinine, and trimethylalanine-N-oxide in the effusions, andcomparing the normalized intensity of the biomarker in the effusions of the subject in need of diagnosis and the normalized intensity in a normal or non-diseased subject,wherein the subject is diagnosed with a malignant disease when the normalized intensity of the biomarker in the effusion of the subject in need of diagnosis is at least 1.5 fold greater than the normalized intensity of the biomarker in the effusion of the normal or non-diseased subject. 2. The method of claim 1, wherein the normalized intensity of the biomarker in the effusion of the subject in need of diagnosis is not greater than fifty fold the normalized intensity of the biomarker in the effusion of the normal or non-diseased subject. 3. The method of claim 1, wherein the effusion is a pleural effusion, pericardial effusion, peritoneal effusion, subdural effusion, or knee effusion. 4. The method of claim 1, wherein the effusion is pleural effusion. 5. The method of claim 1, wherein measuring the normalized intensity of a biomarker in the effusions is measured with nuclear magnetic resonance (NMR) spectroscopy. 6. The method of claim 5, wherein the NMR spectroscopy is proton NMR spectroscopy. 7. The method of claim 1, wherein the biomarker is acetone. 8. The method of claim 1, wherein the biomarker is 3-hydroxybutyrate. 9. The method of claim 1, wherein the normalized intensities of both acetone and 3-hydroxybutyrate are measured. 10. The method of claim 1, wherein the normalized intensity of acetone in the effusion of the normal or non-diseased subject is 0.12%. 11. The method of claim 1, wherein the normalized intensity of 3-hydroxybutyrate in the effusion of the normal or non-diseased subject is 0.26%. 12. The method of claim 1, wherein the subject is determined to have a malignant disease based on the measured normalized intensity of the biomarker. 13. The method of claim 1, wherein the disease prognosis of the subject is determined based on the measured normalized intensity of the biomarker. 14. The method of claim 1, wherein the biomarker is acetone and the median survival value is less than six months when the normalized intensity of acetone is greater than 0.25%. 15. The method of claim 1, wherein the biomar... | The University of Hong Kong | G01N033/00 G01N024/08 G01N033/574 G01R033/465 G01N033/487 A61B005/055
US09808539 | B2 | 2017-11-07 | Hypoosmotic solutions for lymph node detection | 2014-03-10 | Provided are compositions for rapid detection of lymph nodes. The compositions include magnetic particles, such as iron oxide, and a solute present in an amount that results in a hypoosomotic solution. Methods for detecting lymph nodes also are provided. | 1. A hypoosmotic suspension for medical injection comprising: about 13 mg/mL to about 200 mg/mL of superparamagnetic particles; andan osmolyte selected from either about 0.01% w/v to about 0.6% w/v of an inorganic salt or about 0.5% w/v to about 1.5% w/v of a glycolwherein the hypoosmotic suspension has an osmolality of about 80 mOsm/kg to about 160 mOsm/kg. 2. The hypoosmotic suspension of claim 1, wherein the superparamagnetic particles are iron oxide. 3. The hypoosmotic suspension of claim 1, comprising between about 13 mg/ml and 53 mg/ml of superparamagnetic particles. 4. The hypoosmotic suspension of claim 1, further comprising an excipient. 5. The hypoosmotic suspension of claim 1 wherein the superparamagnetic particles are coated. 6. The hypoosmotic suspension of claim 5, wherein the coating comprises dextran. 7. The hypoosmotic suspension of claim 1, wherein the inorganic salt is sodium chloride. 8. The hypoosmotic suspension of claim 1, wherein the suspension is used for the detection of sentinel nodes and comprises about 0.05%-0.3% w/v of the inorganic salt. 9. The hypoosmotic suspension of claim 1, wherein the glycol is propylene glycol. 10. The hypoosmotic suspension of claim 1, wherein the suspension is used for magnetic hyperthermia treatment and comprises about 20 mg/ml-200 mg/ml of the superparamagnetic particles. 11. A method of locating a lymph node in a patient, the method comprising the steps of: providing the hypoosmotic suspension of claim 1;injecting the hypoosmotic suspension into the patient;waiting until the superparamagnetic particles become entrapped in a lymph node; anddetecting the location of the lymph node by detecting the location of the superparamagnetic particles. 12. A method of location a lymph node in a patient, the method comprising the step of: providing the hypoosmotic suspension of claim 1;providing a hypoosmotic suspension comprising superparamagnetic particles, injecting the hypoosmotic suspension into the patient; and detecting a lymph node within 5 minutes of injection by detection the location of the superparamagnetic particles, the detecting sufficient to immediately begin a medical procedure on the lymph node based on the detecting. 13. A method of treating a patient using magnetic hyperthermia, the method comprising the steps of: providing the hypoosmotic suspension of claim 1;injecting the hypoosmotic suspension into the patient; andexposing the patient to an alternating magnetic field. 14. The hypoosmotic suspension of claim 1, wherein the inorganic salt is selected from the group consisting of sodium chloride, potassium chloride, magnesium chloride, ammonium chloride, sodium sulfate, sodium phosphate, potassium phosphate, calcium chloride, magnesium sulfate, potassium acetate, and sodium acetate. | ENDOMAGNETICS LTD. | A61B005/055 A61K049/18 A61K041/00 A61N002/00
US09832412 | B2 | 2017-11-28 | System and method for collection confirmation and sample tracking at the clinical point of use | 2015-01-29 | A system for confirming acquisition of a fluid sample is provided. The system includes a wearable electronic device configured to be worn by a user. The device has a housing, at least one imaging sensor associated with the housing, a data transmission interface for sending data to or receiving data from an external electronic device, a microprocessor for managing the at least one imaging sensor and the data transmission interface, and a program for acquiring and processing images acquired by the at least one imaging sensor. The system further includes a sampling device for collecting a fluid sample in a sample container and at least one identification tag attached to or integrally formed with the sampling device or sample container. The at least one identification tag includes or is associated with a tracking code. | 1. A system comprising: a wearable electronic device configured to be worn by a user comprising:a housing;at least one imaging sensor associated with the housing;a data transmission interface for sending data to or receiving data from an external electronic device;at least one microprocessor for managing the at least one imaging sensor and the data transmission interface; anda program for acquiring and processing images acquired by the at least one imaging sensor;a sampling device for collecting a fluid sample in a sample container; andat least one identification tag attached to or integrally formed with the sampling device or sample container, the at least one identification tag including or being associated with a tracking code,wherein the program processes an image captured by the at least one imaging sensor to identify the at least one identification tag and to determine the tracking code,wherein the program transmits a confirmation that a sample was acquired, and the tracking code, to the external electronic device via the data transmission interface,wherein image capture, image processing, and data transmission are performed automatically without an input or actuation activity by the user, andwherein the program identifies the at least one identification tag in the image captured by processing the image to locate a positional marker on the sampling device and locating the at least one identification tag based on the position of the positional marker, wherein the positional marker points to the at least one identification tag and triggers the wearable electronic device to begin the image processing of the at least one identification tag,wherein, after locating the positional marker by processing the image, the program delays the image processing of the at least one identification tag until the positional marker on the sampling device has been within a field of view of the at least one imaging sensor for a predetermined period of time. 2. The system of claim 1, further comprising a data reporting accessory for providing data to the user. 3. The system of claim 2, wherein the wearable electronic device is a head-worn computer, and wherein the data reporting accessory is a projection prism configured to project a virtual layer to a field of view of the user. 4. The system of claim 3, wherein the virtual layer includes a user interface, the user interface comprising a patient information portion, a sample collection confirmation portion, an identification tag confirmation portion, or any combination thereof. 5. The system of claim 1, wherein image capture, image processing, and data transmission are managed by the program in a hands-free manner. 6. The system of claim 1, wherein the sampling device is a vascular access device, phlebotomy set, renal catheter, tissue sampling device, or any combination thereof. 7. The system of claim 6, wherein the vascular access device is a peripheral IV catheter, a PICC line, or a central line. 8. The sy... | Becton, Dickinson and Company | A61B010/00 H04N005/44 A61M005/142 A61M005/168 A61M005/172 A61M005/315 G06F019/00 A61B005/15 A61B005/157 A61B010/02 A61M005/14 A61M005/42 H04N005/225 H04N007/18 A61B090/96 G06Q050/24 G06Q050/22 A61B090/90 A61M005/20 A61M005/28 A61M005/145 A61M005/158
US09788782 | B2 | 2017-10-17 | Discrimination of apnea type by medical device | 2016-07-18 | Disclosed herein, among other things, are methods and apparatus related to identification of apnea type. One aspect of the present subject matter provides a method for real-time apnea discrimination. The method includes sensing an impedance-based tidal volume signal to monitor a respiratory cycle of a patient, and detecting a reduction in tidal swing using the sensed impendence to detect an apnea event. When the apnea event is detected, a shape of the sensed signal is compared to a stored signal shape to determine whether the apnea event is primarily an obstructive sleep apnea (OSA) event or primarily a central sleep apnea (CSA) event, in various embodiments. | 1. A medical device for apnea discrimination, the device comprising: a sensor configured to sense a parameter related to heart sounds of a patient;a processor adapted to be connected to the sensor, the processor configured to use the sensed parameter to determine whether a detected apnea event is primarily an obstructive sleep apnea (OSA) event or primarily a central sleep apnea (CSA) event; anda memory to store historical data of CSA events and OSA events for the patient,wherein the processor is configured to calculate likelihood for each apnea event based on the historical data and to use the calculated likelihood to determine whether the apnea event is primarily an obstructive sleep apnea (OSA) event or primarily a central sleep apnea (CSA) event. 2. The device of claim 1, wherein the parameter related to heart sounds of the patient includes variability in low frequency components of a spectrum of R-R intervals. 3. The device of claim 1, wherein the parameter related to heart sounds of the patient includes variability in low frequency components of a spectrum of S1 amplitude. 4. The device of claim 1, wherein the parameter related to heart sounds of the patient includes variability in low frequency components of a spectrum of S2 amplitude. 5. The device of claim 1, wherein the parameter related to heart sounds of the patient includes variability in low frequency components of a spectrum of S3 amplitude. 6. The device of claim 1, wherein the parameter related to heart sounds of the patient includes variability in low frequency components of a spectrum of systolic time intervals. 7. The device of claim 1, wherein the processor is configured to use the sensed parameter to determine a mix or ratio of OSA and CSA for the patient. 8. The device of claim 7, further comprising a display configured to present the mix or ratio of OSA and CSA to a clinician for diagnostic assistance. 9. A method for apnea discrimination, the method comprising: sensing a parameter related to heart sounds of a patient using a sensor; andusing a processor connected to the sensor to process the sensed parameter to determine whether a detected apnea event is primarily an obstructive sleep apnea (OSA) event or primarily a central sleep apnea (CSA) event,wherein using the processor to process the sensed parameter includes conducting a time-frequency decomposition of the signal of interest to obtain a spectrogram of the signal over a predetermined time period. 10. The method of claim 9, wherein sensing the parameter related to heart sounds of the patient includes sensing variability in low frequency components of a spectrum of systolic time intervals. 11. The method of claim 10, wherein systolic time intervals include one or more of pre-ejection period (PEP): Q-S1, R-S1, ejection time, diastolic interval and S1-S2 interval. 12. The method of claim 9, wherein using the processor to process the sensed parameter includes using variability along a frequency axis on the sp... | Cardiac Pacemakers, Inc. | A61B005/02 A61B005/00 A61B005/08 A61B007/02 A61B005/053 A61B005/024 A61B007/04
US09792831 | B2 | 2017-10-17 | Incentive spirometer and musical instrument | 2016-07-15 | Systems and methods for teaching proper wind usage for a musical instrument. One embodiment comprises an incentive spirometer apparatus having a tubular body with a bore therethrough. The apparatus has a series of ribs and grooves that are formed in the bore, on the interior surface of the tubular body. The ribs have a first, smaller inner diameter (e.g., ⅙ inch) and the grooves have a second, larger inner diameter (e.g., ⅓ inch). The ribs and grooves may be evenly spaced (e.g., ⅕ inch between successive ribs). When airflow through the tubular body is within a predetermined range, the airflow over the ribs and grooves produces an audible tone from the apparatus. When airflow through the tubular body is outside the predetermined range, no audible tone is produced by the apparatus. | 1. An incentive spirometer apparatus comprising: a tubular body having a bore therethrough, wherein the tubular body has an anterior end and a posterior end; anda series of ribs and grooves that are formed on an interior surface of the tubular body within the bore, wherein the ribs have a first inner diameter and the grooves have a second inner diameter, wherein the first inner diameter is less than the second inner diameter;wherein when airflow through the tubular body is within a predetermined range, the airflow over the series of ribs and grooves produces an audible tone;wherein when airflow through the tubular body is outside the predetermined range, the airflow over the series of ribs and grooves produces no audible tone;wherein the tubular body has a face at the anterior end, wherein the face has an outer rim, wherein the outer rim has a tooth guide edge recess therein, wherein the tooth guide edge recess is adapted to enable a user to position the user's teeth in the tooth guide edge recess and thereby maintain the user's jaws in a desired position, and wherein the tooth guide edge recess is adapted to cause the user's teeth to slip out of the tooth guide edge recess when the user bites down on the apparatus. 2. The apparatus of claim 1, wherein each of the ribs and grooves is evenly spaced from adjacent ones of the series of ribs and grooves. 3. The apparatus of claim 2, wherein the ribs have an inner diameter of about one sixth of an inch, wherein the grooves have an inner diameter of about one third of an inch, and wherein successive ones of the ribs are spaced about one fifth of an inch apart. 4. The apparatus of claim 3, wherein the bore has an anterior throat portion between the anterior end of the tubular body and the series of ribs and grooves, and wherein the anterior throat portion has the second inner diameter. 5. The apparatus of claim 3, wherein the bore has a posterior portion between the posterior end of the tubular body and the series of ribs and grooves, and wherein the posterior portion of the bore has the second inner diameter. 6. The apparatus of claim 1, wherein the tubular body has a first outer diameter at the anterior end, wherein the tubular body has an embouchure recess adjacent to the anterior end, wherein the embouchure recess has a second outer diameter that is less than the first outer diameter, wherein the embouchure recess is adapted to enable a user to position the user's lips within the embouchure recess and thereby hold the tubular body against the user's teeth. 7. The apparatus of claim 1, wherein the tubular body is cylindrical. 8. The apparatus of claim 1, wherein the tubular body has a shape that is not straight. 9. The apparatus of claim 1, wherein the tubular body has an elongated portion between the series of ribs and grooves and the posterior end, wherein the elongated portion has one or more tones holes therein, wherein the apparatus produces different tones when the tones holes are cover... | - | G09B015/06 A61B005/08 G09B015/00
US09801541 | B2 | 2017-10-31 | Remote monitoring of analyte measurements | 2013-03-15 | Methods and apparatus, including computer program products, are provided for remote monitoring. In some example implementations, there is provided a method. The method may include receiving, at a remote monitor, a notification message representative of an event detected, by a server, from analyte sensor data obtained from a receiver monitoring an analyte state of a host; presenting, at the remote monitor, the notification message to activate the remote monitor, wherein the remote monitor is configured by the server to receive the notification message to augment the receiver monitoring of the analyte state of the host; accessing, by the remote monitor, the server, in response to the presenting of the notification message; and receiving, in response to the accessing, information including at least the analyte sensor data. Related systems, methods, and articles of manufacture are also disclosed. | 1. A method comprising: receiving, at a wireless mobile remote monitor, an alert notification message from a server system, the alert notification message representative of a detected event, the event being based on analyte concentration sensor data obtained from a receiver monitoring an analyte state of a host, the analyte concentration sensor data indicating a concentration of an analyte of the host, the wireless mobile remote monitor being remote to the host and the server system;presenting, at the wireless mobile remote monitor, the alert notification message, wherein the wireless mobile remote monitor or a user of the wireless mobile remote monitor is specifically authorized by the host to receive the alert notification message to augment the receiver monitoring of the analyte state of the host, wherein the wireless mobile remote monitor or the user of the wireless mobile remote monitor is authenticated by the server system;receiving, at the wireless mobile remote monitor, user input indicative of acknowledgment of the presented alert notification message;sending an acknowledgment message from the wireless mobile remote monitor to the server system indicating acknowledgement of the presented alert notification message to enable the server system to confirm that the alert notification message has been received;after sending the acknowledgment accessing, by the wireless mobile remote monitor, at least some of the analyte concentration sensor data from the server system; andreceiving, in response to the accessing, information the at least some of the analyte concentration sensor data including historical data, the at least some of the analyte concentration sensor data being generated by the receiver monitoring the analyte state of the host;wherein the event is detected by the server system based on one or more sets of recommended alert rules settings configured by the host using the receiver, at least one of the sets of recommended alert rules settings being modifiable by the user of the wireless mobile remote monitor using the wireless mobile remote monitor. 2. The method of claim 1, wherein the presenting further comprises inhibiting access to one or more applications at the wireless mobile remote monitor until an action at the wireless mobile remote monitor is detected to indicate an interaction by the user with the alert notification message. 3. The method of claim 1, wherein the presenting further comprises presenting the alert notification message as a momentary message on a display at the wireless mobile remote monitor. 4. The method of claim 1, wherein at least one of the wireless mobile remote monitor and the receiver comprise one or more of a mobile station, a wireless terminal, a tablet, a smart phone, a multi-mode wireless device, and a computer. 5. The method of claim 1, wherein the server system comprises at least one processor configured to receive analyte sensor data from the receiver, process the analyte sensor data to dete... | DexCom, Inc. | A61B005/00 A61B005/02 A61B005/145 G06F019/00 H04L029/08
US09821096 | B2 | 2017-11-21 | Tissue sampling, processing and injection syringe device and methods of using the same | 2011-12-08 | Tissue sampling, processing and injection syringe device and methods of filtering and concentrating cellular components of tissue sample in situ within the device. | 1. A tissue sampling, processing and injection syringe device comprising: a tissue collection tube having (i) an interior volume disposed between a distal end opening and a proximal end opening, (ii) a cylindrical outer surface disposed between said distal end opening and said proximal end opening, (iii) a set of micro-pores formed in said tissue collection tube through said cylindrical outer surface between said distal and proximal end openings according to a first spatial arrangement about said cylindrical outer surface, and (iv) a flange having first and second opposite flat side edge surfaces extending about said proximal end opening of said tissue collection tube;a piston and plunger assembly inserted through said proximal end opening and into said interior volume of said tissue collection tube; anda rotatable micro-pore occluder including a flat rectangular flange adapted for engagement with either said first or second opposite flat side edge surfaces;wherein said rotatable micro-pore occluder is configured for rotatable attachment about said tissue collection tube in either of two states, namely:(i) a micro-pore occluded state wherein said rotatable micro-pore occluder is rotated about said tissue collection tube and arranged in said first rotation position and said flange engages with a first one of said opposite flat side edge surfaces extending from said tissue collection tube, and(ii) a micro-pore non-occluded state wherein said rotatable micro-pore occluder is rotated about said tissue collection tube and arranged in a second rotation position and said flange engages with said second one of said opposite flat side edge surfaces extending from said tissue collection tube;wherein said rotatable micro-pore occluder has a cylindrical geometry with a set of micro-pore occluding surfaces spatially aligned with said set of micro-pores when said rotatable micro-pore occluder is rotated about said tissue collection tube and arranged in said micro-pore occluded state;wherein said rotatable micro-pore occluder has a set of apertures spatially aligned with said set of micro-pores when said rotatable micro-pore occluder is rotated about said tissue collection tube and arranged in said micro-pore non-occluded state;wherein said rotatable micro-pore occluder is rotated about said tissue collection tube to said first rotation position and arranged in said micro-pore occluded state so as to occlude said set of micro-pores formed in said tissue collection tube during tissue harvesting operation;wherein said rotatable micro-pore occluder is rotated about said tissue collection tube to said second rotation position and arranged in said micro-pore non-occluded state so as to non-occlude said set of micro-pores formed in said tissue collection tube during in situ tissue processing operations within said tissue collection tube;wherein said distal end opening is adapted to receive and connect to a cap after an aspirated tissue sample has been collected withi... | Rocin Laboratories, Inc. | A61M001/00 A61B010/02 A61M027/00 A61B019/00 A61B018/04 A61B018/18 A61B001/04 A61B005/05
US09788800 | B2 | 2017-10-17 | Notification system of deviation from predefined conditions | 2015-10-02 | The present invention includes a healthcare system for capturing and maintaining patient healthcare records and employing information in the records to proactively alert hospital workers to a condition that places patient health at risk. | 1. A notification system comprising: a patient support having a configuration;a patient support monitoring system, the patient support monitoring system operable to generate multiple notification levels; anda patient's condition or risk factor associated with a patient stored in said patient support monitoring system;wherein, based on the patient's condition or risk factor, said patient support monitoring system is configured to automatically select or determine a set of appropriate patient support configurations or criteria as default configurations or criteria for the patient and to detect and monitor when the patient support configuration or criteria is unsuitable for the patient based on the default configurations or criteria and to generate a notification at one of the notification levels of an unsatisfactory configuration or criteria of the patient support when the patient support is unsuitable for the patient based on the default configuration or criteria based on the patient's condition or risk factor. 2. The notification system according to claim 1, wherein said patient support monitoring system is configured to allow a user to select a notification level from the notification levels. 3. The notification system according to claim 1, wherein the patient support monitoring system is configured to generate a high priority notification level, a medium priority notification level, and a low priority notification level. 4. The notification system according to claim 1, wherein said patient support monitoring system is configured to designate a notification level from the notification levels based on the patient's condition or risk factor. 5. The notification system according to claim 1, wherein said patient support monitoring system is configured to allow a person to select or enter the patient's condition or risk factor in said patient support monitoring system for the patient. 6. The notification system according to claim 5, wherein said patient support monitoring system is configured to allow a person to remotely select or enter the patient's condition or risk factor in said patient support monitoring system. 7. The notification system according to claim 1, wherein the patient support monitoring system comprises a network interface, and the network interface operable to communicate with a hospital network or healthcare system. 8. The notification system according to claim 7, wherein the patient support monitoring system further comprises one or more sensors for detecting the patient support's configuration or criteria, the network interface being operable to send information from the sensor or sensors to the hospital network or healthcare system. 9. The notification system according to claim 1, wherein the notification comprises a mobile alert or a pop-up display. 10. A notification system comprising: a patient support having a configuration;a patient support monitoring system, the patient support monitoring system comprising one or... | Stryker Corporation | G08B005/22 A61B005/00 G08B021/04 A61B005/11 G06F019/00
US09816980 | B2 | 2017-11-14 | Apparatus and corresponding method for sampling and analyzing drugs and respective metabolites in breath air, particularly suitable for performing road drug tests | 2014-03-17 | A portable apparatus to sample and analyze volatile organic compounds in breath air, includes:;a box internally divided in two internal sections, each housing a sample collecting tube or filter in turn containing a stationary phase suitable for retaining determined analytes, as drugs and/or respective metabolites, having specific chemical physical properties, contained in breath air;;a pre-accumulation chamber for receiving breath air from the outside through a breath entrance connection and for feeding the received breath air to the sample collecting filters housed in the two sections, with these two sample collecting filters being connected in parallel to the chamber; and;a tool application kit for removing the two sample collecting filters from the respective sections in the box of the apparatus, in a way that avoids any contamination thereof, and for extracting the analytes that are retained by the stationary phase on the same two sample collecting filters. | 1. An apparatus or device (10, BACT) for use to sample and analyze volatile organic compounds (VOCS) in breath air, particularly suitable for performing road or street drug tests, wherein the apparatus is a portable sampling apparatus and comprises: a box (1) that is divided at the inside in two chambers or sections (2, 3), with each of said two sections (2, 3) housing a respective sample collecting filter (4, 5), in the form of a tube, containing a stationary phase (16) suitable for retaining determined analytes, as drugs and/or respective metabolites, having specific chemical or physical properties, contained in the breath air; anda breath pre-accumulation chamber (18) arranged for receiving a sample of breath air from the outside through a breath entrance connection (6) and for feeding the received breath air to the two sections (2, 3), housing the two sample collecting filters (4, 5) containing the stationary phase (16), with tubes (8, 21) being provided for connecting the pre-accumulation chamber (18) to the sample collecting filters (4, 5) housed in the two sections (2, 3), whereby a parallel connection is realized between said sample collecting filters (4, 5), and with the breath air moving from the pre-accumulation chamber (18) to the two sections (2, 3), while passing through the sample collecting filters (4, 5), whereby the analytes of interest in the air breath are retained by the stationary phase (16) contained in the sample collecting filters (4, 5), and moving from the two sections (2, 3) to the outside through respectively a first and a second expulsion connector (19, 20) in turn connected to a final breath expulsion connector (7);characterized by further comprising a tool application kit (15, 22, 23) for removing the two sample collecting filters (4, 5) from the respective sections (2, 3) in the box (1) of the apparatus (10), in a way that avoids any contamination of said sample collecting filters (4, 5), and for extracting the analytes that are retained by the stationary phase (16) on the same two sample collecting filters (4, 5),wherein said kit (15) is composed of a mechanical tool or instrument (22) suitable for removing the two sample collecting filters (4, 5) from the respective chamber or section (2, 3) in the box (1) of said apparatus (10), while avoiding to touch them with fingers and any other device that can lead to their contamination, and of an analyte extraction solution (23) suitable for extracting and solubilizing the analytes of interest retained by the stationary phase (16) in the two sample collecting filters (4, 5). 2. The apparatus of claim 1, wherein the breath air is moved from the pre-accumulation chamber (18) to the two sections (2, 3) and from the two sections (2, 3) to the outside by means of a pump (17). 3. The apparatus of claim 1, wherein the stationary phase (16) contained in said sample collecting filters (4, 5) is suitable for retaining different compounds depending on its composition, and wherei... | Simone Cristoni Rosario Billetta Claudio Giuliani | G01N033/497 A61B005/097 G01N001/22 G01N030/72
US09820648 | B2 | 2017-11-21 | Image processing apparatus and image processing method | 2014-06-27 | An image processing apparatus includes a generation unit and a display control unit. The generation unit generates a subtraction image using a first region and a second region, the first region being a portion of a first fundus image acquired by performing autofluorescence imaging on a fundus of an eye to be inspected, the second region being a portion of a second fundus image acquired by performing autofluorescence imaging on the fundus of the eye to be inspected at a time different from that of the first fundus image and being at a position corresponding to the first region. The display control unit causes a display unit to display an image generated by superimposing information regarding the subtraction image at a position corresponding to the first and second regions on a fundus image acquired by performing autofluorescence imaging on the fundus. | 1. An image processing apparatus comprising: a region extraction unit configured to extract a blood vessel portion of a first fundus image acquired by performing autofluorescence imaging on a fundus of an eye to be inspected and the blood vessel portion of a second fundus image acquired by performing autofluorescence imaging on the fundus of the eye to be inspected at a time different from that of the first fundus image;an image gradation processing unit configured to perform, using brightness of the blood vessel portion of the first fundus image and brightness of the blood vessel portion of the second fundus image, gradation changing on at least one of the first and second fundus images such that a brightness difference, between the blood vessel portion of the first fundus image and the blood vessel portion of the second fundus image, is reduced;an alignment processing unit configured to align the first fundus image with the second fundus image by positioning the blood vessel portion of the first fundus image and the blood vessel portion of the second fundus image;an image generation unit configured to generate a subtraction image subtracted between different regions of the first and second fundus images after the gradation processing and the alignment processing, the different regions being different from the blood vessel portion of the first fundus image and the blood vessel portion of the second fundus image; anda display control unit configured to cause a display to display the first or second fundus image which superimposes the subtraction image on at least one of the different regions. 2. The image processing apparatus according to claim 1, further comprising: an image superimposition unit configured to superimpose at least one of the subtraction image and a display form illustrating a position of the subtraction image on the fundus image. 3. The image processing apparatus according to claim 2, wherein the image superimposition unit superimposes the display form illustrating the position at the position on the fundus image, and the display control unit causes the display to arrange and display a superimposition image generated by performing superimposition and the subtraction image. 4. The image processing apparatus according to claim 1, wherein the generation unit generates a subtraction image using the first and second fundus images, anda region having an intensity greater than or equal to a certain intensity in the subtraction image generated using the first and second fundus images is generated as the subtraction image. 5. The image processing apparatus according to claim 1, wherein the display control unit causes the display to display the first and second fundus images, and a display form illustrating the second region to be displayed on the second fundus image. 6. The image processing apparatus according to claim 1, wherein the display control unit causes the first and second fundus images to be arranged and displayed along a t... | Canon Kabushiki Kaisha | A61B003/14 A61B003/12 A61B003/00 A61B005/00
US09808671 | B2 | 2017-11-07 | Exercise training system | 2012-11-15 | An exercise training system is provided having a system controller and a plurality of stimulant target units connected via a wireless network to the system controller. Each stimulant target unit has a light source providing light to stimulate a user, a proximity sensor providing an output of a distance between the proximity sensor and an object external to the unit, and means for providing feedback to the user to signal that the unit has been actuated by the user. The system controller includes a program for activating the stimulant target units in a sequence. The reaction by the user to the illumination is registered when the user brings a body part or other object to within a selected distance of the proximity sensor, this distance being programmable by the user. | 1. An exercise training system, comprising: a system controller;a plurality of stimulant target units connected via a wireless network to the system controller, each stimulant target unit including a light source providing light to stimulate a user, a proximity sensor providing an output of a distance between the proximity sensor and an object external to the stimulant target unit, and feedback means for informing the user that the stimulant target unit has been actuated, wherein each stimulant target unit further includes a power source and a stimulant target unit controller that are separate from the power source and the stimulant target unit controller of each of the other stimulant target units, wherein the stimulant target unit controller for each of the plurality of the stimulant target units is wirelessly connected to the system controller and is programmed to receive instructions from said system controller, wherein the selected proximity distance required to actuate each stimulant target unit is programmable by the user;wherein the system controller includes a program for activating the stimulant target units in a sequence, and wherein in the execution of said program, the system controller signals a first of the stimulant target units to illuminate the corresponding light source to stimulate the user, the first stimulant target unit registering an actuation thereof by the user in reaction to the illumination when the user brings a body part or other object to within a selected distance of the corresponding proximity sensor without contacting the first stimulant target unit, the first stimulant target thereafter activating the corresponding feedback means to alert the user that the first stimulant target unit has been actuated;and wherein said first stimulant target unit communicates said registering of said actuation by said first stimulant target unit to the system controller, and said system controller is programmed to signal, in response to communication by said first stimulant target unit of said actuation, a second of the stimulant target units to illuminate the corresponding light source to stimulate the user, the second stimulant target unit registering an actuation thereof by the user in reaction to the illumination when the user brings the body part or other object to within the selected distance of the corresponding proximity sensor without contacting the second stimulant target unit, the second stimulant target thereafter activating the corresponding feedback means to alert the user that the second stimulant target unit has been actuated,wherein each stimulant target unit includes a housing which includes a first shell connected to a second shell via an elastomeric rim which permits the first shell to move along an axis relative to the second shell upon impact,wherein the stimulant target unit controller is mounted to the second shell and is spaced laterally from the elastomeric rim, and wherein the first shell is generally a... | FITLIGHT SPORTS CORP. | A63B024/00 A63B069/00 A63B071/06 G08B005/36 H04B010/114 A61B005/16 G09B019/00 A63B023/02 A63B022/00 A63B102/06 A63B102/02 A63B021/00 A63B102/22
US09788728 | B2 | 2017-10-17 | Endoscopic polarized multispectral light scattering scanning method | 2010-01-22 | Method and system for early detection of precancerous and other abnormal changes in tissue of various organs. The system comprises a combination of endoscopic scanning with light scattering spectroscopy and improves detection of abnormalities that may otherwise remain undetected. The system may include a probe that collects data of quality that is independent of a distance of the probe from the scanned tissue. During endoscopy, tissue of an organ is imaged using polarized multispectral light scattering scanning and results are presented to a user in a manner that allows detecting abnormal morphological and biochemical changes in the tissue. A determination of whether to perform biopsy may be performed while the endoscopy is being performed, which thus provides guided biopsy. An entire surface of the organ may be rapidly scanned and results of the scanning are analyzed with a reduced time delay. | 1. Apparatus for obtaining characteristics of a surface of a hollow organ, the apparatus comprising: a probe configured to scan the surface of the hollow organ, the probe comprising: polarization and collimating optics configured to illuminate at least one portion of the hollow organ with polarized and collimated light having a first polarization, andlight collection optics configured to receive light backscattered by the at least one portion as a result of the illumination;at least one spectrometer configured to generate at least one spectrum from the received light;an imager configured to capture at least one image of the at least one portion; andan analysis unit configured to analyze the at least one spectrum associated with the backscattered light to determine wavelength-dependent values of a reduced scattering coefficient for a sub-epithelial region of the at least one portion and to detect at least one abnormal morphological change of the at least one portion from the wavelength-dependent values. 2. The apparatus of claim 1, wherein the analysis unit is further configured to determine a correspondence between the at least one image and the at least one abnormal morphological change. 3. The system of claim 1, further comprising a user interface configured to present information comprising the at least one image in association with at least one visual representation of the at least one abnormal morphological change, wherein the information is used to determine whether to take a tissue sample from the at least one portion. 4. The apparatus of claim 3, wherein the information comprises the at least one image superimposed with the at least one visual representation of the at least one abnormal morphological change. 5. The apparatus of claim 3, wherein the information is presented as at least one of a color map and a pseudocolor map. 6. The apparatus of claim 1, wherein the at least one spectrum comprises a summation of backscattered light having a same polarization as the first polarization and backscattered light having a second polarization that is perpendicular to the first polarization. 7. The apparatus of claim 6, wherein the at least one abnormal morphological change comprises a degree of dysplasia of the at least one portion. 8. The apparatus of claim 6, wherein the analysis unit is configured to determine a degree of dysplasia of the at least one portion based, at least in part, upon the wavelength-dependent values of the reduced scattering coefficient. 9. The apparatus of claim 1, wherein the probe is configured to be located at a distance from the at least one portion without contacting a surface of the at least one portion. 10. The apparatus of claim 1, wherein the imager is configured to obtain the at least one image of the at least one portion at a time when the probe scans the at least one portion. 11. The apparatus of claim 1, wherein the wavelength-dependent values comprise multiple values of reduced scattering coeffici... | Beth Israel Deaconess Medical Center, Inc. | A61B006/00 A61B005/00
US09814426 | B2 | 2017-11-14 | Mobile wearable electromagnetic brain activity monitor | 2015-01-18 | This invention is a mobile wearable electromagnetic brain activity monitor comprising a frame worn on a person's head, electromagnetic energy sensors which collect data concerning the person's electromagnetic brain activity, and a control unit. The electromagnetic energy sensors can be electroencephalogram (EEG) electrodes. In an example, the frame can be circular. In an example, the frame can loop around the person's head from one ear to the other ear. In an example, the frame can be an eyewear frame which also includes light-transmitting optical members. | 1. A mobile wearable electromagnetic brain activity monitor comprising: a wearable frame which is configured to be worn on a person's head, wherein the wearable frame comprises the frame of a pair of eyeglasses, wherein this wearable frame is configured to completely encircle a portion of the person's head, wherein this wearable frame has a central circular axis which is configured at an anterior acute angle in the range of 0 to 45 degrees with respect to a horizontal plane when the person's head is upright, wherein this wearable frame has two or more full-phase sinusoidal variations around its central circular axis, wherein the wearable frame is configured to have an arcuate wave above an ear which is part of a full-phase sinusoidal variation around the central circular axis of the wearable frame, and wherein the wearable frame is configured to be supported by the person's ears; a plurality of electromagnetic energy sensors which are configured to be held in electromagnetic communication with the person's brain by the wearable frame, wherein these sensors collect data concerning the person's electromagnetic brain activity; and a control unit which is part of the wearable frame with one or more components selected from the group consisting of power source and/or power-transducing component, data transmission and data reception component, data memory component, and data processor. 2. The monitor in claim 1 wherein data from the electromagnetic energy sensors is used to adjust and/or control the light absorption, light reflection, light refraction, light spectrum transformation, focal direction, focal distance, light polarization, and/or parallax view of at least one optical lens or other light-transmitting member. 3. The monitor in claim 1 wherein this monitor further comprises one or more other sensors selected from the group consisting of: accelerometer, inclinometer, gyroscope, strain gauge, or other motion or position sensor; microphone or other sound sensor; thermometer or other temperature sensor; camera or other imaging sensor; optical sensor or optoelectronic sensor; blood pressure sensor; ECG/EKG sensor, heart rate monitor, and/or heart rate sensor; EMG sensor or other muscle activity sensor; GPS sensor, other location sensor, magnetometer, or compass; spectroscopy sensor or other spectral analysis sensor; electrochemical sensor; blood oximetry sensor; piezoelectric sensor; chewing sensor or swallowing sensor; respiration sensor; pressure sensor; galvanic skin response sensor; and taste or odor sensor. 4. The monitor in claim 1 wherein this monitor further comprises one or more interface components selected from the group consisting of: a computer-to-human interface such as a display screen, one or more lights, one or more speakers, and/or one or more tactile actuators; a human-to-computer interface such as a touch screen, one or more touch-activated buttons, microphone and speech-recognition capability, and/or gesture recognition capa... | Medibotics LLC | A61B005/04 A61B005/00 A61B005/0476 A61B005/0205 A61B005/0478 A61B005/021 A61B005/024 A61B005/0404 A61B005/0488 A61B005/053 A61B005/08 A61B005/11 A61B005/1455
US09811638 | B2 | 2017-11-07 | Means and method for improved glycemic control for diabetic patients | 2016-10-03 | A glycemic control system includes a physician processor, remote processor, and a portable telephone having a data input mechanism, a display, and an internal processor for bi-directional communication with the physician's processor and the remote processor. A patient inputs data to the internal processor responsive to input from the physician's processor and then transmits the information to the remote processor where an optimized number of units to be administered is sent back and displayed on the portable telephone. | 1. A method comprising: receiving, at data processing hardware, an insulin sensitivity factor for a patient and a target blood glucose range defined by upper and lower blood glucose limits for the patient from a medical professional computing device associated with an authorized medical professional;storing, by the data processing hardware, the insulin sensitivity and the target blood glucose in memory hardware in communication with the data processing hardware;receiving, at the data processing hardware, a current blood glucose measurement of the patient and a time period of the current blood glucose measurement from a user device controlled by the patient and in communication with the data processing hardware, wherein the user device: receives the current blood glucose measurement and the time period of the current blood glucose measurement from a blood glucose meter in communication with the user device; andtransmits the current blood glucose measurement of the patient and the time period of the current blood glucose measurement to the data processing hardware;in response to receiving the current blood glucose measurement from the user device, obtaining, by the data processing hardware, the insulin sensitivity factor for the patient and the target blood glucose range for the patient from the memory hardware;determining, by the data processing hardware, whether the current blood glucose measurement exceeds a midpoint of the target blood glucose range for the patient; andwhen the current blood glucose measurement exceeds the midpoint of the target blood glucose range for the patient: determining, by the data processing hardware, a correction dose based on a function of the current blood glucose measurement, the mid-point of the target blood glucose range and the insulin sensitivity factor; andtransmitting the correction dose from the data processing hardware to the user device, the correction dose, when received by the user device, causing a user interface executing on the user device to display the correction dose. 2. The method of claim 1, wherein the correction dose is calculated as follows:CD=(BG-Tm)(1700⁢((Tm-60)⁢xS1⁢x⁢⁢24))wherein CD is the correction dose, BG is the current blood glucose measurement, Tm is the mid-point of the target blood glucose range, and S1 is the insulin sensitivity factor. 3. The method of claim 1, further comprising: receiving, at the data processing hardware, one or more patient condition parameters from the user device, wherein the user device is configured to: display a list of patient condition parameters in the user interface executing on the user device;receive, in the user interface, selection inputs from the patient for each patient condition parameter from the list of patient condition parameters; andtransmit the patient condition parameters to the data processing hardware; andstoring, by the data processing hardware, the one or more patient condition parameters in the memory hardware. 4. The method of c... | Aseko, Inc. | G06Q050/00 G06F019/00 G06Q050/22 A61B005/145 A61M005/172
US09799114 | B2 | 2017-10-24 | Method, system and apparatus for quantitative surgical image registration | 2014-11-27 | A method in a computing device for quantitative surgical image registration includes: prior to a surgical procedure, obtaining, using a first imaging modality, a preoperative image of patient tissue and a plurality of preoperative measurements of a material property of the patient tissue. The preoperative measurements correspond to respective points in the preoperative image. The method includes storing the preoperative image and the preoperative measurements, and during the surgical procedure, using a second imaging modality, capturing an intraoperative image of the patient tissue and a second plurality of intraoperative measurements of the material property of the patient tissue. The intraoperative measurements correspond to respective points in the intraoperative image. The method includes comparing the first and second pluralities of measurements to determine a transformation for registering the preoperative image and the intraoperative image; and storing the transformation in association with one of the intraoperative image and the preoperative image. | 1. A method of image registration, comprising: prior to a surgical procedure, obtaining a preoperative image of patient issue at a computing device using a first imaging modality;using the first imaging modality, obtaining a first plurality of preoperative measurements of a material property of the patient tissue at the computing device, each measurement of the first plurality of preoperative measurements corresponding to a respective point in the preoperative image;storing the preoperative image and the preoperative measurements in a memory of the computing device;during the surgical procedure, obtaining an intraoperative image of the patient tissue at the computing device using a second imaging modality;using the second imaging modality, obtaining a second plurality of intraoperative measurements of the material property of the patient tissue at the computing device, each measurement of the second plurality of intraoperative measurements corresponding to a respective point in the intraoperative image;comparing the first and second pluralities of measurements at the computing device to determine a transformation for registering the preoperative image and the intraoperative image;storing the transformation in the memory in association with one of the intraoperative image and the preoperative image;applying the transformation to the intraoperative image; andcontrolling a display connected to the computing device to present the preoperative image overlaid with the intraoperative image. 2. The method of claim 1, wherein obtaining the preoperative measurements comprises deriving the preoperative measurements from the preoperative image; and wherein obtaining the intraoperative measurements comprises deriving the intraoperative measurements from the intraoperative image. 3. The method of claim 1, wherein the first imaging modality and the second imaging modality are selected from magnetic resonance imaging (MRI), computed tomography (CT), ultrasound, optical coherence tomography (OCT), and optical spectroscopy. 4. The method of claim 3, wherein one of the first imaging modality and the second imaging modality is OCT, and wherein the corresponding one of the first and second pluralities of measurements are obtained using polarization-sensitive OCT (PS-OCT). 5. The method of claim 3, wherein one of the first imaging modality and the second imaging modality is optical spectroscopy, and wherein the corresponding one of the first and second pluralities of measurements are obtained using Raman spectroscopy. 6. The method of claim 1, wherein the first and second pluralities of measurements of the material property are selected from a list consisting of flow, stiffness, anisotropy, and metabolite concentration. 7. The method of claim 1, wherein the first imaging modality is different from the second imaging modality. 8. The method of claim 1, wherein the preoperative image is registered to the patient tissue; the method further comprising: determining ... | SYNAPTIVE MEDICAL (BARBADOS) INC. | G06T007/00 A61B034/10 A61B005/00 A61B005/055 A61B006/03 A61B008/08 A61B006/00
US09775474 | B2 | 2017-10-03 | Body drying and inspection apparatus | 2015-04-27 | A body drying and personal inspection apparatus is provided for air drying and for providing the user with several inspection and data input devices. The apparatus comprises a base platform and an upright member with one or more handles to support the user. A blower assembly forces air through the base platform and through an elongated blower tube handled by the user. Further provided is an inspection camera coupled to the blower tube and a foot camera within the base platform, both used to capture images of the user's body while on the platform. A display and processing system is also provided that allows images to be processed and displayed, for input to be received via a user input device, and optionally for data to be transmitted to a third party over a network. The base platform also provides weight measuring capabilities for measuring the weight of the user thereon. | 1. A personal drying and inspection apparatus, comprising: a base platform which is adapted to support a user positioned thereupon;an upright member extending from the base platform;one or more handles extending from the upright member along its length for the user to grip while positioned on the base platform;a blower motor assembly comprising a blower motor and a housing, the blower motor assembly having an air inlet and a first blower outlet disposed at the distal end of an elongated tube extending from the blower motor assembly;the elongated tube extending from the blower motor assembly and being sufficiently flexible for the user to position the first blower outlet towards targeted portions of the user while positioned on the base platform;a second blower outlet between the blower motor assembly and the base platform;one or more platform apertures that allow air to be forced therethrough and from the second blower outlet to direct forced air substantially upwards and onto the user positioned on the base platform;a heating element within the blower motor assembly for heating air forced through the first blower outlet and the second blower outlet;blower motor controls disposed along the upright member for the user to energize and to control the blower motor and heating element of the blower motor assembly. 2. The personal drying and inspection apparatus of claim 1, further comprising: a display screen;a user input device adjacent to the blower motor controls;a body inspection camera connected to the elongated tube, whereby the body inspection camera comprises a lens disposed adjacent to the first blower outlet of the elongated tube such that the user can direct the lens to different body area while positioned on the base platform;a processing system adapted to receive signals from the body inspection camera and to transmit the signals to the display screen for visualizing a feed from the body inspection camera. 3. The personal drying and inspection apparatus of claim 2, wherein: the processing system further comprises a processor and a memory;wherein the processing system is adapted to capture still images from the feed of the body inspection camera;wherein the processing system is further adapted to receive physiological data from the user via the user input device. 4. The personal drying and inspection apparatus of claim 3, wherein: the processing system further comprises a network connectivity device for transmitting and receiving signals over a network; andwherein the processing system is further adapted to transmit the still images and the physiological data over the network. 5. The personal drying and inspection apparatus of claim 2, further comprising: a user weight measuring device within the base platform that is adapted to register a weight of the user when positioned on the base platform. 6. The personal drying and inspection apparatus of claim 1, further comprising: a display screen;a user input device adjacent to the blower m... | - | A47K010/48 A61B005/00 A61B005/145 G01G019/50
US09797858 | B2 | 2017-10-24 | Biosensor, measuring instrument for biosensor, and method of quantifying substrate | 2014-01-29 | A method of measuring a quantity of a substrate contained in sample liquid is provided. This method can reduce measurement errors caused by a biosensor. The biosensor includes at least a pair of electrodes on an insulating board and is inserted into a measuring device which includes a supporting section for supporting detachably the biosensor, plural connecting terminals to be coupled to the respective electrodes, and a driving power supply which applies a voltages to the respective electrodes via the connecting terminals. One of the electrodes of the biosensor is connected to the first and second connecting terminals of the measuring device only when the biosensor is inserted into the measuring device in a given direction and has a structure such that the electrode becomes conductive between the first and second connecting terminals due to a voltage application by the driving power supply. | 1. A method of measuring a quantity of a substrate in a sample liquid by inserting a biosensor into a measuring device, the method comprising: providing the biosensor which includes: an insulating board;an electrode part which includes a counter electrode and a measuring electrode disposed on at least a part of the insulating board;a sample supplying path which supplies the sample liquid to the electrode part; anda reagent layer at least partially disposed in the sample supplying path;providing the measuring device which includes: a supporting section which detachably supports the biosensor;a connecting terminal which applies a voltage to the electrode part; anda driving power supply;applying the voltage to the electrode part by the driving power supply when the biosensor is inserted into the supporting section;detecting a start of supplying the sample liquid to the sample supplying path by detecting that a first output signal from the electrode part reaches a given threshold;determining that a sufficient quantity of the sample liquid is supplied to the sample supplying path by detecting that a second output signal from the electrode part reaches a given threshold within a predetermined period after the start of supplying the sample liquid to the sample supplying path;measuring an elapsed time between the first output signal and the second output signal;measuring the quantity of the substrate of the sample liquid based on a third output signal from the electrode part; andif the elapsed time is within a predetermined time, correcting the measurement of the quantity of the substrate of the sample liquid based on the elapsed time, orif the elapsed time is not within a predetermined time, terminating the measurement of the quantity of the substrate of the sample liquid. 2. The method of claim 1, wherein a compensation value in the correction of the measurement increases when the elapsed time increases. 3. The method of claim 1, wherein the predetermined time is 10 seconds. 4. The method of claim 1, wherein measuring the quantity of the substrate is terminated if a sufficient quantity of the sample liquid has not been supplied when the elapsed time exceeds the predetermined time. 5. The method of claim 1, wherein measuring the quantity of the substrate is temporally halted when the elapsed time is within the predetermined time, during a period after the start of supplying the sample liquid to the sample supplying path and before determining whether the necessary quantity of the sample liquid is supplied to the sample supplying path. | Panasonic Healthcare Holdings Co., Ltd. | G01N027/327 C12Q001/00 G01N033/487 G01N033/543 A61B005/1486 G01N027/26
US09826909 | B2 | 2017-11-28 | Neurogenic baroreflex sensitivity measurement device, neurogenic baroreflex sensitivity measurement program and neurogenic baroreflex sensitivity measurement method | 2014-03-28 | [Problem] To provide a neurogenic baroreflex sensitivity measurement device, neurogenic baroreflex sensitivity measurement program and neurogenic baroreflex sensitivity measurement method capable of easily and objectively measuring neurogenic baroreflex sensitivity that is not dependent on vascular hardness without using blood pressure or pulsations in the diameter of the carotid artery.;[Solution] The neurogenic baroreflex sensitivity measurement device comprises: a pulse wave data-acquiring unit (41) for acquiring pulse wave data of an artery; a normalized pulse wave volume-calculating unit (42) for calculating a normalized pulse wave volume on the basis of the pulse wave data; a pulse interval-acquiring unit (43) for acquiring pulse intervals corresponding to the pulse wave data; a baroreflex series-detecting unit (44) for detecting baroreflex series in which the normalized pulse wave volume and pulse interval both increase or decrease for at least three beats in a row; and a neurogenic baroreflex sensitivity-calculating unit (45) for calculating the slope of a regression line representing the correlation between the normalized pulse wave volume and the pulse interval in a baroreflex series as the neurogenic baroreflex sensitivity, which is an index representing the neurogenic baroreflex function. | 1. A neurogenic baroreflex sensitivity measurement device comprising: a pulse-wave-data acquiring unit that acquires pulse wave data that indicates a volume change in an artery;a normalized-pulse-volume calculating unit that calculates, on the basis of the pulse wave data, a normalized pulse volume obtained by dividing amplitude of an alternating-current component of a pulse wave by an average of a direct-current component of the pulse wave at the same time;a beat-interval acquiring unit that acquires a beat interval corresponding to the pulse wave data;a baroreflex-series detecting unit that detects a baroreflex series in which both of the normalized pulse volume and the beat interval continuously increase or decrease for three beats or more; anda neurogenic-baroreflex-sensitivity calculating unit that calculates, as neurogenic baroreflex sensitivity that is an index indicating a neurogenic baroreflex function, a slope of a regression line indicating a correlation between the normalized pulse volume and the beat interval in the baroreflex series. 2. A non-transitory computer-readable recording medium storing a neurogenic baroreflex sensitivity measurement program that causes a computer to function as: a pulse-wave-data acquiring unit that acquires pulse wave data that indicates a volume change in an artery;a normalized-pulse-volume calculating unit that calculates, on the basis of the pulse wave data, a normalized pulse volume obtained by dividing amplitude of an alternating-current component of a pulse wave by an average of a direct-current component of the pulse wave at the same time;a beat-interval acquiring unit that acquires a beat interval corresponding to the pulse wave data;a baroreflex-series detecting unit that detects a baroreflex series in which both of the normalized pulse volume and the beat interval continuously increase or decrease for three beats or more; anda neurogenic-baroreflex-sensitivity calculating unit that calculates, as neurogenic baroreflex sensitivity that is an index indicating a neurogenic baroreflex function, a slope of a regression line indicating a correlation between the normalized pulse volume and the beat interval in the baroreflex series. 3. A neurogenic baroreflex sensitivity measurement method comprising: a pulse-wave-data acquiring step for acquiring pulse wave data that indicates a volume change in an artery;a normalized-pulse-volume calculating step for calculating, on the basis of the pulse wave data, a normalized pulse volume obtained by dividing amplitude of an alternating-current component of a pulse wave by an average of a direct-current component of the pulse wave at the same time;a beat-interval acquiring step for acquiring a beat interval corresponding to the pulse wave data;a baroreflex-series detecting step for detecting a baroreflex series in which both of the normalized pulse volume and the beat interval continuously increase or decrease for three beats or more; anda neurogenic-baroreflex-... | SAPPORO MEDICAL UNIVERSITY | A61B005/02 A61B005/021 A61B005/00 A61B005/024 A61B005/026
US09802158 | B2 | 2017-10-31 | Ultrafiltration membrane device, bioartificial organ, and related methods | 2016-07-29 | The present invention relates to ultrafiltration. In particular, the present invention provides nanoporous membranes having pores for generating in vitro and in vivo ultrafiltrate, devices and bioartificial organs utilizing such nanoporous membranes, and related methods (e.g., diagnostic methods, research methods, drug screening). The present invention further provides nanoporous membranes configured to avoid protein fouling with, for example, a polyethylene glycol surface coating. | 1. An ultrafiltration system comprising: a) a housing comprising a first chamber and a second chamber, wherein the first chamber and the second chamber are separated by a membrane consisting essentially of a plurality of nanofabricated pores having a width less than 20 nanometers, wherein the distance between nanofabricated pores on said membrane is around 50 nm, wherein the thickness of the membrane is approximately 500 nm, wherein the membrane has a porosity around 20%;b) a first inlet positioned to deliver fluid into the first chamber; andc) a first outlet positioned to remove fluid out of the first chamber. 2. The system of claim 1, wherein said membrane has a surface coating selected from the group consisting of polyethylene glycol, oligosaccharide surfactant polymers, heparin, and hyaluronan. 3. The system of claim 1, wherein the nanofabricated pores are modified with polyethyleneglycol. 4. The system of claim 1, wherein the material of said membrane is selected from the group consisting of silicon, polysilicon, silicon carbide, ultrananocrystalline diamond, diamond-like-carbond (DLC), silicon dioxide, PMMA, SU-8, PTFE, titanium, silica, silicon nitride, polytetrafluorethylene, polymethylmethacrylate, polystyrene, and silicone. 5. The system of claim 1, wherein said nanofabricated pores separate the first chamber from the second chamber. 6. The system of claim 1, wherein the length of said nanofabricated pores is less than 50 pm. 7. The system of claim 1, wherein said plurality of nanofabricated pores have a slit shape. 8. The system of claim 1, wherein said nanofabricated pores are arranged in the center of said membrane. 9. The system of claim 1, wherein said housing comprises a coating, said coating being biocompatible for in vivo use. 10. The system of claim 1, wherein said housing has a length and a width, said length of said housing being less than 300 millimeters and said width of said housing being less than 300 millimeters. 11. The system of claim 1, wherein the first inlet and first outlet are positioned in the housing. 12. The system of claim 1, wherein the inlet and outlet further comprise a conduit. 13. The system of claim 12, wherein the conduit comprises a tubing. 14. The system of claim 13, wherein the tubing is attached to the inlet and outlet by a clamp or threaded connection. 15. The system of claim 1, wherein the inlet and outlet pass through the housing. 16. The system of claim 1, wherein the inlet and outlet comprise a spheroid, elliptical, or slit shape. 17. The system of claim 1, wherein the first chamber comprises a biological fluid. 18. The system of claim 17, wherein the biological fluid is blood. 19. The system of claim 18, wherein the blood enters the first chamber through the inlet. 20. The system of claim 18, wherein blood exits the first chamber through the outlet. 21. The system of claim 1, wherein the second chamber comprises a dialysate. 22. The system of claim 1, wherein the first cha... | THE REGENTS OF THE UNIVERSITY OF MICHIGAN THE CLEVELAND CLINIC FOUNDATION | A61M001/16 B01D061/18 A61B005/145 B01D061/14 B01D061/20 B01D067/00 A61M001/26 G01N001/40
US09808318 | B2 | 2017-11-07 | Robotic arthroplasty system | 2017-06-15 | A robotic system for use in performing an arthroplasty procedure of a patient is provided. The robotic system includes a robotic mechanism, a cutting tool coupled to the robotic mechanism for resecting a portion of a bone of the patient through an incision in the skin of the patient, and a computer coupled to the robotic mechanism and configured to control the robotic mechanism. The robotic system also includes at least one navigation member positionable into engagement with tissue in the patient, an optical system coupled to the computer, a position sensor configured to provide movement information of the robotic mechanism relative to the bone, and an interface coupled to the computer, the interface configured to operate the computer. | 1. A robotic system for use in performing an arthroplasty procedure of a patient, the robotic system comprising: a robotic mechanism, wherein the robotic mechanism is configured to perform the arthroplasty procedure without being coupled to an operating table;a cutting tool coupled to the robotic mechanism for resecting a portion of a bone of the patient through an incision in the skin of the patient, the robotic mechanism configured to control movement of the cutting tool during the resection;a computer coupled to the robotic mechanism and configured to control the robotic mechanism;at least one navigation member positionable into engagement with tissue in the patient;an optical system coupled to the computer, the optical system configured to determine the location of at least one bone relative to other tissue in the patient, wherein the optical system is configured to provide location information of the at least one bone to the computer for used by the robotic mechanism;a position sensor configured to provide movement information of the robotic mechanism relative to the bone; andan interface coupled to the computer, the interface configured to operate the computer. 2. The system of claim 1, wherein the bone is at least one of a femur and tibia. 3. The system of claim 1, further comprising a motor configured to facilitate movement of the robotic mechanism. 4. The system of claim 1, wherein the cutting tool is a saw. 5. The system of claim 4, wherein the saw is one of an oscillating saw and a reciprocating saw. 6. The system of claim 5, further comprising a prosthetic implant. 7. The system of claim 6, wherein the prosthetic implant includes at least one of an implant bearing surface and an arcuate shape. 8. The system of claim 7, further comprising a screw configured to couple to one or more tissues of the patient. 9. A robotic system for use in performing an arthroplasty procedure of a patient, the robotic system comprising: a robotic mechanism configured to perform the arthroplasty procedure without being coupled to an operating table;a prosthetic implant;a saw coupled to the robotic mechanism for resecting a portion of one of the tibia and the femur of the patient through an incision in the leg of the patient, the robotic mechanism configured to control movement of the saw during the resection by use of an electric motor;a computer coupled to the robotic mechanism and configured to control the robotic mechanism;at least one navigation member positionable into engagement with tissue of the patient;an optical system coupled to the computer, the optical system configured to determine a location of at least one bone relative to other tissue in the patient, wherein the optical system is configured to provide location information of the at least one bone to the computer for use by the robotic mechanism;a position sensor configured to provide movement information of the robotic mechanism relative to at least one of the tibia and femur; andan... | P Tech, LLC | A61B017/00 A61B034/30 A61B017/064 A61B034/32 A61B017/02 A61B006/03 A61B034/10 A61B017/16 A61B005/00 A61B017/04 A61B017/14 A61B034/00 A61F002/00 A61B017/88 A61B090/00 A61F002/38 A61F002/30 A61B017/15
US09795328 | B2 | 2017-10-24 | Method and apparatus for providing glycemic control | 2017-01-06 | Methods to provide glycemic control and therapy management based on monitored glucose data, and current and/or target Hb1AC levels are provided. Systems to provide glycemic control and therapy management based on monitored glucose data, and current and/or target Hb1AC levels are provided. Kits to provide glycemic control and therapy management based on monitored glucose data, and current and/or target Hb1AC levels are provided. | 1. A method implemented using one or more processors, comprising: determining, using the one or more processors, a correlation between a mean glucose value information and an HbA1C level associated with a subject;determining, using the one or more processors, a rate of glycation of the subject based at least in part on the determined correlation between the mean glucose value information and the HbA1C level;applying, using the one or more processors, one or more subject specific parameters to the determined correlation between the mean glucose value information and the HbA1C level;determining, using the one or more processors, a target HbA1C level based on determining the rate of glycation of the subject and applying the one or more subject specific parameters to the determined correlation; andproviding, using the one or more processors, a therapy recommendation on a display device operatively coupled to the one or more processors including at least the determined target HbA1C level; wherein the one or more subject specific parameters includes an age of the subject, an activity level of the subject, a medication intake information of the subject, or a living condition of the subject. 2. The method of claim 1, wherein the mean glucose value information includes applying a weighting function to monitored glucose level information over a predetermined time period. 3. The method of claim 2, wherein the weighting function is based on a time of day information associated with the monitored glucose level information. 4. The method of claim 2, wherein the weighting function is based on a time period associated with the monitored glucose level information. 5. The method of claim 2, wherein the predetermined time period includes one of approximately 30 days, approximately 45 days, or approximately 90 days. 6. The method of claim 1, wherein providing the therapy recommendation including the determined target HbA1C level on the display device includes generating a visual output associated with the determined target HbA1C level and one or more of the mean glucose value information, the rate of glycation, HbA1C level, and the subject specific parameters. 7. The method of claim 1, further including storing one or more of the mean glucose value information, the HbA1C level, the determined correlation between the mean glucose value information and the HbA1C level, and the target HbA1C level. 8. The method of claim 1, wherein the rate of glycation is a ratio of the mean glucose value information and the HbA1C level. 9. The method of claim 1, wherein the mean glucose value information is obtained using a glucose sensor that comprises a plurality of electrodes including a working electrode. 10. The method of claim 9, wherein the working electrode comprises a glucose-responsive enzyme and a mediator, wherein at least one of the glucose-responsive enzyme and the mediator is chemically bonded to a polymer disposed on the working electrode, and wherein at least... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/145 G06F019/00 A61B005/11 A61B005/1486
US09820669 | B2 | 2017-11-21 | System and method for electrical impedance spectroscopy | 2013-08-07 | Systems and methods for determining brain health of a subject include or employ an electrical stimulator configured to apply a current to at least one pair of electrodes, and the electrodes are positioned on a skull of the subject to apply the current and to receive brain activity of the subject. The electrical stimulator is configured to apply a current having a waveform according to a Stochastic Gabor Function (SGF). A signal processor is configured to record the brain activity of the subject in the form of spectral electrical impedance data, and a computer system having non-transient computer readable media is programmed and configured to process the spectral electrical impedance data and indicate an impedance change within the brain of the subject. | 1. A system for monitoring the brain of a subject, the system comprising: an electrical stimulator configured to apply a current to at least one pair of electrodes, the electrodes configured to be positioned on a skull of the subject to apply the current and to receive brain activity of the subject, wherein the electrical stimulator is configured to apply a current having a waveform according to a Stochastic Gabor Function (SGF), wherein the waveform includes a dual energy pulse;a signal processor configured to record the brain activity of the subject in the form of spectral electrical impedance data; anda computer system having non-transient computer readable media programmed and configured to process the spectral electrical impedance data and indicate an impedance change within the brain of the subject. 2. The system of claim 1 wherein a pulse of the dual energy pulse is formed of white noise. 3. The system of claim 2 wherein the computer readable media is programmed to subtract images of differing energies represented by the spectral electrical impedance data. 4. The system of claim 3 wherein the subtraction of the images is programmed to distinguish between ischemic and hemorrhagic stroke. 5. The system of claim 2 wherein the dual energy pulse provides impedance change within a depth of greater than about 1 mm in the brain. 6. The system of claim 1 wherein the current comprises white noise. 7. The system of claim 1 wherein the electrical stimulator is configured to apply current frequencies between about 10,000 kHz and 100,000 kHz. 8. The system of claim 1 wherein the electrical stimulator is configured to apply current frequencies between about 30,000 kHz and 500,000 kHz. 9. A method for monitoring the brain of a subject, the method comprising: positioning a plurality of electrodes on a skull of the subject;applying a current to at least one pair of the electrodes, the current having a waveform based on a Stochastic Gabor Function (SGF), the waveform including a dual energy pulse;measuring brain activity of the subject in the form of spectral electrical impedance data and storing the data within a non-transient computer readable media; andprocessing, using a computer system, the data so as to obtain an indication of an impedance change within the brain of the subject. 10. The method of claim 9 wherein a pulse of the dual energy pulse is formed of white noise. 11. The method of claim 10 wherein the impedance change is associated with the indication of stroke. 12. The method of claim 11 wherein the impedance change is associated with distinguishing the stroke between an ischemic and hemorragic stroke. 13. A system for analyzing a brain of a subject, comprising: a plurality of electrodes configured to be coupled to the subject to deliver excitation pulses to a brain of the subject and receive electrical information from the brain of the subject, the excitation pulses including dual energy pulses;a controller operatively connected to... | The General Hospital Corporation | A61B005/04 A61B005/0478 A61B005/053 A61B005/00
US09814534 | B2 | 2017-11-14 | Catheter control system and graphical user interface | 2015-12-22 | A remote workstation for the control of percutaneous intervention devices is provided. The remote workstation includes a control system for remotely and independently controlling at least two percutaneous intervention devices. The control system includes at least one input device to control the percutaneous intervention devices. The control system controls movement of at least one of the percutaneous intervention devices along at least two degrees of freedom. The remote workstation also includes a graphical user interface for displaying icons representative of the operational status of each of the percutaneous intervention devices. | 1. A remote workstation for the control of percutaneous intervention devices comprising: a control system configured to remotely and independently control a first percutaneous intervention device, the control system including at least one input device to control the first percutaneous intervention device, wherein the control system controls movement of the first percutaneous intervention device along at least one degree of freedom;a measurement module configured to allow a user to measure a length of a vasculature structure by aligning a tip of the first percutaneous intervention device controlled by the control system with a distal end of the vasculature structure, withdrawing the tip until it is aligned with a proximate end of the vasculature structure and measuring a distance moved by the first percutaneous intervention device as the first percutaneous intervention device traverses the length of the vasculature structure; anda display configured to provide a first measurement control comprising a first reset button, a first distance indicator and a first stored distance indicator, the first reset button configured to, when actuated, reset a value displayed by the first distance indicator, wherein the first distance indicator indicates a distance moved by the first percutaneous intervention device since the first reset button was actuated. 2. The remote workstation of claim 1, wherein the first percutaneous intervention device is one of a guide wire, a working catheter, a guide catheter, a stent, and an angioplasty balloon. 3. The remote workstation of claim 1, wherein the control system controls movement of the first percutaneous intervention device in a first degree of freedom and in a second degree of freedom, wherein the first degree of freedom and the second degree of freedom include axial movement and rotation. 4. The remote workstation of claim 1, wherein the display includes a first set of icons and the first set of icons are displayed in a first color to indicate that the first percutaneous intervention device is active and are displayed in a second color to indicate that the first percutaneous intervention device is inactive. 5. The remote workstation of claim 4 wherein, when the first percutaneous intervention device is inactive, the control system is configured to hold the inactive first percutaneous intervention device such that the first percutaneous intervention device is not permitted to move axially relative to a patient on which a procedure is being performed. 6. The remote workstation of claim 4, wherein the first set of icons provide at least one of a visual indication of movement of the first percutaneous intervention device and a visual indication of the physical position of the at least one input device. 7. The remote workstation of claim 1, wherein the display displays an icon representative of an amount of contrast agent delivered to a patient during a medical procedure. 8. The remote workstation of claim 1, wher... | CORINDUS, INC. | A61F002/95 A61B005/107 A61B005/00 A61B034/00 A61M025/01 A61B034/37 A61B090/98 A61B034/20 A61B005/0205 A61B017/00 A61M025/10 A61B034/30
US09814397 | B2 | 2017-11-14 | Scale-based systems and methods for monitoring heart function | 2015-12-11 | In accordance with an example embodiment, a body-weight sensing scale includes cardio-based physiological sensing circuitry to detect heart characteristics of a user, and provide outputs indicative of the detected heart characteristics. A processor circuit is arranged with the cardio-based physiological sensing circuitry to process data to provide a noise-reduced cardiogram signal which characterizes functionality/health of the user's heart. | 1. A body-weight sensing scale apparatus comprising: a weighing scale configured and arranged with a housing and a scale platform;user-physiologic sensors, including a first sensor and including a second sensor, configured and arranged to capture a cardiogram signal of a user from at least one foot of the user in real-time while the user is in an upright position and with said at least one foot of the user on the scale platform of the body weight sensing scale apparatus,the user-physiologic sensors including circuitry configured and arranged to detect and output an indication of at least one of noise present in the cardiogram signal and a physiologic characteristic of the user, by sensing and capturing a series of heart-signals from a user while the user stands on the scale platform at a stationary position, the series of heart-signals corresponding to the cardiogram signal based on a plurality of a heartbeats from the user; anda processor circuit configured and arranged to process the cardiogram signal and generate therefrom a noise-filtered cardiac-related signal indicative of a condition of the user's heart by filtering, from the cardiogram signal, noise that is due to user movement. 2. The body-weight sensing scale apparatus of claim 1, wherein the processor circuit includes circuitry configured and arranged to track changes for a cardiac condition, or indicate an abnormal cardiac condition. 3. The body-weight sensing scale apparatus of claim 1, wherein user-physiologic sensors are configured and arranged with an electrocardiogram (ECG) sensor that, when operative, detects an ECG signal from the user. 4. The body-weight sensing scale apparatus of claim 1, wherein user-physiologic sensors are configured and arranged with a photoplethysmograph sensor that, when operative, detects blood-flow pulsations of the user. 5. The body-weight sensing scale apparatus of claim 1, wherein the cardiogram signal is a ballistocardiogram signal and wherein noise-filtered cardiac-related signal is a noise-filtered ballistocardiogram signal. 6. The body-weight sensing scale apparatus of claim 1, wherein the cardiogram signal is a ballistocardiogram (BCG) signal, wherein noise-filtered cardiac-related signal is a noise-filtered ballistocardiogram signal, and wherein the processor circuit is further configured and arranged to process the cardiogram signal by executing an algorithm to process the captured BCG signal. 7. The body-weight sensing scale apparatus of claim 1, wherein the weighing scale is integrated with the housing and the scale platform being part of the weighing scale, and further including communication circuitry located within the housing and communicatively coupled with or integrated as part of the processor circuit, the communication circuitry being configured and arranged to communicate cardio-function data with an external circuit. 8. The body-weight sensing scale apparatus of claim 1, wherein the weighing scale is integrated with the hou... | The Board of Trustees of the Leland Stanford Junior University | A61B005/04 A61B005/0205 A61B005/024 A61B005/029 A61B005/11 A61B005/00 A61B005/0402 A61B005/0488 G01G019/50 A61B005/0456 A61B005/053
US09788895 | B2 | 2017-10-17 | Basket catheter with deflectable spine | 2016-11-04 | A catheter adapted for mapping and/or ablation in the atria has a basket-shaped electrode array with two or more location sensors with a deflectable expander. The catheter has comprises a catheter body, a basket electrode assembly at a distal end of the catheter body, and a control handle at a proximal end of the catheter body. The basket electrode assembly has a plurality of electrode-carrying spines and an expander that is adapted for longitudinal movement relative to the catheter body for expanding and collapsing the assembly via a proximal end portion extending past the control handle that can be pushed or pulled by a user. The expander is also adapted for deflection in responsive to an actuator on the control handle that allows a user to control at least one puller wire extending through the catheter body and the expander. | 1. A catheter comprising: an elongated catheter body having proximal and distal ends and at least one lumen therethrough;an electrode assembly at the distal end of the catheter body, the electrode assembly having proximal and distal ends and comprising a plurality of spines, the spines comprising a plurality of electrodes;an expander having proximal and distal ends, the expander forming a longitudinal axis of the assembly, the spines attached at their proximal ends to the catheter body and at their distal ends to the expander;at least one puller wire that extends through the expander, the puller wire having a distal end anchored in the expander at a location distal of the distal end of the catheter body and proximal of the distal end of the assembly;at least one proximal location sensor and at least one distal location sensor; anda control handle proximal of the catheter body, the control handle having an actuator adapted to move the at least one puller wire,wherein the electrode assembly has an expanded arrangement when the expander is moved proximally along the longitudinal axis relative to the catheter body and the electrode assembly has a collapsed arrangement when the expander is moved distally along the longitudinal axis relative to the catheter body, andwherein the expander is adapted to deflect in at least one direction when the at least one puller wire is moved by the actuator. 2. The catheter of claim 1, further comprising a proximal junction at the distal end of the catheter body, the proximal junction comprising: a tubing having a proximal end and a distal end, the proximal end being mounted over the distal end of the catheter body;a ring inside the distal end of the tubing, the ring having a through-hole; anda tunnel member in the through-hole of the ring,wherein the tunnel member has a lumen through which the expander extends and has longitudinal movement. 3. The catheter of claim 2, wherein the tunnel member has at least one off-axis lumen. 4. The catheter of claim 2, further comprising a distal junction at the distal end of the expander, the distal junction comprising: an outer tubing having a lumen;a puller wire anchor in the lumen of the outer tubing; anda ring in the lumen of the outer tubing,wherein distal ends of spine support members are anchored between the ring and the puller wire anchor. 5. The catheter of claim 4, wherein the at least one proximal location sensor is in the proximal junction and the at least one distal location sensor is in the distal junction. 6. The catheter of claim 1, wherein the spines are mounted in about 360 radial degrees around the expander. 7. The catheter of claim 1, wherein the spines are mounted in about 180 radial degrees around the expander. 8. The catheter of claim 1, wherein at least one spine carries at least one ring electrode. 9. The catheter of claim 1, wherein the expander carries at east one ring electrode distal of the distal end of the catheter body. 10. The catheter of c... | Biosense Webster (Israel) Ltd. | A61B005/042 A61B018/14 A61M025/01 A61M025/00 A61B005/00 A61B017/00 A61B018/00
US09777307 | B2 | 2017-10-03 | Heterocyclic nitrogen containing polymer coated analyte monitoring device and methods of use | 2016-11-09 | The present invention is directed to membranes composed of heterocyclic nitrogen groups, such as vinylpyridine and to electrochemical sensors equipped with such membranes. The membranes are useful in limiting the diffusion of an analyte to a working electrode in an electrochemical sensor so that the sensor does not saturate and/or remains linearly responsive over a large range of analyte concentrations. Electrochemical sensors equipped with membranes described herein demonstrate considerable sensitivity and stability, and a large signal-to-noise ratio, in a variety of conditions. | 1. A system comprising: an electrochemical sensor, comprising: a working electrode comprising: a sensing layer in contact with a conductive material of the working electrode;a diffusion limiting membrane layer disposed over the sensing layer, wherein the diffusion limiting membrane layer comprises a crosslinker and a polymer having the formula:wherein the solid horizontal line represents a polymer backbone and n is an integer from 150 to 15,000; and a counter electrode in electrochemical communication with the working electrode; andan electrochemical sensor control unit comprising: a housing; anda transmitter for transmitting sensor signals from the electrochemical sensor. 2. The system of claim 1, wherein the sensor control unit further comprises a data storage unit for storing data from the electrochemical sensor. 3. The system of claim 1, wherein the sensor control unit further comprises a reference voltage generator. 4. The system of claim 1, wherein the sensor control unit further comprises a temperature probe. 5. The system of claim 1, wherein the sensing layer of the working electrode comprises a glucose-responsive enzyme. 6. The system of claim 1, wherein the sensing layer of the working electrode comprises a redox mediator. 7. The system of claim 6, wherein the redox mediator comprises a complex selected from the group consisting of a ruthenium-containing complex and an osmium-containing complex. 8. The system of claim 6, wherein the redox mediator is non- leachable with respect to the working electrode. 9. The system of claim 6, wherein the redox mediator is immobilized on the working electrode. 10. The system of claim 1, wherein the crosslinker comprises a poly(ethylene glycol). 11. The system of claim 1, wherein the crosslinker is poly(ethylene glycol) diglycidyl ether. 12. The system of claim 1, wherein the diffusion limiting membrane layer limits flux of glucose or lactate thereacross. 13. The system of claim 1, wherein the diffusion limiting membrane layer limits flux of glucose or lactose thereacross in vivo. 14. The system of claim 1, wherein the analyte responsive enzyme is covalently bound to the polymer. 15. The system of claim 1, wherein the redox mediator is covalently bound to the polymer. 16. The system of claim 1, wherein the analyte responsive enzyme is crosslinked to the polymer. 17. The system of claim 1, wherein the redox mediator is covalently crosslinked to the polymer. 18. The system of claim 1, wherein the diffusion limiting membrane layer is covalently bound to the sensing layer. | Abbott Diabetes Care Inc. | G01N027/327 G01N027/333 C12Q001/00 A61B005/1486 A61B005/145 A61B005/1495
US09788798 | B2 | 2017-10-17 | Method of visualizing a bridge therapy process | 2012-06-12 | The present invention provides for a simultaneous graphical representation, a risk of bleeding and a risk of thrombosis providing a visualized bridge therapy process. Furthermore, the present invention provides for a computer-based prediction of the haemostatic situation of the examined blood circulation by using a combination of a biochemical model and a pharmacokinetic model for calculation or another mathematical representation of the blood circulation. | 1. A method of visualizing a bridge therapy process, the method comprising the steps of: receiving coagulation data describing a haemostatic situation of a blood circulation of a patient at a computer having a processor and a display, anddisplaying, on the display of the computer, the coagulation data to a user by means of a graphical representation, the displaying including:displaying a first value measured by a first type of blood test, wherein the first value indicates a first effect of a heparin like drug on the haemostatic situation, andsimultaneously to the displaying of the first value displaying a second value measured by a second type of blood test, wherein the second value indicates a second effect of a vitamin K antagonist type anticoagulant on the haemostatic situation, andwherein the first and second values describe the haemostatic situation of the blood circulation at a first point in time and are simultaneously displayed as a first point within an X and Y coordinate system having the first value for the X axis and the second value for the Y axis. 2. The method according to claim 1, the displaying of the coagulation data further comprising the steps of: displaying, on the display of the computer, a third value measured by the first type of blood test, wherein the third value indicates a third effect of a heparin like drug on the haemostatic situation,simultaneously to the displaying of the third value displaying, on the display of the computer, a fourth value measured by the second type of blood test, wherein the fourth value indicates a fourth effect of a vitamin K antagonist type anticoagulant on the haemostatic situation,wherein the third and fourth values describe the haemostatic situation of the blood circulation at a second point in time and are simultaneously displayed as a second point within the X and Y coordinate system having the third value for the X axis and the fourth value for the Y axis. 3. The method according to claim 1, wherein the first and second type of blood test are respectively and independently chosen from the group consisting of activated partial thrombopiastine time (aPTT) test, anti-Factor 10a test, prothrombin time (PT) test, international normalized ratio (INR) test, a test indicating thrombosis or coagulation, a test indicating bleeding or anti-coagulation, a test indicating the haemostatic function. 4. The method according to claim 1, further comprising the step of: calculating, using the processor of the computer, a prediction of the haemostatic situation of the blood circulation by a mathematical representation of the blood circulation. 5. The method according to claim 1, further comprising the step of: calculating, using the processor of the computer, a prediction of the haemostatic situation of the blood circulation by combining a biochemical model of the blood circulation and a pharmacodynaniical model of the blood circulation. 6. The method according to claim 5, further comprising the step... | Koninklijke Philips N.V. | G06T011/20 A61B005/00 G06F019/26 G06F019/00 A61B005/026 A61B005/02
US09801567 | B2 | 2017-10-31 | Medical image processing device, method for operating the same, and endoscope system | 2015-07-30 | First RGB image signals are inputted. A first B/G ratio and a first G/R ratio are calculated. The first B/G ratio and the first G/R ratio are converted into a second B/G ratio and a second G/R ratio, respectively, through a color information conversion process. Owing to the color information conversion process, a difference between first and second observation areas in a second signal ratio space formed by the second B/G ratio and the second G/R ratio is greater than a difference between the first and second observation areas in a first signal ratio space formed by the first B/G ratio and the first G/R ratio, and a difference between the first and third observation areas in the second signal ratio space is greater than a difference between the first and third observation areas in the first signal ratio space. | 1. A medical image processing device comprising: an input processing unit for performing an input process of a first color image signal;a color information obtaining section for obtaining two or more pieces of first color information from the first color image signal; anda color information converter for performing a color information conversion process to convert the two or more pieces of first color information into two or more pieces of second color information, the color information conversion process making a difference D12y greater than a difference D12x and making a difference 13y greater than a difference 13x, the difference D12x being a difference between a first observation area and a second observation area that differs in position from the first observation area in a first feature space formed by the two or more pieces of first color information, the difference D12y being a difference between the first and second observation areas in a second feature space formed by the two or more pieces of second color information, the difference D13x being a difference between the first observation area and a third observation area that differs in position from the first and second observation areas in the first feature space, the difference D13y being a difference between the first and third observation areas in the second feature space. 2. The medical image processing device according to claim 1, wherein the color information converter has a first color information converter for performing a first color information conversion process as the color information conversion process, and the first color information conversion process makes at least one of saturation and hue of the first observation area in the second feature space different from at least one of saturation and hue of the first observation area in the first feature space. 3. The medical image processing device according to claim 2, wherein the difference D12x before the first color information conversion process represents a difference in saturation, and the difference D13x before the first color information conversion process represents a difference in hue, and the difference D12y after the first color information conversion process represents differences in saturation and hue, and the difference D13y after the first color information conversion process represents a difference in hue. 4. The medical image processing device according to claim 2, wherein the color information converter has a second color information converter for performing a second color information conversion process as the color information conversion process, and the second color information conversion process makes the first observation area in the second feature space identical in saturation and hue to the first observation area in the first feature space. 5. The medical image processing device according to claim 4, wherein the difference D12x before the second color information conversion process represents a d... | FUJIFILM Corporation | A61B005/103 H04N001/56 A61B001/00 A61B001/273 A61B001/06 G06T007/00 H04N009/64 G06T007/90 A61B001/05 A61B001/04 H04N005/225
US09827448 | B2 | 2017-11-28 | Hybrid ultrasound and magnetic resonance imaging device | 2015-03-21 | In order to position and simultaneously operate both a magnetic resonance (MR) coil and an ultrasound transducer adjacent to a target volume while minimizing interference between the MR coil and the ultrasound transducer, hybrid ultrasound and magnetic resonance imaging (MRI) devices are provided. One of the hybrid ultrasound and MRI devices includes a transducer array, a communications interface connected with the transducer array, and a housing that electromagnetically shields and encloses the transducer array and the communications interface. An electrically conducting side of the housing receives a radio frequency (RF) signal from an MRI system and transmits the RF signal to an electrically conducting side of a housing of another of the hybrid ultrasound and MRI devices. | 1. A radio frequency (RF) loop antenna for a magnetic resonance imaging (MRI) system, the RF loop antenna comprising: housings each comprising an electrically conducting side;a plurality of transducer arrays, each transducer array of the plurality of transducer arrays being electromagnetically shielded and enclosed by a corresponding one of the housings and comprising a multi-dimensional array of elements;communications interfaces each being electromagnetically shielded and enclosed by the corresponding one of the housings, and being in communication with a corresponding transducer array of the plurality of transducer arrays; andtuning capacitors each electrically connecting the electrically conducting sides of two of the housings, respectively,wherein the electrical conducting side of the respective housing is configured to receive an RF signal from the MRI system and transmit the RF signal to the electrically conducting side of another housing of the RF loop antenna via at least one of the tuning capacitors. 2. The RF loop antenna of claim 1, further comprising resonant circuits, each of the resonant circuits being physically and electrically connected to the electrically conducting sides of two of the housings and comprising a corresponding one of the tuning capacitors. 3. The RF loop antenna of claim 1, wherein the electrically conducting sides of a subset of the housings form an RF loop, the subset of the housings electromagnetically shielding and enclosing corresponding transducer arrays of the plurality of transducer arrays, wherein the RF loop is operative at a frequency of 64 MHz or 128 MHz, and wherein each transducer array of the corresponding transducer arrays is operative at 1-5 MHz. 4. The RF loop antenna of claim 3, wherein the subset of the housings comprises four housings. 5. The RF loop antenna of claim 1, wherein the subset of housings is a first subset of the housings, wherein electrically conducting sides of the first subset of the housings form a first RF loop, and wherein the electrically conducting sides of a second subset of the housings form a second RF loop. 6. The RF loop antenna of claim 5, wherein the housings comprise a first housing, a second housing, a third housing, a fourth housing, a fifth housing, and a sixth housing, wherein the first subset comprises the first housing, the second housing, the fifth housing, and the sixth housing, and the second subset comprises the third housing and the fourth housing, and wherein the third housing is positioned between the first housing and the fifth housing, and the fourth housing is positioned between the second housing and the sixth housing. 7. The RF loop antenna of claim 1, wherein each of the housings comprises a ground foil, wherein the RF loop antenna further comprises low-pass filters, each of the low-pass filters being electrically connected between the electrically conducting side and the ground foil of the corresponding one of the housings. 8. The RF loop... | Siemens Aktiengesellschaft | A61B005/055 A61N007/02 G01R033/48 A61B008/00 A61B008/08 A61B090/00
US09827137 | B2 | 2017-11-28 | Devices for treating bruxism | 2013-11-05 | The present disclosure relates to devices that aid a user in ceasing bruxism. The devices may be placed between a user's teeth, such as during sleep, and alert the user when the user clenches his/her teeth. In an aspect, a device includes a housing containing a battery power source, a vibration device, and an electrical contact mechanism having a movable element which closes a circuit to actuate the vibration device. A disk-shaped portion adapted to be emplaced in a person's mouth between a cheek and adjacent teeth is disposed around the housing. A resilient bite portion extends from the housing and is adapted to be emplaced between upper and lower teeth. A compressible element is disposed in the bite portion and causes the movable element of the electrical contact to close to actuate the vibration device in response to compression of the bite portion. | 1. A dental device, comprising: an outer wall defining an inner chamber, said outer wall having a disc portion adapted to be emplaced in a person's mouth between a cheek and adjacent teeth, and a tab portion, adapted to be emplaced between upper and lower teeth, said inner chamber extending within both said disc portion and said tab portion;a housing carried within said inner chamber, said housing encapsulating a battery, a vibrator and electrical conductors, said housing including a thinned portion; andan elongate member extending within the portion of said inner chamber defined by said tab portion, wherein said thinned portion deforms based on the upper and lower teeth moving toward each other and activates said vibrator in response to said elongate member being deformed. 2. A dental device, comprising: an outer wall defining an inner chamber, said outer wall having a disc portion adapted to be emplaced in a person's mouth between a cheek and adjacent teeth, and a tab portion, adapted to be emplaced between upper and lower teeth, said inner chamber extending within both said disc portion and said tab portion;a vibrator and electrical conductors carried within said inner chamber; andan actuating member extending within the portion of said inner chamber defined by said tab portion, wherein said electrical conductors activate said vibrator in response to said actuating member being deformed based on the upper and lower teeth moving toward each other. 3. A dental appliance for detecting and treating teeth clenching, comprising: a first housing containing a battery power source, a vibration device connected to said power source, and an electrical contact mechanism having a movable element which closes a circuit to actuate said vibration device;a disk-shaped portion around said first housing and adapted to be emplaced in a person's mouth between a cheek and adjacent teeth;a resilient bite portion extending from said first housing and adapted to be emplaced between upper and lower teeth, said bite portion having an interior area which is compressible, as by upper and lower teeth moving toward each other; anda compressible element disposed in said interior area, said compressible element causing said movable element of said electrical contact to close to actuate said vibration device in response to compression of said bite portion. 4. The dental appliance of claim 3, wherein said compressible element engages said movable element in response to said compression. 5. The dental appliance of claim 4, wherein said compressible element causes said movable element to physically move to close said circuit in response to said compression. 6. The dental appliance of claim 5, wherein said compressible element is an encapsulated medium adapted to bulge toward said movable element to close said circuit in response to said compression. 7. The dental appliance of claim 6, wherein said medium is a gas. 8. The dental appliance of claim 6, wherein said medium is ... | SJS GROUP, LLC | A61B005/00 A61F005/56
US09775517 | B2 | 2017-10-03 | Hand-held medical-data capture-device having optical detection of vital signs from multiple filters and interoperation with electronic medical record systems | 2014-12-29 | In one implementation, an apparatus estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point is described, estimates temperature from a digital infrared sensor and determines vital signs from a solid-state image transducer, or determines vital signs from a solid-state image transducer and estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point; after which the estimated and/or determined information is transmitted to an external database. | 1. A non-touch thermometer to measure a temperature, the non-touch thermometer comprising: a microprocessor;a battery operably coupled to the microprocessor;a single button operably coupled to the microprocessor;a camera operably coupled to the microprocessor and operable to capture at least two images to a memory;wherein the microprocessor includes a pixel-examination-module that is operable to examine pixel values of the at least two images in the memory, a temporal-variation module that is operable to determine temporal variation of the pixel values between the at least two images, wherein the temporal-variation module includes a skin-pixel-identifier that identifies pixel values that are representative of the skin in at least two images, a spatial bandpass filter that is operably coupled to the skin-pixel-identifier and that processes output of the skin-pixel-identifier, a temporal bandpass filter that is operably to a regional facial clusterial module and that is applied to output of the spatial bandpass filter and a temporal-variation identifier that is operable coupled to the temporal bandpass filter and that identifies temporal variation of the output of the temporal bandpass filter, the microprocessor further comprising a signal processing module that is operable to amplify the temporal variation resulting in amplified temporal variation, and a vital-sign generator that is operably coupled to the signal processing module that generates at least one human vital sign from the temporal variation;a wireless communication subsystem that is operably coupled to the microprocessor and that is operable to transmit a representation of the at least one human vital sign. 2. The non-touch thermometer of claim 1, wherein the signal processing module is further configured to amplify variations of the pixel values between the at least two images. 3. The non-touch thermometer of claim 1, wherein the signal processing is temporal processing. 4. The non-touch thermometer of claim 3, wherein the temporal processing is a bandpass filter. 5. The non-touch thermometer of claim 4, wherein the bandpass filter is configured to analyze frequencies over time. 6. The non-touch thermometer of claim 1, wherein applying signal processing includes spatial processing. 7. The non-touch thermometer of claim 1 wherein the wireless communication subsystem transmits to a static IP address in order to reduce IP discovery burden on the non-touch thermometer. 8. The non-touch thermometer of claim 1 wherein the wireless communication subsystem transmits through an IP tunnel. 9. The non-touch thermometer of claim 1 wherein a connection is established and data is pushed from the non-touch thermometer through the wireless communication subsystem, thereafter an external device controls flow of the data between the non-touch thermometer and the external device, wherein the connection further comprises: an authenticated communication channel. 10. The non-touch thermometer of c... | ARC DEVICES, LTD | G01K013/00 A61B005/00 A61B005/01 G06F019/00 A61B005/021 A61B005/026 A61B005/08 A61B005/024 A61B005/0205 G06T003/40 G06T007/00 H04W048/16 H04W076/02 A61B005/0402 A61B005/03 G01J005/00 H04W080/04 G01J005/02 G01J005/10 G08C017/02 H04L029/08 H04W012/06 H04B007/26 H04L012/46 H04N005/33 H04W048/08 G08B005/36 G08B021/18 G06K009/62 G06T007/174 G06T007/269 G06T007/90 A61B005/0245 H04W012/08
US09808168 | B2 | 2017-11-07 | Method and system for non-invasive measurement of cardiac parameters | 2014-11-17 | A method and system are presented for use in assessment of at least one cardiac parameter of an individual. An electrodes arrangement is applied to an individual's body, for applying an electrical field to the body and providing an electrical output indicative of a systolic impedance change and of a velocity of said change during a cardiac cycle. Also provided is additional data indicative of at least of the following conditions of the individual: a value of total peripheral resistance (TPR), a value of cardiac index (CI), and existence of the AHF condition. Data corresponding to these condition is analyzed to determine whether the TPR satisfies a first predetermined condition and/or the CI satisfies a predetermined second condition and/or whether the acute heart failure (AHF) condition is identified, to thereby use the data indicative of the measured electrical output and selectively calculate said at least one cardiac parameter based on either the systolic impedance change data or on said data of the velocity of the impedance change. | 1. An impedance measurement system for assessment of at least one cardiac parameter of an individual, the impedance measurement system comprising: a measurement unit comprising an electrodes arrangement comprising a first and second electrode sets, the first electrode set being configured and operable to attach to one arm at a first lateral side of the individual's body and the second electrode set being configured and operable to attach to one leg at a second lateral side of the individual's body, said electrodes arrangement being configured and operable to apply an electrical field to the body through one of the first and second electrode sets and to obtain an electrical output through the other of the first and second electrode sets, said electrical output comprising measured data indicative of both systolic impedance change, ΔR, and velocity of a systolic impedance change, δR/δt, of the individual during a cardiac cycle; anda control unit comprising:a data input utility configured and operable for receiving input data indicative of at least one condition of the individual, said at least one condition being indicative of whether the individual is sustained by an acute heart failure (AHF), and for receiving said measured data;a processor programmed for analyzing said input data indicative of said at least one condition of the individual, and upon identifying that said input data is indicative of that the individual is sustained by AHF, selecting the data of the velocity of systolic impedance change for calculating the at least one cardiac parameter of the individual, andupon identifying that said input data is indicative of that the individual is not sustained by AHF, either selecting the data of systolic impedance change for calculating the at least one cardiac parameter, or analyzing additional input data indicative of at least one other predetermined condition of the individual in order to determine whether to use the systolic impedance change data or the velocity of systolic impedance change data, said at least one other condition comprising value of at least one of a total peripheral resistance (TPR) and a cardiac index (CI); anda data output utility configured and operable for generating an output indicative of the calculation result. 2. The computer system of claim 1, wherein said at least one cardiac parameter is selected from the following: stroke volume, cardiac output, cardiac index, stroke index, heart rate. 3. The computer system of claim 1, wherein said processor utilizes first measured data of the systolic impedance change and second measured data indicative of at least an individual's blood pressure; in order to determine the individual's TPR. 4. The computer system of claim 1, wherein said CI is determined as CI=CO/BSA, where BSA is the individual's body surface area, and CO is the individual's cardiac output determined from the measured data. 5. The computer system of claim 1, wherein said at least one other predetermined ... | New N.I. Medical (2011) LTD. | A61B005/02 A61B005/04 A61B005/029 A61B005/053 A61B005/0205 A61B005/021 A61B005/0408 A61B005/0432 A61B005/044 A61B005/0295 A61B005/024
US09812846 | B2 | 2017-11-07 | Method and apparatus for performing optical imaging using frequency-domain interferometry | 2016-04-27 | Exemplary apparatus and method are provided. In particular, an electromagnetic radiation can be emitted with, e.g. a light source arrangement. For example, the light source arrangement can include a cavity and a filter, and a spectrum of the electromagnetic radiation can be controlled, e.g., with such cavity and filter, to have a mean frequency that changes (i) at an absolute rate that is greater than about 100 terahertz per millisecond, and (ii) over a range that is greater than about 10 terahertz. Additionally or alternatively, the light source arrangement can include a frequency shifting device which can shift the mean frequency of the electromagnetic radiation. | 1. An apparatus comprising: a light source arrangement configured to emit an electromagnetic radiation, wherein the light source arrangement includes a cavity and a filter which cause a spectrum of the electromagnetic radiation to have a mean frequency that changes (i) at an absolute rate that is greater than about 100 terahertz per millisecond, and (ii) over a range that is greater than about 10 terahertz. 2. The apparatus according to claim 1, wherein the mean frequency changes repeatedly at a repetition rate that is greater than 5 kilohertz. 3. The apparatus according to claim 2, wherein the spectrum has a tuning range whose center is approximately centered at 1300 nm. 4. The apparatus according to claim 2, wherein the spectrum has a tuning range whose center is approximately centered at 850 nm. 5. The apparatus according to claim 2, wherein the spectrum has a tuning range whose center is approximately centered at 1700 nm. 6. The apparatus according to claim 2, wherein the spectrum has an instantaneous line width that is smaller than 100 gigahertz. 7. The apparatus according to claim 1, wherein the light source arrangement further comprises a polygon arrangement which is adapted to receive at least one signal that is associated with the emitted electromagnetic radiation, and at least one of reflect and deflect the at least one signal to a further location. 8. The apparatus according to claim 1, wherein the light source arrangement further comprises a laser resonating system forming an optical circuit and configured to control a spatial mode of the electromagnetic radiation. 9. The apparatus according to claim 8, wherein the light source arrangement causes the electromagnetic radiation to propagate substantially unidirectionally within at least one portion of the resonating arrangement. 10. The apparatus according to claim 9, further comprising an optical circulator which is configured to control a direction of propagation of the electromagnetic radiation within the resonating arrangement. 11. The apparatus according to claim 8, wherein the filter is configured to at least one of transmit or reflect at least one portion of the electromagnetic radiation based on a frequency of the electromagnetic radiation, and wherein the at least one portion has a full-width-at-half-maximum frequency distribution which is less than about 100 GHz. 12. The apparatus according to claim 1, wherein the light source arrangement further comprises a gain medium. 13. An apparatus comprising: a light source arrangement configured to emit an electromagnetic radiation, wherein the light source arrangement includes a cavity and a filter which cause a spectrum of the electromagnetic radiation to have a mean frequency that changes at an absolute rate that is greater than about 100 terahertz per millisecond, wherein the light source arrangement further includes a frequency shifting device which shifts the mean frequency of the electromagnetic radiation. 14. The a... | THE GENERAL HOSPITAL CORPORATION | H01S003/08 H01S005/50 H01S005/00 H01S005/14 H01S005/10 A61B005/00 G01J003/02 G01J003/453 G01N021/47 G02B027/48 G01B009/02 G02B006/36 G02B026/12 G02F001/09 G02F001/11 H01S003/00 H01S003/106 G01J009/02 H01S003/105
US09788716 | B2 | 2017-10-17 | Method and device of testing color blindness | 2015-09-15 | Embodiments of the present invention disclose a method and a device of testing color blindness, which belong to the technical field of test. The method includes: acquiring a plurality of preset coloring schemes, each of which is provided to test one type of color blindness and comprises a base color and a changing color; acquiring color sense test patterns, each of which corresponding to one of the coloring schemes; changing a color of at least a part of the grids in one of color sense test patterns one by one from the base color to the changing color according to the grid color-changing sequence of the corresponding color sense test pattern during displaying the corresponding color sense test pattern, when the corresponding color sense test pattern is used to perform color blindness test on a tested person, and acquiring positions of the eyeballs of the tested person as the color of the grids is changing; obtaining a first fitting degree by fitting an acquired movement trace of the color point to a movement trace of the eyeballs; and, judging which type of color blindness of the tested person is according to the first fitting degree. According to the present invention, dependence analysis is performed on the movement trace of the eyeballs of the tested person and the movement trace of a color point in the corresponding color sense test pattern to obtain in quantitative sense which type of the color blindness of the tested person is, thereby increasing accuracy of the color blindness testing. | 1. A method of testing color blindness, the method comprising: acquiring a plurality of coloring schemes preset, each of which is provided to test one type of color blindness and comprises a base color and a changing color;acquiring color sense test patterns, each of which corresponds to one of the coloring schemes and includes a plurality of grids, each of the plurality of grids storing position coordinates thereof, the base color, the changing color and a grid color-changing sequence;changing a color of at least a part of the grids in one of the color sense test patterns one by one from the base color to the changing color according to the grid color-changing sequence of the corresponding color sense test pattern during displaying the corresponding color sense test pattern, when the corresponding color sense test pattern is used to perform color blindness test on a tested person, and acquiring positions of the eyeballs of the tested person as the color of the grids is changing;acquiring a movement trace of a color point, the movement trace of the color point being determined by the color-changing grid;acquiring a movement trace of the eyeballs based on the positions of the eyeballs of the tested person;obtaining a first fitting degree by fitting the movement trace of the color point to the movement trace of the eyeballs; andjudging which type of color blindness of the tested person is according to the first fitting degree. 2. The method according to claim 1, wherein the step of judging which type of color blindness of the tested person is according to the first fitting degree comprises: comparing the first fitting degree with a first preset numerical value;determining that the tested person does not belong to the type of color blindness tested by the corresponding color sense test pattern when the first fitting degree is greater than or equal to the first preset numerical value; anddetermining that the tested person belongs to the type of color blindness tested by the corresponding color sense test pattern when the first fitting degree is smaller than the first preset numerical value. 3. The method according to claim 1, wherein the method further comprises: making each of the color-changed grids flash for a specify time period at preset frequency, during the color of the at least part of the grids in the corresponding color sense test pattern changing one by one from the base color to the changing color, and acquiring brain wave signals received by the tested person when the color-changed grid flashes;performing a wave filtering process and a power spectrum analysis on the brain wave signals received by the tested person during the each of the color-changed grids flashing, and then considering a frequency value corresponding to a greatest power value in the power spectrum as a frequency value of the brain wave signals received by the tested person when the color-changed grid flashes;acquiring a color point flash curve that is determined by th... | BOE Technology Group Co., Ltd. | A61B003/02 A61B003/06 A61B003/113 A61B003/00 A61B005/0484 A61B005/048
US09779497 | B2 | 2017-10-03 | Measuring glomerular number from kidney MRI images | 2015-09-14 | Measuring the number of glomeruli in the entire, intact kidney using non-destructive techniques is of immense importance in studying several renal and systemic diseases. In particular, a recent Magnetic Resonance Imaging (MRI) technique, based on injection of a contrast agent, cationic ferritin, has been effective in identifying glomerular regions in the kidney. In various embodiments, a low-complexity, high accuracy method for obtaining the glomerular count from such kidney MRI images is described. This method employs a patch-based approach for identifying a low-dimensional embedding that enables the separation of glomeruli regions from the rest. By using only a few images marked by the expert for learning the model, the method provides an accurate estimate of the glomerular number for any kidney image obtained with the contrast agent. In addition, the implementation of our method shows that this method is near real-time, and can process about 5 images per second. | 1. A method of counting glomeruli in an image of a kidney, the method being implemented via execution of computer instructions configured to run at one or more processors and configured to be stored at one or more non-transitory computer-readable media, the method comprising: extracting a second patch centered around each pixel of training images;identifying the second patches that are centered at positive pixels and the second patches that are centered at negative pixels using expert-marked ground truth images;constructing l1 graphs to model inter-class and intra-class relationships;performing local discriminant embedding;analyzing patches corresponding to each pixel within the image;determining, using the patches, whether each pixel within the image belongs to at least one of the glomeruli; andcounting the glomeruli in the image. 2. The method of claim 1, further comprising: using graph-based embedding to exploit similarities between the patches. 3. The method of claim 1, further comprising: considering relations between glomerular and non-glomerular regions. 4. The method of claim 1, further comprising: adopting a non-local subspace approach. 5. The method of claim 1, wherein: the method is robust to measurement noise in the image. 6. The method of claim 1, wherein: projection directions are pre-computed using training patches. 7. The method of claim 1, further comprising: for each patch in the image, obtaining low dimensional embeddings using an inner product operation. 8. The method of claim 1, further comprising: performing glomeruli identification using K-Means clustering. 9. The method of claim 1, further comprising: displaying at least a portion of the glomeruli on a screen. 10. The method of claim 1, wherein: the l1 graphs are made robust to intensity variations across several images. 11. The method of claim 10, wherein: weights in the l1 graphs are obtained based on sparse representations of the patches. 12. A method of counting glomeruli in an image of a kidney, the method being implemented via execution of computer instructions configured to run at one or more processors and configured to be stored at one or more non-transitory computer-readable media, the method comprising: extracting a second patch centered around each pixel of training images;identifying the second patches that are centered at positive pixels and the second patches that are centered at negative pixels using expert-marked ground truth images;constructing l1 graphs to model inter-class and intra-class relationships;performing local discriminant embedding;extracting a first patch centered around each pixel of the image;computing a low-dimensional projection for the each first patch;performing segmentation using K-Means clustering to obtain independent regions; andcounting a total number of the independent regions. 13. The method of claim 12, wherein: the first patches are each 5 pixels by 5 pixels. 14. The method of claim 12, further comprising: normaliz... | ARIZONA BOARD OF REGENTS, A BODY CORPORATE OF THE STATE OF ARIZONA, ACTING FOR AND ON BEHALF OF ARIZONA STATE UNIVERSITY | G06K009/00 G06T007/00 A61B005/055 A61B005/20 A61B005/107 G06K009/62 G06T011/60 A61P013/12
US09808612 | B2 | 2017-11-07 | Conduit connector for a patient breathing device | 2012-08-10 | In an embodiment, a connector or connector assembly for attaching a nasal cannula with a gas delivery hose includes a sensor port for a sensor probe positioned near an end of a nasal cannula, which can allow the sensor probe to be placed closer to the patient's nostrils than previous connector parts allowed. The connector can be configured to advantageously allow the nasal cannula to rotate relative to the gas delivery hose, thereby allowing a patient or healthcare provider to untangle or otherwise straighten the hose or the cannula. The connector assembly can be configured to automatically align locking protrusions on a first component with locking recesses on a second component, where insertion of the second component within the first component causes the second component to rotate relative to the first component, thereby aligning the locking protrusions with associated locking recesses. | 1. A connector assembly for a gas delivery conduit for providing gases to a patient, the connector assembly comprising: a first conduit connector, anda connecting adapter configured to releasably connect the first conduit connector to a second conduit connector, the connecting adapter comprising: a first adapter conduit body; andone or more locking fingers configured to releasably attach to the first conduit connector, the one or more locking fingers protruding longitudinally from the first adapter conduit body; andthe first conduit connector comprising: a first connector conduit body;one or more locking tabs configured to releasably engage with the one or more locking fingers when the connecting adapter is received into a terminal aperture of the first conduit connector, the one or more locking tabs protruding radially inward from an interior surface of the first connector conduit body; andan alignment tab protruding radially inward from the interior surface of the first connector conduit body, the alignment tab configured to self-align the one or more locking fingers with the one or more locking tabs when the connecting adapter is received into the terminal aperture. 2. The connector assembly of claim 1, wherein the first conduit connector further comprises a sensor port configured to receive a sensor probe, the sensor port formed substantially perpendicularly adjacent to the terminal aperture. 3. The connector assembly of claim 2, wherein the sensor port provides a sensor aperture into the first conduit connector, the sensor aperture adjacent to an insertion aperture formed by the one or more locking fingers when the connecting adapter is received into the terminal aperture, the sensor aperture configured to allow the sensor probe to extend into the insertion aperture. 4. The connector assembly of claim 3, wherein the one or more locking fingers extend past the sensor aperture when the connecting adapter is received into the terminal aperture. 5. The connector assembly of claim 1, wherein the first conduit connector further comprises a receptacle configured to receive a sensor probe, the receptacle configured to position the sensor probe to extend into an insertion aperture formed by the one or more locking fingers when the connecting adapter is received into the terminal aperture. 6. The connector assembly of claim 1, wherein engagement of the one or more locking tabs with the one or more locking fingers generates an audible sound. 7. The connector assembly of claim 1, wherein the connecting adapter is configured to provide a quick connect mechanism between the first conduit connector and the second conduit connector. 8. A connector for a gas delivery conduit for providing gases to a patient, the connector comprising: a conduit body comprising a terminal aperture, the terminal aperture configured to receive a connecting adapter into the conduit body;a sensor port formed substantially perpendicularly on the conduit body, the sensor port... | Fisher & Paykel Healthcare Limited | A61M039/10 A61M016/06 A61M016/08 A61M016/16 A61B005/08 A61M016/00 A61M016/10
US09782081 | B2 | 2017-10-10 | Photoacoustic apparatus | 2015-05-12 | A photoacoustic apparatus disclosed herein includes a plurality of receiving elements which receive photoacoustic waves caused by irradiation of light from a light source to a subject and outputs reception signals and a signal data acquisition unit which generates and stores reception signal data by reducing a data amount of reception signals output from the plurality of receiving elements when a region where the directional axes of the plurality of receiving elements are gathered at a point is not included in a region of interest. | 1. A photoacoustic apparatus comprising: a light source;a plurality of receiving elements which receive photoacoustic waves caused by irradiation of light from the light source to a subject and output reception signals;a base material which supports the plurality of receiving elements such that the directional axes of the plurality of receiving elements are gathered at a point;a control unit which determines a position of the base material at each of a plurality of time points when light from the light source is irradiated to the subject;a moving mechanism which moves the base material to the position of the base material determined by the control unit;a signal data acquisition unit which generates and stores reception signal data based on the reception signal output from the plurality of receiving elements;a setting unit which sets a region of interest; andan information acquisition unit which acquires subject information in the region of interest based on the reception signal data stored in the signal data acquisition unit,wherein the signal data acquisition unit generates and stores the reception signal data by reducing a data amount of the reception signals output from the plurality of receiving elements if a region where the directional axes of the plurality of receiving elements are gathered at a point is not included in the region of interest. 2. The photoacoustic apparatus according to claim 1, wherein the signal data acquisition unit does not store the reception signals output from the plurality of receiving elements when the region where the directional axes of the plurality of receiving elements are gathered at a point is not included in the region of interest. 3. The photoacoustic apparatus according to claim 1, wherein the signal data acquisition unit generates and stores the reception signal data by lowering a sampling frequency for the reception signal output from the plurality of receiving elements when the region where the directional axes of the plurality of receiving elements are gathered at a point is not included in the region of interest, as compared to a sampling frequency for the reception signal output from the plurality of receiving elements when the region where the directional axes of the plurality of receiving elements are gathered at a point is included in the region of interest. 4. The photoacoustic apparatus according to claim 1, wherein for a target position within the region of interest, the information acquisition unit acquires the subject information at the target position within the region of interest without using reception signal data, among the reception signal data stored in the signal data acquisition unit, corresponding to reception signals output from the plurality of receiving elements, when the region where the directional axes of the plurality of receiving elements are gathered at a point is not included in the target position. 5. The photoacoustic apparatus according to claim 1, wherein the sign... | Canon Kabushiki Kaisha | A61B005/00
US09782120 | B2 | 2017-10-10 | Cannulated screw probe | 2014-05-02 | Concepts for measuring and monitoring characteristics of osseous tissue are presented. In various embodiments, a monitoring probe may be inserted through a hollow passageway within a cannulated screw during an operation in order to measure characteristics of the osseous tissue in which the cannulated screw is inserted. In various embodiments, the cannulated screw may be inserted into a femoral head, through a physis, and into an epiphysis for stabilization of Slipped Capital Femoral Epiphysis. During surgical treatment thereof, a pressure transducer may be inserted into the osseous tissue through the cannulated screw in order to monitor epiphysis perfusion through the femoral head. | 1. A method for monitoring osseous tissue characteristics, the method comprising the steps of: inserting a cannulated screw such that at least a portion of the cannulated screw extends through a first osseous region and into a second osseous region, wherein the cannulated screw comprises a body and a passageway extending longitudinally from a first end of the body to a second end of the body;inserting a measurement probe into the second osseous region through the passageway of the cannulated screw;measuring osseous tissue characteristics of the second osseous region using the inserted measurement probe;recording waveforms in the measured osseous tissue characteristics via an osseous tissue monitoring device in electronic communication with the measurement probe;measuring a heart rate via a heart rate monitoring device;recording a heart rate waveform via the heart rate monitoring device;comparing, via the osseous tissue monitoring device, the heart rate waveform to an osseous tissue characteristics waveform; anddetermining, via the osseous tissue monitoring device, the presence of perfusion in the second osseous region based on the comparison between the heart rate waveform and the osseous tissue characteristics waveform. 2. The method of claim 1, wherein the second osseous region comprises an epiphysis region located in a femoral head. 3. The method of claim 1, wherein the osseous tissue characteristics comprise at least one of pressure, temperature, or oxygen level. 4. The method of claim 1, further comprising the step of: inserting a guide wire such that at least a portion of the guide wire extends through the first osseous region and into the second osseous region; andafter inserting the cannulated screw, removing the guide wire; andwherein the cannulated screw is positioned based at least in part on the guide wire. 5. The method of claim 1, wherein the heart rate waveform comprises a heart rate period and the osseous tissue characteristics waveform comprises an osseous tissue characteristics waveform period; and the comparing step comprises the step of comparing the heart rate period and the osseous tissue characteristics waveform period and determining if the heart rate period and waveform period are substantially equivalent. 6. The method of claim 1, further comprising the steps of: if the comparison step indicates that perfusion is not present in the second osseous region, performing an operation to establish perfusion in the second osseous region; andmeasuring the osseous tissue characteristics of the second osseous region after performing the operation. 7. A system for monitoring osseous tissue characteristics comprising: a cannulated screw comprising a body having a passageway formed therein extending longitudinally from a first end of the body to a second end of the body, wherein the body includes a head portion proximate the first end and a threaded portion and configured for insertion into osseous tissue;a measurement probe con... | - | A61B005/02 A61B005/00 A61B005/0245 A61B005/026 A61B017/74
US09782094 | B2 | 2017-10-10 | Identifying ambiguous cardiac signals for electrophysiologic mapping | 2015-07-31 | A system and associated method is disclosed for determining whether a signal is ambiguous. The system comprises a plurality of electrodes configured to be located proximate tissue of a patient. A display apparatus comprises a graphical user interface configured to present information to a user. A computing apparatus is coupled to the electrode apparatus. The computing apparatus is configured to calculate a first derivative of a signal acquired from the electrodes, determine a first minimum from the first derivative, determine a second minimum from the second derivative and a second index within a window, calculate a ratio of the first and second derivatives, calculate a difference between a first and second index. In response to determining the ratio and the difference between the first and second index, the display apparatus displays whether the signal is ambiguous. | 1. A system for use in monitoring cardiac signals comprising: an electrode apparatus comprising a plurality of electrodes configured to be located proximate tissue of a patient;a display apparatus comprising a graphical user interface, wherein the graphical user interface is configured to present information for use in assisting a user;computing apparatus coupled to the electrode apparatus and display apparatus, wherein the computing apparatus is configured to determine whether a signal acquired from the electrode apparatus is ambiguous, the computing apparatus configured to:calculate a first derivative of the signal;identify a first minimum value of the first derivative and a first time index associated therewith;identify a second minimum value of the first derivative and a second time index associated therewith, within a time window that also includes the first minimum value;calculate a ratio of the identified first and second minimum values of the first derivative;calculate a difference between the first and second time indexes;identify whether the signal is ambiguous or non-ambiguous based on the ratio and the difference between the first and second time indexes; anddisplay on the display apparatus an indication of the signal as ambiguous or non-ambiguous. 2. A system according to claim 1, wherein the computing apparatus is further configured to compare the difference between the first and second indexes to a first preselected number and, in response to determining that the difference between the first and second time indexes is greater than or equal to the first preselected number, to display on the display apparatus an indication that the signal is potentially ambiguous. 3. A system according to claim 2, wherein the computing apparatus is configured to compare the ratio to a second preselected number and, in response to determining that the ratio is less than or equal to the second preselected number, to display on the display apparatus an indication that the signal is confirmed as being ambiguous. 4. A system according to claim 2, wherein the computing apparatus is configured to compare the ratio to a third preselected number and, in response to determining that the ratio is greater than the third preselected number, to display on the display apparatus an indication that the signal is confirmed as being non-ambiguous. 5. A system according to claim 1, wherein the computing apparatus is further configured to compare the difference between the first and second time indexes to a fourth preselected number and, in response to determining the difference is less than the fourth preselected number, to display on the display apparatus an indication that the signal is non-ambiguous. 6. A system according to claim 1, wherein the computing apparatus is configured to compare the ratio to a fifth preselected number and, in response to determining that the ratio is greater than the fifth preselected number, to display on the display apparatus an ind... | Medtronic, Inc. | A61B005/00 A61B005/04 A61B005/0408 A61B005/0456 A61N001/365
US09795359 | B2 | 2017-10-24 | System and method for measuring fluidics in arteries | 2016-11-18 | An apparatus for performing measurements in an artery includes a Doppler catheter, comprising a Doppler transducer and a wire connected to the Doppler transducer; and an elongated catheter body having a conduit therein for housing the wire of the Doppler catheter. The body has a proximal end and a distal end; wherein the wire is movable in the conduit relative to the catheter body so that the Doppler transducer and the wire are capable of being threaded into said conduit at the proximal end after the distal end of the catheter has been inserted into the artery, until the Doppler transducer emerges outside the conduit at the distal end of said body for performing ultrasound measurements in the artery. Another embodiment employs at least an additional side transducer for measuring the cross-sectional dimension of the artery useful for computing blood flow. | 1. An apparatus for performing measurements in an artery, comprising: a Doppler transducer;a wire, capable of carrying an electrical signal along a length of the wire, the wire having a distal end connected to the Doppler transducer at a distal portion of the wire, and the wire having a proximal end;a first connector having an opening through which the wire would pass during operation of the apparatus;a second connector at the proximal end of the wire, suitable for electrical connections to the wire;a catheter body having a proximal end and a distal end and having a conduit therein for containing portions of the wire;an enclosing sleeve having a first position and a second position wherein, when in the first position, the enclosing sleeve encloses substantially all of the wire save for the distal portion of the wire connected to the Doppler transducer, and wherein, when in the second position, the enclosing sleeve does not enclose a majority of the wire and unenclosed portions of the wire would pass into the conduit of the catheter body;a third connector attached to the catheter body at the proximal end of the catheter body, shaped to mate with the first connector to form a stable connection between the catheter body and the enclosing sleeve;wherein, after the stable connection has been formed, the wire and Doppler transducer are movable in the conduit through the stable connection with movement of the Doppler transducer being from around the proximal end of the catheter body to beyond the distal end of the catheter body, whereby the Doppler transducer can emerge outside the conduit at the distal end of the catheter body for performing ultrasound measurements in the artery; andwherein the enclosing sleeve provides a barrier between the wire and an environment when the wire is withdrawn from the catheter body. 2. The apparatus of claim 1, wherein the enclosing sleeve is shorter than the wire such that when the wire is withdrawn from the conduit, the length of the enclosing sleeve limits an extent by which the wire can be withdrawn from the conduit. 3. The apparatus of claim 1, further comprising a cap for connection to the first connector when the Doppler transducer and wire are not connected to the catheter body, to enclose and protect the Doppler transducer. 4. The apparatus of claim 1, wherein the first connector provides centering for the Doppler transducer for when the Doppler transducer is inserted into the conduit of the catheter body. 5. The apparatus of claim 1, wherein the third connector comprises: a first port that permits the Doppler transducer and the wire to pass there through and be inserted into the conduit;a second port coupled to the proximal end of the catheter body; anda third port providing a clear path enabling blood pressure in the artery near the distal end of the catheter body to be monitored through the second port whether or not the Doppler transducer and the wire are in the conduit. 6. The apparatus of claim 1, wh... | Nadarasa Visveshwara | A61B008/12 A61B008/06 A61B008/08 A61B005/00 A61B008/00 A61B005/021 A61B005/0205 A61B005/15
US09788791 | B2 | 2017-10-17 | Patient monitoring system and method | 2006-06-01 | The present invention relates to a patient monitoring system for monitoring cardio pulmonary performance or the like by means of capacitive measurement. Further the invention relates to a method for monitoring cardio pulmonary performance or the like by means of a capacitive measurement. In order to provide a reliable technique for monitoring cardio pulmonary performance or the like, which technique is particularly suitable for home use, a patient monitoring system (1) for monitoring cardio pulmonary performance or the like by means of capacitive measurement is suggested, said system (1) comprising a number of electrodes (2) arranged in the form of a matrix (3), said electrodes (2) being adapted to be integrated with a bed (4) or the like, each electrode (2) being individually selectable, means (5) for determining a number of electrodes (2) depending on the position of the patient on the bed (4) or the like, said means (5) being adapted to determine the number of electrodes (2) by determining the capacitance of a number of electrodes (2), and means (5) for selecting a number of said determined electrodes (2) for carrying out a capacitive measurement. | 1. A patient monitoring system for monitoring cardio pulmonary performance by means of capacitive measurement, said system comprising: a plurality of electrodes arranged in the form of a matrix, the plurality of electrodes being adapted to be integrated with a bed, each electrode of the plurality of electrodes being individually selectable to perform a capacitive measurement, wherein a sensitivity of each electrode of the plurality of electrodes in the matrix is calibrated to filter out parasitic signals from measured cardiopulmonary electrical signals,a dielectric comprised of an elastic composition, the dielectric having a first side facing the plurality of electrodes and a second side on which one or more counter electrodes are arranged, the second side being opposite the first side, wherein a stress is detected in the dielectric if a patient is situated on the bed, the stress affecting a capacitance between an electrode of the plurality of electrodes and the one or more counter electrodes, anda computer system programmed by computer-readable instructions that, when executed, cause the computer system to: determine a capacitance between each electrode of the plurality of electrodes and the one or more counter electrodes;determine suitable electrodes from the plurality of electrodes based on increases in the capacitances between each electrode of the plurality of electrodes and the one or more counter electrodes;select a subset of the suitable electrodes for cardio pulmonary performance monitoring, wherein the selection is based on a suitable electrode being within a predetermined distance from at least one other suitable electrode, the selected subset of the suitable electrodes comprising two or more electrodes; andmonitor, via the selected subset of the suitable electrodes, cardio pulmonary performance of the patient, wherein the monitoring comprises one or more of electrocardiogram-based monitoring, electromyogram-based monitoring, electroencephalogram-based monitoring, or bio-impedance measurement-based monitoring. 2. The patient monitoring system as claimed in claim 1, wherein the computer system is further caused to detect, with a bridge circuit, the increases in the capacitances between each electrode of the plurality of electrodes and the one or more counter electrodes. 3. The patient monitoring system as claimed in claim 1, wherein determining the capacitance comprises: supplying, with a voltage supply, a voltage signal to the plurality of electrodes. 4. The patient monitoring system as claimed in claim 1, comprising a shielding adapted to be located between the patient and the dielectric. 5. The patient monitoring system as claimed in claim 1, wherein (i) the body of the patient or (ii) an object placed on the dielectric acts as the one or more counter electrodes. 6. The patient monitoring system as claimed in claim 1, wherein the dielectric comprises a textile bed sheet. 7. The patient monitoring system as claimed in claim 1, w... | KONINKLIJKE PHILIPS N.V. | A61B005/08 A61B005/00 A61B005/0408 A61B005/113
US09826934 | B2 | 2017-11-28 | Non-invasive intracranial pressure system | 2012-09-17 | Non-invasive intracranial pressure detection and/or monitoring and use of data with respect thereto. Illustratively, with respect to a method, there can be a method to digitally produce and communicate intracranial pressure data from skull deformation electric signals, the method including: receiving, from at least one sensor, detected skull deformation electric signals at electrical equipment configured to transform and process the skull deformation signals that are received; transforming and processing, by the electrical equipment, the received skull deformation electric signals to produce digital intracranial pressure data; and outputting, by the electrical equipment, the digital intracranial pressure data via an output device operably associated with the electrical equipment to render the digital intracranial pressure data. | 1. An apparatus to digitally produce and communicate intracranial pressure data from skull deformation electric signals bearing intracranial pressure information combined with cardiac and respiratory influences, the apparatus including: at least one non-invasive skull deformation sensor having a pin adapted for direct contact with a patient's head, the pin being coupled to a cantilever bar that transmits forces to a transducer that converts the transmitted forces into said skull deformation electrical signals, the pin and transducer each being coupled to the cantilever bar such that expansion and contraction of the patient's skull causes movement of the pin, producing the transmitted forces that are converted into said skull deformation electric signals;electrical equipment arranged to receive the skull deformation electric signals produced by the at least one non-invasive sensor and configured to transform and process the skull deformation signals that are received to produce digital intracranial pressure data;an output device operably associated with the electrical equipment so as to output the digital intracranial pressure data, wherein the electrical equipment includes: an amplifier operably associated with the sensor and arranged to amplify the skull deformation electric signals that are received from said at least one sensor to produce amplified analog skull deformation signals;an analog-to-digital converter operably associated with the amplifier and arranged to convert the amplified skull deformation signals from analog form into digital form skull deformation electric signals;a processor operably associated with the analog-to-digital converter and arranged and programmed to apply a Fourier Transform, a Fast Fourier Transform, or both on the digital form skull deformation electric signals to produce the digital intracranial pressure data by extracting the digital intracranial pressure data while excluding cardiac and respiratory influences on the basis of frequency; anda memory operably associated with the processor and arranged to store the digital intracranial pressure data in a database. 2. The apparatus of claim 1, wherein said at least one sensor includes said at least one strain gage. 3. The apparatus of claim 2, wherein the output device includes a monitor. 4. The apparatus claim 2, wherein the detected skull deformation electric signals correspond to a human or an animal. 5. The apparatus claim 4, wherein said at least one sensor is noninvasively locatable by a strap, band, hat, or helmet. 6. The apparatus of claim 4, wherein said at least one sensor is noninvasively locatable by a stationary apparatus which substantially fixes the patient's position with respect to said at least one sensor. 7. The apparatus of claim 4, wherein said at least one sensor comprises a plurality of sensors, and the electrical equipment transforms the received skull deformation electric signals from each of said plurality of sensors. 8. The appara... | Braincare Desenvolvimento E Inovação Tecnológica LTDA | A61B005/00 A61B005/03
US09808170 | B2 | 2017-11-07 | Electrode with charge-operated indicator | 2013-03-15 | An electrode includes separate first and second electrical contacts to contact the skin of a subject. A charge-holding structure is electrically connected between the first and second contacts. An indicator is operatively coupled to the charge-holding structure so that the indicator changes visibly in response to a change in the charge stored in the charge-holding structure. The electrode can include a rectifier across the contacts. A container for electrodes includes an electrical supply and a plurality of receptacles for electrodes so that a voltage difference is maintained across conductors of each receptacle (and contacts of an electrode therein) for at least one week. A method of making electrodes includes arranging the contacts over a support, connecting the charge-holding structure between them, arranging the indicator over the support, and charging the charge-holding structure so that the indicator has a first visual appearance. | 1. An electrode, comprising: a support having a subject side and a non-subject side opposite the subject side, wherein when the electrode is removably attached to a skin surface of a subject, the subject side is oriented generally facing the subject;first and second electrical contacts adapted to contact the skin surface of the subject, wherein at least a portion of the support spaces the first and second electrical contacts apart from the non-subject side;a charge-holding structure of the electrode electrically connected, in series, between the first and second electrical contacts and configured to store a charge, wherein: the charge-holding structure is at least partly (1) embedded in the support, (2) disposed on the subject side of the support, or (3) disposed on the non-subject side of the support; andthe charge-holding structure is configured to, in response to an electrical connection between the first electrical contact and the second electrical contact through the subject, dissipate at least some of the charge stored in the charge-holding structure through the subject; andan indicator arranged so that at least a portion of the support spaces the indicator away from the subject side of the support, wherein the indicator is: electrically connected to the charge-holding structure; andconfigured to visibly change state in response to a change in the charge stored in the charge-holding structure. 2. The electrode according to claim 1, wherein: the charge-holding structure comprises spaced-apart first and second conductors disposed at least partly in or on the support and electrically connected to the first and second electrical contacts, respectively; andthe at least one indicator comprises an electronic ink located at least partly between the first and second conductors and controlled at least partly by a magnitude of charge between the first and second conductors. 3. The electrode according to claim 2, wherein: the electronic ink is at least partly (1) embedded in the support, or (2) contained by the support; andat least one of the first conductor or the second conductor is spaced apart from the non-subject side by at least part of the support. 4. The electrode according to claim 3, wherein: the second conductor is spaced apart from the subject side by at least part of the support; andthe electrode further comprises a via through the support that electrically connects the second conductor to the second electrical contact. 5. The electrode according to claim 1, wherein the second electrical contact surrounds the first electrical contact. 6. The electrode according to claim 1, wherein the first and second electrical contacts are positioned so that neither surrounds the other. 7. The electrode according to claim 1, wherein the non-subject side of the support comprises an aperture arranged with respect to the indicator so that, when the electrode is removably attached to the skin surface, the indicator is visible through the aperture. 8. ... | Welch Allyn, Inc. | A61B005/04 A61B018/16 A61B005/0408 A61B005/01 A61B005/00
US09799348 | B2 | 2017-10-24 | Systems and methods for an automatic language characteristic recognition system | 2016-01-15 | In some embodiments, a method of creating an automatic language characteristic recognition system. The method can include receiving a plurality of audio recordings. The method also can include segmenting each of the plurality of audio recordings to create a plurality of audio segments for each audio recording. The method additionally can include clustering each audio segment of the plurality of audio segments according to audio characteristics of each audio segment to form a plurality of audio segment clusters. Other embodiments are provided. | 1. A method of creating an automatic language characteristic recognition system, the method being implemented via execution of computer instructions configured to run at one or more processors and configured to be stored at one or more non-transitory computer-readable media, the method comprising: receiving a plurality of audio recordings;segmenting each of the plurality of audio recordings to create a plurality of audio segments for each audio recording;clustering each audio segment of the plurality of audio segments according to audio characteristics of the each audio segment to form a plurality of audio segment clusters, wherein the audio characteristics of the each audio segment used in the clustering comprise at least one of a pitch of sound in the each audio segment, a duration of the sound in the each audio segment, a rhythm of the sound in the each audio segment, or an organization of the sound in the each audio segment; andgenerating an age-based model in a data store that associates the plurality of audio segment clusters to weightings for specific ages of those represented in the plurality of audio recordings. 2. The method of claim 1 further comprising: transforming the data store to store the plurality of audio segment clusters. 3. The method of claim 1, wherein: clustering the each audio segment of the plurality of audio segments comprises using an unsupervised clustering method. 4. The method of claim 1, wherein: clustering the each audio segment of the plurality of audio segments comprises using K-means clustering. 5. The method of claim 1, wherein: a quantity of the plurality of audio segment clusters is at least 64. 6. The method of claim 1, wherein clustering the each audio segment of the plurality of audio segments comprises: initiating a plurality of center points; anduntil a convergence criterion is met, repeating: for each audio segment of the plurality of audio segments: measuring a distance from the plurality of center points to the audio segment; andassigning the audio segment to a center point of the plurality of center points that is closest to the audio segment; andfor each center point of the plurality of center points: averaging the audio segments assigned to the center point to obtain an average; andreassigning the center point to the average. 7. The method of claim 1 further comprising: transforming the data store to store the plurality of audio segment clusters,wherein: clustering the each audio segment of the plurality of audio segments comprises using an unsupervised clustering method. 8. The method of claim 1 further comprising: transforming the data store to store the plurality of audio segment clusters,wherein: clustering the each audio segment of the plurality of audio segments comprises using K-means clustering. 9. The method of claim 1 further comprising: transforming the data store to store the plurality of audio segment clusters,wherein: clustering the each audio segment of the plurality of audi... | LENA FOUNDATION | G10L015/00 G10L013/00 G10L025/66 G10L017/16 G10L015/02 G10L017/26 A61B005/16 A61B005/00 G10L015/06
US09801686 | B2 | 2017-10-31 | Neural monitor-based dynamic haptics | 2015-03-30 | A computer-assisted surgery system may have a robotic arm including a surgical tool and a processor communicatively connected to the robotic arm. The processor may be configured to receive, from a neural monitor, a signal indicative of a distance between the surgical tool and a portion of a patient's anatomy including nervous tissue. The processor may be further configured to generate a command for altering a degree to which the robotic arm resists movement based on the signal received from the neural monitor; and send the command to the robotic arm. | 1. A computer-implemented method for controlling a surgical system, the method comprising: receiving, from a neural monitor, a signal indicative of a distance between a surgical tool connected to a robotic arm and a portion of a patient's anatomy;receiving a joint angular velocity of one or more joints of the robotic arm;determining a neural monitor gain based on the signal received from the neural monitor;generating a first force value proportional to the neural monitor gain and the joint angular velocity of one or more joints of the robotic arm and a second force value based on a relationship of the surgical tool with a repulsive virtual haptic geometry associated with the patient's anatomy; andgenerating a command to control the surgical system by altering a degree to which the robotic arm resists movement by combining the first force value and the second force value. 2. The computer-implemented method of claim 1, wherein the virtual haptic geometry defines a volume associated with a sensitive portion of the patient's anatomy. 3. The computer-implemented method of claim 2, wherein the virtual haptic geometry is configured to generate haptic feedback forces that repulse the surgical tool away from the sensitive portion of the patient's anatomy associated with the virtual haptic geometry. 4. The computer-implemented method of claim 1, the method further comprising: providing an electrical potential to the surgical tool;measuring an electromyographic signal at another portion of the patient's anatomy innervated by the portion of the patient's anatomy; andgenerating the signal indicative of the distance between the surgical tool and the portion of the patient's anatomy based on the electromyographic signal. 5. The computer-implemented method of claim 2, wherein the portion of the patient's anatomy is a portion of nervous tissue. 6. A computer-assisted surgery system comprising: a robotic arm including a surgical tool;a processor communicatively connected to the robotic arm and configured to: receive, from a neural monitor, a distance between the surgical tool connected to the robotic arm and a portion of a patient's anatomy;receive a joint angular velocity of one or more joints of the robotic arm;determine a neural monitor gain based on the signal received from the neural monitor;generate a first force value proportional to the neural monitor gain and the joint angular velocity of one or more joints of the robotic arm and a second force value based on a relationship of the surgical tool with a repulsive virtual haptic geometry associated with the patient's anatomy; andgenerate a command to control the surgical system by altering a degree to which the robotic arm resists movement by combining the first force value and the second force value. 7. The computer-assisted surgery system of claim 6, further comprising the neural monitor, wherein the neural monitor is configured to: provide an electrical potential to the surgical tool;measure an elec... | MAKO Surgical Corp. | G05B015/00 G05B019/00 A61B019/00 A61B005/04 A61B005/00 A61B005/0488 A61B017/17 A61B017/16 A61B017/86 A61F002/30 G06F003/01 A61B017/00 A61F002/38 A61B005/11 A61F002/46 A61B017/70
US09801568 | B2 | 2017-10-31 | Gait pattern analysis for predicting falls | 2015-01-07 | A method for acquiring gait parameters of an individual is disclosed. The method includes capturing calibration images from foot marker placed on feet or shoes of an individual while an individual is standing still, the calibration images are obtained from a camera worn by the individual, capturing subsequent time-varying images from the foot markers while the individual is walking, and comparing the calibration images to the subsequent time-varying images by a processing unit that is coupled to the camera to determine changes between the initial relative image size of the foot markers and the time-varying images of the foot markers as a function of time to analyze gait of the individual. | 1. A method for acquiring gait parameters of an individual, comprising: capturing calibration images from foot markers placed on feet or shoes of an individual while an individual is standing still, the calibration images are obtained from a camera worn by the individual;capturing subsequent time-varying images from the foot markers while the individual is walking; andcomparing the calibration images to the subsequent time-varying images by a processing unit that is coupled to the camera to determine changes between the initial relative image size of the foot markers and the time-varying images of the foot markers as a function of time to analyze gait of the individual. 2. The method of claim 1, the camera is an integral part of the processing unit. 3. The method of claim 1, the camera physically coupled to the processing unit. 4. The method of claim 1, the camera electronically coupled to the processing unit. 5. The method of claim 4, the coupling is by a wireless channel. 6. The method of claim 1, each of the foot markers includes at least one identifiable feature such that the foot marker can be identified in the images. 7. The method of claim 6, the identifiable feature is a color. 8. The method of claim 7, the color is a solid color. 9. The method of claim 6, the identifiable feature is a pattern. 10. The method of claim 9, the pattern is a checker pattern. 11. A smart gait analysis system, comprising: a camera worn by an individual;a processing unit coupled to the camera;a left foot marker placed on the left shoe or foot of the individual; anda right foot marker placed on the right shoe or foot of the individual,the camera is configured to acquire images from the foot markers as the individual is walking,the processing unit is configured to capture calibration images from the foot marker while an individual is standing still obtained from the camera,capture subsequent time-varying images from the foot markers while the individual is walking, andcompare the calibration images to the subsequent time-varying images to determine changes between the initial relative image size of the foot markers and the time-varying images of the foot markers as a function of time to analyze gait of the individual. 12. The gait analysis system of claim 11, the camera is an integral part of the processing unit. 13. The gait analysis system of claim 11, the camera physically coupled to the processing unit. 14. The gait analysis system of claim 11, the camera electronically coupled to the processing unit. 15. The gait analysis system of claim 14, the coupling is by a wireless channel. | PURDUE RESEARCH FOUNDATION | A61B005/00 A61B005/11
US09811818 | B1 | 2017-11-07 | Wearable personal digital device for facilitating mobile device payments and personal use | 2017-05-19 | Provided is a wearable personal digital device for point of healthcare saliva testing. The wearable personal digital device may comprise a processor, a display, biometric sensors, activity tracking sensors, a memory unit, a communication circuit, a housing, an input unit, a projector, a timepiece unit, a haptic touch control actuator, a band, a mounting clip, a saliva sample insert apparatus, a pinhole, a light emitting diode board, a battery, and a set of sensors. The processor may be operable to receive data from an external device, provide a notification to a user based on the data, receive a user input, and perform a command selected based on the user input. The housing may be adapted to enclose the components of the wearable personal digital device. The band may be adapted to attach to the housing and secure the wearable personal digital device on a user body. | 1. An artificial intelligence (AI) wearable and mobile personal digital system for facilitating mobile device payments, personal saliva testing, personal use, and health care, and being a point of healthcare (POH) saliva testing device for POH saliva testing, the AI wearable and mobile personal digital system comprising a mobile device physically connected with a portable saliva testing device, the portable saliva testing device comprising: a mounting clip connecting the portable saliva testing device to the mobile device in front of a camera of the mobile device;a saliva sample insert apparatus, wherein the saliva sample insert apparatus is configured to receive a saliva sample;a pinhole;a light-emitting diode (LED) board;a battery; anda set of sensors, wherein the set of sensors includes at least a biosensor including a thin-film silicon photonic biosensor, the biosensor including at least a film made of nanoparticles and polymers, wherein the biosensor uses beams of light to detect changes in a composition of the saliva sample of the user placed on the film, each of the saliva sample comprising a biological entity including one or more of a protein, a deoxyribonucleic acid (DNA), and a ribonucleic acid (RNA);wherein the mobile device comprises:a processor being operable to:receive data from an external device associated with POH saliva testing, the external device including a digital device associated with one or more of a health care center, a hospital, an emergency center, a saliva research center, and a deoxyribonucleic acid genetic testing and analysis authority;based on the data, provide a notification to a user;receive a user input;perform a command, the command being selected based on the user input;provide a natural language user interface to communicate with the user, the natural language user interface being operable to sense a user voice and provide a response in a natural language to the user;a near field communication (NFC) unit communicatively coupled to the processor;a display communicatively coupled to the processor, the display including a touchscreen, wherein the display includes a force sensor, wherein the force sensor is operable to sense a touch force applied by the user to the display and calculate coordinates of a touch by the user, and further operable to analyze the touch force and, based on the touch force, select a tap command or a press command based on a predetermined criteria, wherein each of the tap command and the press command is executed based on a duration of the touch of the user using one or more applications running on the mobile device;a projector communicatively coupled to the processor, the projector being operable to project a data onto an arm of the user, the data including a virtual keyboard operable to provide inputs to the processor and one or more of the following: the notification, time data, data requested by the user, a caller name, a text message, a reminder, a social media alert, an email, an... | WORLD AWARD ACADEMY, WORLD AWARD FOUNDATION, AMOBILEPAY, INC. | G06Q020/32 A61B005/00 G06F021/32 H04M015/00 H04M001/2755 G06Q020/00 G06Q020/02 G06K015/00 H04M001/725 G06Q020/16 G06Q020/22
US09788786 | B2 | 2017-10-17 | Systems and methods for automatic detection and testing of images for clinical relevance | 2011-05-04 | Disclosed herein are systems and methods for automatic detection of clinical relevance of images of an anatomical situation. The method includes comparing a first image and a second image and determining whether a difference between the first and second images is at least one of a local type difference and a global type difference. The local type difference is a local difference of the first image and the second image and the global type difference is a global difference between the first image and the second image. The second image is determined as having a clinical relevance if it is determined that the difference between the first image and the second image comprises a local type difference. | 1. A method for automatic detection on clinical relevance of images of an anatomical situation provided by an imaging device, the method comprising: taking a first image of the anatomical situation at a first time;taking a second image of the anatomical situation at a second later time;comparing the first image and the second image and determining a difference between the first image and the second image by separating each of the first image and the second image into a plurality of corresponding image blocks each comprising a plurality of pixels, and comparing at least one of the corresponding image blocks of the first and second image;determining whether the difference between the first image and the second image comprises at least one of a local type difference and a global type difference, wherein the local type difference is a local difference of the first image and the second image and wherein the global type difference is a global difference between the first image and the second image, anddetermining whether the second image has clinical relevance by determining whether the difference between the first image and the second image comprises a local type difference,wherein the local type difference is detected if a first difference of a single image block of the first image and a corresponding single image block of the second image exceeds a first predetermined threshold and wherein the global type difference is detected if a second difference of all of the blocks of the plurality of blocks of the first image and all the blocks of the second image exceeds a second predetermined threshold, wherein the second threshold is larger than the first threshold. 2. The method according to claim 1, wherein the first and second differences are calculated using the same metric. 3. The method according to claim 2, wherein at least one of the threshold and the first threshold is a locally varying threshold. 4. The method according to claim 2, wherein at least one of the threshold, the first threshold and the second threshold is a dynamic threshold. 5. The method according to claim 1, further comprising determining the image block having the maximum local difference and comparing the image block having the maximum local difference with remaining image blocks and determining a local type difference based on the comparison of the image block having the maximum local difference and remaining image blocks. 6. The method according to claim 1, further comprising determining the image block having a significantly increased local difference and comparing the image block having the significantly increased local difference with image blocks in the vicinity of the image block having a significantly increased local difference and determining a local type difference based on the comparison of the image block having the significantly increased local difference with image blocks in the vicinity of the image block having a significantly increased local difference. 7. ... | Stryker European Holdings I, LLC | A61B005/00 G06T007/00 G06T007/254
US09808187 | B2 | 2017-11-07 | Measuring chamber for an optical sensor for determining a concentration of a substance in the tissue fluid of a mammal | 2012-11-30 | A measuring chamber device for an optically operating sensor for determining a concentration of a substance that is contained in tissue fluid of a mammal. The measuring chamber device has a measuring chamber filled with a liquid measuring medium and a wall with better diffusion permeability for the substance than for other constituents of the tissue fluid. A transmitter device for emitting optical radiation into the measuring chamber, and a receiver device for receiving optical radiation that has passed through the measuring chamber are also provided. | 1. A measurement chamber device for an optical sensor for measuring the concentration of a substance in the interstitial liquid of a mammal, wherein the measurement chamber device comprises: a measurement chamber, wherein the measurement chamber is filled with a measurement liquid and comprises a wall, the wall including at least one wall section enabling a better diffusion of the substance than of other components of the interstitial liquid, the measurement chamber, further including a first measurement chamber window section receiving a transmitter unit for emitting optical radiation into the measurement chamber and a second measurement chamber window section receiving a receiver unit for detecting optical radiation transmitted through the measurement chamber, the measurement chamber device further comprising a sterilization envelope which encloses the wall section of the measurement chamber but does not block the optical radiation at the measurement chamber window sections, the sterilization envelope defining a pair of openings in opposing walls thereof corresponding to the measurement chamber windows, wherein each of the pair of openings is sealingly coupled to the measurement chamber around the first measurement chamber window section and the second measurement chamber window section, respectively, wherein the sterilization envelope is filled with measurement liquid to bathe the wall section in measurement liquid, and wherein the sterilization envelope is selectively removable before use of the measurement chamber. 2. The measurement chamber device of claim 1, wherein the sterilization envelope comprises a sheet having an opening thread to remove the enclosing envelope. 3. The measurement chamber device of claim 1, wherein the measurement chamber window sections are glued to the wall and extend into the measurement chamber and cover sections of the wall, at which sections a diffusion permeability of the wall is reduced due to the gluing. 4. The measurement chamber device of claim 1, wherein a mounting portion is provided proximate at least one of the first measurement chamber window section or the second measurement chamber section to mount either the transmitter unit or the receiver unit, wherein the at least one mounting portion includes a flange. 5. The measurement chamber device of claim 4, wherein the at least one mounting portion provides a reservoir for an immersion liquid. 6. The measurement chamber device of claim 1, wherein a calibration chamber is provided which contains a calibration liquid and is connected to the sterilization envelope via a liquid connection, wherein the liquid connection is closed initially and can be opened for calibration purposes to mix the measurement liquid with the calibration liquid. 7. The measurement chamber device of claim 1, wherein a mounting and protection device is attached to the first measurement window section. 8. The measurement chamber device of claim 7, wherein the mounting and prote... | Schildtec GMBH | A61B005/1455 A61B005/145 A61B005/1459 A61B005/1495 G01N021/03 G01N021/27
US09811636 | B2 | 2017-11-07 | Connected health care system | 2014-09-20 | A connected health care system is provided with various devices for tracking and monitoring health statistics and behaviors including oral health, fitness, heart health, bone health, salivary diagnostics, diabetes, and further options that will be explored in the subsequent sections. The connected health care system is comprised of various devices, a data transfer medium (i.e. “smartphone”), and the Cloud, which allows for data transfer between multiple platforms and devices. Some devices comprised in the connected health care system include a toothbrush, a connected surface with sensors, a salivary diagnostic system, a breath analysis system, and further devices that will be explored in the subsequent sections. The connected health care system further provides a user identification system utilizing capacitive coupling of the human body between devices. | 1. A connected health care system for monitoring one or more health statistics of a user, the connected health care system comprising: a cloud computing network having at least one data processing unit that stores and processes data and a transceiver that receives and transmits data;a connected surface having a contact surface that allows the user to stand thereon during normal daily activities, a data processing unit that stores and processes data, a transceiver that receives and transmits data, a weight sensor that measures weight and balance, and a power source; anda connected implement having a handle, a data processing unit, a transceiver, a power source, and a salivary diagnostics system, the salivary diagnostics system further comprising: a collection system that collects at least one saliva sample from the user,a thermal-centrifugal filter system that treats the at least one saliva sample, where the collection system is detachably coupled to the thermal-centrifugal filter system, andan evaluation system that analyzes the at least one saliva sample to detect and/or diagnose a disease or a disease risk characteristic, the thermal-centrifugal filter system and the evaluation system being enclosed within a single unit. 2. The connected health care system of claim 1, further comprising a data transfer medium having a transceiver that receives and transmits data, a data processing unit that stores and processes data, and a display medium that displays data. 3. The connected health care system of claim 1, wherein the connected health implement is a toothbrush that comprises a plurality of bristles, a pressure sensor that detects brushing pressure, and an orientation sensor that detects a position of the toothbrush. 4. The connected health care system of claim 1, wherein the connected surface further comprises at least one additional sensor selected from the group consisting of an oximetry sensor, an ultrasonic sensor that measures bone density, a thermal imaging sensor, a galvanic skin response sensor, a body fat sensor, an electrocardiogram sensor that measures the rate and regularity of heartbeats, a pressure point sensor that identifies pressure points, and a temperature sensor, and any combination thereof. 5. The connected health care system of claim 1, wherein the connected health implement is a breath analysis system having a mouthpiece and a semiconducting oxide gas sensor that detects amounts of acetone in breath. 6. The connected health care system of claim 1, further comprising a user identification system having at least one mobile unit, at least one base unit, and a capacitive coupler that connects the mobile unit to the the base unit such that the mobile unit communicates with the base unit. 7. The connected health care system of claim 6, wherein the mobile unit is selected from the group consisting of a toothbrush, the collection system, the thermal-centrifugal filter system, the evaluation system, the reporting system, a dat... | BEAM IP LAB LLC | G01N009/30 G01N001/10 G06F019/00 G06Q050/22 A61B005/00 A61B005/1455 A61C019/04 A61B005/0205 A61B005/08 A46B015/00 A61C017/06 A61C017/26 A61C017/22 A46B005/00
US09814430 | B1 | 2017-11-14 | System and method for measuring eye movement and/or eye position and postural sway of a subject | 2015-04-17 | A system and method for measuring eye movement and/or eye position and postural sway of a subject is disclosed herein. The system generally includes an eye movement tracking device, a postural sway detection device, and a data processing device operatively coupled to the eye movement tracking device and the postural sway detection device. During the execution of the method, the eye movement and/or eye position of the subject is measured using the eye movement tracking device, and the postural sway of the subject is measured using the postural sway detection device. A method for determining a gaze direction of a subject during a balance test and/or concussion screening test, and a method for assessment of a medical condition of a subject are also disclosed herein. | 1. A method for concurrently measuring the eye movement or eye position and postural sway of a subject, the method comprising the steps of: providing an eye movement tracking device configured to track eye movement or eye position of the subject while the subject performs a balance test or a concussion screening test, the eye movement tracking device being configured to output one or more first signals that are representative of the detected eye movement or eye position of the subject;providing a postural sway detection device, the postural sway detection device configured to detect a postural sway of the subject while the subject performs the balance test or the concussion screening test, the postural sway detection device being configured to output one or more second signals that are representative of the detected postural sway of the subject;providing at least one limb position detection device, the at least one limb position detection device configured to detect a position of one or more limbs of the subject and output one or more third signals that are representative of the detected position of the one or more limbs of the subject;providing a data processing device operatively coupled to the eye movement tracking device, the postural sway detection device, and the at least one limb position detection device, the data processing device configured to receive the one or more first signals that are representative of the detected eye movement or eye position of the subject, the one or more second signals that are representative of the detected postural sway of the subject, and the one or more third signals that are representative of the detected position of the one or more limbs of the subject, the data processing device further configured to determine the eye movement or eye position of the subject using the one or more first signals, the postural sway of the subject using the one or more second signals, and the position of the one or more limbs of the subject using the one or more third signals;positioning the subject in an upright position on a surface;positioning the at least one limb position detection device on the one or more limbs of the subject between a hand and a shoulder of the subject;providing a visual target that is a portion of the one or more limbs of the subject;instructing the subject to displace the one or more limbs while the subject maintains his or her gaze on the visual target;measuring the eye movement or eye position of the subject using the eye movement tracking device, and outputting the one or more first signals that are representative of the detected eye movement or eye position of the subject from the eye movement tracking device;measuring the postural sway of the subject using the postural sway detection device while measuring the eye movement or eye position of the subject, and outputting the one or more second signals that are representative of the postural sway of the subject from the postural sway detection dev... | Bertec Corporation | A61B005/00 G06F019/00 A61B003/113 A61B005/11 A61B008/10 A61B005/0496 A61B008/00
US09820693 | B2 | 2017-11-21 | Measurement of biosignals of a user | 2017-02-03 | An information processing apparatus includes a retaining unit, an image acquisition unit, an image processor, a comparator, and an image generator. The retaining unit is configured to retain a defined attachment position being an attachment position of a biological detection sensor with respect to a body site of a user, the defined attachment position being defined in advance. The image acquisition unit is configured to acquire a capture image captured by an imaging unit. The image processor is configured to extract, from the capture image, a body shape of the user. The comparator is configured to compare the body shape with the defined attachment position. The image generator is configured to generate, based on a comparison result by the comparator, a display image to be displayed on a display unit. | 1. An information processing apparatus, comprising: circuitry configured to: acquire an image of a user with a biosignal detection sensor;acquire information indicating an attachment state of the biosignal detection sensor; andprovide at least one of a first degree of accuracy of an attachment position of the biosignal detection sensor or a second degree of accuracy of the attachment state of the biosignal detection sensor based on the image and the information,wherein the information indicating the attachment state includes at least one of an attachment direction of the biosignal detection sensor or a degree of contact of the biosignal detection sensor. 2. The information processing apparatus according to claim 1, wherein the circuitry is further configured to display the image with a marker located at a position on the image, wherein the position corresponds to a defined attachment position. 3. The information processing apparatus according to claim 1, wherein the circuitry is further configured to display the image with a marker located at a position on the image to generate a display image based on a defined attachment position that is different from a detected attachment position detected from the acquired image, wherein the position corresponds to the defined attachment position different from the detected attachment position. 4. The information processing apparatus according to claim 1, wherein the biosignal detection sensor includes an electroencephalographic electrode, and wherein the circuitry is further configured to determine the attachment state based on contact resistance of the electroencephalographic electrode. 5. The information processing apparatus according to claim 1, wherein the circuitry is further configured to: detect the attachment position of the biosignal detection sensor from the image;compare the detected attachment position with a defined attachment position of the biosignal detection sensor; andprovide the first degree of accuracy based on the comparison. 6. An information processing method, comprising: in an information processing apparatus: acquiring an image of a user with a biosignal detection sensor;acquiring information indicating an attachment state of the biosignal detection sensor; andproviding at least one of a first degree of accuracy of an attachment position of the biosignal detection sensor or a second degree of accuracy of the attachment state based on the image and the information,wherein the information indicating the attachment state includes at least one of an attachment direction of the biosignal detection sensor or a degree of contact of the biosignal detection sensor. 7. A non-transitory computer-readable medium having stored thereon computer-readable instructions, which when executed by a computer, cause the computer to execute operations, the operations comprising: acquiring an image of a user with a biosignal detection sensor;acquiring information indicating an attachment state of the ... | SONY CORPORATION | H04N009/47 H04N007/18 A61B005/00 A61B005/0478 G06T007/73
US09801563 | B2 | 2017-10-31 | Micro-magnetic reporter and systems | 2014-11-05 | The present disclosure describes system and methods for detecting and amplifying weak magnetic fields generated by anatomical structures. The disclosure describes an implantable magnetic reporter system. The magnetic reporter system includes a magnetic reporter. The magnetic reporter includes a platform coupled to a support structure by a plurality of torsional flexures. A magnet is disposed on the platform, and the magnet and platform rotate when exposed to a magnetic field. The rotation of the magnet generates a stronger magnet field that is detectable external to the patient. | 1. A system comprising: a magnetic reporter comprising: a platform coupled to a support structure by a plurality of torsional flexures, wherein the platform is configured to rotate about an axis; anda magnet disposed on the platform;a wound coil disposed proximal to the magnetic reporter;an energy source electrically coupled with the wound coil;a sensor configured to measure a physiological signal;a controller electrically coupled to the sensor and configured to control a flow of current between the energy source and the wound coil based on the measured physiological signal; anda hermetical package forming a seal around the magnetic reporter. 2. The system of claim 1, wherein the controller is configured to enable the current flow from the energy source to the wound coil responsive to detecting a signal above a predetermined threshold. 3. The system of claim 1, wherein the sensor is an implantable electroencephalography electrode. 4. The system of claim 1, wherein the sensor is an ion concentration sensor. 5. The system of claim 1, wherein the energy source is one of a capacitor and a rechargeable battery. 6. The system of claim 1, wherein the wound coil is an energy scavenger configured to recharge the energy source. 7. The system of claim 1, wherein the plurality of torsional flexures suspend the platform over a recess in a capping layer. 8. The system of claim 1, wherein the plurality of torsional flexures comprises at least one pair of torsional flexures coupled to opposite ends of the platform. 9. The system of claim 1, wherein the magnet comprises a polymer binder and a magnetic powder. 10. The system of claim 1, further comprising a second magnet disposed on a second face of the platform. 11. A method comprising: implanting a magnetic reporter system in a patient, the magnetic reporter system comprising: a magnetic reporter comprising a platform coupled to a support structure by a plurality of torsional flexures, wherein the platform is configured to rotate about an axis; and a magnet disposed on the platform;a wound coil disposed proximal to the magnetic reporter;an energy source electrically coupled with the wound coil;a controller to control a flow of current between the energy source and the wound coil;a sensor electrically coupled to the controller and configured to measure a physiological signal; anda hermetical package forming a seal around the magnetic reporter;monitoring a physiological parameter with the magnetic reporter system;measuring a magnetic field generated by the magnetic reporter system in response to a change in the monitored physiological parameter. 12. The method of claim 11, further comprising measuring the magnetic field generated by the magnetic reporter with a magnetic field sensor external to the patient. 13. The method of claim 11, further comprising implanting the magnetic reporter system proximal to a neuron. 14. The method of claim 11, further comprising energizing, responsive to a change in the... | THE CHARLES STARK DRAPER LABORATORY, INC. | A61B005/05 A61B005/0478 A61B005/04 A61B005/00
US09788705 | B1 | 2017-10-17 | Minimally invasive networked surgical system and method | 2016-06-23 | A system for performing non-invasive networked medical procedures including a number of in vivo medical devices, a communication path between at least two of the devices, an ex vivo control unit to control the behavior of the devices, and a wireless communication path between the control unit and at least one of the devices. An associated method for performing non-invasive networked medical procedures is also provided. Further included is a simulation method that utilizes accurate electromagnetic field simulations, using a software based test bench, to determine the maximum allowable transmitted power levels from in vivo devices to achieve a required bit error rates (BER) at an in vivo or ex vivo node (receiver) while maintaining the specific absorption rate (SAR) under a required threshold. | 1. A system for performing networked medical procedures on a subject, comprising: a plurality of in vivo medical devices including an in vivo transmitter;a network providing a communication path between at least two of said plurality of in vivo medical devices, said network permitting receipt and transmission of control signals between said at least two in vivo medical devices over said communication path; andan ex vivo control unit to control a behavior of said plurality of in vivo medical devices, said behavior including spatial parameters of said plurality of in vivo medical devices;said network providing a wireless communication path between said ex vivo control unit and at least one of said plurality of in vivo medical devices. 2. The system of claim 1, further comprising: a second wireless communication path between at least one of said plurality of in vivo medical devices and an ex vivo system. 3. The system of claim 1, wherein said ex vivo control unit is a magnet. 4. The system of claim 1, wherein said ex vivo control unit is an electronic device that generates control signals. 5. The system of claim 1, wherein said ex vivo control unit is an electronic device that generates communication signals. 6. The system of claim 1, wherein each of said plurality of in vivo medical devices are one or more medical devices chosen from the group consisting of an imaging device, a sensor, a power source, and a cutting tool. 7. The system of claim 1, wherein said communication path between said at least two in vivo medical devices is wireless. 8. The system of claim 1, wherein said communication path between said at least two in vivo devices is wired. 9. The system of claim 1, wherein each of said plurality of in vivo medical devices is separately wirelessly addressable and controllable. 10. The system of claim 1, further comprising: a plurality of wireless communication paths between said ex vivo control unit and each of said plurality of in vivo medical devices. 11. The system of claim 1, further comprising: said communication path including a signal having a specific absorption rate less than or equal to 0.08 W/kg. 12. The system of claim 1, further comprising: said positioning of said in vivo transmitter being furthest from organs and tissues within said subject. 13. The system of claim 1, further comprising: said network being a Layer 3 data network. 14. The system of claim 1, further comprising: said behavior of said plurality of in vivo medical devices including spatial parameters of said plurality of in vivo medical devices. 15. The system of claim 1, further comprising: said ex vivo control unit being an expert system. 16. A method of performing networked medical procedures comprising: providing a plurality of in vivo medical devices;establishing a network that provides a communication path between at least two of said plurality of in vivo medical devices, said network permitting receipt and transmission of control signals betwe... | University of South Florida | H04B007/00 A61B001/00 A61B005/00 A61B017/32 H04W076/02
US09795723 | B2 | 2017-10-24 | Manifold for medical waste collection device | 2016-04-29 | Medical waste collection manifolds that that may connect to a vacuum system to create suction, such that liquid, semi-solid, or solid biological or medical waste may be collected are provided. The manifold may be configured to collect hazardous waste in a manner to ensure proper containment. | 1. An interchangeable waste collection manifold for connection to a vacuum device having a receiver for receiving the manifold, comprising: a manifold body removably engageable with the receiver of the vacuum device, the manifold body having proximal and distal ends and defining an internal volume providing an internal fluid path therebetween, and wherein the manifold body has an angular bend disposed along the length thereof, such that the proximal end lies along a first longitudinal axis and the distal end lies along a second longitudinal axis, and wherein the first and second longitudinal axes are angularly offset;a suction port assembly providing at least one suction port having a proximal end external to the manifold body and a distal end extending internal to the manifold body, the at least one suction port providing an inlet fluid path into the internal volume of the manifold body, wherein each of the at least one suction ports are disposed within the proximal end of the manifold body such that all inlet fluid paths into the manifold body are disposed along the first longitudinal axis;a vacuum port disposed within the distal end of the manifold body, and providing at least one fluid outlet from the internal volume of the manifold body, the vacuum port being configured to engage a vacuum inlet integral with the receiver of the vacuum device;a filter element secured within the internal volume of the manifold body such that the internal fluid path between the proximal and distal ends passes therethrough;at least one orientation feature disposed on the outer surface of the manifold body and configured such that when the manifold body is engaged with the vacuum device, the orientation feature interlocks the manifold body within the receiver of the vacuum device and fixes the orientation of the manifold body relative to the vacuum device such that each of the at least one suction ports is angled upward relative to the direction of gravity; andwherein the angular bend offsets the second end of each of the at least one suction ports from the internal fluid path disposed along the second longitudinal axis of the manifold body to prevent reflux of a liquid waste back through the suction port when interlocked within the vacuum device. 2. The manifold of claim 1, wherein the angular offset is greater than 25° and less than 90°. 3. The manifold of claim 2, wherein the angular bend is greater than 30° and less than 60°. 4. The manifold of claim 3, wherein the angular bend is greater than 35° and less than 55°. 5. The manifold of claim 1, wherein the filter element is a basket disposed along the second longitudinal axis and containing a plurality of apertures having at least a first aperture dimension, the plurality of apertures forming a portion of the fluid path through the manifold body. 6. The manifold of claim 5, wherein the basket has at least one overflow-relief aperture having a second aperture dimension larger than the at least first apertu... | 5x Surgical, LLC | A61M001/00 A61B017/16 A61B005/00 A61B010/00
US09776064 | B2 | 2017-10-03 | Protective temperature helmet and shoulder pads, protective temperature helmet and shoulder pad liner | 2015-02-02 | The invention provides head, neck, chest, and spine logical impact protection and thermal-regulation device having a webbing adapted to encapsulate head and neck, wherein the webbing is composed of a plurality of reversibly deformable substrates that are capable of being cooled or heated; wherein the webbing further provides apertures interspersed between the reversibly deformable substrates to allow fluid exchange between the head, neck, chest, and spine and the regions exterior to the webbing; and a temperature indicator, wherein the temperature indicator measures the temperatures of the body surface underlaying the webbing. | 1. A head and neck impact protection and thermal-regulation device comprising: a. a webbing adapted to encapsulate head and neck, wherein the webbing comprises a plurality of reversibly deformable substrates that are capable of being cooled or heated;b. wherein the webbing further comprises apertures interspersed between the reversibly deformable substrates to allow fluid exchange between the head and neck and the regions exterior to the webbing; andc. a temperature indicator, wherein the temperature indicator measures the temperatures of the body surface underlaying the webbing. 2. The device as recited in claim 1 further comprising downwardly depending webbing that extends from the base of the neck to overlay the back, chest, and shoulders of a wearer. 3. The device as recited in claim 1 wherein the temperature indicator further comprises a device selected from the group consisting of a liquid crystal thermometer, temperature sensitive paint, a digital thermometer, a digital thermometer programmed to display specified colors in specified temperature ranges, a digital thermometer programmed to emit specified sounds in specified ranges, a digital thermometer that transmits temperature readings to a receiving device, and combinations thereof. 4. The device as recited in claim 1 wherein the webbing further defines flexible sheeting defining cavities adapted to reversibly receive the deformable substrates. 5. The device as recited in claim 4 wherein the substrates are cooled or heated while residing in the cavities. 6. The device as recited in claim 4 wherein the substrate are cooled or heated before being received by the cavities. 7. The device as recited in claim 1 wherein at least one of the deformable substrates is the temperature indicator. 8. The device as recited in claim 1 further comprising a semi-rigid or rigid material overlaying the webbing. 9. The device as recited in claim 8 wherein the material is reversibly attached to the webbing. 10. The device as recited in claim 8 wherein the material is integrally molded with the webbing. 11. The device as recited in claim 1 wherein the device is adapted to receive a helmet. 12. An upper-body impact protection and thermal-regulation device comprising: a. a webbing adapted to encapsulate the human head and neck, wherein downwardly depending webbing extends from the base of the neck to overlay a wearer's back, chest, and shoulders, the webbing comprising a plurality of reversibly deformable substrates, wherein the substrates are capable of being cooled or heated; andb. wherein the webbing further defines apertures interspersed between the cavities to allow fluid exchange between the head and neck and the regions exterior to the webbing. 13. The device of claim 12 further comprising a temperature indicator, wherein the temperature indicator measures the temperatures of the body surface underlaying the webbing. 14. The device as recited in claim 13 wherein the temperature indicator furt... | - | A41D013/00 A63B071/12 A61B005/01 A63B071/10 A42B003/12 A42B003/28 A61F007/00 A41D013/05
US09819074 | B2 | 2017-11-14 | Monolithically integrated implantable flexible antenna for electrocorticography and related biotelemetry devices | 2014-03-11 | A sub-skin-depth (nanoscale metallization) thin film antenna is shown that is monolithically integrated with an array of neural recording electrodes on a flexible polymer substrate. The structure is intended for long-term biometric data and power transfer such as electrocorticographic neural recording in a wireless brain-machine interface system. The system includes a microfabricated thin-film electrode array and a loop antenna patterned in the same microfabrication process, on the same or on separate conductor layers designed to be bonded to an ultra-low power ASIC. | 1. An apparatus, comprising: a flexible polymer substrate;a thin film antenna; andan array of one or more recording electrodes;wherein the thin film antenna is monolithically integrated with the array of one or more recording electrodes on the flexible polymer substrate to form a conformal monolithic device;wherein said monolithic device comprises an implant configured for performing wireless signal acquisition, power delivery, and communication with an external antenna;wherein the array of one or more recording electrodes comprises a high-density microfabricated ECoG grid array configured to measure one or more physiological characteristics associated with a patient at an implant site within the patient; andwherein the thin film antenna comprises a loop antenna disposed around the ECoG grid array. 2. An apparatus as recited in claim 1, wherein the loop antenna is disposed within the same layer as the ECoG grid array. 3. An apparatus as recited in claim 2, wherein the ECoG grid array and loop antenna layer is bonded to an integrated circuit; the integrated circuit configured for executing one or more of said wireless signal acquisition, power delivery, and communication with the external antenna. 4. An apparatus as recited in claim 3, wherein said integrated circuit is bonded as a flip-chip via an anisotropic conductive film thermo-compression. 5. An apparatus as recited in claim 1, wherein the loop antenna and ECoG grid array comprise a nanoscale metallization pattern configured to allow the monolithic device to conform to a non-planar shape of tissue at the implant site. 6. An apparatus as recited in claim 5, wherein the ECoG grid array and loop antenna are disposed within a same wafer level on the flexible polymer substrate. 7. An apparatus as recited in claim 6, wherein ECoG grid array and loop antenna layer comprise a conductive layer having a total thickness of 200 nm. 8. An apparatus as recited in claim 7, wherein said conductive layer comprises a Pt adhesion layer, a ductile Au core, and a biocompatible Pt passivation layer. 9. An implantable device, comprising: a flexible polymer substrate; andan antenna micro-patterned onto the flexible polymer substrate in the form of a polymer thin-film layer;the polymer thin-film layer further comprising an array of one or more recording electrodes configured to measure one or more physiological characteristics associated with a patient at an implant site within the patient;wherein the array of one or more recording electrodes comprises a high-density microfabricated ECoG grid array; andwherein the antenna comprises a loop antenna disposed around the ECoG grid array. 10. An implantable device as recited in claim 9, wherein the antenna is monolithically integrated with the array of one or more recording electrodes on the flexible polymer substrate to form a conformal application-specific integrated circuit (ASIC). 11. An implantable device as recited in claim 10: said conformal ASIC comprisi... | THE REGENTS OF THE UNIVERSITY OF CALIFORNIA | H01Q007/00 A61B005/00 H01Q001/27 A61B005/04 H01Q001/22
US09811635 | B2 | 2017-11-07 | Immune and oxygen system measuring and drug screening method and apparatus | 2015-08-14 | Method and apparatus for monitoring health as related to immune system function, and for measuring the effects of toxins and other stresses. A method for pre-screening drugs for the pharmaceutical pipeline. A method for using an Immunogram as a research tool. A method for preparing compounds or drugs for treatment, therapy, or cure of diseases. | 1. A method for enhancing an individual's immune system, said method comprising the steps of: (a) providing an apparatus, including (i) a first sensor that is an optical probe,(ii) a second sensor that is a membrane probe, and(iii) a reader processor in electrical connection with and responsive to said first and second sensors, wherein said first and second sensors detect molecular oxygen and nitric oxide and send electrical signals in response to the detection of said molecular oxygen and said nitric oxide to said reader processor, said reader processor reading said electrical signals and recording data related to said electrical signals as a function of time, including as part of said data concentrations of molecular oxygen and nitric oxide, changes in said concentrations, and rates of change in said concentrations;(b) obtaining a sample of blood from the body of an individual, said sample including molecular oxygen and nitric oxide;(c) measuring with said apparatus an initial concentration of said molecular oxygen and said nitric oxide to determine a baseline value for said individual;(d) continuing measurement of the concentration of said molecular oxygen and said nitric oxide over time by said apparatus;(e) stressing said sample;(f) acquiring a first set of data from said sample by said apparatus over the course of a first interval of time following said stressing step to determine a first value for said molecular oxygen and a first value for said nitric oxide, said first value being a decrease in concentration of said molecular oxygen or a decrease in concentration of said nitric oxide with respect to said baseline;(g) repeating said stressing step;(h) acquiring a second set of data from said sample by said apparatus over the course of a second interval of time following said repeating step to determine a second value for said molecular oxygen and said nitric oxide, said second value being any increase in concentration of said molecular oxygen or any increase in concentration of said nitric oxide with respect to said first value;(i) comparing by said apparatus said first and second values to preselected standard values to show the immune system performance capabilities of said individual, wherein steps (c)-(i) are repeated until a sharp increase of said molecular oxygen or said nitric oxide is detected; and(j) administering said blood to said individual following step (i) to enhance said immune system performance capabilities. 2. The method as recited in claim 1, wherein said first sensor only detects molecular oxygen and said second sensor detects molecular oxygen and nitric oxide and sends electrical signals in response to the detection of said molecular oxygen and nitric oxide to said reader processor. 3. The method as recited in claim 1, wherein said stressing step (e) further comprises the step of stirring said sample at a speed and for a period of time sufficient to cause a drop in concentration of said molecular oxygen. 4. The met... | - | G06F019/00 A61B005/145 A61K035/14 A61K033/00 A61K033/22 A61K031/60 G01N033/84 A61B005/1455 G01N033/48
US09794668 | B2 | 2017-10-17 | Smart flexible interactive earplug | 2017-01-23 | An embodiment of the invention provides a resiliently deformable and flexible in-ear sound device having stretchable electronic circuitry. The in-ear sound device may be configured in a variety of ways, including, but in no way limited to a smart earplug, a flexible personal sound amplification product, a personal music player, a “walkie-talkie” and the like. | 1. An in-ear monitoring device, comprising: a deformable and flexible body having a longitudinal axis extending between a distal end and a proximal end;an electronic component package comprising stretchable electronic circuitry, wherein the electronic component package includes at least one sensor configured to measure a data element in proximity of the in-ear monitoring device and wherein the stretchable electronic circuitry is at least one of embedded in or on the deformable and flexible body;a processor in the electronic component package configured to analyze the data element and take an action if the data element exceeds a threshold wherein the at least one sensor and the processor are configured to operate as a measurement device that provides one of an alarm, a stop watch, a calendar, and a notification function and wherein the action if the data element exceeds the threshold related to the measurement device comprises preparing a notification; anda communication module in communication with a remote device that sends the notification to the remote device if the data element exceeds the threshold. 2. The in-ear monitoring device of claim 1 wherein the measurement device operates as an alarm and wherein the at least one sensor is configured as at least one of a thermometer, a pulse rate monitor, a heart rate monitor, a V02 Max monitor, a pulse oximetry monitor, a respiratory rate monitor, a respiratory monitor, an oxygen consumption monitor, a cardiac efficiency monitor, a heart rate variability monitor, a metabolic rate monitor, a blood pressure monitor, an EEG data monitor, a galvanic skin response monitor, an EKG/ECG monitor, a blood analyte monitor, an ambient temperature monitor, a humidity monitor, a motion detector, a GPS locator, a pressure sensor, an altitude sensor, an accelerometer, a gyroscope, and a magnetometer. 3. The in-ear monitoring device of claim 1, further comprising: a speaker located at the proximal end of the deformable and flexible body wherein the notification is also played through the speaker to a wearer of the in-ear monitoring device. 4. The in-ear monitoring device of claim 1, further comprising: a microphone in the electronic component package configured to receive sounds external to the in-ear monitoring device,wherein the processor is further configured to analyze the sounds received by the microphone, separate meaningful sounds from ambient noise, and provide notice information about the meaningful sounds to the remote device via the communication module. 5. The in-ear monitoring device of claim 4, further comprising: a data repository containing identification data pertaining to the meaningful sounds, wherein the processor's analysis of the received sound data from the microphone comprises determining if the received sound data matches the identification data and generating the notice information if the identification data matches the received sound data. 6. The in-ear monitoring device of claim 5 wh... | SmartEar, Inc. | H04R001/10 H04R001/04 H04W004/00 H04R001/02 H04R025/00 H04M001/60 G10L017/22 G10L025/51 A61B005/00 A61B005/0205 A61B005/1455 A61B005/024 A61B005/08 A61B005/021 A61B005/0476 A61B005/0402 A61B005/11 A61B005/053
US09775519 | B2 | 2017-10-03 | Network connectivity unit for hospital bed | 2016-10-24 | A hospital bed is programmable with new firmware that is downloaded to the bed over a network. The firmware is downloaded to the bed automatically from a remote computer device. The remote computer device receives a message from the hospital bed which includes data regarding the version number of the bed's current firmware, and if the version number indicates that the firmware is an outdated version, the remote computer device downloads a new version of the firmware to the bed. | 1. An interface unit for use in a healthcare facility having a computer located remotely from a patient room which has a patient bed therein, the interface unit being located in the patient room and spaced from the patient bed, the interface unit comprising: a first connector which is communicatively coupleable to the hospital bed,a second connector which is communicatively coupleable to the computer, andcircuitry comprising at least two device connectors for connection to external devices, the circuitry storing a unit identification (ID) that is associated with the patient room of the healthcare facility, the unit ID being transmitted by the circuitry through the second connector without transmitting the unit ID through the first connector and the at least two device connectors. 2. The interface unit of claim 1, wherein at least one device connector of the at least two device connectors comprises an Ethernet port for connection of the circuitry to an Ethernet of the healthcare facility. 3. The interface unit of claim 2, wherein the Ethernet port comprises an RJ-45 port. 4. The interface unit of claim 2, wherein the data from the hospital bed is received by the circuitry via the first connector according to a first communication protocol and the circuitry converts the data from the first protocol into an Ethernet protocol. 5. The interface unit of claim 4, wherein the data from the hospital bed is sent to the computer via the second connector according to the first protocol. 6. The interface unit of claim 1, wherein the data from the hospital bed is passed through to the computer via the first and second connectors. 7. The interface unit of claim 6, wherein the data from the hospital bed that is passed through to the computer includes a nurse call signal. 8. The interface unit of claim 6, wherein the data from the hospital bed that is passed through to the computer includes a signal for controlling room lighting. 9. The interface unit of claim 6, wherein the data from the hospital bed that is passed through to the computer includes a signal for controlling entertainment. 10. The interface unit of claim 1, wherein at least one device connector of the at least two device connectors comprises a JTAG connector. 11. The interface unit of claim 1, wherein the first connector is communicatively coupled to the hospital bed via a wired data link. 12. The interface unit of claim 1, wherein the first connector is communicatively coupled to the hospital bed via a wireless data link. 13. The interface unit of claim 1, wherein the second connector is communicatively coupled to the computer via a wired data link. 14. The interface unit of claim 1, wherein the second connector is communicatively coupled to the computer via a wireless data link. 15. The interface unit of claim 1, wherein at least one device connector of the at least two device connectors comprises at least one general purpose input/output (GPIO) connector. 16. The interface unit of... | Hill-Rom Services, Inc. | G08B005/22 A61B005/00 A61B005/11 A61G007/018 G06F019/00 G07C003/00 G08B021/22 G06F009/445 H04L029/08 A61B005/024 A61G012/00
US09782108 | B2 | 2017-10-10 | Bed-leaving sensor and bed-leaving detection method | 2015-09-23 | A bed-leaving sensor for detecting bed-leaving by a user on a bed, including: a sitting position detection member that detects a sitting position of the user based on a detection value of a pressure sensor arranged on the bed; a foot movement detection member for detecting foot movement of the user based on the detection value of the pressure sensor; and a bed-leaving behavior detection member that detects bed-leaving behavior of the user with the sitting position or the foot movement being detected within a preset bed-leaving expectation region as a condition. | 1. A bed-leaving sensor for detecting bed-leaving by a user on a bed, comprising: a sitting position detection member that detects a sitting position of the user based on a detection value of a pressure sensor arranged on the bed;a foot movement detection member for detecting foot movement of the user based on the detection value of the pressure sensor; anda bed-leaving behavior detection member that detects bed-leaving behavior of the user with the sitting position or the foot movement being detected within a preset bed-leaving expectation region as a condition,wherein the sitting position detection member sets a determination area in a periphery of a pressure sensing center of a region in which pressure of a given contact threshold value or greater is detected within a preset sitting position determination region on the pressure sensor provided on the bed-leaving, expectation region, and the sitting position is detected based on fulfilling all the following requirements: a ratio of a total pressure detected within the determination area occupying a total pressure detected within the sitting position determination region is a designated threshold value or greater;a breadth of the region in which the pressure of the contact threshold value or greater is detected within the sitting position determination region is a designated value or greater; andthe total pressure detected within the sitting position determination region is a designated threshold value or greater, andwherein the foot movement detection member detects the foot movement based on fulfilling all the following requirements: a breadth of a region in which pressure of a given contact threshold value or greater has been detected within a preset foot pressure detection region on the pressure sensor provided on the bed-leaving expectation region is within a designated range;a total pressure detected within the foot pressure detection region across a given time is a designated threshold value or greater; anda maximum pressure value detected within the foot pressure detection region is a preset threshold value or greater. 2. The bed-leaving sensor according to claim 1, wherein the bed-leaving expectation region is set to be a foot side region on the bed. 3. The bed-leaving sensor according to claim 1, wherein the pressure sensor is arranged only at a bed-leaving possibility region on the bed. 4. The bed-leaving sensor according to claim 1, wherein the pressure sensing center is a center of gravity of the region in which the pressure of the contact threshold value or greater is detected with the pressure sensor. 5. The bed-leaving sensor according to claim 1, wherein the pressure sensing center is an area center of the region in which the pressure of the contact threshold value or greater is detected with the pressure sensor. 6. The bed-leaving sensor according to claim 1, further comprising a notification member for notifying that the bed-leaving behavior of the user has been detected.... | SUMITOMO RIKO COMPANY LIMITED | G08B021/22 A61B005/11 G08B021/04 G08B029/18 A61B005/103
US09801555 | B2 | 2017-10-31 | Thoracic diagnosis assistance system | 2014-08-28 | According to one implementation, the system includes, a capturing unit, a deforming unit, and a generating unit. The capturing unit captures a dynamic state of a thoracic portion to generate a plurality of frame images. The deforming unit sets a reference point in a position corresponding to each other among the plurality of generated frame images. The deforming unit extracts a lung field region from each of the frame images. The deforming unit deforms a shape of the lung field region so that a distance from the set reference point to an outline of an outer side of the lung field region becomes a certain distance. The generating unit analyzes a dynamic state in the lung field region and generates an analysis result image showing a result of the analysis in a corresponding position in the deformed lung field region. | 1. A thoracic diagnosis assistance system comprising: a capturing unit which captures a dynamic state of a thoracic portion to generate a plurality of frame images;a deforming unit which sets a reference point in a position corresponding to each other among the plurality of generated frame images, extracts a lung field region from each of the plurality of frame images, and deforms a shape of the lung field region in each of the plurality of frame images so that a distance from the set reference point to an outline of an outer side of the lung field region becomes equal; anda generating unit which analyzes a dynamic state in the lung field region based on the plurality of frame images in which a shape of the lung field region is deformed by the deforming unit, and generates an analysis result image showing a result of the analysis in a corresponding position in the deformed lung field region. 2. The thoracic diagnosis assistance system of claim 1, wherein, the generating unit analyzes blood flow in the lung field region, and generates an analysis result image showing an analysis result of the blood flow in a corresponding position in the lung field region. 3. The thoracic diagnosis assistance system of claim 1, wherein, the reference point is a pulmonary hilum in the lung field region. 4. The thoracic diagnosis assistance system of claim 1, wherein, the deforming unit deforms the shape of the lung field region so that the outline of the outer side of the lung field region is an arc shape. 5. The thoracic diagnosis assistance system of claim 1, further comprising, a display control unit which assigns a color or a brightness according to a value of the analysis result in the lung field region of the analysis result image with the shape of the lung field region deformed and displays the analysis result image on a display unit. 6. The thoracic diagnosis assistance system of claim 1, further comprising, a region analysis unit which sets an analysis target region according to a distance from the reference point, the analysis target region set in the lung field region in the analysis result image with the shape of the lung field region deformed, and analyzes the dynamic state in the set analysis target region. 7. The thoracic diagnosis assistance system of claim 1, further comprising, a difference image generating unit which calculates a difference value between values of an analysis result of corresponding positions in two analysis result images with the shape of the lung field region deformed, and generates a difference image; andan extracting unit which extracts at least one of a location in which the difference value exceeds a predetermined threshold value or a location in which the difference value is equal to or less than the predetermined threshold value in the difference image. 8. The thoracic diagnosis assistance system of claim 5, wherein, the display control unit displays on the display unit the analysis result image with the shape of t... | KONICA MINOLTA, INC. | G06K009/00 A61B005/026 A61B005/113 G06T005/00 G06T007/11 A61B005/00 A61B006/00
US09775551 | B2 | 2017-10-03 | Devices and techniques associated with diagnostics, therapies, and other applications, including skin-associated applications | 2015-07-13 | The present invention generally relates to devices and techniques associated with diagnostics, therapies, and other applications, including skin-associated applications, for example, devices for delivering and/or withdrawing fluid from subjects, e.g., through the skin. In some embodiments, the device includes a system for accessing an extractable medium from and/or through the skin of the subject at an access site, and a pressure regulator supported by a support structure, able to create a pressure differential across the skin at at least a portion of the access site. The device may also include, in some cases, a sensor supported by the support structure for determining at least one condition of the extractable medium from the subject, and optionally a signal generator supported by the support structure for generating a signal relating to the condition of the medium determined by the sensor. | 1. A device for withdrawing blood from the skin and/or from beneath the skin of a subject, the device comprising: a recess;a vacuum chamber separate from the recess having an internal pressure less than atmospheric pressure before blood is withdrawn into the device and before the device is applied to the skin of the subject;a storage chamber, separate from the vacuum chamber and the recess, for receiving blood withdrawn into the recess when a negative pressure is applied to the skin of the subject; anda sensor for determining a condition of the blood withdrawn into the device. 2. The device of claim 1, further comprising a microfluidic channel, discrete from the vacuum chamber and the storage chamber and having a smaller cross-sectional area than the cross-sectional area of the storage chamber, wherein the vacuum chamber is in fluid communication with the storage chamber via the microfluidic channel. 3. The device of claim 1, further comprising a plurality of microneedles for accessing the blood from and/or through the skin of the subject. 4. The device of claim 3, further comprising a support structure constructed and arranged to move the plurality of microneedles into contact with the skin, and to withdrawal the plurality of microneedles from the skin after contact with the skin. 5. The device of claim 1, further comprising a reaction entity contained within the storage chamber able to react with an analyte contained within the blood, wherein a product of the reaction entity reacted with the analyte is determinable by the sensor. 6. The device of claim 1, wherein the vacuum chamber does not have a variable volume. 7. The device of claim 1, further comprising a signal generator for generating a signal relating to the blood determined by the sensor. 8. The device of claim 1, further comprising a display able to display a result of the determination of the analyte by the sensor. 9. A device for withdrawing blood from the skin and/or from beneath the skin of a subject, the device comprising: a vacuum chamber having an internal pressure less than atmospheric pressure before blood is withdrawn into the device and before the device is applied to the skin of the subject;a storage chamber, separate from the vacuum chamber;a microfluidic channel, discrete from the vacuum chamber and the storage chamber and having a smaller cross-sectional area than the cross-sectional area of the storage chamber;a support structure constructed and arranged to be positioned proximate the skin of the subject;a plurality of microneedles, associated with the support structure, for accessing blood from and/or through the skin of the subject; anda sensor for determining a condition of the blood withdrawn into the device,wherein the vacuum chamber is in fluid communication with the storage chamber via the microfluidic channel, and wherein the support structure is constructed and arranged to move the plurality of microneedles into contact with the skin, and to withdrawal... | Seventh Sense Biosystems, Inc. | A61B005/157 A61B005/15 A61B005/151 A61B005/145 A61B005/1459 A61B005/1486 A61B005/00
US09788749 | B2 | 2017-10-17 | System and method for detecting bone defects | 2014-03-07 | A system for detecting bone defects includes a dental implant contacting member, at least one electromagnetic source, at least one first magnetic field detecting device, at least one second magnetic field detecting device and a computing device. The dental implant contacting member has a magnetic body. The electromagnetic source is used to generate a variable magnetic field to vibrate the magnetic body. The first magnetic field detecting device is used to detect a magnetic field of the magnetic body and the variable magnetic field, so as to generate first detected data. The second magnetic field detecting device is used to detect the variable magnetic field, so as to generate second detected data. The computing device is used to obtain vibration data based on the difference between the first detected data and the second detected data. | 1. A system for detecting bone defects, the system comprising: a dental implant contacting member having a magnetic body;at least one electromagnetic source for generating a variable magnetic field to vibrate the magnetic body;at least one first magnetic field detecting device configured to detect a magnetic field of the magnetic body and the variable magnetic field, so as to generate first detected data;at least one second magnetic field detecting device configured to detect the variable magnetic field, so as to generate second detected data; andat least one computing device configured to detect bone defects based on vibration data, wherein the vibration data are obtained based on a difference between the first detected data and the second detected data. 2. The system of claim 1, wherein the second magnetic field detecting device is positioned out of a range of the magnetic field of the magnetic body. 3. The system of claim 2, wherein the first magnetic field detecting device and the second magnetic field detecting device are respectively positioned on two opposite ends of the electromagnetic source. 4. The system of claim 2, wherein the first magnetic field detecting device and the second magnetic field detecting device are positioned on the same end of the electromagnetic source. 5. The system of claim 1, further comprising: a frequency-sweep signal generator for controlling the variable magnetic field by using a frequency-sweep signal. 6. The system of claim 1, wherein the computing device comprises: a time-frequency domain converting unit, wherein the first detected data and the second detected data are time-domain data, and the time-frequency domain converting unit is used to convert the first detected data and the second detected data to first frequency-domain data and second frequency-domain data. 7. The system of claim 6, wherein the computing device comprises: a noise reduction unit for reducing noises in the first frequency-domain data and the second frequency-domain data. 8. The system of claim 6, wherein the computing device comprises: a frequency response computation unit for obtaining a frequency response function of the dental implant contacting member based on the first frequency-domain data and the second frequency-domain data, so as to obtain the vibration data comprising a vibration response. 9. The system of claim 1, wherein the first magnetic field detecting device or the second magnetic field detecting device is a Hall-effect sensor, or the first magnetic field detecting device and the second magnetic field detecting device are Hall-effect sensors. 10. The system of claim 1, further comprising: a transient signal generator for controlling the variable magnetic field by using a plurality of transient signals that have different frequencies. 11. A method for detecting bone defects, the method comprising: providing a detector having an electromagnetic source and at least one magnetic field detecting device disposed o... | NATIONAL CENTRAL UNIVERSITY | A61B005/05 A61B005/00 A61C008/00
US09782175 | B2 | 2017-10-10 | Systems, methods and apparatuses for real-time anastomosis guidance and surgical evaluation using optical coherence tomography | 2014-04-24 | A surgical evaluation and guidance system for use in extravacular evaluation and guidance of fluid vessel surgeries, includes an optical coherence tomography engine; a vessel probe optically coupled to the optical coherence tomography engine; and a signal processing and display system adapted to communicate with the optical coherence tomography engine to receive imaging signals therefrom. An optical coherence tomography vessel probe for use in extravacular evaluation and guidance of fluid vessel surgeries, includes a probe body having a proximal end and a distal end; an optical relay system disposed within the probe body; and a vessel adapter at least one of attached to or integral with the probe body at the distal end of the probe body, wherein the vessel adapter defines a groove that is configured to accept one of a fluid vessel and a lymphaticovenous vessel. | 1. A surgical evaluation and guidance system for use in extravascular evaluation and guidance of fluid vessel surgeries, comprising: an optical coherence tomography engine;a vessel probe optically coupled to the optical coherence tomography engine, the vessel probe having a longitudinal axis; anda signal processing and display system adapted to communicate with the optical coherence tomography engine and receive imaging signals therefrom,wherein the vessel probe is configured to accommodate a fluid vessel in a position and orientation to provide a plurality of B-mode optical coherence tomography images displaced along the fluid vessel relative to each other,wherein the signal processing and display system is configured to selectively display three-dimensional structural features of the fluid vessel and fluid flow through the fluid vessel, andwherein the vessel probe comprises a vessel adapter attached to a distal end thereof, the vessel adapter defining a vessel groove to facilitate positioning the fluid vessel, and a distal end of the vessel adapter having an inclined surface with respect to the longitudinal axis to facilitate measurements of the optical coherence tomography images of the fluid vessel, the inclined surface being in a range of 4-21 degrees with respect to the longitudinal axis of the vessel probe. 2. The surgical evaluation and guidance system according to claim 1, wherein the vessel probe is configured to accommodate at least one of blood vessels and lymphatic vessels. 3. A surgical evaluation and guidance system according to claim 1, wherein the surgical evaluation and guidance system comprises Doppler optical coherence tomography, and wherein the inclined surface is configured to facilitate acquisition of Doppler optical coherence tomography signals. 4. A surgical evaluation and guidance system according to claim 1, wherein the signal processing and display system is configured to selectively display three-dimensional structural features of said fluid vessel and fluid flow through the fluid vessel in real time during a procedure for evaluating or repairing the fluid vessel. 5. A surgical evaluation and guidance system according to claim 1, wherein said vessel probe has a size and shape to be hand operable during surgical guidance and evaluation. 6. An optical coherence tomography (OCT) vessel probe for use in extravascular evaluation and guidance of fluid vessel surgeries, comprising: a probe body having a proximal end and a distal end, wherein the probe body defines an optical port, the probe body having a longitudinal axis;an optical relay system disposed within said probe body; anda vessel adapter at least one of attached to or integral with said probe body at said distal end of said probe body, wherein said vessel adapter defines a vessel groove that is configured to accept one of a fluid vessel and a lymphaticovenous vessel to facilitate positioning the fluid vessel, and a distal end of the vessel adapter having an i... | THE JOHNS HOPKINS UNIVERSITY | A61B017/11 A61B005/00 A61B005/02 A61B005/026 A61B090/00
US09808206 | B1 | 2017-11-07 | Data acquisition quality and data fusion for personal portable wireless vital signs scanner | 2014-05-30 | In one embodiment of the invention, an interactive vital signs scanning method is disclosed including concurrently scanning for a plurality of vital signs with a portable vital signs scanner; detecting movement of the portable vital signs scanner during the scanning for the plurality of vital signs; and determining a measure of quality of the scanning for the plurality of vital signs with the portable vital signs scanner. In another embodiment, a method of improving the quality of vital signs data is disclosed including concurrently sensing data from a plurality of vital signs sensors over a period of time with a portable vital signs scanner; determining a plurality of vital sign values for a respective plurality of vital signs in response to the sensed data from the plurality of vital signs sensors over the period of time; and fusing at least two vital sign values of the plurality of vital sign values for the respective plurality of vital signs. | 1. A method comprising: with a portable vital signs scanner having a plurality of different vital signs sensors, receiving first motion data corresponding to a first motion of a user;receiving second motion data corresponding to a second motion of the user, the second motion data indicating that the second motion is smaller than the first motion;determining that a difference between the first motion data and the second motion data is greater than a predetermined threshold;concurrently sensing data from a user with the plurality of different vital signs sensors over a period of time;over the period of time, determining a plurality of different vital sign values for a respective plurality of different vital signs associated with the user in response to the sensed data from one or more of the plurality of different vital sign sensors; andbased on the difference determination between the first motion data and the second motion data, designating at least two vital sign values as valid;fusing together at least two different vital sign values that were designated as valid to improve the quality of at least one vital sign value of a vital sign. 2. The method of claim 1, further comprising: fusing a first vital sign value designated as valid with a second vital sign value designated as valid, the vital signs associated with the user;generating a third vital sign value associated with the user without the need for sensing the third vital sign value directly from the user; anddesignating the third vital sign value as valid. 3. The method of claim 1, further comprising: combining electrocardiogram (ECG) scan data and photoelectric plethysmogram (PPG) scan data together with user biometrics; andderiving a value for blood pressure,wherein the user biometrics are one or more of the set of height, weight, sex, and age of a user. 4. The method of claim 1, wherein scan data for at least two vital sign values of the plurality of vital sign values can be fused together to form a value for respiration rate or heart rate. 5. The method of claim 1, wherein sensed data from one vital sign sensor is used to derive a plurality of different vital sign values. 6. The method of claim 5, wherein sensed data from a pulse oximeter (PulseOX) sensor is used to derive values for blood oxygenation (SpO2) and pulse rate. 7. The method of claim 5, wherein sensed data from an electrocardiogram (ECG) sensor is used to derive heart rate and heart rate variability. 8. The method of claim 1, further comprising: sensing acceleration of the portable vital signs scanner;determining movement of the portable vital signs scanner along a first axis in response to the sensed acceleration;sensing sounds around the portable vital signs scanner; andfusing the movement along the first axis with sounds around the portable vital signs scanner to form a value for respiration rate. 9. The method of claim 1, further comprising: determining pulse rate from photoelectric plethysmogram (PPG) sensor d... | Scanadu, Inc. | A61B005/11 A61B005/00 A61B005/0205 A61B005/0295 A61B005/0402 A61B005/1455
US09794653 | B2 | 2017-10-17 | Methods and apparatus for improving signal quality in wearable biometric monitoring devices | 2015-08-18 | A wearable biometric monitoring device is configured to assess the biometric signal quality of one or more sensors associated with the monitoring device, determine how the user should adjust the device to improve the biometric fit, and instruct the user to wear the biometric monitoring device a certain way. Communicating instructions to a user may include instructing the user to execute a testing regimen while wearing the biometric monitoring device. The testing regimen facilitates an estimation of a signal quality that can be used to provide feedback to the user that he/she needs to adjust the device to improve the biometric fit and the biometric signal quality. | 1. A method of monitoring signal quality of a wearable biometric monitoring device, wherein the biometric monitoring device includes at least one optical sensor configured to detect and/or measure physiological information from a subject wearing the biometric monitoring device, at least one motion sensor, and at least one processor in communication with the at least one optical sensor and the at least one motion sensor, wherein the at least one processor is configured to receive and analyze signals produced by the at least one optical sensor and the at least one motion sensor, the method comprising the following performed via the at least one processor: processing information from the at least one optical sensor and/or the at least one motion sensor to determine if the monitoring device is being worn;initiating a signal quality test of the at least one optical sensor in response to determining that the monitoring device is being worn, comprising: instructing the subject to begin an exercise regimen; andprocessing information from the at least one optical sensor and the at least one motion sensor during the exercise regimen to determine if signal quality produced by the at least one optical sensor is acceptable, comprising calculating a ratio of physiological information to motion information in a signal from the at least one optical sensor, wherein a higher ratio of physiological information to motion information indicates a higher quality of the signal; andcommunicating information to the subject regarding the signal quality during the exercise regimen, the communicating comprising informing the subject that the device is being worn correctly if the signal quality is acceptable, and informing the subject that the device is not being worn correctly if the signal quality is not acceptable. 2. The method of claim 1, wherein communicating information to the subject regarding the signal quality further comprises communicating instructions to the subject to adjust the biometric monitoring device relative to the body of the subject in response to detecting that the signal quality is not acceptable. 3. The method of claim 1, wherein communicating information to the subject regarding the signal quality comprises communicating information to the subject that the signal quality is above a threshold level. 4. The method of claim 1, wherein instructing the subject to begin an exercise regimen comprises sending an audio and/or visual communication to a remote device in communication with the biometric monitoring device. 5. The method of claim 1, wherein communicating information to the subject regarding the signal quality comprises sending an audio and/or visual communication to a remote device in communication with the biometric monitoring device. 6. The method of claim 1, wherein instructing the subject to begin an exercise regimen comprises causing the biometric monitoring device and/or a remote device to vibrate. 7. The method of claim 1, wherein co... | Valencell, Inc. | H04Q009/00 G08B021/18 H04R001/10 A61B005/00 A61B005/024 A61B005/11 A61B005/1455
US09791396 | B2 | 2017-10-17 | Space adaptive reconstruction technique for adaptive electrical capacitance volume tomography | 2015-05-07 | A capacitance system having a capacitance sensor provides a high resolution Spatial-Adaptive Reconstruction Technique (SART) for use with Adaptive Electrical Capacitance Volume Tomography (AECVT). The system is adapted to analyze information provided by an image reconstruction of a first spatial region of an imaging domain of the sensor; provide control signals to the sensor to increase the resolution at a second spatial region of the imaging domain of the sensor based on the prior image reconstruction of the first spatial region; obtain image reconstruction information for the second spatial region of the imaging domain; and combine the image reconstruction information for the first spatial region and the second spatial region to obtain a combined image of the imaging domain. | 1. A system for tomographic imaging in a pipe, tube or other object comprising: a sensor, comprising at least four electrodes, positioned in a predetermined arrangement relative to each other around the pipe, tube, or object; each of the electrodes comprising a plurality of capacitance plate segments, each of the capacitance plate segments of each of the electrodes defining a capacitor when paired with a corresponding capacitance plate segment of one of the other electrodes;a voltage source, arranged to impart a voltage to selected capacitance plate segments, a sensitivity of the sensor being controlled by at least one of: the amount and the distribution of the imparted voltage applied to the selected capacitance plate segments;a measuring circuit, connected to each of the at least four electrodes to detect a capacitance of any of the capacitors between any selected pair of the electrodes though currents induced in the measuring circuit; anda first processing system, in communication with the measuring circuit, the first processing system programmed with instructions for executing on the first processing system to: 1) convert capacitance data collected therefrom into an image; 2) analyze information provided by an image reconstruction of a peripheral region of an imaging domain of the sensor; and 3) provide control signals to the sensor to increase the resolution at a center region of the imaging domain of the sensor based on the prior image reconstruction of the peripheral region; 4) obtain image reconstruction information for the center region of the imaging domain; and 5) combine the image reconstruction information for the peripheral region and the center region to obtain a combined image of the imaging domain. 2. A system for tomographic imaging according to claim 1, wherein the sensitivity of the sensor can be controlled by changing the frequency of the voltage distribution applied to at least one electrode. 3. A system for tomographic imaging according to claim 1, wherein the system is adapted to activate one electrode at a time as a source electrode, one electrode as the detecting electrode, and wherein the detecting electrode is connected to the measuring circuit. 4. A system for tomographic imaging according to claim 1, wherein the system is configured to obtain the capacitance between all the electrodes and to reconstruct an image of the region between the electrodes based on the capacitances obtained. 5. A system for tomographic imaging according to claim 1, wherein each capacitance plate segment of each electrode is individually addressable by a voltage that is connected to each of the capacitance plate segments through a switch. 6. A system for tomographic imaging according to claim 1, wherein the electric field between two electrodes can be focused based on the voltage levels used to address the capacitance plate segments. 7. A system for tomographic imaging according to claim 1, further comprising a plurality of switches and... | TECH4IMAGING LLC OHIO STATE INNOVATION FOUNDATION | G01N027/22 G06T011/00 A61B005/053
US09808621 | B2 | 2017-11-07 | System and method for reversing the effects of paralysis | 2015-08-05 | A method and system for applying therapeutic stimulation to a patient stricken with paralysis or other condition. The inventive method includes the steps of: inserting an acupuncture needle against the spine of the patient and applying an electrical potential to the needle with a signal that mimics natural neural impulses in the body. The needle is located depending on the nature of the patient's condition. In the illustrative application, the needle is inserted against the ligamentum flavum of the patient's spine at L3. In the best mode, the patient's legs are moved during the application of the electrical potential. The inventive system is adapted to minimize a notch reaction and includes an electrode inserted into a tissue of the patient and a circuit for applying an electrical potential to the electrodes with a signal that mimics natural neural impulses in the body. | 1. A system for restoring motor control of a patient including: an elongate acupuncture needle adapted for nonsurgical transcutaneous insertion into a spinal cavity of the patient whereby a portion of said needle is internal to the patient's body at least partially in parallel with and against the dura and ligament tissue and against one of the dura and ligament tissue of the patient's spine in the epidural space in a spinous process to repair spinal cord transection and restore muscle control and a portion of the needle is external to the patient's body;a circuit for applying an electrical potential to the external portion of the needle with a signal that mimics natural neural impulses in the patient's body in order to transmit and signals along the spinal column to restore movement; andmeans including a motor for moving a limb of the patient during the application of the signal. 2. The invention of claim 1 further including means for adjusting the motor in response to a detection of neural impulses during the application of the signal. 3. The invention of claim 1 wherein the system includes a memory. 4. The invention of claim 3 wherein the memory stores a database of signal profiles. 5. The invention of claim 3 wherein the memory includes a program which when executed by the controller applies one or more selected stimulation profiles to a patient through the needle. 6. The invention of claim 5 wherein the program stores data relating to a patient's therapy sessions with respect to the profiles. 7. The invention of claim 6 further including means for detecting and monitoring the patient's neural activity. 8. The invention of claim 7 wherein the program includes code for adjusting the profiles in response to the detected neural activity. 9. The invention of claim 1 wherein the system includes an electroencephalogram. 10. A system for applying therapeutic stimulation to a patient stricken with paralysis including: an elongate flexible acupuncture needle adapted for nonsurgical transcutaneous insertion into a spinal cavity of the patient whereby a portion of said needle is internal to the patient's body at least partially in parallel with and against the dura and ligament tissue and against one of the dura and ligament tissue of the patient's spine in the epidural space in a spinous process to repair spinal cord transection and restore muscle control;a system for applying an electrical potential to a portion of the needle external to the patient's body with a signal that mimics natural neural impulses in the patient's body in order to transmit said signals along the spinal column to restore movement;means for moving a limb of the body during the application of the neural impulses, said means for moving including a motor and means for adjusting the motor in response to a detection of neural impulses during the application of the signal;means for detecting the patient's neural activity during the application of the signal; andmeans for chan... | - | A61N001/36 A61B005/0478 A61N001/05
US09801562 | B1 | 2017-10-31 | Cardiac monitoring and diagnostic systems, methods, and devices | 2015-06-01 | The disclosure herein provides cardiac monitoring and diagnostic system, methods, and devices. A cardiac diagnostic system for detecting potential indicators of cardiac disease comprises: a cardiac monitoring device comprising at least one electrode configured to detect an analog electrical cardiac signal of a user; a cardiac signal filter configured to convert the analog electrical cardiac signal to digital electrocardiogram data; an atypical cardiac sequence detector configured to detect atypical cardiac patterns in the electrocardiogram data that are potentially indicative of cardiac disease; and a notification transmitter configured to generate an alert when a detected atypical cardiac pattern is determined to be indicative of cardiac disease. | 1. A cardiac diagnostic system for detecting potential indicators of cardiac disease, the cardiac diagnostic system comprising: a cardiac monitoring device comprising at least one electrode configured to detect an analog electrical cardiac signal of a user;a cardiac signal filter configured to convert the analog electrical cardiac signal to digital electrocardiogram data; andan atypical cardiac sequence detector configured to analyze the digital electrocardiogram data to detect atypical cardiac patterns that are potentially indicative of cardiac disease,wherein detecting the atypical cardiac patterns comprises: determining a first encoded length of a segment of the digital electrocardiogram data using a first encoder based on typical data;determining a second encoded length of the segment of the digital electrocardiogram data using a second encoder without knowledge of typical data, so that the encoded data is decodable; andcomparing the first and second encoded lengths,wherein the second encoded length being smaller than the first encoded length is indicative of an atypical cardiac pattern. 2. The cardiac diagnostic system of claim 1, further comprising: an atypical cardiac sequence processor configured to analyze detected atypical cardiac patterns in conjunction with historical data to determine whether the detected cardiac patterns are indicative of cardiac disease. 3. The cardiac diagnostic system of claim 2, wherein the atypical cardiac sequence processor is further configured to generate an alert when a detected atypical cardiac pattern is determined to be indicative of cardiac disease. 4. The cardiac diagnostic system of claim 1, wherein the first encoder comprises a typical encoder and the second encoder comprises a universal encoder. 5. The cardiac diagnostic system of claim 1, wherein the first encoder comprises a universal encoder that has been trained and frozen. 6. The cardiac diagnostic system of claim 5, wherein the second encoder comprises the same universal encoder, but not having been frozen. 7. The cardiac diagnostic system of claim 1, wherein the atypical sequence detector is further configured to store a detected pattern in a known patterns database. 8. The cardiac diagnostic system of claim 1, wherein detecting the atypical cardiac patterns further comprises: determining a difference in the first and second encoded lengths;comparing the determined difference to a threshold value; andflagging the segment as atypical when the determined difference is greater than the threshold value. 9. The cardiac diagnostic system of claim 1, further comprising one or more computer processors configured to operate the cardiac signal filter and atypical cardiac sequence detector. 10. A cardiac diagnostic system for detecting potential indicators of cardiac disease, the cardiac diagnostic system comprising: an atypical cardiac sequence detector configured to analyze digital electrocardiogram data to detect atypical cardiac patterns th... | UNIVERSITY OF HAWAII | A61B005/0402 A61B005/04 A61B005/0452 A61B005/00
US09782100 | B2 | 2017-10-10 | Bio-information analysis device | 2015-02-26 | A bio-information analysis device includes a measuring unit that measures biometric waveforms of a subject, and a calculating unit that calculates a matching rate of the biometric waveforms based on shapes and magnitudes of the biometric waveforms measured by the measuring unit. In the bio-information analysis device, the biometric waveforms may be electrocardiogram waveforms. | 1. A bio-information analysis device, comprising: a measuring unit that measures a first shape and a second magnitude of a first biometric waveform of a subject and a second shape and a second magnitude of a second biometric waveform of the subject; anda calculating unit that calculates a similarity rate between the first biometric waveform and the second biometric waveform based on angles formed by first and second space vectors calculated from the first biometric waveform and the second biometric waveform measured by the measuring unit and the magnitudes of the space vectors. 2. The bio-information analysis device according to claim 1, wherein the first biometric waveform and the second biometric waveform are electrocardiogram waveforms. 3. The bio-information analysis device according to claim 2, wherein the first biometric waveform includes a waveform measured when a premature ventricular contraction is occurred and the second biometric waveform includes a waveform measured when a waveform is artificially occurred by pacing. 4. The bio-information analysis device according to claim 3, wherein when calculating the similarity rate between the first biometric waveform and the second biometric waveform, the similarity rate is calculated in such a manner that a point, at which a highest similarity rate is exhibited, of points on the second biometric waveform is overlapped with a highest amplitude point of the first biometric waveform. 5. The bio-information analysis device according to claim 3, wherein when calculating the similarity rate, a stimulation potential waveform component, which is generated upon pacing, of the second biometric waveform is excluded. 6. The bio-information analysis device according to claim 4, wherein when calculating the similarity rate, a stimulation potential waveform component, which is generated upon pacing, of the second biometric waveform is excluded. 7. The bio-information analysis device according to claim 3, wherein upon calculation of a similarity rate with respect to magnitudes of the first biometric waveform and the second biometric waveform, a magnitude of a portion having a phase reversed relative to that of the first biometric waveform, of the second biometric waveform, is changed to zero. 8. The bio-information analysis device according to claim 4, wherein upon calculation of a similarity rate with respect to magnitudes of the first biometric waveform and the second biometric waveform, a magnitude of a portion having a phase reversed relative to that of the first biometric waveform, of the second biometric waveform, is changed to zero. 9. The bio-information analysis device according to claim 5, wherein upon calculation of a similarity rate with respect to magnitudes of the first biometric waveform and the second biometric waveform, a magnitude of a portion having a phase reversed relative to that of the first biometric waveform, of the second biometric waveform, is changed to zero. 10. The bio-i... | NIHON KOHDEN CORPORATION | A61B005/00 A61B005/04 A61B005/0452 A61N001/37 A61N001/365
US09795317 | B2 | 2017-10-24 | Body composition measurement device, body composition measurement method, and correction method in body composition measurement | 2013-10-23 | Provided is a body composition measurement device capable of accurately acquiring visceral fat quantity-related information as body composition-related information of a human body. A body composition measurement device has a voltage detection unit that detects a voltage that occurs at a human-body-mimic resistor in order to measure a resistance value of a compensating resistor unit and detects a voltage that occurs at voltage measurement electrode pair as a result of current that is generated by a current generation unit and that is supplied to the abdomen of the subject via the current supply electrode pair, and a calculation unit that generates body composition-related information for the subject on the basis of the voltage occurring at the voltage measurement electrode pair and the voltage generated at the human-body-mimic resistor, which have been detected by the voltage detection unit. | 1. A body composition measurement device comprising: a current supply electrode pair and a voltage measurement electrode pair that contact an abdomen of a measuring subject;a current generation unit configured to generate current;a voltage detection unit configured to detect voltage generated by a correction resistor unit to measure resistance of the correction resistor unit, wherein the voltage detection unit is configured to detect voltage generated by the voltage measurement electrode pair using current generated by the current generation unit and supplied to the abdomen of the measuring subject through the current supply electrode pair; anda computation unit configured to generate body composition related information based on the voltage generated by the voltage measurement electrode pair and the voltage generated by the correction resistor unit, which are detected by the voltage detection unit,wherein the correction resistor unit includes a plurality of simulation resistor units, and each of the plurality of simulation resistor units includes a human body simulation resistor configured to simulate an abdomen of a measuring subject, andwherein the computation unit is configured to derive an approximate curve using measurement values of the plurality of human body simulation resistors of the plurality of simulation resistor units and generate the body composition related information using the approximate curve. 2. The body composition measurement device according to claim 1, wherein the computation unit configured to generate visceral fat amount related information as the body composition related information. 3. The body composition measurement device according to claim 1, comprising: an operation unit configured to generate a control signal; anda switching unit configured to supply either one of the current supply electrode pair and the correction resistor unit with the current generated by the current generation unit in accordance with the control signal generated by the operation unit, wherein the switching unit is configured to select the voltage of either one of the voltage measurement electrode pair and the correction resistor unit and supplies the voltage to the voltage detection unit in accordance with the control signal. 4. The body composition measurement device according to claim 1, wherein each of the plurality of simulation resistor units includes: a first electrode contact simulation resistor and a second electrode contact simulation resistor configured to simulate contact of the current supply electrode pair and the abdomen of the measuring subject, anda third electrode contact simulation resistor and a fourth electrode contact simulation resistor configured to simulate contact of the voltage measurement electrode pair and the abdomen of the measuring subject. 5. The body composition measurement device according to claim 3, wherein the switching unit includes a first switching circuit and a second switching circuit;the curre... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | A61B005/05 A61B005/053 A61B005/00
US09775533 | B2 | 2017-10-03 | System and method of determining a risk score for triage | 2013-03-08 | The present disclosure provides a system and method of determining a risk score for triage. In particular, a system is provided for providing an assessment of risk of a cardiac event for a patient, for example an incoming patient to a hospital emergency department complaining of chest pain. In the disclosure, the system includes an input device for measuring physiological data based vital signs parameter of the patient, a twelve-lead electrocardiogram (ECG) device for establishing an ECG obtained from results of the electrocardiography procedure, and determining an ECG parameter and a heart rate variability (HRV) parameter therefrom. An ensemble-based scoring system is further provided, establishing weighted classifier based on past patient data and where the vital signs parameter, the ECG parameter and the HRV parameter are compared to corresponding weighted classifiers to determine a risk score. A corresponding method to determine a risk score for triage is also provided. | 1. A system for determining a risk score for a triage, comprising: a first input device for measuring a first input parameter relating to physiological data of a patient, the first input parameter comprising a vital signs parameter;a twelve-lead electrode electrocardiogram (ECG) device, for carrying out an electrocardiography procedure on the patient, and establishing an ECG obtained from results of the electrocardiography procedure, the ECG device comprising an ECG extraction module to extract at least one ECG parameter from the ECG;a heart rate variability (HRV) analysis module for determining a HRV analysis from the ECG, the HRV analysis comprising at least one HRV parameter; andan ensemble-based scoring system, comprising: a plurality of weighted classifiers for providing a risk score calculation, the plurality of weighted classifiers established based on past patient data in a database of accumulated past patient data; andan analysis module for receiving the first input parameter, the at least one HRV parameter, and the at least one ECG parameter which are communicated or transmitted to the ensemble-based scoring system,a sorting module arranged to receive data from the database of accumulated past patient data, and sort the data into a plurality of data sets, each data set corresponding to a classifier, and comprising an imbalanced data set;a sampling module arranged to receive a first imbalanced data set corresponding to a first classifier comprising a first majority data set comprising a first number of data samples, and a first minority data set comprising a second number of data samples, from the sorting module, and extract a first majority data subset comprising a third number of samples from the first majority data set, wherein the third number of samples in the first majority data subset is equal to the second number of samples in the first minority data set;a classifier generation module for establishing the plurality of weighted classifiers based on past patient data, wherein the classifier generation module further comprises a training module arranged to receive the first majority data subset and the first minority data set from the sampling module and build a first classification model to represent the first classifier with the first majority data subset and the first minority data set;an over-sampling module arranged to receive the first majority data subset and the first minority data set from the sampling module, and create a first synthetic data set by applying a process of synthetic over-sampling with replacement on the first majority data subset and the first minority data set; anda validation module arranged to validate the first classification model against the first synthetic data set, and obtain a resultant prediction accuracy of the first classification model representing the importance of the first classifier,wherein the analysis module determines a risk score by comparing the first input parameter, the at least one H... | SINGAPORE HEALTH SERVICES PTE LTD | A61B005/04 A61B005/024 A61B005/0452 G06F019/00 A61B005/0205 A61B005/00 A61B005/0408 A61B005/0456 A61B005/046 A61B005/0468 A61B005/0472 A61B005/1455
US09795365 | B2 | 2017-10-24 | MRI biopsy apparatus incorporating a sleeve and multi-function obturator | 2013-04-24 | An apparatus for use with a minimally invasive medical procedure into human breast tissue includes a cannula and an obturator. The cannula includes an open distal end, a lateral opening proximate to the open distal end, and a longitudinal lumen communicating with the lateral opening and the open distal end. The lumen has a non-circular cross-section. The obturator is sized for insertion into the cannula. The obturator has a distal end extending from the open distal end of the cannula when the obturator is inserted fully into the cannula. The obturator has a recess proximate of the distal end of the obturator. The recess is positioned along a portion of the length of the obturator to align with the lateral opening of the cannula when the obturator is inserted fully into the cannula. | 1. A targeting set for use in a minimally invasive magnetic resonance imaging (MRI) guided procedure for obtaining tissue samples of human breast tissue, the targeting set comprising: (a) a cannula having an open distal end and an open proximal end sized for receiving a tissue biopsy device for obtaining tissue samples from a breast, wherein the cannula is formed of an MIRI compatible matetial, wherein the cannula further defines a lateral opening having a proximal end and a distal end, wherein the distal end of the lateral opening is proximate to the open distal end of the cannula, wherein the cannula further includes a longitudinal lumen communicating with the lateral opening and the open distal end, the lumen having a non-circular cross-section; and(b) an obturator, wherein the obturator is formed of an MRI compatible material, wherein the cannula is sized to receive the obturator for targeting purposes prior to receiving the biopsy device for obtaining tissue samples, the obturator having a tissue piercing distal end extending from the open distal end of the cannula when the obturator is inserted fully into the cannula, and the obturator having a recess proximate to the distal end of the obturator, wherein the recess includes a proximal end and a distal end, wherein the recess is positioned along a portion of the length of the obturator such that the proximal and distal ends of the recess are respectively adjacent to the proximal and distal ends of the lateral opening of the cannula when the obturator is inserted fully into the cannula, wherein the lateral opening of the cannula and the recess of the obturator together define a tissue receiving dug-out, wherein the tissue receiving dug-out is configured to receive tissue thereby providing a reference for locating the lateral opening of the cannula under MRI imaging. 2. The targeting set of claim 1, wherein the longitudinal lumen of cannula is configured to receive a biopsy cutter member. 3. The targeting set of claim 2, wherein the obturator is configured to be inserted within the cannula in lieu of the biopsy cutting member. 4. The targeting set of claim 1, wherein the obturator includes a lumen extending from a proximal end of the obturator and communicating with the obturator recess. 5. The targeting set of claim 4, wherein the obturator includes a marker insert with fluid leak passages positioned in the lumen. 6. The targeting set of claim 1, wherein the recess is configured to receive an MRI visible material to enhance local contrast to provide positive identification of the lateral opening of the cannula. 7. The targeting set of claim 1, wherein a selected one of the cannula and the obturator includes a liquid crystal polymer. 8. The targeting set of claim 1, wherein a selected one of the cannula and the obturator includes a poly ether ketone (PEEK) material. 9. The targeting set of claim 1, further includes a hub coupled to the proximal end of the obturator. 10. The targeting ... | Devicor Medical Products, Inc. | A61B017/34 A61B010/02 A61B005/055 A61B090/17 A61B090/11 A61B010/00 A61B017/00 A61B090/00
US09817440 | B2 | 2017-11-14 | Garments having stretchable and conductive ink | 2015-07-14 | Methods of forming garments having one or more stretchable conductive ink patterns. Described herein are method of making garments (including compression garments) having one or more highly stretchable conductive ink pattern formed of a composite of an insulative adhesive, a conductive ink, and an intermediate gradient zone between the adhesive and conductive ink. The conductive ink typically includes between about 40-60% conductive particles, between about 30-50% binder; between about 3-7% solvent; and between about 3-7% thickener. The stretchable conductive ink patterns may be stretched more than twice their length without breaking or rupturing. | 1. A wearable electronics device, the device comprising: a garment comprising a fabric; andat least one conductive ink pattern on the garment, wherein the conductive ink pattern comprises: a layer of conductive ink having: between about 40-60% conductive particles, between about 30-50% binder; between about 3-7% solvent; and between about 3-7% thickener;a layer of an elastic adhesive on the garment; anda gradient region between the conductive ink and the adhesive having a thickness of greater than about 10 μm, the gradient region comprising a nonhomogeneous mixture of the conductive ink and the adhesive wherein the concentration of conductive ink decreases from a region closer to the layer of conductive ink to the layer of elastic adhesive. 2. The device of claim 1, wherein the thickness of the layer of elastic adhesive is greater than the thickness of the gradient region and the thickness of the layer of conductive ink is less than the thickness of the elastic layer. 3. The device of claim 1, wherein the garment is a compression garment. 4. The device of claim 1, wherein the conductive particles comprise particles of carbon black. 5. The device of claim 1, wherein the conductive particles comprise particles of one or more of: carbon black, mica, graphene, graphite, silver metal powder, copper metal powder, or iron metal powder. 6. The device of claim 1, wherein the binder comprises acrylic binder. 7. The device of claim 1, wherein the solvent comprises propylenyc glycol. 8. The device of claim 1, wherein the thickener comprises polyurethanic thickener. 9. The device of claim 1, wherein the elastic adhesive comprises a thermo-adhesive water-based glue that is electrically insulative. 10. The device of claim 1, further comprising an insulating resin at least partially over the layer of conductive ink. 11. The device of claim 1, wherein the conductive ink pattern comprises a plurality of layers of the conductive ink. 12. The device of claim 1, wherein a resistivity of the conductive trace is less than about 10 Kohms/square. 13. The device of claim 1, wherein a resistivity of the conductive pattern varies with applied stretch. 14. The device of claim 1, wherein the conductive ink pattern is configured to stretch up to 500% of a resting length without breaking. 15. The device of claim 1, wherein conductive ink pattern is formed as a sensor. 16. The device of claim 1, wherein the conductive ink pattern is formed as a trace. 17. The device of claim 1, wherein the conductive ink pattern is configured as an electrode. 18. The device of claim 1, further comprising a conductive wire coupled to the garment and connected at one end to the conductive ink pattern. 19. A wearable electronics device, the device comprising: a garment comprising a compression fabric; andat least one stretchable and conductive ink pattern on the garment having a sheet resistivity of less than about 10 Kohms/square, wherein the conductive ink pattern is stretchable... | L.I.F.E. Corporation S.A. | G06F001/16 A41D001/00 A61B005/00 A41D013/12 D06N003/00 D06N003/04 H05K001/03 H05K001/09 H05K001/02
US09814409 | B2 | 2017-11-14 | Detection of chronic obstructive pulmonary disease exacerbations from breathing patterns | 2014-11-11 | An apparatus may include a sensing circuit and a processor. The sensing circuit is configured to generate a sensed physiological signal, wherein the physiological signal includes respiration information of a subject. The processor includes an end expiratory volume (EEV) module configured to determine a value of EEV of the subject using the sensed physiological signal, and a lung hyperinflation detection module configured to generate an indication of lung hyperinflation of the subject according to the value of EEV and provide the indication to at least one of a user or process. | 1. A system comprising: a first implantable medical device including:a sensing circuit configured to generate a sensed physiological signal representative of respiration of a subject when operatively coupled to electrodes configured to contact the subject;a communication circuit configured to communicate information with a separate device; anda first processor including an end expiratory volume (EEV) module configured to determine a measure representative of EEV of the subject using the sensed physiological signal; anda second medical device including:a communication circuit configured to communicate information with the first medical device; anda second processor including a lung hyperinflation detection module configured to generate an indication of lung hyperinflation of the subject according to the measure of EEV and provide the indication to a user. 2. The system of claim 1, wherein the second device includes a display and the second processor includes a trending module, wherein the trending module is configured to receive values of the measure of EEV, trend the values of the measure of EEV, and display the trend of the measure of EEV with values of a measured respiration parameter of the subject. 3. The system of claim 1, wherein the EEV module is configured to measure a respiration parameter of the subject using the sensed physiological signal or a second sensed physiological signal representative of respiration of the subject, andwherein the second medical device includes a display and the second processor includes a trending module, wherein the trending module is configured to determine a lung hyperinflation index (HII) using the measure of EEV and the measured respiration parameter, trend values of the determined HII of the subject, and display the trend of the determined HII. 4. The system of claim 1, wherein the second medical device includes: a display;a port configured to receive an indication of a patient symptom into the second medical device; andwherein the second processor includes a correlation module configured to correlate measurements of EEV with the patient symptom and display the correlation between the measurements of EEV and the patient symptom. 5. The system of claim 1, wherein the wherein the second processor includes a prescription module configured to generate a recommendation of therapy for the subject according to the indication of lung hyperinflation and provide the indication to at least one of a user or process. 6. The system of claim 1, wherein second medical device is configured to initiate delivery of therapy in response to the indication of lung hyperinflation. 7. The system of claim 1, wherein the first medical device includes a sensing circuit configured to determine a level of physical activity of the subject;wherein the EEV module is configured to measure a respiration parameter of the subject using the sensed physiological signal or a second sensed physiological signal representative of respiratio... | Cardiac Pacemakers, Inc. | A61B005/085 A61B005/053 A61N001/36 A61B005/08 A61B005/00
US09782093 | B2 | 2017-10-10 | Vector-based shock indication | 2015-03-13 | A system for managing care of a person receiving emergency cardiac assistance includes one or more capacitors arranged to deliver a defibrillating shock to a person; one or more electronic ports for receiving a plurality of signals from sensors for obtaining indications of an electrocardiogram (ECG) for the person; and a patient treatment module executable on one or more computer processors using code stored in non-transitory media and to provide a determination of a likelihood of success from delivering a future defibrillating shock to the person with the one or more capacitors, using a mathematical computation applied to a vector value defined by signals from at least two of the plurality of signals. | 1. A defibrillation system for treating cardiac arrhythmia, the system comprising: an external defibrillator having one or more capacitors arranged to deliver a defibrillating shock to a person receiving emergency cardiac assistance;a plurality of electrodes configured to be coupled to the person and acquire a plurality of electrocardiogram (ECG) signals from the heart of the person; anda patient treatment module executable on one or more computer processors using code stored in non-transitory media and configured to execute operations comprising: receiving the plurality of ECG signals acquired by the plurality of electrodes,determining an ECG vector representation of the plurality of ECG signals comprising at least two orthogonal or near orthogonal signals indicative of the ECG electrical vector of the person, wherein the ECG vector representation includes information of both magnitude and direction of the ECG electrical vector,identifying an angle for the ECG vector representation that provides a maximum amplitude of a projection of the ECG vector representation over a sampling interval,determining a likelihood of shock success of delivering a defibrillating shock to the person with the one or more capacitors based on the identified angle of projection for the vector representation, andaffecting control of the external defibrillator based on the determined likelihood of shock success. 2. The system of claim 1, wherein the operations comprise performing one or more Fast Fourier Transforms (FFTs) on the plurality of ECG signals. 3. The system of claim 2, wherein the operations comprise performing one or more amplitude spectrum area calculations on the plurality of ECG signals. 4. The system of claim 1, wherein the operations comprise applying a pre-transform to the plurality of ECG signals so as to decrease an error of determining a likelihood of success from delivering the future defibrillating shock to less than about 10 percent. 5. The system of claim 4, wherein the pre-transform is applied in response to determining that the plurality of signals were not previously orthogonal or near orthogonal. 6. The system of claim 1, wherein the patient treatment module is programmed to apply a mathematical computation to the vector representation by calculating FFT for each of the plurality of ECG signals to create processed values and then combining the processed values. 7. The system of claim 6, wherein combining the processed values comprises determining a root of a sum of the processed values. 8. The system of claim 1, wherein the operations comprise performing a mathematical transform on the plurality of ECG signals from a time domain to a frequency domain on a window of data. 9. The system of claim 8, wherein the window comprises a tapered window. 10. The system of claim 9, wherein the tapered window comprises a Tukey window. 11. The system of claim 10, wherein the tapered window is between about one second and about 2 seconds in width. ... | ZOLL Medical Corporation | A61B005/04 A61B005/00 A61N001/39 A61B005/053
US09826936 | B2 | 2017-11-28 | Body cavity physiological measurement device | 2015-04-03 | Provided herein is a self-contained physiological measuring device adapted for disposition within a patient body cavity, primarily the vagina, for an extended period of time (e.g., 6-48 hours or more). While disposed within the body cavity, the device periodically measures one or more physiological parameters at known locations within the body cavity. In addition to measuring such physiological parameters, the device is operative to store such measurements to memory for subsequent download/processing upon removal of the device from the body cavity and/or upon wireless interrogation. | 1. A self contained physiological monitoring device adapted for temporary disposition within a vaginal cavity free of connection outside of the vaginal cavity, comprising: a toroidal body adapted for insertion into the vaginal cavity free of connection outside of the vaginal cavity;a battery attached to said toroidal body;a memory device attached to said toroidal body; anda control circuit attached to said toroidal and operatively connected to said battery and said memory device, wherein said battery, said memory device, and said control circuit are disposed within the vaginal cavity when said toroidal body is inserted within the vaginal cavity;at least three sensors mounted at different locations about a periphery of said toroidal body, each said sensor being operatively connected to said battery and said control circuit, wherein each said sensor is operative to generate at least one sensor output of at least one physiological parameter for adjacent tissue;at least three accelerometers mounted at different locations about the periphery said toroidal body, said accelerometers being operative to generate vector outputs for receipt by said control circuit;wherein said control circuit is operative to: utilize said vector outputs from said at least three accelerometers to determine an angular orientation of said toroidal body relative to the vaginal cavity wherein a first portion of said toroidal body is identified as being proximate to the posterior fornix of the vaginal cavity and a second portion of said toroidal body is identified as being proximate to the anterior fornix of the vaginal cavity;determine a location of each said sensor about said periphery of said toroidal body relative to different regions of the vaginal cavity based on the angular orientation of the toroidal body relative to the vaginal cavity, wherein said sensors outputs are mapped to the different regions of the vaginal cavity; andstore the at least one sensor output from said sensors to the memory device connected to the toroidal body. 2. The device of claim 1, further comprising: at least one reservoir chamber mounted proximate to the body, wherein the reservoir chamber administers one or more therapeutic agents. 3. The device of claim 2, further comprising: a plurality of reservoir chambers mounted at said different angular locations about said toroidal body. 4. The device of claim 3, wherein said control circuit is operative to administer an agent from at least one of said plurality of reservoir chambers based on said angular orientation. 5. The device of claim 1, wherein said sensors are equally angularly spaced about said toroidal body. 6. The device of claim 5, wherein said accelerometer and said sensors are co-located at said different locations. 7. The device of claim 1, wherein said accelerometers comprise three-axis accelerometers. 8. The device of claim 1, wherein said control circuit is operative to calculate said angular orientation based on an estimated p... | - | A61B005/00 A61B005/11 A61B005/1455 A61B005/01
US09801558 | B2 | 2017-10-31 | Detection and monitoring of abdominal aortic aneurysm | 2011-07-14 | Ruptured Abdominal Aortic Aneurysms (AAA) cause a large number of deaths annually. Ruptures occur even in people who are already diagnosed with AAA and are being monitored. The reason is that the interval between tests is too long because of the need to visit a pathological facility with imaging equipment. It is preferable to estimate the progress of AAA frequently, once detected, in a non-invasive manner, preferably at the subject's home, without the need for the subject to visit a pathological facility. A device is disclosed for detecting a state of a vascular pathology of a subject, comprising a sensor signal unit (103) for providing a signal representative of a blood volume in a body part of a subject, a comparator (107) for comparing the sensor signal with a reference signal, and a user interface (109) for conveying a result based on the comparison to a user of the device. | 1. A device for testing a subject for estimating a progress of a state of an existing abdominal aorta aneurysm of the subject based on a determined change over time, the device comprising: a sensor configured to sense a pulsation of blood volume in a body part of the subject;a memory;a processor configured to: acquire from the sensor an amplified sensor signal representative of the blood volume in the body part of the subject;acquire an electrocardiogram (ECG) signal of the subject via an ECG electrode, the ECG signal indicating an activity of the heart muscle of the subject;analyze a time interval between a peak of the ECG signal and the sensor signal for a number of cycles so as to synchronize the acquisition of the sensor signal with the ECG signal;increase the amplification of a DC component of the sensor signal to a higher amplification for the time period when the ECG signal indicates that the heart muscle of the subject is in a diastolic state, and reduce the amplification of the DC component of the sensor signal for the time period when the ECG signal indicates that the heart muscle of the subject is in a systolic state;compare one or more characteristics of the sensor signal with corresponding one or more characteristics of a reference signal stored in the memory, wherein the reference signal is a corresponding sensor signal previously acquired from a test on the subject a time period ago, timing information of the ECG signal being used to compensate for a difference in heart rates of the subject between acquisitions of the sensor signal and the reference signal; anddetermine a change in the one or more characteristics of the sensor signal over the time period; anda user interface configured to display a result of the determination,wherein the result provides an estimate of the progress of the state of the existing abdominal aorta aneurysm. 2. The device of claim 1, wherein the sensor is a photoplethysmogram sensor. 3. The device of claim 1, wherein the processor is further configured to: determine a change in one or more relationships between the sensor signal and the ECG signal of the subject,wherein the one or more relationships comprise the time interval between a peak of the ECG signal and the sensor signal,and further wherein the result of the determination is based upon the determined change in the one or more relationships between the sensor signal and the ECG signal. 4. A method of testing a subject for estimating a progress of a state of an existing abdominal aorta aneurysm of the subject based on a determined change over time, the method comprising: sensing a pulsation of blood volume in a body part of the subject with a sensor and acquiring an amplified sensor signal representative of the blood volume;acquiring, by a processor, an electrocardiogram (ECG) signal of the subject via an ECG electrode, the ECG signal indicating an activity of the heart muscle of the subject;analyzing, by the processor, a time interval between a... | KONINKLIJKE PHILIPS N.V. | A61B005/02 A61B005/0295 A61B005/024 A61B005/026 A61B005/0285 A61B005/04 A61B005/0402 A61B005/00
US09826959 | B2 | 2017-11-28 | Ultrasonic diagnostic device | 2009-11-02 | An ultrasonic diagnostic device includes: a transmission and reception circuit that processes reception signals output from an ultrasonic probe to generate reception data; a memory that stores the reception data; an image data generation portion that generates B-mode image data and M-mode image data based on the reception data; a detection portion that detects hand movement or body movement in the reception data; a determination portion that detects an amount of change of a relative position between the ultrasonic probe and a body under test in the reception data in which the hand movement or the body movement is detected, and determine an accuracy of the reception data; and an elastic characteristic calculation portion that measures a displacement magnitude of a living body tissue in the reception data determined by the determination portion to be highly accurate, and calculates elastic characteristic by using the displacement magnitude. | 1. An ultrasonic diagnostic device comprising: an ultrasonic probe including a plurality of ultrasonic transducers transmitting and receiving ultrasound to and from a living body tissue moving in synchronization with a pulse of a body under test;a transmission circuit configured to supply a plurality of drive signals to said ultrasonic probe;a reception circuit configured to process a plurality of reception signals output from said ultrasonic probe to generate reception data;a memory configured to store said reception data generated by said reception circuit; anda central processing unit configured to execute software instructions; wherein the software instructions instruct the central processing unit to:generate B-mode image data representing a B-mode image of said living body tissue and M-mode image data representing at least one M-mode image showing temporal changes of said living body tissue for a predetermined period on at least one line of interest set in said B-mode image, based on said reception data stored in said memory;detect an amount of change of a relative position between said ultrasonic probe and said body under test in said reception data corresponding to said at least one M-mode image for said predetermined period;determine whether at least one of hand movement and body movement occurs or not in said reception data corresponding to said at least one M-mode image for said predetermined period, based on said amount of change of said relative position;display said at least one M-mode image together with information indicating whether the at least one of hand movement and body movement occurs or not at any point of time shown in said at least one M-mode image for said predetermined period, on a screen of a display; andmeasure a displacement magnitude of said living body tissue in said reception data corresponding to at least one M-mode image in which hand movement and body movement do not occur while eliminating any M-mode image in which at least one of hand movement and body movement occurs, and calculate an elastic characteristic of an area within said living body tissue by using said displacement magnitude. 2. The ultrasonic diagnostic device of claim 1, wherein said central processing unit is instructed to detect said amount of change of said relative position in said reception data obtained from said living body tissue between (i) said ultrasonic probe which is pressed onto said body under test and (ii) said area of which said elastic characteristic is calculated. 3. The ultrasonic diagnostic device of claim 1, wherein said central processing unit is instructed to determine that the at least one of hand movement and body movement occurs in a case where said amount of change of said relative position is larger than a threshold value. 4. The ultrasonic diagnostic device of claim 3, wherein said threshold value for said amount of change of said relative position is not less than 0.1 mm. 5. The ultrasonic diagnostic device of c... | FUJIFILM Corporation | A61B008/00 A61B008/08 A61B005/02 A61B005/022 A61B005/00
US09782127 | B2 | 2017-10-10 | Photoelectric sphygmograph measurement device | 2011-06-27 | A photoelectric sphygmograph measurement device includes a light emitting part emitting light pulses, a light receiving part having a light receiving element that receives the light pulses and producing a corresponding output signal, and a control part driving the light emitting part to emit the light pulses and performing a pulse wave measurement by using the output signal. The control part drives the light emitting part so that a charge accumulated in the light emitting element converges on a predetermined amount in a case where the charge accumulated in the light receiving element decreases to the predetermined amount or less during a time when driving of the light emitting part is stopped and/or in a case where the light receiving element changes to a first state in which the light receiving element is capable of receiving the light pulses from a second state in which the light receiving element is not capable of receiving the light pulses. | 1. A photoelectric sphygmograph measurement device comprising: a light emitting circuit configured to emit light pulses;a light receiving circuit configured to have a light receiving element that receives the light pulses emitted by the light emitting circuit, to produce an output signal corresponding to the light pulses received by the receiving element, and to output the output signal; anda control circuit configured to drive the light emitting circuit to emit the light pulses and perform a pulse wave measurement based on the output signal of the light receiving circuit,wherein the control circuit is configured to:determine whether a signal level of the output signal of the light receiving circuit becomes less than a predetermined level while emission of the light pulses from the light emitting circuit is stopped by the control circuit; anddrive the light emitting circuit to emit the light pulses until the signal level of the output signal of the light receiving circuit becomes equal to or greater than the predetermined level and less than a first signal level, which is obtained when the light pulses emitted by the light emitting circuit are received directly by the light receiving element, when the control circuit has determined that the signal level of the output signal of the light receiving circuit becomes less than the predetermined level, wherein the control circuit is configured to: determine whether the light receiving element changes to a first state from a second state, the light receiving element being capable of receiving the light pulses in the first state, the light receiving element being incapable of receiving the light pulses in the second state; anddrive the light emitting circuit to emit the light pulses until the signal level of the output signal of the light receiving circuit reaches the predetermined level. 2. The photoelectric sphygmograph measurement device according to claim 1, wherein the control circuit is configured to: cause the light emitting circuit to temporarily stop emitting the light pulses during a time when the light pulses emitted by the light emitting circuit are received directly by the light receiving element:determine whether the signal level of the output signal of the light receiving circuit becomes less than the predetermined level after the light emitting circuit is caused to temporarily stop emitting the light pulses andenable the light emitting circuit to emit the light pulses until the signal level of the output signal becomes equal to or greater than the predetermined level and less than the first signal level when the control circuit has determined that the signal level of the output signal of the receiving circuit becomes less that the predetermined level. 3. The photoelectric sphygmograph measurement device according to claim 1, wherein when the light receiving element changes to the second state from the first state, the control circuit drives the light emitting circuit to emit the light p... | FUJITSU LIMITED | A61B005/02 A61B005/00 A61B005/024
US09808213 | B2 | 2017-11-07 | Image processing apparatus, image processing method, medical image diagnostic system, and storage medium | 2015-08-04 | An image processing apparatus includes: an obtaining unit configured to obtain three-dimensional image data and positions of a plurality of feature points in the three-dimensional image data; a designation unit configured to designate a designation point indicating a portion of interest in the three-dimensional image data; a selection unit configured to select, using information determined by a position of the designation point and the positions of the plurality of feature points, a feature point from the plurality of feature points; a generation unit configured to generate, from the three-dimensional image data, a cross section image on a plane determined based on the position of the designation point and a position of the selected feature point; and a display control unit configured to cause a display unit to display the generated cross section image. | 1. An image processing apparatus comprising: an obtaining unit configured to obtain three-dimensional image data and positions of a plurality of feature points in the three-dimensional image data;a designation unit configured to designate a designation point indicating a portion of interest in the three-dimensional image data;a selection unit configured to select, using information determined by a position of the designation point and the positions of the plurality of feature points, a feature point from the plurality of feature points;a generation unit configured to generate, from the three-dimensional image data, a cross section image on a plane determined based on the position of the designation point and a position of the selected feature point; anda display control unit configured to cause a display unit to display the generated cross section image. 2. The apparatus according to claim 1, wherein said selection unit selects two feature points from the plurality of feature points. 3. The apparatus according to claim 1, wherein said selection unit selects, based on distances between the designation point and the plurality of feature points, a feature point from the plurality of feature points. 4. The apparatus according to claim 3, wherein said selection unit selects a feature point positioned near the designation point. 5. The apparatus according to claim 3, wherein said selection unit selects a feature point positioned away from the designation point by not less than a predetermined distance. 6. The apparatus according to claim 1, wherein said generation unit generates, from the three-dimensional image data, a cross section image on a plane determined by the position of the designation point, a position of a reference point indicating a specific portion in the object, and the position of the feature point selected by said selection unit. 7. The apparatus according to claim 1, wherein said selection unit selects, from the plurality of feature points, a feature point so that a positional relationship between a reference point indicating a specific portion in an object, the designation point, and the feature point can satisfy a predetermined condition. 8. The apparatus according to claim 1, wherein said selection unit selects, based on comparison between a threshold and information determined using the position of the designation point and the positions of the plurality of feature points, a feature point from the plurality of feature points. 9. The apparatus according to claim 8, wherein said selection unit selects, from the plurality of feature points, two feature points so that an area of a plane formed by the designation point and the two feature points will be equal to or larger than a threshold. 10. The apparatus according to claim 8, wherein said selection unit selects, from the plurality of feature points, two feature points so that an angle formed by a line connecting the designation point and a first feature point and a line co... | Canon Kabushiki Kaisha | G06K009/00 A61B006/00 A61B008/00 G06T007/33 A61B008/08 A61B005/00 A61B005/055
US09779088 | B2 | 2017-10-03 | Translation station | 2016-02-12 | An interactive electronic translation and communications process is provided for use in a translation station that provides a mobile or stationary fixed interactive facility for interviews or interrogations to be carried out between two persons speaking in different languages. The process can be assisted by animated virtual characters (avatars) realistically created and displayed on a computer screen to represent ethnic looks from around the globe. The avatars can be lip synchronized to deliver messages to the interviewee in the interviewee's languages and can guide the users and interviewee through a series of questions and answers. Biometric conditions of the interviewee and electronic identification of the interviewee can also be readily accomplished by the novel process. The process is particularly useful for hospitals, law enforcement, military, airport security, transportation terminals, financial institutions, and government agencies. | 1. An interactive electronic communications process for use in a translation station, comprising: providing a computer system with a printer and dual displays operatively connected by hard wire or wireless to a central processing unit (CPU) for interaction between a host comprising an interviewer and a person comprising an interviewee, said CPU selected from the group consisting of a computer, laptop, desktop computer, portable computer, microprocessor, computer system, iPad, tablet computer, wireless computer, wired computer, netbook, electronic communications device, portable networking device, internet communications device, mobile phone, flip phone, camera phone, clamshell device, radio telephone, cellular phone, smart phone, tablet phone, portable media player (PMP), personal digital assistant (PDA), wireless e-mail device, handheld electronic device, mobile electronic device, network communications device, cloud network communications device, electronic subscriber device, and combinations of any of the preceding;electronically recognizing a person's language from a person's conversations selected from the group consisting of voice, speech, voice message, on-line conversations, off-line conversations, verbal communications, discussions, conferences, oral responses, replies, electronic communications, and sound;electronically comparing the person's language to a host language in the CPU;electronically translating the person's language into the host language in the CPU if the person's language and host language are not the same;visually displaying and electronically transmitting an electronic animated avatar from the CPU on one of the dual displays with lip synchronized speech in the interviewee's language, the avatar being modeled from human beings from the interviewee's country;electronically transmitting audible communications in the person's language from the CPU to the person, the audible communications selected from the group consisting of inquiries, electronically synchronized speech, questions, interviews, interrogations, sounds, speech, words, dialog, discussions, menu driven communications, electronics, interactive verbal communications, and combination thereof; andproviding an audit trail by printing said electronic translations and said electronic communications in both the host language and the person's language on the printer. 2. An interactive electronic communications process in accordance with claim 1 including: detecting biometric conditions of the person;the biometric conditions being selected from the group consisting of behavior patterns, temperature changes, pulse readings, eye scan, cornea scan, iris scan, retina scan, facial expressions, mouth expressions, photograph, voice, resonance, pitch, imagery, X-rays, scans, palm reading, nostril observations, finger print analysis, teeth identification, breathing patterns, gait, rhythm, speech, sweating, latency in typing, hesitation in providing information which should be kn... | - | G06T013/40 G06F017/28 G10L015/00 G10L015/25 G10L025/63 G06K009/00 G06K009/62 G06T013/00 G06Q010/10 A61B005/00
US09820654 | B2 | 2017-11-21 | Observation device | 2017-04-29 | An observation device includes a light source unit capable of adjusting a quantity of 390-490 nm illuminating light according to a B1 range and a B2 range having a longer wavelength than the B1 range, and an imaging element that images reflected light from a subject irradiated with illuminating light. The light source unit changes the light quantity ratio of the B1 range to the B2 range depending on whether white light imaging is conducted or narrow band imaging is conducted. The imaging element includes RGB color filters arranged according to pixels and the R color filter transmits more light in an R range and the B1 range than light in the second B range and has a spectral characteristic that satisfies a predetermined conditional formulae. | 1. An observation device comprising: a light source unit capable of adjusting a quantity of illuminating light in a B range that spans from a wavelength of 390 nm to 490 nm according to a first B range and a second B range having a longer wavelength than the first B range; andan imaging element that images reflected light from a subject irradiated with illuminating light emitted from the light source unit,wherein the light source unit changes a ratio of a light quantity in the first B range to a light quantity in the second B range depending on whether white light imaging is conducted or narrow band imaging is conducted, and satisfies conditional formula (4), andthe imaging element includes RGB color filters arranged according to pixels, and the R color filter transmits more light in an R range and the first B range than light in the second B range and has a spectral characteristic that satisfies conditional formulae (1) and (2): 0&lt;Srb2/Srr≦0.1 (1)0.5≦Sbb1/Srb1≦2 (2)Lnb2/Lnb1&lt;Lwb2/Lwb1 (4)where: Srr represents sensitivity of an R pixel to the R range,Srb1 represents sensitivity of the R pixel to the first B range,Srb2 represents sensitivity of the R pixel to the second B range,Sbb1 represents sensitivity of a B pixel to the first B range,Lnb1 represents a light quantity in the first B range of the light source unit during narrow band imaging,Lnb2 represents a light quantity in the second B range of the light source unit during narrow band imaging,Lwb1 represents a light quantity in the first B range of the light source unit during white light imaging, andLwb2 represents a light quantity in the second B range of the light source unit during white light imaging. 2. The observation device according to claim 1, wherein the imaging element satisfies conditional formula (3): Srb1/Srb2≧3 (3). 3. The observation device according to claim 1, wherein the first B range is a wavelength range of 390 nm or more and less than 440 nm, and the second B range is a wavelength range of 440 nm or more and less than 490 nm. 4. The observation device according to claim 1, wherein the color filters are arranged in a Bayer array. 5. The observation device according to claim 1, wherein the light source unit includes at least one solid-state illuminating element having a peak in the first B range and at least one solid-state illuminating element having a peak in the second B range. 6. The observation device according to claim 1, wherein the light source unit includes a solid-state illuminating element that has a peak at a wavelength of 580 nm to 650 nm. | OLYMPUS CORPORATION | A61B005/05 A61B005/00 A61B001/05 A61B001/07 H05B033/08 F21V008/00 H05B037/02 H04N005/225 G02B023/24 G02B005/20
US09808195 | B2 | 2017-11-07 | 2D scanning videokymography system for analyzing vibration of vocal-fold mucosa, and method of analyzing vibration of vocal-fold mucosa using the same | 2013-10-10 | Disclosed is a 2D scanning videokymography system for analyzing a vibration of vocal-fold mucosa, including: a laryngoscope for observing vocal folds; a light source for illuminating the vocal folds; a video camera for recording and storing images observed through the laryngoscope; a computer incorporating an image capture unit for converting a video signal transmitted from the video camera into a digital image signal, a storage unit for storing the digital image signal, a control unit for analyzing the image signal of the storage unit and displaying the analysis results on a monitor, and analysis software for analyzing the image signal of the storage unit; and a monitor for displaying a captured image and analysis results. | 1. A 2D scanning videokymography system for analyzing a vibration of vocal-fold mucosa, comprising: a laryngoscope for observing vocal folds;a light source for illuminating the vocal folds;a video camera for recording and storing images observed through the laryngoscope;a computer comprising an image capture unit for converting a video signal transmitted from the video camera into a digital image signal, a storage unit for storing the digital image signal, a control unit for analyzing the image signal of the storage unit and displaying the analysis results on a monitor, and analysis software for analyzing the image signal of the storage unit; anda monitor for displaying a captured image and analysis results,wherein the computer displays a normalized index including an average vocal-fold width for a longitudinal-axis length of a glottis, a glottis opening ratio which is a ratio of a vocal-fold opening period and a total period, an asymmetric index which is a difference in opening degree of both vocal folds, a basic frequency, a vibration strength, regularity of vibration, a mucosa wave, symmetricity of the vibration, an outer boundary shape, an inner boundary shape, an abnormal cycle, and vibration absence of the vocal-fold mucosa. 2. The system of claim 1, wherein the video camera scans the vocal-fold mucosa with a rolling shutter to acquire 2D scanning videokymography. 3. The system of claim 1, wherein a shutter speed of the video camera is set to 1/1000 seconds or more. 4. The system of claim 1, wherein the light source is a Xenon continuous light source. 5. An analysis method of a vibration of vocal-fold mucosa, which is performed in a 2D scanning videokymography system according to claim 1, the method comprising: (a) storing a 2D scanning kymogram image acquired by photographing the motion of the vocal folds by using a laryngoscope, a light source, and a video camera in a storage unit through an image capture unit;(b) configuring, by a control unit, a predetermined number of image frames stored in the storage unit as a still screen for each frame and displaying the configured frames on a monitor and configuring and displaying a menu screen so as to select a frame which a user intends to observe through a mouse or a keyboard; and(c) displaying, by analysis software, a normalized index for the vocal folds observed by analyzing an image. 6. The method of claim 5, wherein in step (a), an entire part of the vocal folds is photographed without vocalization and thereafter, “i” or “e” is vocalized and the motion of the entire part of a vocal-fold vibration part is photographed. | - | A61B006/00 A61B005/00 A61B001/04 A61B001/267 A61B001/00 A61B005/11 A61B001/06
US09801572 | B2 | 2017-10-31 | Transcutaneous analyte sensor | 2015-06-18 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A system for measuring an analyte in a host, comprising: a sensor electronics housing;an on-skin assembly configured to releasably attach with the sensor electronics housing; anda sensor comprising at least one electrode, the sensor comprising: an in vivo portion adapted for insertion through an exit port of the sensor electronics housing and through an exit site of the host's skin; andan ex vivo portion adapted to remain above the host's skin after sensor insertion and to operably connect to the sensor electronics housing, wherein the ex vivo portion is configured to be bent by attachment of the sensor electronics housing to the on-skin assembly. 2. The system of claim 1, wherein the ex vivo portion is configured to be bent relative to the in vivo portion following attachment of the sensor electronics housing to the on-skin assembly. 3. The system of claim 1, wherein the on skin assembly is adapted for placement over an exit site of the host's skin. 4. The system of claim 1, wherein the exit port is at a location substantially closer to the center than edges of the on skin assembly. 5. The system of claim 1, wherein the exit port comprises a layer of sealing material to reduce or eliminate the migration of bacteria or other contaminants between the exit-site and the housing. 6. The system of claim 4, wherein the layer of sealing material comprises a septum seal. 7. The system of claim 4, wherein the layer of sealing material comprises a silicone seal. 8. The system of claim 1, further comprising an applicator configured to insert the sensor through the exit port and the exit site. 9. The system of claim 1, wherein the sensor comprises a plurality of electrodes. 10. The system of claim 1, wherein the sensor electronics comprise a data sampling portion in data communication with the at least one electrode, wherein the data sampling portion comprises an analog portion configured to measure the current flowing at the at least one electrode and to convert the current measurement to digital values representative of the current, wherein the digital values are acquired at a predetermined time interval. 11. The system of claim 10, wherein the predetermined time interval is adaptively determined based on another parameter of the system. 12. The system of claim 10, wherein the sensor electronics comprise an integrated circuit comprising the data sampling portion. 13. The system of claim 1, wherein the sensor electronics comprise a transmitter portion for transmitting a signal indicating a level of the analyte in the host at a programmable transmission interval. 14. The system of claim 13, wherein the programmable transmission interval is adaptively determined based on another parameter of the system. 15. The system of claim 13, wherein the sensor electronics comprise an integrated circuit comprising the transmitter portion. 16. A system for measuring an analyte in a host, comprising: a sensor electronics housing;an on-skin assembly config... | DexCom, Inc. | A61B005/1473 A61B005/1486 A61B005/1495 A61B005/05 A61M005/14 A61B005/145 A61B005/00 A61B005/15 A61B017/34 A61M005/172 A61M005/158 A61M005/142
US09789252 | B2 | 2017-10-17 | Therapy program modification based on a therapy field model | 2009-01-22 | Techniques for modeling therapy fields for therapy delivered by medical devices are described. Each therapy field model is based on a set of therapy parameters and represents where therapy will propagate from the therapy system delivering therapy according to the set of therapy parameters. Therapy field models may be useful in guiding the modification of therapy parameters. As one example, a processor compares an algorithmic model of a therapy field to a reference therapy field and adjusts at least one therapy parameter based on the comparison. As another example, a processor adjusts at least one therapy parameter to increase an operating efficiency of the therapy system while substantially maintaining the modeled therapy field. | 1. A method comprising: determining, by a processor, a first therapy program that comprises a set of therapy parameters values;generating, by the processor, an algorithmic model of a therapy field based on the first therapy program, the algorithmic model representing where therapy will propagate from a therapy system delivering therapy according to the first therapy program; andautomatically determining, by the processor, a second therapy program that increases an operating efficiency of the therapy system while substantially maintaining the therapy field. 2. The method of claim 1, further comprising controlling, by the processor, the therapy system to deliver therapy to a patient according to the second therapy program. 3. The method of claim 1, wherein the algorithmic model of the therapy field comprises a first algorithmic model of a first therapy field, the method further comprising generating, by the processor, a second algorithmic model of a second therapy field representing where therapy will propagate from the therapy system based on the second therapy program. 4. The method of claim 3, further comprising presenting the first algorithmic model of the first therapy field and the second algorithmic model of the second therapy field to a user. 5. The method of claim 1, wherein automatically determining the second therapy program comprises automatically determining the second therapy program based on at least one of an amplitude-duration curve, a dose-response curve, a strength-duration curve, or a three-dimensional (3D) depolarization field model. 6. The method of claim 5, wherein the at least one amplitude-duration curve, dose-response curve, strength-duration curve, or 3D depolarization field model is configured based on at least one of a patient condition, a desired therapeutic outcome, or a tissue characteristic of a target delivery site. 7. The method of claim 1, wherein the second therapy program increases the operating efficiency of the therapy system by at least one of decreasing power consumption or operating at an efficient amplitude determined based on a voltage multiplier level. 8. The method of claim 1, wherein the algorithmic model of the therapy field comprises a first algorithmic model of a first therapy field, the method further comprising generating a second algorithmic model of a second therapy field based on the second therapy program, wherein substantially maintaining the therapy field comprises maintaining a difference between the first therapy field and the second therapy field below a threshold value, wherein the threshold value defines a maximum allowable change in at least one field characteristic. 9. The method of claim 8, wherein the at least one field characteristic comprises at least one of a centroid, a volume, an amplitude, a width at a defined amplitude, or a charge density. 10. The method of claim 8, wherein the threshold value comprises a weighted threshold of a plurality of field characteristics. ... | Medtronic, Inc. | A61B005/00 A61M005/172 A61N001/372 G06F019/00 A61M005/142 A61N001/36
US09788753 | B2 | 2017-10-17 | Method and system for characterizing cortical structures | 2010-02-25 | A method of identifying brain structures is disclosed. The method comprises imaging the cortex of the brain via inversion recovery magnetic resonance imaging, such that at least two areas are zeroed in terms of their longitudinal relaxation times, and associating the two areas with different brain structures, thereby identifying brain structures. | 1. A method of identifying brain structures, comprising: using magnetic resonance imaging system for imaging the cortex of the brain via inversion recovery magnetic resonance imaging, said imaging comprises applying to the brain energy that is resonantly absorbed in the brain, said energy being applied as a resonance pulse sequence selected such that, for each of at least two areas or regions in the cortex, a longitudinal magnetization of the respective area or region is suppressed compared to a longitudinal magnetization of all other areas or regions, wherein each of these areas or regions is characterized by a different characteristic recovery time;using an image processor for classifying said at least two areas or regions to a laminar pattern based on characteristic recovery time and associating said at least two areas or regions with different brain structures, thereby identifying brain structures; andgraphically displaying said identified laminar patterns on a display. 2. The method of claim 1, wherein said analysis comprises a multi-exponential signal decay fit. 3. The method of claim 2, wherein said fit is performed along a vector perpendicular to the surface of the brain from a white matter region and to a cerebrospinal fluid region. 4. The method of claim 2, wherein said analysis comprises three-dimensional analysis. 5. The method of claim 1, wherein said magnetic resonance imaging is such that at least one of a white matter zone and a cerebrospinal fluid zone is zeroed in terms of its longitudinal relaxation time. 6. The method of claim 1, wherein said at least two areas comprises at least five laminar cortical layers, each corresponding to a different characteristic longitudinal relaxation time. 7. The method of claim 1, further comprising identifying at least one Brodmann region among said at least two areas. 8. The method of claim 1, wherein said magnetic resonance imaging is at a resolution characterized by a unit voxel size of at least 1 cubic millimeter. 9. The method of claim 1, further comprising combining diffusion weighted magnetic resonance imaging with said inversion recovery magnetic resonance imaging. 10. The method of claim 1, further comprising combining functional magnetic resonance imaging with said inversion recovery magnetic resonance imaging. 11. The method of claim 1, further comprising identifying an abnormal laminar pattern in the cortex. 12. The method of claim 11, further comprising associating said abnormal laminar pattern with a brain disease. 13. The method of claim 11, further comprising treating said brain disease. 14. The method of claim 1, further comprising repeating said imaging and said associating at different times for the same cortex so as to identify changes in a laminar pattern of the cortex. 15. The method of claim 14, further comprising assessing brain plasticity based on said changes. 16. The method of claim 14, further comprising assessing changes in brain function based on sai... | Ramot at Tel-Aviv University Ltd. | G06K009/00 A61B005/055 G06T007/00 A61B005/00
US09782095 | B2 | 2017-10-10 | Single-use biomedical sensors | 2011-12-16 | A disposable self-powered biomedical sensor comprises a printed wet electrode on a substrate sheet. The wet electrode is provided with an electrolyte element to enhance the electrical contact with a surface to be measured. Moreover, a printed battery encapsulated in a hermetically sealed compartment is provided on the substrate sheet. The disposable self-powered sensor can be stored within an enclosure or a package which provides a proper atmosphere to prevent the drying of the electrolyte and prolong the shelf life of the sensors. | 1. A biomedical sensor, comprising: a substrate sheet;at least one electrode printed directly on a first surface of the substrate sheet, the at least one electrode being provided with an electrolyte element to enhance electrical contact with a surface to be measured;a hermetically sealed compartment disposed on a second surface of the substrate sheet;at least one printable soft battery disposed within the hermetically sealed compartment, the hermetically sealed compartment comprising: a first humidity-proof material layer between the second surface of the substrate sheet and the battery; anda second humidity-proof material layer on a side of the battery opposite the second surface of the substrate sheet;the biomedical sensor further comprising a venting opening extending through the second humidity proof material layer of the hermetically sealed compartment; anda removable humidity resistant patch over the venting opening. 2. A biomedical sensor according to claim 1, wherein one of the first humidity-proof material layer and the second humidity-proof material layer comprises the substrate sheet itself made of humidity-proof material or chemically treated to be humidity-proof. 3. A biomedical sensor according to claim 1, wherein one of the first humidity-proof material layer and second humidity-proof material layer comprises a humidity-proof material layer covering a top of the printed soft battery. 4. A biomedical sensor according to claim 1 wherein at least one of the first humidity-proof material layer and the second humidity-proof material layer comprises a metal layer or a dielectric layer. 5. A biomedical sensor according to claim 1, wherein the soft battery is manufactured in the same printing process with the remaining components of the electrode and covered by printing or laminating or other relevant process with a humidity-proof layer. 6. A biomedical sensor according to claim 1, wherein the soft battery is manufactured separately, attached on top of an electrode substrate sheet, and covered by printing or laminating or other relevant process with a humidity-proof layer. 7. A biomedical sensor according to claim 1, wherein the soft battery is manufactured separately, encapsulated hermetically with a humidity resistant material, and the hermetically encapsulated soft battery is attached on top of an electrode substrate sheet. 8. A biomedical sensor according to claim 1, wherein the hermetically sealed compartment comprises one or more of the following characteristics: the hermetically sealed compartment is capable of accommodating to a pressure of gases created inside the soft battery during one or more of storage or use of the sensor; orthe hermetically sealed compartment comprises a gas permeable area closed by the removable closure member during storage, the removable closure member being opened when use of the sensor is started so as to vent out gases created inside the soft battery through the gas permeable area. 9. A biomedi... | GENERAL ELECTRIC COMPANY | A61B005/04 A61B005/0408 A61B005/00 A61B005/0478 A61B005/0492
US09814433 | B2 | 2017-11-14 | Creating a vascular tree model | 2014-01-15 | A method for vascular modeling is disclosed. The method, in some embodiments, comprises receiving a plurality of 2-D angiographic images of a portion of a vasculature of a subject, and processing the images to automatically detect 2-D features, for example, paths along vascular extents, which are projected into 3-D to determine homologous features among blood vessels. In some embodiments, projection and/or image registration is iteratively altered to improve feature position matching. Based on 3-D vascular extents and their registration to 2-D images, additional features such as vascular width are optionally determined and added to the model. | 1. A method for construction of a vascular tree model comprising: receiving by image processing circuitry a plurality of 2-D angiographic images showing blood vessel segments of a portion of a vasculature of a subject;extracting automatically by the image processing circuitry, from each of said plurality of 2-D angiographic images, an image feature set comprising 2-D feature positions of said blood vessel segments; andidentifying homologous vascular branches among the 2-D angiographic images using said blood vessel segments, wherein the identifying comprises: adjusting automatically by the image processing circuitry said 2-D feature positions to reduce relative position error in a common 3-D coordinate system to which each said image feature set is back-projectable, andthen associating automatically by the image processing circuitry image features of the image feature sets as representing homologous vascular branches, based on their projection from the planes of the 2-D angiographic images to common blood vessel segment regions in said common 3-D coordinate system. 2. The method of claim 1, wherein the image feature set which is extracted comprises a centerline data set including 2-D centerline positions ordered along said blood vessel segments. 3. The method of claim 2, wherein said adjusting and associating are performed with elements of said centerline data set. 4. The method of claim 2, wherein said adjusting comprises registration of said 2-D images in 3-D space according to parameters which bring said 2-D centerline positions into closer correspondence among their 3-D back-projections. 5. The method of claim 2, comprising estimating a metric of radial vascular width based on values of at least one of said plurality of 2-D angiographic images along lines perpendicular to said ordered 2-D centerline positions. 6. The method of claim 2, wherein said adjusting comprises adjusting of said 2-D feature positions based on projection of said common 3-D coordinate system into the 2-D plane of at least one of said plurality of 2-D angiographic images. 7. The method of claim 1, wherein information required for said associating automatically of image features is entirely provided before review of said images by a human operator. 8. The method of claim 1, further comprising defining a surface corresponding to a shape of the heart of the subject, and using said surface as a constraint for said associating of said image features. 9. The method of claim 1, wherein said images are acquired upon injection of a contrast agent to said vasculature, and further comprising: determining temporal characteristics of the movement of said contrast agent through said vasculature; andconstraining said 2-D feature positions based on said temporal characteristics. 10. The method of claim 1, wherein said portion of said vasculature comprises coronary arteries. 11. The method of claim 1, further comprising: inspecting projections from the blood vessel segment regio... | CathWorks LTD. | G06T017/00 A61B006/00 G06F019/00 G06T011/00 G06T007/00 A61M005/00 A61B005/02 A61B005/026 G06T007/55 A61B006/02 A61B006/03 A61B008/06 A61B005/021 A61B005/00
US09801958 | B2 | 2017-10-31 | Polymer nanoparticle composite and composition for MRI imaging including same | 2014-08-22 | The present invention provides a safe polymer nanoparticle composite with few side effects, and an MRI contrast agent incorporating said polymer nanoparticle composite. The polymer nanoparticle composite is capable of specifically accumulating on a tumor tissue to selectively extract the tissue, exhibiting high contrast even when used in small amounts, and enabling imaging over prolonged periods of time. This polymer nanoparticle composite is characterized by containing a block copolymer that includes a non-charged hydrophilic polymer chain segment and an anionic polymer chain segment, and MnCaP. | 1. A polymer nanoparticle composite comprising: a block copolymer comprising a non-charged hydrophilic polymer chain segment and an anionic polymer chain segment; and MnCaP; wherein the anionic polymer chain segment is a polypeptide having an anionic group on a side chain of the polypeptide; andwherein the non-charged hydrophilic polymer chain segment is selected from the group consisting of polyethylene glycol, poly(2-methyl-2-oxazolineT poly(2-ethyl-2-oxazolineT poly(2-isoprot>yl-2-oxazoline), polyacrylamide, polymethacrylamide, polyvinyl alcohol, polvthvdroxyethyl acrylate) and polv(hvdroxy ethyl methacrylate). 2. The composite according to claim 1, wherein the anionic polymer chain segment is derived from an anionic polymer selected from the group consisting of poly(glutamic acid) and poly(aspartic acid). 3. The polymer nanoparticle composite according to claim 1, which has a configuration in which MnCaP is incorporated into a nanoparticulate particle that is formed with the block copolymer having the non-charged hydrophilic polymer chain segment as a shell part and the anionic polymer chain segment as a core part. 4. The polymer nanoparticle composite according to claim 1, wherein the block copolymer is represented by General Formula (1-a) or (2-a) below:wherein R1 represents a hydrogen atom or an unsubstituted or substituted, linear or branched C1-12 alkyl group, L1 and L2 represent a linking group, R2 each independently represents a methylene group or an ethylene group, R3 each independently represents a hydrogen atom, a protective group of an amino group, a hydrophobic group or a polymerizable group, R4 represents a hydroxyl group or an initiator residue, R5 each independently represents a hydrogen atom or an alkali metal ion, m represents an integer of 5-20,000, and n represents an integer of 2-5,000. 5. The polymer nanoparticle composite according to claim 1, whose average dispersed particle diameter in an aqueous medium is 30 nm-150 nm as measured by a dynamic light scattering method. 6. The polymer nanoparticle composite according to claim 1, wherein the polymer nanoparticle composite releases an Mn2+ion. 7. An MRI contrast composition comprising the polymer nanoparticle composite according to claim 1. 8. The composition according to claim 7, which is used for detecting a tumor. 9. The composition according to claim 8, wherein the tumor is a primary tumor or a metastatic tumor. 10. The composition according to any one of claims 7-9, which is used for assessing grade of malignancy of a tumor. 11. An MRI contrast method comprising a step of administering the polymer nanoparticle composite according to claim 1 to a body of a test animal. 12. The method according to claim 11, which is used for detecting a tumor. 13. The method according to claim 12, wherein the tumor is a primary tumor or a metastatic tumor. 14. The method according to any one of claims 11-13, which is used for assessing grade of malignancy of a tumor. 15. A... | THE UNIVERSITY OF TOKYO | A61B005/055 A61K049/18 A61K049/12 C08K003/32 A61K009/14
US09782113 | B2 | 2017-10-10 | Lancing device | 2014-02-27 | A lancing device includes: a cover section, a part of which is provided with an opening; a main body held on the cover section and configured to be movable in a predetermined direction from the opening relative to the cover section; a lancing unit which is launched in the predetermined direction; and an operating section configured to launch the lancing unit when the operating section is pushed by applying a force in the predetermined direction, wherein the lancing unit is launched when the force is applied to the operating section in the predetermined direction with regulating a movement of the main body in the predetermined direction, and the main body moves, instead of pushing the operating section, by applying the force in the predetermined direction without regulating the movement of the main body in the predetermined direction is disclosed. | 1. A lancing device for puncturing a skin of a subject, comprising: a tubular cover section having a first opening at one end and a second opening at the other end;a tubular main body having one end and the other end opposite to the one end and arranged to directly slide into and out of the first opening of the cover section, such that the one end of the main body is arranged inside of the cover section, the other end of the main body is arranged outside of the cover section, and the main body is configured to be movable in a predetermined direction relative to the cover section;a lancing unit housed in the main body and launched in the predetermined direction from the other end of the main body;an operating section installed to the one end of the main body and configured to be pushed by applying a force in the predetermined direction through the second opening of the cover section;a biasing unit installed in the main body and configured to bias the lancing unit in the predetermined direction; anda locking unit installed in the main body and configured to lock a bias of the biasing unit and to release a lock of the biasing unit;wherein the locking unit releases the lock, the biasing unit biases the lancing unit, the lancing unit is launched and the main body remains at a first position relative to the cover section when the subject presses the other end of the tubular main body against the skin of the subject and applies a force to the operating section in the predetermined direction through the second opening, andwherein when the other end of the tubular main body is not pressed against of the skin of the subject, the main body is configured to move from the first position to a second position relative to the cover section without releasing the lock of the locking unit and the lancing unit is not launched in response to the force is applied to the operating section in the predetermined direction through the second opening. 2. The lancing device of claim 1, wherein the biasing unit is an elastic body. 3. The lancing device of claim 1, wherein the main body is configured to return the first position relative to the cover section after the main body has moved from the first position to the second position relative to the cover section to cause the biasing unit to transition from a non-energized state to an energized state and cause the locking unit to lock the biasing unit and maintain the lock. 4. The lancing device of claim 1, wherein the main body comprises a guide part for guiding the movement in the predetermined direction relative to the cover section between the first position and the second position. 5. The lancing device of claim 1, wherein the cover section is configured to be held by a subject; andthe operating section is configured to be operated by a hand of the subject which holds the cover section. 6. The lancing device of claim 1, wherein the lancing unit is a fine needle tip having a plurality of needles. 7. The lancing devic... | SYSMEX CORPORATION | A61B005/151 A61B005/15 A61B005/145
US09829546 | B2 | 2017-11-28 | Low-temperature cooling apparatus and superconducting quantum interference device sensor module | 2015-04-21 | Provided are a low-temperature cooling apparatus and a superconducting quantum interference device (SQUID) sensor module. The low-temperature cooling apparatus includes an outer container; an inner container disposed inside the outer container, the inner container including a neck portion having a first diameter and a body portion having a second diameter greater than the first diameter; an insert inserted into the neck portion of the inner container; and a plurality of SQUID sensor modules inserted into the body portion of the inner container. Each of the SQUID sensor modules is in the form of a fan-shaped pillar and is fixedly coupled with an inner bottom plate of the inner container. | 1. A low-temperature cooling apparatus comprising: an outer container;an inner container disposed inside the outer container, the inner container including a neck portion having a first diameter and a body portion having a second diameter greater than the first diameter;an insert inserted into the neck portion of the inner container; anda plurality of SQUID sensor modules inserted into the body portion of the inner containerwherein each of the plurality of SQUID sensor modules is in the form of a circular-sector pillar and is fixedly coupled with an inner bottom plate of the inner container;wherein each of the plurality of SQUID sensor modules comprises: an upper sensor-coupling plate having a through-hole;a lower sensor-coupling plate including a protrusion disposed to be aligned with the through-hole;an upper printed circuit board disposed below the upper sensor-coupling plate, the upper printed circuit board having a printed circuit board through-hole aligned with the through-hole; anda SQUID sensor;wherein one end of each SQUID sensor is inserted into the through-hole and the printed circuit board through-hole; andwherein another end of each SQUID sensor has a groove that is disposed on the protrusion of the lower sensor-coupling plate. 2. The low-temperature cooling apparatus as set forth in claim 1, wherein the inner bottom plate includes a partition such that the SQUID sensor module is inserted and aligned with the inner bottom plate. 3. The low-temperature cooling apparatus as set forth in claim 1, wherein the insert further comprises an intermediate connection block, and wherein an interconnection of the SQUID sensor module is electrically connected to the intermediate connection block. 4. The low-temperature cooling apparatus as set forth in claim 1, wherein the upper sensor coupling plate further has an auxiliary through-hole. 5. The low-temperature cooling apparatus as set forth in claim 1, wherein the upper sensor coupling plate further has a trench formed on its bottom surface. 6. The low-temperature cooling apparatus as set forth in claim 1, wherein the SQUID sensor module further comprises a module coupling pillar coupled with the inner bottom plate through a fan-shaped central region. 7. The low-temperature cooling apparatus as set forth in claim 1, wherein the SQUID sensor comprises: a bobbin on which a pick-up coil is wound;a SQUID mounted on the bobbin and electrically connected to the pick-up coil; anda printed circuit board electrically connected to the SQUID and fixed to the bobbin. 8. The low-temperature cooling apparatus as set forth in claim 7, wherein the bobbin comprises: a cylindrical body portion;a cut-out portion connected to the body portion and cut out such that its side surface forms a plane;a square-pillar portion connected to the cut-out portion, the square-pillar portion having a square section; anda cylindrical portion extending from the center of the square-pillar portion,wherein the groove is formed ... | Korea Research Institute of Standards and Science | H05K007/02 G01R033/035 A61B005/05 G01R033/00 A61B005/04
US09789337 | B2 | 2017-10-17 | Combined imaging modalities for radiation treatment planning | 2011-10-07 | A system includes acquisition of a first three-dimensional image of a patient volume using a magnetic resonance imaging scanner, acquisition of a second three-dimensional image of the patient volume using cone beam radiation emitted by the linear accelerator, and generation of a radiation treatment plan based on the first image and the second image. | 1. A method comprising: acquiring a first three-dimensional image of a patient volume using a magnetic resonance imaging scanner;recording a treatment isocenter of the patient volume based on the first three-dimensional image;registering the treatment isocenter of the patient volume with an isocenter of a linear accelerator;acquiring a second three-dimensional image of the patient volume using cone beam radiation emitted by the linear accelerator while the treatment isocenter of the patient volume is registered with the isocenter of the linear accelerator;defining a plurality of sub-regions within the patient volume based on the first three-dimensional image;determining a radiation dose associated with each of the plurality of sub-regions based on the second three-dimensional image, the determining of the radiation dose associated with each of the plurality of sub-regions including: overlaying contours illustrative of the plurality of sub-regions from the first three-dimensional image of the patient volume onto the second three-dimensional image of the patient volume, wherein the contours are aligned with the second three-dimensional image based on the registration of the treatment isocenter of the patient volume with the isocenter of the linear accelerator;calculating the radiation dose associated with each of the plurality of sub-regions using electron density information of the second three-dimensional image; andgenerating a radiation treatment plan based on an overlay of a representation of the calculated radiation dose for each of the plurality of sub-regions onto the first image. 2. The method according to claim 1, further comprising: acquiring a third three-dimensional image using a positron emission tomography scanner,wherein generating the radiation treatment plan based on the first image and the second image comprises generating the radiation treatment plan based on the first image, the second image and the third image. 3. The method according to claim 1, wherein the cone beam radiation emitted by the linear accelerator comprises megavoltage radiation. 4. The method according to claim 1, wherein the cone beam radiation emitted by the linear accelerator comprises kilovoltage radiation. 5. The method according to claim 1, wherein acquiring the second three-dimensional image of the patient volume using cone beam radiation emitted by the linear accelerator comprises: acquiring at least one projection image by positioning an imaging panel of the linear accelerator so that a central ray of cone beam radiation emitted from the linear accelerator does not intersect the center of the imaging panel. 6. A non-transitory computer-readable medium storing processor-executable program code executable to: acquire a first three-dimensional image of a patient volume using a magnetic resonance imaging scanner;record a treatment isocenter of the patient volume based on the first three-dimensional image;register the treatment isocenter of the patient v... | Siemens Medical Solutions USA, Inc. | A61B005/055 A61N005/10
US09802056 | B2 | 2017-10-31 | Accurate cardiac event detection in an implantable cardiac stimulus device | 2013-10-15 | Methods, systems, and devices for signal analysis in an implanted cardiac monitoring and treatment device such as an implantable cardioverter defibrillator. In some illustrative examples, detected events are analyzed to identify changes in detected event amplitudes. When detected event amplitudes are dissimilar from one another, a first set of detection parameters may be invoked, and, when detected event amplitudes are similar to one another, a second set of detection parameters may be invoked. Additional methods determine whether the calculated heart rate is “high” or “low,” and then may select a third set of detection parameters for use when the calculated heart rate is high. | 1. A method of event detection in an implantable cardiac stimulus device having a plurality of electrodes adapted for implantation in a patient, the method comprising: capturing electrical signals from the electrodes;comparing captured electrical signals to a detection threshold that varies with time following a detection profile;when the captured electrical signals cross the detection threshold, declaring a detected event;determining whether either of these conditions is met: a) the detected event is similar to a prior detected event, orb) whether the detected event is separated from an immediately prior detected event by an interval exceeding an interval threshold;selecting from between first and second variants of a selected detection profile for use in subsequent event detection as follows: if b) is met, then selecting the first variant;if b) is not met and a) is met, then selecting the first variant; and else selecting the second variant; anddetecting a subsequent cardiac event using the selected variant of the selected detection profile. 2. The method of claim 1 wherein the selected detection profile is selected by selecting from among a plurality of detection profiles each corresponding to one of: a calculated beat rate range for a patient; anda post-shock state of the device. 3. The method of claim 1 wherein each of the first and second variants are dependent upon an amplitude measure of the prior detected event. 4. The method of claim 3 wherein the first and second variants each include a period in which the detection profile is flat for a fixed period and is calculated as a fixed percentage of the amplitude measure of the prior detected event such that: the first variant has a first fixed period and a first fixed percentage, andthe second variant has a second fixed period that is longer than the first fixed period, and a second fixed amplitude that is higher than the first fixed amplitude. 5. An implantable cardiac stimulus device having a plurality of electrodes adapted for implantation in a patient and operational circuitry coupled to the plurality of electrodes which is configured to perform a method comprising: capturing electrical signals from the electrodes;comparing captured electrical signals to a detection threshold that varies with time following a detection profile;when the captured electrical signals cross the detection threshold, declaring a detected event;determining whether either of these conditions is met: a) the detected event is similar to a prior detected event, orb) whether the detected event is separated from an immediately prior detected event by an interval exceeding an interval threshold;selecting from between first and second variants of a selected detection profile for use in subsequent event detection as follows: if b) is met, then selecting the first variant;if b) is not met and a) is met, then selecting the first variant; and else selecting the second variant; anddetecting a subsequent cardiac event usi... | Cameron Health, Inc. | A61N001/39 A61B005/00 A61B005/0245 A61N001/37 A61N001/365 A61B005/0452
US09782211 | B2 | 2017-10-10 | Preferential volume reduction of diseased segments of a heterogeneous lobe | 2014-10-01 | Methods and apparatus for diagnosing and treating disorders of the lung are provided, which may include any number of features. In one embodiment, a method comprises obtaining diagnostic information relating to a patient's lungs, compiling a list of potential treatment plans for lung volume reduction in the first and second lungs, excluding treatment plans from the list of potential treatment plans that propose treatment of a lung segment that falls within a segment exclusion rule, and identifying at least one preferred treatment plan from the list of potential treatment plans that targets sufficiently diseased lung segments while also targeting a preferred combined volume of the first and second lungs. | 1. A computer-implemented method of treatment planning for assisting a physician perform bronchoscopic lung volume reduction, comprising: obtaining diagnostic information including information derived from images relating to a patient's first and second lungs;compiling, automatically on a processor or electronic controller, a plurality of potential lung segment combinations for lung volume reduction in the first and second lungs wherein each potential lung segment combination includes at least one lung segment from the first lung and at least one lung segment from the second lung;computing, automatically on the processor or electronic controller, a plurality of diagnostic values for each lung segment of said potential lung segment combinations;determining a segment exclusion rule operable to exclude one or more of the potential lung segment combinations based on comparing the diagnostic value from the computing step to a threshold value or range;excluding, automatically on the processor or electronic controller, at least one of the potential lung segment combinations based on said segment exclusion rule; andidentifying, automatically on the processor or electronic controller, at least one preferred segment combination not excluded from the potential lung segment combinations for the bronchoscopic lung volume reduction, wherein the preferred segment combination targets diseased lung segments while also targeting a preferred combined volume of the first and second lungs based on the diagnostic values from the computing step. 2. The method of claim 1 wherein the diagnostic values are selected from the group consisting of a tissue-to-air ratio, a heterogeneity index, an air volume, a mass, and a percentage of a lobe of each segment of the first and second lungs. 3. The method of claim 1 wherein compiling comprises compiling a list of potential treatments that include a first treatment that targets a first segment of a first lobe of the first lung, and a second treatment that targets a second segment of a second lobe of the second lung. 4. The method of claim 3 wherein the second treatment also targets a third segment of the second lobe of the second lung. 5. The method of claim 1 wherein the segment exclusion rule requires that lung segment combinations including lung segments that have a heterogeneity index under 1.2 be excluded. 6. The method of claim 1 wherein the segment exclusion rule requires that lung segment combinations including lung segments that have a mass less than 13 g be excluded. 7. The method of claim 1 wherein the segment exclusion rule requires that lung segment combinations including segments having a tissue mass in grams plus air volume in milliliters greater than 1700 be excluded. 8. The method of claim 1 wherein the segment exclusion rule requires that lung segment combinations segments that have a mass greater than 48 g be excluded. 9. The method of claim 1 wherein the segment exclusion rule requires that lung segment... | UPTAKE MEDICAL TECHNOLOGY INC. | A61B005/08 A61M037/00 A61B018/04 A61B034/10 A61B018/00
US09782114 | B2 | 2017-10-10 | Devices and methods for body fluid sampling and analysis | 2012-08-03 | Described here are meters and methods for sampling, transporting, and/or analyzing a fluid sample. The meters may include a meter housing and a cartridge. In some instances, the meter may include a tower which may engage one or more portions of a cartridge. The meter housing may include an imaging system, which may or may not be included in the tower. The cartridge may include one or more sampling arrangements, which may be configured to collect a fluid sample from a sampling site. A sampling arrangement may include a skin-penetration member, a hub, and a quantification member. | 1. An integrated body fluid sampling meter comprising: a meter housing comprising a contact pad;a tower held inside the meter housing, wherein the tower is holding a disposable cartridge, the disposable cartridge comprising a plurality of sampling arrangements;a force sensor coupled to the tower; anda processor operably coupled to the force sensor, wherein the processor receives data from the force sensor and initiates a testing procedure when a predetermined force is applied to the disposable cartridge via the contact pad. 2. The meter of claim 1 wherein the tower is fixed relative to the meter housing. 3. The meter of claim 2 wherein the force sensor comprises a strain gauge. 4. The meter of claim 3 wherein the meter is configured such that the application of the predetermined force to the disposable cartridge causes a deflection in the tower, and wherein the strain gauge measures the deflection in the tower. 5. The meter of claim 1 further comprising a detector held within the tower. 6. The meter of claim 5 further comprising an optical system, wherein the disposable cartridge comprises a plurality of viewing windows, and wherein the optical system optically connects the detector to one of the plurality of viewing windows. 7. The meter of claim 1 wherein the tower is configured to fit at least partially within a recess of the disposable cartridge. 8. The meter of claim 7 further comprising a rotation element for rotating the disposable cartridge around the tower. 9. The meter of claim 1 further comprising a light generating assembly held within the tower. 10. The meter of claim 9 wherein the light generating assembly is configured to produce light at a plurality of wavelength ranges. 11. The meter of claim 9 wherein the light generating assembly comprises a plurality of light sources. 12. The meter of claim 11 wherein each light source is configured to generate light in a different predetermined wavelength range. 13. The meter of claim 1 wherein the predetermined force comprises a force within a predetermined range. | Intuity Medical, Inc. | A61B005/1455 A61B005/157 A61B005/00 A61B005/151 A61B005/145 G01N015/06 G01N021/03 G01N033/487 G01N021/78 A61B005/15
US09787965 | B2 | 2017-10-10 | Camera system, color conversion device and method employed thereupon, and recording medium for color conversion program | 2015-05-08 | A camera system includes: a database which stores a plurality of stereoscopic color profiles, in which conversion relationships calculated from second image data obtained by photographing a plurality of reference color stereoscopic objects assigned with reference colorimetric values in advance and the reference colorimetric values corresponding to the second image data are associated with a plurality of illumination conditions in photographing; a selection unit which, based on an illumination condition at the time of photographing of a stereoscopic subject, selects a stereoscopic color profile corresponding to the illumination condition; and a color conversion unit which performs color conversion from first image data of a photographed image of the stereoscopic subject to colorimetric values, based on the selected stereoscopic color profile. | 1. A camera system comprising: a photographing unit which photographs a stereoscopic subject to acquire first image data;a database which stores a plurality of stereoscopic color profiles, in which a plurality of conversion relationships calculated from second image data obtained by photographing a plurality of reference color stereoscopic objects assigned with reference colorimetric values in advance by the photographing unit and the reference colorimetric values corresponding to the second image data are associated with a plurality of illumination conditions including illumination geometric conditions in the photographing of the reference color stereoscopic objects;a selection unit which, based on an illumination condition at the time of photographing of the stereoscopic subject, selects a stereoscopic color profile corresponding to the illumination condition from among the plurality of stereoscopic color profiles stored in the database; anda color conversion unit which performs color conversion from the first image data of a photographed image of the stereoscopic subject photographed by the photographing unit to colorimetric values, based on the stereoscopic color profile selected by the selection unit. 2. The camera system according to claim 1, wherein the plurality of reference color stereoscopic objects are reference color stereoscopic objects of three or more colors, andthe plurality of stereoscopic color profiles are respectively created according to the plurality of illumination conditions for the respective reference color stereoscopic objects of three or more colors. 3. The camera system according to claim 1, wherein the reference color stereoscopic objects have at least one of a curved surface, a stereoscopic shape, an uneven surface shape, and a layer structure which are identical or analogous to those of the stereoscopic subject, and a polyhedron composed of a plurality of flat surfaces having different inclinations or a curved surface, and are stereoscopic objects made of a material analogous to the stereoscopic subject. 4. The camera system according to claim 1, wherein the reference color stereoscopic objects include reference color stereoscopic objects having at least one of semi-transparency, non-transparency, a light scattering property, and a color tint of milky white and/or flesh color which are analogous to those of the stereoscopic subject. 5. The camera system according to claim 1, wherein the photographing unit includes a digital camera, andan illumination geometric condition includes geometric positional information of three of a light source, the digital camera, and the stereoscopic subject or the reference color stereoscopic objects. 6. The camera system according to claim 5, wherein each of the reference color stereoscopic objects is a polyhedron composed of a plurality of flat surfaces having different inclinations or a curved surface, andarrangement angle information of the light source is analyzed based on a ... | FUJIFILM Corporation | H04N013/02 H04N013/00 H04N001/60 G01J003/50 H04N009/73 G01J003/52 A61C019/04 G06T011/00 A61B005/103 H04N009/64 G06T007/90 A61B005/00
US09829432 | B2 | 2017-11-28 | Gas measuring apparatus | 2016-09-09 | A gas measuring apparatus includes a cell portion, a light source portion, a detection portion, and a control portion. The cell portion includes a space into which a sample gas containing breath containing a first isotope of carbon dioxide and a second isotope of carbon dioxide is introduced. The light source portion changes a wavelength of the light in a band of 4.345 μm or more and 4.384 μm or less. The detection portion performs an operation including first detection of an intensity of the light passing through the space and second detection of an intensity of the light passing through the space into which the sample gas is not introduced. The control portion calculates a ratio of an amount of the second isotope to an amount of the first isotope based on a result of the first detection and a result of the second detection. | 1. A gas measuring apparatus, comprising: a cell portion that includes a space into which a sample gas containing a first isotope of carbon dioxide and a second isotope of carbon dioxide is introduced;a light source portion that emits light into the space in a wavelength band of 4.345 μm or more and 4.384 μm or less, the wavelength band including at least a first wavelength corresponding to one of absorption peak of the first isotope, a second wavelength corresponding to one of absorption peak of the second isotope, and a third wavelength corresponding to absorption peak of the first isotope or the second isotope, absorption of the first isotope at the second wavelength not corresponding to absorption peak and being lower than absorption of the second isotope;a detection portion that measures a first intensity of the light passing through the space into which the sample gas is introduced and a second intensity of the light passing through the space into which the sample gas is not introduced; anda control portion that calculates a ratio of an amount of the second isotope to an amount of the first isotope in the sample gas based on the first intensity and the second intensity,a range of wavenumber band of the light being 1 cm−1 or less, andwavenumber of the light being swept in the range. 2. The apparatus according to claim 1, wherein the light source portion includesa semiconductor light-emitting element that emits the light by energy relaxation of electrons in subbands in a plurality of quantum wells, anda driving portion that sweeps the wavenumber of the light. 3. The apparatus according to claim 1, wherein the control portion calculates the ratio a plurality of times. 4. The apparatus according to claim 1, further comprising:a gas introduction portion from which the sample gas is introduced;a first pipe; anda second pipe,one end of the first pipe being connected to the gas introduction portion, the other end of the first pipe being connected to an external environment,one end of the second pipe being connected to the first pipe or the gas introduction portion, andthe other end of the second pipe being connected to the cell portion. 5. The apparatus according to claim 4, wherein an amount of the sample gas flowing through the first pipe is larger than an amount of the sample gas flowing through the second pipe. 6. The apparatus according to claim 4, further comprising:a third pipe thorough which air is introduced into the cell portion from the external environment,the third pipe including a filter that reduces an amount of carbon dioxide in the air introduced from the external environment. 7. The apparatus according to claim 1, wherein the detection portion includes a semiconductor sensor element. | Kabushiki Kaisha Toshiba | G01J005/02 G01N021/3504 G01N021/39 G01N033/497 H01S005/343 G01N021/05 H01S005/125 H01S005/34 A61B005/00 A61B005/08 G01N033/00
US09799100 | B2 | 2017-10-24 | Integration of T2-preparation and motion tracking in magnetic resonance imaging | 2014-04-15 | An embodiment in accordance with the present invention provides concurrent measurement of motion during T2-weighted magnetic resonance imaging. The present invention combines T2 preparation, a module used to impart T2 contrast, and motion measurement, tracking, and/or correction. The present invention provides for the expedition of more efficient motion compensation during T2-weighted imaging. The proposed invention can be used to provide a variety of measurements of motion, with no overhead in imaging time. The proposed invention also enables T2 contrast imaging to be executed while a subject is breathing freely, without the additional time cost associated with the standard motion tracking methodologies. | 1. A method of obtaining a corrected image of a subject comprising: using a T2 preparation module comprising RF pulses and gradients to obtain a base magnetic resonance (MR) image of the subject;inserting motion tracking elements into the T2 preparation module;determining motion information from the motion tracking elements; andcorrecting the base MR image of the subject using the motion information. 2. The method of claim 1 further comprising interspersing said motion tracking elements between 2 consecutive RF pulses. 3. The method of claim 1 further comprising using the method for 1D projection of 2D and 3D volumes. 4. The method of claim 1 further comprising for tracking and correction of respiratory displacement. 5. The method of claim 4 further comprising tracking one of the respiratory displacement of the heart and the motion of the heart itself due to cardiac contraction. 6. The method of claim 1 further comprising tracking affine or non-rigid movement of the abdomen. 7. The method of claim 1 further comprising using RF pulses taking the form of at least one of hard pulses, shaped slice-selective pulses, MLEV composite pulses, adiabatic pulses, or any other equivalent pulses for excitation, refocusing, and restoring. 8. The method of claim 1 further comprising using signal conditioning and imaging data acquisition after using the T2 preparation module. 9. The method of claim 1 further comprising using the motion tracking element that comprises a projection. 10. The method of claim 9 further comprising inserting multiple projections along more than one axis. 11. The method of claim 1 further comprising using the motion tracking element comprising radial readouts. 12. The method of claim 1 further comprising using the motion tracking element taking the form of a motion-measurement readout. 13. The method of claim 1 further comprising using a magnetic resonance imaging machine to execute steps of the method. 14. The method of claim 1 further comprising executing steps of the method with a non-transitory computer readable medium. 15. The method of claim 13 further comprising using a computing device networked to the magnetic resonance imaging machine, the computing device being loaded with a non-transitory computer readable medium programmed with steps for executing the method. | The Johns Hopkins University | G06K009/00 G06T005/00 A61B005/055 A61B005/00 G01R033/56 G01R033/567 G01R033/565 G06T007/00 G06T007/207
US09775576 | B2 | 2017-10-03 | Systems and methods for applying scaling laws of tree structures | 2014-03-11 | Methods are provided for determining flow resistance within at least a portion of a vessel, whereby at least a partial biological tree image is obtained and resistance is calculated therefrom based upon at least the length of a defined vessel portion and a diameter of a stem vessel at a location proximal to a vessel crown. Methods are also provided for determining vessel volume, whereby at least a partial biological tree image is obtained and cumulative crown volume is calculated based on at least the diameter and length of the stem segment. | 1. A method for diagnosing and treating disease in a patient's biological tree, the method comprising the steps of: using a processor to produce an image showing a vasculature of at least part of a patient's biological tree, wherein the processor is operably connected to a storage medium capable of receiving and storing the image;calculating patient resistance based upon at least a length of a patient vessel portion and a diameter of a patient stem vessel proximal to the vessel crown from the patient's biological tree image;obtaining a model biological tree image and comparing the calculated patient resistance to a calculated model resistance calculated from the model biological tree image to determine the extent of vessel and/or organ disease; andtreating the patient based on the determined extent of vessel and/or organ disease. 2. The method of claim 1, wherein the calculated model resistance is based upon at least a length of a model vessel portion and a diameter of a model stem vessel proximal to the vessel crown from the model biological tree image. 3. The method of claim 1, wherein the patient resistance is a resistance of a stem segment, and wherein the length of the patient vessel portion is a length of a stem. 4. The method of claim 1, wherein the patient resistance is a resistance of a crown, and wherein the length of the patient vessel portion is a crown length. 5. The method of claim 1, wherein the patient resistance is a total resistance of the biological tree, and wherein the length of the patient vessel portion is a cumulative biological tree vessel length. 6. The method of claim 1, wherein the patient resistance is the resistance of blood within a vessel portion. 7. The method of claim 1, wherein the patient vessel portion comprises a vessel selected from the group consisting of a blood vessel, a bile duct, a lung, and a neuron. 8. The method of claim 1, wherein the patient resistance is the resistance of a gas within a vessel portion. 9. The method of claim 8, wherein the gas comprises breathable air, and wherein the vessel portion comprises a portion of a lung. 10. The method of claim 1, wherein the extent of vessel and/or organ disease is indicative of a disease selected from the group consisting of diffuse coronary artery disease, hypertension, atherosclerosis, heart failure, myocardial ischemia, myocardial infarction, an edematous disorder, and postural tachycardia syndrome. 11. A method for diagnosing and treating disease in a patient's biological tree, the method comprising the steps of: using a processor to produce an image showing a vasculature of at least part of a patient's biological tree, wherein the processor is operably connected to a storage medium capable of receiving and storing the image;calculating patient cumulative crown volume based upon at least a diameter of a patient stem segment and a length of a patient stem segment from the patient's biological tree image;obtaining a model biological tree ima... | DTherapeutics, LLC | G06F019/00 A61B005/00 A61B006/00 A61B005/02 A61B005/026 A61B005/055 A61B006/03
US09808232 | B2 | 2017-11-07 | Dilation system | 2012-11-01 | A dilation system for accessing a surgical target site to perform surgical procedures. In one version, the dilation system includes a plurality of open sided dilators having non-circular cross sections and configured to slidably and non-rotationally engage with one another in a sequential manner. In another version, the dilation system includes an expandable dilator movable from a closed condition to an expanded condition. | 1. A dilation system, comprising: a first dilator having an outer surface, a proximal end, a tapered distal end, and a bore extending from the proximal end to the tapered distal end, the first dilator having a first guide and a second guide formed on the proximal end thereof, the bore sized to receive a stimulating probe;a second dilator having a proximal end, a tapered distal end, an intermediate section positioned between the proximal end of the second dilator and the tapered distal end of the second dilator, and a generally C-shaped transverse cross section with an outer surface and an inner surface, the inner surface of the second dilator at the tapered distal end of the second dilator being sized and configured to be matingly received over the outer surface of the first dilator at the tapered distal end of the first dilator, the inner surface of the second dilator being configured along at least a portion of the intermediate section of the second dilator to be laterally inserted into engagement with the first guide of the first dilator, and the inner surface of the second dilator being configured along the proximal end of the second dilator to engage with the first guide of the first dilator in a way that the second dilator is slidably retained on the first dilator as the second dilator is slid along the first dilator from the proximal end of the first dilator toward the distal end of the first dilator; anda third dilator having a proximal end, a tapered distal end, an intermediate section positioned between the proximal end of the third dilator and the tapered distal end of the third dilator, and generally C-shaped transverse cross section with an outer surface and an inner surface, the inner surface of the third dilator at the tapered distal end of the third dilator being sized and configured to be matingly received over the outer surface of the first dilator at the tapered distal end of the first dilator, the inner surface of the third dilator being configured along at least a portion of the intermediate section of the third dilator to be laterally inserted into engagement with the second guide of the first dilator, and the inner surface of the third dilator being configured along the proximal end of the third dilator to engage with the second guide of the first dilator in a way that the third dilator is slidably retained on the first dilator as the third dilator is slid along the first dilator from the proximal end of the first dilator toward the distal end of the first dilator. 2. The dilation system of claim 1, wherein the first guide has a first size, wherein the second guide has a second size, and wherein the first size is different from the second size. 3. The dilation system of claim 1, wherein the first guide and the second guide extend radially outwardly in opposing directions. 4. The dilation system of claim 1, wherein each of the first guide and the second guide has a tapered proximal end. 5. The dilation system of claim 1,... | DePuy Synthes Products, Inc. | A61B017/02 A61B005/04 A61B017/34 A61B017/00
US09778187 | B2 | 2017-10-03 | Correction for osmotic pressure variations in chemo-optical sensor spots | 2014-05-30 | The present invention relates to a method for optically determining the concentration of a gas. The method includes using at least two luminescent dyes, the first being in-sensitive to the concentration of a gas with respect to the luminescence response (reference dye) and the second being sensitive to the concentration of a gas with respect to the luminescence response (indicator dye) the dyes show different luminescence decay times so that the resultant phase angle is indicative for the concentration of a gas, wherein the detected luminescent amplitude of the reference dye at a first moment in time is utilized to correct for sensitivity changes after the first moment. | 1. A method for optically determining concentration of a gas, the method comprising: contacting a gas with a reference dye and an indicator dye, the reference dye being in-sensitive to the concentration of the gas with respect to a luminescence response and the indicator dye being sensitive to the concentration of the gas with respect to the luminescence response,measuring the luminescence decay times of the reference dye and the indicator dye when in contact with the gas;measuring a resultant phase angle of the decay times, the resultant phase angle being indicative of the concentration of the gasmeasuring a luminescent amplitude of the reference dye at a first moment in time; andcorrecting a measurement of luminescence response of the indicator dye at a second moment in time based on the measured luminescent amplitude of the reference dye at the first moment in time. 2. The method of claim 1, wherein the dyes are excited by a single light source. 3. The method of claim 1, wherein the dyes are excited simultaneously. 4. The method of claim 1, wherein the dyes are provided in a sensing layer of a chemo-optical sensor unit. 5. The method of claim 4, wherein the chemo-optical sensor unit comprises, adjacent to the sensing layer, at least one gas-permeable layer, and wherein the method futher comprises: passing the gas through the gas-permeable layer prior to the sensing layer. 6. The method of claim 5, further comprising preventing by the gas-permeable layer, light from passing through the gas-permeable layer. 7. The method of claim 5, further comprising disposing a contact medium at least between the gas-permeable layer and a surface layer on which the chemo-optical sensor is to be applied. 8. The method of claim 1, comprising: measuring a phase behavior of luminescent responses of the reference and indicator dyes by a single detector;obtaining a luminescent response vector {right arrow over (F)} based on the measured phase behavior responses, wherein the luminescence response vector {right arrow over (F)} is independent of a total intensity of both luminescent dyes;de-composing the measured luminescent response vector {right arrow over (F)} into an imaginary part (β) of vector {right arrow over (F)} and a real part (α) of vector {right arrow over (F)}, wherein the imaginary part (β) reflects the reference dye and the real part (α) is a summation of a real part of the reference dye and of the indicator dye according to a formula (I), wherein A is an amplitude of the luminescent response vector {right arrow over (F)}, iφ is an imaginary phase shift w(t) is frequency of a light modulation, and Re({right arrow over (F)}) and Im({right arrow over (F)}) are real and imaginary parts of the luminescent response vector {right arrow over (F)} respectively, and wherein formula (I) comprises: {right arrow over (F)}=A eiφ=Re({right arrow over (F)})+i Im({right arrow over (F)})=A cos(ωt)+i A sin(ωt)=α+iβcompensating the measured luminescent response of ... | KONINKLIJKE PHILIPS N.V. | G01N021/64 G01N033/497 G01N033/84 A61B005/1468 G01N021/77
US09788779 | B2 | 2017-10-17 | Seizure detection based on work level excursion | 2014-01-30 | We report a method of determining an occurrence of an epileptic convulsive seizure in a patient, comprising: receiving body data from a patient during a first time period, determining a work level relating to said first time period at least based partially upon said body data; determining whether said work level exceeds an extreme work level threshold; performing a responsive action, in response to a determination that said work level exceeds said extreme work level threshold. We also report a medical device system configured to implement the method. We also report a non-transitory computer readable program storage unit encoded with instructions that, when executed by a computer, perform the method. | 1. A medical device, comprising: a body data module configured to receive body data from a patient;a work level data module configured to determine a first work level for the patient during a first time period, and a second work level for the patient during a second time period;a work level excursion module configured to:determine whether the second work level exceeds the first work level by an amount greater than a first work level excursion threshold which indicates that the second work level has a current body signal value which is incommensurate to an expected body signal value range for the second work level; anddetect a work level excursion based upon the determination that the current body signal value for the second work level is incommensurate to the expected body signal value range for the second work level which indicates an epileptic seizure event; anda controller configured to perform at least one of generating a warning of the epileptic seizure event, providing a therapy for the epileptic seizure event, or logging an occurrence of the epileptic seizure event, in response to detecting the work level excursion. 2. The medical device of claim 1, wherein the first time period is at a first time of day on a current day and the second time period is at the first time of day on at least one prior day. 3. The medical device of claim 1, wherein the first time period and the second time period are contiguous, partially overlapping, or fully overlapping. 4. The medical device of claim 1, wherein the first work level and the second work level are each a work level magnitude or a work level rate of change. 5. The medical device of claim 1, further comprising: at least one sensor, each configured to collect at least one body signal from the patient relating to a work level of the patient. 6. The medical device of claim 5, wherein the at least one sensor comprises at least one of: an autonomic data acquisition unit configured to acquire at least one of heart beat data, blood pressure data, respiration data, or blood gas data;a neurological data acquisition unit configured to acquire kinetic data selected from the group consisting of accelerometer data and inclinometer data;an endocrine data acquisition unit configured to acquire endocrine data,a metabolic data acquisition unit configured to acquire metabolic data;a tissue stress marker data acquisition unit configured to acquire tissue stress marker data; ora physical fitness data acquisition unit configured to acquire data relating to the patient's physical fitness. 7. The medical device of claim 1, further comprising a patient state data unit configured to determine the patient state based upon the body data. 8. The medical device of claim 1, wherein the medical device is configured to confirm the pathological state based on at least one of cardiac data of the patient, kinetic data of the patient, the patient's responsiveness, or a patient's awareness. 9. The medical device of claim 1, w... | Flint Hills Scientific, L.L.C. | A61B005/00 A61B005/1455 A61B005/0205 A61B005/11 A61B005/024 A61B005/08 A61B005/145
US09823120 | B2 | 2017-11-21 | UV dosimetry system with optimal sun exposure prediction | 2014-12-18 | A UV exposure dosimetry system includes at least one UV sensor that accurately measures the UV irradiance intensity. The UV dosimetry system integrates the measured UV irradiance intensity over time to calculate the real-time UV dosage and the vitamin D production by taking into account factors comprising UV sensor location, body surface area, clothing coverage, and sunscreen usage. Based on the measurement, the system can predict the time remaining to skin burn and the time remaining to reach daily goal of vitamin D production. The system can also estimate UV intensity for a time in the future at a geographic location based on the forecast UV index data, and predict UV dose and vitamin D generation for the user corresponding to user defined scenarios. The UV dosimetry system supports multi-user control through an advanced and user friendly input and output interface. | 1. A device for optimizing a user's UV exposure, comprising: a UV sensing unit comprising a circuitry for measuring an irradiating UV intensity;a communication unit for retrieving forecast UV index data;a computing unit for estimating UV intensity for a time in the future at a geographic location based on the retrieved forcast UV index data, predicting UV dose and vitamin D generation for the user at the geographic location corresponding to each of a plurality of user-defined scenarios based on the estimated UV intensity, calculating a composite score based on the predicted UV dose and vitamin D generation for each of the plurality of user-defined scenarios, wherein the composite score assesses a degree of UV balance between a risk of UV exposure and a benefit of UV exposure; and the computing unit is further adapted to calculate a present UV dose and a present vitamin D generation for the user based on the measured irradiating UV intensity;and an output unit for displaying a ranked list of user-defined scearios according to the calculated composite score for each of the plurality of user-defined scenarios, wherein a scenario ranked higher on the list corresponds to a higher degree of UV balance and a more optimal UV exposure condition; and the output unit is further adapted to display the present UV dose and the present vitamin D generation; wherein the device results in the user being treated with the optimal UV exposure condition. 2. The device of claim 1, wherein each of the user-defined scenarios specifies an exposed body surface area of the user at the time in the future, wherein the exposed body surface area is a function of the user's height, weight, and clothing coverage. 3. The device of claim 2, wherein each of the user-defined scenarios specifies a hypothetical start time of UV exposure by the user at the time in the future. 4. The device of claim 3, wherein each of the user-defined scenarios specifies a hypothetical duration of UV exposure by the user at the time in the future. 5. The device of claim 3, wherein the device is adapted to predict a time duration that is needed for the accumulated UV dose from the hypothetical start time to reach a user-programmable UV dose limit. 6. The device of claim 3, wherein the device is adapted to predict a time duration that is needed for the accumulated vitamin D generation from the hypothetical start time to reach a user-programmable goal of vitamin D. 7. The device of claim 1, wherein estimating UV intensity for a time in the future at a geographic location is performed by means of temporal and spatial interpolation of the retrieved forecast UV index data. 8. The device of claim 1, wherein the risk of UV exposure is expressed as a function of the predicted UV dose and a user-programmable maximum UV dose value, and the benefit of UV exposure is expressed as a function of the predicted vitamin D generation and a user-programmable minimum vitamin D value. 9. The device of claim 1, wherei... | Ultra, Inc. | G01N033/48 G01N033/50 G01J001/42 G01T001/02 G06F001/16 G01J001/02 G01J001/04 A61B005/00 G01J001/44
US09818189 | B2 | 2017-11-14 | Method of classification of organs from a tomographic image | 2012-02-17 | The present invention relates to a method for classification of an organ in a tomographic image. The method comprises the steps of receiving (102) a 3-dimensional anatomical tomographic target image comprising a water image data set and a fat image data set, each with a plurality of volume elements, providing (104) a prototype image comprising a 3-dimensional image data set with a plurality of volume elements, wherein a sub-set of the volume elements are given an organ label, transforming (106) the prototype image by applying a deformation field onto the volume elements of the prototype image such that each labeled volume element for a current organ is determined to be equivalent to a location for a volume element in a corresponding organ in the target image, and transferring (108) the labels of the labeled volume elements of the prototype image to corresponding volume elements of the target image. | 1. A method for classification of organs in a tomographic image, the method comprising the steps of: receiving a 3-dimensional anatomical tomographic target image comprising a water image data set and a fat image data set, each of the water image data set and the fat image data set with a plurality of volume elements,providing a prototype image comprising a 3-dimensional image data set with a plurality of volume elements, wherein a sub-set of the volume elements of the prototype image are given an organ label,transforming the prototype image by applying a deformation field onto the volume elements of the prototype image such that each labeled volume element for a current organ is determined to be equivalent to a location for a volume element in a corresponding organ in the target image,transferring the labels of the labeled volume elements of the prototype image to corresponding volume elements of the target image, wherein similarly labeled volume elements in the target image define an organ volume representing a current organ,extending said organ volume to the entire organ based on properties in the target image,calibrating the target image based on labeled organs in the water image data set, wherein each labeled volume element in the target image is given a predetermined water value, andsubtracting said predetermined water value with that labeled volume element's fat value in the fat image data set. 2. The method according to claim 1, wherein the step of extending the organ volume comprises detecting an edge of the organ in the target image. 3. The method according to claim 1, further comprising a step of manually entering a label to, or removing a label from, a volume element in the target image, wherein said manual step is followed by a redefinition of the organ volume referring to the present label. 4. The method according to claim 1, wherein the step of transforming the prototype image comprises applying a non-rigid co-registration of the prototype image onto the target image to increase the similarity between the prototype image and the target image, and wherein the non-rigid co-registration provides said deformation field representing a deformed prototype image. 5. The method according to claim 1, wherein said tomographic image is received from a magnetic resonance (MR) scanner or a computed tomography (CT) scanner. 6. The method according to claim 1, further comprising the step of removing the label of a first labeled volume element in the target image based on an analysis of said first labeled volume element compared to a similarly labeled second volume element in the target image. 7. The method according to claim 6, wherein said analysis comprises an analysis of said first labeled volume element in the water image data set and in the fat image data set compared to the similarly labeled second volume element in the water image data set and the fat image data set. 8. The method according to claim 6, wherein the step of removing th... | ADVANCED MR ANALYTICS AB | G06T007/00 A61B005/055 A61B006/03 A61B006/00 G06T007/33 G06T007/11 G06T007/174
US09782116 | B2 | 2017-10-10 | Stability-assessing system | 2015-05-01 | A system for assessing the stability of a person relative to a horizontal surface during an assessment session is provided. The system includes a platform having a stage upon which the person stands. A sensing configuration senses when the person is unbalanced on the stage, and conveys this information to a processor. The processor converts this information into meaningful graphics for viewing on a display. | 1. A stability-assessing device for assessing the stability of a person relative to a horizontal surface, the device comprising: a platform having a generally planar top surface, the platform having a perimeter;a base disposed below and supporting the platform, the base having a generally non-planar bottom surface such that the platform will rotate in response to a weight imbalance on the platform;a sensing configuration integrated into the device, wherein the sensing configuration is configured to sense orientation and/or motion of the platform relative to the horizontal surface and to generate signals representative of the sensed orientation and/or motion of the platform relative to the horizontal surface, wherein the platform is electronically divided into zones around the perimeter of the platform; anda visual indicator configuration integrated into the device, wherein the visual indicator configuration is configured to provide selective visual indications to one or more of the electronically defined zones. 2. The stability-assessing device of claim 1, wherein the sensing configuration is configured to sense orientation and/or motion of each zone of the platform. 3. The stability-assessing device of claim 1, wherein the sensing configuration includes one or more sensors associated with each zone and configured to sense orientation and/or motion of each zone of the platform relative to the horizontal surface. 4. The stability-assessing device of claim 3, wherein the sensors include pressure sensors, force sensors, acceleration sensors, tilt sensors and/or bump sensors. 5. The stability-assessing device of claim 1, wherein the platform is electronically divided into at least four zones. 6. The stability-assessing device of claim 1, wherein the platform is electronically divided into at least eight zones. 7. The stability-assessing device of claim 1, wherein the sensor configuration is configured to detect when a portion of the perimeter of the platform corresponding to one of the electrically defined zones contacts the horizontal surface. 8. The stability-assessing device of claim 7, wherein the sensor configuration is configured to detect the force with which a portion of the perimeter of the platform corresponding to one of the electrically defined zones contacts the horizontal surface. 9. The stability-assessing device of claim 1, wherein the sensor configuration includes a plurality of sensors within the platform configured to detect a person's weight distribution when the user is on the platform. 10. A stability-assessing device for assessing the stability of a person relative to a horizontal surface, the device comprising: a platform having a generally planar top surface, the platform having a perimeter;a base disposed below and supporting the platform, the base having a generally non-planar bottom surface such that the platform will rotate in response to a weight imbalance on the platform;a sensing configuration integrated into ... | - | A61B005/00 A61B005/11
US09804189 | B2 | 2017-10-31 | Upper body motion measurement system and upper body motion measurement method | 2015-03-25 | An upper body motion measurement system 1 has a plurality of inertia sensor units 2, each of which incorporating an angular velocity sensor 4 and an acceleration sensor 5. The plurality of the inertia sensor units 2 is attached to places that are different from each other on the upper body of a subject P. Based on the detection outputs of the angular velocity sensor 4 and the acceleration sensor 5, the attitude of each of the inertia sensor units 2 is estimated, and the acceleration thereof is further estimated. The angular acceleration of the upper body of the subject P is estimated based on the estimated accelerations of the plurality of the inertia sensor units 2. | 1. An upper body motion measurement system comprising: at least one central processing unit comprisinga plurality of inertia sensor units, each of which includes a set of an angular velocity sensor that detects an angular velocity and an acceleration sensor that detects an acceleration and which are attached to places that are different from each other on an upper body of a subject;a sensor unit attitude estimating unit configured to estimate an attitude of each of the inertia sensor units in a motion environment of the subject based on a detection value of an angular velocity and a detection value of an acceleration indicated by detection outputs of the angular velocity sensor and the acceleration sensor of each of the inertia sensor units attached to the upper body of the subject;a sensor unit acceleration estimating unit configured to estimate an acceleration of each of the inertia sensor units observed in a global coordinate system set in the motion environment of the subject based on a detection value of an acceleration indicated by a detection output of the acceleration sensor of each of the inertia sensor units attached to the upper body of the subject, and the attitude of each of the inertia sensor units estimated by the sensor unit attitude estimating unit; andan upper body angular acceleration estimating unit configured to estimate an angular acceleration of the upper body of the subject observed in the global coordinate system based on a set of accelerations estimated by the sensor unit acceleration estimating unit on each of the plurality of the inertia sensor units. 2. The upper body motion measurement system according to claim 1, wherein the plurality of the inertia sensor units includes at least two inertia sensor units attached to the upper body with an interval provided in a vertical direction of the upper body along a trunk axis of the subject, andthe angular acceleration estimated by the upper body angular acceleration estimating unit includes at least one of the angular acceleration of the upper body in a roll direction of the subject and the angular acceleration in a pitch direction of the subject. 3. The upper body motion measurement system according to claim 2, wherein an upper inertia sensor unit of the two inertia sensor units is attached to the upper body of the subject at a location on an upper side from a boundary between a thoracic vertebra and a lumbar vertebra of the subject, and a lower inertia sensor unit of the two inertia sensor units is attached to the upper body of the subject at a location on a lower side from the boundary between the thoracic vertebra and the lumbar vertebra of the subject. 4. The upper body motion measurement system according to claim 1, wherein the acceleration of each of the inertia sensor units observed in the global coordinate system or the angular acceleration of the upper body of the subject observed in the global coordinate system is defined as a target state amount, and the senso... | HONDA MOTOR CO., LTD. KYOTO UNIVERSITY | A61B005/00 G01P015/18 G01C021/16 A61B005/107 A61B005/11
US09816882 | B2 | 2017-11-14 | Electronic skin, preparation method and use thereof | 2014-01-28 | The invention provides a piezoresistive electronic skin, a preparation method and a use thereof. The piezoresistive electronic skin uses carbon nanotube film as the conductive layer and uses materials provided with micro-nano patterns, such as polydimethylsiloxane, polyethylene terephthalate, polyvinyl alcohol, polyvinyl formal, polyethylene, and so on, as the substrate, enabling the substrate has advantages of high flexibility and being pliable, and it needs low operating voltage and little power consumption, but has high sensitivity and short response time. More importantly, the invention uses the patterned flexible substrate as the basis, greatly improving the sensitivity of electronic skin reacting to tiny applied force from outside. The invention also provides a capacitive electronic skin and a preparation method thereof. Further, the invention also provides a use of the piezoresistive electronic skin or the capacitive electronic skin on speech recognition, pulse detection, medical robot, etc. | 1. A piezoresistive electronic skin, characterized by comprising: a plurality of overlapped flexible substrates;two conductive layers respectively arranged on two microstructure-patterned contact surfaces of adjacent flexible substrates and being contacted with each other, wherein contact areas of said conductive layers have non-planar microstructure; andconductive electrodes electrically connected with said two conductive layers;wherein said conductive layer comprises a carbon nanotube film which comprises networks formed by cross-linked carbon nanotubes;wherein said plurality of overlapped flexible substrates are made of one or a combination of more of the following materials: polydimethylsiloxane (PDMS) film, polyvinylidene fluoride (PVDF), polyvinyl fluoride (PVF), polyvinyl chloride (PVC), polyethylene (PE), polypropylene (PP), polystyrene (PS), polymethyl methacrylate (PMMA), nylon (Nylon), polycarbonate (PC), polyurethane (PU), polytetrafluoroethylene (PTFE), polyethylene terephthalate (PET, PETE), and high-polymer rubber materials. 2. The piezoresistive electronic skin according to claim 1, characterized in that at least one surface of said polydimethylsiloxane film is provided with patterns, the sizes of said patterns are between 0.1 and 500 μm, and the unit amount of said patterns in a square centimeter is between 1 and 1012. 3. The piezoresistive electronic skin according to claim 2, characterized in that said patterns are molded by coating said polydimethylsiloxane onto a template and solidifying it, said template is anyone of a silicon substrate provided with microstructure, a glass substrate provided with microstructure, a metallic substrate provided with microstructure, a plastic substrate, fabric or silk article provided with microstructure, and a bio-organ provided with microstructure. 4. The piezoresistive electronic skin according to claim 1, characterized in that the light transmittance of said carbon nanotube film is between 50% and 97%, and the sheet resistance is between 102 Ω/sq and 107 Ω/sq. 5. The piezoresistive electronic skin according to claim 4, characterized in that the thickness of said carbon nanotube film is between 10 nm and 500 nm, said carbon nanotube film is made of one type or a combination of more types of single-walled carbon nanotubes, double-walled carbon nanotubes and multi-walled carbon nanotubes, wherein said single-walled carbon nanotubes are metallic single-walled carbon nanotubes, semiconductor-type carbon nanotubes or hybrid single-walled carbon nanotubes containing both metallic single-walled carbon nanotubes and semiconductor-type carbon nanotubes. 6. The piezoresistive electronic skin according to claim 1, characterized in that said conductive layer is made of one or a combination of more of any conductive metal among copper, silver and gold, and semiconductor materials. 7. The piezoresistive electronic skin according to claim 1, characterized in that said conductive layer is attached wit... | SUZHOU INSTITUTE OF NANO-TECH AND NANO-BIONICS (SINANO), CHINESE ACADEMY OF SCIENCES | G01L001/18 G06F003/041 G01L001/14 A61B005/11 A61B005/024 A61B005/00 A61B034/30 A61B001/00 A61B003/16 A61B005/0205 A61B005/021 B29C041/02 B29C041/42 B29C065/00 B29K083/00 B29L009/00 B29L031/34 B29L031/00
US09814857 | B2 | 2017-11-14 | Method and corresponding kit for administering an adductor canal block | 2016-01-21 | Apparatus for administering certain nerve blocks includes a sheath constructed from a flexible ultrasound echogenic material, a more rigid introducer/dilator for introducing the sheath into the patient, and an ultrasound echogenic catheter for inserting through the sheath once the distal end of the sheath is in place adjacent the nerve(s) to be blocked and the introducer/dilator has been withdrawn. The catheter has provisions at its proximal end for connecting to a source of local anesthetic. Methods for use of this apparatus are also described. | 1. A method for administering an adductor canal block (ACB) to a patient, the method comprising: identifying the patient's sartorius muscle,advancing a sheath and a needle sheathed within the sheath having an open distal end through which a distal end of the needle projects to the base of the sartorius muscle,injecting a first amount of a local anesthetic into the sartorius muscle, wherein the first amount of the local anesthetic is injected at the base of the satorius muscle,identifying the bottom of the sartorius fascia,advancing the sheath and the sheathed needle through the sartorius fascia,administering a second amount of a local anesthetic under the sartorius muscle,removing the needle, leaving the sheath in the patient's adductor canal,inserting a catheter into the sheath and advancing a distal end of the catheter through the sheath to the adductor canal,removing the sheath while leaving the catheter in place,connecting a proximal end of the catheter to a source of local anesthetic, andcommencing anesthesia, wherein a third amount of local anesthetic is delivered into the patient's adductor canal. 2. The method of claim 1 wherein identifying the patient's sartorius muscle comprises identifying the patient's sartorius muscle using ultrasound. 3. The method of claim 1 wherein connecting the proximal end of the catheter to the source of local anesthetic, and commencing anesthesia comprises connecting the proximal end of the catheter to a constant infusion pump to deliver the local anesthetic postoperatively. 4. The method of claim 1, wherein the catheter is a multi-orifice catheter and anesthesia is delivered locally from the catheter while the catheter is left in place. | Avent, Inc. | A61M019/00 A61B046/00 A61M005/42 A61M025/06 A61B001/313 A61M005/142 A61B046/20 A61M005/46 A61M005/158 A61M005/172 A61M025/00 A61B017/34 A61B005/11 A61B005/00 A61B008/08 A61M029/00
US09788811 | B2 | 2017-10-17 | Imaging system of microbubble therapy and image evaluation method using the same | 2014-06-12 | An imaging system of microbubble therapy cooperated with an ultrasound device for monitoring a cavitation on microbubbles in a vessel of an affected part is disclosed in the present invention, in which the cavitation is occurred by applying an ultrasound to disrupt the microbubbles. The system comprises an image acquiring module and a controlling module. The image acquiring module comprises at least one magnetic resonance device for acquiring a plurality of magnetic resonance images of the cavitation, and the controlling module provided for controlling an acquiring time of the magnetic resonance device and an irradiation time of the ultrasonic device through a controlling mode. An image evaluation method using the same is also disclosed herein and comprises steps as the following. First, injecting the microbubbles into the vessel of the affected part is performed. And then, a plurality of magnetic resonance images by a magnetic resonance device and in an acquiring time is acquired. The microbubbles are irradiated for an irradiation time by an ultrasound. Finally, changes of the magnetic resonance images will be monitored, in which an irradiation path of the ultrasound may be perpendicular to a direction of flow in the vessel and the irradiation time is within the acquiring time. | 1. An imaging system of microbubble therapy coordinated with an ultrasonic device for monitoring a cavitation on microbubbles in a vessel of an affected part, wherein the cavitation is created by applying an ultrasound to disrupt the microbubbles, comprising: an image acquiring module comprising at least one magnetic resonance device for acquiring a plurality of magnetic resonance images of the cavitation; anda controlling module provided for controlling an acquiring time of the magnetic resonance device and an irradiation time of the ultrasonic device through a controlling mode, wherein the controlling mode comprises a consecutive mode, an intermittent mode and a burst mode; anda computing module for providing a relation diagram between time and an acquiring signal intensity, wherein the acquiring signal intensity is a contrast value of the magnetic resonance images acquired, respectively, at a specific time and at an initial time. 2. The imaging system according to claim 1, wherein the microbubbles comprise drugs for treating the affected part. 3. The imaging system according to claim 1, wherein the magnetic resonance images are gradient echo magnetic resonance images. 4. The imaging system according to claim 3, wherein the gradient echo magnetic resonance images are acquired by adopting a flip angle of 0-90 degrees. 5. The imaging system according to claim 1, wherein the acquiring signal intensity is a mean contrast value of several positions of the magnetic resonance images acquired, respectively, at the specific time and at the initial time of several positions. 6. The imaging system according to claim 1, the computing module provides a revising value that is fed back to the controlling module for adjusting the magnetic resonance images taken. 7. The imaging system according to claim 1, wherein the consecutive mode is performed to apply continuous ultrasound pulses for the irradiation time within the acquiring time. 8. The imaging system according to claim 1, wherein the intermittent mode is performed to apply several times of continuous ultrasound pulses spaced at intervals for the irradiation time within the acquiring time. 9. The imaging system according to claim 1, wherein the burst mode is performed to apply ultrasound pulses for the irradiation time with a duty cycle within the acquiring time. 10. An image evaluation method of microbubble therapy for monitoring a cavitation of microbubbles in a vessel of an affected part, comprising steps: injecting the microbubbles into the vessel of the affected part;acquiring a plurality of magnetic resonance images by a magnetic resonance device and in a acquiring time;irradiating the microbubbles for an irradiation time by an ultrasound, wherein the irradiation time is within the acquiring time; andmonitoring changes of the magnetic resonance images, and providing a relation diagram between time and an acquiring signal intensity, wherein the acquiring signal intensity is a contrast value ... | NATIONAL TSING HUA UNIVERSITY | G01V003/00 A61B008/06 G01R033/48 A61B005/055 A61B008/08 G01R033/56 A61B005/00 A61N007/00 A61B090/00
US09782590 | B2 | 2017-10-10 | Brain stimulation system including diagnostic tool | 2014-10-16 | A system for treating a patient comprises a stimulator for stimulating brain tissue, a controller for setting stimulation parameters and a diagnostic tool for measuring patient parameters and producing diagnostic data. The stimulation parameters comprise test stimulation parameters and treatment stimulation parameters. The stimulator delivers test stimulation energy to the brain tissue based on at least one test stimulation parameter and delivers treatment stimulation energy to the brain tissue based on at least one treatment stimulation parameter. One or more treatment stimulator parameters are determined based on the diagnostic data produced by the diagnostic tool The system is constructed and arranged to treat a neurological disease or a neurological disorder. Methods of treating a neurological disease or neurological disorder are also provided. | 1. A system for treating a patient comprising: a stimulator for stimulating brain tissue;a controller for setting stimulation parameters comprising at least one test stimulation parameter of the stimulator and at least one treatment stimulation parameter of the stimulator; anda diagnostic tool for measuring at least one patient parameter and producing diagnostic data representing the at least one measured patient parameter;wherein the controller is configured to analyze the diagnostic data and compare the diagnostic data to a threshold characteristic of an adverse event or a desired event;wherein the stimulator is constructed and arranged to deliver test stimulation energy to the brain tissue based on the at least one test stimulation parameter and to deliver treatment stimulation energy to the brain tissue based on the at least one treatment stimulation parameter;wherein the at least one treatment stimulator parameter is determined based on the diagnostic data;wherein the controller is configured to set the at least one treatment stimulation parameter based on a safety margin based on the comparison of the diagnostic data to the threshold characteristic; andwherein the system is constructed and arranged to treat at least one of a neurological disease or a neurological disorder. 2. The system according to claim 1, wherein the at last one of a neurological disease or a neurological disorder comprises Alzheimer's Disease. 3. The system according to claim 1, wherein the brain tissue stimulated comprises a fornix. 4. The system according to claim 1, wherein the system is constructed and arranged to determine the at least one treatment stimulation parameter to at least one of prevent or reduce the adverse event. 5. The system according to claim 1, wherein the system is constructed and arranged to determine the at least one treatment stimulation parameter to improve treatment of the at least one of a neurological disease or a neurological disorder. 6. The system according to claim 1, wherein at least one of the stimulation parameters comprises a signal voltage ranging between 0.1 Volts and 10.0 Volts. 7. The system according to claim 1, wherein the controller is constructed and arranged to set at least one stimulation parameter based on a threshold at which the adverse event is detected by the diagnostic tool. 8. The system according to claim 7, wherein the adverse event comprises an event selected from the group consisting of: undesirable heart rate; undesirable respiration rate; undesirable sweating; undesirable hallucinations; undesirable tingling; flushing; an undesirable psychiatric effect; an undesirable cognitive effect; unpleasant generalized warming; undesirable perceptions described as déjà vu; seizure; synchronized neuronal firing pattern; undesired neural response time; undesired brain state; undesired theta phase; undesired p300 amplitude; and combinations thereof. 9. The system according to claim 7, wherein the at least one stimul... | Functional Neuromodulation, Inc. | A61N001/36 A61N002/00 A61N005/06 A61M005/172 A61B005/00 A61M005/142 A61B005/0402 A61B005/0476 A61B005/055 A61N001/372 A61N001/05 A61N007/00
US09801528 | B2 | 2017-10-31 | Guidewire | 2013-10-21 | [Problem] To provide a guidewire that has good operability and that can perform procedures efficiently.;[Solution to Problem] A guidewire is for guiding insertion of a medical appliance into a subject. The guidewire includes an inserting portion and an observation optical system. The inserting portion has a first part and a second part that has higher flexibility than the first part, and the inserting portion is to be inserted into the subject. The observation optical system is provided to the inserting portion, and the observation optical system is provided for observing inside the subject. | 1. A guidewire that is for guiding insertion of a medical appliance into a subject, comprising: an inserting portion to be inserted into the subject, the inserting portion having a first part and a second part that has higher flexibility than the first part; andan observation optical system that is for observing inside the subject, the observation optical system being provided to the inserting portion,a helical slit being formed in at least a part of the second part, the helical slit having two or more types of pitch widths,the slit being provided with curved incisions at predetermined intervals, the curved incisions each having bending parts. 2. A guidewire according to claim 1, wherein the inserting portion hasa tip portion configuring a tip anda long shaft portion having the first part and the second part, the shaft portion being connected to the tip portion, andthe second part is provided to the tip portion side, in the shaft portion. 3. A guidewire according to claim 1, wherein a hydrophilic coating layer is formed in a periphery of the slits. 4. A guidewire according to claim 1, wherein the observation optical system is configured includingan image pickup device that captures an image inside the subject, the image pickup device being provided to a tip of the inserting portion,an objective lens provided in front of an imaging area of the image pickup device, andan illuminating optical system that is for illuminating inside the subject. 5. A guidewire according to claim 1, wherein the observation optical system is configured includingan image guide fiber that transmits an optical image inside the subject,an objective lens provided in front of an incidence plane of the image guide fiber,an image pickup device that captures an optical image that has been transmitted from the image guide fiber, the image pickup device being provided to a rear end of the inserting portion, andan illuminating optical system that is for illuminating inside the subject. 6. A guidewire according to claim 4, wherein the image pickup device is a CMOS sensor. 7. A guidewire according to claim 4, wherein the illuminating optical system includes a light guide fiber that guides light from a light source to inside the subject. 8. A guidewire according to claim 7, wherein a plurality of the light guide fibers are provided, andthe plurality of the light guide fibers are arranged in a peripheral edge of the objective lens. 9. A guidewire according to claim 4, wherein the illuminating optical system is provided to a tip of the inserting portion and the illuminating optical system includes an LED light source that irradiates light to inside the subject. 10. A guidewire according to claim 1, wherein at least two of the bending parts are formed and the two bending parts are bent in different directions with respect to the slit. | FUJIKURA LTD. | A61B006/00 A61B001/012 A61B001/00 A61B001/04 A61B001/06 A61B001/07 A61B005/00 A61M025/09
US09776013 | B2 | 2017-10-03 | System for analyzing energy delivered to ECG device from defibrillator | 2015-06-10 | A system for analyzing the energy delivered to ECG device from the defibrillator is disclosed. The system includes an energy analyzing unit configured to measure the energy delivered to the ECG device from the defibrillator, wherein the energy gets diverted to the ECG device during operation of the defibrillator; and a presentation unit capable of presenting the measured energy in the ECG device. | 1. A system for analyzing energy delivered to an ECG device from a defibrillator, the system comprising: an energy analyzing unit configured to measure the energy diverted to the ECG device from the defibrillator;wherein the diverted energy is a portion of a total energy outputted by the defibrillator and is diverted to the ECG device rather than delivered as a therapeutic pulse to a subject; anda presentation unit configured to present the diverted energy measurement. 2. The system of claim 1, wherein the energy analyzing unit is further configured to measure a recovery time of the ECG device from an energy overload from the defibrillator. 3. The system of claim 1, further comprises an energy setting engine configured to set a predefined threshold for the energy diverted to the ECG device based on user input. 4. The system of claim 3, wherein the energy analyzing unit is further configured to: determine if the energy diverted to the ECG device is within the predefined threshold; andsend an alarm signal if the diverted energy measurement exceeds the predefined threshold. 5. The system of claim 1, wherein the energy analyzing unit is further configured to set a defibrillator mode for the ECG device, wherein in the defibrillator mode the energy analyzing unit measures the diverted energy to the ECG device. 6. The system of claim 5, wherein the energy analyzing unit is configured to assign a key for setting the ECG device in the defibrillator mode. 7. An ECG device operating with a defibrillator, the ECG device comprising: an energy analyzing unit configured to measure the energy diverted from the defibrillator, wherein the diverted energy is a portion of a total energy outputted by the defibrillator and is diverted to the ECG device rather than delivered as therapeutic pulse to a subject; anda display device configured to present the diverted energy measurement. 8. The ECG device of claim 7, wherein the energy analyzing unit is further configured to measure a recovery time of the ECG device from an energy overload from the defibrillator. 9. The ECG device of claim further comprises an energy setting engine configured to set a predefined threshold for the diverted energy received from the defibrillator based on user input. 10. The ECG device of claim 9, wherein the energy analyzing unit is further configured to: determine if the diverted energy received from the defibrillator is within the predefined threshold; andsend an alarm signal if the diverted energy exceeds the predefined threshold. 11. The ECG device of claim 7, wherein the energy analyzing unit is further configured to set a defibrillator mode, wherein the energy analyzing unit measures the diverted energy received from the defibrillator when the ECG device is in the defibrillator mode. 12. The ECG device of claim 11, wherein the energy analyzing unit is further configured to assign a key for setting the ECG device in the defibrillator mode. 13. The ECG device of claim 7, where... | General Electric Company | A61N001/39 A61B005/0402 A61B005/044 A61B005/00
US09776000 | B2 | 2017-10-03 | Programming systems for eliciting evoked responses in a cochlear implant patient and performing predetermined actions in accordance with the evoked responses | 2016-11-30 | An exemplary system includes 1) a programming device configured to be located external to a cochlear implant patient and communicatively coupled to a cochlear implant system associated with the patient, 2) a programming interface device communicatively coupled to the programming device and configured to be located external to the patient, and 3) a receiver communicatively coupled directly to the programming interface device. The programming device directs at least one of the cochlear implant system and the receiver to apply stimulation to the patient, records an evoked response that occurs in response to the stimulation, and performs a predetermined action in accordance with the evoked response. Corresponding systems and methods are also disclosed. | 1. A system comprising: a programming device configured to be located external to a patient and communicatively coupled to a cochlear implant system associated with the patient, the cochlear implant system comprising a sound processor configured to located external to the patient, a cochlear implant configured to be implanted within the patient, and an electrode array configured to be located within the patient;a programming interface device communicatively coupled to the programming device and configured to be located external to the patient; anda receiver communicatively coupled directly to the programming interface device;wherein the programming device directs, during a period of time after an implantation procedure in which the electrode array is inserted into the patient, at least one of the cochlear implant system and the receiver to apply stimulation to the patient,records an evoked response that occurs in response to the stimulation,compares the evoked response to an additional evoked response previously recorded by the programming device during the period of time,determines, based on the comparison, that a residual hearing status of the patient has deteriorated during the period of time,increases, in response to the determination that the residual hearing status of the patient has deteriorated during the period of time, an intensity level of additional stimulation provided by the cochlear implant system in order to present audio content to the patient subsequent to the period of time,determines that the residual hearing status of the patient improves subsequent to the period of time, anddecreases, in response to the determination that the residual hearing status of the patient improves subsequent to the period of time, the intensity level of the additional stimulation provided by the cochlear implant system. 2. The system of claim 1, wherein the stimulation comprises electrical stimulation, and wherein the programming device directs the cochlear implant system to apply the electrical stimulation to the patient by way of at least one electrode included in the electrode array. 3. The system of claim 1, wherein the stimulation comprises acoustic stimulation, and wherein the programming device directs the receiver to apply the acoustic stimulation to the patient. 4. The system of claim 1, wherein the stimulation comprises electrical stimulation and acoustic stimulation, and wherein the programming device concurrently directs the cochlear implant system to apply the electrical stimulation to the patient by way of at least one electrode included in the electrode array and the receiver to apply the acoustic stimulation to the patient. 5. The system of claim 1, wherein the programming device records the evoked response using an electrode included in the electrode array. 6. The system of claim 1, further comprising an extra-cochlear electrode communicatively coupled to at least one of the programming interface device and the cochlear implant... | Advanced Bionics AG | A61N001/36 A61B005/04 A61B005/12 A61N001/05 H04R025/00 A61N001/372 H04R001/10
US09808162 | B2 | 2017-11-07 | Pulse wave sensor and semiconductor module | 2015-10-22 | A pulse wave sensor includes: a white LED emitting white light to a human body; a G sensor converting, into a first electrical signal, green light included in light emitted from the white LED and reflected within the human body; an R sensor converting, into a second electrical signal, red light included in the light emitted from the white LED and reflected within the human body; and an arithmetic control unit configured to generate a signal showing a heart rate based on a level difference between the first electrical signal and the second electrical signal. Therefore, a distance between the G sensor and the R sensor does not have to be increased, so that an apparatus can be reduced in size. | 1. A pulse wave sensor comprising: a light emitting unit configured to emit light including green light and red light to a living body;a first light receiving unit configured to convert, into a first electrical signal, the green light included in light emitted from said light emitting unit and reflected within said living body;a second light receiving unit configured to convert, into a second electrical signal, the red light included in the light emitted from said light emitting unit and reflected within said living body; andan arithmetic control unit configured to detect information related to a pulse wave in said living body based on a level difference between said first electrical signal and said second electrical signal,said arithmetic control unit including: a subtracter configured to generate a third electrical signal showing the level difference between said first electrical signal and said second electrical signal,a first frequency component analysis unit configured to obtain a frequency spectrum of said first electrical signal, anda second frequency component analysis unit configured to obtain a frequency spectrum of said third electrical signal, andwherein the arithmetic control unit is configured to compare the frequency spectrum of said first electrical signal and the frequency spectrum of said third electrical signal to obtain the information related to the pulse wave in said living body based on a comparison result. 2. The pulse wave sensor according to claim 1, wherein said first electrical signal includes a pulse wave component that changes in accordance with the pulse wave in said living body, and a body motion component that changes in accordance with a body motion of said living body, andsaid second electrical signal includes said body motion component. 3. The pulse wave sensor according to claim 2, wherein said arithmetic control unit amplifies or attenuates at least one of said first electrical signal and said second electrical signal such that a level of said body motion component of said first electrical signal and a level of said body motion component of said second electrical signal are identical to each other, and then, obtains the level difference to generate said third electrical signal. 4. The pulse wave sensor according to claim 1, wherein said first frequency component analysis unit applies fast Fourier transform to said first electrical signal to obtain the frequency spectrum of said first electrical signal, andsaid second frequency component analysis unit applies fast Fourier transform to said third electrical signal to obtain the frequency spectrum of said third electrical signal. 5. The pulse wave sensor according to claim 1, wherein said first frequency component analysis unit applies discrete Fourier transform to said first electrical signal to obtain the frequency spectrum of said first electrical signal, andsaid second frequency component analysis unit applies discrete Fourier transform to said third ele... | Rohm Co., Ltd. | A61B005/02 A61B005/024 G01J001/44 G01J001/42 A61B005/00
US09782558 | B2 | 2017-10-10 | Minimally invasive lung volume reduction devices, methods, and systems | 2014-08-06 | A lung volume reduction system is disclosed comprising an implantable device adapted to be delivered to a lung airway of a patient in a delivery configuration and to change to a deployed configuration to bend the lung airway. The invention also discloses a method of bending a lung airway of a patient comprising inserting a device into the airway in a delivery configuration and bending the device into a deployed configuration, thereby bending the airway. | 1. A system for treating a patient having a lung with an airway, the system comprising: a lung implant having an elongate length with a proximal and distal end for deployment within the airway of the patient, the lung implant transitionable between a delivery configuration and a deployed configuration, the deployed configuration configured to impart a force on the airway sufficient to bend an axis of the airway so as to locally compress lung tissue;a delivery catheter with a lumen for delivering the lung implant to the airway of the lung; anda loading cartridge comprising a lumen and configured to contain the lung implant for delivery into the delivery catheter, the loading cartridge constraining the implant to the delivery configuration and the lung implant configured to elastically recover toward a bent configuration to impart the force on the airway that bends the axis of the airway into a folded configuration; anda grasper configured to couple with a proximal portion of the lung implant to pull the lung implant into the loading cartridge and to push the lung implant from the loading cartridge into the delivery catheter for delivery into the airway. 2. The system of claim 1, wherein the loading cartridge is configured to couple with a proximal end of the delivery catheter so that the lumen of the loading cartridge and the lumen of the delivery catheter form a continuous lumen. 3. The system of claim 2, wherein the loading cartridge is configured to couple to the proximal end of the delivery catheter via a locking hub connection. 4. The system of claim 3, wherein the locking hub connection comprises a luer lock hub. 5. The system of claim 1, wherein the grasper comprises a set of jaws for releasably locking onto the proximal portion of the lung implant. 6. The system of claim 1, further comprising a bronchoscope with a working channel and wherein the delivery catheter is configured to deliver the lung implant to the airway through the working channel of the bronchoscope. 7. The system of claim 6, further comprising an x-ray guidance system for imaging a position of the lung implant within the lung. 8. The system of claim 1, wherein the loading cartridge is terminated with an open end. 9. The system of claim 1, wherein the loading cartridge is terminated with a hub. 10. The system of claim 1, wherein the loading cartridge constrains the lung implant to a diameter that is equal to or less than 18 mm. 11. The system of claim 1, wherein the lung implant is disposed within the loading cartridge. 12. The system of claim 1, further comprising an actuation plunger. 13. The system of claim 12, wherein the actuation plunger is operable to lock the grasper to the lung implant. 14. The system of claim 12, wherein the actuation plunger is operable to release the lung implant from the grasper. 15. A method of treating a patient having a lung with an airway, the method comprising: positioning a distal end of a delivery catheter to the airway of ... | PneumRx, Inc. | A61B017/08 A61M016/20 A61B017/12 A61M016/04 A61B001/018 A61B001/04 A61B001/267 A61B005/06 A61B006/03 A61B006/12 A61B006/00 A61M037/00 A61F002/04 A61B090/00 A61M016/08 A61B017/00
US09788746 | B2 | 2017-10-17 | Diagnostic testing headband | 2014-07-23 | A diagnostic headband has an elongated body with two ends. A mechanism couples the two ends with one another. At least one first lead is coupled with the body. At least one second lead extends from the body. The leads couple with lead pads on the user. An electrical harness is coupled with the body. The electrical harness electrically couples with the at least one first and second leads. | 1. A diagnostic testing headband comprising: an elongated continuous flexible body with two ends, the body comprising a cover layer and a foam insert sandwiched within the cover layer, a mechanism coupled with the body for connecting the two ends;at least one first lead stationarily coupled with the body for coupling with conventional lead pads contacting the user's body when the headband is placed on the user's head, the at least one first lead including a mechanism for electrically coupling with the conventional lead pads;at least one second lead extending from the body for coupling with the conventional lead pads on the user, the at least one second lead including a wire to enable a lead cap to hang from the body and be moved to couple with the conventional lead pads at different positions on the user, the lead cap including a mechanism for electrically coupling with the conventional lead pads; andan electrical harness coupled with the body such that a portion of the electrical harness is positioned underneath the cover layer and adjacent the foam insert to extend between the first and second leads, the electrical harness electrically coupled with the at least one first lead and with the at least one second lead wire and the electrical harness including a connector for coupling with a testing apparatus. 2. The diagnostic testing headband of claim 1, wherein a plurality of first leads is coupled with the body. 3. The diagnostic testing headband of claim 2, wherein the plurality of first leads is color coded. 4. The diagnostic testing headband of claim 1, wherein a plurality of second leads extend from the body. 5. The diagnostic testing headband of claim 4, wherein the plurality of second leads is color coded. 6. The diagnostic testing headband of claim 1, wherein the mechanism for connecting the two ends of the body is a hook and loop fastener. | Good Sleep LLC | A61B005/0478 A61B005/00 A61B005/01 A61B005/024 A61B005/0408
US09795831 | B2 | 2017-10-24 | Monitoring fitness using a mobile device | 2015-12-29 | Athletic performance monitoring and tracking may provide multiple ways in which to track athletic movement and activity. In one example, an athletic monitoring device may include or be associated with multiple types of movement sensors and switch between the sensors or use both depending on various factors including type of workout. Workouts may also be tagged with various parameters including mood, weather, terrain, athletic equipment used and the like. In one or more examples, the parameters may be automatically determined based on location. User workouts and accomplishments may also be celebrated through messages from celebrities, family, friends and other users. In some cases, the messages may be triggered by various conditions. Coaching may also be provided to the user to help improve workouts and overall athletic performance. Running routes may also be automatically tracked, stored and shared. | 1. An apparatus comprising: a processor; andmemory storing computer readable instructions that, when executed, cause the apparatus to: receive, from a first sensor, athletic performance data for an athletic activity performed by a user and store the athletic performance data;receive, from a second sensor, geographic location data associated with the athletic activity;determine an attribute of the geographic location data, wherein the attribute includes information other than identification of geographic locations associated with the athletic activity performed;tag the athletic performance data with the attribute of the geographic location data;generate a route map showing the geographic location data, wherein the route map further includes a first marker at a first geographic location and a second marker at a second geographic location, wherein the first marker and the second marker are placed upon a determination that a first value of the attribute at the first geographic location is different than a second value of the attribute at the second geographic location; andresponsive to receiving a user input interacting with the first marker at the first geographic location, providing the first value of the attribute. 2. The apparatus of claim 1, wherein the attribute includes weather information. 3. The apparatus of claim 1, wherein the attribute includes terrain information. 4. The apparatus of claim 1, wherein the attribute includes elevation information. 5. The apparatus of claim 1, wherein the attribute is retrieved from a remote network device. 6. The apparatus of claim 1, wherein the first marker identifies at least one metric of the athletic performance data, and wherein a simultaneously provisioning of the first value of the attribute of the geographic location data and the at least one metric is provided via a user selection. 7. The apparatus of claim 6, wherein the at least one metric comprises pace and wherein generating the route map comprises: selecting different colors for different portions of the route map depending on the pace of the user during each of the different portions. 8. The apparatus of claim 1, wherein the first marker on the route map indicates a highest elevation of the route map and the second marker on the route map indicates a lowest elevation of the route map. 9. The apparatus of claim 1, wherein a difference between the first value of the attribute of the first marker at the first geographical location and the second value of the attribute of the second marker at the second geographic location is above a threshold value. 10. An apparatus comprising: a processor; andmemory storing computer readable instructions that, when executed, cause the apparatus to: receive, from a first sensor, geographic location data of a user during an athletic activity;identify at least a portion of a route on which the user is performing an athletic activity;determine at least one attribute of the at least one portion of the rou... | NIKE, Inc. | A63B024/00 A63B071/06 G01S019/19 G06F001/16 G06F019/00 G06F003/01 G01S019/48 G06F003/16 G09B019/00 A61B005/00 H04M001/725 G06Q050/00
US09775649 | B2 | 2017-10-03 | System and method for identifying a landmark | 2011-12-12 | A system (1010, 1110) for identifying a landmark is disclosed. The system includes a field generator (1016, 1116) for generating a magnetic field, an orthopedic implant (1030, 1130) located within the magnetic field, the implant having at least one landmark (1028, 1128), a removable probe (1029, 1129) with a first magnetic sensor (1026, 1126), a landmark identifier (1016, 1116) with a second magnetic sensor (1020, 1120) and a processor (1012, 1112) for comparing sensor data from the first and second sensor and using the set distance to calculate the position of the landmark identifier relative to the at least one landmark. The system allows for blind targeting of one or more landmarks. | 1. A probe for guiding placement of an identifier relative to an orthopaedic implant, comprising: an elongated element that includes an anti-rotation feature, the anti-rotation feature comprising a non-circular cross-sectional geometry of at least a portion of the elongated element, wherein the elongated element is configured to engage a stop at any of a plurality of points along a length of the elongated element,wherein the elongated element comprises a plurality of detents,wherein at least a portion of the elongated element has a longitudinal axis that has a natural curvature, andwherein the elongated element is configured to limit rotation of the elongated element relative to the stop when the elongated element is received through a slot in the stop; anda magnetic sensor coupled to the elongated element. 2. The probe of claim 1, wherein: the elongated element has a proximal end and a distal end;the magnetic sensor is coupled to the distal end of the elongated element; andthe elongated element is a flexible elongated element. 3. The probe of claim 1, wherein the anti-rotation feature comprises the elongated element having a cross-section perpendicular to the longitudinal axis that is at least one of rectangular, oval, or square. 4. The probe of claim 1, wherein the detents in the plurality of detents correspond to particular insertion depths of the elongated element within a channel of the orthopaedic implant. 5. The probe of claim 4, wherein one of the detents indicates a position on the elongated element corresponding to an appropriate depth of insertion of the elongated element to target a landmark of an orthopaedic implant having a particular length. 6. The probe of claim 1, wherein the elongated element has a longitudinal extent from the magnetic sensor to a proximal end where the elongated element terminates, and the non-circular cross-sectional geometry of the elongated element extends continuously along the entire longitudinal extent of the elongated element. 7. The probe of claim 1, wherein the anti-rotation feature is configured to limit rotation of the elongated element within a cannulation having a circular cross-section. 8. The probe of claim 1, wherein the orthopaedic implant is an intramedullary nail having a particular length, and at least one of the plurality of detents on the elongated element indicates a position on the elongated element corresponding to an appropriate depth of insertion of the elongated element to target a landmark of the intramedullary nail having the particular length. 9. The probe of claim 1, further comprising a memory device storing calibration data for the magnetic sensor. 10. The probe of claim 1, wherein at least one of the plurality of the detents is located at a position where a length of the elongated element between the stop and the magnetic sensor is dimensioned to place the magnetic sensor at a known placement relative to a landmark of the orthopaedic implant when the probe is inserted... | Smith & Nephew, Inc. | A61B017/17 A61B017/70 A61B005/06 A61B090/00
US09784678 | B2 | 2017-10-10 | Method for improving fluorescence image contrast | 2005-12-23 | There is provided an improved method for enhancing fluorescence images of an object, such as a biological tissue, by selectively eliminating or reducing unwanted fluorescence from fluorophores other than the fluorophore of interest. The method is based on the measurement of the lifetime of fluorophores while preserving information related to the fluorescence intensity of the fluorophore of interest. | 1. A method for improving optical imaging fluorescence contrast of a region of interest (ROI) of an object containing two or more fluorophore species, said method comprising: exciting the fluorophore species with a pulsed light source at a fluorescence excitation wavelength;collecting by a time-resolved photo detector temporal point spread function of the (ROI) of the object, temporal point spread function comprising time and amplitude information for a plurality of pixels wherein each pixel corresponds to a detection point of the time-resolved photo detector;generating by a computer a Continuous Wave fluorescence intensity image by performing a time integration of the collected temporal point spread function of the fluorescence signal for each pixel;improving, by the computer, contrast of each of the plurality of pixels of the Continuous Wave fluorescence intensity image by multiplying the Continuous Wave fluorescence intensity of each pixel by a weighting factor, the weighting factor being representative of a fraction contribution of each fluorophore decay time to the Continuous Wave fluorescence intensity image, the weighting factor being calculated by:fijk,CW=αijk⁢τijk∑k⁢αijk⁢τijkwhere: i,j represents each pixel; k is an index of the fluorophore species;αijk represents an amplitude of a kth fluorophore; andτijk represents a fluorescence lifetime of the kth fluorophore; andgenerating, using the computer, an optical image using the plurality of pixels with improved contrast. 2. The method as claimed in claim 1 further comprising: deriving by the computer a contribution fraction for at least one of said fluorophore species for each of the pixels; andmultiplying by the computer said weighted fluorescence intensity by said contribution fraction for each of the pixels. 3. The method as claimed in claim 2 wherein said contribution fraction is determined by modeling said fluorescence signal with a multi-exponential decay function by the computer. 4. The method as claimed in claim 1 wherein said Indicator function for each pixel is defined by boundaries which are function of said at least one predetermined fluorescence lifetime. 5. The method as claimed in claim 4 wherein said boundaries are also function of an error associated with said calculated fluorescence lifetime. 6. The method as claimed in claim 2 further comprising a step of: multiplying by the computer said fluorescence intensity of each pixel by a preliminary weighting factor prior to multiplying said fluorescence intensity by said weighting factor or said contribution fraction, whereby said preliminary weighting factor being an Indicator function of said calculated fluorescence lifetime and two predetermined fluorescence lifetimes corresponding to expected lifetimes of said two or more fluorophore species. 7. The method as claimed in any one of claims 1-3, 4-5 and 6 wherein said fluorescence species comprises a fluorophore that is distributed between a free state and a bound state. ... | SOFTSCAN HEALTHCARE GROUP LTD. | G01N033/50 G01N021/64 A61B005/00 G06T005/00
US09820684 | B2 | 2017-11-21 | Method and apparatus for a fluid sampling device | 2005-06-03 | Methods and apparatus are provided for manufacturing an analyte detecting device. In one embodiment, the apparatus comprises a housing; a penetrating member driver; a cartridge containing a plurality of penetrating members; a display on the cartridge; a linear slider on the housing, the slider coupled to a rod; and the rod moving with the slider, the rod having at least one roller. The device uses the linear motion of the slider to rotate the cartridge, punch open a new cavity and load a new penetrating member. | 1. A fluid sampling device comprising: a housing;a cartridge positioned in the housinga plurality of body fluid sensing strips integrated into a multi-strip roll;at least one penetrating member positioned in the cartridge;a gripper engaged to a penetrating member;a drive assembly coupled to the cartridge;a member for moving the cartridge to bring an unused penetrating member into position for launch, the member clearing the gripper and drive assembly to lift the gripper and drive assembly clear so that the cartridge can rotate; anda sterile blood barrier with an offset pad to maintain an air gap between a patients' finger and the sterile blood barrier, wherein the offset pad is a multi-layer laminate, the sterile blood barrier including a bend relief and one or more location features to position the sterile blood barrier. 2. The device of claim 1 wherein a cam surface is aligned parallel to a slider. 3. The device of claim 1 wherein a linear motion of a cam rotates the cartridge and moves a plunger to break the sterility barrier on the cartridge. 4. The device of claim 1 wherein a cam surface comprises a linear strip of maternal with at least two raised portions and two depressed portions. 5. A fluid sampling device comprising; a housing;a cartridge defining a plurality of cavities, said cartridge sized to fit within said housing; a plurality of body fluid sensing strips integrated into a multi-strip roll;a plurality of penetrating members at least partially contained in said cavities of the cartridge wherein each of a penetrating member is slidably movable, to extend outward from said cartridge to penetrate tissue; said cavities each having a lateral opening providing access to an elongate portion of the penetrating member;a gripper configured to be engaged with each of a penetrating member;a drive assembly coupled to the cartridge;a sterile blood barrier with an offset pad to maintain an air gap between a patients' finger and the sterile blood barrier, wherein the offset pad is a multi-layer laminate, the sterile blood barrier including a bend relief and one or more location features to position the sterile blood barrier; anda member for moving the cartridge to bring an unused penetrating member into position for launch, the member clearing the gripper and drive assembly to lift the gripper and drive assembly clear so that the cartridge can rotate. 6. A device comprising: a housing;a penetrating member driver;a cartridge containing a plurality of penetrating members;a plurality of body fluid sensing strips integrated into a multi-strip roll;a display on said cartridge;a linear slider on the housing, said slider coupled to a rod;said rod moving with said slider, said rod having at least one roller using a linear motion of the slider to rotate the cartridge, punch open a new cavity and load an unused penetrating member;a gripper configured to be engaged with each of a penetrating member;a driver coupled to the cartridge;a member for moving t... | SANOFI-AVENTIS DEUTSCHLAND GMBH | A61B005/00 A61B005/157 A61B005/151 A61B005/15
US09827389 | B2 | 2017-11-28 | Blood glucose regulation through control of breathing | 2009-03-05 | A respiratory treatment apparatus (1) provides blood glucose monitoring and breathing control based on detected blood glucose information. In an example embodiment, a flow generator provides a flow of breathable gas at a pressure above atmospheric to a patient interface according to a pressure treatment control protocol such as a CPAP, APAP, bi-level CPAP, etc. A detector determines a blood glucose condition indicator with one or more sensors that are used to sense physiological information. In response to signals from the sensors, a controller, such as a digital signal processor, controls adjustments to the flow of breathable gas provided by the flow generator. The adjustments are determined by the controller based on the detected blood glucose indicator and/or changes thereto. | 1. A device for treating sleep disordered breathing and diabetes comprising: at least one sensor detecting a physiological signal and transmitting physiological data to a controller; anda controller to control an air flow generator to provide pressurized air to a patient according to a first treatment protocol, and wherein said controller is adapted to determine a blood glucose level from the physiological data and to change to controlling the air flow generator to provide pressurized air according to a second treatment protocol based on the determined blood glucose level, wherein the first treatment protocol controls pressure to generally remain constant over a breathing cycle of the patient, and wherein the second treatment protocol controls the pressure to vary over the breathing cycle of the patient and comprises initiating target ventilation control. 2. The device of claim 1, wherein the controller compares the determined blood glucose level to one or more thresholds. 3. The device of claim 2 wherein the one or more thresholds comprises a minimum threshold value and a maximum threshold value. 4. The device of claim 3, wherein the minimum threshold value is within a range of about 2 to 3.5 mmol/L. 5. The device of claim 3, wherein the maximum threshold value is within a range of about 5.6 to 9 mmol/L. 6. The device of claim 3, wherein said controller activates an alarm, when blood glucose level is below the minimum or above the maximum threshold values. 7. A device for treating sleep disordered breathing: wherein said device includes a patient interface, an air flow generator configured to provide pressurized gas and operated by a controller and at least one sensor connected to the controller, wherein said controller determines glucose levels from a detected breathing pattern of a patient and then adjusts a target ventilation parameter of the air flow generator based on the determined glucose levels. 8. A method in a controller of a respiratory treatment device for treating sleep disordered breathing and diabetes, the method comprising: sensing a physiological signal of a patient;determining blood glucose level from the physiological signal; andchanging a first treatment protocol for controlling a flow generator to provide pressurized air to a second treatment protocol for controlling the flow generator to provide pressurized air based on the blood glucose level, wherein the first treatment protocol controls pressure to generally remain constant over a breathing cycle of the patient, and wherein the second treatment protocol controls the pressure to vary over the breathing cycle of the patient and comprises initiating target ventilation control. 9. The method of claim 8, wherein said method further includes analyzing the determined blood glucose level with respect to a minimum and maximum threshold value. 10. A respiratory treatment apparatus comprising: a flow generator to provide a flow of breathable air at a pressure above atmosph... | ResMed Limited | A61M016/00 A61B005/145 A61B005/00 G06F019/00
US09788747 | B2 | 2017-10-17 | Headgear with displaceable sensors for electrophysiology measurement and training | 2014-12-12 | A method and system provides for headgear usable for electrophysiological data collection and analysis and neurostimulation/neuromodulation or brain computer interface for clinical, peak performance, or neurogaming and neuromodulation applications. The headgear utilizes dry sensor technology as well as connection points for adjustable placement of the bi-directional sensors for the recoding of electrophysiology from the user and delivery of current to the sensors intended to improve or alter electrophysiology parameters. The headgear allows for recording electrophysiological data and biofeedback directly to the patient via the sensors, as well as provide low intensity current or electromagnetic field to the user. The headgear can further include auditory, visual components for immersive neurogaming. The headgear may further communication with local or network processing devices based on neurofeedback and biofeedback and immersive environment experience with balance and movement sensor data input. | 1. An apparatus for electrophysiology measurement, the apparatus comprising: a cross-bar having a first track disposed therein; anda plurality of sensors attached to the cross-bar via the first track for contacting engagement with a head of a user for receiving electrophysiology measurements therefrom. 2. The apparatus of claim 1 further comprising: a first side-bar in contacting engagement with the cross-bar; andat least one sensor attached to the first side-bar for contacting engagement with the head of the user. 3. The apparatus of claim 2 further comprising: at least one hinge mechanism engaging the first side-bar to the cross-bar, whereby the at least one hinge mechanism allows for the articulation of the first side-bar relative to the cross-bar. 4. The apparatus of claim 2 further comprising: a second side-bar in contacting engagement with the cross-bar. 5. The apparatus of claim 4 further comprising: at least one hinge mechanism engaging the first side-bar and the second side-bar to the cross-bar, whereby the at least one hinge mechanism allows for the articulation of the first side-bar and the second side-bar relative to the cross-bar. 6. The apparatus of claim 1 further comprising: at least one processing device receiving the electrophysiology measurements from the head of the user, the processing device in operative communication with an external processing system for electrophysiology measurements. 7. The apparatus of claim 6, wherein the operative communication is performed by at least one of: wireless communication and wired communication. 8. The apparatus of claim 1 further comprising: an audio receiving device operative to receive audio input from the user. 9. The apparatus of claim 1 further comprising: a lens operative to display a visual display thereon, wherein the lens is disposed outward from the cross-bar by an arm. 10. The apparatus of claim 1, wherein the sensors are dry sensors operative to perform transcranial current stimulation. 11. The apparatus of claim 10, wherein the transcranial stimulation includes at least one of: transcranial alternating current stimulation and transcranial direct current stimulation. 12. The apparatus of claim 1, further comprising: at least one speaker disposed in an earpiece, operative to provide audio output to the user in conjunction with the plurality of sensors receiving electrophysiology measurements. | EVOKE NEUROSCIENCE, INC. | A61N001/00 A61B005/0482 A61B005/0478 A61N002/00 A61B005/00 A61N001/36 A61N001/20 A61N001/04
US09801561 | B2 | 2017-10-31 | System and method for electrocardiogram analysis and optimization of cardiopulmonary resuscitation and therapy delivery | 2016-12-30 | The system and method provide for electrocardiogram analysis and optimization of patient-customized cardiopulmonary resuscitation and therapy delivery. An external medical device includes a housing and a processor within the housing. The processor can be configured to receive an input signal for a patient receiving chest compressions and to select at least one filter mechanism and to apply the filter mechanism to the signal to at least substantially remove chest compression artifacts from the signal. A real time dynamic analysis of a cardiac rhythm is applied to adjust and integrate CPR prompting of a medical device. Real-time cardiac rhythm quality is facilitated using a rhythm assessment meter. | 1. A system for treating a cardiac arrest patient, the system comprising: an EGG monitor; anda mechanical chest compression device configured to automatically deliver a series of chest compressions and ventilation pauses, wherein the ECG monitor is configured to filter the ECG signal to remove the compression artifact and produce a filtered ECG signal, wherein the ECG monitor is further configured to display the filtered signal when the mechanical chest compression device is delivering chest compressions and displays the unfiltered ECG signal during ventilation pauses, wherein the ECG monitor includes a filter configured to perform the filtering of the ECG signal by attenuating a plurality of regularly spaced bands, and wherein the ECG monitor includes a filter protection switch configured to switch to a protected mode during pauses in the chest compressions to protect the filter during the pauses in the chest compressions. 2. The system of claim 1, wherein an input to a numerator delay line of the filter is taken from an output of the numerator delay line when the filter protection switch is in the protected mode. 3. The system of claim 2, wherein the filter protection switch is further configured to switch away from the protected mode to allow the filter to operate normally responsive to a resuming of the chest compressions after the ventilation pauses. 4. The system of claim 3, wherein the input to the numerator delay line of the filter is the ECG signal when the filter protection switch is no longer in the protected mode. 5. The system of claim 1, wherein the filter comprises a comb filter. 6. The system of claim 1, wherein the ECG monitor is configured to identify the ventilation pauses. 7. The system of claim 6, wherein the ECG monitor is further configured to automatically switch from displaying the filtered ECG signal to displaying the unfiltered ECG signal responsive to identifying the ventilation pauses. 8. The system of claim 1, wherein an input to a numerator delay line of the filter is the ECG signal. 9. The system of claim 1, wherein the mechanical chest compression device is configured to transmit a signal to the ECG monitor indicative of the ventilation pauses. 10. The system of claim 1, wherein the ECG monitor is configured to change a quality factor (Q) of the filter responsive to an event. 11. The system of claim 10, wherein the event includes an attachment of a monitoring lead to the ECG monitor, a delivery of a shock, or a beginning of the delivery of chest compressions. 12. The system of claim 10, wherein the filter has a first value of Q for a first period of time and a second value of Q for a second period of time. 13. The system of claim 12, wherein the ECG monitor is configured to temporarily reduce the Q of the filter from the first value to the second value responsive to the event. 14. The system of claim 13, wherein the ECG monitor is further configured to restore the Q of the filter from the second value... | Physio-Control, Inc. | A61B005/04 A61B005/044 A61N001/39 A61B005/046 A61H031/00 A61B005/00
US09775506 | B2 | 2017-10-03 | System and method for detecting and classifying severity of retinal disease | 2013-07-10 | A method of detecting, and classifying severity of, a retinal disease using retinal images includes at least one of receiving, retrieving or generating reference data that includes information concerning occurrences of key image features for each of a plurality of retinal disease and disease severity conditions; receiving a retinal image of an individual; processing the retinal image of the individual to identify occurrences of each of a plurality of distinguishable image features throughout at least a region of interest of the retinal image; identifying which ones of the identified occurrences of the plurality of distinguishable image features of the retinal image of the individual correspond to the key image features of the reference data; calculating, based on the identifying, a number of occurrences of each of the key image features in the retinal image of the individual; and determining at least one of a likelihood of a presence of a retinal disease or a likelihood of developing a retinal disease based on a comparison of the number of occurrences of each of the key image features in the retinal image of the individual to the reference data. | 1. A method of detecting, and classifying severity of, a retinal disease using retinal images, comprising: at least one of receiving, retrieving or generating reference data comprising information concerning occurrences of key image features for each of a plurality of retinal disease and disease severity conditions;receiving a retinal image of an individual;processing said retinal image of said individual to identify occurrences of each of a plurality of distinguishable image features throughout at least a region of interest of said retinal image;identifying which ones of said identified occurrences of said plurality of distinguishable image features of said retinal image of said individual correspond to said key image features of said reference data;calculating, based on said identifying, a number of occurrences of each of said key image features in said retinal image of said individual; anddetermining at least one of a probability of a presence of a retinal disease or a probability of developing a retinal disease based on a comparison of said number of occurrences of each of said key image features in said retinal image of said individual to said reference data,wherein said processing said retinal image of said individual to identify occurrences of each of said plurality of distinguishable image features comprises applying a plurality of image operators to said retinal image. 2. The method of claim 1, further comprising determining a probability of a severity of said retinal disease based on said comparison of said number of occurrences of each of said key image features in said retinal image of said individual to said reference data. 3. The method of claim 1, wherein said plurality of image operators includes at least one of a Laplacian of Gaussian (LoG), Difference of Gaussian (DoG), Determinant of Hessian (DoH), SURF (Speeded Up Robust Features), SIFT (Scale Invariant Feature Transform), Morphological-based feature detector, Multi-scale feature detector, or Corner detector. 4. The method of claim 1, wherein said retinal image of said individual is an optical image from a fundus camera. 5. The method of claim 1, wherein said retinal image of said individual is an optical coherence tomography image. 6. The method of claim 1, further comprising receiving non-image information regarding said individual, and wherein said determining at least one of said probability of said presence of said retinal disease or said probability of developing said retinal disease is further based on said non-image information regarding said individual. 7. The method of claim 1, wherein said generating said reference data comprises generating said reference data based on processing a plurality of expert-classified retinal images that have been classified according to at least one of risk of developing a retinal disease, retinal disease type and retinal disease severity by an expert. 8. The method of claim 7, wherein said generating said reference data comprises... | THE JOHNS HOPKINS UNIVERSITY | A61B003/14 A61B003/00 A61B003/10 A61B003/12 A61B005/00 G06T007/00
US09810639 | B2 | 2017-11-07 | System for physiologic parameter examination and method for test strip recognition and interpretation | 2016-02-26 | A system for physiologic parameter examination includes a housing and a test-strip tray. The housing includes a carrying cavity disposed on a top surface of the housing and including a window. The carrying cavity is configured to receive a portable electronic device having an image-capturing component. The image-capturing component corresponds to the window. The test-strip tray is movably disposed in the housing and located below the carrying cavity. The test-strip tray is configured to carry at least one test strip. The image-capturing component is adapted to capture a test reaction of the test strip through the window. A method for test strip recognition and interpretation applicable to the system is also provided. | 1. A system for physiologic parameter examination, comprising: a housing comprising a carrying cavity, wherein the carrying cavity is disposed on a top surface of the housing and comprises a window, the carrying cavity is configured to receive a portable electronic device having an image capturing component, a bottom surface and a side surface of the portable electronic device are disposed in the carrying cavity, and the image capturing component corresponds to the window; anda test-strip tray movably disposed in the housing and located below the carrying cavity, wherein the test-strip tray is configured to receive at least one test strip, and the image capturing component is adapted to capture a test reaction of the test strip through the window. 2. The system for physiologic parameter examination as claimed in claim 1, wherein an inner wall of the test-strip tray is a black surface or covered with a light-absorbing coating. 3. The system for physiologic parameter examination as claimed in claim 1, wherein the test-strip tray and the housing jointly form a closed space. 4. The system for physiologic parameter examination as claimed in claim 1, wherein a number of the at least one test strip is plural, and the test strips are configured to examining a plurality of physiologic parameters respectively. 5. The system for physiologic parameter examination as claimed in claim 4, wherein colors of the test strips are different from each other. 6. The system for physiologic parameter examination as claimed in claim 4, wherein the test-strip tray comprises a plurality of positioning grooves configured to receive and position the test strips respectively. 7. The system for physiologic parameter examination as claimed in claim 6, wherein the positioning grooves comprise a plurality of first positioning grooves and a second positioning groove, the first positioning grooves are parallel to each other, and the second positioning groove intersects the first positioning grooves. 8. The system for physiologic parameter examination as claimed in claim 1, wherein the housing further comprises a connecting cavity disposed on the top surface and comprising an electrical connector, a bottom portion of the portable electronic device is configured to be inserted into the connecting cavity to be electrically connected to the electrical connector. 9. The system for physiologic parameter examination as claimed in claim 8, wherein the connecting cavity is located in the carrying cavity. 10. The system for physiologic parameter examination as claimed in claim 8, wherein when the portable electronic device is electrically connected to the system for physiologic parameter examination, the portable electronic device is configured to perform a data transmission with the system for physiologic parameter examination. 11. The system for physiologic parameter examination as claimed in claim 1, further comprising a plurality of device trays, wherein each of the device tray... | HTC Corporation | G01N021/00 G01N021/84 A61B005/0402 A61B005/021 A61B005/01 A61B005/00
US09775981 | B2 | 2017-10-03 | Medical connector | 2014-03-03 | A medical connector for use in a fluid pathway. A valve member with sealing rings helps preclude undesired accumulation of fluid within the connector. A branched connector includes a fluid diverter extending away from a port of the branched connector. The fluid diverter is configured to divert fluid flowing through the branched connector and into a medical connector attached thereto, flushing a distal portion of the medical connector. | 1. A three-way stopcock adapted for flushing a needleless connector on one port of the stopcock, the three-way stopcock comprising: a stopcock body comprising a first port, a second port, a third port, and a connecting region connecting the first port, the second port, and the third port, wherein the third port is positioned between the first and second ports;a fluid director positioned at least partially within the connecting region, the fluid director configured to selectively place one or more of the first port, the second port, and the third port in fluid communication with another of the first port, the second port, and the third port;a fluid diverter extending away from the connecting region at the third port and having a proximal end and a distal end positioned further from the third port than the proximal end; andthe needleless connector at the third port and at least partially surrounding the fluid diverter, the needleless connector comprising: a connector housing;a compressible seal positioned at least partially within the connector housing and having an interior cavity and a slit on a top of the seal that extends through the top and into the interior cavity;an internal projection member positioned at least partially within the compressible seal, the internal projection member having walls that define an internal cavity that encompasses the fluid diverter, an opening at a proximal end of the internal projection member, an interior height from the opening to a most distal surface of the walls that define the internal cavity, and at least one distal opening at a distal end of the internal projection member, the at least one distal opening having a proximal surface and a distal surface;wherein the fluid diverter is adjacent the walls of the internal cavity of the internal projection member to substantially bifurcate the internal cavity of the internal projection member at the proximal end of the internal projection member and wherein the fluid diverter has a distal tip that extends within the internal projection member to at least the proximal surface of the at least one distal opening and the distal tip is positioned proximal the distal surface of the at least distal opening. 2. The three-way stopcock of claim 1, wherein the fluid diverter bifurcates the internal cavity of the internal projection member at the proximal end of the internal projection member. 3. The three-way stopcock of claim 1, wherein the fluid diverter substantially bifurcates at least about half of the internal cavity of the internal projection member. 4. The three-way stopcock of claim 3, wherein the fluid diverter substantially bifurcates at least about three quarters of the internal cavity of the internal projection member. 5. The three-way stopcock of claim 1, wherein the compressible seal has a plurality of sealing rings on an interior surface thereof, the plurality of sealing rings configured to contact and seal against the internal projection member. 6. The... | ICU Medical, Inc. | A61B005/00 A61M039/22 A61B005/15
US09808377 | B2 | 2017-11-07 | Fluid extraction or filtration device, associated materials and methods | 2013-10-25 | A fluid extraction or filtration device for removing fluid from a body, the device comprising an array of microneedles for contacting fluid in said body and an absorbent gel matrix in fluid communication with said array of microneedles. There is further provided a combined fluid extraction and sampling device and methods employing the device(s), for example in determining the level of a target species in a sample of fluid, transdermal dialysis and renal replacement therapy. There is further provided an absorbent gel matrix for use in the treatment of uraemia and a haemodialysis filter comprising an absorbent gel matrix. Use of such a matrix is described in the context of methods including haemodialysis, transdermal dialysis and gastrointestinal dialysis. | 1. A method of treating fluid overload employing a device comprising an array of microneedles for contacting fluid in a body of a patient and an absorbent matrix in fluid communication with said array of microneedles, the method comprising contacting said fluid with said array of microneedles such that said fluid flows from said body to said absorbent matrix via said microneedles; andremoving fluid from the body of the patient at a rate of at least 10 ml/day and up to 1000 ml/day. 2. The method according to claim 1 wherein the fluid overload results from renal failure or cardiac failure. 3. The method according to claim 1 wherein said method treats salt and water overload in said patient. 4. The method according to claim 1 wherein fluid is removed from the body of the patient at a rate of at least 50 ml/day. 5. The method according to claim 1 wherein fluid is removed from the body of the patient at a rate of at least 100 ml/day. 6. The method according to claim 1 wherein the method further comprises use of a vacuum suction device to accelerate the flow of fluid from said body to said absorbent matrix via said microneedles. 7. The method according to claim 1 wherein the method further comprises using reverse iontophoresis to aid the flow of fluid from said body to said absorbent matrix via said microneedles. 8. The method according to claim 1 wherein the absorbent matrix is a gel. 9. The method according to claim 1 wherein an osmotic gradient is established between the fluid and the absorbent matrix such that said fluid flows, by osmosis, from said body to said absorbent matrix via said microneedles. 10. The method according to claim 1 wherein the absorbent matrix is dry prior to contacting the fluid. 11. The method according to claim 1 wherein the absorbent matrix exhibits an ability to entrap up to 30 times its own weight of water. 12. The method according to claim 1 wherein the absorbent matrix exhibits an ability to entrap water of up to 90% w/w. 13. The method according to claim 1 wherein said fluid comprises a target species, said patient being in need of having said target species removed from said fluid and said target species being retained in said absorbent matrix. 14. The method according to claim 13 wherein said target species is selected from the group consisting of water, a uraemic toxin, a metabolic product, a salt and an ion. 15. The method according to claim 13 wherein said target species is selected from the group consisting of Retinol Binding Protein, beta-2-Microglobulin, Parathyroid hormone, Adrenomedullin, Atrial Natriuretic Peptide, Asymmetric dimethylarginine, Indole-3-Acetic acid, Uric acid, Homocysteine, Creatine, Creatinine, P-Cresol, Oxalate, Urea and Phosphate. 16. The method according to claim 1 wherein the method is used for treatment of uraemia; and wherein said patient suffers from uraemia; and wherein said fluid that flows from said body to said absorbent matrix via said microneedles contains one or... | Renephra Limited | A61M001/00 A61F013/15 A61B005/145 A61B005/151 A61B005/00 A61M001/14 A61M037/00 A61B005/15 A61B005/1486
US09820667 | B2 | 2017-11-21 | Method and system for detecting heartbeat irregularities | 2013-10-17 | There is a method and system for detecting heartbeat irregularities comprising the steps of receiving a dataset representative of at least one waveform, the at least one waveform indicative of a subject's heart activity over a predetermined period of time; identifying from the data representative of at least one waveform, a plurality of peaks, each peak corresponding to a heartbeat; identifying from the predetermined period of time the time occurrence of each peak; calculating the difference (duration) between the time occurrence of each peak with its adjacent peak; determining the difference between each duration; classifying the absolute value of the difference into one of at least three intermediate categories; wherein each intermediate category comprises a specified range such that the absolute value is categorized into the intermediate category if it falls between the range; the intermediate categories further providing an indication of whether the subject has heartbeat irregularity. | 1. A method for detecting heartbeat irregularities comprising the steps of: a. receiving a dataset representative of at least one waveform, the at least one waveform indicative of a subject's heart activity over a predetermined period of time;b. identifying from the data representative of at least one waveform, a plurality of peaks, each peak corresponding to a heartbeat;c. identifying from the predetermined period of time the time occurrence of each peak;d. calculating the duration difference (Pn−(n+1)) between the time occurrence of each peak with its adjacent peak;e. determining the difference (Dn)between each duration difference (Pn −(n+1)) calculated in step d; andf. classifying the absolute value (¦Dn¦) of the difference (Dn) into one of at least three intermediate categories; wherein the first intermediate category of the at least three intermediate categories has a specified range of between 0 to 5; the second intermediate category of the at least three intermediate categories has a specified range of between 6 to 11; and the third intermediate category of the at least three intermediate categories has a specified range of 12 and above; andwherein if all of the classified absolute values ¦Dn¦occurs in the first intermediate category, with no occurrences in the second and third categories, the at least one waveform is categorized as ‘Regular heartbeat’; wherein if there are a number of occurrences ¦Dn¦ in all three Intermediate categories, the at least one waveform is categorized as ‘Irregularly irregular’ heartbeat; and wherein in all other cases the at least one waveform is categorized as ‘Regularly irregular’ heartbeat. 2. A method according to claim 1, wherein the waveform indicative of the subject's heart activity over a period of time is an arterial pulse waveform, an ECG signal or a time series of obtained camera frames captured based on variations in finger skin colour and brightness that occur due to blood pulsation. 3. A method according to claim2, wherein where the at least one waveform is an arterial pulse waveform, the peaks are determined and identified based on the identification of dicrotic notches as well as the gradient of the upstroke and downstroke identified on the at least one waveform. 4. A method according to claim 2, wherein where the at least one waveform is a time series of obtained camera frames, the method further comprises a conversion step before step c. 5. A method according to claim 4, wherein the conversion step includes the step of accounting for the discrepancy in sampling rates across different mobile devices using the following mathematical expression: f(HRn)=(60 seconds×S)/Δtn−(n+1) wherein f(HRn) denotes beats per minute of each heartbeat, S denotes the sampling rate of the captured waveform; and tn−(n+1) denotes time units in milliseconds between each peak. 6. A mobile device having camera and flash capabilities, the mobile device operable to obtain a time series of obtained camera frames captu... | HEALTHSTATS INTERNATIONAL PTE LTD | A61B005/04 A61B005/046 A61B005/021 A61B005/024 A61B005/0432 A61B005/0456 A61B005/0468 A61B005/00
US09814488 | B2 | 2017-11-14 | Surgical probe incorporating a dilator | 2013-03-15 | A surgical probe and a method for forming and enlarging an access opening through a psoas muscle to provide for minimally invasive lateral approach for surgical access to a lumber intervertebral disc. A distal end portion of the probe is equipped with an electrode useful for confirming proper location of the probe and includes an inflatable dilator body for enlarging an access opening through tissue adjacent to a spinal column. The probe includes a cannula through which a K wire can be extended to anchor the probe to a patient. | 1. A probe for use in a minimally invasive surgical procedure, comprising: (a) an elongate body including a longitudinally disposed cannula extending along a central longitudinal axis within the elongate body, the elongate body having a proximal end part and a distal end portion, the distal end portion in an initial condition having a first transverse dimension;(b) an expansible member associated with the distal end portion, and operable as a dilator, to expand the distal end portion radially to a larger second transverse dimension, wherein the expansible member expands uniformly circumferentially about the central longitudinal axis;(c) an electrode attached to and exposed on an exterior surface of the distal end portion of the cannula, wherein the electrode extends out of an opening formed on a distal most end of the elongate body; and(d) an electrical conductor interconnected with the electrode and extending along the elongate body through an annular space around an exterior surface of the cannula and to a location adjacent the proximal end part, the electrical conductor having a portion thereof available to receive an electrical signal,wherein the electrode is configured to receive the electrical signal and is disposed on a side of the distal end portion in order to determine a location of at least one nerve based upon the orientation of the electrode, andwherein the expansible member is configured to be positioned based upon the location of the at least one nerve to dilate tissue without damaging the at least one nerve. 2. The probe of claim 1 wherein the electrode extends radially along a tip of the distal end portion of the cannula. 3. The probe of claim 1 wherein the expansible member is a hollow flexible and elastic body surrounding the cannula and capable of radial expansion when a quantity of a fluid under pressure is introduced into the expansible member. 4. The probe of claim 1 wherein the cannula extends through the elongate body as a continuous tube, to the distal end. 5. The probe of claim 3 including a conduit communicating with an interior of the expansible member and extending from the expansible member along the main body and having a proximal end available to be connected to a source of fluid under pressure. 6. The probe of claim 1 wherein the cannula is formed of an electrically insulative material. 7. The probe of claim 1 wherein the electrical conductor is insulated. 8. The probe of claim 1 wherein a proximal end of the expansible member is tapered and attached to an exterior surface of the elongate body. 9. The probe of claim 1 wherein the expansible member is in the form of a balloon or bladder that is generally cylindrical in shape. 10. The probe of claim 1 wherein the expansible member has a first generally cylindrical shape when in a relaxed condition and a second generally cylindrical shape when inflated, the second generally cylindrical shape having a greater diameter than the first generally cylindrical sha... | GLOBUS MEDICAL, INC. | A61B017/34 A61B017/16 A61N001/05 A61B017/00 A61B017/32 A61B005/00
US09788750 | B2 | 2017-10-17 | Seizure prediction | 2007-04-30 | Seizure prediction systems and methods include measuring impedance within a brain of a patient to determine whether the brain is in a state indicative of a possibility of an onset of a seizure. In some embodiments, the measured impedance is compared to a predetermined threshold in order to determine whether the brain is in a state indicative of a possibility of a seizure. In other embodiments, a trend of the impedance measurements is correlated to a template. In other embodiments, a frequency component of a waveform of the impedance measurement amplitudes over time is correlated to frequency components of a template waveform. Upon detecting a state in which a seizure is likely to occur, a seizure indicator may be generated, which, in some embodiments, may be used to activate therapy delivery to the patient or, in other embodiments, activate an alarm. | 1. A method comprising: receiving, with a processor, impedance measurements indicative of an impedance of a brain of a patient, wherein the impedance of the brain is indicative of a physiological state of the brain;determining, with the processor, a trend in the impedance measurements over time;determining, with the processor, that the brain is in a state indicative of a possibility of a future seizure based on the trend; andcontrolling, with the processor, delivery of therapy to the patient based on determining that the brain is in the state indicative of the possibility of the future seizure. 2. The method of claim 1, further comprising measuring the impedance of the brain of the patient with an impedance sensing module, wherein measuring the impedance of the brain of the patient comprises delivering a stimulation current in a range of about 500 nanoamps to about 10 microamps across tissue in one or more regions of the brain of the patient. 3. The method of claim 1, further comprising, with the processor, comparing an amplitude of at least one of the impedance measurements to a predetermined threshold and determining that the brain is in the state indicative of the possibility of the future seizure based on the comparison. 4. The method of claim 3, wherein comparing the amplitude of the at least one of the impedance measurements to the predetermined threshold value comprises comparing an average amplitude of a plurality of the impedance measurements over a period of time to the predetermined threshold value. 5. The method of claim 1, further comprising: with the processor, comparing the trend to a template; andbased on the comparison of the trend to the template, determining the trend substantially correlates to the template, wherein determining that the brain is in the state indicative of the possibility of the future seizure comprises determining that the brain is in the state indicative of the possibility of a future seizure based on determining the trend substantially correlates to the template. 6. The method of claim 5, wherein determining the trend substantially correlates to the template comprises determining a first frequency component of the trend substantially correlates to a second frequency component of the template. 7. The method of claim 1, wherein the trend indicates a rate of change of the impedance measurements over time. 8. The method of claim 1, wherein controlling delivery of therapy comprises at least one of controlling delivery of a pharmaceutical agent or controlling delivery of electrical stimulation to tissue of the patient. 9. The method of claim 1, further comprising, with the processor, controlling the therapy to inhibit an onset of the future seizure. 10. The method of claim 1, further comprising, with the processor, controlling the therapy to reduce severity or duration of the future seizure. 11. The method of claim 1, wherein the state indicative of the possibility of a future seizure is a first state, t... | Medtronic, Inc. | A61B005/04 A61B005/053 A61B005/0478 A61B005/00 A61B005/024 A61B005/0402 A61B005/0488 A61N001/08 A61N001/36
US09777053 | B2 | 2017-10-03 | Methods, arrangements and systems for obtaining information associated with an anatomical sample using optical microscopy | 2007-02-08 | Arrangements and methods are provided for obtaining information associated with an anatomical sample. For example, at least one first electro-magnetic radiation can be provided to the anatomical sample so as to generate at least one acoustic wave in the anatomical sample. At least one second electro-magnetic radiation can be produced based on the acoustic wave. At least one portion of at least one second electro-magnetic radiation can be provided so as to determine information associated with at least one portion of the anatomical sample. In addition, the information based on data associated with the second electro-magnetic radiation can be analyzed. The first electro-magnetic radiation may include at least one first magnitude and at least one first frequency. The second electro-magnetic radiation can include at least one second magnitude and at least one second frequency. The data may relate to a first difference between the first and second magnitudes and/or a second difference between the first and second frequencies. The second difference may be approximately between −100 GHz and 100 GHz, excluding zero. | 1. An apparatus, comprising: at least one first arrangement configured to provide at least one first electro-magnetic radiation to a live biological sample so as to generate at least one acoustic wave in the biological sample, wherein at least one second electro-magnetic radiation is produced from the biological sample based on the at least one acoustic wave;at least one second arrangement configured to receive at least one portion of the at least one second electro-magnetic radiation;at least one third arrangement configured to determine information associated with a biomechanical property of at least one portion of the biological sample based on a measurement of at least one of a magnitude or frequency of a spectral peak in a spectrum of the at least one second electro-magnetic radiation; andan electro-magnetic radiation source configured to provide at least one third electro-magnetic radiation to the biological sample,wherein the at least one first electro-magnetic radiation includes at least one first frequency, wherein the at least one second electro-magnetic radiation includes at least one second frequency of the spectral peak, wherein the at least one third electro-magnetic radiation includes at least one third frequency, and wherein the second frequency is different from the first frequency and is different from the third frequency and is determined as a function of the first and third frequencies. 2. The apparatus according to claim 1, wherein the biological sample is at least one of an organ, a tissue or cell. 3. The apparatus, according to claim 1, further comprising at least one fourth arrangement configured to generate an image of the at least one portion of the biological sample as a function of data associated with the at least one second electro-magnetic radiation. 4. The apparatus according to claim 3, further comprising at least one fifth arrangement configured to translate the at least one first electro-magnetic radiation across the biological sample. 5. The apparatus according to claim 4, wherein the at least one fifth arrangement includes at least one lens, and wherein the at least one lens focuses the at least one first electro-magnetic radiation and collects the at least one second electro-magnetic radiation. 6. The apparatus according to claim 3, wherein the at least one first electro-magnetic radiation includes at least one first frequency, wherein the at least one second electro-magnetic radiation includes at least one second frequency, and wherein the data relates to a difference between the first and second frequencies. 7. The apparatus according to claim 6, wherein the second difference is between −100 GHz and 100 GHz, excluding zero. 8. The apparatus according to claim 1, wherein the at least one second arrangement includes a spectral filter which facilitates a determination of the spectrum of the at least one second electro-magnetic radiation. 9. The apparatus according to claim 4, wherein the at least one f... | The General Hospital Corporation | A61B005/00 C07K014/705 G01J003/44 G01N021/47 G01N021/63
US09784922 | B2 | 2017-10-10 | Electro-optical connector with hot electrical contact protection | 2014-07-18 | A cable connector, including a housing having a base and a lip, which surrounds the base and defines an aperture configured to receive a mating plug. The cable connector also includes a plurality of electrical contacts enclosed by the housing and configured to convey electrical signals, the electrical contacts having respective first proximal and first distal ends, the first proximal ends being implanted in the base so that the first distal ends are recessed within the aperture at a first distance from the base. The cable connector additionally includes one or more optical fiber terminals containing end portions of respective optical fibers configured to convey optical signals and having respective second proximal and second distal ends, the second proximal ends being implanted in the base so that the second distal ends are recessed within the aperture at a second distance from the base, which is greater than the first distance. | 1. A cable connector, comprising: a housing comprising a base, a lip that surrounds the base and defines an aperture configured to receive a mating plug, and a plurality of structures arranged in the aperture that protrude from the base;a plurality of electrical contacts enclosed by the housing and configured to convey electrical signals, the electrical contacts having respective first proximal and first distal ends, the first proximal ends being implanted in the base so that the first distal ends are recessed within the aperture at a first distance from the base; andone or more optical fiber terminals containing end portion of an optical fiber configured to convey optical signals;wherein the plurality of structures protrude to a second distance from the base, which is greater than the first distance, the plurality of structures comprising:a central structure disposed centrally within the plurality of electrical contacts;a plurality of structures regularly spaced around the central structure and arranged to partially surround the plurality of electrical contacts;at least one of the plurality of structures regularly spaced around the central structure containing the one or more optical fiber terminal;wherein the arrangement of the plurality of structures is such that when a finger enters the aperture it will press against at least one of the plurality of structures and thereby be prevented from touching the plurality of electrical contacts. 2. The cable connector according to claim 1, wherein the electrical signals are selected from a list consisting of radio frequency (RF) energy and sensor measurements. 3. The cable connector according to claim 1, wherein the optical signals are selected from a list consisting of optical radiation and data. 4. The cable connector according to claim 1, wherein the housing and the one or more optical fiber terminals comprise a non-conductive material. 5. The cable connector according to claim 1, wherein the one or more optical fiber terminals comprise a conductive material which is grounded. 6. The cable connector according to claim 1, wherein the one or more optical fiber terminals comprise one or more male optical fiber terminals, and wherein the plurality of electrical contacts comprise female contact sockets, and comprising a mating plug configured for insertion into the housing, the mating plug comprising a corresponding male contact pin for each of the plurality of female contact sockets, and a corresponding female optical fiber terminal for each of the one or more male optical fiber terminals. 7. The cable connector according to claim 6, wherein a combination of a given male optical fiber terminal and the corresponding female plug optical fiber terminal comprises an optical fiber connector selected from a list consisting of a ferrule connector, a biconical connector, an expanded beam connector and a multi-fiber connector. 8. A method for forming a connector, comprising: providing a housing comprising... | BIOSENSE WEBSTER (ISRAEL) LTD | G02B006/38 H01R009/03 H01R013/516 H01R024/00 A61B005/00 A61B018/14 H01R107/00 A61B018/00
US09795324 | B2 | 2017-10-24 | System for monitoring individuals as they age in place | 2014-12-05 | A computer-implemented method, and related system, for monitoring the wellbeing of an individual by providing eyewear that includes at least one sensor for monitoring the motion of the user. In various embodiments, the system receives data generated by the at least one sensor, uses the data to determine the user's movements using the received data, and compares the user's movements to previously established movement patterns of the user. If the system detects one or more inconsistencies between the user's current movements as compared to the previously established movement patterns of the user, the system may notify the user or a third party of the detected one or more inconsistencies. The system may similarly monitor a user's compliance with a medical regime and notify the user or a third party of the user's compliance with the regime. | 1. A computer-implemented method of monitoring the wellbeing of an individual, the method comprising the steps of: a. providing a user with computerized eyewear comprising at least one sensor for monitoring the motion of the user;b. receiving, by one or more processors, a pre-defined command from the user, for the at least one sensor to generate a first set of data identifying one or more user-defined movement patterns for the user;c. in response to receiving the pre-defined command, collecting, by one or more processors, via the at least one sensor, the first set of data;d. using, by one or more processors, the first set of data to determine one or more typical movement patterns of the user;e. generating, by one or more processors, an established one or more user-defined movement patterns for the user based on the first set of data generated by the at least one sensor after the user with the computerized eyewear provides the pre-defined command for the at least one sensor to generate the first set of data identifying the one or more user-defined movement patterns for the user;f. receiving, by one or more processors, a second set of data generated by the at least one sensor after the established one or more user-defined movement patterns have been generated;g. at least partially in response to receiving the second set of data generated by the at least one sensor, determining, by one or more processors, the user's movements using the received second set of data;h. at least partially in response to determining the user's movements, comparing, by one or more processors, the user's movements based on the second set of data to the previously established one or more user-defined movement patterns for the user;i. detecting, by one or more processors, one or more inconsistencies between the current user's determined movements as compared to the previously established one or more user-defined movement patterns;j. at least partially in response to detecting the one more inconsistencies, notifying, by one or more processors, at least one recipient of the one or more inconsistencies, where the at least one recipient is a recipient selected from a group consisting of: the user and a third party. 2. The computer-implemented method of claim 1, wherein the at least one sensor comprises at least one sensor selected from a group consisting of: a. a motion sensor;b. an accelerometer;c. a gyroscope;d. a geomagnetic sensor;e. a global positioning system sensor;f. an impact sensor;g. a microphone;h. a forward facing camera;i. a heart rate monitor;j. a pulse oximeter;k. a blood alcohol monitor;l. a respiratory rate sensor; andm. a transdermal sensor. 3. The computer-implemented method of claim 2, wherein the at least one sensor comprises at least one sensor selected from a group consisting of: a motion sensor, an accelerometer, a global positioning sensor, a gyroscope, and a forward facing camera. 4. The computer-implemented method of claim 2, wherein the method fur... | Vision Service Plan | A61B005/11 A61B005/00 A61B003/11 A61B005/0402 A61B005/0476 A61B005/103 A61B005/1171 A61B005/16 A61B007/04 G09B005/00 A61B005/1455 G06K009/00 G06K009/62 G08B021/04 A63B024/00 G09B005/06 G09B019/00 G06F019/00 A61F002/76 A61B005/0205 A61B005/024 A61B005/053 A61B005/08 G02C011/00
US09795298 | B2 | 2017-10-24 | System and method for transmitting, receiving and analysing partograph information | 2012-12-13 | There is provided a system of transmitting partograph information and analyzing the same comprising a client device adapted to receive partograph information as inputs; the inputs transmittable to a partograph processing server via a wireless communication network; wherein the partograph processing server is configured to receive the partograph information; conform the partograph information to a standardized format and disseminate the partograph information to at least one computer device. Preferably, the partograph processing server further comprises an analyzer to analyze the partograph information for one or more abnormality condition and provides an alert to the at least one computer device and the client device. Alternatively, the analyzer is installed in the client device in the form of a software application. | 1. A system for transmitting, receiving and analysing partograph information comprising: a mobile device adapted to receive partograph information as inputs; the partograph information transmittable to a partograph processing server via a communication network;wherein the partograph processing server is arranged to: access a database to receive profile information of a patient, conform the partograph information to a standardized partograph format, and disseminate the partograph information to at least one computer device, andwherein the partograph processing server comprises an analyser configured to analyse and use, from the partograph information transmittable to the partograph processing server, cervical dilation per unit time information and/or fetal station per unit time information as input for detection of at least one abnormality condition. 2. A system according to claim 1, wherein the database comprises security settings to maintain privacy of the patient. 3. A system according to claim 1, wherein the analyser comprises a detection algorithm configured to use cervical dilatation per unit time or fetal station per unit time as inputs for detection of at least one of the following abnormality conditions: Protracted active phase; Secondary arrest of cervical dilatation; prolonged deceleration phase; Failure of descent; protracted descent; and Arrest of descent. 4. A system according to claim 1, wherein the communication network is a telecommunication network. 5. A system according to claim 4, wherein the partograph information may be transmitted as Short Messaging Service (SMS) message(s), Unstructured Supplementary Service Data (USSD) messages, and/or via cellular data network such as GPRS, and 3G. 6. A system according to claim 1, wherein an alert in the form of a notification message is generated when the at least one abnormality condition is determined. 7. A system according to claim 6, wherein the alert is a SMS or USSD message. 8. A system for transmitting, receiving and analysing partograph information comprising: a mobile device adapted to receive partograph information as inputs; the partograph information transmittable to a partograph processing server via a communication network;wherein the partograph processing server is arranged to: receive profile information of a patient, conform the partograph information to a standardized partograph format, and disseminate the partograph information to at least one computer device, andwherein the partograph processing server comprises an analyser configured to analyse and use, from the partograph information transmittable to the partograph processing server, cervical dilation per unit time information and/or fetal station per unit time information as input for detection of at least one abnormality condition. | EINNOVATIONS HOLDINGS PTE. LTD. | A61B005/00 G06F019/00
US09826940 | B1 | 2017-11-28 | Optical tracking of heart rate using PLL optimization | 2015-05-26 | Systems and methods are provided for determining a pulse rate based on a photoplethysmographic measurement of blood in a portion of subsurface vasculature. A plurality of samples of the photoplethysmographic signal are obtained and a plurality of digital phase-locked loops are operated to generate respective oscillating signals based on the plurality of samples. Respective correlations between the plurality of samples and each of the generated oscillating signals are determined and used to select a subset of the plurality of phase-locked loops. The frequencies of the selected oscillating signals generated by the selected phase-locked loops are determined and used to determine a cardiovascular pulse rate. This method can be performed a plurality of times for a respective plurality of overlapping sets of the obtained plurality of samples to determine pulse rates over time. | 1. A method comprising: receiving a plurality of samples of a signal, wherein the signal is related to a cardiovascular pulse;generating, based on the plurality of samples, a plurality of oscillating signals using a plurality of digital phase-locked loops, wherein each digital phase-locked loop has a different respective set of operational parameters;comparing the plurality of samples to each of the generated oscillating signals;selecting two or more of the generated oscillating signals based on the comparison;determining a plurality of frequencies based on the selected oscillating signals, wherein each frequency corresponds to a respective one of the selected generated oscillating signals; anddetermining a pulse rate based on the determined plurality of frequencies. 2. The method of claim 1, wherein generating, based on the plurality of samples, the plurality of oscillating signals using the plurality of digital phase-locked loops comprises: receiving an artifact signal;filtering the received plurality of samples of the signal based on the artifact signal to generate a plurality of filtered samples; andgenerating, based on the plurality of filtered samples, the plurality of oscillating signals using the plurality of digital phase-locked loops. 3. The method of claim 2, wherein filtering the received plurality of samples of the signal based on the artifact signal comprises applying a least mean squares filter to the received plurality of samples, wherein the artifact signal is an input to the least mean squares filter, and wherein the plurality of filtered samples comprise a difference between the received plurality of samples of the signal and an output of the least mean squares filter. 4. The method of claim 1 wherein receiving a plurality of samples of a signal comprises: illuminating a portion of subsurface vasculature; andreceiving light emitted from the portion of subsurface vasculature in response to illumination, wherein the plurality of samples of a signal comprise measurements of the received light at a plurality of respective points in time. 5. The method of claim 1, further comprising: partitioning the plurality of samples of the signal into a plurality of overlapping subsets of the plurality of samples, wherein the plurality of overlapping subsets of the plurality of samples comprises a first subset of the plurality of samples and a second subset of the plurality of samples, and wherein the first subset of the plurality of samples and the second subset of the plurality of samples overlap; andwherein generating, based on the plurality of samples, a plurality of oscillating signals using the plurality of digital phase-locked loops comprises: generating, based on the first subset of the plurality of samples, a first plurality of oscillating signals using the plurality of digital phase-locked loops; and generating, based on the second subset of the plurality of samples, a second plurality of oscillating signals using the plurality of... | Verily Life Sciences LLC | A61B005/02 A61B005/00 A61B005/0205 A61B005/11 A61B005/024
US09775548 | B2 | 2017-10-03 | Heart rate sensor with high-aspect-ratio photodetector element | 2016-06-24 | Blood oxygenation sensors including high-aspect-ratio photodetector elements are discussed herein. Such high-aspect-ratio photodetector elements may provide improved signal-strength-to-power-consumption performance for blood oxygenation sensors incorporating such photodetector elements as compared with blood oxygenation sensors incorporating, for example, square photodetector elements. | 1. An apparatus comprising: a light source;one or more discrete, high-aspect-ratio (HAR) photodetectors, each HAR photodetector having a die with a photosensitive area and configured to provide an output signal that indicates the total amount of light that is incident on the photosensitive area of the die of the HAR photodetector, the photosensitive area of each HAR photodetector having: a first edge having a first length, anda first width in a direction perpendicular to the first edge; andcontrol logic, the control logic communicatively connected with the light source and each HAR photodetector and configured to: cause the light source to emit light,obtain one or more measured light intensity measurements from the one or more HAR photodetectors based on the output signal from each HAR photodetector, anddetermine a heart rate measurement based, at least in part, on the one or more light intensity measurements,wherein: the ratio of the first length to the first width of each photodetector element is substantially between 3:1 to 5:1,the light source is mounted on a surface of a substrate, andeach photosensitive area faces away from the surface of the substrate on which the light source is mounted. 2. The apparatus of claim 1, wherein: the first edge of the photosensitive area of each HAR photodetector, when viewed from a direction perpendicular to the substrate, is perpendicular to, and intersects with, a corresponding axis radiating out from a center of the light source and parallel to the substrate, andthe one or more HAR photodetectors and the light source are supported by the substrate. 3. The apparatus of claim 1, wherein: the first edge of the photosensitive area of each HAR photodetector, when viewed from a direction perpendicular to the substrate, is transverse to, and intersects with, a corresponding axis radiating out from a center of the light source and parallel to the substrate, andthe one or more HAR photodetectors and the light source are supported by the substrate. 4. The apparatus of claim 1, wherein: the light source includes a plurality of light-emitting devices. 5. The apparatus of claim 4, wherein: the plurality of light-emitting devices includes at least two light-emitting devices that predominantly emit light of different wavelengths from one another. 6. The apparatus of claim 4, wherein: there are a plurality of HAR photodetectors arranged in a pattern, andthe plurality of light-emitting devices is collocated at a center point of the pattern of HAR photodetectors. 7. The apparatus of claim 1, wherein the ratio of the first length to the first width of the photosensitive area of each HAR photodetector is substantially between 3:1 to 3.5:1. 8. The apparatus of claim 1, wherein the ratio of the first length to the first width of the photosensitive area of each photodetector is substantially between 3.5:1 to 5:1. 9. The apparatus of claim 1, wherein the photosensitive area of each HAR photodetector: has a first length be... | Fitbit, Inc. | A61B005/1455 A61B005/024 A61B005/00
US09801640 | B2 | 2017-10-31 | Lateral spondylolisthesis reduction cage | 2015-10-22 | An intervertebral fusion device comprising inferior and superior fusion cage devices that provide an ability to correct spondylolisthesis via in-situ adjustment. | 1. A method of installing an intervertebral fusion device, wherein the device comprises: a) an upper endplate having a proximal endwall, a distal endwall, an anterior wall, a posterior wall, an upper surface for contacting an upper vertebral body, teeth extending from the upper surface for enhancing securement, a lower surface, and a throughhole running from the upper surface to the lower surface and adapted for promoting fusion,b) a lower endplate having a proximal endwall, a distal endwall, an anterior wall, a posterior wall, a lower surface for contacting a lower vertebral body, teeth extending from the lower surface for enhancing securement, an upper surface, and a throughhole running from the lower surface to the upper surface and adapted for promoting fusion,c) a core component having i) an upper endportion engaged with the upper endplate and ii) a lower endportion engaged with the lower endplate, andwherein the proximal and distal endwalls define an insertion orientation,wherein the device has a collapsed configuration and an expanded configuration,wherein the upper endplate does not contact the lower endplate in the expanded configuration,wherein each of the upper endplate and the upper endportion of the core moves away from the lower surface of the lower endplate during expansion from the collapsed configuration to the expanded configuration,wherein the upper endplate moves away from the lower endplate in a height direction during expansion from the collapsed configuration to the expanded configuration;wherein the method comprises the steps of: inserting the device into a disc space in the insertion orientation and in an initial insertion direction, andexpanding the inserted device from the collapsed configuration to the expanded configuration, wherein the anterior and posterior walls of the upper endplate move in a direction that is substantially transverse to the initial insertion direction and substantially transverse to the height direction during expansion from the collapsed configuration to the expanded configuration. 2. The method of claim 1, wherein the distal endwalls of the upper and lower endplates define a taper for ease of insertion into the disc space. 3. The method of claim 1, wherein the device further comprises: d) a pivot at a proximal endportion of the device. 4. The method of claim 1, wherein an external geometry of the upper and lower endplates is lordotic. 5. The method of claim 1, wherein the upper surface of the upper endplate and the lower surface of the lower endplate are domed. 6. The method of claim 1, wherein the anterior walls of the upper and lower endplates have a convex curve and wherein inserting the device comprises aligning the convexly curved anterior walls with the convex shape of the anterior portion of the disc space. | DePuy Synthes Products, Inc. | A61F002/44 A61B017/16 A61B017/17 A61M029/02 A61F002/46 A61B005/00 A61B017/02 A61B017/34 A61M025/00 A61M025/01
US09808201 | B2 | 2017-11-07 | Sensors for prosthetic heart devices | 2015-08-13 | Prosthetic heart devices may be implanted into the heart with a sensor coupled to the device, the sensor being configured to measure physiological data, such as blood pressure, in the heart. Devices that may employ such sensors include prosthetic heart valves and occlusion devices, although sensor systems may be deployed in the heart separate from other implantable devices. The sensors may include a body with different configurations for attaching to the implantable device, such as apertures for sutures or fingers for connecting to structures of the implantable device. The sensors may provide data that allow a determination of aortic regurgitation or other information indicative of function of the implantable device and patient health during and after implantation of the device. | 1. A prosthetic heart valve system comprising: a prosthetic heart valve including: a stent extending along a longitudinal axis from an outflow portion to an inflow portion and having an expanded condition and a collapsed condition; anda valve assembly mounted to the stent; anda first sensor configured to measure physiological data, the first sensor including a body extending along a body axis and a plurality of apertures extending through the body on opposite sides of the body axis,wherein a first suture attaches the first sensor to the stent via at least some of the plurality of apertures so that, in the expanded condition of the stent, the longitudinal axis of the stent is not parallel to the body axis of the first sensor, and in the collapsed condition of the stent, the longitudinal axis of the stent is parallel to the body axis of the first sensor. 2. The prosthetic heart valve system of claim 1, further comprising: a second sensor configured to measure physiological data, the second sensor including a body and a plurality of apertures extending through the body of the second sensor and adapted to receive at least one suture therethrough for attaching the second sensor to the stent. 3. The prosthetic heart valve system of claim 2, wherein the first sensor is attached to the inflow portion of the stent and the second sensor is attached to the outflow portion of the stent with a second suture. 4. The prosthetic heart valve system of claim 2, wherein the first and second sensors are each configured to measure blood pressure and each include an induction coil disposed within the respective body and a capacitor in electrical communication with the induction coil. 5. The prosthetic heart valve system of claim 1, wherein the plurality of apertures includes four apertures arranged in a rectangular pattern on one end of the body. 6. The prosthetic heart valve system of claim 1, wherein the body of the first sensor includes two lateral projections, the plurality of apertures including at least one aperture extending through each of the two projections. 7. The prosthetic heart valve system of claim 6, wherein the plurality of apertures includes two apertures extending through each of the two projections. 8. The prosthetic heart valve system of claim 6, wherein the plurality of apertures includes at least two additional apertures extending through the body exclusive of the projections. 9. The prosthetic heart valve of claim 6, wherein the inflow portion of the stent includes a first circumferential row of cells, each cell being defined by a plurality of struts, the at least one aperture in each of the two projections being positioned on opposite sides of a selected strut of one of the cells. 10. The prosthetic heart valve of claim 9, wherein a fastener extends from the at least one aperture in a first of the two lateral projections to the at least one aperture in a second of the two lateral projections, the selected strut being positioned betwee... | St. Jude Medical, Cardiology Division, Inc. | A61F002/24 A61B005/00 A61F002/844 A61B005/0215 A61B005/02 A61B005/026 A61B017/00 A61B017/12
US09808205 | B2 | 2017-11-07 | Abnormality prediction device, abnormality prediction system, abnormality prediction method, biological information measuring device, biological information measuring system, and warning notification method | 2015-08-26 | An abnormality prediction device includes a detection unit that detects detection information including at least one of biological information of a user and body motion information regarding a body motion of the user, a premonitory symptom determination unit (control unit) that determines whether a premonitory symptom has occurred in the user on the basis of the detection information detected by the detection unit and a premonitory pattern which is a pattern of the detection information according to the premonitory symptom of abnormality occurring in the user, and an information output unit (control unit) that, when the premonitory symptom determination unit determines that the premonitory symptom has occurred, outputs notification information indicating that occurrence of abnormality in the user is predicted. | 1. An abnormality prediction device comprising: a detection unit that detects detection information including at least one of biological information of a user and body motion information regarding a body motion of the user;a premonitory symptom determination unit that determines whether a premonitory symptom has occurred in the user on the basis of the detection information detected by the detection unit and a premonitory pattern which is a pattern of the detection information according to the premonitory symptom of abnormality occurring in the user;an information output unit that, when the premonitory symptom determination unit determines that the premonitory symptom has occurred, outputs notification information indicating that occurrence of abnormality in the user is predicted;a premonitory pattern storage unit that stores the premonitory pattern;a premonitory pattern updating unit that updates the premonitory pattern stored in the premonitory pattern storage unit, on the basis of the detection information detected prior to an occurrence timing of the abnormality when the abnormality has occurred in the user; andan abnormality determination unit that determines whether abnormality has occurred in the user,wherein when the abnormality determination unit determines that abnormality has occurred, the premonitory pattern updating unit updates the premonitory pattern. 2. The abnormality prediction device according to claim 1, further comprising a notification unit that notifies the user of the notification information which is input from the information output unit. 3. The abnormality prediction device according to claim 1, further comprising an operation unit that receives an input operation from the user, wherein when an operation signal indicating that the abnormality has occurred is input from the operation unit, the premonitory pattern updating unit updates the premonitory pattern. 4. The abnormality prediction device according to claim 3, wherein the operation unit is configured to be able to input either an elapsed time since the abnormality has occurred in the user or an occurrence time of the abnormality, andwherein the premonitory pattern updating unit updates the premonitory pattern on the basis of the detection information detected before either a starting time of the elapsed time or the occurrence time which is input by the operation unit. 5. The abnormality prediction device according to claim 1, wherein the premonitory symptom determination unit determines that the premonitory symptom has occurred in the user when a pattern based on the detection information detected by the detection unit in a state where the abnormality has not occurred in the user is similar to the premonitory pattern rather than a normal pattern which is a pattern of the detection information according to a state where the premonitory symptom and the abnormality have not occurred in the user, and determines that the premonitory symptom has not occurred in the ... | SEIKO EPSON CORPORATION | G08B023/00 A61B005/00 G06F019/00 A61B005/044 A61B005/0402 A61B005/0456 A61B005/0476 A61B005/0488
US09820652 | B2 | 2017-11-21 | Multi-photon microscope having an excitation-beam array | 2014-12-08 | A two-photon imaging system capable of imaging of an image region in real time is presented. The imaging system comprises a source of excitation light that provides the excitation light as a plurality of laser beamlets. The plurality of laser beamlets is collectively scanned by a single-axis scanner along a first direction in the focal plane of the image region and oriented such that neither the rows nor columns are aligned with the first direction. As a result, each laser beamlet scans a different sub-region of the image region and the plurality of sub-regions are simultaneously scanned. As a result, the entirety of the image region is scanned in the same amount of time required to scan one image sub-region. | 1. An imaging system comprising: a source operative for providing a first plurality of light signals the first plurality of light signals being arranged such that they form a two-dimensional array of foci at an image region that defines a first plane;a single-axis scanner operative for linearly scanning the first plurality of light signals along a first direction in the first plane such that each light signal of the first plurality thereof interrogates a different one of a plurality of sub-image regions, wherein the first plurality of light signals is arranged as the two-dimensional array having a first periodicity in a second direction in the first plane and a second periodicity in a third direction in the first plane, and wherein the two-dimensional array is oriented such that the second direction and first direction form a first non-zero, acute angle, θ, within the first plane, and further wherein the plurality of sub-image regions are arranged in a two-dimensional arrangement within the image region; anda detector operative for detecting a second plurality of light signals, wherein each of the second plurality of light signals is generated in response to absorption of optical energy from at least one of the first plurality of light signals. 2. The imaging system of claim 1, wherein at least one of the first periodicity and second periodicity enables an image resolution for the imaging system that is less than or equal to two microns. 3. The imaging system of claim 1, wherein at least one of the first periodicity and second periodicity enables an image resolution for the imaging system that is less than or equal to one micron. 4. The imaging system of claim 1, wherein the two-dimensional array includes N first light signals in the second direction, and wherein θ=arctan(1/N). 5. The imaging system of claim 1 wherein the source comprises an ultrashort-pulsed regenerative fiber laser amplifier. 6. The imaging system of claim 1 wherein the first plurality of light signals have substantially the same intensity. 7. The imaging system of claim 1 wherein the detector comprises a camera and an image intensifier. 8. The imaging system of claim 1 wherein the imaging system is characterized by a point-spread function, and wherein the imaging system further comprises a processor that is operative for reconstructing an image of the first plane based on the position of a first light signal of the first plurality of light signals within its corresponding sub-region and the point-spread function. 9. A method for forming an image of an image region, the method comprising: providing a plurality of first light signals as a two-dimensional array having a first periodicity in a first direction in a first plane and a second periodicity in a second direction in the first plane;orienting the two-dimensional array such that the first direction and a third direction form a first non-zero, acute angle, θ, within the first plane;scanning the plurality of first lig... | The Board of Trustees of the Leland Stanford Junior University | A61B005/00 G02B021/00 G01N021/64
US09808180 | B2 | 2017-11-07 | Devices and methods for catheter localization | 2014-04-08 | A medical device can be localized by providing at least three non-colinear localization elements (e.g., electrodes) thereon. Once placed in a non-ionizing localization field, three adjacent localization elements, at least one of which will typically be a spot electrode, may be selected, and the non-ionizing localization field may be used to measure their locations. A cylinder is defined to fit the measured locations of the selected localization elements. The cylinder is rotationally oriented using the measured location of a spot electrode. Location and rotational attitude information may be used to construct a three-dimensional representation of the medical device within the localization field. The electrodes may be provided on the medical device or on a sheath into which the medical device is inserted. The invention also provides systems and methods for identifying and calibrating deflection planes where the medical device and/or sheath are deflectable. | 1. A system for localizing a medical device within a body wherein the medical device includes an elongate body, the system comprising: the elongate body having a diameter and at least three non-colinear localization electrodes mounted on the elongate body;a position determination processor configured to determine positions of the at least three non-colinear localization electrodes within a non-ionizing localization field; anda graphical representation modeling processor configured to: receive as input the determined positions of the at least three non-colinear localization electrodes; andconstruct a three-dimensional representation of the elongate body using the determined positions of the at least three non-colinear localization electrodes by defining a cylinder having a central axis connecting the determined positions of two non-colinear localization electrodes of the at least three non-colinear localization electrodes and a diameter equal to the diameter of the elongate body; anddisplay the three-dimensional representation of the elongate body in order to localize the medical device within the body. 2. The system according to claim 1, wherein the at least three non-colinear localization electrodes comprises at least three non-colinear spot electrodes. 3. The system according to claim 2, wherein the at least three non-colinear spot electrodes comprises at least three pairs of diametrically opposed spot electrodes. 4. The system according to claim 3, wherein each of the at least three pairs of diametrically opposed spot electrodes has a unique circumferential positioning. 5. The system according to claim 3, wherein each of the at least three pairs of diametrically opposed spot electrodes are in uniformly spaced along a length of the elongate body. 6. The system according to claim 1, wherein: the elongate body comprises an inner layer and an outer layer; andthe at least three non-colinear localization electrodes are sandwiched between the inner layer and the outer layer. 7. The system according to claim 6, wherein the inner layer is bonded to the outer layer. 8. The system according to claim 1, wherein the graphical representation modeling processor is further configured to construct the three-dimensional representation of the elongate body using the determined positions of the at least three non-colinear localization electrodes by rotationally orienting the cylinder using the determined position of a third non-colinear localization electrode of the at least three non-colinear localization electrodes. | St. Jude Medical, Atrial Fibrillation Divison, Inc. | A61B005/05 A61B005/06 A61M025/00 A61M025/01 A61B005/00 A61B008/12 A61B017/00 A61B034/20
US09795342 | B2 | 2017-10-24 | System and method for predicting acute cardiopulmonary events and survivability of a patient | 2016-08-03 | A method of producing an artificial neural network capable of predicting the survivability of a patient, including: storing in an electronic database patient health data comprising a plurality of sets of data, each set having at least one of a first parameter relating to heart rate variability data and a second parameter relating to vital sign data, each set further having a third parameter relating to patient survivability; providing a network of nodes interconnected to form an artificial neural network, the nodes comprising a plurality of artificial neurons, each artificial neuron having at least one input with an associated weight; and training the artificial neural network using the patient health data such that the associated weight of the at least one input of each artificial neuron is adjusted in response to respective first, second and third parameters of different sets of data from the patient health data. | 1. A patient survivability prediction system comprising: a patient worn device comprising a first input to receive a first set of parameters relating to heart rate variability data of a patient, anda second input to receive a second set of parameters relating to vital sign data of the patient;a memory module to store the first set of parameters and the second set of parameters;at least one processor coupled to the memory module and configured to analyze the first set of parameters and the second set of parameters and output a prediction on the survivability of the patient over the next 72 hours based on the first set of parameters and the second set of parameters; anda display coupled to the at least one processor to display the prediction on the survivability of the patient over the next 72 hours. 2. The patient survivability prediction system of claim 1, wherein at least one of the memory module, the at least one processor, and the display is situated at a location remote from the patient-worn device. 3. The patient survivability prediction system of claim 1, further comprising a third input to receive a third set of parameters including at least one of demographic information of the patient and medical history information of the patient, wherein the at least one processor is further configured to analyze the first set of parameters, the second set of parameters, and the third set of parameters, and output the prediction on the survivability of the patient over the next 72 hours based on the first set of parameters, the second set of parameters, and the third set of parameters. 4. The patient survivability prediction system of claim 3, wherein the at least one of the demographic information and the medical history information includes both the demographic information and the medical history information, and wherein the demographic information includes at least one of an age, a gender, and an ethnicity of the patient, and the medical history information includes at least one of a Glascow Coma Scale score, a history of any medical conditions including diabetes, high blood pressure, and myocardial infarction, and any specific medications the patient is using. 5. The patient survivability prediction system of claim 3, wherein the demographic information includes at least one of an age and a gender of the patient. 6. The patient survivability prediction system of claim 1, wherein the vital sign data includes at least one of a blood oxygen saturation level of the patient, a temperature of the patient, a respiration rate of the patient, a pulse rate of the patient, a systolic blood pressure of the patient, a diastolic blood pressure of the patient, a Glasgow Coma Scale score of the patient, and a pain score of the patient. 7. The patient survivability prediction system of claim 1, wherein the heart rate variability data includes at least one of time domain data, frequency domain data, and geometric domain data. 8. The patient survivability pred... | SINGAPORE HEALTH SERVICES PTE LTD. NANYANG TECHNOLOGICAL UNIVERSITY | A61B005/00 A61B005/145 G06N003/08 G06F019/00 G06N003/04 A61B005/0205 A61B005/0402 A61B005/01 A61B005/04 A61B005/0456 A61B005/1455 A61B005/024 A61B005/021 A61B005/08
US09788740 | B2 | 2017-10-17 | Flexible neural interfaces with integrated stiffening shank | 2016-07-13 | A neural interface includes a first dielectric material having at least one first opening for a first electrical conducting material, a first electrical conducting material in the first opening, and at least one first interconnection trace electrical conducting material connected to the first electrical conducting material. A stiffening shank material is located adjacent the first dielectric material, the first electrical conducting material, and the first interconnection trace electrical conducting material. | 1. A neural interface device adapted to be inserted into tissue, comprising: a neural interface body, wherein said neural interface body includesan insertion end anda body portion extending from said insertion end;an upper layer made of a first biocompatible dielectric material having at least one first opening for a first electrical conducting material;said upper layer extending over said insertion end and said body portion of said neural interface body wherein said at least one first opening for a first electrical conducting material is located in said upper layer extending over said insertion end;a first electrical conducting material in said first opening;an electrode positioned over said at least one first opening and over said first electrical conducting material in said first opening;at least one first interconnection trace electrical conducting material in said upper layer connected to said first electrical conducting material, wherein said at least one first interconnection trace electrical material is located in said upper layer extending over said insertion end and said body portion of said neural interface body;a lower layer made of a second dielectric material wherein said second dielectric material is biocompatible, wherein said lower layer extends under said insertion end and said body portion of said neural interface body; anda stiffening shank between said upper layer and said lower layer wherein said stiffening shank is adjacent said first dielectric material, said first electrical conducting material, and said first interconnection trace electrical conducting material, wherein said stiffening shank is made of a material that provides the neural interface device with sufficient stiffening that the neural interface device is adapted to be inserted into tissue, and wherein said stiffening shank is located only in said insertion end of said neural interface body. 2. A neural interface device adapted to be inserted into tissue of claim 1 wherein said first electrical conducting material in said first opening is made of metal. 3. A neural interface device adapted to be inserted into tissue of claim 1 wherein said first electrical conducting material in said first opening is made of a conducting polymer. | Lawrence Livermore National Security, LLC | A61B005/04 A61B005/00 A61N001/05 A61N001/36
US09792686 | B2 | 2017-10-17 | System and method for acquiring optoacoustic data and producing parametric maps using subband acoustic compensation | 2016-11-17 | A method is disclosed for generating sinograms by sampling transducers acoustically coupled with a surface of a volume after a pulse of light, each transducer being associated with a channel in an optoacoustic imaging system, and processing at least two multi-channel sinograms, each corresponding to a different one of the at least two different predominant wavelengths, to create at least two processed sinograms. The processing step includes a step of sub-band acoustic compensation. Image reconstruction is performed based upon the processed sinograms to generate at least two optoacoustic images. Image post processing is performed on the optoacoustic images to generate post-processed images. A parametric map is generated based upon information contained in the post-processed images, and an ultrasound image is generated using the transducers. The parametric map and the ultrasound image are co-registered and a co-registered image is displayed. | 1. A method for creating and displaying images in an optoacoustic imaging system, comprising: generating a plurality of multi-channel sinograms, each said channel generated by sampling a transducer acoustically coupled with a surface of a volume for a predetermined period of time after a pulse of light having a predominant wavelength selected from at least two different predominant wavelengths;processing at least two of the plurality of multi-channel sinograms using sub-band acoustic compensation, each of the plurality of multi-channel sinograms corresponding to a different one of the at least two different predominant wavelengths to create at least two processed sinograms, wherein the sub-band acoustic compensation comprises: filtering each of the at least two multi-channel sinograms to create a filtered representation of each of the sinograms;demodulating each said filtered representation of each of the sinograms in the frequency domain;applying a narrowband simplification to each said filtered representation of each of the sinograms to create a processed filtered representation of each of the sinograms, wherein the narrowband simplification is applied in the time domain to the demodulated filtered representations of each of the sinograms;creating two processed images from information contained in each of the processed sinograms; anddisplaying at least one of the two processed images. 2. The method as specified in claim 1, further comprising displaying both of the two processed images. 3. The method as specified in claim 1, wherein the applying a narrowband simplification comprises computing a time shift between two demodulated narrow-band signals in the time domain. 4. The method as specified in claim 1, wherein the filtering comprises using a filter bank of sub-band filters to define a plurality of sub-bands in the frequency domain. 5. The method as specified in claim 4, further comprising reassembling the filtered representation of each of the sinograms for each sub-band filter after applying the narrowband simplification. 6. The method as specified in claim 4, further comprising compensating at least one of the plurality of sub-bands based on a depth dependent compensation factor associated with the corresponding sub-band filter. 7. The method as specified in claim 6, further comprising identifying a range of frequencies for each of the plurality of sub-band filters. 8. The method as specified in claim 4, further comprising compensating each of the plurality of sub-bands for frequency dependent attenuation. 9. The method as specified in claim 8, wherein the compensation for at least one of the plurality of sub-bands is different from the compensation for at least one other of the plurality of sub-bands. 10. The method as specified in claim 1, further comprising perfoiiiiing image reconstruction based upon the at least two processed sinograms to generate at least two optoacoustic images. 11. The method as specified in claim 10, fur... | Seno Medical Instruments, Inc. | G06K009/00 G06T007/00 A61B005/00 A61B008/08 A61B008/00 A61B005/145 G06T011/00
US09814423 | B2 | 2017-11-14 | Method and system for monitoring pain of users immersed in virtual reality environment | 2015-08-26 | The present disclosure provides a method and system for monitoring a change in intensity of pain experienced by one or more users using a virtual reality environment. The method includes receiving one or more virtual content corresponding to the virtual reality environment on a virtual reality device, displaying the one or more virtual content corresponding to the virtual reality environment on the virtual reality device to the one or more users during one or more treatments provided to the one or more users, assessing the change in the intensity of pain experienced by the one or more users during the one or more treatments, generating a pain profile for each of the one or more users and determining a co-relation between the change in the intensity of pain experienced by the one or more users and the one or more virtual content viewed by the one or more users. | 1. A computer-implemented method for monitoring a change in intensity of pain experienced by one or more users using a virtual reality environment, the one or more users wearing a virtual reality device, the computer-implemented method comprising: receiving, with a processor, one or more virtual content corresponding to the virtual reality environment on the virtual reality device, wherein the virtual reality device is associated with a pain monitoring application;displaying, with the processor, the received one or more virtual content corresponding to the virtual reality environment on the virtual reality device to the one or more users during one or more treatments provided to the one or more users;assessing, with the processor, a change in an intensity of pain experienced by the one or more users during the one or more treatments, the assessing being performed on a set of pain monitoring scales using a plurality of bio-markers associated with the one or more users, the set of pain monitoring scales being determined and utilized based on an age category for each of the one or more users and the plurality of bio-markers are measured by a plurality of bio-sensors;generating, with the processor, a pain profile for each of the one or more users, the generated pain profile showing the change in the intensity of pain experienced by the one or more users and the pain profile for each of the one or more users is viewed by the one or more users on a corresponding plurality of communication devices; anddetermining, with the processor, a co-relation between the change in the intensity of pain experienced by the one or more users during the one or more treatments and the one or more virtual content viewed by the one or more users, the co-relation representative of an effect of the one or more virtual content on the intensity of pain of the one or more users; andstoring, with the processor, the one or more virtual content viewed by the one or more users, characteristics of the one or more users, the generated pain profile of the one or more users and the plurality of bio-markers associated with the one or more users. 2. The computer-implemented method as recited in claim 1, further comprising categorizing, with the processor, the one or more users based on the change in the intensity of pain of the one or more users and characteristics of the one or more users, wherein the characteristics of the one or more users utilized for the categorizing comprises age, gender, mentality, Body Mass Index, family history or behavior of the one or more users during the one or more treatments. 3. The computer-implemented method as recited in claim 1, wherein the one or more virtual content comprises one or more video content, one or more image content, one or more text content and one or more graphic content. 4. The computer-implemented method as recited in claim 1 , wherein the plurality of bio-markers comprises heart rate, heart rate variability (HRV), blood flow, blo... | - | A61B005/00 A61B005/0205 G06T019/00
US09808326 | B2 | 2017-11-07 | 3D dentofacial system and method | 2015-03-18 | A reference standard or reference system for dental feature location and orientation can include evaluating the facial appearance of one or more subjects based on aesthetic criteria and selecting a subset of subjects. The reference standard further includes constructing 3D representations, (e.g., a 3D virtual model), of the facial and dental features of the subset of subjects, from photographs of the face and mouth of each subject and determining the location and/or orientation of one or more dental features for each subject. The location and/or orientation values for each subject can be used to produce an average value and a standard deviation that forms the basis for a reference standard as part of reference system for evaluating patients. The method and system further includes constructing 3D representations (e.g., a 3D virtual model) of the facial and dental features of a patient from photographs of the face and mouth of the patient and determining the location and/or orientation of one or more dental features of the patient in order to compare them to the reference standard and develop a treatment plan for the subject based on differences between the patient's measurements and reference standard. The reference standard can use the pupils (e.g., in the natural head position orientation) as a landmark for registration and scaling of the reference standard to the patient under evaluation. | 1. A method for determining a reference standard diagnostic tool comprising: selecting a set of subjects meeting a first set of aesthetic criteria;determining a 3 dimensional (3D) representation of a head and mouth of each subject from two or more photographic images of the head and mouth of each subject, the 3D representation including an indication a central point of a left pupil and a right pupil of each subject;wherein the 3D representation of the head and mouth is determined by indexing a first photographic image to a second photographic image by marking an area in the first photographic image and searching for an area in the second photographic image having a substantially same curvature as the marked area in the first photographic image;identifying at least one dental feature and at least one facial feature in the 3D representation of the head and mouth of each subject;identifying at least one landmark on the at least one dental feature and at least one landmark on the at least one facial feature in the 3D representation of the head and mouth of each subject;defining in the 3D representation of the head and mouth of each subject a mid-axial plane extending horizontally through the central point of both pupils of each subject;defining in the 3D representation of the head and mouth of each subject a mid-coronal plane extending vertically through the central point of both pupils of each subject;defining in the 3D representation of the head and mouth of each subject a left sagittal plane extending vertically through the left pupil of each subject;defining in the 3D representation of the head and mouth of each subject a right sagittal plane extending vertically through the right pupil of each subject;determining a first dental landmark distance from at least one of the mid-axial plane, the mid-coronal plane, the left sagittal plane, and the right sagittal plane to the at least one landmark on the at least one dental feature; anddetermining a first facial landmark distance from at least one of the mid-axial plane, the mid-coronal plane, the left sagittal plane, and the right sagittal plane to the at least one landmark on the at least one facial feature. 2. The method according to claim 1 wherein the at least one dental feature includes a tooth and the at least one landmark is a point along a facial axis of a clinical crown of the tooth. 3. The method according to claim 1 wherein the at least one facial feature includes a location on the face of the subject and the at least one landmark is a point midway between the pupils of the subject. 4. The method according to claim 1 wherein the first dental landmark distance is determined as a perpendicular distance from one of the mid-axial plane, the mid-coronal plane, the left sagittal plane, and the right sagittal plane. 5. The method according to claim 1 wherein the first facial landmark distance is determined as a perpendicular distance from one of the mid-axial plane, the mid-coronal plane, the ... | President and Fellows of Harvard College | G06K009/00 A61C007/00 A61B005/00 A61B034/10
US09805166 | B1 | 2017-10-31 | Home healthcare management system and hardware and software | 2015-07-10 | A business method is described comprising doctorless testing, diagnosis, and treatment of patient diagnosed health maladies. Antithetical to the current regulatory and general public healthcare practices, in the doctorless method the patient has the equivalent of a clinical laboratory and physician's clinic at home. The herein disclosed microscope, blood tests, and physiological monitoring are patient selected, patient interpreted, and acted upon by the patient without physician support. The invention herein disclosed is a multi-element bundled comprehensive solution that lowers yearly primary healthcare costs to that of a cell phone, and will be purchased at retail stores such as Wal-Mart. | 1. A doctorless business method for self-collection of whole blood diagnostic information comprising: a digital microscope having x, y, and z axes automated movement means and wavelength limited resolution,said movement means moving a self-collected whole blood specimen, that is imaged by said digital microscope,said digital microscope collecting images directly from a camera chip that is an attached and integrated into said digital microscope,wherein said camera is not an iPhone camera,said images being of whole blood nuclei and constituents from said specimen,said image data from said collected images, processed to identify said nuclei and constituents,the diagnostic information from said processing of said image data provided to said specimen donor, andsaid specimen donor is not a physician. 2. A doctorless business method for self-collection of whole blood diagnostic information comprising: a digital microscope having x, y, and z axes automated movement means and wavelength limited resolution,said movement means moving a self-collected whole blood specimen, that is imaged by said digital microscope,said digital microscope collecting images directly from a camera that is an attached and integrated into said digital microscope,wherein said camera is not an iPhone camera,said images being of whole blood nuclei and constituents from said specimen,said image data from said collected images, processed to identify said nuclei and constituents,the diagnostic information from said processing of said image data provided to said specimen donor,said specimen donor is not a physician, andbilling for said information. | - | G06F019/00 G06K009/00 A61B005/15 A61B010/00
US09801956 | B2 | 2017-10-31 | PSMA-targeted NIR dyes and their uses | 2015-11-12 | The present disclosure relates to prostate specific membrane antigen (PSMA) targeted compounds conjugated to near-infra red (NIR) dyes and methods for their therapeutic and diagnostic use. More specifically, this disclosure provides compounds and methods for diagnosing and treating diseases associated with cells and/or vasculature expressing prostate specific membrane antigen (PSMA), such as prostate cancer and related diseases. The disclosure further describes methods and compositions for making and using the compounds, methods incorporating the compounds, and kits incorporating the compounds. | 1. A compound having the formula: B-X-Y-Z,wherein B comprises a compound capable of binding to prostate specific membrane antigen (PSMA),X is selected from the group consisting of polyethylene glycol (PEG), N-amino-dPEG2-acid and polyethylene amine (PEA),Y comprises at least one amino acid, or derivative thereof, andZ comprises a near-infra red (NIR) dye. 2. The compound of claim 1, wherein B is selected from the group consisting of a small molecule, ligand, inhibitor, agonist, and a derivative thereof. 3. The compound of claim 1, wherein B is 2-[-3-(1,3-dicarboxypropyl)-ureido]pentanedioic acid (DUPA) or derivative thereof. 4. The compound of claim 1, wherein X is a length from 7 atoms to 14 atoms. 5. The compound of claim 1, wherein Y comprises aromatic amino acids. 6. The compound of claim 1, wherein Z is selected from the group consisting of LS288, IR800, SP054, S0121, KODAK, S2076, S0456, 7. A compound having the formula: B-X-Y-Z,wherein B comprises DUPA or a derivative thereof;X is selected from the group consisting of, PEG, N-amino-dPEG2-acid, and PEA;Y is selected from the group consisting of phenylalanine-tyrosine, histidine-tyrosine, phenylalanine-arginine-tyrosine, and histidine-tyrosine;and Z comprises S0456. 8. A compound having the structural formula:or a pharmaceutically acceptable salt thereof, or isotopes thereof, wherein:R1 represents a hydrogen or SO3H;R2 represents a hydrogen, or CH3, or C3H6SO3−, or C3H6SO3H or C4H8SO3−, or C4H8SO3H or C3H6N+(CH3)3;R3, and R5 each represents a carbon, optionally one or more sharing bonds, or oxygen, or sulfur, or nitrogen R4 represents a carbon with optionally one or more sharing bonds;R6 represents nitrogen, oxygen, or sulfur or no atom (direct C—C bond between aromatic ring and vinyl ring);R7 is optional and when present represents electron donating aromatic substitution group;R8 is optional and when present represents linkers with aromatic amino acids such as Phe, Trp, His, Tyr, or derivative of them, and/or cationic amino acids such Arg, Lys, or derivative of them, and/or anionic amino acids such as Asp, Glu or derivative of them, and/or unnatural amino acids of aromatic/cationic/anionic acids or derivative;R9 represents a PEG, N-amino-dPEG2-acid, PEA, or derivative of them;R10 represents a CO2H, PO3H2, SO3H, CH2SO3H, CH2CONHCH2SO3H, CH2CONHCH2CH2SO3H;R11 represents CO2H, SO3H, CH2CONHCH2SO3H, CH2CONHCH2CH2SO3H; andR12 represents independently represents a hydrogen, a methyl group, CH2COOH, a CH2 and may optionally represent each a CH2 sharing a bond. 9. The compound of claim 8 wherein said compound is selected from the group consisting of:Chiral CenterCompound1*2*3*4*35LLLL114LLLD115LLDL116LLDD117LDLL118LDLD119LDDL120LDDD121DLLL122DLLD123DLDL124DLDD125DDLL126DDLD127DDDL128DDDDor racemic mixture thereof. 10. The compound of claim 1 wherein the amino acid is selected from the group consisting of tyrosine, cysteine, lysine or a derivative thereof, or phenylalanine-tyrosine, phenyla... | On Target Laboratories, LLC | A61B005/00 A61B008/00 A61B010/00 A61K049/00 C07D209/20 C07K005/062 C07K005/083 G01N021/64 G01N033/574 C07K005/087 C07K005/11 C07K007/06 C07K005/09 C07K005/078 C07K005/065 C07K005/072 A61K049/16
US09829398 | B2 | 2017-11-28 | Pressure sensor, microphone, ultrasonic sensor, blood pressure sensor, and touch panel | 2015-05-26 | According to an embodiment, a pressure sensor includes a substrate, a support part, a flexible membrane part, and a magnetoresistive element. The support part is adhered on the substrate by using a first adhesive material with a first Young's modulus and a second adhesive material with a second Young's modulus different from the first Young's modulus. The membrane part is supported by the support part. The magnetoresistive element is provided on the membrane part, and includes a first magnetic layer, a second magnetic layer, and a spacer layer provided between the first magnetic layer and the second magnetic layer. | 1. A pressure sensor comprising: a membrane comprising a peripheral edge, the peripheral edge including a first edge portion, a second edge portion, a third edge portion and a fourth edge portion, the first edge portion being opposed to the second edge portion, the third edge portion being opposed to the fourth edge portion;a support member supporting a part of the peripheral edge of the membrane;a first adhesive member comprising a first adhesive material with a first Young's modulus;a second adhesive member comprising the first adhesive material, the first adhesive member and the second adhesive member being arranged in a first direction;a third adhesive member comprising a second adhesive material with a second Young's modulus;a fourth adhesive member comprising the second adhesive material, the third adhesive member and the fourth adhesive member being arranged in a second direction, the second direction intersecting with the first direction;a substrate, the first adhesive member being provided between the first edge portion and the substrate, the second adhesive member being provided between the second edge portion and the substrate, the third adhesive member being provided between the third edge portion and the substrate, the fourth adhesive member being provided between the fourth edge portion and the substrate; anda magnetoresistive element provided on the membrane part, the magnetoresistive element including a first magnetic layer, a second magnetic layer, and a spacer layer provided between the first magnetic layer and the second magnetic layer. 2. The sensor according to claim 1, wherein a length of the first edge portion and the second edge portion is shorter than a length of the third edge portion and the fourth edge portion, respectively, and the second Young's modulus is lower than the first Young's modulus. 3. The sensor according to claim 1, wherein the magnetoresistive element is provided on the third edge portion. 4. The sensor according to claim 3, comprising a plurality of magnetoresistive elements provided on the third edge portion. 5. A pressure sensor comprising: a membrane comprising a peripheral edge, the peripheral edge including a first edge portion, a second edge portion, a third edge portion and a fourth edge portion, the first edge portion being opposed to the second edge portion, the third edge portion being opposed to the fourth edge portion;a support member supporting a part of the peripheral edge of the membrane;a first adhesive member comprising a first adhesive material with a first coefficient of thermal expansion;a second adhesive member comprising the first adhesive material, the first adhesive member and the second adhesive member being arranged in a first direction;a third adhesive member comprising a second adhesive material with a second coefficient of thermal expansion;a fourth adhesive member comprising the second adhesive material, the third adhesive member and the fourth adhesive member being ar... | Kabushiki Kaisha Toshiba | G01H011/00 G01L001/22 A61B005/021 A61B005/00 A61B005/05 G01L009/00 G01L001/12
US09820674 | B2 | 2017-11-21 | Method and apparatus for determining cardiac performance in a patient with a conductance catheter | 2016-06-27 | An apparatus for determining cardiac performance in the patient includes a plurality of electrodes adapted to be placed in the patient and in communication with a heart chamber of the patient for measuring conductance and blood volume in the heart chamber of the patient; and a processor for determining instantaneous volume of the ventricle based on measurement of complex admittance at a plurality of frequencies to identify mechanical strength of the chamber, the processor in communication with the electrodes. A method for determining cardiac performance in a patient. | 1. An apparatus for determining cardiac performance in a patient comprising: a plurality of electrodes adapted to be placed in the patient and in communication with a heart chamber of the patient for measuring conductance and blood volume in the heart chamber of the patient; anda processor for determining instantaneous volume of the ventricle based on measurement of complex admittance at a plurality of frequencies to identify mechanical strength of the chamber, the processor in communication with the electrodes. 2. An apparatus as described in claim 1 wherein the processor produces a plurality of desired wave forms at desired frequencies for the electrodes. 3. An apparatus as described in claim 2 wherein the processor produces the plurality of desired wave forms at desired frequencies simultaneously, and the processor separates the plurality of desired wave forms at desired frequencies the processor receives from the electrodes. 4. An apparatus as described in claim 3 wherein the plurality of electrodes is disposed on a conductance catheter to measure a least one segmental volume of the heart chamber. 5. An apparatus as described in claim 4 wherein the pressure sensor is in contact with the conductance catheter to measure ventricular pressure in the chamber. 6. An apparatus as described in claim 5 including a pressure sensor for measuring instantaneous pressure of the heart chamber in communication with the processor. 7. An apparatus as described in claim 6 wherein the plurality of electrodes includes intermediate electrodes to measure and instantaneous voltage signal from the heart, and outer electrodes to which a current is applied from the processor. 8. An apparatus as described in claim 7 wherein the pressure sensor is disposed between the intermediate electrodes and the outer electrodes. 9. An apparatus as described in claim 8 wherein the processor includes a computer with a signal synthesizer which produces the plurality of desired wave forms at desired frequencies and a data acquisition mechanism for receiving and separating the plurality of desired wave forms at desired frequencies. 10. A method for determining cardiac performance in a patient comprising the steps of: measuring conductance and blood volume in a heart chamber of the patient with a plurality of electrodes disposed in the patient and in communication with the heart chamber; anddetermining instantaneous volume of the ventricle based on measurement of complex admittance at a plurality of frequencies to identify mechanical strength of the chamber with a processor, the processor in communication with the electrodes. 11. A method as described in claim 10 including the step of producing a plurality of desired wave forms at desired frequencies for the electrodes disposed on a conductance catheter. 12. A method as described in claim 11 wherein the producing step includes the step of producing the plurality of desired wave forms at desired frequencies simultaneously, and i... | Board of Regents, The University of Texas System | A61B005/053 A61B005/0215 A61B005/029 A61B005/02
US09808538 | B2 | 2017-11-07 | PSMA-targeted NIR dyes and their uses | 2015-11-10 | The present disclosure relates to prostate specific membrane antigen (PSMA) targeted compounds conjugated to near-infra red (NIR) dyes and methods for their therapeutic and diagnostic use. More specifically, this disclosure provides compounds and methods for diagnosing and treating diseases associated with cells and/or vasculature expressing prostate specific membrane antigen (PSMA), such as prostate cancer and related diseases. The disclosure further describes methods and compositions for making and using the compounds, methods incorporating the compounds, and kits incorporating the compounds. | 1. A method of optical imaging of a biological tissue that expresses PSMA, the method comprising: (a) contacting the biological tissue with a composition a compound having the formula: B—X—Y—Z, wherein B comprises a compound capable of binding to prostate specific membrane antigen (PSMA),X, is selected from the group consisting of polyethylene glycol (PEG), N-amino-dPEG2-acid and polyethylene amine (PEA),Y comprises at least one amino acid, or derivative thereof, andZ comprises a near-infra red (NIR) dye,(b) allowing time for the compound in the composition to distribute within the biological target;(c) illuminating the tissue with an excitation light of a wavelength absorbable by the compound; and(d) detecting the optical signal emitted by the compound. 2. The method of claim 1, wherein the signal emitted by the compound is used to construct an image. 3. The method of claim 1, wherein the biological tissue is selected from the group consisting of prostate cells, prostate cancer cells, bladder cancer cells, pancreatic cancer cells, liver cancer cells, lung cancer cells, kidney cancer cells, sarcoma cells, breast cancer cells, brain cancer cells, neuroendocrine carcinoma cells, colon cancer cells, testicular cancer cells and melanoma cells. 4. A method of optical imaging of a biological tissue that expresses PSMA, the method comprising: (a) contacting the biological tissue with a composition a compound having the structural formula:or a pharmaceutically acceptable salt thereof, or isotopes thereof, wherein:R1 represents a hydrogen or SO3H;R2 represents a hydrogen, or CH3, or C3H6SO3−, or C3H6SO3H or C4H8SO3−, or C4H8SO3H or C3H6N+(CH3)3;R3, and R5 each represents a carbon, optionally one or more sharing bonds, or oxygen, or sulfur, or nitrogenR4 represents a carbon with optionally one or more sharing bonds;R6 represents nitrogen, oxygen, or sulfur or no atom (direct C—C bond between aromatic ring and vinyl ring);R7 is optional and when present represents electron donating aromatic substitution group;R8 is optional and when present represents linkers with aromatic amino acids such as Phe, Trp, His, Tyr, or derivative of them, and/or cationic amino acids such Arg, Lys, or derivative of them, and/or anionic amino acids such as Asp, Glu or derivative of them, and/or unnatural amino acids of aromatic/cationic/anionic acids or derivative;R9 represents a PEG, N-amino-dPEG2-acid, PEA, or derivative of them;R10 represents a CO2H, PO3H2, SO3H, CH2SO3H, CH2CONHCH2SO3H, CH2CONHCH2CH2SO3H;R11 represents CO2H, SO3H, CH2CONHCH2SO3H, CH2CONHCH2CH2SO3H; andR12 represents independently represents a hydrogen, a methyl group, CH2COOH, a CH2 and may optionally represent each a CH2 sharing a bond,(b) allowing time for the compound in the composition to distribute within the biological target;(c) illuminating the tissue with an excitation light of a wavelength absorbable by the compound; and(d) detecting the optical signal emitted by the compound. 5. The method of c... | On Target Laboratories, LLC | A61B005/00 A61B008/00 A61B010/00 A61K049/00 C07D209/20 C07K005/062 C07K005/083 G01N021/64 G01N033/574 C07K005/087 C07K005/11 C07K007/06 C07K005/09 C07K005/078 C07K005/065 C07K005/072 A61K049/16
US09788787 | B2 | 2017-10-17 | Patient-mounted micro vein enhancer | 2015-05-06 | The present invention is a Miniature Vein Enhancer that includes a Miniature Projection Head. The Miniature Projection Head may be operated in one of three modes, AFM, DBM, and RTM. The Miniature Projection Head of the present invention projects an image of the veins of a patient, which aids the practitioner in pinpointing a vein for an intravenous drip, blood test, and the like. The Miniature projection head may have a cavity for a power source or it may have a power source located in a body portion of the Miniature Vein Enhancer. The Miniature Vein Enhancer may be attached to one of several improved needle protectors, or the Miniature Vein Enhancer may be attached to a body similar to a flashlight for hand held use. The Miniature Vein Enhancer of the present invention may also be attached to a magnifying glass, a flat panel display, and the like. | 1. A portable vein enhancing device, for use in imaging of one or more veins beneath a target surface on a patient for a venipuncture procedure, said portable vein enhancing device comprising: a vein imager, said vein imager comprising: a first laser configured to emit a beam of light at a first wavelength;a scanner configured to transmit said beam of light from said first laser in a pattern onto the target surface;a photo detector configured to detect said first wavelength of light after being reflected from the target surface as a contrasted vein image; said photodetector further configured to output a signal representing said contrasted image;a second laser configured to receive said signal from said photodetector and to use said signal to emit a beam of light at a second wavelength, said scanner further configured to transmit said beam of light emitted at said second wavelength from said second laser to project said contrasted image onto the target surface; anda base member comprising a transparent material and extending downward from said vein imager, said vein imager configured to be mounted to a top portion of said base member and to emit and project said first and second beams of light downward onto the target area within a perimeter of said base member, a bottom surface of said base member comprising a first curved contour on a first side and a second curved contour on a second side opposite said first side, said first and second curved contours of said bottom surface respectively configured to be received upon the arm of the patient to support said vein imager at a position to overlie the target area for imaging and projecting thereon, said bottom surface comprising a third side, said third side having a concave contour configured to provide access to the target area. 2. The portable vein enhancing device according to claim 1, further comprising: A piece of hook material fixedly secured to said first side of said base; andA strap having a first end fixedly secured to said second side of said base, said strap comprising a piece of loop material fixedly secured thereto; said strap thereby configured to releasably secure said vein enhancing device to the arm of the patient using said hook and loop materials. 3. The portable vein enhancing device according to claim 2, wherein said transparent material of said base member comprises a clear plastic. 4. The portable vein enhancing device according to claim 2, wherein said transparent material of said base member comprises glass. 5. The portable vein enhancing device according to claim 2, further comprising a control panel configured with means for providing visual enhancements to said projected image. 6. The portable vein enhancing device according to claim 5, further comprising a combiner configured to combine said beam of light from said first laser and said beam of light from said second laser into a single beam of light; and wherein said scanner is configured to scan said combined fir... | AccuVein, Inc. | A61B006/00 A61B005/00 A61M005/42 A61M005/46 A61B017/34 A61B019/00 A61B090/00
US09820665 | B2 | 2017-11-21 | Remote interfacing of extended wear electrocardiography and physiological sensor monitor | 2016-09-02 | A remotely-interfaceable extended wear electrocardiography and physiological sensor monitor recorder is provided. A sealed housing forms on a bottom surface a cavity shaped to accommodate an upward projection of a battery compartment formed on a non-conductive receptacle of a disposable extended wear electrode patch and includes a set of electrical contacts that protrude from the bottom surface and are arranged in alignment with electrical pads provided on the non-conductive receptacle. Electronic circuitry is provided within the sealed housing and includes an externally-powered micro-controller operable to execute under micro programmable control. An electrocardiographic front end circuit is electrically interfaced to the micro-controller. A wireless transceiver electrically interfaces with the micro-controller and externally-powered flash memory is electrically interfaced with the micro-controller and operable to store samples of the electrocardiographic signals. | 1. A remotely-interfaceable extended wear electrocardiography and physiological sensor monitor recorder, comprising: a sealed housing forming on a bottom surface a cavity shaped to accommodate an upward projection of a battery compartment formed on a non-conductive receptacle of a disposable extended wear electrode patch and comprising a set of electrical contacts that protrude from the bottom surface and are arranged in alignment with electrical pads provided on the non-conductive receptacle;a patient-operable tactile-feedback button provided on a top surface of the sealed housing; andelectronic circuitry comprised within the sealed housing, comprising: an externally-powered micro-controller operable to execute under micro programmable control;an electrocardiographic front end circuit electrically interfaced to the micro-controller and operable to sense electrocardiographic signals through electrocardiographic electrodes provided on the disposable extended wear electrode patch;a wireless transceiver electrically interfaced with the micro-controller and operable to wirelessly interface with an external wireless-enabled device to communicate samples of the electrocardiographic signals; andexternally-powered flash memory electrically interfaced with the micro-controller and operable to store the samples of the electrocardiographic signals. 2. An electrocardiography and physiological sensor monitor recorder according to claim 1, further comprising at least one of: a seal coupling circumferentially surrounding the set of electrical contacts. 3. An electrocardiography and physiological sensor monitor recorder according to claim 2, wherein the seal coupling securely mates with a moisture-resistant seal formed on the non-conductive receptacle. 4. An electrocardiography and physiological sensor monitor recorder according to claim 1, wherein a battery is provided in the battery compartment of the non-conductive receptacle of the extended wear electrode patch and powers the electronic circuitry. 5. An electrocardiography and physiological sensor monitor recorder according to claim 1, further comprising: a feedback buzzer utilized by the microcontroller. 6. An electrocardiography and physiological sensor monitor recorder according to claim 1, further comprising: an external physiological sensor provided with the circuitry and comprising at least one of an SpO2 sensor, blood pressure sensor, temperature sensor, respiratory rate sensor, glucose sensor, airflow sensor, and volumetric pressure sensor. 7. An electrocardiography and physiological sensor monitor recorder according to claim 1, wherein the sealed housing comprises a rounded isosceles trapezoidal-like shape. 8. An electrocardiography and physiological sensor monitor recorder according to claim 1, further comprising: a least one of a retention detent and a tension detent each molded along a top edge of a top surface of the sealed housing. 9. An electrocardiography and physiological sensor moni... | Bardy Diagnostics, Inc. | A61B005/00 A61B005/0432 A61B005/0408 A61B005/0205 A61B005/145 A61B005/021 A61B005/03 A61B005/0452 A61B005/11 A61B005/1455 A61B005/08 A61B005/087
US09791539 | B2 | 2017-10-17 | System and method for multi-level border control within sites | 2016-07-15 | A system and method for keeping restricted objects within a confined area that forms part of an institutional site. The confined area having lockable exits, a safe area where the restricted objects are designated to stay, a buffer area adjacent to the first area, and a lock area situated between the buffer area and the lockable exits. The system detects whether the restricted object is in the buffer area or in the lock area and checks whether the restricted object is authorized to leave the safe area. Upon recognizing that the restricted object is in the buffer area and is not authorized to leave the safe area, the system sends a message to a tag of a staff member, instructing to move the restricted object back to the safe area; upon recognizing that the restricted object is in the lock area and is not authorized to leave the safe area, the system sends a locking signal to lockable exits. | 1. A system for keeping a restricted object within a confined area that forms part of an institutional site, the confined area having: (i) one or more lockable exits which can be remotely locked to prevent exit of the restricted object from said confined area, (ii) a safe area where the restricted object is designated to stay, (iii) a buffer area adjacent to the safe area, and (iv) a lock area situated between the buffer area and the one or more lockable exits, the system comprising: at least one beacon, each beacon of the at least one beacon includes a wireless transmitter configured to recurrently transmit a short-range wireless signal; andat least one tag, each tag of the at least one tag being a device that is either a personal tag borne by a person, an asset tag attached to an asset, or a fixed tag that is fixed at a location that is either at the border of or within the buffer area or lock area, each tag of the at least one tag including: a receiver configured to receive short-range wireless signals transmitted from the at least one beacon,a network communication device configured to at least wirelessly communicate with other tags, anda processor programmed to: based on a short-range wireless signal sent from the at least one beacon and received via the receiver, identify whether the restricted object is in the buffer area or in the lock area,check whether the restricted object is authorized to leave the safe area,upon recognizing that the restricted object is in the buffer area and is not authorized to leave the safe area, send a message to a tag of a selected staff member, the message instructing to move the restricted object back to the safe area prior to the restricted object reaching the lock area, andupon recognizing that the restricted object is in the lock area and is not authorized to leave the safe area, send a locking signal to at least one lockable exit of the one or more lockable exits to prevent exit of the restricted object from said confined area. 2. The system of claim 1, wherein the at least one processor is further programmed to: identify one or more staff members that are currently in vicinity of the restricted object; andselect the selected staff member from the one or more staff members that are currently in vicinity of the restricted object. 3. The system of claim 1, wherein the at least one tag is further programmed to: wait for acknowledgement from a tag of a staff member of the at least one staff member; andif no acknowledgement is received within a predefined period of time: send a message to a tag of another selected staff member within the institutional site. 4. The system of claim 1, wherein either: (1) the at least one beacon includes multiple different beacons that are fixed at locations within the institutional site, and the at least one tag is a tag attached to the restricted object; or(2) the at least one beacon is a beacon attached to the restricted object, and the at least one tag includes multiple ta... | TECHIP INTERNATIONAL LIMITED | A61B005/11 G01S005/02 G07C009/00 G08B021/02 G08B021/22 G01S001/68 G08B013/24 G06K017/00
US09795305 | B2 | 2017-10-24 | Systems and methods for determining patient temperature | 2016-02-05 | A temperature probe includes a handle and a shaft extending from the handle. The shaft includes a distal end, a proximal end, and a tip at the distal end. The temperature probe also includes a capacitance sensor disposed on one of the handle and the shaft, the capacitance sensor configured to measure a change in capacitance when positioned proximate a conductor. The temperature probe further includes a temperature sensor disposed on the shaft, the temperature sensor configured to measure a body cavity temperature of a patient. | 1. A temperature probe, comprising: a shaft having a distal end, a proximal end, and a tip at the distal end;a first sensor disposed on the shaft, the first sensor configured to generate a first signal indicative of an identity of a measurement site;a second sensor disposed on the shaft, the second sensor configured to generate a second signal indicative of a temperature associated with the measurement site; anda controller in communication with the first sensor and the second sensor, the controller configured to: receive the first signal,receive the second signal,determine the identity of the measurement site based on the first signal,select an operating mode based on the identity of the measurement site, andestimate a core temperature of a patient based on the operating mode and the second signal. 2. The temperature probe of claim 1, wherein the identity of the measurement site is selected from a group consisting of a mouth, an axilla, a rectum, and an ear drum. 3. The temperature probe of claim 1, wherein the first sensor comprises a capacitance sensor configured to determine a change in capacitance as the capacitance sensor approaches the measurement site, the first signal being indicative of the change in capacitance, andthe controller being configured to determine the identity of the measurement site based at least in part on the change in capacitance. 4. The temperature probe of claim 3, wherein the controller is configured to determine the identity of the measurement site based on a correlation between the change in capacitance and a known change in capacitance associated with the measurement site. 5. The temperature probe of claim 1, wherein the first sensor comprises a capacitance sensor configured to determine a capacitance of the measurement site, the first signal being indicative of the capacitance. 6. The temperature probe of claim 5, wherein the controller is configured to determine the identity of the measurement site by: determining differences between the capacitance and a plurality of known capacitance values, each value of the plurality of known capacitance values being associated with a respective potential measurement site of the patient, andselecting one of the respective potential measurement sites of the patient corresponding to a smallest determined difference. 7. The temperature probe of claim 1, wherein the probe further includes an optical sensor in communication with the controller and configured to determine a change in an amount of radiation received by the optical sensor as the optical sensor approaches the measurement site, the controller being configured to determine the identity of the measurement site based at least in part on the change in the amount of radiation. 8. The temperature probe of claim 7, wherein the optical sensor comprises a photodiode configured to determine a change in an amount of light received by the photodiode, the controller being configured to determine the identity of the measure... | Welch Allyn, Inc. | A61B006/00 A61B005/01 A61B005/00 A61B005/06 A61B034/20
US09832301 | B2 | 2017-11-28 | Systems and methods for adjusting power levels on a monitoring device | 2016-02-23 | A system and method is provided for health management, for performing a parameter measurement in a fluid sample, including a mobile communications device with an application adapted to perform health management for one or more health related conditions; a strip adapted to absorb a fluid sample and to produce a signal indicative of the parameter level in the sample; and a fluid testing apparatus for obtaining an analog measurement of the fluid parameter, the fluid testing apparatus including an adaptor adapted to connect the strip to the mobile communications device to thereby allow delivery of the produced signal or a correlated signal to the mobile communications device for obtaining a measurement of the fluid parameter displayed on the mobile communications device, wherein the fluid testing apparatus relies on the mobile communications device for at least one of power supply and display means, and wherein the system includes a power level adjustment means associated with the device audio jack and/or the adaptor, for providing a suitable power output to operate the fluid testing apparatus. | 1. A health management system for performing a parameter measurement in a fluid sample, comprising: a. a mobile communications device with an application to perform health management for one or more health related conditions;b. a strip adapted to absorb a fluid sample and to produce a signal indicative of the parameter level in the sample; andc. a fluid testing apparatus for obtaining an analog measurement of the fluid parameter, the fluid testing apparatus including an adaptor to connect the strip to the mobile communications device to thereby allow delivery of the produced signal or a correlated signal to the mobile communications device for obtaining a measurement of the fluid parameter displayed on the mobile communications device, wherein the fluid testing apparatus relies on the mobile communications device for at least one of power supply and display means;and wherein the system includes a power level adjustment means associated with the device audio jack, the adaptor, or a combination thereof, for providing an electrical power output to operate the fluid testing apparatus when the power level adjustment means is connected to the mobile communications device,wherein the power level adjustment means is configured to dynamically modify in real time the electrical power output from the mobile communications device. 2. The system of claim 1, wherein the power level adjustment means are adapted to modify the electrical power output from the communications device in calculated increments. 3. The system of claim 1, wherein the power level adjustment means are configured to reduce printed circuit board heating of an audio-powered fluid resting apparatus, resulting from excessive electrical power input. 4. The system of claim 1, wherein the power level adjustment means are configured to increase battery life of the mobile communications device. 5. The system of claim 1, wherein the power level adjustment means are configured to reduce an output volume of electrical power provided by the mobile communications device. 6. The power level adjustment means of claim 1, wherein the power level adjustment means include a file with instructions to execute commands to enable execution of instructions to maximize efficiency of an output of electrical power for powering the FTA connected to the mobile device. 7. The power level adjustment means of claim 1, wherein the power level adjustment means is connected to a memory in the mobile communications device having stored therein power level adjustment data. 8. A blood testing apparatus for performing glucose measurement in a blood sample comprising: a. a glucose strip adapted to absorb a blood sample and to produce a signal indicative of the glucose level in the blood sample; andb. an adaptor adapted to connect the glucose strip to a mobile communications device to thereby allow delivery of the produced signal or a correlated signal to the mobile communications device for obtaining a measurement of the g... | LABSTYLE INNOVATION LTD. | H04W052/02 H04M001/725 A61B010/00 A61B005/145
US09801601 | B2 | 2017-10-31 | Method and system for performing multi-bone segmentation in imaging data | 2015-12-29 | A computer implemented method for performing bone segmentation in imaging data of a section of a body structure is provided. The method includes: Obtaining the imaging data including a plurality of 2D images of the section of the body structure; and performing a multiphase local-based hybrid level set segmentation on at least a subset of the plurality of 2D images by minimizing an energy functional including a local-based edge term and a local-based region term computed locally inside a local neighborhood centered at each pixel of each one of the 2D images on which the multiphase local-based hybrid level set segmentation is performed, the local neighborhood being defined by a Gaussian kernel whose size is determined by a scale parameter (σ). | 1. A computer implemented method for performing bone segmentation in imaging data of a section of a body structure, the method comprising: Using a processor,Obtaining the imaging data including a plurality of 2D images of the section of the body structure; andPerforming a multiphase local-based hybrid level set segmentation on at least a subset of the plurality of 2D images by minimizing an energy functional including a local-based edge term and a local-based region term computed locally inside a local neighborhood centered at each pixel of each one of the 2D images on which the multiphase local-based hybrid level set segmentation is performed, the local neighborhood being defined by a Gaussian kernel whose size is determined by a scale parameter (σ);Generating a 3D volume including a plurality of 3D subvolumes from the segmented blobs of the secondary segmented image data; andAssociating an anatomical component to each one of the 3D subvolumes using the anatomical knowledge data relative to the section of the body structure of the imaging data. 2. The computer implemented method of claim 1, wherein the local neighborhood is circular and performing the multiphase local-based hybrid level set segmentation further comprises: for each pixel of the 2D images, dynamically changing region descriptors based on a position of a center of the local neighborhood. 3. The computer implemented method of claim 1, further comprising: selecting a value of λ to adjust a performance of the multiphase local-based hybrid level set segmentation with λ being greater than 0 and smaller than 1, wherein λ multiplies the local-based edge term and (1−λ) multiplies local-based region term. 4. The computer implemented method of of claim 1, wherein the 2D images include two phases and the energy functional is: 2-phase(φ,c,b)=(1−λ)region(φ,c,b)+λedge(φ)+μp(φ)where λ is greater than or equal to 0 and smaller than or equal to 1, μ is a positive constant, b is a bias field accounting for intensity inhomogeneity, and c is a vector representing intensity-based constant values in disjoint regions, edge(φ)=νg(φ)+αg(φ)wherein ν and α are normalization constants,ℒℊ⁡(ϕ)⁢=Δ⁢∫ℊσ,τ⁢⁢δɛ⁡(ϕ)⁢∇ϕ⁢d⁢⁢x,⁢𝒜ℊ⁡(ϕ)⁢=Δ⁢∫ℊσ,τ⁢⁢ℋɛ⁡(-ϕ)⁢d⁢⁢x,⁢ℊσ,τ⁢=Δ⁢11+𝒻σ,τ,⁢σ,τ⁡(x)=∫Kσ⁡(y-x)⁢uτ⁡(y)⁢d⁢⁢y,⁢uτ⁢=Δ⁢∇Gτ*I2with Gτ being a Gaussian kernel with a standard definition τ and I being the image, Kσ, a kernel function computed by means of a truncated Gaussian function of the scale parameter (σ), ρ being the radius of the local circular neighborhood:Kσ⁡(u)={1a⁢e-u2/2⁢σ2,u≤ρ0,otherwise⁢⁢and⁢⁢Hɛ⁡(ϕ)=12⁡[1+2π⁢arctan⁡(ϕɛ)]with ε being a parameter; and region(φ,c,b)=∫(Σi=1N(∫Kσ(y−x)¦I(x)−b(y)ci¦2dy)i(φ(x))dx i is a membership function of each region Ωi, and is defined as: 1(φ)=ε(φ)2(φ)=1−ε(φ)wherein p is a regularisation term: p(φ)=∫p(¦∇φ¦)dx and the minimization of the hybrid energy functional is carried out by gradient descent method:∂ϕ∂t=-∂F2⁢_phase∂ϕ. 5. The computer implemented method of of claim 1, wh... | LABORATOIRES BODYCAD INC. | G06K009/00 A61B006/00 G06T007/11 G06T007/12 G06T007/174 G06T007/136 A61B005/00 A61B005/055 A61B006/03
US09775578 | B2 | 2017-10-03 | Unmapped region visualization | 2013-08-12 | A method, including capturing, from an imaging system, a three-dimensional (3D) image of a body cavity, and using the captured 3D image to construct a simulated surface of the body cavity. A probe having a location sensor is inserted into the body cavity, and in response to multiple location measurements received from the location sensor, multiple positions are mapped within respective regions of the body cavity so as to generate respective mapped regions of the simulated surface. Based on the simulated surface and the respective mapped regions, one or more unmapped regions of the simulated surface are delineated, and the simulated surface of the body cavity is configured to indicate the delineated unmapped regions. | 1. A method, comprising: capturing a three dimensional image of at least part of a body cavity;constructing a simulated surface of at least part of said body cavity using said three dimensional image, and showing the simulated surface on a display;after constructing said simulated surface, inserting a probe having a location sensor into said body cavity;with the probe inserted inside said body cavity, mapping at least a portion of said body cavity corresponding to said simulated surface, the mapping process comprising: 1. moving a tip of the probe to a position in the body cavity where the tip is positioned against a body cavity tissue, and determining that the tip is positioned against the body cavity tissue;2. determining a position of the tip of the probe while it is against the body cavity tissue using the location sensor in the probe;3. mapping said position of the tip of the probe while it is against the body cavity tissue with respect to the simulated surface, and visually marking said position of the tip of the probe as a mapped region of the simulated surface on the display by overlaying one of a plurality of different mapped visual designs on a corresponding area of the simulated surface, 3a. wherein mapped regions remain marked with one of said plurality of mapped visual designs at least until completion of the mapping process;3b. wherein areas of the simulated surface which remain unmapped are visually indicated as unmapped by overlaying a first visual design on corresponding areas of the simulated surface unless and until the unmapped areas are mapped; and3c. wherein said mapping and visual marking takes place without intervention from an operator;4. moving the tip of the probe to a different position in the body cavity where the tip is positioned against body cavity tissue; and5. repeating steps 1-4 a plurality of times; wherein at least some regions of the simulated surface which have been visually marked as being mapped on the display during said mapping process are further modified during the course of the mapping process to visually indicate changing mapping resolution; wherein mapping resolution corresponds to a number of different positions in a region of the simulated surface which have been mapped by the mapping process;wherein said plurality of mapped visual designs comprise a plurality of visually distinguishable mapping patterns each corresponding to different respective mapping resolutions; andwherein the display is modified during the course of the mapping process to reflect changes in mapping resolution in regions of the simulated surface as additional tip positions are mapped;the method further comprising presenting a legend on the display, the legend comprising the first visual design and the plurality of mapped visual designs, the plurality of mapped visual designs comprising a range of mapping patterns corresponding to different mapping resolutions. 2. The method according to claim 1, wherein the 3D image comprise... | BIOSENSE WEBSTER (ISRAEL) LTD. | A61B005/055 A61B006/00 G01R033/34 A61B006/03 G01R033/48 A61B005/00 A61B005/06 A61B005/042 A61B034/20 A61B017/00 A61B034/00 A61B090/00
US09814392 | B2 | 2017-11-14 | Visual tracking and annotaton of clinically important anatomical landmarks for surgical interventions | 2010-11-01 | A visual tracking and annotation system for surgical intervention includes an image acquisition and display system arranged to obtain image streams of a surgical region of interest and of a surgical instrument proximate the surgical region of interest and to display acquired images to a user; a tracking system configured to track the surgical instrument relative to the surgical region of interest; a data storage system in communication with the image acquisition and display system and the tracking system; and a data processing system in communication with the data storage system, the image acquisition and display system and the tracking system. The data processing system is configured to annotate images displayed to the user in response to an input signal from the user. | 1. A visual tracking and annotation system for surgical intervention in an eye, comprising: a surgical instrument comprising a surgical tool suitable for use in the eye and an optical sensor attached to the surgical tool, wherein the surgical tool has a reference portion and the optical sensor has an end fixed relative to said reference portion such that light reflected by said reference portion can be detected by the optical sensor along with light reflected by tissue that is proximate to or in contact with a distal end of said surgical tool while in use in order to provide a distance of the distal end of the surgical tool to the tissue;an image acquisition and display system arranged to obtain a plurality of image streams of a surgical region of interest and of said surgical instrument proximate to said surgical region of interest and to display said plurality of image streams to a user;a tracking system configured to track a position of said surgical instrument relative to said surgical region of interest based on processing said plurality of image streams;a data storage system in communication with said image acquisition and display system and said tracking system; anda data processing system in communication with said data storage system, said image acquisition and display system and said tracking system,wherein said data processing system is configured to generate an annotation from a first image of said plurality of image streams displayed to said user during said surgical intervention in response to an input signal from said user, the annotation being associated with the position of the surgical instrument and a location of the tissue relative to the surgical region of interest,wherein the data processing system is further configured to determine the location of the tissue based on the position of the surgical instrument and the distance of the distal end of the surgical tool to the tissue;wherein the tracking system is further configured to acquire a location of the annotation relative to at least one corresponding anatomical landmark location acquired from said plurality of image streams,wherein said data storage system is configured to store said annotation and said location of said annotation relative to the corresponding anatomical landmark location, wherein said data processing system is further configured to retrieve said annotation and said location of said annotation relative to the corresponding anatomical landmark location from said data storage system, andwherein said data processing system is further configured to overlay said annotation onto a second image of said plurality of image streams having at least one of a different direction and position from said first image based on said location of said annotation relative to said corresponding anatomical landmark and processing of said first image. 2. A visual tracking and annotation system for surgical intervention according to claim 1, further comprising a user input device ... | THE JOHNS HOPKINS UNIVERSITY | A61B005/00 A61B003/10 A61B034/20 A61B003/12 A61B008/08 A61B003/13 A61B090/00 A61B005/06 A61B005/107 G06F017/24 A61F009/007 G02B021/00
US09801681 | B2 | 2017-10-31 | Catheters and methods for intracardiac electrical mapping | 2013-01-24 | A method and system capable of identifying ectopic foci, rotors, or conduction pathways involved in reentrant arrhythmias within cardiac tissue, and then treating identified ectopic foci, rotors, and/or pathways with either lethal or sub-lethal temperatures. The system includes a medical device having one or more mapping elements and one or more treatment elements, and a computer programmable to identify ectopic foci and rotors based at least in part on signals received from the one or more mapping elements at one or more locations. | 1. A medical system comprising: a medical device defining a distal end including at least one array arm a plurality of treatment elements coupled to the at least one array arm, each treatment element having a mapping element protruding therefrom, at least a portion of each mapping element having an insulative layer that electrically isolates the mapping element from the treatment element; and a control unit in communication with the medical device, the control unit being configured to deliver electroporation energy to the plurality of electroporation treatment elements and being programmed to: assess a morphology of an electrical conduction of interest within cardiac tissue based on monophasic action potential signals received from the mapping elements; assess a directionality of the electrical conduction of interest based on signals received from the mapping elements; identify a source of the electrical conduction of interest based on the morphology and directionality of the electrical conduction of interest; determine and communicate to a user one or more optimal treatment sites; and disrupt the electrical conduction of interest by activating the plurality of treatment elements in at least one of the one or more optimal treatment sites, activating the plurality of treatment elements including delivering electroporation energy from the control unit to the plurality of electroporation treatment elements. 2. The system of claim 1, wherein the plurality of mapping elements are arranged in a radially symmetrical pattern. 3. The system of claim 1, wherein the one or more tissue areas are composed of cells and wherein the plurality of mapping elements selectively transmit energy capable of at least one of ablation and electroporation of the cells. 4. The system of claim 1, wherein the one or more tissue areas are composed of cells and wherein the treatment element is configured to additionally transmit energy capable of at least one of adjusting the temperature of the one or more tissue areas to a sub-lethal temperature and ablating the one or more tissue areas. 5. The system of claim 1, wherein the medical device defines a longitudinal axis, each of the at least one array arm including a first portion and a second portion, the first portion being in a plane that is substantially orthogonal to the longitudinal axis of the device. 6. The system of claim 1, wherein each of the mapping elements records monophasic action potentials, such that at least one of depolarization timing, repolarization timing, action potential morphology, and action potential duration is unambiguously determined. 7. The system of claim 1, wherein each mapping element has a curved surface. 8. The system of claim 1, wherein each of the plurality of treatment elements includes a first portion that is at least partially wrapped around one of the at least one array arm and a second portion that extends a distance from the at least one array arm. 9. The system of claim 8, wher... | Medtronic Ablation Frontiers LLC | A61B005/04 A61B018/14 A61B018/18 A61B005/00 A61B005/0408 A61B005/042 A61N001/32 A61B018/00 A61B018/24 A61B018/02
US09797909 | B2 | 2017-10-24 | Sensor | 2016-09-26 | A sensor for detecting and/or quantifying the amount of analyte in a sample, the sensor including: a sensing region; anda barrier layer including a reactive oxygen species (ROS)-quenching, analyte-permeable membrane having an ROS-quenching agent adsorbed thereto; wherein the sensor is adapted so that the sample enters the sensing region of the sensor through the barrier layer. | 1. A sensor comprising: a sensing region;a porous barrier layer; anda reactive oxygen species (ROS)-quenching agent formed on or in the barrier layer;wherein the sensing region of the sensor defines one or more openings to allow an analyte in a bodily fluid sample to enter the sensing region;wherein the barrier layer is disposed across the one or more openings to allow the analyte to enter the sensing region through the barrier layer such that the analyte comes into contact with the ROS-quenching agent;wherein the barrier layer has an ROS-quenching activity sufficient to quench a solution of H2O2 having a concentration of 10 ppm; andwherein the sensor is configured such that fluid passing through the barrier layer is in contact with the ROS-quenching agent throughout the time the fluid is moving through the barrier layer. 2. The sensor of claim 1, wherein the ROS-quenching agent comprises platinum. 3. The sensor of claim 2, wherein the ROS-quenching agent comprises platinum in the form of nanoparticles. 4. The sensor of claim 1, wherein an entire sensing region, or an entire surface of the sensor that comes into contact with the sample, is coated or sheathed with the barrier layer. 5. The sensor of claim 1, wherein the barrier layer comprises polypropylene hollow fibres. 6. The sensor of claim 1, wherein the sensing region is provided on a support material. 7. The sensor of claim 1, wherein the barrier layer is provided over the entire sensing region or is provided on only part of the sensing region. 8. A sensor comprising: a sensing region; anda porous barrier layer comprising an ROS-quenching agent;wherein the sensing region of the sensor defines one or more openings to allow an analyte in a bodily fluid sample to enter the sensing region; andwherein the barrier layer is disposed across the one or more openings to allow the analyte to enter the sensing region through the barrier layer such that the analyte comes into contact with the ROS-quenching agent;wherein the barrier layer has an ROS-quenching activity sufficient to quench a solution of H2O2 having a concentration of at least 10 ppm; andwherein the sensor is configured such that fluid passing through the barrier layer is in contact with the ROS-quenching agent throughout the time the fluid is moving through the barrier layer. 9. The sensor of claim 8, wherein the barrier layer is impregnated with the ROS-quenching agent. 10. The sensor of claim 8, wherein the barrier layer comprises platinum in the form of nanoparticles. 11. The sensor of claim 8, wherein the barrier layer is applied directly onto the sensing region. 12. The sensor of claim 8, wherein the barrier layer substantially does not cause depletion of analyte. 13. The sensor of claim 8, wherein sensing region comprises a receptor and a fluorophore that are bound to one another and to a polymeric matrix. 14. The sensor of claim 8, wherein the sensing region is provided on a support material. 15. The sensor of claim 8... | Lightship Medical Limited | G01N033/66 G01N033/542 G01N033/84 G01N021/64 G01N021/77 A61B005/145 A61B005/1459 A61B005/1473 G01N033/543
US09789287 | B2 | 2017-10-17 | Head mounted IV holder assembly and method of use | 2014-07-18 | A head mounted catheter holder assembly is disclosed comprising a head mount assembly and an one or more catheter holders. Said one or more catheter holders comprises a first catheter holder. Said one or more catheter holders attach to said head mount assembly. Said head mount assembly releasably attaches to a head of a user. Said one or more catheter holders releasably hold a one or more catheter tubes. | 1. A head mounted catheter holder assembly comprising: a head mount assembly, one or more catheter holders, one or more catheter tubes, and a catheter;said one or more catheter holders comprises a first catheter holder;said one or more catheter holders are attached to said head mount assembly;said head mount assembly is configured to releaseably attach to a head of a user;said one or more catheter holders releaseably hold said one or more catheter tubes;said one or more catheter tubes each comprise said catheter at a first end of said one or more catheter tubes andan IV bag at a second end of said one or more catheter tubes;said one or more catheter holders hold said one or more catheter tubes between said first end and said second end;said catheter is configured to attach to an insertion point on a neck of said user;said head mounted catheter holder assembly is configured to relieve stress on said insertion point of said user by supporting said one or more catheter tubes between said insertion point and said IV bag;said insertion point with said first end of said one or more catheter tubes are downstream of said first catheter holder and said IV bag and said second end of said one or more catheter tubes are upstream of said first catheter holder;said first catheter holder is positioned above said insertion point;said head mount assembly consists of a headband;said headband is configured to fit around said head of said user;said one or more catheter holders attach to said headband;said headband comprises a nonirritant band configured to limit irritation of skin of said user of said headband;said headband comprises an elastic material configured to move without becoming dislodged to isolate tugs and pulls on said catheter when said catheter tube is moved upstream of said one or more catheter holders;said head mount assembly is configured to securely hold a portion of said catheter tube above said insertion point securely enough that tape is not required to hold said catheter in said insertion point;said one or more catheter holders comprise said first catheter holder and a second catheter holder;said head mount assembly comprises a first side and a second side;said first catheter holder attaches to said first side of said head mount assembly; andsaid second catheter holder attaches to said second side of said head mount assembly. 2. Said head mounted catheter holder assembly of claim 1 wherein: said headband comprises a variable circumference configured to securely attaching saidhead mounted catheter holder assembly to said head of said user;said headband comprises a first fastener having a first portion and a second portion;said headband comprises a strip of rectangular fabric having a width being longer than its height;said headband comprises a first end, a second end, an outside surface and an inside surface;said first portion attaches to said headband at said first end of said headband on said outside surface;said second portion attaches to s... | - | A61M025/02 A61B005/0215 A61B005/00
US09775565 | B1 | 2017-10-03 | Device and system for monitoring operator biometric condition and blood alcohol presence to prevent driving of a vehicle by an alcohol or otherwise impaired operator | 2015-11-20 | A new system monitors ethanol alcohol levels of a vehicle operator by collecting sweat from the operator's hands, and detecting the presence, if any, of ethanol in the sweat. The system can then be used, if ethanol is present, to take action, such as immobile vehicle disablement, in the event of intoxication caused by an impermissibly high levels of ethanol of the operator. The system includes devices, referred to as pods in the description, which are sweat-collecting devices that are attached to the steering wheel of the vehicle. If the measurable ethanol in collected moisture from the operator's hands exceeds a preset threshold, the system could be configured to warn the operator to park the vehicle thereafter to disable operation of the vehicle, but if the operator does not so discontinue operation of the vehicle hazard warning lights and audible warnings within and without the vehicle will be activated alerting near vehicles of a dangerous vehicle being operated in close proximity. Likewise the new system monitors for pulse rate and oxygen levels of the operator can be used to recommend operator action when the pulse rate and/or oxygen levels are outside of the normal range for an operator. When the pulse rate and or oxygen levels of the system are outside of the normal parameters for an operator, the system will warn the operator to park the vehicle and thereafter to disable operation of the immobile vehicle, but if the operator does not so discontinue operation of the vehicle the hazard warning lights, on board video, and audible warnings within and without the vehicle will be activated alerting near vehicles of a dangerous vehicle being operated in close proximity, and alarms will be sent to real time recording and monitoring devices. | 1. A vehicle operator monitoring system for installation in a vehicle and responsive to sensing of ethanol in the body of an operator of the vehicle, comprising: an ethanol vapor collection device oriented to collect moisture epidermally present on the palms of hands of the vehicle operator, for determining from said moisture the presence of ethanol in the body of the operator,a control module separate from the ethanol vapor collection device, the control module being configured to continuously control one or more vehicle functions including disabling the vehicle after a predetermined time, in accordance with the presence of ethanol in the body of the operator,the ethanol vapor collection device comprising a layer of moisture pervious material for receiving the moisture,a source of partial vacuum for drawing the moisture through the pervious material, so as to cause the received moisture, with any ethanol in the moisture, to be deposited on the pervious material,a moisture vapor extraction conduit for receiving the moisture vapor, with any ethanol in the moisture vapor, and providing it to the control module for ethanol analysis from which to determine blood alcohol content (BAC) of the operator,the control module providing a signal concerning at least one operation of the vehicle in response to the determination of a determined level of blood alcohol content of the vehicle operator, that signal including an announcement to the operator of a period of time in which the operator has to park the vehicle safely wherein the control module controls a digital image camera aimed at the operator and wherein the camera Is configured to take a digital image of the operator if the operator lifts either hand from a respective ethanol vapor collection device, and, wherein the control module monitors the BAC ethanol level periodically in normal operation, but, more frequently than, in said normal operation if tire BAC trends higher, but less frequently than in said normal operation, if the BAC trends lower. 2. A system as set forth in claim 1 wherein the moisture epidermally present on the palms of hands of a vehicle operator is palmar sweat of the operator. 3. A system as set forth in claim 1 wherein the porous material is a layer of high-efficiency particulate air filtration media. 4. A system as set forth in claim 1 wherein a palm shield having apertures therein overlies the porous material, such that received moisture will pass through the apertured palm shield and will be deposited as received moisture upon the filter material, and the received moisture will include any ethanol in the moisture resulting from the consumption of ethanol by the operator. 5. A system as set forth in claim 2 wherein the ethanol vapor collection device includes a contact surface for receiving the hands of the operator when driving. 6. A system as set forth in claim 5 wherein the ethanol vapor collection device includes an enclosure enclosing the moisture pervious material.... | - | A61B005/00 A61B005/1477 G06F007/00 G06F017/00 B60W040/08 A61B010/00 A61B005/145
US09801608 | B2 | 2017-10-31 | Apparatus and method for measuring intracranial pressure | 2015-08-19 | An apparatus for measuring intracranial pressure constituted of: a transmitter arranged to transmit a first acoustic signal through a first cranial point; a receiver arranged to receive a second acoustic signal from a second cranial point; and a control circuitry, wherein the control circuitry is arranged to: extract from the detected second acoustic signal a first set of frequency components associated with the transmitted first acoustic signal; extract from the detected second acoustic signal a second set of frequency components associated with intracranial processes; and determine intracranial pressure responsive to the extracted first set of frequency components and the extracted second set of frequency components. | 1. An apparatus for measuring intracranial pressure comprising: a transmitter arranged to transmit a first acoustic signal through a first cranial point;a receiver arranged to receive a second acoustic signal from a second cranial point, the second acoustic signal comprising: the first acoustic signal after having traveled from the first cranial point to the second cranial point; anda third acoustic signal comprising acoustic vibrations of intracranial processes, anda control circuitry,wherein said control circuitry is arranged to: extract from said received second acoustic signal a first set of frequency components, said first set of frequency components associated with said transmitted first acoustic signal;extract from said received second acoustic signal a second set of frequency components, said second set of frequency components associated with the third acoustic signal;determine intracranial pressure responsive to said extracted first set of frequency components and said extracted second set of frequency components; andoutput said determined intracranial pressure. 2. The apparatus of claim 1, wherein the intracranial processes comprise the vascular system of the brain and the respiratory cycle. 3. The apparatus of claim 1, wherein said transmitted first acoustic signal exhibits a predetermined dominant frequency, said first set of frequency components comprising said predetermined dominant frequency, and wherein said second set of frequency components comprises frequencies of the acoustic vibrations of the intracranial processes. 4. The apparatus of claim 3, wherein said predetermined dominant frequency is 621 Hz. 5. The apparatus of claim 1, wherein said control circuitry is further arranged to: calculate a mean of peak to peak values of said extracted first set of frequency components;calculate a mean of standard deviations of a plurality of predetermined portions of said extracted first set of frequency components;calculate an average of said calculated first frequency component set peak to peak value mean and said calculated first frequency component set standard deviation mean;determine a severity index responsive to said calculated average; andoutput an indicator of said determined severity index. 6. The apparatus of claim 1, wherein said control circuitry is further arranged to: calculate a first function of a mean of peak to peak values of said extracted first set of frequency components;calculate a second function of a mean of standard deviations of a plurality of predetermined portions of said extracted first set of frequency components; andcalculate a sum of said calculated first function and second function,wherein said determination of intracranial pressure by said control circuitry is responsive to said calculated sum. 7. The apparatus of claim 1, wherein said control circuitry is further arranged to: calculate a third function of a mean of peak to peak values of said extracted second set of frequency components;calcula... | HEADSENSE MEDICAL LTD. | A61B008/08 A61B005/0205 A61B008/00 A61B005/03 A61B005/00
US09820668 | B2 | 2017-11-21 | Insertion of medical devices through non-orthogonal and orthogonal trajectories within the cranium and methods of using | 2014-03-10 | An elongated device adapted for insertion, including self-insertion, through the body, especially the skull is disclosed. The device has at least one effector or sensor and is configured to permit implantation of multiple functional components through a single entry site into the skull by directing the components at different angles. The device may be used to provide electrical, magnetic, and other stimulation therapy to a patient's brain. The lengths of the effectors, sensors, and other components may completely traverse skull thickness (at a diagonal angle) to barely protrude through to the brain's cortex. The components may directly contact the brain's cortex, but from there their signals can be directed to targets deeper within the brain. Effector lengths are directly proportional to their battery size and ability to store charge. Therefore, longer angled electrode effectors not limited by skull thickness permit longer-lasting batteries which expand treatment options. | 1. A method of effecting or collecting data about a target site in a brain, the method comprising: creating an opening at an entry site in an outer surface of a patient's skull, the skull having a thickness defined between the outer surface and an inner surface; inserting a first shaft through the opening into the skull at the entry site along a first angle such that at least a portion of the first shaft extends within the skull between the outer and inner surfaces, the portion extending within the skull between the inner and outer surfaces diverging away from the opening of the outer surface as the first shaft extends from the outer surface to the inner surface, the first shaft having coupled thereto at least one component including an effector or a sensor, wherein the at least one component is inserted into the skull and extends inside the skull generally parallel to an inner surface of the skull, and extends between the skull and the brain; and inserting a second shaft through the opening into the skull at the entry site along a second angle such that at least a portion of the second shaft extends within the skull between the outer and inner surfaces, the portion extending within the skull between the inner and outer surfaces diverging away from the opening of the outer surface as the second shaft extends from the outer surface to the inner surface, the second angle different from the first angle, the second shaft having coupled thereto at least one component including an effector or a sensor, wherein the at least one component is inserted into the skull and extends inside the skull generally parallel to the inner surface of the skull, and extends between the skull and the brain. 2. The method of claim 1, further comprising providing biofeedback to the patient based on data collected about the target site. 3. The method of claim 1, wherein effecting the target site comprises stimulating the target site in the brain. 4. The method of claim 1, wherein effecting the target site comprises stimulating the dura to stimulate the trigeminal nerve. 5. The method of claim 1, wherein upon insertion of the first shaft, an angle between the first shaft and skin is different from an angle between the first shaft and an inner cortex of the skull. 6. The method of claim 1, wherein upon insertion of the first shaft, an angle between the first shaft and an outer cortex of the skull is different from an angle between the first shaft and an inner cortex of the skull. 7. The method of claim 1, further comprising monitoring brainwaves to collect data, and using the collected data for biofeedback. 8. The method of claim 1, wherein the shafts are inserted percutaneously, and are removed after a duration of time. 9. The method of claim 1, further comprising inflating a balloon cuff disposed around the shafts to stabilize the components within the brain, and deflating the balloon cuff prior to removal of the shafts from the skull. 10. The method of claim 1, wh... | - | A61B005/0478 A61B017/16 A61N001/05 A61B005/00 A61B005/0482 A61N001/36
US09802007 | B2 | 2017-10-31 | Methods and apparatus for lancet actuation | 2013-11-18 | A lancet driver is provided wherein the driver exerts a driving force on a lancet during a lancing cycle and is used on a tissue site. The driver comprises of a drive force generator for advancing the lancet along a path into the tissue site, and a manual switch for a user interface input. | 1. An agent injection device including electrical components, comprising: a housing; the housing including a first housing and a second housing, where the first housing being larger than the second housing, the first and second housings having dissimilar shapes, the first housing configured to house a battery, and the second housing configured to house the electrical components; a needle with an interior channel, the needle positionable in the housing; an agent container coupled to the needle and positioned at the housing;an actuator processor coupled to a position sensor;a human interface processor that is lower power and slower than the actuator processor, the human interface processor configured to run intermittently to conserve power for the human interface processor, the human interface processor configured to control power for the actuator processor and is a timer for the actuator processor so that the actuator processor does not remain on for long periods of time;a human interface coupled to the human interface processor that allows a user to adjust a sampling setting; and a controllable drive force generator coupled to the actuator processor and configured to be coupled to the needle to advance the needle along a path into a tissue site to a selected depth of the agent. 2. The device of claim 1, wherein the needle has a controllable velocity. 3. The device of claim 1, wherein the agent container is collapsible. 4. The device of claim 3, wherein the agent container is collapsible in response to a signal from the actuator processor. 5. The device of claim 1, further comprising: a cassette that includes a plurality of needles and a plurality of agent containers. 6. The device of claim 5, wherein the actuator processor provides a series of metered dose of the agent for long-term medication needs. 7. The device of claim 1, wherein the needle is disposable in the housing. 8. The device of claim 1, wherein the needle has a distal end with three facets. 9. A method for delivering an agent from an agent injection device that includes electrical components, comprising: providing the agent injection device that includes a housing, a needle, an agent container coupled to the needle, an actuator processor and a controllable drive force generator coupled to the actuator processor, the housing including a first housing and a second housing, where the first housing being larger than the second housing, the first and second housings having dissimilar shapes, the first housing configured to house a battery, and the second housing configured to house the electrical components; indicating a low battery output and a device error condition output via a user interface; andusing the controllable drive force generator to advance the needle along a path into a tissue site to a selected depth for delivery of the agent at the selected depth; using the actuator processor including a memory that stores profiles and in response to the actuator processor selecti... | Sanofi-Aventis Deutschland GmbH | A61B005/00 A61M005/46 A61B005/15 A61B005/157 A61M005/145 A61M005/172 A61M005/20 A61M005/32 A61B005/151
US09802025 | B2 | 2017-10-31 | Over-the-needle expandable catheter | 2014-06-17 | The present disclosure is directed to an over-the-needle (OTN) catheter assembly having an expandable catheter. More specifically, the OTN catheter assembly includes a catheter having an expandable body with a proximal end and a distal end coaxially mounted onto a needle. The body defines a lumen extending from the proximal end to the distal end. Further, the body is expandable between a compressed length and an expanded length. The needle is configured within the lumen of the catheter and has a first length. The first length is longer than the compressed length of the catheter and shorter than the expanded length of the catheter. In addition, the body of the catheter is configured to expand from the compressed length to the expanded length when heat is applied to the catheter such that the distal end of the catheter extends past a distal end of the needle. | 1. An over-the-needle catheter assembly, comprising: a catheter comprising an expandable body having a proximal end and a distal end, said body defining a lumen extending from said proximal end to said distal end, said body comprising a coil spring constructed of a shape-memory material, said body expandable between a compressed length and an expanded length via said coil spring;a needle configured within the lumen of the catheter, said needle comprising a first length, said first length being longer than said compressed length of said catheter and shorter than said expanded length of said catheter, wherein a distal end of said coil spring extends past a distal end of said needle when said body is expanded to said expanded length,wherein said body of said catheter is configured to expand from said compressed length to said expanded length when heat is applied to said coil spring such that said distal end of said body of said catheter expands past said distal end of said needle. 2. The catheter assembly of claim 1, further comprising a heat application assembly configured to apply heat to said catheter such that said body of said catheter expands from said compressed length to said expanded length. 3. The catheter assembly of claim 2, wherein said heat application assembly comprises at least one of a nerve stimulation assembly, one or more battery devices, temperature-controlled water, an ultrasound device, or a vibration device. 4. The catheter assembly of claim 1, wherein said coil spring is embedded within a wall of said catheter. 5. The catheter assembly of claim 1, wherein said coil spring fits within said lumen of said catheter. 6. The catheter assembly of claim 1, wherein said shape-memory material comprises at least one of a shape-memory polymer or a shape-memory alloy. 7. The catheter assembly of claim 6, wherein said shape-member material comprises the shape-memory polymer, wherein said shape-memory polymer comprises at least one of block copolymers, thermoplastic polymers, or thermosetting polymers. 8. The catheter assembly of claim 6, wherein said shape-member material comprises the shape-memory alloy, wherein said shape-memory alloy comprises Nitinol. 9. A method for using an over-the-needle catheter assembly to provide treatment to a targeted site within a patient, the method comprising: providing an expandable catheter having a proximal end and distal end, the catheter coaxially mounted onto a needle, said catheter comprising a coil spring constructed of a shape-memory material;inserting simultaneously the catheter and the needle into the patient until the distal end of the needle reaches the targeted site;applying heat to the coil spring of the catheter such that the catheter expands from a compressed length to an expanded length, wherein a distal end of said coil spring extends past the distal end of said needle when said catheter is expanded to said expanded length;removing the needle from the catheter while the catheter ... | Avent, Inc. | A61M025/00 A61M025/06 A61B005/00 A61M025/01 A61N001/05
US09801566 | B2 | 2017-10-31 | Automatic identification of instruments used with a surgical navigation system | 2013-11-22 | A system and apparatus taught to determine the identification and selected information relating to surgical instruments near a reader. The information can be stored on a member operable to transmit the information to a reader at a selected time. The information can be used in a navigation system to assist in navigation of the instrument relative to a patient. | 1. A system to determine information related to a selected device used with a surgical navigation system for navigating a procedure relative to an anatomy, the system comprising: an information tag having a memory operable to store, receive, or transmit a selected information;an information tag reader configured to read the selected information from the memory of the information tag, write the selected information to the memory of the information tag, or combinations thereof at least during an operative procedure;a surgical instrument associated with the information tag wherein the information tag includes the selected information specific to the surgical instrument;a navigation system;a communication system interconnecting the information tag reader and the navigation system wherein the selected information read from the information tag is transmitted to the navigation system;wherein the information tag reader is further configured to write the selected information to the memory of the information tag during or after a use of the surgical instrument including a change from a previous calibration information including size of the surgical instrument, tip location of the surgical instrument, or orientation of the surgical instrument; anda display operable to display a representation of the surgical instrument based upon the information read by the information tag reader from the memory of the information tag including the change in calibration information when previously written, wherein the selected information read by the information tag reader from the memory is read during the operative procedure to ensure that the representation of the surgical instrument is correct. 2. The system of claim 1, wherein the memory of the information tag includes a flash memory storage system interconnected with at least one of a transmission system, a receiving system, or combinations thereof. 3. The system of claim 2, wherein the information tag reader is operable to receive data transmitted from the information tag. 4. The system of claim 2, wherein the information tag reader includes a radio frequency transceiver operable to receive data from the information tag. 5. The system of claim 1, wherein the surgical instrument includes a first portion and a plurality of a second portion; wherein each of the plurality of second portions are operable to be interconnected with the first portion at a selected time. 6. The system of claim 5, wherein the information tag reader system is operable to automatically obtain information from the information tag via the information tag reader and transmit the information via the communication system to the navigation system. 7. The system of claim 6, wherein at least one of the information tag reader or the navigation system is configured to compare the read selected information to a pre-operative surgical plan and determine that at least the surgical instrument is appropriate for a selected portion of the pre-operative su... | Medtronic Navigation, Inc. | A61B005/05 A61B005/06 A61B090/00 A61B034/20 A61B090/90 A61B090/98 A61B034/10 A61B034/00
US09821156 | B2 | 2017-11-21 | Apparatus for detecting and localizing insulation failures of implantable device leads | 2016-02-26 | The invention relates to a method and apparatus for diagnosis of conductor anomalies, such as insulation failures, in an implantable medical device, such as an implantable cardioverter defibrillator (ICD), a pacemaker, or a neurostimulator. Insulation failures are detected and localized by identifying changes in electrical fields via surface (skin) potentials. Small variations in potential are detected along the course of the electrode near the site of insulation failure. | 1. An apparatus for detecting insulation failures of an implantable defibrillation lead in electrical communication with the heart of a patient, the apparatus comprising: an AC signal generator operably coupleable to the implantable defibrillation lead and configured to output a generated signal having a voltage that is below a pacing threshold and having a frequency that is higher than a cardiac cell capture frequency above which cardiac cell capture occurs in response to stimulation;an AC amplifier having one input connected to a plurality of recording sites and one input connected to a reference electrode, the AC amplifier amplifying skin surface signals from a plurality of recording sites on a surface of the skin of the patient, the recording sites being located at one or more external locations in the region of the defibrillation lead within the body of the patient;a filter connected to the amplifier output and to the output of the AC signal generator, the filter configured to produce a filtered signal passing signal characteristics of the generated signal; andcircuitry operably coupleable to the filter to receive the filtered signal and configured to generate a human or machine readable representation of an electrical potential of the filtered signal at each one of the plurality of recording sites, the human or machine readable representation indicating presence of an insulation failure based at least in part on an analysis of a local maximum voltage spike. 2. The apparatus of claim 1, wherein the circuitry operably coupleable to the filter comprises a modulator configured to output a modulated representation of the filtered signal, the modulated representation of the filter signal being configured for processing by an electrocardiogram (ECG) device. 3. The apparatus of claim 2, wherein the modulated representation of the filtered signal is a lower frequency representation of the filtered signal. 4. The apparatus of claim 3, wherein the modulated representation of the filtered signal has a frequency of 100 Hz. 5. The apparatus of claim 1, wherein the circuitry operably coupleable to the filter comprises a digital meter configured to display the filtered signal. | Lambda Nu Technology LLC | A61N001/08 A61N001/37 A61N001/39 A61B005/0424 A61B005/053 A61B005/00
US09823894 | B2 | 2017-11-21 | Controlling audio tempo based on a target heart rate | 2015-08-12 | A method for controlling an audio output comprises playing a first audio file having a first tempo, measuring a first heart rate of a user, determining whether the first heart rate of the user is greater than a target heart rate, and playing a second audio file having a second tempo, the second tempo is slower than the first tempo, responsive to determining that the first heart rate of the user is greater than the target heart rate. | 1. A method for controlling an audio output, the method comprising: receiving a target heart rate for a user;calculating a warmup heart rate for the user based at least in part on the target heart rate;starting a warmup timer;playing a first audio file having a first tempo, wherein the first tempo corresponds to the warmup heart rate;determining whether the warmup timer has expired;responsive to determining that the warmup timer has not expired, measuring a first heart rate of the user,determining whether the first heart rate of the user is greater than the warmup heart rate, andresponsive to determining that the first heart rate of the user is greater than the warmup heart rate, playing a second audio file having a second tempo, the second tempo being slower than the first tempo; andresponsive to determining that the warmup timer has expired, playing a third audio file having a third tempo, wherein the third tempo corresponds to the target heart rate,measuring a second heart rate of a user,determining whether the second heart rate of the user is greater than the target heart rate,determining a time remaining in the third audio file,playing a fourth audio file having a fourth tempo, the fourth tempo being slower than the third tempo, responsive to determining that the second heart rate of the user is greater than the target heart rate and responsive to the time remaining in the third audio file being greater than a threshold,completing the playing the third audio file before playing the fourth audio file, responsive to the time remaining in the third audio file being less than the threshold;determining whether an acceleration of the user is below an acceleration threshold, andresponsive to determining that the acceleration of the user is below the acceleration threshold, pausing the playing the third audio file or pausing the playing the fourth audio file. 2. The method of claim 1, further comprising playing a fifth audio file having a fifth tempo, the fifth tempo being faster than the third tempo, responsive to determining that the second heart rate of the user is not greater than the target heart rate. 3. The method of claim 1, further comprising: after playing the third audio file having the third tempo, determining whether the user is moving;pausing the play of the third audio file and starting a timer responsive to determining that the user is not moving;determining whether the timer has expired and whether the user has resumed moving; andresuming playing the third audio file having the third tempo responsive to determining that the user has resumed moving and that the timer has not expired. 4. The method of claim 1, wherein the first audio file includes music having the first tempo. 5. A system for outputting audio comprising: a processor operative to: receive a target heart rate for a usercalculate a warmup heart rate for the user based at least in part on the target heart rate;start a warmup timer;play a first audio file having a firs... | INTERNATIONAL BUSINESS MACHINES CORPORATION | G06F017/00 G06F003/16 G05B015/02 A61B005/11
US09814413 | B2 | 2017-11-14 | Long-term implantable monitoring system and methods of use | 2015-07-24 | Methods and systems include a long-term implantable ultra-filtrate monitoring system that uses micro-porous membranes to produce an ultra-filtrate of tissue interstitial fluid or blood plasma. The ultra-filtrate is transported through a sensor to detect a level of analyte in the ultra-filtrate. The long-term implantable fluid monitoring system thus includes a first porous catheter, a second porous catheter, a sensor configured to measure an amount of analyte in fluid, and a pump configured to move fluid through the first porous catheter to the sensor and from the sensor through the second porous catheter. | 1. A method of measuring a level of analyte in body fluid, comprising: implanting a monitoring system into a body of a patient, the implanted monitoring system having a first multi-layer porous membrane, a second multi-layer porous membrane, and a sensor, the monitoring system further comprising a catheter, the first multi-layer porous membrane forming a wall of the catheter;creating an ultra-filtrate from a body fluid by imparting a pressure differential to filter a portion of the body fluid through the first multi-layer porous membrane of the implanted monitoring system, the first multi-layer porous membrane having an outer layer and an inner layer, the outer layer having larger pores than the inner layer;measuring an amount of analyte in the ultra-filtrate with the sensor of the implanted monitoring system; andreturning the ultra-filtrate to the body fluid through the second multi-layer porous membrane. 2. The method of claim 1, further comprising calibrating the monitoring system based upon an external measurement of the analyte. 3. The method of claim 1, wherein the implanting step comprises implanting the catheter into subcutaneous vascular tissue, the method further comprising promoting ingrowth of the subcutaneous vascular tissue into the pores of the outer layer. 4. The method of claim 1, wherein the implanting step comprises implanting at least a portion of the catheter into a blood vessel. 5. The method of claim 1, wherein the implanting step comprises implanting the catheter into a blood vessel and implanting a central module connected to the porous catheter into subcutaneous tissue. 6. The method of claim 1, wherein the implanting step comprises implanting the monitoring system between a vein and an artery as a vascular shunt. 7. The method of claim 1, wherein the sensor is an optical sensor. 8. The method of claim 1, wherein the pressure differential pulls the ultra-filtrate into the sensor and returns the ultra-filtrate to the body. 9. The method of claim 8, further comprising maintaining a constant pressure differential despite dynamic changes in hydrostatic pressure or oncotic pressure. 10. The method of claim 1, wherein a rate of creating the ultra-filtrate is between 0.5 μL/min and 20 μL/min. 11. The method of claim 1, further comprising displaying the measured amount of analyte. 12. The method of claim 1, wherein the analyte is glucose and wherein the body fluid is interstitial fluid or blood. 13. The method of claim 12, further comprising activating an alarm for hyperglycemia based upon the amount of analyte, hypoglycemia based upon the amount of analyte, or a rapid rate of change in the amount of analyte. 14. The method of claim 5, wherein the first multi-layer porous membrane includes a third layer configured to provide structural support to the membrane. 15. The method of claim 1, wherein the inner layer of the first multi-layer porous membrane comprises a plurality of extensions configured to extend into the... | THOMAS JEFFERSON UNIVERSITY | A61B005/00 A61B005/145 A61B005/1473 A61M005/172 A61M027/00 A61B005/1455 A61B005/1459 A61B005/1495 A61M005/14
US09804245 | B2 | 2017-10-31 | Magnetic resonance imaging apparatus and magnetic resonance imaging method | 2008-06-26 | An IR pulse is applied to a tag region B that is disposed at the upstream side of the ascending aorta relative to a tag region A at a timing with a second predetermined delay time TD2 (for example, 600 ms) from the application time of an IR pulse to the tag region A to thereby perform tagging. By this tagging, it is possible to suppress the MR signals derived from the substantial portions and the blood within the tag region B. Subsequently, an imaging scan is performed after a predetermined time lapse TIA (for example, 1200 ms) from the application time of the IR pulse to the tag region A or after a predetermined time lapse TIB (for example, 600 ms) from the application time of the IR pulse to the tag region B. | 1. A magnetic resonance (MR) imaging apparatus comprising: static and gradient magnetic field generators, at least one RF coil for transmitting and/or receiving RF signals to/from a patient when disposed within said magnetic field generators and at least one programmed computer system and associated memory connected to control the gradient magnetic field generators and RF transmitter and receiver circuits coupled to said at least one RF coil;said at least one computer system being configured (a) to accept an input individually designating at least one of position, shape and size for each of a plurality of tag regions, (b) to accept an input individually of a time period from an RF tag pulse to an imaging pulse for each of the plurality of tag regions, and (c) to respectively set the period for each of the plurality of tag regions, the plurality of tag regions being at least partially spatially located within an MR imaging region;configured to apply the RF tag pulse to each tag region based on the respectively corresponding set period, apply the imaging pulse to the MR imaging region after applying the plurality of RF tag pulses to the plurality of tag regions, and acquire MR echo signals from the MR imaging region; andconfigured to reconstruct an image of the MR imaging region based on said acquired MR echo signals,wherein said acquired MR echo signals emanating from at least one of said MR tag regions are of different intensity compared to acquired MR echo signals emanating from elsewhere in said MR imaging region. 2. The magnetic resonance imaging apparatus according to claim 1, wherein the plurality of tag regions are all set within the imaging region. 3. The magnetic resonance imaging apparatus according to claim 1, wherein: the at least one computer system is configured to accept an input preventing the plurality of tag regions from overlapping with each other. 4. The magnetic resonance imaging apparatus according to claim 1, wherein: the at least one computer system is configured to accept an input causing at least portions of the plurality of tag regions to overlap with each other. 5. The magnetic resonance imaging apparatus according to claim 1, wherein: the at least one computer system is configured to accept an input designating at least one of the plurality of tag regions to include another tag region. 6. The magnetic resonance imaging apparatus according to claim 1, further comprising: said at least one computer system being configured to designate a tag pulse flip angle or time lapse from application of the tag pulse to start of image data acquisition for each of the plurality of tag regions. 7. The magnetic resonance imaging apparatus according to claim 1, wherein: the at least one computer system is configured to apply a tag pulse to a first of the plurality of tag regions and subsequently to a second of the plurality of tag regions, and to acquire MR imaging data from the imaging region, after said tag pulses have been appli... | TOSHIBA MEDICAL SYSTEMS CORPORATION | A61B005/05 G01R033/563 A61B005/055 G01R033/20 G01R033/54 G01R033/48 G01R033/483 G01R033/56
US09814422 | B2 | 2017-11-14 | Compositions for solubilizing cells and/or tissue | 2012-03-28 | Solubilizing compositions are provided that include a 3-(alkyl dimethyl ammonia) propane sulfonate zwitterionic surfactant and a polyethylene glycol alkyl ether nonionic surfactant, provided that the zwitterionic surfactant and the nonionic surfactant are not simultaneously and respectively 3-(decyl dimethyl ammonia) propane sulfonate and polyoxyethylene (4) lauryl ether. Also provided are methods for solubilizing cells and/or tissue of a subject in vivo and methods for recovering proteins from skin cells of a subject in vivo using the solubilizing compositions. | 1. A method for solubilizing skin cells and/or skin tissue of a subject in vivo, the method comprising: contacting the skin cells and/or skin tissue of the subject in vivo with a buffered composition effective to solubilize the skin cells and/or skin tissue of the subject in vivo, the buffered composition comprising:a 3-(alkyl dimethyl ammonia) propane sulfonate zwitterionic surfactant; anda polyethylene glycol alkyl ether nonionic surfactant,the buffered composition being characterized by a pH of between about 7 and about 9,provided that the zwitterionic surfactant and the nonionic surfactant are not simultaneously and respectively 3-(decyl dimethyl ammonia) propane sulfonate and polyoxyethylene (4) lauryl ether. 2. The method of claim 1, further comprising subjecting the skin cells and/or skin tissue to energy. 3. The method of claim 1, further comprising collecting from the solubilized skin cells and/or skin tissue one or more of: a small organic molecule, a protein, DNA, RNA, and a lipid. 4. The method of claim 1, further comprising collecting at least one of cytosolic proteins, nuclear proteins, and surface proteins from the solubilized skin cells and/or skin tissue. 5. The method of claim 1, wherein the buffered composition further comprises a protease inhibitor. 6. The method of claim 1, wherein the skin cells comprise at least one of endothelial cells and keratinocyte cells. 7. The method of claim 1, the buffered composition comprising one or more of: phosphate-buffered saline, Tris-buffered saline, Tris-hydrochloride, and ethylene diamine tetraacetic acid. 8. The method of claim 1, wherein the composition further comprises a buffer solution, the pH of the composition in the buffer solution being about 8.8. 9. The method of claim 1, the buffered composition characterized by a total (w/v) concentration of the zwitterionic and nonionic surfactants of one or more of: between about 0.01% and about 10%; between about 0.01% and about 5%; between about 0.1% and about 2%; between about 0.1% and about 0.5%; and about 1%. 10. The method of claim 1, the buffered composition characterized by a concentration ratio of the zwitterionic surfactant to the nonionic surfactant of between about 2:3 and about 3:2. 11. A method for recovering proteins from skin cells of a subject in vivo, the method comprising: contacting the skin cells of the subject in vivo with a buffered composition effective to provide solubilized proteins from the skin cells of the subject in vivo, the buffered composition comprising:a 3-(alkyl dimethyl ammonia) propane sulfonate zwitterionic surfactant; anda polyethylene glycol alkyl ether nonionic surfactant, the buffered composition being characterized by a pH of between about 7 and about 9, provided that the zwitterionic surfactant and the nonionic surfactant are not simultaneously and respectively 3-(decyl dimethyl ammonia) propane sulfonate and polyoxyethylene (4) lauryl ether; andsubjecting the solubilized proteins to a ... | THE REGENTS OF THE UNIVERSITY OF CALIFORNIA DX BIOSCIENCES, INC. | A61B005/00 C07K001/14 A61K009/06 A61K031/205 A61K031/765 A61K009/00 C07K016/18 C12N001/06
US09801618 | B2 | 2017-10-31 | Laparoscopic surgical scaffold | 2013-07-12 | A collapsible scaffolding can collapse for deployment to a laparoscopic surgical site through a trocar or the like, and can expand to provide a surface for organ retraction within a body cavity. In the expanded state, the scaffolding may assist a surgical procedure in a variety of ways, such as by providing a rigid structure upon which to secure retracted organs or surgical tools such as lights, cameras, and so forth. | 1. A device comprising: a primary member having: a front end and a back end;a front hinge and a back hinge, that collectively articulate in a plane of articulation;an ancillary member movably connected to the primary member via the front hinge and the back hinge; whereinthe primary member and ancillary member are: constructed from a surgical material; andindividually flexible in a direction perpendicular to the plane of articulation to a degree substantially conforming to a human abdomen; and further whereinthe device is configured to transition between a first state and a second state by articulating the hinges, in which the first state is characterized by minimizing a distance between the primary member and the ancillary member and the second state is characterized by maximizing the distance between the primary member and the ancillary member, in which a maximum device cross section in the first state is suitable for insertion through a trocar in a laparoscopic procedure on a human; anda number of independent flexural structures in each of the primary member and ancillary member to facilitate flexure in the direction perpendicular to the plane of articulation, each independent flexural structure including an opposing pair of axial protrusions extending toward one another across a gap within the plane of articulation when the flexural structure is unflexed, wherein the opposing pair of axial protrusions impose a flexural limit at a point where a deflection of a corresponding one of the independent flexural structures causes the opposing pair of axial protrusions to come in contact across the gap. 2. The device of claim 1, wherein each one of the flexural structures includes a top layer, a flexible middle layer, and a bottom layer, in which the top and bottom layers each include a plurality of gaps that facilitate a deflection of the flexible middle layer and inhibit the deflection beyond an upward and a downward flexural limit according to a size of one or more corresponding top or bottom ones of the plurality of gaps. 3. The device of claim 1, further including an electrical coupling site. 4. The device of claim 3, further comprising a light source electrically coupled to the electrical coupling. 5. The device of claim 4, wherein the light source is movably coupled to at least one of the primary member and the ancillary member. 6. The device of claim 3, further comprising a camera electrically coupled to the electrical coupling site. 7. The device of claim 6, wherein the camera is movably coupled to at least one of the primary member and the ancillary member. 8. The device of claim 6, further comprising a plurality of cameras, each camera having a corresponding field of view, wherein the field of view of any camera in the plurality overlaps with at least one other field of view of at least one other camera in the plurality. 9. The device of claim 1, further comprising an attachment point for anchoring an organ. 10. The device of claim ... | Vicarious Surgical Inc. | A61B017/02 A61B001/06 A61B005/00 A61B001/32 A61B001/00 A61B001/313 A61B090/30 A61B090/50 A61B017/00
US09801993 | B2 | 2017-10-31 | Sensor clip assembly for an optical monitoring system | 2011-11-17 | Systems and sensor clip assemblies for optically monitoring blood flowing through a blood chamber are provided. A sensor clip assembly includes emitters and photodetectors positioned on opposing arms, a signal conditioning circuit for conditioning raw analog signals generated by the photodetectors while the sensor clip assembly is fastened to a blood chamber, and an analog-to-digital converter for converting the conditioned analog signals to raw digital data. The sensor clip assembly may output the raw digital data to an external device and receive synchronized control signals from the external device, or the sensor clip assembly may include a microcontroller for performing calculations on the raw digital data and providing synchronized control signals internally. Parameters of blood flowing through the blood chamber such as hematocrit, oxygen saturation, and change in blood volume may be calculated from the raw digital data derived from the raw analog signals generated by the photodetectors. | 1. A sensor clip assembly for optically monitoring extracorporeal blood flowing through a blood chamber, the sensor clip assembly comprising: a housing;a first photodetector, configured to, in a first time interval, generate a first current corresponding to a first detected signal based on light of a first wavelength from a first emitter of the sensor clip assembly that has passed through extracorporeal blood in the blood chamber;a second photodetector, configured to, in a second time interval, generate a second current corresponding to a second detected signal based on light of a second wavelength from a second emitter of the sensor clip assembly that has passed through extracorporeal blood in the blood chamber;a signal processing unit within the housing, comprising a processor and a signal conditioning circuit, the signal conditioning circuit being configured to convert the first current to a first voltage and apply a first gain to the first voltage, and configured to convert the second current to a second voltage and apply a second gain to the second voltage; andan analog-to-digital converter, configured to convert the first voltage with the first gain applied to a first digital signal and to convert the second voltage with the second gain applied to a second digital signal, wherein the first digital signal corresponds to a hemoglobin level of the extracorporeal blood and the second digital signal corresponds to a water concentration level of the extracorporeal blood;wherein the processor is configured to synchronize operation of the signal conditioning circuit with operation of the first emitter and the second emitter of the sensor clip assembly such that the first gain is only applied to the first voltage while the first emitter is on and the second gain is only applied to the second voltage while the second emitter is on, and to calculate a hematocrit value based on a ratio between the first and second digital signals;wherein the sensor clip assembly further comprises: an output port configured to facilitate transmission of the calculated hematocrit value to an external computing device;wherein the signal processing unit is further configured to, in response to determining an insufficient number of bits on the analog-to-digital converter being activated, adjust a resolution level for signals received by the analog-to-digital converter by increasing the first gain and/or the second gain. 2. The sensor clip assembly of claim 1, wherein the signal conditioning circuit further comprises: at least one transimpedance amplifier corresponding to each photodetector for converting current to voltage; andat least one digitally-controllable trimpot corresponding to each photodetector for applying a gain. 3. The sensor clip assembly of claim 1, wherein the sensor clip assembly further comprises: at least one shroud for blocking ambient light from being received at the first photodetector and/or the second photodetector. 4. The sensor clip assembly of... | Fresenius Medical Care Holdings, Inc. | A61B005/1455 A61M001/36 A61B005/145 F16B002/22
US09788757 | B2 | 2017-10-17 | Breathing biofeedback device | 2006-12-21 | A breathing biofeedback device, having a microphone configured to acquire sounds of a user's breathing; a controller communicatively connected with the microphone, the controller processing the signals acquired by the microphone to produce an output signal, the controller processing the signal whereby the microphone signal is first pre-amplified to a voltage level that can be processed by an audio envelope detector circuit, the envelope detector signal is then fed into the analog-to-digital converter input of the controller allowing it to constantly sample the input volume level, the controller then controlling the output volume level fed to the headphones utilizing a digitally controlled variable-gain amplifier, wherein the output signal is not modified in any manner from the original input, except in volume; and a pair of earphones connected with the controller and configured to convey the output signal to the user. | 1. A breathing biofeedback device, comprising: a microphone configured to convert breathing related sound signals into a first output signal; a preamplifier amplifying the first output signal to generate a second output signal;a variable gain amplifier operable to amplify the second output signal in accordance with a gain level signal to generate a third output signal;a controller comprising an audio envelope detector operable to receive the second output signal from the preamplifier, wherein the controller is operable to generate the gain level signal in response to an analysis of an output of the audio envelope detector and wherein the gain level signal increases in response to an increase in a volume level of the second output signal and decreases in response to a decrease in the volume level of the second output signal;and earphones operable to convert the third output signal into breathing sound signals output to a user, wherein the controller further comprises a state machine configured to detect a type of breath cycle and further wherein the controller is configured to operate in one of several functional modes based on the type of breath cycle detected, wherein the functional mode is selected from a group consisting of: a basic mode for enabling a user to set one of an output volume and effects preset to the breathing sound signals output to the user, an anti-snore mode operable to detect breath cycles associated with snoring, a stress reduction mode operable to add additional sounds to the breathing sound signals output to the user, and a fitness training mode operable to add a sound of a pulse to the breathing sound signals output to the user. 2. The device of claim 1 further comprising a user interface operatively coupled with the controller. 3. The device of claim 1 wherein in said anti-snore mode the controller is further configured to detect the wave frequency and volume of a snore, and further wherein the gain level increases as the detected volume increases and the gain level decreases as the detected volume decreases. 4. The device of claim 1 wherein in said stress reduction mode the controller is configured to add background sounds operable to enhance relaxed breathing patterns to the breathing sound signals output to the user. 5. The device of claim 1 wherein in said fitness training mode the controller is configured to add the sound of a pulse along with the user's breath, wherein the pulse is operable to provide a cue for the user to breathe at a certain rate. 6. A method for determining breathing biofeedback comprising: converting breathing related sound signals into a first output signal;amplifying the first output signal using a preamplifier to generate a second output signal;generating a third output signal by amplifying the second output signal in accordance with a gain level signal using a variable gain amplifier;using a digital signal processor (DSP) comprising an audio envelope detector operable to receive the se... | BREATH RESEARCH, INC. | A61B005/08 A61B007/00 A61B005/00 A61M021/00
US09808199 | B2 | 2017-11-07 | Two part EEG monitor with databus and method of communicating between the parts | 2012-08-17 | A personal wearable EEG monitor comprises a base part (1) having signal processing means (23), and an electrode part (2) with at least two electrodes (11, 12) for measuring an EEG signal of a person. The electrode part (2) comprises means for converting the EEG signal into a digital signal. The EEG monitor comprises a databus for transferring data between the base part (1) and the electrode part (2) and for providing power from one part to the other. The databus is adapted for application of two electrical wires. The invention further provides a method for communicating between two parts of an EEG monitor. | 1. A personal EEG monitor, comprising a base part having a signal processor, an electrode part with at least two electrodes for measuring an EEG signal of a person, said electrode part having an analog-to-digital converter for converting the EEG signal into a digital signal, and a two-wire databus for transferring data between said base part and said electrode part over two wires with said base part and electrode part located adjacent one ear, and for providing power over said two wires from one part to the other; and for the application of at least four different states in respective time slots, where a first state is during a first period of time and is for transfer of power, a second state is during a second period of time and is for transmission of signal from said base part to said electrode part, a third state is during a third period of time and is for transmission of signal from said electrode part to said base part; and a fourth state is during a fourth period of time and is set to low in order for the first state to start with a rising edge, with no power being transferred between said base part and electrode part during said second and third states, and wherein data to be transferred includes at least first and second values, and said databus represents said first value with a first signal level and represents said second value with a second signal level. 2. A personal EEG monitor, comprising a base part having a signal processor,an electrode part with at least two electrodes for measuring an EEG signal of a person, said electrode part having an analog-to-digital converter for converting the EEG signal into a digital signal, anda two-wire databus for transferring data between said base part and said electrode part over two electrical wires with said base part and electrode part located adjacent one ear, and for also providing power from one part to the other over said two wires, wherein data to be transferred includes at least first and second values, and said databus represents said first value with a first signal level and represents said second value with a second signal level different from the first signal level. 3. The EEG monitor according to claim 2, wherein two different states of the databus are applied in different time slots, including at least a first state for transfer of power without transfer of data, and a second state for transmission of signal from said electrode part to said base part without transfer of power. 4. The EEG monitor according to claim 3, wherein said databus is configured such that said first state for transfer of power takes up at least 50% of the time on the databus. 5. The EEG monitor according to claim 4, wherein said first state for transfer of power takes up at least 70% of the time on the databus. 6. The EEG monitor according to claim 2, wherein at least three different states of the databus are applied in different time slots, where a first state is for transfer of power, a second state is... | Widex A/S | A61B005/0476 A61B005/0478 A61B005/00 G06F001/26 H04R025/00 H04B003/54 H02J007/34
US09788774 | B2 | 2017-10-17 | Methods and apparatus for selectively occluding the lumen of a needle | 2014-09-18 | A fluid transfer device for parenterally transferring fluid to and/or from a patient includes a housing, a needle, and an occlusion mechanism. The housing defines a fluid flow path and is couplable to a fluid reservoir. The needle has a distal end portion that is configured to be inserted into the patient and a proximal end portion that is configured to be fluidically coupled to the fluid flow path of the housing, and defines a lumen therebetween. The occlusion mechanism selectively controls a fluid flow between the needle and the fluid flow path. The occlusion mechanism includes an occlusion member that is movable between a first configuration where the lumen of the needle is obstructed during insertion into the patient and a second configuration where the lumen of the needle is unobstructed after the needle has been inserted into the patient allowing fluid transfer to or from the patient. | 1. A method for transferring fluid to or from a patient using a parenteral transfer device, the device including a housing defining a fluid flow path, a needle having a lumen in fluid communication with the fluid flow path and being configured for insertion into the patient, and an occlusion mechanism operable to selectively control fluid flow to or from a patient through the needle lumen and the fluid flow path, the occlusion mechanism includes an occlusion member disposed about at least a portion of the needle and configured to selectively obstruct the lumen of the needle during insertion into the patient, the method comprising: disposing the occlusion member in a first position relative to the needle prior to insertion into the patient to place the occlusion mechanism in a first configuration in which the lumen of the needle is obstructed to prevent at least one of tissue, bodily fluid, and contaminants from entering the lumen;inserting the needle into the patient, the occlusion mechanism being in the first configuration as the needle is inserted into the patient;after the needle has been inserted into the patient, rotating the occlusion member from the first position to a second position to place the occlusion mechanism in a second configuration in which the lumen of the needle is unobstructed and the occlusion member is disposed outside of the fluid flow path; andtransferring a volume of a fluid substantially free of tissue and contaminants between the patient and a fluid reservoir in fluid communication with the fluid flow path when the occlusion mechanism is in the second configuration. 2. The method of claim 1, wherein the occlusion member is manually rotated from the first position to the second position. 3. The method of claim 1, wherein the occlusion member is mechanically rotated from the first position to the second position. 4. The method of claim 1, wherein the occlusion member is automatically rotated from the first position to the second position. 5. The method of claim 1, wherein the needle includes a recessed portion fluidically isolated from the needle lumen, the recessed portion configured to receive and retain at least one of tissue, bodily fluid, and contaminants when the needle is inserted into the patient. | Magnolia Medical Technologies, Inc. | A61B005/15 A61M005/158 A61M039/22 A61M005/168 A61M005/32
US09788721 | B2 | 2017-10-17 | Low-power activity monitoring | 2015-07-22 | Some examples include a secure low-power sensor platform able to be used for activity monitoring or other purposes. For instance, a sensor system may sense sparse signals using compressed sensing. As one example, compressed sensing may be used to obtain sparse signals from analog sensor signals received from one or more sensors. The data obtained by compressed sensing may be sent subsequently to a computing device to reconstruct the sensor signal. Examples herein enable lower power usage, lower cost, and lower weight than conventional devices, while also enabling processing advantages, such as less overall data to process and lower data storage utilization. | 1. A sensor block comprising: an amplifier to receive a sensor signal from a sensor and output an amplified signal;a compressed sensing analog-to-digital conversion block coupled to the amplifier, the compressed sensing analog-to-digital conversion block including a random number generator that generates random values to enable an analog-to-digital converter (ADC) to take random samples of the amplified signal based on the random values, and output compressed sensing data comprising the random samples;a post processor coupled to the compressed sensing analog-to-digital conversion block to receive the compressed sensing data and remove noise from the compressed sensing data by active thresholding; anda communication interface coupled to the post processor to receive the compressed sensing data from the post processor and send the compressed sensing data to a computing device, wherein the sensor block further utilizes the random number generator for encrypting the compressed sensing data prior to sending to the computing device. 2. The sensor block as recited in claim 1, wherein the ADC takes the random samples converting the amplified signal from an analog signal to a digital signal based on the random values generated by the random number generator to generate a sparse random sampling of the amplified signal. 3. The sensor block as recited in claim 1, wherein: the random number generator is a linear feedback shift register; andthe ADC is a switched capacitor ADC. 4. The sensor block as recited in claim 1, wherein the sensor block further utilizes the random number generator for encrypting the compressed sensing data prior to sending to the computing device by using the random number generator to generate a randomization matrix for encrypting the compressed sensing data, wherein the computing device is able to read the encrypted compressed sensing data using a corresponding randomization matrix. 5. The sensor block as recited in claim 1, wherein the communication interface comprises a transmitter and an antenna for wirelessly transmitting the compressed sensing data to the computing device. 6. The sensor block as recited in claim 1, wherein the communication interface comprises a wired connection for transmitting the compressed sensing data to the computing device. 7. The sensor block as recited in claim 1, wherein the sensor block is associated with an activity monitor and is integrated with at least one of: an accelerometer;a gyrometer;a compass;a proximity sensor;electrocardiogram sensor;a heart rate monitor;a pressure sensor; ora GPS receiver. 8. The sensor block as recited in claim 1, further comprising a storage for receiving the compressed sensing data prior to sending to the computing device, wherein at least one of the storage, the post-processor, or the communication interface cycle between an active state and a quiescent state based at least in part on at least one of: an occurrence of sampling of the sensor signal by the compres... | BIOBIT, INC. | A61B005/00 H03M001/12 G06F007/58 H03M001/46 A61B005/11 H04M001/725
US09804035 | B2 | 2017-10-31 | Thermometer with age specific feature selection | 2016-06-14 | A thermometer has temperature sensing tip, a processor taking temperature readings and determining a sensed temperature reading of the living being from the temperature sensing tip. The thermometer also includes a display and a backlight for lighting the display. The backlight is activated upon a command from the processor and the processor determines whether to activate the backlight based upon the temperature readings. The method embodiment can includes the steps of using the processor to monitor a temperature change indicated by a temperature sensing element. The processor then detects a temperature decrease and activates a first color light emitting element to backlight a display if the temperature decrease exceeds or equals a predetermined threshold. The thermometer is operable in one of a plurality of selectable operating modes, and the predetermined threshold is dependent upon the selected operating mode. For example, operating modes may depend on patient age range or measurement location. Age range may include infant, toddler and adult. | 1. A thermometer for use with a living being, comprising: a housing having an associated display;a temperature sensor located in or on the housing;a processor provided within the housing and in signal communication with the temperature sensor, the processor operable to determine a temperature of the living being based on a signal provided by the temperature sensor;a memory coupled to the processor, the memory configured to store information related to a plurality of operating modes;a switch located in or on the housing and operably connected to the processor, the switch configured to select one operating mode from among the plurality of operating modes, to produce a selected operating mode, wherein the selected operating mode comprises a plurality of temperature ranges;a plurality of output indicators located in or on the housing and operable upon a command from the processor, at least a first output indicator associated with the display and comprising graphic or text, at least a second output indicator comprising colored light; anda plurality of completion indicators located in or on the housing and operable upon a command from the processor, at least a first completion indicator comprising a lighting element, at least a second completion indicator comprising an audible indicator. 2. The thermometer of claim 1, wherein each operating mode corresponds to a different range of ages of the living being. 3. The thermometer of claim 1, wherein each of the plurality of temperature ranges of the selected operating mode comprises a different range of temperatures. 4. The thermometer of claim 1, wherein at least one output indicator is associated with an operating mode. 5. The thermometer of claim 1, wherein at least one output indicator indicates the selected operating mode. 6. The thermometer of claim 1, wherein at least one output indicator is associated with at least one temperature range of the selected operating mode. 7. The thermometer of claim 1, wherein at least one output indicator is associated with one of the plurality of temperature ranges of the selected operating mode. 8. The thermometer of claim 1, wherein at least one output indicator is configured to indicate a temperature range of the selected operating mode that a determined temperature of the living being is within. 9. The thermometer of claim 1, wherein at least one completion indicator is activated prior to determination of a final temperature of the living being. 10. The thermometer of claim 1, wherein at least one completion indicator is activated after determination of a final temperature of the living being. 11. A thermometer for use with a living being, comprising: a housing;a temperature sensor located in or on the housing;a processor provided within the housing and in signal communication with the temperature sensor, the processor operable to determine a temperature of the living being based on a signal provided by the temperature sensor;a memory coupled to the proc... | Helen of Troy Limited | G01K001/00 G01K007/00 G01K013/00 A61B005/01 G01K001/02
US09791246 | B2 | 2017-10-17 | Stimulant target unit and accessory for a stimulant target unit | 2015-08-28 | A stimulant target unit including: a housing, at least one stimulation source configured to provide a stimulus to stimulate a user, a contact sensor configured to detect an impact to the stimulant target unit and triggered by an impact that is within a selected range of impact forces, a feedback system configured to inform the user the stimulant target unit has been actuated in response to the stimulus, a controller programmed to register triggering of the contact sensor, a protective cover configured to retain the stimulant target unit and a mounting structure configured to connect the stimulant target unit to a support structure. The mounting structure includes a vibration isolator configured to, over the selected range of impact forces, inhibit transmission of force from an impact within the selected range of impact forces on the support structure to the contact sensor sufficiently to prevent the triggering of the contact sensor. | 1. A stimulant target unit, comprising: a housing;at least one stimulation source coupled to the housing, the stimulation source configured to provide a stimulus to stimulate a user;a contact sensor coupled to the housing and configured to detect an impact to the stimulant target unit;a feedback system configured to inform the user that the stimulant target unit has been actuated in response to the stimulus;a controller programmed to register triggering of the contact sensor;a protective cover configured to be positioned in front the housing and to transmit some impact force to the housing during some impacts; anda mounting structure configured to connect the stimulant target unit to a support structure, wherein the mounting structure includes a vibration isolator configured such that wherein, over a selected range of impact forces, the contact sensor is configured to be triggered by an impact that is within the selected range of impact forces, and wherein the vibration isolator inhibits transmission of force from an impact within the selected range of impact forces on the support structure to the contact sensor sufficiently to prevent the triggering of the contact sensor,wherein the housing includes a first shell connected to a second shell via an elastomeric rim,and wherein the rim has opposing first and second grooves and the first shell has a first axially extending flange along a periphery thereof,the second shell has a second axially extending flange along a periphery thereof,the first flange is disposed in said first groove and the second flange is disposed in the second groove, andthe first flange ordinarily not bottoming in the first groove and the second flange ordinarily not bottoming in said second groove to thereby enable the first shell to move axially relative to the second shell upon impact. 2. A stimulant target unit according to claim 1, wherein the at least one stimulation source includes at least one of an auditory source to generate sound and a light source configured to generate light to be transmitted through the protective cover. 3. A stimulant target unit according to claim 2, wherein the light source includes a light emitting diode. 4. A stimulant target unit according to claim 1, wherein the vibration isolator includes at least one of a resilient member and a damping member. 5. A stimulant target unit according to claim 4, wherein the resilient member includes a spring. 6. A stimulant target unit according to claim 4, wherein the damping member includes a deformable member configured to absorb at least a portion of impact force on the support structure. 7. A stimulant target unit, comprising: a housing;at least one stimulation source coupled to the housing, the stimulation source configured to provide a stimulus to stimulate a user;a contact sensor coupled to the housing and configured to detect an impact to the stimulant target unit;a feedback system configured to inform the user that the stimulant target unit ha... | FITLIGHT SPORTS CORP. | A63B069/00 F41J005/04 F41J005/14 A63B063/00 A63B071/06 A63B024/00 F41J001/10 F41J005/056 F41J001/01 A61B005/16
US09805266 | B2 | 2017-10-31 | System and method for video content analysis using depth sensing | 2016-01-26 | A method and system for performing video content analysis based on two-dimensional image data and depth data are disclosed. Video content analysis may be performed on the two-dimensional image data, and then the depth data may be used along with the results of the video content analysis of the two-dimensional data for tracking and event detection. | 1. A video content analysis method comprising: capturing a video sequence that includes a plurality of frames, each frame including a video image;for each frame, receiving two-dimensional (2D) image data of the video image and also receiving depth data associated with the image data;analyzing the 2D image data, and based on an analysis of the 2D image data without the depth data detecting one or more objects depicted in the video sequence as potential human beings;using the depth data along with the one or more detected objects to classify at least a first object of the one or more detected objects as a person to be tracked, wherein a volume of the one or more detected objects is used to classify at least the first object as a person to be tracked;performing tracking on at least the first classified object; andperforming event detection analysis on the first classified object,wherein the volume is determined by using the depth data along with the 2D image data to determine a plurality of convex hull slices on different Z-planes, and by summing areas of the plurality of convex hull slices. 2. The video content analysis method of claim 1, further comprising: using the depth data along with the one or more detected objects to additionally classify at least a second object of the one or more detected objects as an object not to be tracked. 3. The video content analysis method of claim 1, wherein: classifying the first object as an object to be tracked includes classifying the first object as an object above a predetermined height or volume threshold. 4. The video content analysis method of claim 1, wherein analyzing the image data to detect one or more objects depicted in the video sequence includes detecting at least one blob that corresponds to the one or more objects. 5. The video content analysis method of claim 1, further comprising: classifying at least the first object of the one or more detected objects as a person to be tracked by using the depth data associated with the one or more detected objects and without analyzing depth data associated with a portion of the video image that is not part of the one or more detected objects. 6. The video content analysis method of claim 1, wherein the depth data is determined by a single depth sensor. 7. The video content analysis method of claim 1, wherein analyzing the 2D image data to detect one or more objects depicted in the video sequence as potential human beings includes performing two-dimensional (2D) analysis on the image data to perform motion and change detection. 8. The video content analysis method of claim 7, wherein: analyzing the 2D image data to detect one or more objects depicted in the video sequence as potential human beings further includes, based on the motion and change detection, detecting at least one blob that corresponds to the one or more objects; andusing the depth data along with the one or more detected objects to classify at least the first object of the one or mor... | AVIGILON FORTRESS CORPORATION | G06K009/00 A61B005/00 H04N007/18 A61B005/107 A61B005/11 A61B005/1171 G08B013/196 G08B021/04 G06T007/00 G06T007/73 G06T007/50 G06T007/55 G06T007/579 G06T007/246 G06T007/62
US09808084 | B2 | 2017-11-07 | Technique for adjusting the posture of a seated person | 2014-06-19 | A posture adjustment system includes a set of sensors coupled to a chair on which a person may sit. The posture adjustment system gathers data from the set of sensors and generates a posture model that reflects a posture associated with the seated person. The posture adjustment system then determines corrections to the posture of the person that could, potentially, improve their posture. The posture adjustment system then indicates those corrections to the person, or, alternatively, applies a set of adjustments to the chair to cause the person to assume a new posture that reflects the posture corrections. | 1. A non-transitory computer-readable medium storing program instructions that, when executed by a processing unit, cause the processing unit to effect modifications to the posture of a seated person, by performing the steps of: acquiring sensor data that reflects a current posture associated with the seated person;determining at least one modification to make to the current posture to change the current posture to a modified posture;generating one or more chair adjustments to be applied to a chair in which the seated person is sitting to change the current posture to the modified posture based on the at least one modification and a chair-posture mapping indicating changes in position or orientation of each section of the chair and corresponding changes in posture of a person sitting in the chair; andapplying the one or more chair adjustments to the chair in which the seated person is sitting to automatically change the current posture of the seated person to the modified posture. 2. The non-transitory computer-readable medium of claim 1, further comprising the step of generating a first posture model that includes a mathematical description of the current posture. 3. The non-transitory computer-readable medium of claim 2, wherein the step of determining the at least one modification comprises: comparing the first posture model to a second posture model;identifying one or more differences between the first posture model and the second posture model; anddetermining a modification to the first posture model that causes the first posture model to more closely approximate the second posture model. 4. The non-transitory computer-readable medium of claim 2, wherein the step of determining the at least one modification comprises: comparing the first posture model to a set of preferences associated with the seated person that reflects a preferred posture of the seated person; anddetermining a modification to the first posture model that causes the first posture model to more closely approximate the preferred posture. 5. The non-transitory computer-readable medium of claim 2, wherein the step of determining the at least one modification comprises: comparing the first posture model to a set of posture recommendations received from a medical professional that reflects a recommended posture for the seated person; anddetermining a modification to the first posture model that causes the first posture model to more closely approximate the recommended posture. 6. The non-transitory computer-readable medium of claim 1, wherein generating the one or more chair adjustments comprises: determining that a first position of the section of the chair corresponds to the current posture;determining that a second position of the section of the chair corresponds to the modified posture;comparing the first position to the second position to determine a position difference; anddetermining an adjustment to the section of the chair based on the position difference and the c... | HARMAN INTERNATIONAL INDUSTRIES, INCORPORATED | G06F019/00 A47C001/00 A47C031/12 A61B005/103 G05B013/04 A61B005/00 A61B005/11 B60N002/02 A61B003/113
US09812002 | B2 | 2017-11-07 | Wireless communication authentication for medical monitoring device | 2016-04-20 | Methods, systems, and devices for short-range low-power wireless communication of analyte information are provided. In some implementations, short-range low-power wireless communication of analyte information may include receiving an electromagnetic wireless communication signal and harvesting energy from the electromagnetic wireless communication signal. In some implementations, short-range low-power wireless communication of analyte information may include enabling capabilities associated with an external sensor in response to detecting the external sensor. In some implementations, short-range low-power wireless communication of analyte information may include detecting an analyte sample; determining an analyte concentration associated with the detected analyte sample; and transmitting an indication of the analyte concentration to an external device. | 1. An apparatus, comprising: a housing;a memory unit having computer executable instructions stored thereon;a wireless communication interface;anda processor configured to execute the computer executable instructions stored in the memory unit, wherein execution of the computer executable instructions causes the processor to: cause a search for an external sensor while one or more user interface (UI) capabilities of the apparatus are disabled, the one or more UI capabilities being associated with the external sensor, wherein the search for the external sensor includes wireless transmission of a signal adapted to detect the external sensor, andin response to detection of the external sensor, enable one or more disabled UI capabilities associated with the external sensor and cause display of an analyte concentration associated with a measurement made by the external sensor. 2. The apparatus of claim 1, wherein the wireless communication interface is configured to perform wireless communication according to a Bluetooth or a Bluetooth Low-Energy communication protocol. 3. The apparatus of claim 2, wherein the transmission of the signal occurs according to the Bluetooth or Bluetooth Low-Energy communication protocol. 4. The apparatus of claim 1, wherein the transmission of the signal occurs for a determined period of time. 5. The apparatus of claim 4, wherein execution of the computer executable instructions further causes the processor to terminate the transmission of the signal in response to not detecting the external sensor within the determined period of time. 6. The apparatus of claim 1, wherein the transmission of the signal occurs in response to powering on the apparatus. 7. The apparatus of claim 1, wherein the external sensor includes a portion positioned through a skin surface of a user in fluid contact with a bodily fluid of the user. 8. The apparatus of claim 1, further comprising a monitor adapted to analyze fluid on a test strip. 9. The apparatus of claim 8, wherein capabilities associated with the monitor adapted to analyze fluid on the test strip are enabled while the one or more UI capabilities associated with the external sensor are disabled. 10. A method for enabling one or more user interface (UI) capabilities of an apparatus, wherein the UI capabilities are associated with an external sensor, the method comprising: searching for the external sensor while the one or more UI capabilities associated with the external sensor are disabled, wherein searching for the external sensor includes transmitting a signal adapted to detect the external sensor; andin response to detecting the external sensor, enabling one or more disabled UI capabilities associated with the external sensor and causing display of an analyte concentration associated with a measurement made by the external sensor. 11. The method of claim 10, further comprising performing, by a wireless communication interface of the apparatus, wireless communication accordi... | ABBOTT DIABETES CARE INC. | G08C017/02 H04Q009/00 A61B005/00 A61B005/145 H02J003/38 H04B005/00 H04W004/00
US09820682 | B2 | 2017-11-21 | Blood sample optimization system and blood contaminant sequestration device and method | 2016-04-27 | Blood sample optimization systems and methods are described that reduce or eliminate contaminates in collected blood samples, which in turn reduces or eliminates false positive readings in blood cultures or other testing of collected blood samples. A blood sample optimization system can include a blood sequestration device located between a patient needle and a sample needle. The blood sequestration device can include a sequestration chamber for sequestering an initial, potentially contaminated aliquot of blood, and may further include a sampling channel that bypasses the sequestration chamber to convey likely uncontaminated blood between the patient needle and the sample needle after the initial aliquot of blood is sequestered in the sequestration chamber. | 1. A blood sequestration device comprising; an inlet port;an outlet port;a sequestration chamber and a sampling channel, each of the sequestration chamber and the sampling channel being connected with the inlet port via a junction, the sequestration chamber having a vent that includes an air permeable blood barrier, the sampling channel having a distal end connected with the outlet port; anda housing that houses and defines the inlet port, the outlet port, the sequestration chamber, and the sampling channel,the sequestration chamber, the sampling channel and the junction being sized and configured such that a first portion of blood flows into the sequestration chamber toward the air permeable blood barrier for sequestration therein, and a second portion of blood bypasses the sequestration chamber and the first portion of blood sequestered therein and is directed into the sampling channel toward the outlet port. 2. The blood sequestration device in accordance with claim 1, wherein the air permeable blood barrier includes an air permeable material that seals at least partially upon contact with blood. 3. The blood sequestration device in accordance with claim 1, wherein the sequestration chamber includes a sequestration channel defined in the housing. 4. The blood sequestration device in accordance with claim 1, wherein the housing comprises a bottom member mated with a top member, each of the bottom member and the top member defining a part of each of the inlet port, the outlet port, the sequestration chamber, and the sampling channel. 5. A blood sequestration device connected with a blood sampling pathway having a patient needle and a sample collection device, the blood sequestration device comprising; an inlet port connected with the patient needle;a sequestration chamber and a sampling channel, each of the sequestration chamber and the sampling channel being connected with the inlet port via a junction, the sequestration chamber having a vent that includes an air permeable blood barrier;the sequestration chamber and the junction being configured to direct a first portion of blood into the sequestration chamber toward the air permeable blood barrier;the sampling channel having a distal end connected with an outlet port, the sampling channel and the junction being configured to allow a second portion of blood bypasses the sequestration chamber and the first portion of blood sequestered therein, and is directed into the sampling channel toward the outlet port, the outlet port being connected with the sample collection device; anda housing that houses and defines the inlet port, the outlet port, the sequestration chamber, and the sampling channel. 6. The blood sequestration device in accordance with claim 5, wherein the inlet port is connected with the patient needle via a first portion of the blood sampling pathway. 7. The blood sequestration device in accordance with claim 6, wherein the outlet port is connected with the sample collection d... | Kurin, Inc. | A61B005/00 B65D081/00 A61B005/15 A61M039/04 A61B005/153 A61B005/154 A61M039/02 A61M039/06
US09802014 | B2 | 2017-10-31 | Apparatus and method for improved assisted ventilation | 2015-06-25 | Devices and methods for allowing for improved assisted ventilation of a patient by allowing blind insertion of a device that can allow ventilation regardless of whether the device is positioned within a trachea or an esophagus and allow for timed delivery of ventilations based on a condition of a thoracic cavity to increase the amount and efficiency of blood flow during a resuscitation procedure. | 1. A method for artificially ventilating an individual, the method comprising: inserting a ventilation device within a respiratory opening of the individual and positioning a working end of the ventilation device within a body passageway of the individual, the ventilation device having a pressure sensor configured to detect pressure changes within the body passage;delivering a bolus of air into an airway of the individual at a pre-determined rate;delaying delivery of the bolus of air when the sensor detects a pressure change in the air within the body passage, where the pressure change within the body passage results from a chest compressionwhere the ventilation device includes a working end for positioning in the body passageway of the individual, where the working end includes a first opening fluidly coupled to a first lumen and a second opening fluidly coupled to a second lumen, where the second opening is located along the ventilation device proximal to the first opening; andwhere inserting the ventilation device within the respiratory opening further comprises:drawing suction through the first opening to induce collapse of the body passageway and maintaining the suction for a period of time;monitoring a fluid parameter to determine whether the body passageway collapses; automatically ventilating the individual through the second lumen upon detecting collapse of the body passageway while maintaining suction to maintain the collapse of the body passageway; andautomatically ventilating the individual through the first lumen upon failure to detect collapse of the body passageway;wherein delivery of a bolus of air during automatically ventilating the individual occurs at a first timing;measuring a condition of a thoracic cavity to determine a change in the thoracic cavity; andaltering the timing of the delivery of the bolus of air during automatically ventilating the individual upon detecting the change in the condition of the thoracic cavity. 2. The method of claim 1, where delivering the bolus of air into the airway of the individual occurs after a pre-determined delay. 3. The method of claim 1, where the sensor is configured to intermittently detect pressure changes within the body passageway. 4. The method of claim 1, further comprising, after altering the pre-determined rate, the ventilation device resumes delivering the bolus of air at the pre-determined rate if the pressure sensor fails to detect a pressure range within a first period of time. 5. The method of claim 1, where measuring the condition of the thoracic cavity comprises measuring the condition of the thoracic cavity using the ventilation device. 6. The method of claim 1, where measuring the condition of the thoracic cavity comprises measuring a change in a compression of a chest of the patient. 7. The method of claim 6, where measuring the change of the compression of the chest of the patient comprises observing a force applied to the ventilation device by the body passage... | Colabs, Inc. | A61M016/04 A61M016/00 A61M016/06 A61B005/00 A61B005/087 A61M016/20 A61N001/39 A61N001/05 A61N001/362
US09782498 | B2 | 2017-10-10 | Selective delivery molecules and methods of use | 2014-01-30 | Disclosed herein is a selective delivery molecule comprising: (a) an acidic sequence (portion A) which is effective to inhibit or prevent the uptake into cells or tissue retention, (b) a molecular transport or tissue retention sequence (portion B), and (c) a linker between portion A and portion B, and (d) cargo moieties (portion DA and DB). | 1. A selective delivery molecule having the structure of SDM-41: 2. A tissue sample comprising a selective delivery molecule having the structure of SDM-41: 3. The tissue sample of claim 2, wherein the tissue sample is cancerous. 4. The tissue sample of claim 3, wherein the cancerous tissue is: breast cancer tissue, colorectal cancer tissue, squamous cell carcinoma tissue, skin cancer tissue, prostate cancer tissue, melanoma tissue, thyroid cancer tissue, ovarian cancer tissue, cancerous lymph node tissue, cervical cancer tissue, lung cancer tissue, pancreatic cancer tissue, head and neck cancer tissue, or esophageal cancer tissue. 5. The tissue sample of claim 3, wherein the cancerous tissue is breast cancer tissue. 6. The tissue sample of claim 3, wherein the cancerous tissue is colorectal cancer tissue. 7. The tissue sample of claim 3, wherein the cancerous tissue is squamous cell carcinoma tissue. 8. The tissue sample of claim 3, wherein the cancerous tissue is skin cancer tissue. 9. The tissue sample of claim 3, wherein the cancerous tissue is cancerous lymph node tissue. 10. A method of delivering Cy5 and Cy7 to a tissue of interest, comprising contacting the tissue of interest with a selective delivery molecule having the structure of SDM-41: 11. The method of claim 10, wherein the tissue of interest is cancerous. 12. The method of claim 11, wherein the cancerous tissue is: breast cancer tissue, colorectal cancer tissue, squamous cell carcinoma tissue, skin cancer tissue, prostate cancer tissue, melanoma tissue, thyroid cancer tissue, ovarian cancer tissue, cancerous lymph node tissue, cervical cancer tissue, lung cancer tissue, pancreatic cancer tissue, head and neck cancer tissue, or esophageal cancer tissue. 13. The method of claim 11, wherein the cancerous tissue is breast cancer tissue. 14. The method of claim 11, wherein the cancerous tissue is colorectal cancer tissue. 15. The tissue sample of claim 11, wherein the cancerous tissue is squamous cell carcinoma tissue. 16. The tissue sample of claim 11, wherein the cancerous tissue is skin cancer tissue. 17. The tissue sample of claim 11, wherein the cancerous tissue is cancerous lymph node tissue. 18. A method of visualizing a tissue of interest in an individual in need thereof, comprising: (a) administering to the individual a selective delivery molecule having the structure of SDM-41:and (b) visualizing at least one of Cy5 or Cy7 of SDM-41. 19. The method of claim 18, wherein the tissue of interest is cancerous. 20. The method of claim 19, wherein the cancerous tissue is: breast cancer tissue, colorectal cancer tissue, squamous cell carcinoma tissue, skin cancer tissue, prostate cancer tissue, melanoma tissue, thyroid cancer tissue, ovarian cancer tissue, cancerous lymph node tissue, cervical cancer tissue, lung cancer tissue, pancreatic cancer tissue, head and neck cancer tissue, or esophageal cancer tissue. 21. The method of claim 19, wherein the cancerous tiss... | AVELAS BIOSCIENCES, INC. | A61B010/00 A61B005/00 A61B008/00 A61K049/00 A61K047/48
US09775528 | B2 | 2017-10-03 | Photoplethysmography sensor apparatus and method | 2015-06-19 | The present invention relates to a photoplethysmography sensor apparatus (200, 300), a photoplethysmography sensor method (400), and a photoplethysmography sensor computer program product. It is proposed to measure a photoplethysmographic signal without ambient light interference. Ambient light signals are rejected by subtraction of a compensation current at the input of a transimpedance amplifier (240, 340). The compensation current is controlled via a closed loop, without interfering with the low duty cycle operation of a photoplethysmography LED (210, 310). | 1. A photoplethysmography sensor apparatus comprising: a light source unit configured to generate a first light signal;a photodetector unit configured to observe a second light signal, where said second light signal is indicative of an absorption of said first light signal in a target, wherein said photodetector unit is further configured to output a photodetector output signal in response to observing said second light signal; andan ambient light compensating feedback loop comprising a track-and-hold unit configured to receive said photodetector output signal;wherein said photoplethysmography sensor apparatus is configured to compare an output signal of said track-and-hold unit to a reference signal (VR), and wherein said ambient light compensating feedback loop comprises a first amplifying unit configured to provide a compensation current based on said comparison to the track-and-hold unit. 2. The photoplethysmography sensor apparatus of claim 1, further comprising a second amplifying unite for amplifying a sum of the compensation current and the photodetector output signal and for delivering an amplified signal to an input of the track-and-hold unit. 3. The photoplethysmography sensor apparatus of claim 2, wherein said second amplifying unit comprises a transimpedance amplifier. 4. The photoplethysmography sensor apparatus of claim 1, wherein said first amplifying unit comprises a trans-conductance amplifier, wherein said trans-conductance amplifier comprises a resistor and/or a voltage controlled current source. 5. The photoplethysmography sensor apparatus of claim 1, wherein said photoplethysmography sensor apparatus comprises an analog-to-digital converter configured to receive an output signal from said second amplifying unit. 6. The photoplethysmography sensor apparatus of claim 1, wherein said photoplethysmography sensor apparatus further comprises a light source driver unit configured to control said light source unit. 7. The photoplethysmography sensor apparatus of claim 1, wherein said photoplethysmography sensor apparatus further comprises a synchronous detector. 8. The photoplethysmography sensor apparatus of claim 7, wherein said photoplethysmography sensor apparatus further comprises a light source driver unit configured to control said light source unit and wherein said synchronous detector is configured to operate synchronously with said light source driver unit. 9. The photoplethysmography sensor apparatus of claim 7, wherein said photoplethysmography sensor apparatus comprises an analog-to-digital converter configured to receive the output signal from said synchronous detector. 10. The photoplethysmography sensor apparatus of claim 7, wherein said synchronous detector comprises a synchronous rectifier followed by a low pass filter. 11. The photoplethysmography sensor apparatus of claim 10, wherein said synchronous rectifier multiplies the signal with +1 or −1. 12. The photoplethysmography sensor apparatus of claim 1... | KONINKLIJKE PHILIPS N.V. | A61B005/08 A61B005/024 A61B005/00
US09804114 | B2 | 2017-10-31 | Sensor head for use with implantable devices | 2016-03-02 | The present invention provides a sensor head for use in an implantable device that measures the concentration of an analyte in a biological fluid which includes: a non-conductive body; a working electrode, a reference electrode and a counter electrode, wherein the electrodes pass through the non-conductive body forming an electrochemically reactive surface at one location on the body and forming an electronic connection at another location on the body, further wherein the electrochemically reactive surface of the counter electrode is greater than the surface area of the working electrode; and a multi-region membrane affixed to the nonconductive body and covering the working electrode, reference electrode and counter electrode. In addition, the present invention provides an implantable device including at least one of the sensor heads of the invention and methods of monitoring glucose levels in a host utilizing the implantable device of the invention. | 1. A sensor for use in a glucose measuring device, the sensor comprising: a plurality of electrodes, wherein the plurality of electrodes comprises a first electrode and a second electrode, wherein the first electrode comprises an electrochemically reactive surface;a non-conductive body located between the first electrode and the second electrode; anda multi-region membrane covering the plurality of electrodes, wherein the multi-region membrane comprises an immobilized enzyme domain comprising an enzyme in at least a portion thereof, wherein the immobilized enzyme domain covers and contacts the first electrode among the plurality of electrodes and does not cover and contact any other electrode of the plurality of electrodes. 2. The sensor of claim 1, wherein the multi-region membrane further comprises a resistance domain. 3. The sensor of claim 1, wherein the multi-region membrane comprises a silicone-containing copolymer. 4. The sensor of claim 3, wherein the multi-region membrane further comprises an acrylate polymer. | DexCom, Inc. | A61B005/05 G01N027/327 A61B005/145 A61B005/1486 C12Q001/00 G01N033/487
US09775560 | B2 | 2017-10-03 | Methods and apparatus for monitoring and encouraging health and fitness | 2016-03-16 | Methods and apparatus are provided for monitoring and encouraging health and fitness. In accordance with a first aspect, an apparatus is provided that is adapted to assist in weight loss and exercise. The apparatus comprises a personal digital assistant (PDA) having computer program code adapted to assist in at least one of calorie counting, meal selection, meal suggestion, weight monitoring, weight loss or gain monitoring, fat consumption monitoring, sugar consumption monitoring and salt consumption monitoring. The PDA also includes computer program code adapted to display historical data regarding at least one of calorie counting, meal selection, meal suggestion, weight monitoring, weight loss or gain monitoring, fat consumption monitoring, sugar consumption monitoring and salt consumption monitoring. Numerous other embodiments are provided, as are methods, systems and computer program products. | 1. A system comprising: a wearable monitor that monitors biometric information of a user as the user exercises, wherein the wearable monitor includes a wrist band and communicates biometric information wirelessly; andan application executable on a web-enabled, handheld electronic device associated with the user, the application including computer program code operative to direct the web-enabled, handheld electronic device to: receive biometric information measured by the wearable monitor worn by the user as the user exercises;communicate exercise level information determined from the received biometric information to a remotely-located server;communicate geographic information about the user to the remotely-located server;receive input from the user indicating a diet and exercise goal;provide the diet and exercise goal to the remotely-located server;obtain product information for a product determined by the remotely-located server to be consistent with the diet and exercise goal, such determination being based on the exercise level information provided to the remotely-located server; andpresent the product information to the user on the web-enabled, handheld electronic device as being consistent with the diet and exercise goal. 2. The system of claim 1 wherein the application executable on the web-enabled, handheld electronic device includes computer program code that facilitates purchase of the product by the user. 3. The system of claim 1 wherein the application executable on the web-enabled, handheld electronic device includes computer program code that directs the web-enabled, handheld electronic device to communicate food intake information to the remotely-located server. 4. The system of claim 3 wherein the application executable on the web-enabled, handheld electronic device includes computer program code that directs the web-enabled, handheld electronic device to monitor food intake information of the user. 5. The system of claim 1 wherein the application executable on the web-enabled, handheld electronic device includes computer program code that directs the web-enabled, handheld electronic device to provide inspirational messages to the user during exercise. 6. The system of claim 1 wherein the application executable on the web-enabled, handheld electronic device includes computer program code that directs the web-enabled, handheld electronic device to provide inspirational messages to the user prior to meal time. 7. The system of claim 1 wherein the application executable on the web-enabled, handheld electronic device includes computer program code that directs the web-enabled, handheld electronic device to provide inspirational messages to the user automatically. 8. The system of claim 1 wherein the application executable on the web-enabled, handheld electronic device includes computer program code that provides a video game and uses monitored exercise level information to affect the video game. 9. A system comprising: a weara... | - | A63B024/00 A61B005/00 G06F019/00 G06Q010/08 G06Q030/06 G09B019/00 G09B005/00 A61B005/11 G09B005/02
US09810537 | B2 | 2017-11-07 | Method of previewing off-road trails and viewing associated health requirements and related system | 2016-03-17 | A method of previewing off-road trails and viewing associated health requirements includes receiving selection of an off-road trail from a user, and displaying an overview of the selected off-road trail in response to receiving selection of the selected off-road trail from the user, the overview of the selected off-road trail indicating latitude and longitude coordinates associated with the selected off-road trail, elevation data associated with the selected off-road trail, and recommended minimum health guidelines associated with the selected off-road trail. | 1. A method of previewing off-road trails and viewing associated health requirements, the method comprising: a navigation device receiving selection of an off-road trail from a user; andthe navigation device displaying an overview of the selected off-road trail in response to receiving selection of the selected off-road trail from the user, the overview of the selected off-road trail indicating latitude and longitude coordinates associated with the selected off-road trail, elevation data associated with the selected off-road trail, and recommended minimum health guidelines that indicate what kind of fitness statistics the user is recommended to have in order for the user to attempt the selected off-road trail. 2. The method of claim 1, further comprising: receiving real-time fitness data from a fitness monitoring device worn by the user, the fitness monitoring device containing at least one sensor for sensing the real-time fitness data over time;recording the real-time fitness data for the user as the user travels on the selected off-road trail;recording video with a camera of a navigation device as the user travels on the selected off-road trail;recording latitude and longitude coordinates and elevation data with the navigation device as the user travels on the selected off-road trail; andsynchronizing, with the navigation device, the real-time fitness data, the recorded video, and the recorded latitude and longitude coordinates and elevation data generated as a result of the user traveling on the selected off-road trail to produce synchronized trail data. 3. The method of claim 2, further comprising: storing the synchronized trail data in an online server; andupdating the overview of the selected off-road trail using the stored synchronized trail data. 4. The method of claim 2, further comprising: monitoring the user's real-time fitness data while the user travels on the selected off-road trail;generating with the navigation device an alternate route for the user to follow instead of the selected off-road trail in response to determining that the user's real-time fitness data indicates that the user is not qualified to travel on the selected off-road trail; andproviding navigational assistance to the user for guiding the user to follow the alternate route. 5. The method of claim 2, further comprising issuing a warning if the user's current speed is higher than recommended for the user based on the user's real-time fitness data while the user travels on the selected off-road trail. 6. The method of claim 1, further comprising allowing the user to review the overview of the selected off-road trail while the user travels on the selected off-road trail. 7. The method of claim 1, wherein displaying the overview of the selected off-road trail comprises showing a video of the selected off-road trail, the overview of the selected off-road trail displaying latitude and longitude coordinates associated with segments of the selected off-road trail, ... | MITAC INTERNATIONAL CORP. | G01C021/00 G01C021/20 A61B005/00 A61B005/11
US09775559 | B2 | 2017-10-03 | Staged rhythm detection system and method | 2014-04-15 | A medical device and associated method for detecting and treating tachyarrhythmias acquires a cardiac signal using electrodes coupled to a sensing module. During an initial detection process, a shockable cardiac rhythm is detected by a processing module configured to compare the cardiac signal to a first set of detection criteria. By analyzing the cardiac signal, the processing module establishes at least one patient-specific detection threshold during the initial detection process. Upon establishing the at least one patient-specific detection threshold, the initial detection process is stopped, and a next detection process is started which includes comparing the cardiac signal to a second set of detection criteria including the at least one patient-specific detection threshold. In some embodiments, user programming of tachyarrhythmia detection parameters is not required. | 1. A medical device for detecting and treating tachyarrhythmia, comprising: a sensing module configured to sense cardiac events in response to a cardiac signal acquired via at least some of a plurality of electrodes; anda processing module configured to:perform an initial detection process comprising detecting a shockable cardiac rhythm by comparing the cardiac signal to a first set of detection criteria;establish, by analyzing the cardiac signal, at least one patient-specific detection threshold during the initial detection process; andupon establishing the at least one patient-specific detection threshold, stop the initial detection process and start a next detection process comprising comparing the cardiac signal to a second set of detection criteria comprising the at least one patient-specific detection threshold,the first set of detection criteria relating to a first set of cardiac signal features, and the second set of detection criteria relating to a second set of cardiac signal features different than the first set of cardiac signal features. 2. The device of claim 1, wherein the first set of detection criteria comprises only population-based detection thresholds. 3. The device of claim 1, wherein comparing the cardiac signal to a first set of detection criteria comprises determining a gross morphology feature of the cardiac signal over a predetermined time interval independent of a timing of a cardiac cycle. 4. The device of claim 1, wherein establishing the at least one patient-specific detection threshold comprises determining a beat feature of the cardiac signal representative of a cardiac cycle. 5. The device of claim 1, wherein the processing module is further configured to determine a beat feature of a cardiac cycle of the cardiac signal during the next detection process that is not determined during the initial detection process and compare the beat feature to one of the patient-specific detection thresholds. 6. The device of claim 1, wherein the processing module is further configured to: detect a concerning cardiac rhythm during a first stage of the next detection process; anddiscriminate the concerning cardiac rhythm as one of a shockable rhythm and a non-shockable rhythm during a second stage of the next detection process. 7. The device of claim 6, wherein the processing module is further configured to update the at least one patient-specific detection threshold during the first stage when a concerning cardiac rhythm is not being detected. 8. The device of claim 6, wherein the processing module is configured to discriminate the concerning cardiac rhythm by: determining a cardiac cycle feature and in a first comparison compare the cardiac cycle feature to one of the patient-specific thresholds; anddetermining a gross morphology feature across a cardiac signal segment and in a second comparison compare the gross morphology feature to a population-based threshold. 9. The device of claim 8, wherein the processing module is... | Medtronic, Inc. | A61B005/00 A61B005/04 A61B005/0464 A61N001/37 A61N001/38 A61N001/372 A61N001/362 A61N001/39
US09808185 | B2 | 2017-11-07 | Movement measure generation in a wearable electronic device | 2015-09-18 | Aspects of generating movement measures that quantify motion data samples generated by a wearable electronic device are discussed herein. For example, in one aspect, an embodiment may obtain motion data samples generated by the set of motion sensors in a wearable electronic device. The wearable electronic device may then generate a first movement measure based on a quantification of the first set of motion data samples. The wearable electronic device may also generate a second movement measure based on a quantification of the second set of the motion data samples. The wearable electronic device may then cause the non-transitory machine readable storage medium to store the first movement measure and the second movement measure as time series data. | 1. A wearable electronic device to be worn by a user, the wearable electronic device comprising: a set of one or more motion sensors to generate motion data samples that represent motion of the wearable electronic device, the motion data samples include a plurality of first motion data samples generated during a first time interval and a plurality of second motion data samples generated during a second time interval;a set of one or more processors coupled to the set of motion sensors; anda non-transitory machine readable storage medium coupled to the set of one or more processors and having stored therein instructions, which when executed by the set of one or more processors, cause the set of one or more processors to: obtain the motion data samples generated by the set of motion sensors,generate a first movement measure based on a combination of the first motion data samples into a single numerical value representative of the motion of the wearable electronic device during the first time interval,generate a second movement measure based on a combination of the second motion data samples into a single numerical value representative of the motion of the wearable electronic device during the second time interval,cause the non-transitory machine readable storage medium to store the first movement measure and the second movement measure as time series data,derive a first feature value corresponding to the first time interval based on the single numerical values of the first and second movement measures,derive a second feature value corresponding to the second time interval based on at least the single numerical value of the second movement measure,assign a first sleep state to the user of the wearable electronic device for the first time interval based on the first feature value, andassign a second sleep state to the user of the wearable electronic device for the second time interval based on the second feature value. 2. The wearable electronic device of claim 1, wherein the instructions, when executed by the set of one or more processors, also cause the set of one or more processors to: determine levels of activity of the user for each of the first and second time intervals, wherein the determination of the level of activity for the first time interval is based on a comparison of the first feature value with a threshold value, and the levels of activity are used in the assigning of the sleep state. 3. The wearable electronic device of claim 1, wherein the generating of the first and second movement measures is based on a statistical measure of the corresponding motion data samples along an axis of motion of the set of one or more motion sensors. 4. The wearable electronic device of claim 3, wherein each of the statistical measures is one or more of a quantile, an inter-quantile range, a measure of dispersion, and a measure of entropy. 5. The wearable electronic device of claim 3, wherein each of the statistical measures is derived from a frequen... | Fitbit, Inc. | A61B005/11 A61B005/00 A61B005/0205 A61B005/024 A61B005/053 A61B005/08
US09813597 | B2 | 2017-11-07 | Gaze tracker and method for tracking gaze thereof | 2015-11-18 | The present disclosure provides a gaze tracker and a gaze tracking method. The gaze tracker includes: an image acquirer to acquire eye images of a user using two or more cameras and two or more lightings; and a gaze calculator which detects a lighting reflection point from each eye image to extract a cornea central point, detects a pupil central point from each eye image, and calculates a gaze vector. In particular, when lighting reflection points are detected from each eye image, the gaze calculator calculates vectors for connection between each lighting and a corresponding lighting reflection point and detects an intersection point between the calculated vectors. When only one lighting reflection point is detected, the gaze calculator extracts the cornea central point using distance information between a pre-stored lighting reflection point and the cornea central point and the one lighting reflection point. | 1. A gaze tracker comprising: an image acquirer configured to acquire eye images of a user using at least two cameras and at least two lightings; anda processor configured to calculate a gaze to detect a lighting reflection point from each of the eye images and to extract a cornea central point, the processor configured to detect a pupil central point from each of the eye images and calculate a gaze vector using the cornea central point and the pupil central point,wherein, when at least two lighting reflection points are detected from each of the eye images, the processor is configured to calculate vectors for connection between each lighting and the lighting reflection point corresponding thereto and to detect an intersection point between the calculated vectors as the cornea central point, and when only one lighting reflection point is detected from one or more eye images among the eye images, the processor is configured to extract the cornea central point using distance information between a pre-stored lighting reflection point and the cornea central point and the one lighting reflection point. 2. The gaze tracker according to claim 1, wherein the processor is configured to calculate distances from the cornea central point to the at least two lighting reflection points and to accumulate and store calculated distance information. 3. The gaze tracker according to claim 1, wherein the processor is configured to convert 2D coordinates of the lighting reflection point and the pupil central point into 3D coordinates. 4. The gaze tracker according to claim 1, wherein the processor is configured to determine a lighting that forms the one lighting reflection point at a cornea among the at least two lightings using a face direction vector. 5. The gaze tracker according to claim 4, wherein the processor is configured to calculate a vector for connection between the one lighting reflection point and the determined lighting and extract the cornea central point using the calculated vector and the distance information. 6. The gaze tracker according to claim 1, wherein the processor is configured to calculate similarity between lighting reflection points in an eye image from which the at least two lighting reflection points are detected and an eye image from which one lighting reflection point is detected among the eye images, and to determine a lighting that forms the one lighting reflection point at a cornea. 7. The gaze tracker according to claim 1, wherein the processor is configured to obtain the distance information by calculating an average distance of distance information for a predetermined recent period among distance information items from a pre-stored cornea central point to a lighting reflection point. 8. A method for tracking gaze of a gaze tracker comprising at least two cameras and at least two lightings, the method comprising: acquiring eye images of a user using the at least two cameras;detecting a lighting reflection point from each ... | HYUNDAI MOTOR COMPANY | G06K009/00 H04N005/225 A61B003/113 A61B003/14 A61B003/00 A61B005/18 H04N013/02 A61B003/15
US09801544 | B2 | 2017-10-31 | Monitor subject monitoring device and method, and monitor subject monitoring system | 2014-04-28 | Disclosed are a monitor subject monitoring apparatus, a monitor subject monitoring method, and a monitor subject monitoring system configured to determine a first body state of a subject, based on an image data of the subject, and to determine a second body state of the subject, based on physiological data of the subject, wherein the subject monitoring apparatus, the subject monitoring method and the subject monitoring system are operable, when it is determined that one of the first and second body states of the subject is not abnormal, to re-perform a corresponding one of the determinations, and then determine whether or not a notification indicating that the subject is in an abnormal state should be issued, based on a result of the re-performed determination about the one of the first and second body states. | 1. A subject monitoring apparatus comprising: an image data input unit for inputting image data of a subject to be monitored;a physiological data input unit for inputting physiological data of the subject;a first body state determination unit operable, based on the image data of the subject input from the image data input unit, to determine a first body state of the subject as being one of a normal state and an abnormal state;a second body state determination unit operable, based on the physiological data of the subject input from the physiological data input unit, to determine a second body state of the subject as being one of the normal state and the abnormal state; and,a notification control unit operable to:in a case in which each of the first and second body states of the subject determined by the first and second body state determination units, respectively, is the abnormal state, issue a notification indicating that the subject is in the abnormal state, andin a case in which one of the first and second body states of the subject, determined by the first and second body state determination units, respectively, is not the abnormal state, cause whichever of the first and second body state determination units that made the not abnormal state determination to re-perform the determination, and then determine whether or not the subject's abnormal state-indicating notification should be issued, based on a result of the re-performed determination. 2. The subject monitoring apparatus as recited in claim 1, wherein the notification control unit is operable, in response to ascertaining a situation where the body state of the subject determined by one of the first and second body state determination units is a given first abnormal state, and the body state of the subject determined by a remaining one of the first and second body state determination units is a second abnormal state corresponding to the first abnormal state, as the situation where each of the body states of the subject determined by respective ones of the first and second body state determination units is an abnormal state, to issue the subject's abnormal state-indicating notification. 3. The subject monitoring apparatus as recited in claim 2, wherein the notification control unit is operable, in response to ascertaining a situation where the body state of the subject determined by one of the first and second body state determination units is the first abnormal state, and the body state of the subject determined by a remaining one of the first and second body state determination units is not the second abnormal state corresponding to the first abnormal state, as the situation where one of the first and second body states, determined by a corresponding one of the first and second body state determination units, is not an abnormal state, to cause the remaining one of the first and second body state determination units to re-perform the determination about the corresponding body state of ... | Konica Minolta, Inc. | G08C019/22 A61B005/00 G08B021/02 G08B025/04 G08B021/04
US09801545 | B2 | 2017-10-31 | Method and apparatus for providing rolling data in communication systems | 2015-07-30 | Methods and systems for providing data communication in medical systems are disclosed. | 1. A glucose monitoring device, comprising: an in vivo glucose sensor that monitors glucose level; andsensor electronics operatively coupled to the glucose sensor to receive and process signals from the glucose sensor, the sensor electronics generating a plurality of data packets for communication to a remote location, each generated data packet including one of a plurality of glucose monitoring device data, each glucose monitoring device data assigned with a communication time slot such that the sensor electronics sequentially communicates the plurality of generated data packets based on the assigned communication time slot, each glucose monitoring device data further including a communication count corresponding to a frequency of the sensor electronics data communications to the remote location, and wherein each generated data packet includes glucose data and the corresponding glucose monitoring device data assigned with the respective communication time slot. 2. The device of claim 1, wherein the sensor electronics communicates the plurality of generated data packets to the remote location such that each communicated data packet does not include all of the plurality of glucose monitoring device data. 3. The device of claim 1, wherein the communication time slots assigned to each glucose monitoring device data are non-overlapping. 4. The device of claim 1, wherein the glucose monitoring device data includes operational mode data including type of data communication. 5. The device of claim 1, wherein the glucose monitoring device data includes glucose sensor calibration data. 6. The device of claim 1, wherein the glucose monitoring device data includes glucose sensor temperature data. 7. The device of claim 1, wherein the glucose monitoring device data includes glucose monitoring device power supply status data of sensor electronics power supply. 8. The device of claim 7, wherein the power supply status data includes a count that increments with each communication of the generated data packet to the remote location and that corresponds to the remaining life of the power supply. 9. The device of claim 1, wherein the sensor electronics increments the communication count by an integer value with each subsequent communication. 10. The device of claim 1, wherein the glucose data communicated to the remote location with each generated packet includes real time and historical glucose data. 11. A glucose monitoring device, comprising: an in vivo glucose sensor that monitors glucose level; andsensor electronics operatively coupled to the glucose sensor to receive from the glucose sensor, the sensor electronics including: a memory;a communication unit; anda processor operatively coupled to the memory, and the communication unit, the memory having stored therein a plurality of glucose monitoring device data each assigned with a communication time slot, each glucose monitoring device data further including a communication count corresponding to a ... | Abbott Diabetes Care Inc. | A61B005/1473 H04W072/12 H04W012/02 A61B005/00 H04W072/04 A61B005/145 G06F019/00 A61B005/15
US09807361 | B2 | 2017-10-31 | Three-dimensional display device, three-dimensional image processing device, and three-dimensional display method | 2014-06-05 | A three-dimensional display device for displaying a three-dimensional image, including: a gaze point obtaining unit which obtains a position of a gaze point of a viewer; a fusional area determination unit which determines a fusional area where binocular fusion is allowed, based on the obtained position of the gaze point; a correction unit which corrects the three-dimensional image so as to suppress display of an object that is included in the three-dimensional image outside the fusional area; and a display unit which displays the corrected three-dimensional image. | 1. A three-dimensional display device for displaying a three-dimensional image, comprising: a gaze point obtaining unit configured to obtain a position of a gaze point of a viewer;a fusional area determination unit configured to determine a fusional area where binocular fusion is allowed, based on the obtained position of the gaze point;a correction unit configured to correct the three-dimensional image so as to suppress display of an object which is included in the three-dimensional image outside the fusional area;a display unit configured to display the corrected three-dimensional image; anda fusional area information storage unit configured to store fusional area information which indicates positions of a plurality of gaze points in a depth direction of the three-dimensional image and a plurality of fusional areas corresponding to the positions of the plurality of gaze points in the depth direction,wherein the fusional area determination unit is configured to refer to the fusional area information to determine the fusional area that corresponds to the obtained position of the gaze point. 2. The three-dimensional display device according to claim 1, wherein the correction unit is configured to correct the three-dimensional image by removing an object which is included in the three-dimensional image and located on a side closer to the viewer than the fusional area is. 3. The three-dimensional display device according to claim 1, wherein the correction unit is configured to correct the three-dimensional image by blurring an object which is included in the three-dimensional image and located on a side farther away from the viewer than the fusional area is. 4. The three-dimensional display device according to claim 1, further comprising a viewpoint changing unit configured to change a viewpoint of the three-dimensional image so that a display position of an object which is included in the three-dimensional image and located at the position of the gaze point does not change in a depth direction,wherein the correction unit is configured to correct the three-dimensional image the viewpoint of which has been changed. 5. The three-dimensional display device according to claim 4, wherein changing the viewpoint is a process of rotating the three-dimensional image about the position of the gaze point. 6. A three-dimensional display device for displaying a three-dimensional image, comprising: a gaze point obtaining unit configured to obtain a position of a gaze point of a viewer;a fusional area determination unit configured to determine a fusional area where binocular fusion is allowed, based on the obtained position of the gaze point;a correction unit configured to correct the three-dimensional image so as to suppress display of an object which is included in the three-dimensional image outside the fusional area; anda display unit configured to display the corrected three-dimensional image,wherein the three-dimensional image includes a plurality of bl... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | H04N013/00 G06F003/01 G06T019/00 A61B005/00 G09G003/00 H04N013/04 A61B005/0496 A61B005/02
US09795333 | B2 | 2017-10-24 | Blood sampling devices | 2012-01-06 | A blood sampling device includes a housing and a lancet supporting a lancet needle or tip, the lancet and needle being urged or urgeable in a pricking direction. A trigger mechanism is moveable to a fire position for releasing the lancet to travel in the pricking direction. A removable safety cap initially covers the lancet needle. The lancet has an initial position in which movement of the trigger mechanism to the fire position is blocked and an intermediate position in which the trigger mechanism may be moved to the fire position. The intermediate position is obtainable only on removal of the safety cap. | 1. A blood sampling device comprising: a housing (20);a lancet body (30) movable within the housing (20), the lancet body (30) having an end that includes a sharp tip (36), the lancet body (30) having an initial pre-firing position, an intermediate armed position, and a final fired position with the sharp tip (36), after having protruded from a proximal end of the housing (20), is rebounded within the housing with the lancet body (30) forward of the intermediate armed position;an urging mechanism (40) that urges the lancet body and the sharp tip (36) to move, within the housing, forwardly in a pricking direction from the initial pre-firing position, to the intermediate armed position, and then to the final fired position;a trigger mechanism (22) comprised of a trigger member (24) moveable from an initial rest position to a fire position where the trigger mechanism releases the lancet body (30) from the intermediate armed position to travel in the pricking direction to the final fired position, wherein,i) with the lancet body (30) in the initial pre-firing position, the trigger member (24) is engaged with an element (35) of the lancet body (30) such that the element (35) of the lancet body blocks movement of the trigger member (24) from the rest position to the fire position thereby restraining movement of the trigger member (24) and preventing operation of the trigger mechanism, andii) with the lancet body (30) in the intermediate armed position, the trigger member (24) is no longer engaged with the element (35) of the lancet body (30) and operation of the trigger member (24) moves the trigger member (24) from the initial rest position to the fire position where the trigger mechanism (22) releases the lancet body (30) from the intermediate armed position to have the urging mechanism move the lancet body (30) and the sharp tip (36) forwardly in the pricking direction to the final fired position; anda removable safety cap (32) that at least partially covers the sharp tip (36), the safety cap (32) being at least partially located within the housing (20), the safety cap (32) holding the lancet body (30) in the initial pre-firing position and restraining the lancet body (30) from moving, within the housing, in the pricking direction,wherein removal of the safety cap (32) exposes the sharp tip (36) and allows the urging mechanism (40) to move the lancet body (30) from the initial pre-firing position to the intermediate armed position where the device is armed and ready for use by operation of the trigger member (24),wherein the lancet body (30) can only move into said intermediate armed position when the safety cap has been removed,wherein the intermediate armed position of the lancet body (30) is maintained by engagement of a pawl (28) of the trigger mechanism (22) with a latch surface (37) of the lancet body, and operation of the trigger member (24) moving into the fire position lifts the pawl (28) from the latch surface (37) thereby allowing the urg... | OWEN MUMFORD LIMITED | A61B005/151 A61B005/15
US09788894 | B2 | 2017-10-17 | Catheter with distal section having side-by-side loops | 2016-10-17 | The catheter allows mapping and/or ablation of the area around two or more PV ostia at the same time, with a single placement of a distal section of the catheter having a 2D configuration resembling an infinity or lazy 8 symbol. The catheter has an elongated catheter body, a distal section having at least a flexible elongated member with shape memory, the member being configured to assume a 2D configuration resembling an infinity symbol, and at least one electrode mounted on the member. The 2D configuration resembles a first loop and a second loop, wherein the first and second loops are side-by-side, generally extending in a common plane. | 1. A catheter comprising: an elongated catheter body;a distal section distal the catheter body and having a flexible elongated member with shape memory, the member being configured to assume a 2D configuration comprising a first loop in a first direction and a second loop in a second direction opposite the first direction, the 2D configuration terminating in a free distal end of the member; andat least one electrode mounted on the member. 2. The catheter of claim 1, wherein the first and second loops are side-by-side, generally extending in a common plane. 3. The catheter of claim 1, further comprising a deflection section proximal of the distal section. 4. The catheter of claim 3, wherein the 2D configuration and the deflection section are generally in a common plane. 5. The catheter of claim 3, wherein the 2D configuration and the deflection section are generally perpendicular to each other. 6. The catheter of claim 3, wherein the 2D configuration has a common location of the first and second loops. 7. The catheter of claim 6, wherein the common location of the first and second loops is generally aligned with a longitudinal axis of the deflection section. 8. The catheter of claim 6, wherein the common location of the first and second loops is offset from a longitudinal axis of the deflection section. 9. The catheter of claim 1, wherein the first and second loops have different sizes and/or shapes. 10. A catheter comprising: an elongated catheter body having a longitudinal axis;a distal section distal the catheter body and having at least first and second flexible elongated members with shape memory, the at least first and second flexible elongated members being configured to assume a 2D configuration resembling an infinity symbol, wherein the first elongated member is configured to form a first loop terminating in a first free distal end of the first elongated member and the second elongated member is configured to form a second loop terminating in a second free distal end of the second elongated member, the first and second loops being side-by-side in a common plane and having an intersection, the first and second free distal ends being generally at the intersection, and the common plane of the first and second loops being at an angle relative to the longitudinal axis of the catheter body; andat least one electrode mounted on each of the first and second members. 11. The catheter of claim 10, wherein the first and second loops have different sizes and/or shapes. 12. A catheter comprising: an elongated catheter body;a distal section distal the catheter body and having a flexible elongated member with shape memory, the member being configured to assume a 2D configuration with a first loop in a first direction and a second loop in the first direction, the 2D configuration terminating in a free distal end of the member; andat least one electrode on the member. 13. The catheter of claim 12, wherein the first and second loops are side-by... | Biosense Webster (Israel) Ltd. | A61B005/042 A61B018/14 A61B005/00 A61B018/00
US09810587 | B2 | 2017-11-07 | Composite sensor fibres and applications therefor | 2012-08-02 | A composite sensor fiber comprising a filamentary core (12) and an outer layer (30) encapsulating an intermediate sensor layer (20) of a detectably thermoresponsive material. Also disclosed are a method of making composite sensor fibers, a dressing comprising a fabric having a matrix of the fibers, and a method of monitoring a wound that utilizes the dressing. | 1. A composite sensor fibre comprising a filamentary core and an outer layer encapsulating an intermediate sensor layer of a flowable and detectably thermochromic liquid crystalline material between the filamentary core and the outer layer, the thermochromic liquid crystalline material comprising at least one selected from cholesteryl oleyl carbonate, cholesteryl nonanoate, cholesteryl 2, 4-dichlorobenzoate, cholesteryl benzoate, cholesteryl propionate, cholesteryl oleate, and cholesteryl chloride, wherein the outer layer is transparent and the filamentary core is configured to provide a contrast for facilitating the observation of temperature or temperature gradient changes in the thermochromic liquid crystalline material, and wherein the temperature or temperature gradient changes comprise changes in color or color gradient. 2. A composite sensor fibre according to claim 1 wherein the thermochromic liquid crystalline material is a mixture of liquid crystalline compounds selected to achieve a predetermined temperature indication accuracy over a predetermined temperature range. 3. A composite sensor fibre according to claim 1 wherein the filamentary core is of a cross-section selected from generally circular and generally trilobal sections. 4. A composite sensor fibre according to claim 1, wherein the outer layer has a melting temperature below that of the filamentary core. 5. A composite sensor fibre according to claim 1 wherein the filamentary core is a monofilament or multifilament fibre. 6. A composite sensor fibre according to claim 1, wherein the filamentary core is polyester ether ketone (PEEK). 7. A fibrous material formed from or incorporating composite sensor fibres according to claim 1. 8. A fibrous material according to claim 7 comprising a fabric. 9. A clothing item that includes fibrous material according to claim 8. 10. A method of making a composite sensor fibre according to claim 1, comprising continuously delivering the filamentary core from a supply thereof and co-extruding the thermochromic liquid crystalline material about the core and the outer layer about the thermochromic liquid crystalline material at a temperature that facilitates the coextrusion but at which the core remains substantially solid and the outer layer does not inhibit delivery of the thermochromic liquid crystalline material. 11. A method according to claim 10, wherein the thermochromic liquid crystalline material is a mixture of liquid crystalline compounds selected to achieve a predetermined temperature indication accuracy over a predetermined temperature range. 12. A method according to claim 11, wherein the mixture of liquid crystalline compounds is fed by a pump at an elevated temperature selected to enable the mixture to flow at the elevated temperature. 13. A dressing that includes a fabric having a matrix of thermochromic composite sensor fibres according to claim 1, which fibres indicate localised body temperature changes adjacent the d... | Commonwealth Scientific and Industrial Research Organisation | G01K011/00 G01K011/12 D01F008/00 G01K013/00 A61B005/00 D01D011/06 A61B005/01 A61F013/00 D04H003/005 A61F013/84
US09788761 | B2 | 2017-10-17 | Motion correction for magnetic resonance angiography (MRA) with 3D radial acquisitions | 2014-02-27 | A magnetic resonance imaging (MRI) system, method and/or computer readable medium is configured to effect magnetic resonance angiography (MRA) images with reduced motion artifacts includes acquiring a plurality of k-space data sets by traversing a plurality of radial trajectories in three-dimensional (3D) k-space, generating a plurality of 3D MR images derived from k-space populated by the k-space data sets, aligning the 3D MR images with respect to each other, determining one or more motion parameters for the object based upon the aligning, modifying values of k-space data sets using the determined one or more motion parameters, generating a motion-corrected 3D MR image from a combination of acquired k-space data sets including the modified values. | 1. A magnetic resonance imaging (MRI) system for effecting magnetic resonance (MR) images with reduced motion artifacts, said MRI system comprising: an MRI gantry including a static magnetic field coil, gradient magnetic field coils, at least one radio frequency (RF) coil configured to transmit RF signals into an imaging volume and/or to receive nuclear magnetic resonance (NMR) RF signals from an object located in the imaging volume; andat least one digital data processor coupled to receive digitized MR data from said at least one RF coil, said at least one processor being configured to:acquire a plurality of k-space data sets, each k-space data set being acquired by traversing a plurality of radial trajectories of locations in three-dimensional (3D) k-space;generate a plurality of 3D MR images, each 3D MR image derived from a k-space populated by a respective one of the k-space data sets;align the 3D MR images with respect to each other;determine one or more motion parameters for the object based upon the aligning;modify, using the determined one or more motion parameters, values of at least one of the acquired k-space data sets;generate a motion-corrected 3D MR image by constructing it from a set of corrected sets of acquired k-space data including the modified values; andoutput the motion-corrected 3D MR image to a display, a data storage, or to a data transmission port. 2. The MRI system of claim 1, further comprising: a sequence controller configured to apply an MR sequence comprising a plurality of labeling pulses to the object,wherein each of the acquired k-space data sets is acquired after a respective application of the plurality of labeling pulses. 3. The MRI system of claim 2, wherein the MR sequence comprises at least one background suppression radio frequency pulse preceding the plurality of labeling pulses. 4. The MRI system of claim 2, wherein the MR sequence includes a pseudo-continuous arterial spin labeling (PCASL) sequence. 5. The MRI system of claim 1, further comprising an interface for configuring the MR sequence performed by the sequence controller. 6. The MRI system of claim 1, wherein respective ones of the radial trajectories include a center of k-space. 7. The MRI system of claim 6, wherein said each of the radial trajectories include the origin location of k-space. 8. The MRI system of claim 1, further comprising an interface for configuring one or more parameters for 3D radial acquisitions, wherein for each k-space data set, the traversed plurality of radial trajectories correspond to the configured one or more parameters. 9. The MRI system of claim 8, wherein the configured one or more parameters include at least one pattern for three dimensional radial acquisitions, wherein for each k-space data set, the traversed plurality of radial trajectories correspond to the at least one pattern. 10. The MRI system of claim 1, wherein each of the k-space data sets under samples k-space. 11. The MRI system of claim 1,... | TOSHIBA MEDICAL SYSTEMS CORPORATION | G01R033/48 A61B005/11 G01R033/565 A61B005/026 A61B005/00 G01R033/563
US09820810 | B2 | 2017-11-21 | Medical device for use in bodily lumens, for example an atrium | 2015-05-15 | A medical system includes an expandable structure and a first flexible circuit. The expandable structure is sized to be received in a chamber of a heart and is selectively moveable between a delivery configuration and an expanded configuration. The expandable structure includes a plurality of metal strips. The first flexible circuit includes a first flexible electrically insulative substrate and a plurality of electrically conductive traces patterned on the first flexible electrically insulative substrate. A portion of the first flexible circuit is patterned to form at least a first transducer element, and at least a part of the first flexible circuit is at least positioned between two of the plurality of metal strips that contact the first flexible circuit when the expandable structure is in the expanded configuration. | 1. A medical system comprising: an expandable structure sized to be received in a chamber of a heart, the expandable structure selectively moveable between a delivery configuration in which the expandable structure is sized for percutaneous delivery to the chamber and an expanded configuration in which the expandable structure is sized too large for percutaneous delivery to the chamber, the expandable structure comprising a plurality of metal strips; anda first flexible circuit comprising a first flexible electrically insulative substrate and a plurality of electrically conductive traces patterned on the first flexible electrically insulative substrate, wherein:a portion of the first flexible circuit is patterned to form at least a first transducer element, andat least a part of the first flexible circuit is at least positioned between two of the plurality of metal strips that contact the first flexible circuit when the expandable structure is in the expanded configuration. 2. The medical system of claim 1, wherein the plurality of electrically conductive traces are patterned on a portion of a surface of the first flexible electrically insulative substrate, and at least part of the surface of the first flexible electrically insulative substrate is overlapped by one of the two of the plurality of metal strips when the expandable structure is in the expanded configuration. 3. The medical system of claim 1, wherein respective portions of the two of the plurality of metal strips overlap one another when the expandable structure is in the expanded configuration, and a surface of the first flexible electrically insulative substrate on which the plurality of electrically conductive traces are patterned on is positioned between the respective overlapping portions of the two of the plurality of metal strips when the expandable structure is in the expanded configuration. 4. The medical system of claim 1, wherein the plurality of electrically conductive traces are patterned on a portion of a surface of the first flexible electrically insulative substrate, and wherein respective portions of the two of the plurality of metal strips overlap one another when the expandable structure is in the expanded configuration, and at least part of the surface of the first flexible electrically insulative substrate is positioned between the respective overlapping portions of the two of the plurality of metal strips when the expandable structure is in the expanded configuration. 5. The medical system of claim 1, wherein the two of the plurality of metal strips cross one another at a crossing point when the expandable structure is in the expanded configuration. 6. The medical system of claim 5, wherein the at least part of the first flexible circuit is positioned between the two of the plurality of metal strips at the crossing point when the expandable structure is in the expanded configuration. 7. The medical system of claim 5, wherein one of two of the plurality of m... | KARDIUM INC. | A61B018/14 A61B005/06 A61B005/00 A61B018/18 A61B018/20 A61B018/00 A61B018/02
US09808640 | B2 | 2017-11-07 | Method and apparatus for discriminating tachycardia events in a medical device using two sensing vectors | 2014-04-10 | A method and medical device for detecting a cardiac event that includes sensing cardiac signals from a plurality of electrodes forming a first sensing vector sensing a first interval of the cardiac signal during a predetermined time period and a second sensing vector simultaneously sensing a second interval of the cardiac signal during the predetermined time period, identifying each of the first interval and the second interval as being one of shockable and not shockable in response to first processing of the first interval and the second interval and in response to second processing of one or both of the first interval and the second interval, the second processing being different from the first processing, and determining whether to deliver therapy for the cardiac event in response to identifying each of the first interval and the second interval as being one of shockable and not shockable in response to both the first processing and the second processing of the first interval and the second interval. | 1. A method of detecting a cardiac event in a medical device, comprising: sensing cardiac signals from a plurality of electrodes, the plurality of electrodes forming a first sensing vector sensing a first interval of the cardiac signal during a predetermined time period and a second sensing vector simultaneously sensing a second interval of the cardiac signal during the predetermined time period;identifying each of the first interval and the second interval as being one of shockable or not shockable in response to a first processing of the first interval and the second interval;identifying each of the first interval and the second interval as being one of shockable or not shockable in response to a second processing of the first interval and the second interval, the second processing being different from the first processing and comprising a beat-based analysis identifying individual beats within the first interval and the second interval;determining whether to deliver therapy for the cardiac event in response to identifying each of the first interval and the second interval as being one of shockable or not shockable in response to both the first processing of the first interval and the second interval and the second processing of the first interval and the second interval; anddelivering therapy to treat the cardiac event based on identifying each of the first interval and the second interval as being shockable in response to both the first processing and the second processing. 2. The method of claim 1, wherein sensing cardiac signals includes the first sensing vector sensing a plurality of first intervals during the predetermined time period, the plurality of first intervals including the first interval, and the second sensing vector sensing a plurality of second intervals during the predetermined time period, the plurality of second intervals including the second interval, and wherein determining whether to deliver therapy for the cardiac event comprises: determining whether a predetermined number of the plurality of first intervals and the plurality of second intervals are identified as shockable in response to the first processing; anddetermining whether a predetermined number of the plurality of first intervals and the plurality of second intervals are identified as shockable in response to the second processing. 3. The method of claim 1, wherein the first processing comprises a gross morphology analysis of the first interval and the second interval. 4. The method of claim 1, wherein the first processing comprises: determining whether both the first interval and the second interval are corrupted by noise;determining whether both the first interval and the second interval are within one of a VF shock zone and a VT shock zone; anddetermining, in response to one of the first interval and the second interval not being corrupted by noise and the other of the first interval and the second interval being corrupted by noise, whether the one of th... | Medtronic, Inc. | A61B005/04 A61N001/39 A61B005/00 A61B005/042 A61B005/0452 A61B005/046 A61B005/0464
US09791267 | B2 | 2017-10-17 | Determining three-dimensional information from projections or placement of two-dimensional patterns | 2015-05-12 | Embodiments of a shape measurement system and related methods are disclosed. In some embodiments, the system places a two-dimensional initial pattern, which can be implemented as a standalone molding or can be attached to a light source or printed on an outfit, onto a surface of a three-dimensional object. The system captures a transformed version of the initial pattern in two dimensions that is distorted due to the varying depth of the surface. The system then analyzes the transformed pattern and derives three-dimensional information regarding the target object. The analysis, which can incorporate a calibration process, can rely on the projection nature of the light source, the isomorphism/non-isomorphism of the initial pattern, and other factors. | 1. A method of measuring a three-dimensional (3D) target object using two-dimensional (2D) patterns, comprising: selecting an initial two-dimensional pattern;laying the initial pattern on a 3D target object for viewing from a viewpoint to obtain a 2D transformed pattern;computing first depth information regarding the target object based on the transformed pattern;identifying a region of the target object that is hidden from the viewpoint or has a complexity above a predetermined threshold, based on the transformed pattern and first depth information;computing second depth information different from the first depth information regarding the region of the target object; andderiving a 3D measurement of the target object by combining the first depth information and the second depth information and storing the derived 3D measurement in a computer memory. 2. The method of claim 1, wherein the initial pattern is implemented by a standalone molding, the laying comprising:selecting a light source; andpositioning the light source, the molding, and the target object such that light illuminated from the light source shines through the molding before forming the transformed pattern on a surface of the target object. 3. The method of claim 1, wherein the initial pattern is printed on an outfit made of a fabric that stretches and conforms to track a surface of the target object in the outfit, the laying comprising:placing the target object in the outfit so that as the fabric stretches and conforms, the initial pattern stretches into the transformed pattern. 4. The method of claim 3, further comprising: selecting a 2D calibration pattern;laying the calibration pattern on a flat portion of surface of the target object for view from the viewpoint to obtain a 2D stretch pattern; andcomputing a stretch of the stretch pattern from the calibration pattern,wherein the first depth information is computed based on the stretch. 5. The method of claim 3, further comprising wherein the laying includes obtaining multiple 2D transformed patterns corresponding to the target object at multiple positions over a period of time,wherein computing first depth information, identifying a region of the target object, and computing second depth information is performed for each of the multiple transformed patterns, andwherein the first and second depth information is used to track a movement of the target object. 6. The method of claim 1, wherein computing first depth information comprises: selecting a first portion of the initial pattern;determining a second portion of the transformed pattern corresponding to the first portion; andcalculating specific depth information for a region of the target object corresponding to the second portion based on an area of the first portion, an area of the second portion, a length or diameter of the first portion, a curvature of the second portion, or a centroid of the second portion. 7. The method of claim 6, further comprising capturing a port... | - | G01B011/25 A61B005/107 G06T007/00 A41H001/02 A61B005/00 A41H003/04
US09830441 | B2 | 2017-11-28 | Breath actuation of electronic and non-electronic devices for preventing unauthorized access | 2014-08-26 | A mobile device is presented including an input module for receiving a plurality of breath samples from a user and a breath analysis module for performing a chemical analysis of the plurality of breath samples, the chemical analysis involving identification and selection of a plurality of uncommon molecules and uncommon organic compounds for deriving distinguishing breath characteristics and using such breath information to create an initial chemical breath profile associated with the user. The mobile device further includes a breath authorization module for allowing or preventing access to the mobile device in response to a comparison result derived from comparing at least one subsequently created chemical breath profile with the initial chemical breath profile. | 1. A mobile device comprising: an input module for receiving a plurality of breath samples from a user; anda breath analysis module for performing a chemical analysis of the plurality of breath samples, the chemical analysis involving identification and selection of a plurality of uncommon molecules and uncommon organic compounds for deriving distinguishing breath characteristics and using such breath information to create an initial chemical breath profile associated with the user,wherein the initial chemical breath profile of the user is shared with other mobile devices having chemical analysis capabilities controlled by the same user. 2. The mobile device according to claim 1, further comprising a breath authorization module for allowing or preventing access to the mobile device in response to a comparison result derived from comparing at least one subsequently created chemical breath profile with the initial chemical breath profile. 3. The mobile device according to claim 1, further comprising a breath sampling module for sampling the plurality of breath samples. 4. The mobile device according to claim 1, wherein an authorized user of the mobile device is prompted to reset a previously entered initial chemical breath profile associated with the user, the resetting involving inputting a plurality of new breath samples to create a new initial chemical breath profile for the user. 5. The mobile device according to claim 1, wherein the initial chemical breath profile of the user is monitored for deviations, and upon detection of the deviations, an updated chemical breath profile is created for the user by prompting the user to re-enter new breath samples. 6. The mobile device according to claim 1, wherein the initial chemical breath profile associated with the user is stored in the mobile device. 7. The mobile device according to claim 1, wherein the breath analysis module measures a concentration of each of the plurality of uncommon molecules and uncommon organic compounds selected. 8. The mobile device according to claim 1, further comprising a tracking module for recording breath input activity. 9. The mobile device according to claim 1, wherein a plurality of unique initial chemical breath profiles are created for association with the user. 10. A method of accessing a mobile device, the method comprising: receiving a plurality of breath samples from a user;performing a chemical analysis of the plurality of breath samples, the chemical analysis involving identification and selection of a plurality of uncommon molecules and uncommon organic compounds for deriving distinguishing breath characteristics;creating an initial chemical breath profile associated with the user based on such breath information; andsharing the initial chemical breath profile of the user with other mobile devices having chemical analysis capabilities controlled by the same user. 11. The method according to claim 10, further comprising allowing or preventing acces... | - | G01N033/50 G01N033/497 G06F021/32 A61B005/117 A61B005/00
US09814860 | B2 | 2017-11-14 | Accelerometer-based contact sensing assembly and system | 2013-01-14 | A catheter system including an accelerometer-based sensing assembly is provided. In particular the present teachings relate to an accelerometer based assembly used to determine contact between a catheter and surrounding proximate tissue, such as cardiac tissue. An embodiment of such a system may, for example, be used for visualization, mapping, ablation, or other methods of diagnosis and treatment of tissue and/or surrounding areas. | 1. A catheter comprising: a shaft having a proximal portion, and a distal portion;an electrode having a main body extending along an axis and including a tip and an axially opposite base wherein the base is configured to be coupled to said distal portion of the shaft, and wherein the base of said electrode includes a lumen having a proximally-facing opening;a plurality of accelerometers coupled to said electrode along said base and spaced circumferentially around said axis to thereby couple said accelerometers to said electrode such that when said electrode is displaced said accelerometers experience acceleration of said electrode, said accelerometers being positioned to detect said respective acceleration with respect to said axis, said accelerometers each being configured to generate a respective acceleration signal indicative of a directionality of the acceleration of said electrode relative to said axis when said electrode is subjected to an external applied force thereto by tissue in contact with said electrode. 2. The catheter of claim 1, wherein said accelerometer signals are representative of the magnitude of said acceleration. 3. A catheter comprising: a shaft having a proximal portion, and a distal portion;an electrode having a main body extending along an axis and including a tip and an axially opposite base wherein the base is configured to be coupled to said distal portion of the shaft, and wherein the base of said electrode includes a lumen having a proximally-facing opening;an accelerometer disposed in the lumen to thereby couple said accelerometer to said electrode such that when said electrode is displaced said accelerometer experiences an acceleration of said electrode, said accelerometer being configured to generate a signal indicative of the acceleration of said electrode when said electrode is subjected to an external applied force thereto by tissue in contact with said electrode, further comprising a plurality of accelerometers positioned circumferentially around said axis producing respective accelerometer signals indicative of a directionality of the acceleration of said electrode relative to said axis wherein said plurality of accelerometers that are positioned circumferentially are radially offset from said axis by a radial distance. 4. The catheter of claim 3, wherein said accelerometers are located eccentrically with respect to said axis. 5. The catheter of claim 3, wherein said plurality of accelerometers are configured to provide respective signals indicative of at least one of a magnitude, and an orientation of said electrode acceleration relative to said axis. 6. The catheter of claim 1, wherein said accelerometers are electrically isolated from the electrode. 7. The catheter of claim 1, wherein said accelerometers are thermally isolated from said electrode. 8. The catheter of claim 1, wherein said accelerometers are configured to continually monitor the displacement of said electrode. 9. A catheter system,... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B018/14 A61M025/00 A61B005/06 A61N001/365 A61B034/20 A61M025/01
US09775536 | B2 | 2017-10-03 | Method for constructing a stress-pliant physiological electrode assembly | 2014-08-19 | A method for constructing a stress-pliant physiological electrode assembly is provided. An electrode backing is formed from a stretchable woven textile material compatible to contact the skin on at least one surface. A pair of flexile wires is provided to serve as electrode circuit trace and electrode signal pickup. At least one of the flexile wires is sewn into the textile material which provides a stress-pliant malleability. Each of the flexile wires has an electrically-contacting area functioning for electric signal pickup. The electrically-contacting area may be sewn into the woven textile or affixed to the woven textile via conductive adhesives. The stress-pliant physiological electrode assembly is applicable for a wide array of physiological monitors, including ECG monitors, and especially is suitable for long-term wear. The method disclosed is both environmentally friendly and low-cost. | 1. A method for constructing a stress-pliant physiological electrode assembly, comprising: forming an electrode backing from a stretchable woven textile material, the electrode backing having a proximal electrode backing end and a distal electrode backing end, wherein the proximal electrode backing end and the distal electrode backing end each comprise a side adapted to contact skin of a patient;providing a circuit board and affixing the circuit board only to a side of the proximal electrode backing end that is opposite to the contact side of the proximal electrode backing end, wherein the contact side of the proximal electrode backing end is adapted to contact the skin of the patient;interlacing a flexile wire through the stretchable woven textile material, wherein one wire end of the flexile wire is positioned on the contact side of the distal electrode backing end;forming an electrocardiographic electrode signal pickup from the flexile wire end positioned on the contact side of the distal electrode backing end; andterminating other end of the flexile wire as an electrical connection to connect to the circuit board;interlacing a further flexile wire other than the flexile wire through the stretchable woven textile material, wherein one wire end of the further flexile wire is positioned on the contact side of the proximal electrode backing end;forming a further electrocardiographic electrode signal pickup from the further flexile wire end positioned on the contact side of the proximal electrode backing end; andterminating other end of the further flexile wire as other electrical connection to connect to the circuit board. 2. The method according to claim 1, further comprising: interlacing the flexile wire in a spiral pattern adapted to pick up an electrode signal. 3. The method according to claim 1, wherein the flexile wire comprises a conductor core and at least one layer of an insulator surrounding the conductor core. 4. The method according to claim 3, further comprising: removing at least a portion of the at least one layer of the insulator surrounding the one wire end of the flexile wire. 5. The method according to claim 4, wherein the removing of the at least the portion of the insulator is performed by at least one of heating, a laser, cutting, abrasion, and stripping. 6. The method according to claim 5, further comprising: shaping the midsection of the electrode backing between the distal electrode backing and the proximal electrode backing ends into a narrow isthmus tapering inwards from both sides of the midsection. 7. The method according to claim 6, wherein the physiological monitor comprises an electrocardiography monitor. 8. The method according to claim 7, wherein the electrically conductive adhesive comprises hydrocolloid. 9. The method according to claim 8, further comprising: removing at least a portion of the layer of the insulator surrounding the one flexile wire end. 10. The method according to claim 7, wherein the ... | BARDY DIAGNOSTICS, INC. | A61B005/0408
US09820910 | B2 | 2017-11-21 | Biological measurement apparatus and biological stimulation apparatus | 2015-03-12 | A biological measurement apparatus of this invention includes (i) a micro piezoelectric element for vibrating a nerve cell of a subject without coming in contact with the nerve cell, (ii) an electromagnetic wave antenna for receiving an electromagnetic wave generated by the nerve cell vibrated by the micro piezoelectric element, and (iii) a computer for measuring an electric charge of the nerve cell based on the electromagnetic wave received by the electromagnetic wave antenna. Further, this biological measurement apparatus includes an electromagnetic wave antenna for emitting an electromagnetic wave to a nerve cell. This configuration provides an apparatus capable of measuring an electrical activity of a nerve cell in a living organism in real time and three-dimensionally at a spatial resolution of a nerve cell size. Moreover, this configuration provides an apparatus capable of individually giving electrical stimulations to any desired cells in the subject. | 1. A biological stimulation apparatus comprising: one or more micro piezoelectric elements each for externally vibrating a nerve cell of a subject noninvasively; andan electromagnetic wave antenna for emitting an electromagnetic wave to the nerve cell vibrated by the one or more micro piezoelectric elements,wherein the electromagnetic wave has a frequency equal to a vibration frequency of the one or more micro piezoelectric elements,wherein the one or more micro piezoelectric elements are each configured to emit a sound wave so as to vibrate the nerve cell,wherein the sound wave emitted by the one or more micro piezoelectric elements is a pulse-like wave; anda computer that is programmed to execute instructions that cause the electromagnetic wave to be emitted via the electromagnetic wave antenna, only at the moment when the pulse-like wave reaches the nerve cell; and wherein the one or more micro piezoelectric elements are provided on a surface of a substrate which surface faces the subject, the substrate being configured to cover at least part of a surface of the subject. 2. The apparatus of claim 1, wherein the pulse-like wave is an AM-modulated pulse-like wave. | OKAYAMA PREFECTURE | A61B005/00 A61B005/02 A61H001/00 A61H023/02 A61B005/04 A61B005/055 A61B005/05 A61N005/00
US09808203 | B2 | 2017-11-07 | Method for motion correction of emission computed tomography data by way of magnetic resonance tomography data | 2014-05-20 | A method includes introducing the examination object into an examination region of a combination device; recording emission computed tomography data over a measurement period and storing detection events and detection instants associated therewith; measuring magnetic resonance data of at least two subregions of the examination region at at least two instants during the recording period of the emission computed tomography data and storing the magnetic resonance data and the recording instants; determining motion information describing a motion of a region of the examination object at a first instant relative to the position at a second instant from the magnetic resonance data recorded at the first instant and the second instant, for each subregion; determining a motion model describing motion of the examination object, for the entire object, from information for the subregions; and calculating motion-corrected emission tomography data from detection events, detection instants and the motion model. | 1. A method for motion correction of emission computed tomography data in a combination device for magnetic resonance tomography and emission computed tomography, the method comprising: recording emission computed tomography data over a recording period and storing associated detection events and detection instants;performing seperate magnetic resonance measurements for at least two subregions of an examination region of an examination object to record magnetic resonance data of the at least two subregions at at least two recording instants during the recording period of the emission computed tomography data;storing the recorded magnetic resonance data and the at least two recording instants;determining, for each of the at least two subregions, at least one item of motion information describing a motion of a region of the examination object at a first instant relative to a position at a second instant, from magnetic resonance data recorded at the first instant and the second instant;determining a motion model describing motion of the examination object, for the entire examination object, from the items of motion information for the at least two subregions; andcalculating motion-corrected emission tomography data from the detection events, the detection instants and the motion model. 2. The method of claim 1, wherein the magnetic resonance data is recorded one- or two-dimensionally for at least one of the at least two subregions. 3. The method of claim 1, wherein the magnetic resonance data for at least one of the at least two subregions is recorded one-dimensionally using a two-dimensional high-frequency pulse. 4. The method of claim 1, wherein a time interval between two recordings of magnetic resonance data of a first of the at least two subregions differs from a time interval between two recordings of magnetic resonance data of at least a second of the at least two subregions. 5. The method of claim 1, wherein spatial resolutions of the seperate magnetic resonance measurements are different in at least two of the subregions. 6. The method of claim 1, wherein the seperate magnetic resonance measurements of at least two of the at least two subregions are performed at different instants. 7. The method of claim 1, further comprising: exiting at least two subvolumes for the seperate magnetic resonance measurements simultaneously. 8. The method of claim 7, wherein the exciting comprises: using at least two local coils to perform magnetic resonance measurements for the at least two subregions simultaneously, wherein coupling between the at least two subregions and the at least two local coils is based on respective distances between the at least two local coils and the at least two subregions. 9. The method of claim 7, wherein a simultaneous magnetic resonance measurement is performed for at least two of the at least two subregions by overlaying a second readout gradient. 10. The method of claim 1, wherein a time interval between two magneti... | Siemens Aktiengesellschaft Westfaelische Wilhems-Universitaet Muenster | A61B005/00 A61B005/055 A61B006/00 A61B006/03 G06T007/00 G01R033/48 G01R033/565 G01R035/00
US09792681 | B2 | 2017-10-17 | System and method for medical image analysis and probabilistic diagnosis | 2014-04-11 | Methods and systems for obtaining a probabilistic diagnosis based on medical images of a patient are disclosed. In exemplary embodiments, such methods include the steps of scanning some or all of a patient to obtain a medical image; evaluating the medical image for one or more designated key features; assigning discrete values to the one or more designated key features to form a patient scan key feature pattern; and transmitting the values of the one or more designated key features to a processor programmed to match the patient scan key feature pattern to one or more known disease-specific key feature patterns to create a probabilistic diagnosis and transmit the probabilistic diagnosis to a user. Associated systems include a medical imaging device that is capable of producing a medical image of some or all of a patient and a processor programmed to create a probabilistic diagnosis. | 1. A method for obtaining a probabilistic diagnosis of a patient, the method comprising: a medical imaging device scanning some or all of a patient to obtain a medical image;identifying one or more designated key features in the medical image;assigning discrete values to the one or more designated key features in the medical image to form a patient scan key feature pattern;transmitting the patient scan key feature pattern to a processor;the processor implementing an expert system to match the patient scan key feature pattern to one or more known disease specific key feature patterns to create a probabilistic diagnosis; andthe processor transmitting the probabilistic diagnosis to a user. 2. The method according to claim 1, wherein the identifying step is performed by a human. 3. The method according to claim 1, wherein the identifying step is performed by a computer. 4. The method according to claim 1, wherein the values assigned to key features are quantitative. 5. The method according to claim 1, wherein the values assigned to key features are semi-quantitative. 6. A method for obtaining a probabilistic diagnosis of a patient, the method comprising: a medical imaging device scanning some or all of a patient to obtain a medical image;transmitting the medical image to a processor programmed to create a probabilistic diagnosis by:identifying one or more designated key features in the medical image;assigning quantitative values to the one or more designated key features in the medical image to form a patient scan key feature pattern;implementing an expert system that matches the patient scan key feature pattern to one or more known disease specific key feature patterns to create a probabilistic diagnosis; andtransmitting the probabilistic diagnosis to a user. 7. A system for obtaining a probabilistic diagnosis of a patient, the system comprising: a medical imaging device adapted to produce a medical image of some or all of a patient; anda processor that implements an expert system to create a probabilistic diagnosis by matching a patient scan key feature pattern comprising one or more quantitative or semi-quantitative values corresponding to designated key features extracted from the medical image to one or more known disease-specific key feature patterns and transmitting the probabilistic diagnosis to a user. 8. The system according to claim 7, wherein the expert system creates a probabilistic diagnosis by extracting the one or more quantitative or semi-quantitative values corresponding to designated key features from the medical image before matching the patient scan key feature pattern. 9. The system according to claim 8, wherein the processor is programmed to extract quantitative values corresponding to one or more designated key features by: identifying one or more designated key features in the medical image; andassigning quantitative values to the one or more designated key features to form a patient scan key feature pattern. 10. The... | The Trustees of the University of Pennsylvania | G06K009/00 G06T007/00 G06F019/00 G06K009/46 G06K009/52 G06K009/62 G06N007/00 A61B005/00
US09808549 | B2 | 2017-11-07 | System for detecting sterile field events and related methods | 2015-12-23 | A system is for monitoring a sterile field associated with a medical procedure. The system may include a sensor being adjacent an area where the medical procedure is to be performed. The sensor is configured to monitor at least the area, a patient for the medical procedure, and a medical technician for the medical procedure. The system may include a processor coupled to the sensor and configured to detect a sterile field event, and an associated location for the sterile field event, and an output device coupled to the processor and configured to generate an alert indicator when the sterile field event is detected, the alert indicator also including the associated location. | 1. A system for monitoring a sterile field associated with a medical procedure, the system comprising: at least one sensor being adjacent an area where the medical procedure is to be performed;said at least one sensor configured to monitor at least the area, a patient for the medical procedure, and a medical technician for the medical procedure;a processor coupled to said at least one sensor and configured to generate a spatial model of the area, the spatial model comprising a plurality of elements,generate a contamination probability value for each element of the spatial model, anddetect a sterile field event, and an associated location for the sterile field event; andan output device coupled to said processor and configured to generate an alert indicator when the sterile field event is detected, the alert indicator also including the associated location. 2. The system of claim 1 wherein said at least one sensor comprises an image sensor configured to generate at least one of image data and depth data for the sterile field event. 3. The system of claim 1 further comprising at least one tag coupled to at least one of a device associated with the medical procedure, the patient, and the medical technician; and wherein said at least one sensor is configured to detect a location of the at least one of the device associated with the medical procedure, the patient, and the medical technician via an associated tag. 4. The system of claim 1 wherein said output device comprises a display; and wherein said processor is configured to generate indicators of elements having respective contamination probability values exceeding a threshold using said display. 5. The system of claim 1 wherein said processor is configured to iteratively update the contamination probability value for each element of the spatial model during the medical procedure. 6. The system of claim 1 wherein said processor is configured to: generate a gap value between adjacent elements in the plurality of elements of the spatial model; andwhen one of the adjacent elements is contaminated and the gap value is less than a threshold gap value, indicate the other element also as contaminated. 7. The system of claim 1 wherein said processor is configured to detect the sterile field event comprising at least one of a violation and a risk of violation of the sterile field. 8. A system for monitoring a sterile field associated with a medical procedure, the system comprising: at least one sensor being adjacent an area where the medical procedure is to be performed;said at least one sensor configured to monitor at least the area, a patient for the medical procedure, and a medical technician for the medical procedure;at least one tag coupled to at least one of a device associated with the medical procedure, the patient, and the medical technician;said at least one sensor is configured to detect a location of the at least one of the device, the patient, and the medical technician via an associate... | UNIVERSITY OF CENTRAL FLORIDA RESEARCH FOUNDATION, INC. | H04N005/228 H04N007/18 A61L002/28 G06K009/00 A61B005/00
US09808177 | B2 | 2017-11-07 | Systems and methods for automated voxelation of regions of interest for magnetic resonance spectroscopy | 2016-03-04 | A system and method for automating an appropriate voxel prescription in a uniquely definable region of interest (ROI) in a tissue of a patient is provided, such as for purpose of conducting magnetic resonance spectroscopy (MRS) in the ROI. The dimensions and coordinates of a single three dimensional rectilinear volume (voxel) within a single region of interest (ROI) are automatically identified. This is done, in some embodiments by: (1) applying statistically identified ROI search areas within a field of view (FOV); (2) image processing an MRI image to smooth the background and enhance a particular structure useful to define the ROI; (3) identifying a population of pixels that define the particular structure; (4) performing a statistical analysis of the pixel population to fit a 2D model such as an ellipsoid to the population and subsequently fit a rectilinear shape within the model; (5) repetiting elements (1) through (4) using multiple images that encompass the 3D ROI to create a 3D rectilinear shape; (6) a repetition of elements (1) through (5) for multiple ROIs with a common FOV. A manual interface may also be provided, allowing for override to replace by manual prescription, assistance to identify structures (e.g. clicking on disc levels), or modifying the automated voxel (e.g. modify location, shape, or one or more dimensions). | 1. A method for obtaining information relating to a region of interest, the method comprising: accessing a plurality of electronic magnetic resonance imaging (MRI) images of an area that includes a region of interest in an intervertebral disc of a spine;for each of the plurality of electronic MRI images: automatically processing the electronic MRI image, using one or more computer processors, to emphasize pixels associated with two opposite borders between the intervertebral disc and two superiorly and inferiorly adjacent vertebral bodies, respectively;automatically identifying, using the one or more computer processors, a population of pixels in the electronic MRI image associated with the borders; andautomatically fitting, using the one or more computer processors, a two dimensional rectilinear shape inside a cross-section of the region of interest in the intervertebral disc and between the borders shown by the electronic MRI image based on the population of pixels;automatically generating, using the one or more computer processors, world coordinates that define a three dimensional selected volume comprising a voxel that fits inside of the region of interest in the intervertebral disc and excludes the two superiorly and inferiorly adjacent vertebral bodies associated with the borders based on one or more of the two dimensional rectilinear shapes; andscanning the voxel with a magnetic resonance spectroscopy (MRS) system in communication with the one or more computer processors to provide an MRS spectrum of chemical constituents within the voxel. 2. The method of claim 1, comprising calculating, by the one or more computer processors, a two dimensional model that approximates the region of interest based on the population of pixels. 3. The method of claim 2, wherein calculating the two dimensional model comprises applying an expectation maximization algorithm for estimating parameters of one or more Gaussian distributions for the population of pixels. 4. The method of claim 2, comprising orienting, using the one or more computer processors, the two dimensional rectilinear shape based on the orientation of the two dimensional model. 5. The method of claim 2, wherein the two dimensional model comprises an ellipsoid shape having a semi-major axis and a semi-minor axis, and wherein the two dimensional rectilinear shape comprises a rectangle having a length and a width that are oriented based on the orientation of the semi-major axis and/or the semi-minor axis of the ellipsoid shape. 6. The method of claim 1, further comprising outputting information relating to the region of interest based on the world coordinates, the information comprising one or more of a location, an orientation, a shape, an area, and a volume of the region of interest. 7. The method of claim 1, wherein the plurality of electronic MRI images are of slices substantially parallel to, and spaced apart from, each other. 8. The method of claim 1, further comprising analyzing a ... | NOCIMED, INC. | A61B005/055 G01R033/485 G01R033/54 G06K009/34 G06T007/12 A61B005/00 G01R033/483 G06T007/11 G06T007/62
US09788910 | B2 | 2017-10-17 | Instrument-mounted tension sensing mechanism for robotically-driven medical instruments | 2016-06-21 | A tension mechanism for a robotically-controlled medical device measures the tension applied to an actuation tendon to provide feedback to a robotic controller. In one embodiment, the device comprises an elongated instrument, an elongated member, and a base. The elongated member is coupled to the distal end of the elongated instrument, configured to actuate the distal end of the elongated instrument in response to tension in the elongated member. The base is located at the proximal end of the elongated instrument, and comprises a first redirect surface that redirects the elongated member. The first redirect surface is coupled to a lever element that is configured to exert a reactive force on a sensor in response to tension in the elongated member. | 1. A medical device for performing minimally-invasive surgery comprising: an elongated instrument;a first elongated member coupled to the distal end of the elongated instrument, configured to actuate the distal end of the elongated instrument in response to tension in the first elongated member;a base located at the proximal end of the elongated instrument, the base comprising: a first redirect surface that redirects the first elongated member, the first redirect surface comprising a first rotatable body; anda first lever element that is coupled to the first redirect surface and configured to exert a reactive force on a first sensor in response to tension in the first elongated member, the first lever element coupled to the first redirect surface at a pivot point at a first location of the first lever element, the first lever element in contact with the first sensor at a second location of the first lever element, the pivot point of the first lever element offset from a rotation axis of the first rotatable body, wherein an axis of rotation of the first rotatable body is located between the pivot point and the first sensor. 2. The medical device of claim 1, wherein the first redirect surface is low friction. 3. The medical device of claim 1, wherein the base further comprises a second rotatable body, wherein the elongated member is threaded around the second rotatable body. 4. The medical device of claim 3, wherein rotational motion of the second rotatable body is configured to cause tension in the first elongated member. 5. The medical device of claim 3, wherein the second rotatable body comprises splines that receive rotational motion through a sterile interface from a robotic drive mechanism. 6. The medical device of claim 5, wherein the second rotatable body is a male connector. 7. The medical device of claim 5, wherein the second rotatable body is a female connector. 8. The medical device of claim 3, wherein a path of the first elongated member runs from the distal end of the elongated instrument past the second rotatable body, around the first rotatable body, and around the second rotatable body. 9. The medical device of claim 3, wherein a path of the first elongated member is perpendicular to an axis of rotation of the first rotatable body and an axis of rotation of the second rotatable body. 10. The medical device of claim 3, further comprising a second elongated member coupled to the distal end of the elongated instrument, configured to actuate the distal end of the elongated instrument in response to tension in the second elongated member. 11. The medical device of claim 10, the base further comprising: a second redirect surface that redirects the second elongated member, the second redirect surface comprising a third rotatable body; anda fourth rotatable body located between the second redirect surface and the distal end of the elongated instrument, wherein the second elongated member is threaded around the fourth rotatable bo... | Auris Surgical Robotics, Inc. | A61B034/30 A61M025/01 A61B005/00 G01L005/04
US09818283 | B2 | 2017-11-14 | Alert generation correlating between head mounted imaging data and external device | 2015-01-05 | A method of alerting a vehicle's driver. The method comprises receiving an azimuth value indicating a gazing direction of a head of the driver from a primary device placed on the driver, receiving a vehicle bearing value indicating a vehicle moving direction of the vehicle driven by the driver, calculating an angular difference between the azimuth value and the vehicle bearing value, generating an alert when the angular difference deviates from an angular threshold, and automatically presenting the alert to the driver in response to the angular difference. | 1. A method of alerting a vehicle's driver, comprising: with a processor: measuring head movements of the driver by a wearable mount placed on the driver;calculating an azimuth value indicating a horizontal orientation facing direction of a head of said driver;receiving a vehicle bearing value indicating a current vehicle moving direction of a vehicle driven by the driver;calculating an angular difference between the azimuth value and the vehicle bearing value;receiving a first set of image data, the first set of image data captured by a first image capture module coupled to the wearable mount and directing outwardly from the driver to capture a driver's perspective;receiving a second set of image data, the second set of image data captured by a second image capture module mounted on the vehicle and oriented toward the current vehicle moving direction to capture a vehicle's perspective;processing the first and second sets of image data to determine a comparison result, the comparison result determined according to image similarity between the first and second sets of image data;generating an alert according to said comparison result and when the angular difference deviates from an angular threshold, said angular threshold selected from a plurality of predefined angular thresholds according to a traffic condition; andautomatically presenting the alert via an alert module to the driver. 2. The method of claim 1, wherein said generating further comprises generating an alert when the angular difference deviates from the angular threshold and the comparison result deviates from an image similarity threshold; andautomatically presenting the alert to the driver. 3. The method of claim 1, wherein the angular threshold is dynamically modified based on the first and second sets of image data received. 4. The method of claim 1, wherein the vehicle bearing value is generated by a geographical positioning system (GPS); and wherein the vehicle bearing value is transmitted by one or more of a circuit coupled to the vehicle and a mobile device to the wearable mount placed on the driver. 5. The method of claim 1, further comprising: automatically presenting the alert to the driver in response to the angular difference while the angular difference is still present. 6. The method of claim 1, wherein said automatically presenting the alert further comprises providing a suggestion to said driver to navigate to a rest location, said rest location including at least one of a restaurant, a coffee shop, a rest stop, and a parking lot. 7. A system of alerting a driver of a vehicle, comprising: a wearable mount comprising a processor, said wearable mount configured to be placed on a head of said driver measuring head movements of said driver and calculating an azimuth value indicating a horizontal orientation facing direction of the head of said driver;a first image sensor coupled to said wearable mount, the first image sensor directing outwardly from said driver to ... | iOnRoad Technologies Ltd. | G08B023/00 G08B021/06 A61B005/18 B60K028/06 G06K009/00 B60C009/00 H04N007/18 A61B005/00
US09826928 | B2 | 2017-11-28 | Apparatus and method for measuring thermal conductivity in burns | 2014-12-24 | There are provided an apparatus and method for measuring thermal conductivity capable of easily and accurately obtaining an extent of thermal damage of a targeted tissue. The apparatus for measuring thermal conductivity in burns includes a thermal paper stacking member having a plurality of sheets of thermal paper stacked thereon to form layers, and a pressing member configured to press the stacking member so that the thermal paper is stacked and maintained in a closely adhered state. Here, an extent of thermal damage of the targeted tissue according to the depth can be calculated as an extent of thermal damage of the stacking member according to the depth. | 1. An apparatus comprising: a plurality of sheets of thermal paper;a thermal paper stacking member having the thermal paper stacked thereon to form layers; anda plate having the stacking member loaded thereon,wherein the apparatus is able to measure an extent of burning in a state in which the thermal paper is stacked, wherein adhesive oil is interposed between the stacked thermal paper sheets. 2. The apparatus of claim 1, further comprising a pressing member configured to press the thermal paper stacking member so that the plurality of thermal paper sheets are stacked and maintained in a closely adhered state. 3. The apparatus of claim 1, wherein the plate is made of a heat-resistant silicone material. 4. The apparatus of claim 3, wherein the plate is formed on a plastically deformable lower plate. 5. The apparatus of claim 1, wherein the adhesive oil is at least one selected from the group consisting of oil and thermal grease. 6. The apparatus of claim 2, wherein the pressing member is selected from the group consisting of a magnetic material, a clip, a band, and a clamp. 7. The apparatus of claim 1, wherein the stacking member has reference holes formed therethrough to form reference points for three-dimensional (3D) modeling. 8. A method comprising: a thermal paper stacking operation of determining a number of stacked sheets of thermal paper constituting a thermal paper stacking member according to thermal conductivity of a targeted tissue;a preparation operation of preparing the thermal paper stacking member having a plurality of thermal paper sheets and adhesive oil, wherein the adhesive oil is interposed between the stacked thermal paper sheets such that the thermal paper sheets closely adhere together, wherein upon application of heat, the thermal paper stacking member is configured to emulate a thermal distribution pattern into skin tissues;a heating operation of applying heat to be tested to a top surface of the stacking member; anda thermal conductivity estimation operation of estimating thermal conductivity in the tissue by converting an extent of thermal damage according to a thermal depth of the thermal paper piled as the respective layers of the stacking member into an extent of thermal damage according to a depth of the targeted tissue. 9. The method of claim 8, wherein the preparation operation comprises: an upper plate temperature regulation operation of regulating a temperature of an upper plate configured to be placed below the stacking member to maintain a temperature similar to that of the targeted tissue; anda lower plate attachment/pressing operation of, when measurement of a curved surface is required, plastically deforming a lower plate so that the lower plate has the curved surface and attaching the upper plate to the lower plate. 10. The method of claim 8, further comprising: a 3D reconstitution operation of three-dimensionally reconstituting a thermal image formed on the thermal paper piled as the respective ... | Catholic University Industry Academic Cooperation Foundation | G01K011/12 A61B005/00 G01N025/18
US09810686 | B1 | 2017-11-07 | Urinalysis cassette and system | 2016-09-14 | A urinalysis cassette and system that includes a toilet, into which a urinalysis cassette is inserted. The cassette is comprised of a roll of color-change-reagent strips. Each strip includes a plurality of strip sections wherein an absorbent material is imbued with color-change reagent, or onto which one or more color-change reagents are dispensed. The color-change-reagent strips are consecutive and contiguous on the roll. A cassette includes two spools arranged such that when the spools turn, the reagent strip moves from a first (source spool) to a second (waste spool). The section of the reagent strip that is between the two spools may be exposed to a urine specimen. The absorbent-material strips that have been exposed to both a urine specimen and a color-change reagent may undergo a chemical reaction. Selective-lighting illumination may be applied to absorbent-material strips that have undergone such chemical reactions to measure reflectivity of certain color wavelengths. This measurement may be expressed in terms of a digital readout, which may be displayed on a user interface. | 1. A toilet, comprising: a urinalysis cassette, wherein a roll of contiguous strips subject to one or more color-change reagents is rolled onto a plurality of spools;a first spool for distributing a pre-wound roll of contiguous strips;a second spool for receiving and storing strips distributed from the first spool;a light emitter, wherein light projects onto the strips;a light-wavelength sensor, wherein colors on the strips are detectable;a sensor, optical signal, and/or indicator to detect when the roll of contiguous strips are entirely transferred from the first spool to the second spool;a urine-specimen-capture slit disposed within the toilet bowl;a nozzle, wherein the nozzle is disposed within the urine-specimen capture slit;a housing compartment for the urinalysis cassette located below a toilet bowl, the housing compartment comprising an orifice for receiving a urine sample, wherein the orifice is adjacent to the nozzle, wherein the nozzle is in fluid connection with the urine specimen-capture-slit and the orifice; anda wireless communications network for transmitting urinalysis results as a digital readout. 2. The toilet of claim 1, wherein the roll of contiguous strips is comprised of one or more absorbent-material segments on a surface of the contiguous strips. 3. The toilet of claim 1, wherein the roll of contiguous strips comprises an exposed strip segment that traverses a gap situated between the first spool and the second spool. 4. The toilet of claim 3, wherein the urinalysis cassette comprises a mechanism by which the first and second spools synchronously rotate a previously-specified set of degrees such that once the urinalysis of the exposed strip segment is complete, a new segment of the roll of contiguous strips traverses the gap situated between the first spool and the second spool. 5. The toilet of claim 1, further comprising an electrical power source. 6. The toilet of claim 1, further comprising one or more light sources such as light-emitting diodes, photodiodes, and/or cameras, which comprise a broadband source, a tunable narrowband filter, and/or several narrowband emitters. 7. The toilet of claim 1, further comprising a temperature control system for controlling ambient temperature near the roll of contiguous strips during chemical reactions. 8. The toilet of claim 1, further comprising a heat source to locally heat a portion of the roll of contiguous strips that is received onto the second spool such that a strip segment previously exposed to urine evaporates any moisture retained in the strip segment. 9. The toilet of claim 1, further comprising a desiccant, which is applied to a strip segment previously exposed to urine such that a portion of the roll of contiguous strips that is received onto the second spool is dried. 10. The toilet of claim 1, wherein the urinalysis cassette comprises a sanitation and/or purification system for sanitizing segments of the roll of contiguous strips that were exposed to urin... | - | G01N021/00 G01N033/52 A61B005/20 A61B005/00 A61B010/00
US09776008 | B2 | 2017-10-03 | Posture state responsive therapy delivery using dwell times | 2009-04-30 | Techniques related to classifying a posture state of a living body are disclosed. One aspect relates to sensing at least one signal indicative of a posture state of a living body. Posture state detection logic classifies the living body as being in a posture state based on the at least one signal, wherein this classification may take into account at least one of posture and activity state of the living body. The posture state detection logic further determines whether the living body is classified in the posture state for at least a predetermined period of time. Response logic is described that initiates a response as a result of the body being classified in the posture state only after the living body has maintained the classified posture state for at least the predetermined period of time. This response may involve a change in therapy, such as neurostimulation therapy, that is delivered to the living body. | 1. A medical device system, comprising: a sensor configured to sense at least one signal indicative of a posture state of a living body;posture state detection logic configured to classify the living body as being in a posture state based on the at least one signal and to determine whether the living body is classified in the posture state for at least a predetermined period of time; andresponse logic configured to initiate a response as a result of the living body being classified in the posture state, the response being initiated only after the living body has maintained the classified posture state for at least the predetermined period of time. 2. A medical device system, comprising: a sensor configured to sense at least one signal indicative of a posture state of a living body;posture state detection logic configured to classify the living body as being in a posture state based on the at least one signal and to determine whether the living body is classified in the posture state for at least a predetermined period of time;response logic configured to initiate a response as a result of the living body being classified in the posture state, the response being initiated only after the living body has maintained the classified posture state for at least the predetermined period of time; anda therapy delivery module configured to deliver therapy to the living body, and wherein the response is a change in the therapy delivered by the therapy delivery module. 3. The system of claim 2, wherein the posture state detection logic comprises at least one of a posture classifier configured to classify the posture of the living body and an activity classifier configured to classify the activity state of the living body, and wherein the classification of the living body as being in a posture state is based on at least one of the posture classification and the activity state classification. 4. The system of claim 2, wherein the posture state detection logic comprises dwell time logic having one or more timers, each configured to measure a respective dwell time that commences at the time the living body is classified in the posture state; and wherein upon expiration of each of the one or more timers, the response logic initiates a respective response as a result of the body being classified in the posture state only if the living body has not exited the posture state during the respective dwell time. 5. The system of claim 2, wherein the posture state detection logic comprises dwell time logic configured to automatically select a dwell time as the predetermined period of time, selection being based on the posture state in which the living body is classified. 6. The system of claim 5, wherein the dwell time logic is further configured to select the dwell time based on a posture state in which the living body was previously classified. 7. The system of claim 2, wherein the posture state detection logic comprises dwell time logic configured to automatically... | MEDTRONIC, INC. | A61N001/00 A61N001/365 A61B005/11 A61B005/00
US09775529 | B2 | 2017-10-03 | Body-worn pulse oximeter | 2009-09-14 | The invention provides a body-worn system that continuously measures pulse oximetry and blood pressure, along with motion, posture, and activity level, from an ambulatory patient. The system features an oximetry probe that comfortably clips to the base of the patient's thumb, thereby freeing up their fingers for conventional activities in a hospital, such as reading and eating. The probe secures to the thumb and measures time-dependent signals corresponding to LEDs operating near 660 and 905 nm. Analog versions of these signals pass through a low-profile cable to a wrist-worn transceiver that encloses a processing unit. Also within the wrist-worn transceiver is an accelerometer, a wireless system that sends information through a network to a remote receiver, e.g. a computer located in a central nursing station. | 1. A method for determining an oxygen saturation value from a patient, comprising the following steps: (a) positioning a first radiation source, a second radiation source, and a radiation detector on the base of the thumb such that the first radiation source and second radiation source are positioned proximally to one another and the radiation detector is positioned at an included angle of between 35° and 55°, wherein the included angle is the angle formed between a first plane drawn parallel to a planar surface comprising the emitting surface of the first and second radiation sources and a second plane drawn parallel to a planar surface of a circuit board comprising the receiving surface of the radiation detector, and such that a first radiation emitted by the first radiation source and a second radiation emitted by the second radiation source pass through the thumb tissue and are received by the princeps pollicis artery before reaching the radiation detector;(b) measuring a first time-dependent signal by detecting the first radiation emitted by the first radiation source after the first radiation passes through the thumb tissue;(c) measuring a second time-dependent signal by detecting the second radiation emitted by the second radiation source after the second radiation passes through a portion of the patient;(d) measuring a time-dependent motion signal with at least one motion sensor;(e) measuring a time-dependent electrical signal with an electrical sensor comprising at least two electrodes;(f) processing the time-dependent electrical signal to determine at least one reference point;(g) processing the time-dependent motion signal to determine if a motion-related event that affects at least one of the first and second time-dependent signals has occurred that would render the first and second time-dependent signals unreliable for calculation of SpO2;(h) if the motion event that would render the first and second time-dependent signals unreliable for calculation of SpO2 has not occurred, processing both the first and second time-dependent signals from the reference point to determine processed forms of the first and second time-dependent signals; and(i) determining the oxygen saturation value from the processed forms of the first and second time-dependent signals. 2. The method of claim 1, wherein step (g) further comprises processing the time-dependent motion signal to determine a motion-related event selected from the patient's posture, activity level, and degree of motion. 3. The method of claim 2, wherein the motion-related event is posture, and step (g) further comprises determining a parameter corresponding to orientation of the patient's torso. 4. The method of claim 3, wherein the parameter corresponding to orientation of the patient's torso is a vector. 5. The method of claim 4, wherein step (g) further comprises processing signals corresponding to three axes of a motion sensor to determine the vector. 6. The method of claim 5, wher... | SOTERA WIRELESS, INC. | A61B005/1455 A61B005/0245 A61B005/00 A61B005/021 A61B005/022 A61B005/0452 A61B005/11
US09788725 | B2 | 2017-10-17 | Medical image display apparatus and method | 2013-07-08 | The present invention discloses a medical image fusion apparatus and method. The medical image fusion apparatus comprises: a display unit configured to display a plurality of medical images in layers in one window on a screen, wherein the arrangement direction of the plurality of medical images is different from the extension direction of the plane on which each medical image exists; an operation detection unit configured to detect an operation of selecting the medical images to be fused from the plurality of medical images; and a fused image generation unit configured to generate a fused image of the medical images to be fused according to the selection operation, wherein the display unit is further configured to display the fused image in a predetermined region in the window. | 1. A medical image display apparatus, comprising: processing circuitry configured to cause a first display region on a display to display a plurality of modified medical images that are generated by reducing a length, in a predetermined direction of the first display region, of a plurality of medical images of substantially an identical region of a subject, wherein partial areas of the modified medical images overlap with each other in such a manner as to arrange the modified medical images in the predetermined direction;detect a selection operation to select at least two of the displayed modified medical images to be fused, from the plurality of modified medical images; andgenerate a fused image from the at least two medical images corresponding to the at least two selected modified medical images selected by the selection operation,wherein the processing circuitry is further configured to cause the fused image to be displayed in a second display region on the display, which is different from the first display region on the display. 2. The medical image display apparatus according to claim 1, wherein the processing circuitry is further configured to cause the display to display the fused image in the first display region, wherein a partial area of the fused image is overlapped with at least one of the plurality of modified medical images. 3. The medical image display apparatus according to claim 2, wherein the processing circuitry is further configured to cause the display to display the fused image in the foreground in the first display region. 4. The medical image display apparatus according to claim 1, wherein the processing circuitry is further configured to cause the display to display the fused image and the modified medial images corresponding to the medical images from which the fused image is generated, wherein both of the fused image and the medical images are each displayed with an identical identifier. 5. The medical image display apparatus according to claim 1, wherein the processing circuitry is further configured to detect a cancelling operation to cancel a display of the fused image in the second display region, and cancel the display of the fused image when the cancelling operation is detected. 6. The medical image display apparatus according to claim 5, wherein the processing circuitry is further configured to detect a drag-and-drop operation to move the fused image out of the second display region as the cancelling operation. 7. The medical image display apparatus according to claim 1, wherein the processing circuitry is further configured to detect a layer interval changing operation to change an interval between two layers of a plurality of layers of the modified medical images displayed in the first display region, andcause the display to display in the first display region the two layers of the modified medical images at the interval changed by the layer interval changing operation when the layer interval changing op... | Toshiba Medical Systems Corporation | G09G005/30 A61B005/00 G06T019/00 G06T011/00 A61B006/00
US09827388 | B2 | 2017-11-28 | Methods and apparatus for detecting and treating respiratory insufficiency | 2010-03-31 | Devices and systems provide methods of detecting a severity change in respiratory insufficiency (RI) or chronic obstructive pulmonary disease (COPD) condition of a patient. In an example embodiment, a detection monitoring device determines one or more severity change indicators based on a measure of supplied pressure or other representative measure determined by the device. The supplied pressure may optionally be determined during pressure treatment that satisfies a target ventilation. The supplied pressure or representative data may be compared to one or more thresholds that are selected to represent a change in the condition of the RI or COPD patient such as an exacerbation of a prior condition. Results of the comparisons may trigger one or more warnings or messages to notify a patient or physician of a pending change to the patient's RI or COPD condition so that the patient may more immediately seek medical attention to treat the condition. | 1. A device for monitoring a patient to evaluate a respiratory insufficiency condition of the patient comprising: a transducer to generate a signal representative of a pressure delivered by a flow generator; anda processor, coupled with the transducer, the processor configured to control an evaluation of a severity change indicator based on detecting, using a measure of determined pressure from the signal, a change in the measure of determined pressure from a plurality of treatment sessions, the evaluation comprising a detection of a continuous rise in the measure of determined pressure successively through the plurality of treatment sessions, the indicator representing information about a change in severity of a respiratory insufficiency condition of the patient,wherein the respiratory insufficiency condition is attributable to an insufficiency condition of the patient's lungs. 2. The device of claim 1 further comprising the flow generator configured to generate a flow of breathable gas at a pressure above atmospheric pressure to a patient interface, wherein the processor is configured to control the flow generator. 3. The device of claim 2 wherein the flow generator is configured to generate the flow of breathable gas at a pressure above atmospheric pressure to maintain a target ventilation, and wherein the measure of determined pressure comprises a measure of pressure generated by the flow generator in order to maintain the target ventilation. 4. The device of claim 3 wherein the target ventilation is a minute ventilation and the measure of determined pressure is an average pressure. 5. The device of claim 3 wherein the target ventilation is an estimate of alveolar ventilation. 6. The device of claim 3 wherein the evaluation of the severity change indicator comprises a threshold comparison of the continuous rise in the measure of determined pressure successively through the treatment sessions. 7. The device of claim 6 wherein the threshold comparison involves the measure of determined pressure from one of the plurality of treatment sessions and a prior measure of determined pressure used to maintain the target ventilation during a prior treatment session of the plurality of treatment sessions. 8. The device of claim 6 wherein the threshold comparison represents an indication of a repeated rise in the measure of determined pressure over a period of days. 9. The device of claim 6 wherein the threshold comparison involves a difference between the measure of determined pressure from one of the plurality of treatment sessions and a prior measure of determined pressure used to maintain the target ventilation during a prior treatment session of the plurality of treatment sessions. 10. The device of claim 6 wherein the threshold comparison includes a ratio of the measure of determined pressure from one of the plurality of treatment sessions and a prior measure of determined pressure used to maintain the target ventilation during a prior treat... | ResMed Limited | A61B005/08 A61M016/00 A61B005/00
US09826929 | B2 | 2017-11-28 | Method and device for stuttering alleviation | 2013-01-17 | A device for stuttering alleviation is disclosed, comprising a speech sensor, configured to output signals indicative of speech, a processing unit configured to detect stuttering, log stuttering and/or produce stimulation indication based on stimulation rules Said device may further comprise a remote server and a server user interface, configured to allow the speech therapist access to the processing unit. Further provided is a method for accelerating the learning procedure for obtaining a permanent fluent speech, comprising receiving and analyzing speech parameters, determining whether a negative reinforcement is required and executing the negative reinforcement. | 1. A speech training device comprising: a speech sensor comprising a microphone, said speech sensor configured to output a signal indicative of speech; anda processing unit configured to: receive the signal indicative of speech;detect stuttering in the signal indicative of speech by searching for prolongations, repetitions, blocking and stress in the speech;determine a level of stuttering based on an analysis of the detected stuttering in the signal indicative of speech; andproduce a stimulation indication by applying stimulation rules on the determined level of stuttering, wherein the stimulation comprises a visual stimulation indicative of the determined level of stuttering. 2. The device according to claim 1, wherein the speech sensor further comprises a vocal cord compartment, a breathing compartment or a combination thereof. 3. The device according to claim 2, wherein the vocal cord compartment comprises a vocal cord activity sensor. 4. The device according to claim 3, wherein the vocal cord activity sensor comprises electromyograph (EMG), electroglottograph (EGG) or a combination thereof. 5. The device according to claim 2 wherein the breathing compartment comprises a lung activity sensor. 6. The device according to claim 1, wherein the processing unit is a standalone unit. 7. The device according to claim 1, wherein the processing unit is an application configured to run on a computer or a mobile device. 8. The device according to claim 7, wherein the mobile device comprises operating system, comprising iOS (Apple), Android (Google), Windows (Microsoft), QNX (RIM), Symbian (Accenture), or Linux. 9. The device according to claim 1, wherein the processing unit is configured to detect stuttering in a real time. 10. The device according to claim 1, wherein the processing unit is configured to allow determination of stuttering detection parameters. 11. The device according to claim 10, wherein the stuttering detection parameters comprise threshold length of prolongation, threshold number of repetitions, threshold length of blocking, threshold level of stress or any combination thereof. 12. The device according to claim 1, wherein the processor is further configured to log the detected stuttering; wherein stuttering logging comprises saving and/or outputting the detected stuttering. 13. The device according to claim 12, wherein the stuttering log comprises word spoken, phoneme spoken, stuttering mode, stuttering severity, stuttering confidence level, date, time of day, location or any combination thereof. 14. The device according to claim 12, wherein the stuttering log may be used for analysis by a speech therapist. 15. The device according to claim 1, further comprising a stimulator; wherein the stimulator is a standalone unit. 16. The device according to claim 1, further comprising a stimulator wherein the stimulator and the speech sensor form one unit. | NINISPEECH LTD. | A61F005/58 A61B005/00 A61B005/0488 A61B005/08 A61B007/04
US09826913 | B2 | 2017-11-28 | Non-invasive intracranial pressure monitoring system and method thereof | 2013-07-11 | A system which includes a first sensor placed proximate to a perfusion field of an artery receiving blood which emanates from the cranial cavity is configured to monitor pulsations of the artery receiving blood which emanates from the cranial cavity artery. A second sensor placed proximate to a perfusion field of an artery which does not receive blood emanating from the cranial cavity and approximately the same distance from the heart as the first sensor configured to monitor pulsations of the artery which does not receive blood emanating from the cranial cavity. A third sensor placed distally from a heart is configured to monitor pulsations of a distal artery. A processing system responsive to signals from the first, second, and third sensors is configured to determine intracranial pressure. | 1. A non-invasive intracranial pressure monitoring system comprising: a first sensor adapted to be placed proximate to a perfusion field of an artery receiving blood which emanates from the cranial cavity configured to monitor pulsations of the artery receiving blood which emanates from the cranial cavity artery;a second sensor adapted to be placed proximate to a perfusion field of an artery which does not receive blood emanating from the cranial cavity and placed approximately the same distance from the heart as the first sensor configured to monitor pulsations of the artery which does not receive blood emanating from the cranial cavity;a third sensor adapted to be placed distally from the heart configured to monitor pulsations of a distal artery; anda processing subsystem responsive to signals from the first, second, and third sensors configured to determine an indication of intracranial pressure by determining a first correlation between signals from the first sensor to signals from the third sensor and determining a second correlation between signals from the second sensor to signals from the third sensor and combining the determined first and second correlations. 2. The system of claim 1 in which the first sensor is placed on the forehead. 3. The system of claim 1 in which the second sensor is placed on the temple. 4. The system of claim 1 in which the second sensor is placed on the ear. 5. The system of claim 1 in which the third sensor is placed distally on a finger. 6. The system of claim 1 in which the third sensor is placed distally on a hand. 7. The system of claim 1 in which the third sensor is placed distally on a forearm. 8. The system of claim 1 further including a display coupled to the processing subsystem configured to display the intracranial pressure. 9. A non-invasive intracranial pressure monitoring system comprising: a first sensor adapted to be placed proximate to the supraorbital artery configured to monitor pulsations of the supraorbital artery;a second sensor adapted to be placed proximate to a branch of the external carotid artery and placed approximately the same distance from the heart as the first sensor configured to monitor pulsations of the external carotid artery;a third sensor adapted to be placed distally from the heart configured to monitor pulsations of a distal artery; anda processing subsystem responsive to signals from the first, second, and third sensors configured to determine an indication of intracranial pressure determining a first correlation between signals from the first sensor to signals from the third sensor and determining a second correlation between signals from the second sensor to signals from the third sensor and combining the determined first and second correlations. 10. The system of claim 9 in which the first sensor is adapted to be placed on the forehead. 11. The system of claim 9 in which the second sensor is adapted to be placed on the temple. 12. The system of claim 9 in ... | Vivonics, Inc. | A61B005/03 A61B005/00 A61B005/024
US09826945 | B2 | 2017-11-28 | Fail safe radiation concealment mechanism | 2011-06-09 | An aspect of an embodiment of the invention, relates to an imaging capsule for scanning inside a living body with a fail-safe radiation mechanism that prevents the emission of radiation from the imaging capsule until the imaging capsule is instructed to emit radiation and power is available to activate a motor to unblock the emission of radiation. Optionally, when power is not available the imaging capsule automatically, blocks the emission of radiation. | 1. An imaging capsule for scanning inside a living body, with a fail-safe radiation mechanism, comprising: an encasement for holding and protecting elements forming the imaging capsule; wherein the encasement and elements are configured to be swallowed by a user and flow through a gastrointestinal tract of the user without damaging the elements;the elements comprising:a power source;a radiation source comprising a radioisotope that emits Xrays and/or gamma rays;a rotatable disk with the radiation source mounted on the rotatable disk and wherein the rotatable disk has a collimator structure allowing the emission of radiation from the radiation source from one or more locations on the circumference of the rotatable disk;an outer ring surrounding the circumference of the rotatable disk and configured to rotate relative to the rotatable disk; the outer ring includes areas with a material that blocks radiation and areas with a material that does not block radiation; wherein the areas that block radiation on the outer ring are positioned so that there exists a rest position of the outer ring relative to the rotatable disk in which the areas that block radiation on the outer ring coincide with the one or more locations on the circumference of the rotatable disk that do not block the emission of radiation;a motor powered by the power source for rotating the rotatable disk relative to the outer ring; wherein the rotatable disk and outer ring are initially in the rest position blocking the emission of radiation until the motor is activated to rotate the rotatable disk and allow the emission of radiation; anda spring coupling the rotatable disk to the outer ring, wherein the spring is configured to automatically return the rotatable disk and the outer ring to the rest position when the motor is deactivated, thus providing a fail-safe radiation mechanism. 2. An imaging capsule according to claim 1, further comprising flaps extending from the outer ring and an encasement with an inner lining enclosing the imaging capsule, wherein the flaps are in contact with the inner lining of the encasement and are held by a force that prevents the outer ring from rotating responsive to the torque of the spring and the rotation of the rotatable disk. 3. An imaging capsule according to claim 2, wherein the force between the flaps and the inner lining is a friction force. 4. An imaging capsule according to claim 2, wherein the force between the flaps and the inner lining is an electromagnetic force. 5. An imaging capsule according to claim 2, wherein the force between the flaps and the inner lining is controllable. 6. An imaging capsule according to claim 5, wherein if the motor is deactivated and the force between the flaps and the inner lining is turned off, the outer ring will rotate to return the rotatable disk and outer ring to the rest position. 7. An imaging capsule according to claim 5, wherein if the motor is deactivated and the force between the flaps and the... | CHECK-CAP LTD. | A61B006/00 A61B005/00 A61B001/00 A61B001/04 A61B006/06
US09802080 | B2 | 2017-10-31 | Group performance monitoring system and method | 2016-02-05 | The present invention provides a group monitoring device for monitoring a plurality of individuals engaged in an athletic activity, the device including a display configured to display, during an athletic activity: a metric relating to each of a plurality of individuals engaged in the athletic activity, and a status of a system component used to monitor the athletic activity. The group monitoring device may also include an input configured to allow manipulation of the display. | 1. A group monitoring device for monitoring a plurality of individuals engaged in an athletic activity, the device comprising: a display configured to display, during an athletic activity: a metric relating to each of a plurality of individuals engaged in the athletic activity, anda plurality of system component statuses, each relating to a different one of a plurality of system components used to monitor the athletic activity; andan input configured to allow manipulation of the display,wherein the display is configured to display the plurality of statuses simultaneously,wherein each of the plurality of system components is remote from the display, andwherein the plurality of system components comprises at least one of a base station, an individual monitor, an object monitor, and a sensor. 2. The device of claim 1, wherein at least one of the system components is a base station. 3. The device of claim 1, wherein at least one of the system components is an individual monitor. 4. The device of claim 1, wherein at least one of the system components is an object monitor. 5. The device of claim 1, wherein at least one of the system components is a sensor in communication with an individual monitor. 6. The device of claim 1, wherein at least one of the system components is a sensor in communication with an object monitor. 7. The device of claim 1, wherein at least one of the statuses is an indication of remaining battery power of a battery of at least one of the system components. 8. The device of claim 7, wherein the at least one of the system components is a base station. 9. The device of claim 7, wherein the at least one of the system components is an individual monitor coupled to an individual engaged in the athletic activity. 10. The device of claim 7, wherein the at least one of the system components is an object monitor coupled to a sports object. 11. The device of claim 7, wherein the status is displayed in response to the status passing a system alert threshold. 12. The device of claim 1, wherein at least one of the statuses is an indication of a number of remote monitors docked at a base station. 13. The device of claim 12, wherein the status is displayed in response to the status passing a system alert threshold. 14. The device of claim 1, wherein at least one of the statuses is an indication of a GPS signal strength at a remote monitor. 15. The device of claim 14, wherein the remote monitor is an individual monitor coupled to an individual engaged in the athletic activity. 16. The device of claim 14, wherein the remote monitor is an object monitor coupled to a sports object. 17. The device of claim 14, wherein the status is displayed in response to the status passing a system alert threshold. 18. The device of claim 1, wherein at least one of the statuses is an indication of proper operation of at least one of the system components. 19. The device of claim 18, wherein the at least one of the system components is an individ... | adidas AG | A63F009/24 A63B024/00 G06F019/00 G06Q010/06 G06F011/30 G06F011/34 A61B005/00 A61B005/0205 A61B005/11 A61B005/024
US09782109 | B2 | 2017-10-10 | Apparatus for measuring the luminance and temperature of a light source of a spectrophotometric device | 2012-07-06 | A system for measuring the luminance and temperature of a light source element of a spectrophotometric device is provided. The system includes a photodiode, a luminance measuring device, a temperature measuring device, a first switch, and a controller. The photodiode is operable to receive light signals emitted by the light source element and passing through a subject's body tissue. The luminance measuring device is operable to measure luminance of the light signals received by the photodiode. The temperature measuring device is operable to measure the temperature of the photodiode. The first switch is operable to connect the photodiode to the luminance measuring device or the temperature measuring device. The controller is operable to control the connection of the first switch. | 1. A system for measuring the luminance and temperature of a light source element of a spectrophotometric device, the system comprising: a photodiode configured to receive light signals emitted by the light source element and passing through a subject's body tissue, and to produce first signals;a luminance measuring device configured to measure luminance of the light signals received by the photodiode, using the first signals from the photodiode;a temperature measuring device configured to measure the temperature of the photodiode, using the first signals from the photodiode;a first switch configured to create a first communication path for the first signals to travel from the photodiode to the luminance measuring device or a second communication path for the first signals to travel from the photodiode to the temperature measuring device;a controller configured to control the first switch to create the first communication path for the first signals to travel from the photodiode to the luminance measuring device or the second communication path for the first signals to travel from the photodiode to the temperature measuring device. 2. The system of claim 1, further comprising: a second switch configured to create a third communication path between the spectrophotometric device and the luminance measuring device or a fourth communication path between the spectrophotometric device and the temperature measuring device;wherein the controller is further configured to control the second switch to create the third communication path between the spectrophotometric device and the luminance measuring device or the fourth communication path between the spectrophotometric device and the temperature measuring device. 3. The system of claim 2, wherein the controller is configured to simultaneously control the first and second switches. 4. The system of claim 1, wherein the luminance measuring device comprises a photovoltaic circuit. 5. The system of claim 1, wherein the luminance measuring device comprises a photoconductive circuit. 6. The system of claim 1, wherein the temperature measuring device comprises a circuit configured to simulate a constant current source. | CAS Medical Systems, Inc. | A61B005/1455
US09782079 | B2 | 2017-10-10 | Device for detecting and illuminating the vasculature using an FPGA | 2015-05-28 | A laser based vascular illumination system utilizing a FPGA for detecting vascular positions, processing an image of such vasculature positions, and projecting the image thereof onto the body of a patient. | 1. A laser-based vasculature illumination system comprising: a first laser configured to output a beam of light at an infrared wavelength configured to image subcutaneous veins;a second laser configured to output a beam of light at a red wavelength, to be co-axially aligned with said beam of infrared light;an x-direction mirror configured to reflect said coaxial beam of light, and to be pivotable about a first axis, in a first direction and in a second direction;an x-direction mirror driver configured to drive said x-direction mirror to oscillate about said first axis, to cyclically reflect said coaxial beam of light in a line, in both said first and second directions;a y-direction mirror configured to reflect said line of light received from said x-direction mirror, and to be pivotable about a second axis, in a third direction and a fourth direction;a y-direction mirror driver configured to drive said y-direction mirror to oscillate about said second axis;a field programmable gate array (FPGA) configured to control said x-direction mirror driver and said y-direction mirror driver to control said oscillations about said first and second axes to form a raster pattern of said red and infrared wavelengths of light;a feedback means configured to detect a position of said x-direction mirror and said y-direction mirror, and to signal said positions to said FPGA;a photodiode configured to receive a vasculature image formed from said infrared light of said raster pattern, said photodiode further configured to convert said vasculature image into an analog signal;wherein said second laser driver is further configured to receive said analog signal, and to drive said red laser to project said vasculature image using said analog signal;a line correlator configured to measure a shift in convergence between said line of light in said first direction and said second direction, for each said oscillation of said X-direction mirror;a mirror convergence control configured to receive said measured shift in convergence from said line correlator, and to adjust said control of said first mirror driver by said FPGA, for said line in said first direction to converge with said line in said second direction. 2. The laser-based vasculature illumination system according to claim 1, further comprising: a second feedback means, said second feedback means configured to detect said output of said first and second lasers, and to signal said detected output to said FPGA, said FPGA further configured to thereby control said output of said first and second lasers. 3. The laser-based vasculature illumination system according to claim 2, further comprising a microprocessor configured to redundantly receive and monitor said feedback signal for said first and second lasers, and said feedback signal for said X-direction mirror and Y-direction mirror, in conjunction with said FPGA. 4. The laser-based vasculature illumination system according to claim 1, further comprising: an analog-to-... | AccuVein, Inc. | A61B006/00 A61B005/00 A61B005/02 A61B005/107
US09825471 | B2 | 2017-11-21 | Resonant power transfer systems with protective algorithm | 2013-07-29 | Systems for tuning a wireless power transfer system are provided, which may include any number of features. In one embodiment, a TET system includes a receive resonator is adapted to be implanted in a human patient and is configured to receive wireless power from a transmit resonator. The system can include a controller configured to identify if a foreign object is interfering with the transmission of power or generating an induced voltage in the receive resonator. The controller can also be configured to control the transmit resonator to phase match with the foreign object. Methods of use are also provided. | 1. A method of protecting a wireless power transfer system from interference by an object that is foreign to the wireless power transfer system, comprising the steps of: transmitting wireless power from a transmit resonator to a receive resonator implanted in a human patient;detecting, by the receive resonator, the object that is interfering with the transmission of power or generating an induced voltage in the receive resonator; andcontrolling the transmit resonator to phase match with the object. 2. The method of claim 1 further comprising indicating to the patient that the foreign object is interfering with transmission of power. 3. The method of claim 1 wherein the detecting step further comprises halting the transmission of power, then determining if the receive resonator receives a foreign wireless power signal. 4. The method of claim 1 wherein the detecting step further comprises detecting a return energy transmitted from the receive resonator to the transmit resonator, and comparing the return energy to an expected return energy value. 5. The method of claim 1 wherein the detecting step further comprises detecting a wireless power signal in the receive resonator in approximately the same frequency band as an operating frequency of the wireless power transfer system. 6. A wireless power transfer system, comprising: a transmit resonator configured to transmit wireless power;a receive resonator adapted to be implanted in a human patient and configured to receive wireless power from the transmit resonator; anda controller connected to the transmit resonator and configured to detect a foreign transmitter that is interfering with the transmission of power or generating an induced voltage in the receive resonator, the controller configured to control the transmit resonator to phase match with the foreign transmitter. 7. The system of claim 6 further comprising a receive controller connected to the receive resonator. 8. The system of claim 7, wherein the receive controller is configured to detect the induced voltage in the receive resonator. 9. The system of claim 6, wherein the controller comprises a signal generator configured to adjust a transmit frequency of the transmit resonator to compensate for changes in coupling between the transmit resonator and the receive resonator. 10. The system of claim 6, wherein the controller comprises an amplifier configured to adjust a gain applied to transmissions from the transmit resonator to compensate for changes in coupling between the transmit resonator and the receive resonator. 11. The system of claim 6, wherein the controller is further configured to turn off the transmit resonator when energy losses due to interference by the foreign transmitter are large. 12. The system of claim 6, wherein the controller is further configured to initiate an alarm to the human patient upon detecting the foreign transmitter is interfering. 13. The system of claim 6, wherein the controller is further conf... | THORATEC CORPORATION | H02J005/00 H02J017/00 A61B005/00
US09801579 | B2 | 2017-10-31 | Arousal-level determining apparatus and arousal-level determining method | 2014-09-08 | An arousal-level determining apparatus includes a generating unit, a calculating unit, an identifying unit, and a determining unit. The generating unit generates heartbeat-interval variation data, which indicates changes in heartbeat interval, on the basis of heartbeat signals indicating subject's heartbeats. The calculating unit applies a band-pass filter, which allows passage of a certain range of frequencies, to each frequency band in the heartbeat-interval variation data while changing the frequency band, and calculates spectral density with respect to each frequency band applied with the band-pass filter. The identifying unit identifies a feature point corresponding to a spectral density peak in the calculated spectral densities in the frequency bands. The determining unit determines subject's arousal level on the basis of the identified feature point. | 1. An arousal-level determining apparatus comprising: a memory; anda processor coupled to the memory executes a process configured to:receive heartbeat signals for a subject;generate heartbeat-interval variation data, which indicates changes in heartbeat intervals, on a basis of the heartbeat signals indicating the subject's heartbeats;apply a plurality of times a band-pass filter, which allows passage of a frequency band of a certain range of frequencies, to the heartbeat-interval variation data, the frequency band being changed for each of the plurality of times, and calculate spectral densities with respect to each frequency band, by overlapping a range of frequencies within a first frequency band of the band-pass filter with a second frequency band of the band-pass filter that is nearest to the first frequency band, to detect peaks respectively of the calculated spectral densities;identify as a feature point a peak in the calculated spectral densities based on the detected peaks; anddetermine the arousal level of the subject on a basis of a spectral density and a frequency of the identified feature point, the arousal level inversely indicating a degree of sleepiness of the subject. 2. The arousal-level determining apparatus according to claim 1, wherein the process is configured to apply the band-pass filter that allows passage of a range of frequencies in which the feature point shifts with respect to a change in arousal level of the subject. 3. The arousal-level determining apparatus according to claim 1, wherein the process is configured to apply the band-pass filter to the heartbeat-interval variation data within a frequency range of a HF component. 4. The arousal-level determining apparatus according to claim 1, wherein the process is configured to calculate spectral densities from the heartbeat-interval variation data, and, when no local maximum point exists in the calculated spectral densities, apply the band-pass filter to the heartbeat-interval variation data. 5. The arousal-level determining apparatus according to claim 1, wherein the determining includes determining a change in the arousal level based on a change in a spectral density and a frequency of the identified feature point and in a width of a pulse made up of the detected peaks, and determining a risk level based on the determined arousal level and change therein; andthe process is further configured to output information or an operation to alert the subject to the determined risk level, the information or operation including a message displayed on a display unit or an alarm. 6. A non-transitory computer-readable recording medium having stored therein an arousal-level determining program causing a computer to execute a process comprising: receiving heartbeat signals for a subject;generating heartbeat-interval variation data, which indicates changes in heartbeat intervals, on a basis of the heartbeat signals indicating the subject's heartbeats;applying a plurality of t... | FUJITSU LIMITED | A61B005/18 A61B005/0402 A61B005/00 A61B005/024 A61B005/0245 A61B005/04
US09795311 | B2 | 2017-10-24 | Method for intraluminal pressure detection for diverticular disease | 2016-11-18 | A pressure detection tablet is disclosed. The tablet is a capsule including a dye encapsulated within an inner coating, in which the inner coating is configured to rupture at a pressure equal to or greater than a pressure threshold. The tablet further includes an outer coating surrounding the capsule, in which the outer coating is configured to provide delivery of the capsule to a targeted site. The dye may be a non-metabolized, water soluble dye, such that a ruptured capsule releases the dye which can be visually detected in the urine. In some embodiments, the disclosure provides a method of making the tablet, using the tablet to detect peak pressure in the colon and reducing the risk of developing a colonic disorder. | 1. A method for estimating a peak pressure within the colon of a subject, the method comprising: providing a pressure detection tablet comprising a capsule, wherein the capsule comprises: a water-soluble and non-metabolizable dye encapsulated within an inner coating, wherein the inner coating is configured to rupture at a pressure equal to or higher than a pressure threshold, andan outer coating surrounding the capsule, wherein the dye is absorbed by the colon and excreted by the kidneys in urine upon rupture of the inner coating;orally administering the pressure detection tablet to the subject;observing a color of the subject's urine after a test period; andestimating the peak pressure within the colon of the subject during the test period based on the color of the subject's urine. 2. The method of claim 1, wherein the outer coating is configured to provide targeted delivery of the capsule to the colon. 3. The method of claim 1, wherein a normal urine color is indicative of the peak pressure within the colon being lower than the pressure threshold and a urine color attributable to the dye is indicative of the peak pressure within the colon being equal to or higher than the pressure threshold. 4. The method of claim 1, wherein the inner coating comprises a pressure-sensitive polymer, and wherein the pressure threshold of the inner coating is selectable based on variations in at least one of a thickness, density and elasticity of the inner coating. 5. The method of claim 1, wherein the outer coating comprises a pH sensitive polymer that is insoluble at the acidic pH of the stomach and proximal small intestine and comprises a copolymer of EUDRAGIT L and EUDRAGIT S polymers and formulated to dissolve at a neutral pH, or comprises azo polymers of styrene and 2-hydroxy methyl acrylate formulated to dissolve by bacterial degradation, or comprises a combination thereof. 6. The method of claim 1, wherein the tablet comprises at least two capsules surrounded by the outer coating and wherein the at least two capsules comprise different dyes and are configured to rupture at equal to or higher than a different pressure threshold from one another. 7. The method of claim 1, wherein the dye is water-soluble and non-metabolizable. 8. The method of claim 1, wherein the outer coating comprises a pH sensitive polymer that is insoluble at the acidic pH of the stomach and proximal small intestine. 9. The method of claim 1, wherein the pressure detection tablet comprises at least two capsules surrounded by the outer coating, and wherein the at least two capsules comprise different dyes and are configured to rupture at equal to or higher than a different pressure threshold from one another. 10. The method of claim 1, further comprising administering a treatment to the subject if the peak pressure within the colon is equal to or above a pressure threshold, wherein the treatment is selected to reduce the peak pressure within the colon to below the pressure thres... | EMPIRE TECHNOLOGY DEVELOPMENT LLC | A61K009/48 A61B005/03 G01L001/00 G01N033/493 G01M003/02 A61K047/32 A61K047/38 A61B005/00 A61K009/00 A61K009/20
US09775518 | B2 | 2017-10-03 | Hand-held medical-data capture-device having a digital infrared sensor with no analog readout ports and optical detection of vital signs through variation amplification and interoperation with electronic medical record systems without specific discovery protocols or domain name service | 2014-12-29 | In one implementation, an apparatus estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point is described, estimates temperature from a digital infrared sensor and determines vital signs from a solid-state image transducer, or determines vital signs from a solid-state image transducer and estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point; after which the estimated and/or determined information is transmitted to an external database. | 1. A device to measure a temperature of a human, the device comprising: a microprocessor;a battery operably coupled to the microprocessor;a camera operably coupled to the microprocessor and providing at least two images to the microprocessor;a digital infrared sensor operably coupled to the microprocessor, the digital infrared sensor having only digital readout ports, the digital infrared sensor having no analog sensor readout ports; anda wireless communication subsystem that is operably coupled to the microprocessor and that is operable to transmit a representation of the temperature of the human,wherein the microprocessor is operable to receive from the digital readout ports a digital signal that is representative of an infrared signal detected by the digital infrared sensor and the microprocessor is operable to determine the temperature of the human from the digital signal that is representative of the infrared signal and the microprocessor including a cropper that is operable to receive the at least two images and operable to crop the at least two images to exclude a border area of the images, generating at least two cropped images, the microprocessor also including a temporal-variation-amplifier of the at least two cropped images that is operable to generate a temporal variation of the at least two cropped images, the microprocessor also including a vital-sign generator that is operably coupled to the temporal-variation-amplifier that is operable to generate at least one vital sign from the temporal variation,wherein the device does not support specific discovery protocols or domain name service. 2. The device of claim 1, wherein the temporal-variation-amplifier further comprises: skin-pixel-identifier. 3. The device of claim 1, wherein the temporal-variation-amplifier further comprises: a regional facial clusterial module. 4. The device of claim 1, wherein the temporal-variation-amplifier further comprises: a first frequency filter. 5. The device of claim 1 further comprising: a storage device that is operably coupled to the vital-sign generator and that is operable to store the at least one vital sign in a volatile memory. 6. The device of claim 1 further comprising: a storage device that is operably coupled to the vital-sign generator and that is operable to transmit the at least one vital sign to another apparatus. 7. The device of claim 1 wherein the wireless communication subsystem transmits to a static Internet Protocol address in order to reduce Internet Protocol discovery burden on the device. 8. The device of claim 1 wherein the wireless communication subsystem transmits through an Internet Protocol tunnel. 9. The device of claim 1 wherein a connection is established and data is pushed from the device through the wireless communication subsystem, thereafter an external device controls flow of the data between the device and the external device, wherein the connection further comprises: an authenticated communication channel... | ARC DEVICES, LTD | G08B023/00 A61B005/00 A61B005/01 G06F019/00 A61B005/021 A61B005/026 A61B005/08 A61B005/024 A61B005/0205 G06T003/40 G06T007/00 H04W048/16 H04W076/02 A61B005/0402 A61B005/03 G01J005/00 H04W080/04 G01J005/02 G01J005/10 G08C017/02 H04L029/08 H04W012/06 H04B007/26 H04L012/46 H04N005/33 H04W048/08 G08B005/36 G08B021/18 G01K013/00 G06K009/62 G06T007/174 G06T007/269 G06T007/90 A61B005/0245 H04W012/08
US09788600 | B2 | 2017-10-17 | Customized footwear, and systems and methods for designing and manufacturing same | 2013-12-19 | The invention relates to devices and methods for designing and manufacturing customized footwear, and components thereof. An example method includes a method of designing at least a portion of a sole of an article of footwear customized for a user. The method includes the steps of determining at least one input parameter related to a user, analyzing the at least one input parameter to determine at least one performance metric of a foot of the user, and determining at least one customized structural characteristic of at least a portion of a sole of an article of footwear for the user based on the performance metric. | 1. A method of manufacturing at least a portion of a sole of an article of footwear customized for at least one of a user and an athletic motion, the method comprising the steps of: determining input parameters related to the user, the input parameters comprising pressure distribution measurements over at least a portion of a sole of a foot of the user and force distribution measurements between at least a portion of the sole of the foot of the user and a ground surface during at least a portion of a ground contact phase of the athletic motion;analyzing the input parameters by multiplying the pressure distribution measurements and the force distribution measurements to determine at least one performance metric relating at least in part to at least one angle and at least one magnitude of an interaction between the foot of the user and the around surface at a plurality of locations on the foot of the user;determining at least one customized structural characteristic of at least a portion of the sole of the article of footwear for the user based on the at least one performance metric; andmanufacturing for the user at least the portion of the article of footwear with the customized structural characteristic. 2. The method of claim 1, wherein the at least one input parameter comprises at least one of a pressure, a friction, a force, a position, a velocity, an acceleration, a rotational velocity, a rotational acceleration, a change in shape of at least a portion of the foot, a mass distribution, an energy distribution, a stress, a strain, a time over which an event occurs, a ground contact time, an initial footstrike location, a toe-off position, a pronation characteristic, and a supination characteristic. 3. The method of claim 2, wherein the at least one input parameter comprises at least one of: (i) a pressure distribution over at least a portion of the sole of the foot of the user during at least a portion of the ground contact phase of the athletic motion, (ii) a change in at least one of shape and position of at least a portion of the foot of the user during at least a portion of the ground contact phase of the athletic motion, and (iii) a force distribution between at least a portion of the sole of the foot of the user and a ground surface during at least a portion of the ground contact phase of the athletic motion. 4. The method of claim 1, wherein the at least one input parameter further comprises at least one physiological structural characteristic of the foot, the physiological structural characteristic comprising at least one of a shape, a size, a wear pattern, an injury profile, a bone structure, a sweat gland distribution, a temperature, a moisture level, a circulation metric, a muscle activity, a friction metric, a blood pressure, a heart rate, a volume change, a hydration level, a perspiration level, a ligament structure, a toe shape, a toe distribution, a heel shape, and an arch shape. 5. The method of claim 1, wherein the at least... | New Balance Athletics, Inc. | G06F017/50 A43B013/14 A43B013/18 A43B005/02 A43B005/06 A43B013/22 A43C013/04 A43D001/02 A61B005/103 A43B013/12 B29C044/08 B33Y080/00 A43B007/14 A43B013/04 A43B013/20 A61B005/107 A61B005/00
US09820697 | B2 | 2017-11-21 | Lesion determination apparatus, similar case searching apparatus, lesion determination method, similar case searching method, and non-transitory computer-readable storage medium | 2015-03-10 | A lesion determination apparatus includes: a number-of-images acquiring unit configured to acquire and record, on occurrence of a direction change that is either a first state in which a second operation instruction is received immediately after receipt of a first operation instruction or a second state in which the first operation instruction is received immediately after a receipt of the second operation instruction, a number of one-directionally viewed images that is either a number of tomography images continuously viewed in response to continuous first operation instructions including the first operation instruction in the first state or a number of tomography images continuously viewed in response to continuous second operation instructions including the second operation instruction in the second state; and a locality determining unit configured to determine whether the object part is a focal lesion or a diffuse lesion based on a change in the number of one-directionally viewed images. | 1. A lesion determination apparatus comprising: an image set acquiring unit configured to acquire a tomography image set containing a plurality of tomography images of an object part;an operation instruction receiving unit configured to receive a first operation instruction to view tomography images included in the plurality of tomography images in a first order in which a tomography image position number for identifying each of the tomography images increases, and a second operation instruction to view tomography images included in the plurality of tomography images in a second order in which the tomography image position number decreases;a number-of-images acquiring unit configured to acquire and record, on occurrence of a direction change that is either a first state in which the second operation instruction is received immediately after receipt of the first operation instruction or a second state in which the first operation instruction is received immediately after receipt of the second operation instruction, a number of one-directionally viewed images that is either a number of tomography images continuously viewed in response to continuous first operation instructions including the first operation instruction contained in the first state or a number of tomography images continuously viewed in response to continuous second operation instructions including the second operation instruction contained in the second state; anda locality determining unit configured to determine whether the object part is a focal lesion or a diffuse lesion based on a change in the number of one-directionally viewed images,wherein the tomography image position number is assigned such that a tomography image of a tomographic slice plane closer to a particular portion of the object part has a smaller tomography image position number. 2. The lesion determination apparatus according to claim 1, wherein the locality determining unit determines that the object part is a focal lesion when the number of one-directionally viewed images has monotonically decreased with occurrence of the direction change, and otherwise determines that the object part is a diffuse lesion. 3. The lesion determination apparatus according to claim 2, wherein the locality determining unit excludes the number of one-directionally viewed images from consideration for determination when the number of one-directionally viewed images is equal to or smaller than a predetermined number. 4. The lesion determination apparatus according to claim 1, wherein the locality determining unit sets, as a predetermined threshold value, the number of one-directionally viewed images at a third direction change counted from a start of displaying the tomography image set, and determines that the object part is a focal lesion when the number of one-directionally viewed images at each of a fourth and subsequent direction changes is equal to or smaller than the predetermined threshold value, and otherwise determines tha... | PANASONIC CORPORATION | G06K009/00 A61B005/00 A61B005/055 A61B005/08 A61B006/00 A61B006/03 A61B008/13 A61B008/08 G06F019/00
US09827425 | B2 | 2017-11-28 | Methods and systems of electrode polarity switching in electrical stimulation therapy | 2014-09-03 | Methods for electrically stimulating body tissues to improve function or reduce symptoms provide an electrical stimulation system having two or more electrodes that are capable of being switched independently from a hyperpolarizing (depolarizing) state to a hypopolarizing state. Multiple combinations of hyperpolarizing electrodes and hypopolarizing electrodes are created by polarity switching to determine a polarity configuration having the best performance as determined by symptom reporting and clinical diagnostic tests. Polarity switching is triggered manually or is programmed to be switched automatically. Determining the configuration providing electrical stimulation resulting in the greatest benefit allows the system to be operated with one or more electrodes in a hypopolarizing state, thereby reducing energy requirements, tissue tolerance, and tissue fatigue. | 1. A system for electrically stimulating body tissue, comprising: a pulse generator;a controller; andat least two electrodes operatively connected to said pulse generator and said controller, wherein the controller is programmed to deliver a hyperpolarizing pulse stream through a first one of said two electrodes, said hyperpolarizing pulse stream being sufficient to increase a polarization of said body tissue to a level greater than a resting potential of said body tissue, and is programmed to deliver a hypopolarizing pulse stream through a second one of said two electrodes, said hypopolarizing pulse stream being sufficient to decrease a polarization of said body tissue to a level less than a resting potential of said body tissue, and wherein the controller is configured to switch polarities of said two electrodes such that said first one of said two electrodes delivers the hypopolarizing pulse stream and said second one of said two electrodes delivers the hyperpolarizing pulse stream, wherein said switching of the polarity of said first and second electrodes is adapted to automatically occur when a pH level in a lower esophagus of the patient measures above or below a predetermined threshold value. 2. The system for electrically stimulating body tissue of claim 1 comprising three electrodes. 3. The system for electrically stimulating body tissue of claim 1 comprising four electrodes. 4. The system for electrically stimulating body tissue of claim 1 comprising more than four electrodes. 5. The system for electrically stimulating body tissue of claim 1, wherein said pulse generator and said controller are contained within a metal housing configured as a return electrode, wherein said metal housing is adapted to be implanted a distance further away from a target tissue relative to a distance of said first and second electrodes from said target tissue. 6. The system for electrically stimulating body tissue of claim 5, wherein said metal housing is implanted in a subcutaneous tissue in an abdomen of a patient, in a gastrointestinal tract of a patient, or proximate a gastrointestinal tract of a patient within an abdominal cavity. 7. The system for electrically stimulating body tissue of claim 5, wherein said metal housing comprises hermetic titanium. 8. The system for electrically stimulating body tissue of claim 5, wherein a portion of said metal housing is coated with an electrically insulating material. 9. The system for electrically stimulating body tissue of claim 5, wherein said metal housing also comprises biocompatible plastic with an exposed electrode at the surface. 10. A system for electrically stimulating body tissues comprising: an electrode; anda pulse generator, wherein said pulse generator is operatively connected to one electrode, wherein said one electrode is adapted to be implanted proximate a target tissue, wherein said implantable pulse generator comprises a metal housing configured as a return electrode, wherein said imp... | EndoStim, Inc. | A61N001/36 A61B005/00
US09808368 | B2 | 2017-11-07 | Methods for performing bariatric surgery using gastric sizing systems and instruments | 2013-11-06 | Systems and methods for effecting bariatric procedures are disclosed. Each system includes an instrument, a control valve and, optionally, a suction controller. The instrument is in the form of an elongated, flexible sizing tube having a distal end portion including a plurality of apertures. The control valve enables suction to be applied to the patient's stomach via the apertures to drain gastric fluids and to bring adjacent portions of the patient's stomach into engagement the sizing tube to provide a visually perceptible delineation line along which a portion of the stomach may be resected, sealed and tested. | 1. A method of sizing the stomach of a patient for a bariatric procedure, the stomach of the patient having a lesser curvature and a greater curvature, said method comprising: a) providing a non-expandable instrument for introduction into the stomach of the patient, said instrument comprising a sizing tube in the form of an elongated tubular member having a longitudinal axis extending along a length thereof, a hollow interior, a distal end portion and a proximal end portion, said distal end portion having a plurality of apertures disposed around the periphery of said sizing tube, said plurality of apertures extending along said longitudinal axis and with said apertures directed in a plurality of different radial directions from said longitudinal axis, said plurality of apertures being in fluid communication with said hollow interior of said elongated tubular member;b) introducing said sizing tube into the stomach of the patient in any orientation about said longitudinal axis so that said elongated tubular member extends along the majority of the lesser curvature of the patient's stomach and directly engages the lesser curvature of the patient's stomach, whereupon plural ones of said plurality of apertures face towards the lesser curvature of the patient's stomach irrespective of the particular orientation of said sizing tube about said longitudinal axis in the patient's stomach and with said plural ones of said plurality of apertures that face towards the lesser curvature of the patient's stomach extending along the lesser curvature of the patient stomach;c) connecting said sizing tube to a source of suction;d) applying controlled suction into the interior of the patient's stomach via (i) said plural ones of said plurality of apertures facing the lesser curvature of the patient's stomach and (ii) others of said plurality of apertures facing away from the lesser curvature of the patient's stomach to pull the lesser curvature of the patient's stomach into engagement with the portion of said sizing tube facing the lesser curvature of the patient's stomach and without any portion of said instrument engaging the greater curvature of the patient's stomach, and contemporaneously therewith said controlled suction applied via said others of said plurality of apertures pulls portions of the patient's stomach adjacent thereto towards portions of said sizing tube not facing the lesser curvature of the patient's stomach to anchor said sizing tube in place thereat in said any orientation about said longitudinal axis and without the use of any expandable member to hold said sizing tube in place and to produce a suction-created visually perceptible delineation line on the exterior of the patient's stomach along a portion of the periphery of said sizing tube facing the greater curvature of the patient's stomach;e) visualizing said suction-created visually perceptible delineation line with a laparoscope from outside of the patient's stomach; andf) maintaining said... | Boehringer Laboratories, Inc. | A61F005/00 A61B005/107 A61B017/072 A61M025/00 A61B017/30 A61B017/32 A61B017/00 A61J015/00
US09786060 | B2 | 2017-10-10 | Device for controlling the anesthetization of an animal to be slaughtered | 2015-08-11 | A device for controlling the anesthetization of an animal to be slaughtered. The device includes a stimulation unit via which a stimulation stimulus is given to the animal. The device includes an image capturing unit that has an image capturing range. The image capturing unit is constructed as a 3D-camera. One part of a surface of the animal is optically detected in the image capturing range and range pixels are collected in the area of the surface. The range pixels are provided and transmitted as pixel data. An evaluation unit is connected to the image capturing unit and collects the pixel data provided by the image capturing unit. The evaluation unit determines an active movement of the animal from the detected pixel data and provides and transmits the control result in the event of a determined movement of the animal. | 1. An apparatus for checking the stunning of an animal for slaughter, comprising: a stimulation unit for delivering a stimulation stimulus to the animal for slaughter;an image capture unit having an image capture region, said image capture unit being a 3D camera, said image capture unit optically capturing a section of a surface of the animal for slaughter in the image capture region and capturing image points in the section of the surface, the image points being provided as transmittable image point data; andan evaluation unit connected to the image capture unit and capturing the image point data provided by the image capture unit, said evaluation unit ascertaining an active movement of the animal for slaughter from the captured image point data and in the event of an ascertained active movement of the animal for slaughter, said evaluation unit providing a check result in a form capable of being output. 2. The apparatus as claimed in claim 1, wherein said 3D camera is a time-of-flight camera. 3. The apparatus as claimed in claim 1, wherein said stimulation unit is configured to deliver a contactless stimulation stimulus. 4. The apparatus as claimed in claim 3, wherein said stimulation unit is configured to deliver a thermal stimulation stimulus. 5. The apparatus as claimed in claim 4, wherein the thermal stimulation stimulus is a jet of steam. 6. The apparatus as claimed in claim 1, wherein the provided check result controls external units. 7. The apparatus as claimed in claim 1, wherein the provided check result is associated with a data record relating to the animal for slaughter. | CSB-System AG | A22B003/00 G06T007/20 G01S017/89 A61D007/04 A61B005/00 A61B005/11
US09808191 | B2 | 2017-11-07 | Application of electrochemical impedance spectroscopy in sensor systems, devices, and related methods | 2017-03-08 | A diagnostic Electrochemical Impedance Spectroscopy (EIS) procedure is applied to measure values of impedance-related parameters for one or more sensing electrodes. The parameters may include real impedance, imaginary impedance, impedance magnitude, and/or phase angle. The measured values of the impedance-related parameters are then used in performing sensor diagnostics, calculating a highly-reliable fused sensor glucose value based on signals from a plurality of redundant sensing electrodes, calibrating sensors, detecting interferents within close proximity of one or more sensing electrodes, and testing surface area characteristics of electroplated electrodes. Advantageously, impedance-related parameters can be defined that are substantially glucose-independent over specific ranges of frequencies. An Application Specific Integrated Circuit (ASIC) enables implementation of the EIS-based diagnostics, fusion algorithms, and other processes based on measurement of EIS-based parameters. | 1. A method of calibrating a sensor during a period of sensor transition, the method comprising: defining an electrochemical impedance spectroscopy (EIS)-based sensor status vector (V) for each one of a plurality of sensor current (Isig)-blood glucose (BG) pairs;defining an Isig buffer to store Isig values of said plurality of Isig-BG pairs;defining an BG buffer to store BG values of said plurality of Isig-BG pairs;monitoring the status vectors for the plurality of Isig-BG pairs over time;detecting when there is a difference between a first status vector for a first Isig-BG pair and a subsequent status vector for a subsequent Isig-BG pair;if a magnitude of said difference is smaller than a predetermined threshold, calibrating said sensor based on values of said Isig-BG pairs from a first set of Isig and BG buffers; andif the magnitude of said difference is larger than a predetermined threshold, calibrating said sensor based on values of said Isig-BG pairs from a second set of Isig and BG buffers, wherein said second set of Isig and BG buffers contains at least the values of Isig and BG from said subsequent Isig-BG pair. 2. The method of claim 1, wherein said first set of Isig and BG buffers contain values of Isig and BG, respectively, prior to said period of sensor transition. 3. The method of claim 1, wherein said second set of Isig and BG buffers contain values of Isig and BG, respectively, after start of said period of sensor transition. 4. The method of claim 1, wherein, for each Isig-BG pair, the sensor status vector includes Nyquist slope and at least one of real impedance and imaginary impedance as elements of the vector. 5. The method of claim 1, wherein, for each Isig-BG pair, the sensor status vector includes 1 kHz real impedance as an element of the vector. 6. The method of claim 1, wherein, for each Isig-BG pair, the sensor status vector includes 1 kHz imaginary impedance as an element of the vector. 7. The method of claim 1, wherein, for each Isig-BG pair, the sensor status vector includes Nyquist slope as an element of the vector. 8. The method of claim 1, wherein, for each Isig-BG pair, the elements of the sensor status vector include 1 kHz real impedance, 1 kHz imaginary impedance, Nyquist slope, and Nyquist R2. 9. The method of claim 1, wherein the sensor status vector includes a plurality of EIS-based elements, and wherein the plurality of EIS-based elements are substantially glucose-independent. | MEDTRONIC MINIMED, INC. | A61B005/1495 G01N027/416 G01N027/02 G01N033/66 G01N033/96 A61B005/145 A61B005/053 A61B005/00
US09795303 | B2 | 2017-10-24 | Medical hyperspectral imaging for evaluation of tissue and tumor | 2015-12-07 | Apparatus and methods for hyperspectral imaging analysis that assists in real and near-real time assessment of biological tissue condition, viability, and type, and monitoring the above overtime. Embodiments of the invention are particularly useful in surgery, clinical procedures, tissue assessment, diagnostic procedures, health monitoring, and medical evaluations, especially in the detection and treatment of cancer. | 1. A multispectral or hyperspectral medical imaging system comprising: a plurality of lights, each respective light in the plurality of lights configured to emit radiation at a particular spectral band in a plurality of spectral bands used by a diagnostic protocol module;a first stage optic configured to receive radiation projected, from one or more respective lights in the plurality of lights, off of an in vivo tissue located at a region of interest on a subject;one or more bandpass filters in optical communication with the first stage optic;an imaging sensor for recording an image of the radiation projected off of the in vivo tissue;a diagnostic protocol module adapted to detect a particular characteristic of the in-vivo tissue; adiagnostic processor configured to: switch one or more respective light in the plurality of lights on and off based on the diagnostic protocol module,instruct the imaging sensor to record a plurality of images of the region of interest based on the diagnostic protocol module, each respective image in the plurality of images corresponding to radiation emitted at a particular spectral band by one or more respective light in the plurality of lights,co-register corresponding pixels in each image in the plurality of images;derive a coefficient for oxyhemoglobin and a coefficient for deoxyhemoglobin from each set of co-registered pixels in the plurality of images;generate, based on the plurality of images, a two-dimensional index characterizing the particular characteristic of the in vivo tissue, wherein the two-dimensional index includes a first scale to represent a concentration of oxyhemoglobin in the in vivo tissue and a second scale to indicate a concentration of deoxyhemoglobin in the in vivo tissue;display the coefficient for oxyhemoglobin and the coefficient for deoxyhemoglobin from at least one set of co-registered pixels in the plurality of images independently with the two-dimensional index on a video display. 2. The multispectral or hyperspectral medical imaging system of claim 1, wherein the imaging system is portable. 3. The multispectral or hyperspectral medical imaging system of claim 2, wherein the imaging system weighs less than 25 pounds. 4. The multispectral or hyperspectral medical imaging system of claim 1, wherein the plurality of lights is configured in a circular array about the first stage optic. 5. The multispectral or hyperspectral medical imaging system of claim 1, wherein at least one respective light in the plurality of lights is configured to emit radiation in the visible spectrum. 6. The multispectral or hyperspectral medical imaging system of claim 5, wherein at least one respective light in the plurality of lights is configured to emit radiation in the near-infrared spectrum. 7. The multispectral or hyperspectral medical imaging system of claim 1, wherein at least one respective light in the plurality of lights is configured to emit radiation in the near-infrared spectrum. 8. The mul... | Hypermed Imaging, Inc. | A61B005/00 G01J003/02 G01J003/10 G01J003/28 G01N021/21 G01N021/31 A61B005/02 G01N021/35 G01N021/359
US09795808 | B2 | 2017-10-24 | Devices and techniques for cutting and coagulating tissue | 2015-03-13 | Various embodiments are directed to an apparatus and method of driving an end effector coupled to an ultrasonic drive system of a surgical instrument. The method comprises generating at least one electrical signal. The at least one electrical signal is monitored against a first set of logic conditions. | 1. An ultrasonic surgical system for treating tissue, the system comprising: an ultrasonic transducer;an end effector coupled to the ultrasonic transducer;a switch having a first position and a second position; anda generator in communication with the switch and the ultrasonic transducer, wherein:in the first switch position, the generator is configured to provide a continuous drive signal to the ultrasonic transducer; andin the second switch position, the generator is configured to provide a pulsed drive signal to the ultrasonic transducer. 2. The system of claim 1, further comprising: a generator console, wherein the generator is positioned within the generator console;a device housing, wherein the ultrasonic transducer and the switch are positioned at least partially within the device housing; anda cable providing an electrical connection between the generator and the ultrasonic transducer. 3. The system of claim 1, wherein the generator and the ultrasonic transducer are positioned within a common housing. 4. The system of claim 1, wherein the continuous drive signal is at a power level less than a maximum power level for the system. 5. The system of claim 1, wherein the pulsed drive signal is at a maximum power level for the system. 6. The system of claim 1, wherein the continuous drive signal has a first power and the second pulsed drive signal has a second power higher than the first power. 7. The system of claim 1, wherein with respect to the pulsed drive signal, the generator is configured to provide: a first pulse amplitude for a first pulse time; anda second pulse amplitude for a second pulse time. 8. The system of claim 1, wherein the generator is further configured to: monitor the continuous drive signal against a first set of logic conditions; andtrigger a pulsing response when the first set of logic conditions is met, wherein with respect to the pulsing response, the generator is configured to provide: a first pulse amplitude for a first pulse time; anda second pulse amplitude for a second pulse time. 9. The system of claim 1, wherein the end effector and the ultrasonic transducer are configured to resonate at a resonant frequency, and wherein with respect to the continuous drive signal, the generator is configured to: set the continuous drive signal to a first frequency for a first time period, wherein the first frequency is different from the resonant frequency; andafter the first time period, set the continuous drive signal to the resonant frequency. 10. An ultrasonic surgical system for treating tissue, the system comprising: an ultrasonic transducer;an end effector coupled to the ultrasonic transducer;a switch having a first position and a second position; anda generator in communication with the switch and the ultrasonic transducer, wherein in the first switch position, the generator is configured to provide a continuous drive signal to the ultrasonic transducer and in the second switch position, the generator is conf... | Ethicon LLC | H01L041/09 A61N007/00 A61B017/32 B06B001/02 G05B015/02 A61B005/053 A61B005/00 A61B017/00 A61B018/00
US09820661 | B2 | 2017-11-21 | Biological information detection apparatus | 2014-03-11 | A biological information detection apparatus includes a detection unit which has a light receiving unit receiving light from a subject, a light transmitting member which is provided on a housing surface side in contact with the subject of the biological information detection apparatus, transmits light from the subject, and has a convex portion in contact with the subject to give a pressing force when measuring biological information of the subject, and a pressing force suppression unit which is disposed in periphery of the convex portion on or above the housing surface and suppresses the pressing force given to the subject by the convex portion. A groove portion is provided between the convex portion of the light transmitting member and the pressing force suppression unit. | 1. A biological information detection apparatus comprising: a housing surface;a detection unit which has a light receiving unit configured to receive light from a subject;a light transmitting member which is provided on the housing surface configured to be in contact with the subject of the biological information detection apparatus, configured to transmit light from the subject, and has a convex portion configured to be in contact with the subject to give a pressing force; anda pressing force suppression unit which is provided around an entire circumference of the convex portion to surround the convex portion above the housing surface and configured to suppress the pressing force given to the subject by the convex portion,wherein a groove portion is provided between the convex portion of the light transmitting member and the pressing force suppression unit, andwherein a height of the convex portion, a height of the pressing force suppression unit, and a height of a bottom surface of the groove portion are respectively HA, HB, and HC, HA>HB>HC, wherein the height of the convex portion extends along a direction orthogonal to the housing surface. 2. The biological information detection apparatus according to claim 1, wherein the difference between the height HA of the convex portion and the height HC of the bottom surface of the groove portion is Δh2, Δh2>0.5 mm. 3. The biological information detection apparatus according to claim 1, wherein the groove portion is provided over the entire circumference of the convex portion. 4. The biological information detection apparatus according to claim 1, wherein the bottom surface of the groove portion is a surface of the light transmitting member. 5. The biological information detection apparatus according to claim 1, wherein the light transmitting member has the convex portion at least a part of which is configured to protrude toward the subject, anda body portion which is provided on the downward side of the convex portion, configured to be opposite to the subject, andthe bottom surface of the groove portion is a surface of the body portion. 6. The biological information detection apparatus according to claim 5, wherein the body portion is formed to extend from the convex portion downward of a cover member of the housing surface, and a pressing force suppression surface of the pressing force suppression unit is a surface of the cover member. 7. The biological information detection apparatus according to claim 1, further comprising: a load mechanism having a load configured to generate the pressing force given to the subject by the convex portion, wherein an amount of change in the pressing force with respect to the load of the load mechanism is defined as the amount of change in pressing force, andwherein the amount of change in pressing force in a second load range in which the load of the load mechanism is greater than FL1 becomes smaller than the amount of change in pressing force in a first load... | Seiko Epson Corporation | A61B005/02 A61B005/024 A61B005/00 A61B005/11
US09801585 | B2 | 2017-10-31 | Electrocardiogram noise reduction | 2014-12-31 | A catheterization system that includes an electrophysiologic (EP) catheter which has a lumen receiving an electrically conductive fluid delivered by a hydraulic line that is acted upon by a peristaltic pump advantageously avoids noise in intracardiac ECG signal recordings by using an electrical connection to short triboelectrical charge carried by the conductive fluid in the hydraulic line to an existing analog ground in the system. In one embodiment, the electrical connection includes an electrically conductive wire housed in the control handle and configured to provide electrical connection between the fluid and a pin on a printed circuit board housed in the control handle that is electrically connected to the analog ground. In another embodiment, the electrical connection shorts the electrically conductive fluid proximal of the control handle of the catheter. | 1. A catheterization system, comprising: a catheter having a flexible catheter shaft with a distal portion adapted for insertion into a heart of a subject, the catheter shaft having a lumen for passing an electrically conductive fluid therethrough to exit the catheter at the distal portion;a control handle at a proximal end of the catheter, the control handle providing an isolated ground of the system;an electrical connection having first and second ends, the electrical connection linking the electrically conductive fluid and the isolated ground provided by the control handle, the first end of the electrical connection being in contact with the isolated ground provided by the control handle, and the second end of the electrical connection being in contact with the electrically conductive fluid at a location proximal to the distal portion of the catheter shaft. 2. The catheterization system of claim 1, wherein the second end of the electrical connection is in contact with the electrically conductive fluid in the lumen of the catheter shaft. 3. The catheterization system of claim 1, wherein the second end of the electrical connection is in contact with the electrically conductive fluid inside the control handle. 4. The catheterization system of claim 3, wherein the electrical connection includes an electrically conductive wire that is housed in the control handle and configured to provide electrical connection between the electrically conductive fluid and the isolated ground provided by the control handle. 5. The catheterization system of claim 1, wherein the second end of the electrical connection in contact with the electrically conductive fluid is proximal of the control handle. 6. The catheterization system of claim 5, wherein the electrical connection includes an electrically conductive wire having a divergence with one arm in electrical connection with the isolated ground provided by the control handle and another arm terminating inside a luer connector, the luer connector being configured for connection with the lumen of the catheter shaft for passing the electrically conductive fluid, and the luer connector being proximal of the control handle, the another arm adapted for electrical connection with the fluid passing through the lumen of the catheter shaft at a location proximal of the control handle. 7. The catheterization system of claim 1, wherein the isolated ground includes a pin located on a printed circuit board housed in the control handle. 8. The catheterization system of claim 7, wherein the electrical connection links the electrically conductive fluid and the pin. 9. The catheterization system of claim 8, wherein the electrical connection includes an electrically conductive wire extending between the electrically conductive fluid and the pin. 10. The catheterization system of claim 7, wherein the electrical connection linking the electrically conductive fluid and the pin is located inside the control handle. 11. The cath... | BIOSENSE WEBSTER (ISRAEL) LTD. | A61B005/00 A61B005/0402 A61B005/042 A61B018/14
US09782076 | B2 | 2017-10-10 | Smart messages and alerts for an infusion delivery and management system | 2011-07-18 | Method and system for providing diabetes management is provided. | 1. A method of managing hypoglycemia in real time performed by an electronic therapy device, comprising: detecting a pre-sleep time period comprising a time period prior to when a user enters a sleep cycle and entering a pre-sleep mode in response thereto;receiving, automatically without user input and during the pre-sleep time period, a predetermined number of consecutive analyte related level readings for the pre-sleep time period, wherein the predetermined number of consecutive analyte related level readings includes two or more consecutive analyte related level readings;retrieving, automatically without user input and during the pre-sleep time period, a predetermined target analyte level for the pre-sleep time period from a memory;comparing, automatically without user input and during the pre-sleep time period, each of the predetermined number of consecutive analyte related level readings for the pre-sleep time period to the predetermined target analyte level;in response to one or more of the predetermined number of consecutive analyte related level readings being greater than the predetermined target analyte level, exiting the pre-sleep mode and waiting for a subsequent pre-sleep time period; andin response to all of the predetermined number of consecutive analyte related level readings being equal to or lower than the predetermined target analyte level, generating and outputting an alert. 2. The method of claim 1, further including measuring an analyte related level of the user during the pre-sleep time period to determine the predetermined number of consecutive analyte related level readings. 3. The method of claim 2, wherein measuring the analyte related level includes using one or more of a discrete blood glucose meter or a continuous analyte monitoring system. 4. The method of claim 1, further including outputting the alert on a medication delivery unit. 5. The method of claim 1, wherein the alert includes one or more of an audible alert, a visual alert, or a tactile alert. 6. The method of claim 5, wherein the visual alert includes one or more of an alphanumeric output display, a graphical output display, an icon display, a video output display, a color display or an illumination display. 7. The method of claim 1, wherein the pre-sleep time period begins at approximately a predetermined number of hours after the last meal of the day, and terminates at the beginning of entering the sleep cycle. 8. The method of claim 1, wherein generating the alert includes generating a treatment recommendation. 9. A system for managing hypoglycemia in real time, comprising: a processor unit coupled to a non-transitory memory, the non-transitory memory comprising a plurality of instructions that, when executed, cause the processor unit to: detect a pre-sleep time period comprising a time period prior to when a user enters a sleep cycle and enter a pre-sleep mode in response thereto;automatically receive, without user input and during the pre-sle... | ABBOTT DIABETES CARE INC. | A61B005/00 A61B005/145 G06F019/00
US09808574 | B2 | 2017-11-07 | Insertion device | 2008-06-25 | An insertion apparatus and a method for use with a device for delivery of a therapeutic fluid into a body of a patient and/or for sensing of a bodily analyte are disclosed. The apparatus includes a housing adapted for loading therein at least one cannula cartridge unit having a protective member. The protective member accommodates at least one penetrating cartridge having a subcutaneously insertable element and a penetrating member. The apparatus includes a displacement mechanism capable of protracting the penetrating cartridge towards the body of the patient, where protraction of the penetrating cartridge results in insertion of the subcutaneously insertable element into the body of the patient. | 1. A subcutaneous insertion apparatus comprising: a housing configured to receive at least one cartridge unit, the at least one cartridge unit comprising a protective member containing therein at least one subcutaneously insertable element and at least one penetrating member, the housing including: a first housing opening adapted for at least loading therethrough the at least one cartridge unit into the housing; anda second housing opening adapted for at least protracting therethrough the at least one insertable element and the at least one penetrating member; anda displacement mechanism contained within the housing and configured to at least protract the at least one insertable element and the at least one penetrating member through the second housing opening toward the body of the patient for subcutaneous insertion of the at least one insertable element,wherein the displacement mechanism is configured to be primed for activation prior to the housing receiving the at least one cartridge unit; andthe insertion apparatus is configured to couple to a skin-securable cradle that comprises a tubular protrusion emerging upwardly from the cradle that defines therein a well, the well comprising a plurality of penetration tunnels for directing the subcutaneous insertable element toward a plurality of angles with respect to a plane of a bottom housing opening, and the insertion apparatus is configured to align the subcutaneous insertable element with any of the plurality of penetration tunnels of the well, so as to insert the subcutaneous insertable element through the aligned tunnel. 2. The insertion apparatus according to claim 1, wherein the displacement mechanism is further configured to retract the penetrating member into the protective member subsequent to the subcutaneous insertion of the insertable element. 3. The insertion apparatus according to claim 1, wherein the displacement mechanism comprises a rod for manually forcing the at least one insertable element and the at least one penetrating member toward the body of the patient. 4. The insertion apparatus according to claim 1, wherein the displacement mechanism includes a loadable spring, wherein unloading of the spring results in protraction of the at least one insertable element and the at least one penetrating member toward the body of the patient. 5. The insertion apparatus according to claim 4, wherein the displacement mechanism further comprises a second spring, wherein unloading of the second spring results in retraction of the penetrating member into the protective member subsequent to subcutaneous insertion of the insertable element. 6. The insertion apparatus according to claim 1, further comprising a user-actuated button for actuating the displacement mechanism when the cartridge unit is loaded therein. 7. The insertion apparatus according to claim 1, further comprising a safety means for preventing inadvertent actuation of the displacement mechanism. 8. The insertion apparatus... | Roche Diabetes Care, Inc. | A61M005/142 A61B005/00 A61M005/158 A61M005/42 A61B005/145 A61M005/14
US09795327 | B2 | 2017-10-24 | Measuring apparatus and measurement method | 2010-09-13 | Disclosed is a measuring device for obtaining numerical information concerning a substance present in the interstitial subcutaneous fluid, the device being equipped with a sensor unit that outputs signals in accordance with the numerical information concerning the substance and an arithmetic unit (control unit) that receives the signals outputted from the sensor unit and arithmetically processes the signals. The sensor unit is equipped with a sensor, some of which is punctured into the skin. The arithmetic unit (control unit) has been disposed so as to be separate from the sensor unit. | 1. A measuring apparatus that measures numeric information about a substance contained in at least one carrier selected from subcutaneous interstitial fluid and blood, comprising: a sensor section having a sensor outputting a signal in accordance with the numeric information, a support film that supports the sensor, and an adhesive film that affixes the sensor section to the skin; anda computing section that is disposed on a garment of a user utilizing said measuring apparatus and receives the signal outputted by the sensor and performs computational processing based on the signal,the sensor comprising a deployed portion deployed under the skin and a base portion disposed over the surface of the skin, andthe computing section being disposed so as to be isolated from the sensor section, whereinwhen the deployed portion extends not through the adhesive film but through the support film and the base portion is disposed on the support film, the sensor is supported by the support film and the deployed portion is deployed under the skin,the sensor section comprises, as the adhesive film, two or more pieces of adhesive film having a first adhesive layer on one side, andthe two or more pieces of adhesive film are strip-like in shape and are capable of adhesion to both said sensor section and the skin through the medium of the first adhesive layer of the adhesive film in respectively different locations on an upper face of said sensor section,wherein the computing section communicates with the sensor section via wireless communication,wherein the sensor section includes a water impermeable film having a second adhesive layer on one side, and the water impermeable film is formed so as to prevent the ingress of moisture to the base portion when it covers the base portion with the second adhesive layer facing the base portion, the second adhesive layer of the water impermeable film being configured to be in contact with the skin, andthe sensor section includes a signal transmitting unit connected to the sensor through a wire, and the signal transmitting unit is not covered by the water impermeable film. 2. The measuring apparatus according to claim 1, further comprising a transmitting section wirelessly transmitting the outcome of the computational processing performed by the computing section to an external location, and a housing containing both the computing section and the transmitting section. 3. The measuring apparatus according to claim 2, further comprising a receiver that receives the outcome of the computational processing transmitted by the transmitting section and displays numeric information about the substance based on the outcome of the computational processing. 4. The measuring apparatus according to claim 1, further comprising an amplifier circuit that amplifies the signal outputted by the sensor, wherein the amplifier circuit is provided in at least one location selected from the sensor section and a portion electrically connected to the... | Arkray, Inc. | A61B005/05 A61B005/145 A61B005/1486 A61B005/00
US09826927 | B2 | 2017-11-28 | Blood lancet device | 2012-09-07 | A blood lancet device comprising a housing having one end formed as an opening; a sliding sleeve having one end extending from the opening to the outside of the housing; a blood lancet including a needle body and a needle tip arranged at one end of the needle body, the sliding sleeve being bushed at the outside of the housing, said blood lancet adapted to move relative to the sliding sleeve so as to expose the needle tip at the opening; and a biasing member having one end securely fixed with another end of needle body. | 1. A blood lancet device, comprising: a housing provided with an opening at one end of the housing and guiding portions on an inner surface of the housing;a sliding sleeve having one end extending through said opening to an outside of the housing and the other end engaging with the guiding portions of the housing, the sliding sleeve being adapted to slide along an axial direction of said housing inside the housing;a blood lancet including a needle body and a needle tip disposed at one end of said needle body, said blood lancet adapted to move relative to said sliding sleeve so as to expose the needle tip at said opening; anda biasing member having one end engaged with the other end of the needle body,wherein said sliding sleeve is provided with at least one slot passing completely through a wall of the sliding sleeve, and the at least one slot extends toward said one end of the sliding sleeve from the other end of the sliding sleeve along a longitudinal axis of the sliding sleeve;wherein when the sliding sleeve is located at a first position within the housing, the other end of said sliding sleeve is pressed radially and inwardly by the guiding portions of the housing and elastically deformed radially and inwardly, such that the needle body is abutted against the other end of the sliding sleeve and is prevented from moving toward said opening of said housing;wherein when the sliding sleeve is moved to a second position within the housing, a pressing force applied on the other end of the sliding sleeve by the guiding portions is reduced or removed, such that said needle body is allowed to move toward said opening along said sliding sleeve under the action of the biasing member;wherein the other end of the sliding sleeve is configured to slide on the guiding portions and along a longitudinal axis of the housing during the first position to reach the second position;wherein said at least one slot includes two guiding slots extending toward said one end of the sliding sleeve from the other end of said sliding sleeve, said two guiding slots dividing said other end of the sliding sleeve into two guiding ends which can elastically deform toward each other;wherein said needle body also includes two lugs arranged in opposite, said two lugs are adapted to move along said two guiding slots under the action of the biasing member;wherein engagements between said guiding portions and said guiding ends are adapted to press two guiding ends so as to reduce the width of the guiding slots at said other end of the sliding sleeve to prevent said two lugs from moving toward said one end of said sliding sleeve along said two guiding slots under the action of the biasing member;wherein a tip end of each guiding end has a protrusion protruding radially and outwardly, and said protrusion is adapted to engage with said guiding portions; andwherein during movement of the sliding sleeve from the first position to the second position, the protrusion is continuously engaged a... | Tianjin Huahong Technology Co., Ltd. | A61B005/151 A61B005/15
US09775539 | B2 | 2017-10-03 | Medical device implantation imaging | 2013-02-08 | Implantation imaging techniques are presented herein to provide a surgeon with visual feedback during implantation of an implantable medical device in a recipient. The implantation imaging techniques may include the generation of a magnetic field that induces a voltage at a coil positioned in the recipient. The induced voltage is used to determine the orientation of the coil positioned in the recipient relative to direction of the magnetic field. | 1. A system comprising: a first external coil arranged about a first axis and configured to generate a magnetic field within a surgical region;an intra-cochlear stimulating assembly configured to be inserted into a cochlea of recipient located in the surgical region, wherein the stimulating assembly comprises a plurality of assembly coils spaced along the length of the stimulating assembly, and wherein the plurality of assembly coils are each configured to detect the magnetic field generated by the first external coil and to generate a voltage based on the detected magnetic field; anda hardware processor configured to determine, based on the voltages detected at the plurality of assembly coils, an angular orientation of each of the plurality of assembly coils with respect to the first axis and, based on the determined angular orientations, extrapolate a shape of the stimulating assembly within the cochlea. 2. The system of claim 1, wherein the first external coil is configured to generate a magnetic field having a uniform strength across the surgical region to induce a voltage at the assembly coil. 3. The system of claim 1, wherein the plurality of assembly coils have a non-uniform spacing within the stimulating assembly. 4. The system of claim 1, further comprising: first, second, and third pairs of external coils each arranged about first, second, and third axes, respectively,wherein the plurality of assembly coils are each configured to detect the magnetic fields generated by the first, second, and third pairs of external coils and are each configured to generate three voltages based on the detected magnetic fields, and wherein the hardware processor is configured to use the three generated voltages from each of the plurality of assembly coils to determine an angular orientation of the corresponding assembly coil with respect to each of the first, second, and third axes and, based on the determined angular orientations, extrapolate the shape of the stimulating assembly within the cochlea. 5. The system of claim 1, further comprising: a second external coil arranged about a second axis and configured to generate a second magnetic field within the surgical region. 6. The system of claim 5, further comprising: a third external coil arranged about a third axis and configured to generate a third magnetic field within the surgical region. 7. The system of claim 1, further wherein the elongate assembly comprises a first assembly coil located at a distal end of the stimulating assembly, and wherein the hardware processor is configured to capture information about the magnetic field when the first assembly coil is positioned at a pre-selected location of the cochlea. 8. The system of claim 7, wherein the pre-selected location is an opening in the cochlea. 9. The system of claim 7, wherein the hardware processor is further configured to determine a location of each of the plurality of assembly coils with respect to the pre-selected location of t... | Cochlear Limited | A61B005/05 A61B005/06 A61N001/36 A61N001/05 A61B005/00 A61M025/01 A61B017/00 A61B034/20
US09788812 | B2 | 2017-10-17 | Needle guide with selectable aspects | 2012-06-22 | A needle guide assembly for inserting a needle into the body of a patient in order to access a subcutaneous target, such as a vessel, is disclosed. In one embodiment, the needle guide assembly comprises a needle guide body that is configured to at least indirectly and removably attach to an image producing device, such as an ultrasound probe. The needle guide body defines at least first and second elongate guide channels. Each guide channel defines a unique insertion angle with respect to a longitudinal axis of the ultrasound probe. Further, each guide channel is configured to accept needles of differing gauges. In other embodiments other needle guide assemblies are disclosed that include multiple guide channels for inserting a needle at a variety of insertion angles into the patient's body. Related methods are also disclosed. In yet other embodiments, needle guide assemblies including needle stop features are disclosed. | 1. An assembly, comprising: an ultrasound probe; anda needle guide assembly, comprising: a plurality of guide channels, connected to a needle guide body, the needle guide body designed for attachment to the ultrasound probe, in each of the guide channels a needle can be temporarily disposed, the guide channels for guiding the needle into a patient; anda needle stop feature slidably engaged at a distal end of each of the plurality of guide channels, the needle stop feature remaining slidably engaged while being linearly movable between a first position for blocking the distal end of each of the guide channels and a second position where the needle stop feature enables needle passage through the guide channels. 2. The needle guide assembly as defined in claim 1, wherein the needle stop feature includes a plurality of shutters for covering the distal ends of each of the guide channels when the needle stop feature is in the first position. 3. The needle guide assembly as defined in claim 2, wherein the shutters in the first position are configured so as to engage only a portion of the needle proximal to the needle distal end when the needle is disposed in one of the guide channels so as to not blunt a distal tip of the needle. 4. The needle guide assembly as defined in claim 2, wherein the plurality of shutters covers only a portion of the distal ends of each of the guide channels when the needle stop feature is in the first position. 5. The needle guide assembly as defined in claim 2, wherein the plurality of shutters entirely covers the distal ends of each of the guide channels when the needle stop feature is in the first position. 6. The needle guide assembly as defined in claim 2, wherein the plurality of shutters extend outwardly along one edge of the needle stop feature. 7. The needle guide assembly as defined in claim 1, wherein the plurality of guide channels are arranged on one side of the needle guide assembly. 8. The needle guide assembly as defined in claim 7, wherein the plurality of guide channels includes three distinct guide channels arranged on the same side of the needle guide assembly. 9. The needle guide assembly as defined in claim 1, further comprising a plurality of front faces proximate proximal ends of the plurality of guide channels, wherein each of the plurality of front faces includes a contour shaped to direct a tip of a needle into a guide channel of the plurality of guide channels. 10. The needle guide assembly as defined in claim 9, wherein each of the plurality of front faces is positioned at a different angle, and wherein the contour of each of the plurality of front faces is concavely shaped. 11. The needle guide assembly as defined in claim 9, wherein each of the plurality of front faces is positioned facing outwardly away from the ultrasound probe when attached to the ultrasound probe. 12. The needle guide assembly as defined in claim 11, wherein the press surfaces extend from a main body portion of the... | C. R. Bard, Inc. | A61B019/00 A61B008/08 A61B017/34 A61B005/15 A61M005/46 A61B005/153 A61B005/155 A61B008/00 A61B010/02 A61M005/32 A61B006/12 A61M005/42
US09801573 | B2 | 2017-10-31 | Endoscope system, processor device, and method for operating endoscope system | 2014-02-04 | A color image sensor captures a reflected image of narrowband light having a wavelength range in which an extinction coefficient varies with a change in an oxygen saturation level of hemoglobin in blood. Thereby a first blue signal, a first green signal, and a first red signal are obtained. The color image sensor captures a reflected image of white light. Thereby a second blue signal, a second green signal, and a second red signal are obtained. Only an oxygen saturation level, out of two or more types of biological functional information including a blood volume and the oxygen saturation level, is obtained based on the first blue signal, the second green signal, and the second red signal. The oxygen saturation level is visualized to produce an oxygen saturation image. | 1. An endoscope system comprising: a light source that sequentially applies white light and blue narrowband light in which extinction coefficient varies significantly in accordance with a change in an oxygen saturation level of hemoglobin in blood, to a subject;a color image sensor that receives an imaging reflection light from the subject illuminated by the blue narrowband light in a first frame to obtain a first image signal, and receives an imaging reflection light from the subject illuminated by the white light in a second frame to obtain a second image signal, and a third image signal, the first image signal corresponding to first reflection light including a first wavelength range in a blue region, an extinction coefficient varying in accordance with an oxygen saturation level of hemoglobin in blood in the first wavelength range, the second image signal corresponding to second reflection light including a second wavelength range in a red region, the third image signal being used for standardizing the first and the second image signals, the third image signal corresponding to third reflection light including a third wavelength range in a green region, the third wavelength range differing from the first and the second wavelength ranges;a processing circuitry configured for:obtaining a first standardization signal ratio and a second standardization signal ratio, the first standardization signal ratio being obtained by standardizing the first image signal with the third image signal, the second standardization signal ratio being obtained by standardizing the second image signal with the third image signal;calculating the oxygen saturation level from a combination of the first and second standardization signal ratios with referring to a correlation stored in a memory, the correlation being a predefined relation between the oxygen saturation and the combination of the first and second standardization signal ratios; andproducing an oxygen saturation image in which the oxygen saturation level is visualized in pseudo color, with referring to a color table stored in a memory, anda display that displays the oxygen saturation image. 2. The endoscope system of claim 1, further comprising an oxygen saturation calculator that obtains the oxygen saturation level not dependent on the blood volume, based on the first image signal, second image signal and the third image signal, and an oxygen saturation image generator for producing the oxygen saturation image based on the oxygen saturation level obtained by the oxygen saturation calculator. 3. The endoscope system of claim 1, wherein the color image sensor images the subject, illuminated with blue narrowband light, with a color image sensor to obtain the first image signal, and the color image sensor images the subject, illuminated with white light, with the color image sensor to obtain the second image signal and the third image signal. 4. The endoscope system of claim 3, wherein the white light is pseud... | FUJIFILM Corporation | A61B005/1455 A61B005/145 A61B001/00 A61B001/06 G06T007/00 A61B005/00 A61B005/1459 G06T011/00
US09820689 | B2 | 2017-11-21 | Method and system for monitoring and displaying physiological conditions | 2014-10-14 | Providing an indication of a state of awareness for a patient includes acquiring an EEG signal; filtering the EEG signal to generate a first frequency band of signals; and filtering the EEG signal to generate a second frequency band of signals. Providing the indication also includes concurrently displaying the first and second frequency bands of signals in a single display window such that at least a first portion of the first frequency band overlaps, spatially the second frequency band. | 1. A method for providing an indication of a state of awareness for a patient, comprising: acquiring an EEG signal;filtering the EEG signal to generate a first frequency band of signals;filtering the EEG signal to generate a second frequency band of signals; andconcurrently displaying the first and second frequency bands of signals in a single display window such that at least a first portion of the first frequency band overlaps, spatially the second frequency band by overlaying one of the first and second frequency band of signals over the other of the first and second frequency band of signals on a same abscissa axis and a same ordinate axis. 2. The method of claim 1, wherein: the first frequency band of signals ranges in amplitude between a first negative value and a second positive value; andthe second frequency band of signals ranges in amplitude between a third negative value and a fourth positive value. 3. The method of claim 1, wherein: the first frequency band of signals ranges in amplitude between zero and a first positive value; andthe second frequency band of signals ranges in amplitude between zero and a second positive value. 4. The method of claim 1, wherein a respective amplitude display scale for each of the first and second frequency bands of signals is different from one another. 5. The method of claim 1, further comprising: filtering the EEG signal to generate a third frequency band of signals; andconcurrently displaying the first, second, and third frequency bands of signals such that at least a second portion of the first frequency band overlaps, spatially the second frequency band and the third frequency band by overlaying the first, second and third frequency band of signals over one another on the same abscissa and ordinate axes. 6. The method of claim 3, wherein the single display window comprises a top half and a bottom half and the first and second frequency bands are displayed in the top half. 7. The method of claim 6, further comprising: filtering the EEG signal to generate a third frequency band of signals;filtering the EEG signal to generate a fourth frequency band of signals; andconcurrently displaying the third and fourth frequency bands of signals in the bottom half of the single display window such that at least a first portion of the third frequency band overlaps, spatially the fourth frequency band by overlaying one of the third and fourth frequency band of signals over the other of the third and fourth frequency band of signals on the same abscissa a second set of abscissa and ordinate axes. 8. The method of claim 7, wherein: the third frequency band of signals ranges in amplitude between zero and a first negative value; andthe fourth frequency band of signals ranges in amplitude between zero and a second negative value. 9. The method of claim 1, further comprising: selecting a range of frequencies for the first frequency band based on one or more attributes of the EEG signal. 10. The method of clai... | - | A61B005/048 A61B005/00 A61B005/04 A61B005/16
US09775865 | B2 | 2017-10-03 | Large intestine contents' diagnosis and acoustic characterization | 2016-04-26 | Systems, kits and methods are provided, which analyze the large intestine content and utilize acoustic signals detected during delivery of water into the large intestine and drained large intestine contents to derive large intestine characteristics such as microbiotal analysis. Systems may include a water delivery unit including a water supply and a nozzle connected thereto, configured to introduce water controllably into a patient's large intestine, and an analysis unit that provides information about the drained contents using optical examination or biological assays. The information may be related to acoustic analysis of signals from acoustic sensors that are attachable to a patient's abdomen. A variety of sensor configurations, positioning options, analysis strategies and large intestine characteristics are presented. | 1. A method of large intestine diagnosis and treatment, comprising: introducing water controllably into a patient's large intestine,draining, by gravity, the introduced water with contents of the patient's large intestine,relating at least two portions of the drained contents with at least two sections of the large intestine, after the draining and according to at least one of: time of drainage of the portions, and characteristics of the drained content,deriving microbiome characteristics of the at least two sections of the large intestine by analyzing the corresponding at least two portions of the drained contents, andapplying a treatment to the large intestine according to the derived microbiome characteristics of the at least two sections. 2. The method of claim 1, further comprising determining a large intestine cleanliness level according to the drained contents of the patient's large intestine. 3. The method of claim 1, wherein the relating is carried out optically at a transparent drainage pipe through which the contents are drained. 4. The method of claim 1, wherein the deriving is carried out using at least most of the drained contents. 5. The method of claim 1, further comprising deriving large intestine characteristics by analyzing parameters of the drained water and contents. 6. The method of claim 1, wherein applying the treatment comprises transplanting a microbiota transplant (MT) into the patient's large intestine, the MT selected according to the derived microbiome characteristics of the at least two sections of the large intestine. 7. The method of claim 6, wherein the MT is a fecal microbiota transplant (FMT). 8. The method of claim 7, wherein the FMT is an autologous FMT derived from prior large intestine drainage. 9. The method of claim 1, further comprising: modifying a pressure of the introduced water,measuring a pressure of the drained water and contents, andderiving large intestine characteristics by correlating the measured drainage pressure with the pressure of introduced water according to a specified model. 10. The method of claim 1, further comprising: sensing acoustic signals at a patient's abdomen, andderiving large intestine characteristics by analyzing the sensed acoustic signals associated with the water introduction. 11. The method of claim 1, further comprising storing a plurality of the drained contents' samples from a plurality of the patients, associated with the derived microbiome characteristics of the drained contents samples. 12. The method of claim 1, further comprising: selecting a microbiota transplant (MT) according to the derived microbiome characteristics of the at least two sections of the large intestine, andtransplanting the MT into the patient's large intestine. 13. A method of large intestine diagnosis and treatment, comprising: introducing water controllably into a patient's large intestine,draining, by gravity, the introduced water with contents of the patient's large intestine,... | HYGIEACARE, Inc. | A61K035/74 A61K035/37 A61K009/00 C12Q001/04 A61B005/00 A61B007/00 A61M003/02 A61M001/00 A61K035/00
US09801607 | B2 | 2017-10-31 | Evaluating arterial pressure, vasomotor activity and their response to diagnostic tests | 2014-03-04 | Method and system for evaluating arterial pressure waves, vascular properties, as well as for diagnostic, physiological and pharmacological testing using various combinations of the following data acquisition and processing steps (some of the steps are optional): 1. Perturbing arterial pressure from its steady state. 2. Measuring the dynamics of at least one parameter related to the passage of arterial pressure waves along blood vessels. 3. Characterizing the magnitude and functional relation of changes in parameters described above in relation to changes in blood pressure during its displacement from and/or return to the steady state. 4. Classifying (comparing) the individual functional relation described above with a databank of parameters/functional relations for different states of vasomotor activity. | 1. A method for dynamical evaluation of at least one indicator selected from arterial pressure, cardiac output, and vascular properties, said method comprising: placing at least two sensors, at least one of which contains an accelerometer adapted for placement on the surface of an individual's torso and is adapted for registering at least one torso-surface movement related to cardiovascular mechanical activity,collecting data from said at least two sensors over multiple cardiac cycles; andprocessing said data, wherein said processing is adapted to determine at least two parameters of cardiovascular activity and is further adapted to derive at least one parameter of cardiovascular mechanical activity related to the arterial pressure in the central arteries. 2. A method as set forth in claim 1 that includes preprocessing said physiological signals to separate said signals from noise and interference using at least one method selected from filtering using fixed filters, filtering using adaptive filters, pattern recognition, orthogonal linear decomposition, and wavelet decomposition. 3. A method as set forth in claim 1 in which said processing further includes calibrating said parameters with respect to at least one reference value selected from: i. systolic pressure;ii. diastolic pressure;iii. mean arterial pressure;iv. hear rate;v. vascular properties;vi. arterial-pressure-wave (pulse) transit time (PTT);vii. arterial-pressure-wave velocity;viii. an individual's baseline values of said parameters;ix. characteristics of displacement of said parameters from baseline values;x. the magnitude, range, speed, time length, and pattern of temporal changes in said parameters;xi. the magnitude, range, speed, time length, and pattern of spatial changes in said parameters with respect to different anatomical locations;xii. augmentation index (AI);xiii. a functional relationship between at least two reference values;xiv. at least one measure of similarity with at least one typical value for an individual;xv. at least one measure of difference with an individual's baseline (typical) values;xvi. at least one measure of similarity with at least one typical pattern for an individual;xvii. at least one measure of similarity with at least one typical value for a population;xviii. at least one measure of difference with at least one typical pattern for a population;xix. at least one measure of similarity with at least one typical pattern for a population;xx. a functional relationship between said parameters and at least two reference points within at least one reference signal selected from: a. systolic pressure;b. diastolic pressure;c. mean arterial pressure:d. heart rate;e. arterial-pressure-wave (pulse) transit time;f. arterial-pressure-wave velocity; andg. vascular properties:xxi. a functional relationship between said parameters and at least two reference values within at least two said reference signals. 4. A method as set forth in claim 1 which includes track... | - | G01N033/48 A61B008/04 A61B005/021 A61B005/02 A61B005/0456
US09789878 | B2 | 2017-10-17 | Driver rest recommendation | 2015-04-20 | Methods and systems for monitoring a driver of a vehicle using a driver state monitoring (DSM) system. The DSM system includes a sensing unit that senses biological information of a user of a vehicle. The DSM system also includes a controller that determines a drowsy state of the user based on the sensed biological information. The controller also determines a current condition of traffic and outputs information recommending rest for the user based on the determined drowsy state of the user and the determined current condition of traffic. The controller monitors the traffic and determines a rest state of the user, and outputs an alarm based on monitoring the traffic and the determined rest state of the user. | 1. A driver state monitoring (DSM) system comprising: a sensing unit configured to sense biological information of a user of a vehicle; anda controller configured to: determine a drowsy state of the user based on the sensed biological information;determine a traffic condition to a destination;output information recommending rest for the user based on the determined drowsy state of the user and the determined traffic condition to the destination;monitor the traffic condition to the destination and a rest state of the user based on the user stopping the vehicle in accordance with the information recommending rest;generate a time at which to output an alarm based on monitoring the traffic condition to the destination and the rest state of the user; andoutput the alarm at the generated time. 2. The DSM system of claim 1, wherein the controller is further configured to determine a driving time to the destination, and wherein providing output information recommending rest for the user is based on the driving time and the traffic condition to the destination. 3. The DSM system of claim 1, wherein the controller is further configured to: detect that the user is in the rest state;determine an updated traffic condition; andchange the time at which the alarm is to be output based on the updated traffic condition. 4. The DSM system of claim 3, wherein the controller is further configured to, based on analyzing the updated traffic condition, change the time at which to output the alarm based on determining that the updated traffic condition has improved above a predetermined level. 5. The DSM system of claim 3, wherein the controller is further configured to: re-sense the biological information of the user based on determining that the traffic condition has changed as a result of analyzing the updated traffic condition; anddetermine the time at which to output the alarm further based on the re-sensed biological information. 6. The DSM system of claim 1, wherein the controller is further configured to: re-sense the biological information of the user based on the output of the alarm; andoutput updated information recommending rest for the user based on a result of re-sensing the biological information of the user. 7. The DSM system of claim 6, wherein the updated information recommending rest for the user further comprises information associated with an updated drowsy state of the user determined based on the re-sensed biological information of the user. 8. The DSM system of claim 1, wherein the biological information is associated with a heartbeat spectrum of the user, and the controller is further configured to classify the rest state of the user as any one of a plurality of sleep stages based on a result of analyzing the heartbeat spectrum. 9. The DSM system of claim 8, wherein the controller is further configured to determine the time at which to output the alarm based on the classified sleep stage. 10. The DSM system of claim 1, wherein the contro... | LG Electronics Inc. | B60R025/10 B60Q001/00 G08B023/00 G08B013/00 B60R021/26 A61B005/00 A61B005/04 B60W040/08 A61B005/18 B60K028/06 G08B021/06 G06F003/01 A61B005/024 G06F003/0354
US09775569 | B2 | 2017-10-03 | Electronic fetal monitoring applications and display | 2016-08-16 | Systems, methods, and computer-readable media are provided for displaying on a user display device alarm-triggering events detected by the medical device based on the location of the user display device. The method comprises receiving alarm-triggering patient data from a medical device and determining the physical location of first and second user display devices. The method further comprises determining that the physical location of the first user display device is designated to have an alarm displayed, and that the physical location of the second user display device is not designated to have an alarm displayed. In response to these determinations, the content of the alarm is displayed on the first user display device. | 1. Non-transitory computer storage media having computer-executable instructions embodied thereon that, when executed, cause a computing device to perform a method of displaying on a user display device alert-triggering events detected by a medical device based on a location of the user display device, the method comprising: receiving from a medical device patient data which triggers an alert having associated alert content;determining a physical location of first and second user display devices, wherein the physical location of the first and second user display devices is independent from users logged onto the first and second user display devices;determining that the physical location of the first user display device is designated to have the alert content displayed;in response to the determination that the determination that the physical location of the first user display device is designated to have the alert content displayed, displaying the alert and the alert content on the first user display device, wherein the alert content comprises one or more selectable visual representations of the patient data;receiving, via a user interface on the first user display device, a user selection of at least one selectable visual representation of the patient data of the one or more selectable visual representations of the patient data;determining that the physical location of the second user display device is not designated to have the alert content displayed; andin response to the determination that the physical location of the second user display device is not designated to have the alert content displayed, displaying the alert without the alert content on the second user display device. 2. The media of claim 1, wherein the first user display device is located at a nurse manager's station and the second user display device is located in an operating room. 3. The media of claim 1, wherein the first user display device is located in a patient's room and the second user display device is located in an operating room. 4. The media of claim 1, wherein the patient is a maternity patient and the medical device patient data is for the maternity patient and her unborn baby. 5. The media of claim 1, further comprising receiving, from one or more additional medical devices, patient data which triggers one or more additional alerts having alert content associated with the one or more additional alerts. 6. The media of claim 5, further comprising displaying a number representing a number of the one or more additional alerts on the first user display device and the second user display device. 7. A computerized method carried out by a computing device having one or more processors for displaying on a user display device alert-triggering events detected by a medical device based on a location of the user display device, the method comprising: receiving from a medical device patient data which triggers an alert;determining a physical location of a first user dev... | Cerner Innovation, Inc. | A61B005/00 A61B005/0444
US09827372 | B2 | 2017-11-28 | Integrated insulin delivery system with continuous glucose sensor | 2016-11-28 | Systems and methods for integrating a continuous glucose sensor 12, including a receiver 14, a medicament delivery device 16, a controller module, and optionally a single point glucose monitor 18 are provided. Integration may be manual, semi-automated and/or fully automated. | 1. A method for diabetes management, the method comprising: receiving measured glucose data from a continuous glucose monitor;receiving external data; andcontrolling insulin delivery to a host, using a processor, by iteratively evaluating the received glucose data and the received external data relative to a predetermined glucose boundary and selectively controlling the insulin delivery responsive to the iteratively evaluating, wherein the external data comprises one or more of time of day, meals, sleep, calories, exercise, or sickness, whereby outside influences are considered in therapy instructions. 2. The method of claim 1, wherein the controlling insulin delivery includes selectively turning a controller on or off and selectively turning the insulin delivery off when the glucose data meets a predetermined criterion relative to the predetermined glucose boundary. 3. The method of claim 2, wherein the predetermined criterion for selectively turning the controller off comprises the host's glucose falling below the glucose boundary. 4. The method of claim 2, wherein the selectively turning the controller on or off comprises selectively turning the insulin delivery on when the glucose data meets a predetermined criterion relative to the predetermined glucose boundary. 5. The method of claim 4, wherein the predetermined criterion for selectively turning the controller module on comprises the host's glucose surpassing the glucose boundary. 6. The method of claim 1, wherein the predetermined glucose boundary comprises a set point. 7. The method of claim 6, wherein the predetermined glucose boundary comprises more than one set point. 8. The method of claim 7, wherein the more than one set point comprises a first set point to select an on instruction and a second set point to select an off instruction. 9. The method of claim 1, wherein the glucose data is generated by one or more auxiliary sensors. 10. The method of claim 9, wherein the one or more auxiliary sensors comprises an auxiliary glucose sensor. 11. The method of claim 1, further comprising adaptively adjusting the glucose boundary. 12. The method of claim 11, further comprising requesting validation from a user prior to applying an adaptive adjustment to the glucose boundary. 13. The method of claim 1, wherein the controlling comprises controlling insulin delivery at a faster rate or a slower rate based on the glucose data. 14. The method of claim 1, wherein the iteratively evaluating includes iteratively evaluating internally derived data, where the internally derived data comprises insulin on-board information, and wherein the iteratively evaluating comprises iteratively evaluating the insulin on board information. 15. The method of claim 1, wherein the external data is user-entered. 16. The method of claim 1, wherein the method is fully automated or operates in a fully automated mode. 17. The method of claim 1, wherein the external data is exercise, and wherein the iterati... | DexCom, Inc. | A61M005/172 A61B005/145 A61M005/142 A61M011/00 G06F019/00 A61M005/20 A61M035/00 A61B005/00 A61M005/00 A61M005/178 A61M005/30 A61M015/00
US09779352 | B1 | 2017-10-03 | Systems and methods for individual identification and authorization utilizing conformable electronics | 2016-08-10 | An identification device includes, but is not limited to, a deformable substrate; a sensor assembly including one or more identity sensors configured to generate identity sense signals associated with at least one physical characteristic of an individual subject; circuitry including a comparison module configured to compare the identity sense signals to reference data indicative of one or more physical characteristics associated with an identity of at least one individual; and an adhesive coupled to a surface of the deformable substrate and configured to adhere the deformable substrate to a skin surface, the adhesive configured to transition from an adhesive state to a non-adhesive state responsive to one or more energy signals from the circuitry responsive to a correspondence between the identity sense signals and the physical characteristics associated with the identity being below a threshold correspondence. | 1. An identification device, comprising: a deformable substrate configured to conform to a skin surface of a body portion of an individual subject;a sensor assembly coupled to the deformable substrate, the sensor assembly including one or more identity sensors configured to generate one or more identity sense signals associated with at least one physical characteristic of the individual subject;circuitry operably coupled to the sensor assembly and configured to receive the one or more identity sense signals associated with the at least one physical characteristic of the individual subject, the circuitry including a comparison module configured to compare the one or more identity sense signals generated by the sensor assembly to reference data indicative of one or more physical characteristics associated with an identity of at least one individual; andan adhesive coupled to a surface of the deformable substrate and configured to adhere the deformable substrate to the skin surface, the adhesive configured to transition from an adhesive state to a non-adhesive state responsive to one or more energy signals from the circuitry, the circuitry configured to generate the one or more energy signals responsive to a correspondence between the one or more identity sense signals and the one or more physical characteristics associated with the identity of the at least one individual being below a threshold correspondence. 2. The identification device of claim 1, wherein the adhesive includes an electroactive polymer actuatable by the one or more energy signals from the circuitry to transition from the adhesive state to the non-adhesive state. 3. The identification device of claim 1, wherein the adhesive includes a thermo-responsive adhesive actuatable by a change in temperature of the adhesive responsive to the one or more energy signals from the circuitry to transition from the adhesive state to the non-adhesive state. 4. The identification device of claim 1, wherein the adhesive includes a magnetic-responsive adhesive responsive to the one or more energy signals from the circuitry to transition from the adhesive state to the non-adhesive state. 5. The identification device of claim 1, wherein the adhesive includes a light-responsive adhesive responsive to the one or more energy signals from the circuitry to transition from the adhesive state to the non-adhesive state. 6. The identification device of claim 1, wherein the adhesive includes a reversible microstructure configured to transition between the adhesive state and the non-adhesive state responsive to the one or more energy signals from the circuitry. 7. The identification device of claim 6, wherein the reversible microstructure is configured to reversibly transition between the adhesive state and the non-adhesive state responsive to the one or more energy signals from the circuitry. 8. The identification device of claim 1, further including: at least one reservoir configured to release an adhere... | ELWHA LLC | G06K019/077 A61B005/00 G06K019/04
US09782172 | B2 | 2017-10-10 | Electromechanical drive and remote surgical instrument attachment having computer assisted control capabilities | 2016-01-11 | A medical tool comprising an electromechanical driver and a surgical instrument attachment for use in invasive surgery, including a handle coupled to a flexible sheath which is in turn coupled to a surgical attachment. The handle of the driver includes the electromechanical driver and at least one processor element which controls the actions of the electromechanical driver, and therefore the application elements of the surgical attachment, based on information relayed between the processor element and remotely activatable sensor assemblies in the surgical instrument attachment. | 1. A surgical system, comprising: a handheld electromechanical driver, including: a handle portion;at least one motor supported in the handle portion; anda processor in communication with the at least one motor; anda surgical attachment, selectively coupleable to the handheld electromechanical driver, including: an anvil portion;a staple, blade, sensor, and reservoir (SBSR) portion coupleable to the anvil portion; anda memory unit included within the SBSR portion configured to store identity information and status information of the surgical attachment,wherein, when the surgical attachment is coupled to the handheld electromechanical driver, the processor of the handheld electromechanical driver queries the memory unit of the surgical attachment regarding the identity information and status information. 2. The surgical system according to claim 1, wherein the surgical attachment is an interchangeable surgical attachment. 3. The surgical system according to claim 1, wherein the surgical attachment is an anastomosing, resecting, and stapling attachment. 4. The surgical system according to claim 1, wherein the identity information includes at least one of size, diameter, shape, length, and stapling arrangement of the surgical attachment. 5. The surgical system according to claim 1, wherein the status information includes at least one of functionality and usage history of the surgical attachment. 6. The surgical system according to claim 1, wherein the SBSR portion includes drive elements for moving the anvil portion. 7. The surgical system according to claim 1, wherein the SBSR portion includes motive elements for driving a blade and staples through tissue. 8. The surgical system according to claim 1, wherein the identity information and the status information are displayed on a remote display coupled to the handle portion. 9. The surgical system according to claim 1, wherein the identity information and the status information are displayed on a display integrated on the handle portion. 10. The surgical system according to claim 1, wherein the processor establishes when the handheld electromechanical driver is activatable based on the identity information and the status information received. 11. A surgical attachment for use in a surgical system having an electromechanical driver including a processor, the surgical attachment including: an anvil portion;a staple, blade, sensor, and reservoir (SBSR) portion coupleable to the anvil portion; anda memory unit included within the SBSR portion configured to store identity information and status information of the surgical attachment,wherein, when the surgical attachment is coupled to the electromechanical driver, the processor queries the memory unit of the surgical attachment regarding the identity information and status information. 12. The surgical attachment according to claim 11, wherein the surgical attachment is an interchangeable anastomosing, resecting, and stapling attachment. 13. Th... | Covidien LP | A61B017/068 A61B017/115 A61B005/00 A61B017/072 A61B017/11 A61B005/1459 A61B017/00 A61B017/29
US09806299 | B2 | 2017-10-31 | Cathode for thin film microbattery | 2014-07-24 | A battery comprising an anode comprising anode material in contact with a metal anode current collector. The battery further comprises a cathode comprising cathode material in contact with a cathode current collector comprising a transparent conducting oxide (TCO). The battery further comprises an electrolyte with a pH in a range of 3 to 7. | 1. A battery comprising: a first polymer flex substrate;an anode comprising an anode material in contact with a metal anode current collector, wherein the metal anode current collector of the anode is adhered to a surface of the first polymer flex substrate by a first adhesion metal layer;a cathode comprising a conducting cathode material in contact with a cathode current collector comprising a transparent conducting oxide (TCO);a second polymer flex substrate, wherein the transparent conducting oxide TCO is adhered to a surface of the second polymer flex substrate by a second adhesion metal layer;an electrolyte with a pH in a range of 3 to 7 located between the anode and the cathode; anda polymer bondable seal material surrounding the conducting cathode material, the electrolyte and the anode material. 2. The battery of claim 1, wherein the transparent conducting oxide (TCO) comprises one or more of the following: indium tin oxide (ITO), indium zinc oxide (IZO), Al-doped zinc oxide (AZO), and Ga-doped zinc oxide (GZO). 3. The battery of claim 1, wherein the conducting cathode material is MnO2. 4. The battery of claim 1, wherein the electrolyte is an aqueous solution. 5. The battery of claim 1, wherein the electrolyte is a gel. 6. The battery of claim 1, wherein the electrolyte comprises one or more of the following: zinc chloride, zinc acetate, and ammonium chloride. 7. The battery of claim 1, wherein the battery has a volume less than 1 cubic mm. 8. The battery of claim 1, wherein a thickness of the cathode current collector is in a range of 20 nm to 200 nm. 9. The battery of claim 1, wherein the battery has a flexibility defined by a radius of curvature of between about 4 mm and about 20 mm. 10. The battery of claim 1, wherein the battery has a thickness of less than about 150 microns. 11. The battery of claim 1, wherein the battery has an open circuit voltage of between 1.4V and 1.8V and has a discharge capacity of between 10 uA-hours and 200 uA-hours. 12. The battery of claim 1, wherein the conducting cathode material is one or more of the following: electroplated nickel hydroxide; and a mixture MnO2 with a binder. 13. The battery of claim 1, wherein the anode material is electroplated Zn on the anode current collector that comprises one or more of the following: Sn, In, Ni, and an alloy thereof. 14. The battery of claim 1, wherein the battery has a shape of an arc. 15. The battery of claim 1, wherein a potential for an onset of hydrogen evolution as measured on a cyclic voltammetry curve is less than −1.6V relative to Hg/HgO for a KOH solution. 16. The battery of claim 1, wherein the first and second adhesion metal layers are selected from Ti, Ta and TiW. 17. The battery of claim 1, wherein the first adhesion metal layer is present only on a portion of the first polymer flex substrate, and wherein the second adhesion layer is present only on a portion of the second polymer flex substrate. 18. The battery of claim 17, wherei... | International Business Machines Corporation | H01M006/04 H01M002/02 G02C007/08 A61B005/145 H01M004/38 H01M004/50 H01M004/66 H01M010/04 C25D005/02 H01M004/42 H01M004/48 C25D001/04 C25D001/22 H01M004/139 H01M004/02 C25D003/22 C25D003/54 C25D003/56 C25D005/10 C25D005/50
US09826918 | B2 | 2017-11-28 | Method and device for detecting the surface structure and properties of a probe | 2015-08-28 | A method for detecting the surface structure and composition of a sample by means of a scanning unit, in particular for detecting traces, which are induced by contact of the skin of the human body on the surface of an object or absorbed by means of a trace carrier. The sample and the scanning unit are moved in relation to one another. The sample surface is irradiated line-by-line using a light beam or laser beam emitted from the scanning unit. The light beam or laser beam reflected from the sample surface is detected, and a digital image of the topography of the sample surface and the intensity of the reflected light beam or laser beam is generated from deviations of the reflected light beam or laser beam from the emitted light beam or laser beam to illustrate the composition of the sample surface. | 1. A method for detecting the surface structure and composition of a sample by means of a scanning unit wherein the sample and the scanning unit are moved in relation to one another, whereina sample surface is irradiated line-by-line using a laser beam emitted from the scanning unit, having predefined laser beam diameter in a step width corresponding to the laser beam diameter, and the reflected laser beam is detected and analyzed coaxially in relation to the emitted laser beam, the laser beam reflected from the sample surface is detected and a digital image of the topography of the sample surface and the intensity of the reflected laser beam is generated from deviations of the reflected laser beam from the emitted laser beam and a run time, which is dependent on the distance of the scanning unit from the sample surface, of the laser beam, which is emitted from the scanning unit and reflected from the sample surface, is detected and analyzed to prepare a distance image corresponding to the topography of the sample surface. 2. The method as claimed in claim 1, wherein the laser beam is emitted at a wavelength which corresponds to a wavelength range of an absorption spectrum of a composition of the sample surface. 3. The method as claimed in claim 1, wherein the emitted laser beam is collimated to a laser beam diameter less than or equal to 1 mm. 4. The method as claimed in claim 1, wherein a thickness of internal layers of the sample, which differ from external layers, is analyzed. 5. The method as claimed in claim 1, wherein a phase shift between the laser beam emitted from the scanning unit and the laser beam reflected from the sample surface is detected and analyzed to ascertain the topography of the sample surface. 6. The method as claimed in claim 5, wherein the laser beam emitted from the scanning unit is sinusoidally modulated and, to determine the phase shift between the laser beam emitted from the scanning unit and the laser beam reflected from the sample surface, the reflected laser beam detected by the scanning unit is correlated with a reference signal synchronous with the emitted laser beam. 7. The method as claimed in claim 1, wherein an absorption-related deviation of the laser beam reflected from the sample surface from the laser beam emitted from the scanning unit is analyzed to prepare an intensity image corresponding to a chemical substance on and in the sample surface. 8. The method as claimed in claim 7, wherein the sample surface is scanned serially point-by-point using the modulated laser beam and image elements of a digital image are emulated from the distance and intensity measurements arranged in a matrix. 9. The method as claimed in claim 1, wherein the reflected laser beam is detected by means of an RGB image recording unit to ascertain color values of the scanned sample surface. 10. A device for detecting the surface structure and composition of a sample, comprising: a sample receptacle,a scanning unit compris... | Juergen Marx | A61B005/103 G01S017/89 G01N021/3563 A61B005/1172 A61B005/1455 A61B005/00 G01B011/24 G01S017/02 G01S017/42 G01S007/48 H04N005/33 H04N005/225 H04N003/09 G01N021/47 G01N021/84 G01S007/481
US09775999 | B2 | 2017-10-03 | System comprising a cochlear stimulation device and a second hearing stimulation device and a method for adjustment according to a response to combined stimulation | 2013-08-09 | There is provided a system comprising a first device for neural stimulation of a cochlea of a patient's ipsilateral ear; a second hearing stimulation device selected from the group consisting of a device for neural stimulation of the cochlea of the patient's contralateral ear, a device for acoustic stimulation of the patient's ipsilateral ear and a device for acoustic stimulation of the patient's contralateral ear and comprising means for stimulation of the respective ear; and a fitting device for adjusting at least one of the first stimulation device and the second stimulation device according to a response of the patient to combined stimulation by the first stimulation device and the second stimulation device. | 1. A system comprising: a first stimulation device configured to provide neural stimulation of a cochlea of a patient's ipsilateral ear;a second stimulation device selected from the group consisting of a device configured to provide neural stimulation of the cochlea of the patient's contralateral ear, a device configured to provide acoustic stimulation of the patient's ipsilateral ear, and a device configured to provide acoustic stimulation of the patient's contralateral ear; anda fitting device configured to adjust at least one of the first stimulation device and the second stimulation device according to a response of the patient to combined stimulation by the first stimulation device and the second stimulation device;the first stimulation device comprising a sound processor configured to generate a neural stimulation signal from an input audio signal; anda cochlear implant stimulation arrangement comprising a plurality of stimulation channels configured to stimulate the ipsilateral cochlea at various stimulation sites according to the neural stimulation signal, with each stimulation channel being attributed to a certain one of the stimulation sites;the fitting device comprising a unit configured to cooperate with the first stimulation device and with the second stimulation device in order to generate, in a synchronized manner, a probe neural stimulation signal to be supplied to the cochlear implant stimulation arrangement and a second probe stimulation signal be supplied to the second stimulation device, with the probe neural stimulation signal and the second stimulation signal being adapted to be recognized by the patient as carrying a target speech information in case that the patient is able to correctly fuse the stimuli resulting from the probe neural stimulation signal and the second stimulation signal in the auditory cortex, with the target speech information being divided onto the probe neural stimulation signal and the second stimulation signal, and with the probe neural stimulation signal and the second stimulation signal not carrying said target speech information when perceived separately by the patient,a unit for recording a perceptual behavioral response of the patient to the synchronized neural stimulation of the user's hearing with the probe neural stimulation signal and the second stimulation signal, anda unit for programming the first stimulation device and/or the second stimulation device according to the recorded perceptual response. 2. The system of claim 1, wherein the target speech information is a CVC (consonant-vowel-consonant). 3. The system of claim 2, wherein the target speech information is a word or a phoneme. 4. The system of claim 3, wherein the word or phoneme is synthesized. 5. The system of claim 1, wherein the target speech information is divided onto the probe neural stimulation signal and the second stimulation signal by allocation a first spectral region to the second stimulation signal only and a seco... | Advanced Bionics AG | A61N001/36 A61B005/12 H04R025/00 A61B005/16
US09801554 | B2 | 2017-10-31 | Sensor for determining body parameters during exercise | 2015-01-28 | A noninvasive light sensor for detecting heart beat signals has a circular support member engageable circumferentially with a body part of a person. There are a plurality of light emitters and light detectors located around a circumference of the circular support member for respectively emitting light signals into different areas of tissue surrounding the body part, and receiving reflected light signals from the different areas of tissue surrounding the body part. | 1. A device for detecting heart beat signals, comprising: a circular support member engageable circumferentially with a body part of a person,three light emitters located and spaced at 120 degree intervals about a the circumference of the circular support member for emitting light signals into different areas of tissue surrounding the body part;a plurality of light detectors located respectively about the circumference of the circular support member for detecting reflected light signals from the different areas of tissue surrounding the body part and for providing correspondingly different detected signals, each different detected signal comprising both a heart beat signal and a movement signal; andat least one signal processor configured to compare the different detected signals and to separate the heart beat signals from the movement signals, and to determine heart rate from the separated heart beat signals. 2. The device of claim 1 wherein the light emitters comprise both red and infra-red light emitters. 3. The device of claim 1 wherein the light detectors consist of three light detectors located 120 degrees apart about the circumference of the support member. 4. The device of claim 1 wherein the light emitters and light detectors are located in pairs, each pair consisting of one of the emitters and one of the detectors. 5. The device of claim 1 wherein the support member is one of an ear bud insertable within an ear canal and a band locatable about a wrist or arm. 6. The device of claim 1 wherein the body part is an ear canal, the circular support member is an ear piece insertable within the ear canal along an ear canal axis, the light emitters are located about a periphery of the ear piece for emitting light signals perpendicular to the ear canal axis and into the different areas of tissue surrounding the ear canal, and the light detectors are located about the periphery of the ear piece for detecting the reflected light signals from the different areas of tissue surrounding the ear canal. 7. The device of claim 6 wherein the ear piece is resiliently deformable for expansion within the ear canal, such that the ear piece is firmly locatable within the ear canal. 8. The device of claim 6 wherein the ear piece comprises an inner layer and an outer layer having a plurality of openings, and the light emitters and light detectors are located with the inner layer beneath the openings of the outer layer. 9. The device of claim 8 wherein the inner layer comprises foam and the outer layer comprises rubber. 10. The device of claim 6 wherein the ear piece has an outer surface for mating with the ear canal and a plurality of openings in the outer surface, and the light emitters and light detectors are located within the openings. 11. The device of claim 10 wherein the light emitters and light detectors are located below the outer surface of the ear piece. 12. The device of claim 10 wherein the openings have sides surrounding the light emitter... | Well Being Digital Limited | A61B005/02 A61B005/0205 A61B005/024 A61B005/00 A63B024/00 A63B071/06 A61B005/1455 G05B015/02 G06F003/16
US09788744 | B2 | 2017-10-17 | Systems for monitoring brain activity and patient advisory device | 2008-07-28 | A patient advisory device (“PAD”) and its methods of use. The PAD may be configured to alert the patient about an estimated brain state of the patient. In preferred embodiments, the PAD is adapted to alert the patient of the patient's brain state, which corresponds to the patient's propensity of transitioning into an ictal brain state, e.g., having a seizure. Based on the specific type of alert, the patient will be made aware whether they are highly unlikely to have a seizure for a given period of time, an elevated propensity of having a seizure, a seizure is occurring or imminent, or the patient's brain state is unknown. | 1. A system for monitoring a patient's brain activity, the system comprising: an electrode for detecting physiological signals from the patient;a communication assembly coupled to the electrode, the communication assembly configured to wirelessly transmit a data signal that is encoded with data that is indicative of the detected physiological signals; andan external patient advisory device configured to receive the data signal from the communication assembly, the patient advisory device comprising: a processing assembly comprising one or more brain state algorithms that process the data signal to estimate the patient's brain state;a brain state indicator configured to provide the estimated brain state, wherein the brain state indicator indicates: when the estimated brain state is a contra-ictal brain state comprising a protected state in which the patient has a risk of transitioning from a non-ictal state to an ictal state within an upcoming time period below a first risk level; andwhen the estimated brain state is a pro-ictal brain state comprising a prediction state in which the patient has a risk of transitioning from a non-ictal state to an ictal state within an upcoming time period above a second risk level, wherein the second risk level is higher than the first risk level; andan output assembly comprising a first indicator that provides an indication of a status of the patient advisory device, a second indicator that provides an indication of a status of the communication assembly, and a failure alert indicator that provides an indication that a failure has occurred with the patient advisory device or the communication assembly. 2. The system of claim 1 wherein the output assembly further comprises a third indicator that provides an indication when there is a problem with transmission of the data signal from the communication assembly or receipt of the data signal at the patient advisory device. 3. The system of claim 1 wherein the communication assembly is programmed to substantially continuously transmit the data signal to the patient advisory device. 4. The system of claim 1 wherein the patient advisory device comprises a memory assembly for storing the data signal. 5. The system of claim 1 wherein the brain state indicator comprises a plurality of brain state indicators that differentiate between the contra-ictal state, a pro-ictal state, and an ictal state. 6. The system of claim 1, the system further comprising: a therapy delivery device, wherein the therapy delivery device is adapted to automatically initiate therapy to the patient when the patient has an elevated propensity of seizure. 7. The system of claim 1 wherein the output assembly further comprises a communication indicator that provides an indication that the patient advisory device is out of communication range with the communication assembly. 8. The system of claim 1 wherein the output assembly further comprises a data indicator that provides an indication that data... | CYBERONICS, INC. | A61B005/04 A61B005/0478 A61B005/042 A61B005/0432 A61B005/044 A61B005/0448 A61B005/0476 A61B005/0496 A61B005/00
US09792408 | B2 | 2017-10-17 | Method and apparatus to detect transponder tagged objects and to communicate with medical telemetry devices, for example during medical procedures | 2009-10-27 | The presence or absence of objects tagged with transponders may be determined in an environment in which medical procedures are performed via an interrogation and detection system which includes a controller and a plurality of antennas positioned along a patient support structure. The antennas may be positioned along an operating table, bed, mattress or pad, sheet, or may be positioned on a drape, or shade. Respective antennas may successively be activated to transmit interrogation signals. Multiple antennas may be monitored for responses from transponders to the interrogation signals. For example, all antennas other than the antenna that transmitted the most recent interrogation signal may be monitored. Antennas may be responsive to force, a signal indicative of such force being produced. A wireless physiological condition monitor may detect patient physiological conditions and wirelessly transmit signals indicative of such. | 1. A system for use in a medical facility, the system comprising: a number of antennas; anda controller communicatively coupled to the number of antennas, the controller: selectively causes at least one of the number of antennas to emit wideband interrogation signals;provides wireless communications of at least one of data or instructions with at least one piece of dedicated medical diagnostic, therapeutic or monitoring telemetry equipment via the at least one of the number of antennas used to emit the wideband interrogation signals, the at least one piece of dedicated medical diagnostic, therapeutic or monitoring telemetry equipment having at least one respective active radio; anddetects a response signal, if any, from a number of transponders physically detects coupled to a number of objects used in medical procedures via the at least one of the number of antennas used to emit the wideband interrogation signals and used to provide wireless communications, the response signal responsive to the emitted wideband interrogation signals. 2. The system of claim 1 wherein at least a plurality of the number of antennas are carried by a patient support structure that carries a patient during a medical procedure. 3. The system of claim 2 wherein at least one of the number of antennas is a handheld antenna that inductively cooperates with the plurality of antennas carried by the patient support structure. 4. The system of claim 2 wherein at least one of the number of antennas is a handheld antenna, and one of the handheld antenna or the plurality of antennas carried by the patient support structure transmits wideband interrogation signals and the other one of the handheld antenna or the plurality of antennas carried by the patient support structure receives response signals, if any, to the wideband interrogation signals. 5. The system of claim 2 wherein the number of antennas includes at least three loop antennas spaced along a longitudinal axis of the patient support structure, and the patient support structure is one of an operating table, a bed, a mattress, or a sheet. 6. The system of claim 1 wherein at least a plurality of the number of antennas are carried by a shade that extends from a light fixture positioned relatively above a patient support structure that carries a patient during a medical procedure. 7. The system of claim 1 wherein the controller receives data signals from the at least one piece of dedicated medical diagnostic, therapeutic or monitoring telemetry equipment via at least one of the number of antennas, the data signals representing physiological characteristics of a patient measured by the at least one piece of dedicated medical diagnostic, therapeutic or monitoring telemetry equipment. 8. The system of claim 1 wherein the controller receives data signals from the at least one piece of dedicated medical diagnostic, therapeutic or monitoring telemetry equipment via at least one of the number of antennas, the data signals rep... | Covidien LP | A61N001/00 G06F019/00 A61B005/00 A61B090/96 A61B090/90 A61B090/94 A61B090/98 A61B005/03 A61B005/1455 A61B017/00
US09819754 | B2 | 2017-11-14 | Methods, systems and devices for activity tracking device data synchronization with computing devices | 2015-01-15 | Methods, devices and system are provided. One method includes capturing activity data associated with activity of a user via a device. The activity data is captured over time, and the activity data is quantifiable by a plurality of metrics. The method includes storing the activity data in storage of the device and, from time to time, connecting the device with a computing device over a wireless communication link. The method defines using a first transfer rate for transferring activity data captured and stored over a period of time. The first transfer rate is used following startup of an activity tracking application on the computing device The method also defines using a second transfer rate for transferring activity data from the device to the computing device for display of the activity data in substantial-real time on the computing device. | 1. A method, comprising, capturing activity data associated with activity of a user via a device, the activity data being captured over time, the activity data quantified by a plurality of metrics;storing the activity data in storage of the device;connecting the device with a computing device over a wireless communication link when the device is within range of communication and while an activity tracking application on the computing device is closed;using a first transfer rate for transferring activity data captured and stored over a period of time; andusing a second transfer rate for transferring activity data from the device to the computing device for display of the activity data in substantial-real time on the computing device when the activity tracking application is open, the method is executed by a processor. 2. The method of claim 1, wherein the first transfer rate is selected for transferring the activity data from the storage of the device to the computing device. 3. The method of claim 1, wherein the second transfer rate is selected to transfer activity data, upon or substantially after being produced in the device, to the computing device for display by the activity tracking application on the computing device in substantial real-time. 4. The method of claim 1, wherein the first transfer rate is set at a scaled-up connection interval, while the second transfer rate is set at a scaled-down connection interval, the first transfer rate being higher than the second transfer rate. 5. The method of claim 1, wherein the first transfer rate for transferring activity data captured and stored over the period of time is for one or more background transfers of activity data from the device to the computing device, the background transfers occurring at predefined intervals. 6. The method of claim 1, further comprising, detecting opening of the activity tracking application on the computing device;upon opening the activity tracking application, initiating transfer of firmware updates, using the first transfer rate, to the device;wherein the firmware updates being triggered in response to installation instructions obtained from a server by the activity tracking application. 7. The method of claim 1, wherein the plurality of metrics include one or more of step count metrics, or stair count metrics, or distance traveled metrics, or active time metrics, or calories burned metrics, or heart rate metrics, or sleep metrics, or combinations two or more thereof. 8. The method of claim 1, further comprising, transferring the activity data to an activity management server to update the plurality of metrics associated with the activity of the user, the activity management server, from time to time, synchronizing activity data and settings with the computing device. 9. The method of claim 8, further comprising, receiving setting configurations from the activity management server, as obtained from the computing device, the setting configurations being r... | Fitbit, Inc. | G06F017/40 G06F019/00 G06M003/00 G08C017/02 H04L029/08 A61B005/00 G01D021/02 A61B005/22 G11B027/10 G01C005/00 G01P015/00 A61B005/0205 A61B005/11 G01C022/00 A61B005/024
US09775549 | B2 | 2017-10-03 | Temperature insensitive in vivo analyte devices, methods and systems | 2015-06-11 | Disclosed herein are membrane structures for use in analyte sensors, where the membrane structures exhibit low temperature sensitivity. | 1. A membrane structure comprising: an enzyme layer; anda membrane disposed proximate the enzyme layer, wherein the membrane comprises: a polymer comprising a heterocyclic nitrogen containing component; anda monoamine polyetheramine component that is covalently bonded to the nitrogen of the heterocyclic nitrogen containing component of the polymer. 2. The membrane structure according to claim 1, wherein the heterocyclic nitrogen containing component comprising a heteroaromatic ring system. 3. The membrane structure of claim 1, wherein the polymer comprises a heterocyclic nitrogen containing component is a poly(4-vinylpyridine-co-styrene) polymer. 4. The membrane structure of claim 3, wherein the poly(4-vinylpyridine-co-styrene) polymer comprises a compound of the formula:wherein x and y are each positive integers. 5. The membrane structure of claim 4, wherein the ratio of x and y is from 1:1 and 5:1. 6. The membrane structure of claim 1, wherein the monoamine polyetheramine component comprises a polyethylene oxide component and polypropylene oxide component. 7. The membrane structure of claim 6, wherein the molar ratio of the polypropylene oxide component and polyethylene oxide component is 1:35 to 35:1. 8. The membrane structure of claim 1, wherein the monoamine polyetheramine component has a molecular weight of 500 daltons to 5000 daltons. 9. The membrane structure of claim 1, wherein the monoamine polyetheramine component comprises a compound of the formula:wherein n and m are each positive integers. 10. The membrane structure of claim 1, wherein the membrane is disposed on the enzyme layer. | Abbott Diabetes Care Inc. | A61B005/1455 A61B005/05 A61B005/1486 A61B005/00 A61B005/145
US09814399 | B2 | 2017-11-14 | Biological information detection apparatus | 2014-03-11 | A biological information detection apparatus includes a detection unit which has a light receiving unit receiving light from a subject, a light transmitting member which is provided on a housing surface side in contact with the subject of the biological information detection apparatus, transmits light from the subject, and comes into contact with the subject when measuring biological information of the subject, and a diaphragm unit which is provided between the light transmitting member and the detection unit, between the light transmitting member and the subject, or inside the light transmitting member, and narrows light from the subject in an optical path between the subject and the detection unit. The detection unit includes a light emitting unit which emits light to the subject, and the biological information detection apparatus has a light shielding unit which is provided between the light receiving unit and the light emitting unit. | 1. A biological information detection apparatus comprising: a housing surface;a detection unit which has a light emitting unit which is configured to emit light to a subject and a light receiving unit configured to receive light from the subject, wherein the biological information detection apparatus is configured to be attached to the subject externally;a light transmitting member which is provided on the housing surface configured to be in contact with the subject of the biological information detection apparatus, transmits light from the light emitting unit and the subject, and configured to be placed adjacent to and comes into contact with skin of the subject while measuring biological information of the subject;a first diaphragm provided adjacent to or integral with the light transmitting member, wherein the first diaphragm is provided on light receiving unit side, wherein the first diaphragm is provided between the light transmitting member and the light receiving unit, between the light transmitting member and the subject, or inside the light transmitting member, wherein the first diaphragm includes an opening that is configured to narrows light from the subject in an optical path between the subject and the light receiving unit, wherein: an angle between a first line, which connects a first end portion of a light transmitting region of the light transmitting member and a first end portion of the light receiving unit as an end portion away from the first end portion of the light transmitting region out of two end portions of the light receiving unit, and an optical axis of the light receiving unit is θr,an angle between a second line, which connects the first end portion of the light receiving unit and an end portion on an opening side of the first diaphragm, and the optical axis is θa, andθa is less than θr; anda processing unit for processing an output of the detection unit. 2. The biological information detection apparatus according to claim 1, further comprising: a second diaphragm adjacent to the first diaphragm, wherein the second diaphragm is provided on light emitting unit side, the second diaphragm includes an opening that narrows light from the light emitting unit. 3. The biological information detection apparatus according to claim 2, wherein an area of the opening of the second diaphragm provided on the light emitting unit side is smaller than an area of the opening of the first diaphragm provided on the light receiving unit side. 4. The biological information detection apparatus according to claim 1, further comprising: a light shielding unit which is provided between the light receiving unit and the light emitting unit. 5. The biological information detection apparatus according to claim 4, wherein: a height of the light shielding unit in a direction orthogonal to the housing surface is H1, a height of a lower surface which is the surface of the first diaphragm on the detection unit side is H2, and H1 is greater than H2. ... | Seiko Epson Corporation | A61B005/02 A61B005/024 A61B005/00
US09804112 | B2 | 2017-10-31 | Selective characterization of material under test (MUT) with electromagnetic impedance tomography and spectroscopy | 2014-11-19 | A method of extracting complex impedance from selected volumes of the material under test (MUT) combined with various embodiments of electrode sensor arrays. Configurations of linear and planar electrode arrays provide measured data of complex impedance of selected volumes, or voxels, of the MUT, which then can be used to extract the impedance of selected sub-volumes or sub-voxels of the MUT through application of circuit theory. The complex impedance characteristics of the sub-voxels may be used to identify variations in the properties of the various sub-voxels of the MUT, or be correlated to physical properties of the MUT using electromagnetic impedance tomography and/or spectroscopy. | 1. A system comprising: an array of electrodes for communicating with a surface and a subsurface beneath the surface;a signal generator operably connected with the array of electrodes; andat least one computing device operably connected with the signal generator and the array of electrodes, the at least one computing device configured to: instruct the signal generator to transmit a first set of tomographic signals from a first subset of the array of electrodes into the surface and the subsurface;obtain a first return signal from the array of electrodes after the transmitting of the first set of tomographic signals;instruct the signal generator to transmit a second set of tomographic signals from a second subset of the array of electrodes into the surface and the subsurface; based upon the first return signal, the second subset of the array of electrodes including at least one electrode not included in the first subset of the array of electrodes;obtain a second return signal from the array of electrodes after the transmitting of the second set of tomographic signals; andcombine the first return signal and the second return signal to determine a characteristic of at least one of the surface or the subsurface. 2. The system of claim 1, wherein the array of electrodes includes a planar array of electrodes. 3. The system of claim 1, wherein a number of electrodes in the array of electrodes is equal to a number of combined layers, n, of the surface and the subsurface, plus one, n+1. 4. The system of claim 1, wherein a center-to-center spacing between adjacent electrodes in the array of electrodes is related to the minimum thickness of the layers of the surface and the subsurface to be measured. 5. The system of claim 1, wherein the array of electrodes includes a linear array of equally spaced electrodes configured to operate at a single frequency, and wherein the array of electrodes includes a single transmitting electrode positionally fixed in the linear array. 6. The system of claim 5, wherein the array of electrodes includes a plurality of linear sub-arrays of electrodes including an equal center to center spacing between adjacent electrodes within each sub-array, wherein a spacing between each sub-array of electrodes is at least as great as the spacing between the adjacent electrodes within each sub-array. 7. The system of claim 6, wherein the array of electrodes includes a planar array of electrodes including a plurality of linear sub-arrays of electrodes, wherein each of the plurality of sub-arrays of electrodes includes a single transmitting electrode positionally fixed in each sub-array. 8. The system of claim 7, wherein a center-to-center spacing between adjacent electrodes in each of the plurality of sub-arrays is equal to a minimum thickness of the layer to be detected in the subsurface, and wherein each of the plurality of sub-arrays is separated from an adjacent sub-array by a distance at least as great as the spacing between the ad... | TransTech Systems, Inc. | G01R027/02 G01V003/08 A61B005/04 G01N027/02 A61B005/053 G01R033/32 A61B005/0478 G01N022/04 G01N022/02
US09801698 | B2 | 2017-10-31 | Estimating a surface texture of a tooth | 2015-10-23 | Embodiments for estimating a surface texture of a tooth are described herein. One method embodiment includes collecting a sequence of images utilizing multiple light conditions using an intra-oral imaging device and estimating the surface texture of the tooth based on the sequence of images. | 1. A computing device implemented method for estimating a surface texture of a tooth, comprising: collecting a sequence of images utilizing multiple light conditions using an intra-oral imaging device, wherein the multiple light conditions include capturing different light intensities from multiple light sources, and wherein each light source from the multiple light sources generate light having a different light intensity than each of the other light sources;comparing the light intensities from at least two of the images within the sequence of images to one another; andestimating the surface texture of the tooth based on the comparison of the light intensities from the at least two images within the sequence of images, wherein the estimated surface texture is added to a dental model providing dental detail. 2. The method of claim 1, wherein the method further includes: turning different light sources on and off to move an illumination site of the different light sources while collecting the sequence of images. 3. The method of claim 1, wherein the method further includes: turning different light sources on and off independently to move an illumination site of the different light sources while collecting the sequence of images. 4. The method of claim 1, wherein the method further includes: turning different light sources on and off to move an illumination site of the different light sources from at least a first position to a second position while collecting the sequence of images. 5. The method of claim 1, wherein the multiple light conditions include light generated from light sources in different locations. 6. The method of claim 1, wherein the multiple light conditions include light generated from a first combination of light sources to create a first light condition and a second, different combination of light sources to generate a second light condition. 7. The method of claim 1, wherein the multiple light conditions include light generated from multiple light sources having different frequencies. 8. The method of claim 1, wherein the multiple light conditions include light generated from multiple light sources having different light spectrums. 9. The method of claim 1, wherein the multiple light conditions include light generated from a moving light source. 10. The method of claim 1, wherein the method includes comparing the collected sequence of images to a dental model of the tooth. 11. The method of claim 10, wherein comparing the collected sequence of images includes automatically computing a three-dimensional alignment between the sequence of images and the dental model. 12. A non-transitory computing device readable medium having executable instructions that can be executed by a processor to cause a computing device to perform a method for estimating a surface texture of a tooth, comprising: receiving a sequence of scanned images of a tooth from a dental device, the sequence including a number of individual images, each in... | Align Technology, Inc. | G06K009/00 A61C009/00 A61B005/00 A61B001/00 A61B001/06 A61B001/24 A61C001/08
US09826914 | B2 | 2017-11-28 | Functional analysis of neurophysiological data | 2014-04-17 | A method for functional analysis of neurophysiological data by decomposing neurophysiological data and EEG signal to form a plurality of signal features. The signal features may then optionally be analyzed to determined one or more patterns. | 1. A method of determining a brain activity pattern, the method comprising: using an EEG or brain imaging tool for obtaining neurophysiological data from a subject, while or after said subject is performing a task;using a data processor communicating with said EEG or brain imaging tool for:decomposing the data into a plurality of waveforms;identifying within said waveforms three-dimensional data patterns according to latency, amplitude and frequency of peaks of said waveforms; anddetermining brain activity patterns based on said data patterns; anddisplaying a graphic presentation of said activity patterns, said graphic presentation presenting brain regions that are active while or after said subject is performing said task. 2. The method according to claim 1, further comprising decomposing the data into a plurality of peaks. 3. The method according to claim 1, wherein the data are decomposed to a plurality of overlapping sets of waveforms. 4. The method according to claim 1, further comprising comparing said data patterns to a previously determined pattern. 5. The method according to claim 1, further comprising searching through a database of previously determined data patterns and selecting from said database a pattern closest to at least one of said identified data patterns. 6. The method according to claim 1, further comprising obtaining neurophysiological data also from at least one additional subject for a particular behavioral process, and associating said brain activity patterns with said behavioral process. 7. The method according to claim 1, wherein said neurophysiological data comprise data pertaining to a spontaneous brain activity. 8. The method according to claim 1, further comprising clustering said waveforms according to at least one of said latency, said amplitude and said frequency of peaks of said waveforms, to provide a plurality of clusters. 9. The method according to claim 1, further comprising determining a causality relation among data pattern components, and utilizing said relation to determine an activity network among said data patterns. 10. The method according to claim 9, further comprising determining a brain network activity (BNA) correlation to said activity network. 11. The method according to claim 10, wherein said determination of said BNA correlation comprises determining synchronization, or lack thereof, between a plurality of areas of the brain. 12. The method according to claim 10, wherein said identifying said data patterns comprises identifying source localizations for said BNA. 13. The method according to claim 12, further comprising comparing said data patterns to a previously determined pattern and correcting said source localizations based on said comparison. 14. The method according to claim 1, wherein said neurophysiological data comprises EEG signals. 15. The method according to claim 1, wherein said decomposing the data is effected by a wavelet transform. 16. The method according to c... | Elminda Ltd. | A61B005/04 A61B005/0476 A61B005/0482 A61B005/00
US09788731 | B2 | 2017-10-17 | Optical endoluminal far-field microscopic imaging catheter | 2012-02-21 | An optical endoluminal far-field microscopic imaging catheter comprises a light generating system, a first light delivery conduit for propagating light generated by the light generating system and a light distributor configured to redirect light propagated by the delivery conduit into a direction of an object to be imaged. A discriminator is configured for capturing light reflected from the object incident on a window of the discriminator from a particular direction and transmitting only the light captured from the particular direction to a second light delivery conduit. A drive mechanism is configured to sweep the window through a plurality of directions in a predictable pattern for matching each light capture event in the window with a direction of the window during the event. An analyzer matches the direction of the window with an associated light capture event and generate a visible image based on a mosaic of the captured light. | 1. An imaging system, comprising: a light generating system;a fiber optic cable for propagating light generated by the light generating system;a means for redirecting light propagated by the fiber optic cable into a direction of an object to be imaged;a directional light discriminator configured for capturing light reflected from the object incident on a window of the discriminator from a particular direction and transmitting only the light captured from the particular direction to a second fiber optic cable or into the same first fiber optic cable carrying a net signal that corresponds to multiple wavelengths of light, wherein the directional light discriminator comprises a selective window, the selective window comprising:a transparent core having a first refractive index;a transparent cladding around the transparent core having a second refractive index; andan opaque housing around the cladding;wherein the second refractive index is between 0.001 and 0.030 less than the first refractive or is between 0.001 and 0.030 greater than the first refractive index;a drive motor coupled to the window, configured for sweeping the window through a plurality of directions in a pattern for matching each light capture event in the window with a direction of the window during the event;an optical to electrical conversion module configured for converting an optical signal corresponding to an light capture event into an electrical signal capable of being reconstructed into a visible image; andan optical light tube configured to receive a rotating optical light guide through a central bore, the optical light tube having channels distributed within a periphery of the optical light tube configured for receiving the fiber optic cable;an analyzer coupled to the fiber optic cable, configured to match the direction of the window with an associated light capture event and generate a visible image based on a mosaic of the captured light;wherein the optical light tube is configured to translate in a linear direction while maintaining a fixed rotational position;wherein the drive motor is configured to spin the directional light discriminator during light capture events; andwherein the analyzer comprises a processor operable to reconstruct an image from reflected light data received from the directional light discriminator, by performing a reconstruction process comprising: receiving rotation speed data of the directional light discriminator;receiving light capture event rate data;assigning a pixel value based on a signal generated by the light capture event;assigning a first set of consecutive pixels to a first linear array based on the rotation speed data and the light capture event rate data;assigning a second set of consecutive pixels to a second linear array adjacent to the first linear array based on the rotation speed data and the light capture event rate data; andrepeating the steps of assigning consecutive pixels to subsequent linear arrays based on the rotation ... | Jaywant Philip Parmar | A61B006/00 A61B005/00 A61B001/00 A61B001/04 A61B001/07
US09788734 | B2 | 2017-10-17 | Defibrillator display | 2016-05-23 | An external defibrillator system is provided. The system includes: a graphical display; one or more sensors for obtaining data regarding chest compressions performed on a patient; and a controller configured to display on the graphical display numeric values for depth and/or rate of the chest compressions based upon the data from the one or more sensors. A method for using an external defibrillator including the steps of: obtaining data regarding chest compressions performed on a patient; and displaying on a graphical display screen of the defibrillator numeric values for depth and/or rate of the chest compressions based upon the data is also provided. | 1. An external defibrillator system comprising: a display;an accelerometer comprised in a housing configured to be placed on a patient's sternum for obtaining data comprising at least one of acceleration, velocity or displacement of the sternum as a result of chest compressions performed on the patient;a processor configured to determine sternal displacement and compression rate from the data and to compare the sternal displacement to a predetermined maximum acceptable sternal displacement value and a predetermined minimum acceptable sternal displacement value for sternal displacement; anda controller configured to:display on the display one or more numeric values of a unit length of the determined sternal displacement during chest compressions, the displayed values having at least two significant figures and a resolution of less than an inch, provide a visual indication of when the determined sternal displacement is above the predetermined maximum acceptable sternal displacement value and when the determined sternal displacement is below the predetermined minimum acceptable sternal displacement value, the visual indication comprising highlighting of the numerical value of the unit length of the determined sternal displacement, anddisplay on the display one or more numeric values for compression rate. 2. The external defibrillator system of claim 1, wherein the highlighting of the numerical value of the unit length of the determined sternal displacement comprises changing a color of the numerical value of the unit length of the determined sternal displacement. 3. The external defibrillator system of claim 1, wherein the processor is further configured to compare the compression rate to a predetermined maximum acceptable rate value and a predetermined minimum acceptable rate value and provide a visual indication of when the compression rate is above the predetermined maximum acceptable rate value or below the predetermined minimum acceptable rate value. 4. The system of claim 1, further comprising one or more electrodes for receiving information about the patient's heart activity, and wherein the controller is configured to display on the display, with the numeric values, an electrocardiogram trace of the patient. 5. The system of claim 4, wherein displaying the electrocardiogram trace comprises moving the electrocardiogram trace laterally across the display. 6. The system of claim 4, wherein the controller is further configured to display an unfiltered electrocardiogram, determine that chest compressions have commenced on the patient, and subsequently display a filtered electrocardiogram. 7. The system of claim 4, further comprising one or more sensors for obtaining CO2 data for the patient, wherein the controller is configured to display one or more numbers representing the CO2 data along with the electrocardiogram. 8. The system of claim 1, wherein the controller is configured to determine the numeric value of the unit length of the dete... | ZOLL MEDICAL CORPORATION | A61N001/39 A61B005/0205 A61H031/00 G06F019/00 G09B019/00 A61B005/044 A61B005/11 A61B005/145 A61B005/00 A61N001/04 G09B023/28 A61B005/083 A61B005/0402 A61B005/08 A61B005/0408
US09814428 | B2 | 2017-11-14 | Method and system for providing analyte monitoring | 2015-08-22 | Methods and apparatuses for determining an analyte value are disclosed. | 1. A glucose monitoring system, comprising: an in vivo glucose sensor that generates signals corresponding to a monitored glucose level; anda glucose data processing unit that determines noise filtered glucose data based on noise filtered estimates of glucose data generated from the signals corresponding to the monitored glucose level, and that compensates for a signal condition in the noise filtered glucose data to increase accuracy of the generated glucose data, the glucose data processing unit further including an output unit to output information corresponding to the noise filtered glucose data, the output unit including a display to display the information corresponding to the noise filtered glucose data; wherein the glucose data processing unit determines a fast rate estimate and a slow rate estimate based on the noise filtered estimates of the glucose data to identify the signal condition in the noise filtered glucose data. 2. The system of claim 1, wherein the glucose data processing unit includes a filter that clips the signal condition compensated noise filtered glucose signals to be within a predetermined range. 3. The system of claim 2, wherein the clipped signal condition compensated noise filtered glucose signals are not less than the noise filtered signals. 4. The system of claim 2, wherein the clipped signal condition compensated noise filtered glucose signals are not greater than a predetermined safety multiple of the noise filtered signals. 5. The system of claim 1, wherein the glucose data processing unit identifies the signal condition in the noise filtered glucose data based on a spike in the glucose data rate of change that exceeds a predetermined threshold. 6. The system of claim 1, wherein spikes in the fast rate estimate and the slow rate estimate are associated with the signal condition in the noise filtered glucose data. 7. The system of claim 1, further including an infusion device configured to communicate with the glucose data processing unit and to receive the information corresponding to the noise filtered glucose data. 8. The system of claim 7, wherein the infusion device determines a medication therapy profile based on the information corresponding to the noise filtered glucose data. 9. The system of claim 1, wherein the glucose data processing unit includes a data computing device. 10. The system of claim 1, further including sensor electronics that generates the glucose data based on the signals from the in vivo glucose sensor. 11. The system of claim 10, wherein the sensor electronics communicates the generated glucose data to the glucose data processing unit over a wireless communication link. 12. The system of claim 10, wherein the sensor electronics communicates the generated glucose data to the glucose data processing unit using a radio frequency communication protocol. 13. The system of claim 1, wherein the output unit includes a mobile telephone. 14. The system of claim 1, wherein the glucos... | Abbott Diabetes Care Inc. | A61B005/1473 A61B005/1486 G06Q050/24 A61B005/1468 A61B005/00 A61B005/145 G06F019/00
US09808171 | B2 | 2017-11-07 | Utilization of electrode spatial arrangements for characterizing cardiac conduction conditions | 2014-05-07 | A system and method are provided for determining electrophysiological data. The system comprises an electronic control unit that is configured to receive electrical signals from a set of electrodes, receive position and orientation data for the set of electrodes from a mapping system, compensate for position and orientation artifacts of the set of electrodes, compose cliques of a subset of neighboring electrodes in the set of electrodes, determine catheter orientation independent information of a target tissue, and output the orientation independent information to a display. The method comprising receiving electrogram data for a set of electrodes (80), compensating for artifacts in sensor positions in the mapping system (81), resolving the bipolar signals into a 3D vector electrogram in the mapping system coordinates (82), manipulating observed unipolar voltage signals and the tangent component of the e-field to estimate the conduction velocity vector (83), and outputting the catheter orientation independent information (84). | 1. A system for determining electrophysiological data, comprising: an electronic control unit configured to: receive electrogram data for a set of electrodes,compose a clique of at least three neighboring electrodes from the set of electrodes, wherein the electrogram data comprises a plurality of bipole signals taken between the at least three neighboring electrodes in the clique;receive position and orientation information for the set of electrodes from a mapping system;derive an E-field from the plurality of bipole signals of the clique, wherein bipole signals taken between the at least three neighboring electrodes in the clique overdetermine the E-field;determine catheter orientation independent information of a tissue based on the electrogram data and position and orientation information from the set of electrodes, wherein the catheter orientation independent information comprises a local tangent E-field and an orthogonal unit vector in the direction along which a wavefront is propagating (a); andoutput the orientation independent information to a user or process. 2. The system according to claim 1, wherein the orientation independent information further comprises at least one of a 2D vector electric field, a 3D vector electric field, a normal electric field, the eccentricity of the tangent electric field loop, derived normal and tangent bipolar electrogram signals, a wavefront conduction velocity, or a unit direction vector. 3. The system according to claim 1, wherein the electronic control unit is further configured to compensate for artifacts in the position and orientation information received from the mapping system. 4. The system according to claim 1, wherein the electronic control unit is further configured to resolve the electrogram data into a 3D vector electrogram in a coordinate system of the mapping system. 5. The system according to claim 1, wherein the catheter orientation independent information further comprises a wavefront conduction velocity and a 2D vector electric field. 6. The system according to claim 1, wherein the electronic control unit is further configured to resolve the electrogram data into a 2D vector electrogram in a coordinate system of the mapping system. 7. The system according to claim 6, wherein the catheter orientation independent information further comprises a conduction velocity magnitude. 8. The system according to claim 1, wherein the orientation independent information further comprises a single phase positive waveform. 9. The system according to claim 1, wherein the orientation independent information further comprises a local estimate of the conduction velocity vector. 10. The system according to claim 9, wherein the local estimate of the conduction velocity vector occurs on a beat-by-beat basis. 11. A system for determining electrophysiological data, comprising: an electronic control unit configured to: receive electrical signals from a set of electrodes;receive position and orientation ... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B005/04 A61B005/042 A61B005/0452 A61B005/06 A61B018/14 A61B005/00 A61B017/00 A61B018/00 A61B034/20
US09788765 | B2 | 2017-10-17 | Zwitterion surface modifications for continuous sensors | 2013-02-27 | Devices are provided for measurement of an analyte concentration, e.g., glucose in a host. The device can include a sensor configured to generate a signal associated with a concentration of an analyte; and a sensing membrane located over the sensor. The sensing membrane comprises a diffusion resistance domain configured to control a flux of the analyte therethrough. The diffusion resistance domain comprises one or more zwitterionic compounds and a base polymer comprising both hydrophilic and hydrophobic regions. | 1. A device for measurement of an analyte concentration, the device comprising: a sensor configured to generate a signal associated with a concentration of an analyte; anda sensing membrane located over the sensor, wherein the sensing membrane comprises an outermost domain and an enzyme-containing domain, wherein the enzyme-containing domain comprises one or more zwitterionic compounds. 2. The device of claim 1, wherein the surface of the sensing membrane comprises surface-active end groups configured for cross-linking with the one or more zwitterionic compounds, and wherein the one or more zwitterionic compounds are cross-linked to the surface-active end groups. 3. The device of claim 1, wherein the one or more zwitterionic compounds comprise a betaine compound or derivative thereof. 4. The device of claim 3, wherein the betaine compound or derivative thereof is a hydrolyzable cationic ester of a betaine. 5. The device of claim 4, wherein the hydrolyzable cationic ester of a betaine is a cationic poly(carboxybetaine) ester. 6. The device of claim 1, wherein the one or more zwitterionic compounds comprise at least one compound selected from the group consisting of a carboxyl betaine compound, a sulfo betaine compound, a phosphor betaine compound, and derivatives thereof. 7. The device of claim 1, wherein the one or more zwitterionic compounds comprise one or more compounds selected from the group consisting of cocamidopropyl betaine, oleamidopropyl betaine, octyl sulfobetaine, caprylyl sulfobetaine, lauryl sulfobetaine, myristyl sulfobetaine, palmityl sulfobetaine, stearyl sulfobetaine, betaine (trimethylglycine), octyl betaine, phosphatidylcholine, glycine betaine, poly(carboxybetaine), poly(sulfobetaine), and derivatives thereof. 8. The device of claim 1, wherein the one or more zwitterionic compounds comprise one or more compounds selected from the group consisting of poly(carboxybetaine), poly(sulfobetaine), and derivatives thereof. 9. The device of claim 1, wherein the sensor membrane comprises a bioprotective domain located most distal to the sensor, and wherein the bioprotective domain comprises at least one polymer selected from the group consisting of epoxies, polyolefins, polysiloxanes, polyethers, acrylics, polyesters, carbonates, and polyurethanes, and wherein the polyurethane is a polyurethane copolymer. | DexCom, Inc. | A61B005/145 A61B005/1486
US09785131 | B2 | 2017-10-10 | Device and method for contactless control of a patient table | 2015-05-07 | The invention relates to a device for contact-free control of a patient table includes a camera system for detecting at least one gesture of a user of the device, a processing unit for converting the at least one detected gesture into at least one command for controlling the patient table, and an execution unit for executing the at least one command for controlling the patient table. The device can furthermore include a monitoring system for monitoring the control of the patient table and an authentication system for authenticating the user. | 1. A device for contact-free control of a patient table that is movably mounted on a table base so as to be movable relative to a patient receiving opening of a medical examination apparatus, said device, comprising: a camera that detects at least one gesture made by a person that corresponds in direction or shape to a desired movement of said patient table relative to said opening, and that converts said gesture into image data that represents said at least one gesture;a computer processor in communication with said camera that receives said image data from said camera and is configured to convert said image data representing said at least one gesture into at least command for controlling movement of the patient table;an execution unit in communication with said computer processor that is provided with said command and that is configured to generate and emit a control signal to execute said at least one command that produces a movement of the patient table, relative to said patient-receiving opening, corresponding to said at least one command; anda monitoring system comprising a projector configured to project, at a projection site within a field of view of said person, a projected image of said patient table executing said movement, and thereby allowing said person to alter said movement executed by said table, if said movement executed by said table does not correspond to said desired movement, by making another gesture that is detected by said camera and converted into further image data that are converted into at least one further command provided to said execution unit, which alters said movement of said patient table. 2. A device as claimed in claim 1 comprising an authentication system configured to authenticate the person making such at least one gesture. 3. A device as claimed in claim 2 wherein said authentication system is configured to automatically authenticate said person by at least one optical marker associated with said person. 4. A device as claimed in claim 3 wherein said optical marker is an accessory worn by said person. 5. A device as claimed in claim 2 wherein said authentication system is configured to authenticate said person by detecting an authentication gesture made by said person. 6. A device as claimed in claim 2 wherein said authentication system is configured to authenticate said person by reading a chip card associated with said person. 7. A device as claimed in claim 2 wherein said authentication system is configured, after authenticating said person, is configured to orient said camera to a location of said person. 8. A device as claimed in claim 1 wherein said patient table and said camera is situated in a different room from the patient table. 9. A device as claimed in claim 1 wherein said camera is integrated into the patient table. 10. A device as claimed in claim 1 wherein said camera is situated remotely from the patient table, as a component of a monitoring system that monitors control of the pa... | Siemens Aktiengesellschaft | A61B006/04 G05B015/02 A61G013/02 A61B005/00 A61B034/00
US09795332 | B2 | 2017-10-24 | Method and pulse oximeter apparatus using chemical heating | 2012-11-26 | Method and apparatus for providing reliable blood oxygen (SaO2) and heart rate measurements includes a chemical energy heating source in conjunction with a harness that is adapted to secure the chemical energy heating source and a pulse oximeter probe proximate to a region of the body which is to be warmed prior to measurement. Preferably, the chemical energy heating source is in the form a mixture including a metal powder, which releases heat at a predetermined rate via oxidation of the metal powder when exposed to the atmosphere. The apparatus may be designed to be reusable or disposable and can be used in a transmission or reflectance mode, or both. | 1. A flexible self-heated pulse oximeter apparatus, comprising: a) a flexible harness including first and second sections joined by a bridge section;b) a probe removably secured to said bridge section and including a lead for transmitting a signal to a processor;c) said harness being adapted to hold said probe in a fixed position relative to a measurement site on the body so as to effectively envelope and position the extremity of the measurement site directly adjacent said probe, and including a pocket in one of said first and second sections thereof;d) a self-contained source of chemical heating mixture;e) said pocket being adapted to accommodate said self-contained source of chemical heating mixture therein; andf) said one of said first and second sections being made of an air permeable material so as to allow the air to come in contact with said chemical heating mixture and release heat. 2. The apparatus of claim 1, wherein: a) said chemical heating mixture comprises a metal. 3. The apparatus of claim 1, wherein: a) said chemical heating mixture comprises a metal mixture. 4. The apparatus of claim 3, wherein: a) the metal mixture comprises iron. 5. The apparatus of claim 3, wherein: a) the metal mixture comprises iron, activated charcoal, and vermiculite. 6. The apparatus of claim 2, further comprising: a) an air impermeable container for containing the flexible self-heated pulse oximeter. 7. The apparatus of claim 6, wherein: a) the metal mixture comprises a disposable mixture. 8. The apparatus of claim 1, wherein: a) said harness comprises a disposable harness. 9. The apparatus of claim 1, wherein: a) the air permeable material comprises a cloth material. 10. The apparatus of claim 1, wherein: a) said probe is adapted to function in transmission or reflectance mode, or both. 11. A non-invasive method of measuring a blood oxygen level (SaO2) or pulse rate of a subject, comprising the steps of: a) providing a flexible self-heated pulse oximeter apparatus, comprising: i) a flexible harness including first and second sections joined by a bridge section;ii) a probe removably secured to the bridge section and including a lead for transmitting a signal to a processor;iii) the harness being adapted to hold the probe in a fixed position relative to a measurement site on the body so as to effectively envelope and position the extremity of the measurement site directly adjacent the probe, and including a pocket in one of the first and second sections thereof;iv) a self-contained source of chemical heating mixture;v) the pocket being adapted to accommodate said self-contained source of chemical heating mixture therein; andvi) the one of the first and second sections being made of an air permeable material so as to allow the air to come in contact with the chemical heating mixture and release heat;b) exposing the chemical heating mixture to air;c) positioning the pulse oximeter apparatus at a measurement site on the subject so as to effectivel... | U.S. Department of Veterans Affairs | A61B005/1455 A61B005/1491 A61B005/00
US09826922 | B2 | 2017-11-28 | System and method for automatically discovering, characterizing, classifying and semi-automatically labeling animal behavior and quantitative phenotyping of behaviors in animals | 2016-03-11 | A method for studying the behavior of an animal in an experimental area including stimulating the animal using a stimulus device; collecting data from the animal using a data collection device; analyzing the collected data; and developing a quantitative behavioral primitive from the analyzed data. A system for studying the behavior of an animal in an experimental area including a stimulus device for stimulating the animal; a data collection device for collecting data from the animal; a device for analyzing the collected data; and a device for developing a quantitative behavioral primitive from the analyzed data. A computer implemented method, a computer system and a nontransitory computer readable storage medium related to the same. Also, a method and apparatus for automatically discovering, characterizing and classifying the behavior of an animal in an experimental area. Further, use of a depth camera and/or a touch sensitive device related to the same. | 1. A method for automatically discovering, characterizing and classifying the behavior of an animal in an experimental area, comprising: (a) using a 3D depth camera to obtain a video stream having a plurality of images of the experimental area with the animal in the experimental area, the images having both area and depth information;(b) removing background noise from each of the plurality of images to generate processed images having light and dark areas;(c) determining contours of the light areas in the plurality of processed images;(d) extracting morphometric parameters from both area and depth image information within the contours to form a plurality of multi-dimensional data points, each data point representing the posture of the animal at a specific time, including identifying a derivative of the spine outline;(e) clustering the data points at each specific time to output a set of clusters that are segmented form each other so that each cluster represents an animal behavior;(f) assigning each cluster a label that represents an animal behavior; and(g) outputting a visual representation of the set of clusters and corresponding labels. 2. The method of claim 1 wherein step (b) comprises: (b1) using the 3D depth camera to obtain a video stream having a plurality of baseline images of the experimental area without an animal present;(b2) generating a median of the baseline images to form a baseline depth image;(b3) subtracting the baseline depth image from each of the plurality of images obtained in step (a) to produce a plurality of difference images;(b4) performing a median filtering operation on each difference image to generate a filtered difference image; and(b5) removing image data that is less than a predetermined threshold from each of the plurality of filtered difference images to generate the processed images. 3. The method of claim 1, wherein step (c) comprises determining contours of all light regions in each processed image with a contour detection algorithm and tracking each contour with a Kalman filter. 4. The method of claim 1, wherein the label for each cluster is requested by a user interface after displaying video data representing each cluster. 5. The method of claim 1, wherein the behavior comprises a quantitative behavior primitive. 6. An apparatus for automatically discovering, characterizing and classifying the behavior of an animal in an experimental area, comprising: a 3D depth camera that generates a video stream having a plurality of images of the experimental area with the animal in the experimental area, the images having both area and depth information;a data processing system having a processor and a memory containing program code which, when executed: (a) removes background noise from each of the plurality of images to generate processed images having light and dark areas;(b) determines contours of the light areas in the plurality of processed images;(c) extracts morphometric parameters from both area and dep... | PRESIDENT AND FELLOWS OF HARVARD COLLEGE | A61B005/11 G06K009/00 G06T007/00 A01K067/00 A01K029/00 A61B005/00 G06T007/20 H04N013/02 G06F019/18
US09812509 | B2 | 2017-11-07 | Sensor | 2016-05-17 | According to one embodiment, a sensor includes a light emitter and a light sensor. The light emitter includes a first electrode, a second electrode, and a first light emitting layer. The second electrode is light-transmissive. The first light emitting layer is provided between the first electrode and the second electrode. The light sensor includes a third electrode, a fourth electrode, a fifth electrode, a first photoelectric conversion layer, and a second photoelectric conversion layer. the fourth electrode is light-transmissive. The fifth electrode is provided between the third electrode and the fourth electrode. The fifth electrode is light-transmissive. The first photoelectric conversion layer is provided between the third electrode and the fifth electrode. The second photoelectric conversion layer is provided between the fourth electrode and the fifth electrode. | 1. A sensor, comprising: a light emitter including a first electrode,a second electrode, the second electrode being light-transmissive, anda first light emitting layer provided between the first electrode and the second electrode; anda light sensor including a third electrode,a fourth electrode, the fourth electrode being light-transmissive,a fifth electrode provided between the third electrode and the fourth electrode, the fifth electrode being light-transmissive,a first photoelectric conversion layer provided between the third electrode and the fifth electrode, anda second photoelectric conversion layer provided between the fourth electrode and the fifth electrode, whereinthe first photoelectric conversion layer includes a third organic substance,the second photoelectric conversion layer includes a fourth organic substance,a wavelength of light absorbed by the third organic substance is different from a wavelength of light absorbed by the fourth organic substance, anda difference between an absorption peak wavelength of the first photoelectric conversion layer and an absorption peak wavelength of the second photoelectric conversion layer is 50 nm or more. 2. The sensor according to claim 1, wherein the first light emitting layer includes a first organic substance. 3. The sensor according to claim 1, wherein the light emitter further includes a second light emitting layer provided between the first light emitting layer and the second electrode,the first light emitting layer includes a first organic substance,the second light emitting layer includes a second organic substance, anda peak wavelength of a first light radiated from the first organic substance is different from a peak wavelength of a second light radiated from the second organic substance. 4. The sensor according to claim 3, wherein a difference between the peak wavelength of the first light and the peak wavelength of the second light is 50 nm or more. 5. The sensor according to claim 3, wherein the light emitter further includes a sixth electrode provided between the first light emitting layer and the second light emitting layer, the sixth electrode being light-transmissive. 6. The sensor according to claim 1, wherein a reflectance of the first electrode is higher than a reflectance of the second electrode. 7. The sensor according to claim 1, wherein at least a portion of the light emitter and at least a portion of the light sensor overlap in a first direction, the first direction being from the second electrode toward the first electrode. 8. The sensor according to claim 7, wherein the light emitter is provided on a first substrate, the first substrate being light-transmissive,the light sensor is provided on a second substrate, the second substrate being light-transmissive, andat least a portion of the first substrate and at least a portion of the second substrate are provided between the light emitter and the light sensor. 9. The sensor according to claim 7, wherein the lig... | Kabushiki Kaisha Toshiba | H01L029/08 H01L027/28 H01L051/50 H01L051/52 A61B005/024 A61B005/1455 G01N021/27
US09824446 | B2 | 2017-11-21 | Evaluating electromagnetic imagery by comparing to other individuals' imagery | 2014-03-14 | A computer-based method and system assists users performing medical imaging studies, such as echocardiograms. While the user manipulates the imaging device on an individual, the system automates image processing, relieving the user from the need to simultaneously operate an application and an imaging device. The system assesses an organ system, then through comparison with other imaging data sets, generates an evaluative assessment of the organ system. The system becomes increasingly image-aware and content-aware during an evaluation, and by sequential comparisons against progressive data sets, refines an evaluative assessment. | 1. A computer-based method of imaging an organ system or an anatomic structure in an individual, the method comprising the steps of: while an imaging device is being used on an individual:a) in a recognition module, executable by a processor, automating identification of and recognizing an organ system or an anatomic structure from imaging data acquired through the imaging device;b) in an assessment module, executable by the processor, assessing the identified and recognized organ system or anatomic structure of the individual to produce an assessment data set;c) in an evaluation module, executable by the processor, comparing the assessment data set of step b) to one or more comparison data sets in a multi-dimensional classification space to generate an evaluative assessment of the organ system or anatomic structure, wherein the evaluative assessment comprises a classification of the imaging data;d) in a recommendation module, executable by the processor, recommending to a user a modification of a view, an angle, an image mode or a combination thereof; ande) refining the evaluative assessment of step c) by repeating steps b), c), and d). 2. The method of claim 1, wherein step e) further comprises optimizing a captured imagery, recording a captured imagery, or a combination thereof, wherein the imaging data is fit into one or more multi-dimensional health classification spaces. 3. The method of claim 1, wherein the organ system is the individual's cardiovascular system, brain, or a combination thereof. 4. The method of claim 1, wherein the method further comprises animating the motion of the organ system or anatomic structure, creating one or more freeze-frames of the organ system or anatomic structure, abstracting the shape, boundaries, color, textures, or a combination thereof of the organ system or anatomic structure. 5. The method of claim 1, wherein the imaging device comprises one or more sensors. 6. The method of claim 5, wherein the one or more sensors is a magnetic sensor, gyroscope, accelerometer, electrocardiography sensor, electromyography sensor, camera, RGB sensor, green light sensor, red light sensor, blue light sensor, motion sensor, near infrared camera, infrared camera, thermal camera, GPS or WiFi. 7. The method of claim 1, further comprising ranking the individual on a spectrum based on one or more metrics, one or more diagnoses, or a combination thereof, wherein the spectrum is an ordered group of other individuals. 8. The method of claim 1, further comprising identifying metrics of the individual based upon the evaluative assessment, wherein the metrics are saved and tracked over a period of time. 9. The method of claim 1, further comprising triggering an alert, wherein the alert is generated based on the evaluative assessment. 10. The method of claim 9, wherein the alert is sent to the individual, a member of the individual's family, an individual's health proxy, a healthcare provider, one or more persons designated ... | - | G06T007/00 G06F019/00 A61B005/00
US09811951 | B2 | 2017-11-07 | Systems and methods for managing fault codes | 2016-02-10 | Various embodiments of the present invention provide systems and methods for managing fault codes triggered by one or more vehicles during operation. In general, various embodiments of the invention involve recording and analyzing fault codes triggered during a particular time period while a vehicle is in operation. As a result of the analysis, various embodiments of the invention may set a state for each of the identified fault codes, the state indicating a level of action to address the identified fault code. In particular embodiments, the states may be one of a caution state indicating one or more components or sub-systems of the vehicle should be monitored, a critical state indicating one or more components or sub-systems of the vehicle should be repaired, or an environmental state indicating failure or potential failure of one or more components or sub-systems of the vehicle may affect one or more environmental conditions. | 1. A system for managing fault codes associated with a vehicle, the system comprising: a memory storing a plurality of parameters and a plurality of fault codes, wherein each parameter is associated with a fault code and said plurality of parameters comprise: a threshold value and a consecutive value associated with the particular fault code, the threshold value defining a threshold number of times the particular fault code is generated during the predefined time period the vehicle was in operation and the consecutive value defining a number of consecutive predefined time periods during which the threshold value was met or exceeded; andone or more computer processors configured to: receive fault code information retrieved from the vehicle, the fault code information comprising one or more of said plurality of fault codes, the one or more fault codes generated during a particular time period the vehicle was in operation;retrieve parameters stored in the memory associated with the one or more fault codes; analyze the fault code information to identify each of the one or more fault codes signaling failure or potential failure of one or more components or sub-systems of the vehicle based at least in part on the parameters associated with the one or more fault codes in the fault code information; andset a state for each identified fault code, based at least in part on the parameters associated with each identified fault code, the state indicating a level of action associated with the identified fault code, wherein the state for the particular fault code is also set by the processors in response to the fault code information showing that the particular fault code was generated a number of times meeting or exceeding the threshold value during a number of consecutive time periods meeting or exceeding the consecutive value. 2. The system of claim 1, wherein the state comprises one of a caution state indicating the one or more components or sub-systems of the vehicle should be monitored, a critical state indicating the one or more components or subsystems of the vehicle should be repaired, or an environmental state indicating the failure or potential failure of the one or more components or sub-systems of the vehicle affects one or more environmental conditions. 3. The system of claim 1, wherein the one or more computer processors are configured to send one or more notifications to one or more parties for each of the identified fault codes, the one or more notifications comprising the state for each of the identified fault codes. 4. The system of claim 1, wherein the one or more computer processors are configured to initiate a work order for repair of at least one of the one or more components or sub-systems of the vehicle. 5. The system of claim 1, further comprising an interface configured to provide access to one or more procedures for diagnosing the one or more components or sub-systems associated with each of the identified fault codes. 6. The sy... | United Parcel Service of America, Inc. | G07C005/08 G07C005/02 G01M015/05 G07C005/00 G05B023/02 G06Q050/22 A61B005/00 G06F017/10 G06F019/00 G01M017/00 G06F007/00 G06F015/00 B64F005/60
US09775499 | B2 | 2017-10-03 | Augmented visualization for a surgical robot using a captured visible image combined with a fluorescence image and a captured visible image | 2014-07-16 | An endoscope with an optical channel is held and positioned by a robotic surgical system. A capture unit captures (1) a visible first image at a first time and (2) a visible second image combined with a fluorescence image from the light at a second time. An image processing system receives (1) the visible first image and (2) the visible second image combined with the fluorescence image and generates at least one fluorescence image. A display system outputs an output image including an artificial fluorescence image. | 1. A surgical system comprising: an endoscope held and positioned by the surgical system, the endoscope comprising: an optical channel for transporting light from tissue;a capture unit coupled to the optical channel, the capture unit being configured to capture: (1) at a first time, a first frame from the light, the first frame including a first visible image, the first visible image comprising a first plurality of visible color components; and(2) at a second time different from the first time, a second frame from the light, the second frame including a combination image, the combination image being a combination of a second visible image and a first fluorescence image, the combination image comprising a second plurality of color components, wherein one color component of the second plurality of color component includes one color component that is captured as a combination of the first fluorescence image and a visible color component, image and other color components of the second plurality of color components are captured as visible color components different from the visible color component of the one color component;an image processing system coupled to said capture unit, the image processing system comprising: a fluorescence image with artifacts generator coupled to receive the first and second frames, the fluorescence image with artifacts generator being configured to create a second fluorescence image including artifacts from the first frame and the second frame; anda fluorescence image extractor coupled to the fluorescence image with artifacts generator, the fluorescence image extractor being configured to create a third fluorescence image based on the second fluorescence image including artifacts; anda display system coupled to the image processing system, the display system being configured to output an image including an artificial fluorescence image, the artificial fluorescence image being based on the third fluorescence image. 2. The surgical system of claim 1, the output image further comprising: a blend of the first visible image and the artificial fluorescence image. 3. The surgical system of claim 1, the display system further comprising: a blend module coupled to receive the first frame and to receive the artificial fluorescence image; anda display unit coupled to the blend module. 4. The surgical system of claim 3, wherein the display system further comprises: a display mode select coupled to the blend module, the display mode select having a plurality of selectable modes, wherein upon selection of a first mode of the plurality of selectable modes; the display unit outputs only the first visible image; andwherein upon selection of a second mode of the plurality of selectable modes, the display unit outputs the image including the artificial fluorescence image. 5. The surgical system of claim 1, wherein the endoscope further comprises an illumination channel, and wherein the surgical system further comprises: a combination li... | Intuitive Surgical Operations, Inc. | A61B001/04 A61B001/06 A61B001/00 A61B001/313 A61B005/00 H04N013/02 A61B001/05 A61B090/30 A61B034/37 G01N021/64 A61B034/30 A61B090/00
US09820692 | B2 | 2017-11-21 | Wearable electrochemical sensors | 2013-05-10 | Methods, structures, devices and systems are disclosed for fabricating and implementing electrochemical biosensors and chemical sensors. In one aspect, a method of producing an epidermal biosensor includes forming an electrode pattern onto a coated surface of a paper-based substrate to form an electrochemical sensor, the electrode pattern including an electrically conductive material and an electrically insulative material configured in a particular design layout, and attaching an adhesive sheet on a surface of the electrochemical sensor having the electrode pattern, the adhesive sheet capable of adhering to skin or a wearable item, in which the electrochemical sensor, when attached to the skin or the wearable item, is operable to detect chemical analytes within an external environment. | 1. A method of producing an epidermal biosensor, comprising: forming an electrode pattern onto a coated surface of a paper-based substrate to form an electrochemical sensor, the electrode pattern including an electrically conductive material and an electrically insulative material configured in a particular design layout;attaching an adhesive sheet on a surface of the electrochemical sensor having the electrode pattern, the adhesive sheet capable of adhering to skin or a wearable item, wherein the adhesive sheet includes an outer coating layer on an external surface of the adhesive sheet not in contact with the electrode pattern;removing the outer coating layer from the adhesive sheet to enable adhesion of the electrochemical sensor to the skin or the wearable item via the adhesive sheet; andremoving the paper-based substrate from the electrochemical sensor to expose the electrode pattern to the external environment,wherein the electrochemical sensor, when attached to the skin or the wearable item, is operable to detect chemical analytes within an external environment. 2. The method of claim 1, wherein the outer coating layer includes polyvinyl alcohol (PVA). 3. The method of claim 1, wherein the coated surface includes a release agent material including cellulose acetate. 4. The method of claim 1, wherein the adhesive sheet includes polydimethylsiloxane (PDMS). 5. The method of claim 1, wherein the forming includes performing screen printing or inkjet printing the electrode pattern onto the coated surface of the paper-based substrate. 6. The method of claim 1, wherein the electrically insulative material includes a nonconductive ink including polyethylene terephthalate (PET) or polytetrafluoroethylene (PTFE). 7. The method of claim 1, wherein the electrically conductive material includes a conductive ink including at least one of gold, platinum, nickel, silver, or silver chloride. 8. The method of claim 1, wherein the electrode pattern further includes carbon fiber segments dispersed within the electrically conductive material of the electrode pattern. 9. The method of claim 1, wherein the electrode pattern further includes an electrically semi-conductive material. 10. The method of claim 9, wherein the electrically semi-conductive material includes a semi-conductive ink including at least one of amorphous carbon, carbon black, or graphite. 11. The method of claim 9, wherein the electrode pattern further includes carbon fiber segments dispersed within the electrically semi-conductive material of the electrode pattern. 12. A method of producing an epidermal biosensor, comprising: forming an electrode pattern onto a coated surface of a paper-based substrate to form an electrochemical sensor, the electrode pattern including an electrically conductive material and an electrically insulative material configured in a particular design layout;attaching an adhesive sheet on a surface of the electrochemical sensor having the electrode pattern,... | The Regents of the University of California | B29C065/52 B32B037/12 B32B037/26 B32B038/10 B32B038/14 B32B038/18 B32B043/00 A61B005/00 A61B005/145 A61B005/1477 A61B005/1486 B32B038/00
US09788819 | B2 | 2017-10-17 | Introducer for biopsy device | 2015-04-30 | A biopsy system comprises an introducer and a biopsy device. The introducer comprises a cannula and at least two leaves. The cannula comprises an ovular cross-section, an open distal end and an open proximal end. The ovular cross-section defines a lumen extending between the open distal end and the open proximal end. The ovular cross-section of the cannula includes at least two flat sides. The at least two leaves extend distally from the open distal end. At least a portion of each of the at least two leaves extend from a respective flat side of the at least two flat sides of the cannula. The biopsy device comprises a body and an elongate needle. The elongate needle extends distally from the body. The elongate needle includes an ovular cross-section defining two straight sides. The cannula of the introducer is configured to slidably receive the elongate needle of the biopsy device. | 1. A biopsy system, comprising: (a) an introducer, wherein the introducer includes: (i) a cannula including an oval-shaped cross-section, an open distal end and an open proximal end, wherein the oval-shaped cross-section defines a lumen extending between the open distal end and the open proximal end, wherein the oval-shaped cross-section of the cannula includes at least two flat sides and two non-flat sides, and(ii) at least two leaves extending distally from the open distal end of the cannula, wherein at least a portion of each of the at least two leaves extends from a respective flat side of the at least two flat sides of the cannula, wherein the at least two leaves are positioned relative to the cannula to define an open space between the at least two flat sides of the cannula; and(b) a biopsy device, wherein the biopsy device includes: (i) a body, and(ii) an elongate needle extending distally from the body, wherein the elongate needle includes a closed distal end and an oval-shaped cross-section, wherein the closed distal end comprises a sharp tip, wherein the oval-shaped cross-section defines two straight sides;wherein the cannula of the introducer is configured to slidably receive the elongate needle of the biopsy device, wherein the cannula is configured to form at least two openings between the cannula and the elongate needle along the two non-flat sides when the elongate needle is slidably received within the cannula. 2. The biopsy system of claim 1, wherein the at least two leaves of the introducer consists of two leaves. 3. The biopsy system of claim 1, wherein the introducer includes four leaves. 4. The biopsy system of claim 3, wherein the four leaves are arranged in a first leaf pair and a second leaf pair, wherein each leaf pair is associated with each flat side of the oval-shaped cross-section of the cannula. 5. The biopsy system of claim 4, wherein each leaf of the first leaf pair is separated by a first gap extending between each leaf of the first leaf pair, wherein the second leaf pair is separated by a second gap extending between each leaf of the second leaf pair. 6. The biopsy system of claim 1, wherein the cannula of the introducer further includes a top portion and a bottom portion, wherein the top and bottom portions are positioned away from the at least two leaves. 7. The biopsy system of claim 6, wherein each of the top and bottom portions define a chamfered portion. 8. The biopsy system of claim 7, wherein each chamfered portion extends distally from the distal end of the cannula of the introducer. 9. The biopsy system of claim 8, wherein each chamfered portion extends transversely between the at least two leaves. 10. The biopsy system of claim 1, wherein the biopsy device and the introducer further include complementary latching features configured to selectively secure the introducer to the biopsy device. 11. The biopsy system of claim 10, wherein the complementary latching features comprise a slot and a re... | Devicor Medical Products, Inc. | A61B005/00 A61B010/02 A61B017/34 A61B090/00
US09802040 | B2 | 2017-10-31 | Body motion detection device and human body stimulation apparatus comprising said device | 2013-09-25 | The human body stimulation apparatus comprises a left detection unit, a right detection unit, a determination unit, a control unit, a right stimulation unit and a left stimulation unit. The detection units generate a detection signal according to a walking motion of a human body. The determination unit divides the walking motion of one walking cycle into multiple walking stages on the basis of the detection signal. The determination unit determines that the walking motion is stopped on the basis of at least one duration time of the walking stages. The control unit controls the magnitude of the stimulation that the right stimulation unit and the left stimulation unit apply to the human body on the basis of the determination results of the determination unit. | 1. A body motion detection device that detects movement of a human body, the body motion detection device comprising: a detection unit that generates a detection signal that changes according to a walking movement of the human body; andan evaluation unit that separates the walking movement of one walking cycle into a plurality of walking phases based on the detection signal generated by the detection unit, wherein the evaluation unit determines whether or not the walking movement is still based on a duration time of at least one of the walking phases, wherein:the one walking cycle includes a stance phase and a swing phase that defined as specific walking phases; andthe evaluation unit determines divided walking phases within the stance phase or the swing phase based on the detection signal, and determines whether or not the walking movement is still based on a duration time of at least one of the divided walking phases. 2. The body motion detection device according to claim 1, wherein the evaluation unit determines that the walking movement is still when the duration time of at least one of the divided walking phases is longer than a still determination time. 3. The body motion detection device according to claim 2, wherein the evaluation unit includes a measurement unit that measures the duration time of each of the divided walking phases; andthe evaluation unit determines whether or not the walking movement is still based on a comparison of the duration time of at least one of the divided walking phases measured by the measurement unit with the still determination time. 4. The body motion detection device according to claim 3, further comprising a time setting unit that sets the still determination time based on the duration time of at least one of the walking phases measured by the measurement unit. 5. The body motion detection device according to claim 4, wherein the time setting unit sets the still determination time based on at least one of the duration times measured by the measurement unit before a current walking cycle. 6. The body motion detection device according to claim 4, wherein the time setting unit includes a determination time defining unit that sets a relationship of the duration time of each of the divided walking phases and the still determination time, wherein the determination time defining unit outputs the still determination time based on the relationship when the duration time of the divided walking phase is received from the measurement unit. 7. The body motion detection device according to claim 1, wherein the detection unit includesa right body detector including a first right body sensor that generates an output signal corresponding to a movement of a first right body region and a second right body sensor that generates an output signal corresponding to a movement of a second right body region, anda left body detector including a first left body sensor that generates an output signal corresponding to a movement... | Panasonic Intellectual Property Management Co., Ltd. | A61B005/00 A61N001/36 A61B005/11 A61N001/04
US09820696 | B1 | 2017-11-21 | Cuffless blood pressure measurement using handheld device | 2014-10-17 | A method and system for determining blood pressure are disclosed. The method comprises determining a plurality of heart sounds using a microphone of a handheld device and determining a pulse wave using a camera of the handheld device, wherein the plurality of heart sounds and the pulse wave are utilized to determine the blood pressure. The system includes a processor, a memory device coupled to the processor, and an application coupled to the memory device. The system further comprises a microphone coupled to the processor, wherein the microphone is utilized to determine a plurality of heart sounds and a camera coupled to the processor, wherein the camera is utilized to determine a pulse wave, further wherein the application, when executed by the processor, causes the processor to determine the blood pressure using the plurality of heart sounds and the pulse wave. | 1. A system for determining blood pressure, wherein the system includes a processor, a memory device coupled to the processor, and an application coupled to the memory device, wherein the system further comprises: a microphone coupled to the processor, wherein the microphone is utilized to determine a plurality of heart sounds; anda camera coupled to the processor, wherein the camera is utilized to determine a pulse wave, further wherein the application, when executed by the processor, causes the processor to: determine an ejection time (ET), a vascular transit time (VTT), and a heart rate from any of the plurality of heart sounds and the pulse wave; andperform regression analysis on received user-specific data, the ET, the VTT, and the heart rate to determine the blood pressure. 2. The system of claim 1, wherein to determine the ET further comprises to: determine the ET as a time difference between a first and a second heart sound of the plurality of heart sounds. 3. The system of claim 1, wherein to determine the VTT further comprises to: determine the VTT as a time difference between a first heart sound of the plurality of heart sounds and the pulse wave. 4. The system of claim 1, wherein to determine the heart rate further comprises to: determine the heart rate as an average of the reciprocal of a time difference between any of a first heart sound of one heartbeat and a first heart sound of an adjacent heartbeat, a second heart sound of one heartbeat and a second heart sound of an adjacent heartbeat, and a pulse wave of one heartbeat and a pulse wave of an adjacent heartbeat. 5. The system of claim 1, wherein the application, when executed by the processor, further causes the processor to: utilize calibration values from a reference device, the determined ET, the determined VTT, and the determined heart rate to determine the blood pressure. 6. The system of claim 1, wherein the system further comprises: another microphone coupled to the processor, wherein the another microphone is utilized to perform noise cancellation; andanother camera coupled to the processor, wherein the another camera is utilized to perform common mode noise cancellation. | - | A61B005/02 A61B005/00 A61B005/021 A61B007/04
US09808143 | B2 | 2017-11-07 | Optogenetic probe | 2013-06-18 | An optogenetic probe, an optogenetic system, and a method for fabricating an optogenetic probe are provided. The optogenetic probe has a proximal and a distal end, and includes an elongated body made of a body glass material and extending longitudinally between the proximal and distal ends. The optogenetic probe also includes at least one optical channel, each including an optical channel glass material having a refractive index larger than a refractive index of the body glass material, so as to guide light therealong. The optogenetic probes also includes at least one electrical channel, each including an electrical channel structure having an electrical conductivity larger than the electrical conductivity of the body glass material, so as to conduct electricity therealong. The optogenetic probe further includes at least one fluidic channel, each adapted for transporting fluid therealong. Each optical, electrical and fluidic channel extends longitudinally within the elongated body. | 1. An optogenetic probe having a proximal and a distal end, the optogenetic probe comprising: an elongated body made of a body glass material having a refractive index and an electrical conductivity, the elongated body extending longitudinally between the proximal and distal ends of the optogenetic probe;at least one optical channel extending longitudinally within the elongated body and comprising an optical channel glass material having a refractive index larger than the refractive index of the body glass material, so as to guide light therealong;at least one electrical channel extending longitudinally within the elongated body and comprising an electrical channel structure having an electrical conductivity larger than the electrical conductivity of the body glass material, so as to conduct electricity therealong; andat least one fluidic channel extending longitudinally within the elongated body and adapted for transporting fluid therealong. 2. The optogenetic probe according to claim 1, wherein the body glass material comprises a glass compound selected from the group consisting of AgPO3—WO3, TeO2, GeO2—PbO, Ga2O3, Sb2O3, SiO2, NaPO3—PbF2, NaPO3—WO3, NaPO3—Nb2O5, and a combination thereof. 3. The optogenetic probe according to claim 1, wherein the optical channel glass material of each optical channel comprises a glass compound selected from the group consisting of: (AgPO3)1-x-y(WO3)xAy, where x≦0.30 and y≦0.70;(TeO2)1-xAx, where x≦0.30;(GeO2)1-x-y(PbO)xAy, where x≦0.60 and y≦0.40;(Ga2O3)1-xAx, where x≦0.25;(Sb2O3)1-xAx, where x≦0.25;(SiO2)1-xAx, where x≦0.40;(NaPO3)1-x-y(PbF2)xAy, where x≦0.30 and y≦0.50;(NaPO3)1-x-y(WO3)xAy, where x≦0.30 and y≦0.50;(NaPO3)1-x-y(Nb2O5)xAy, where x≦0.30 and y≦0.50; anda combination thereof,where A is an additive. 4. The optogenetic probe according to claim 3, wherein A comprises an additive selected from the group consisting of a transition metal oxide, an alkali metal oxide, a pnictogen metal oxide, an oxide of the elements of group 13 of the periodic table, an oxide of the elements of group 14 of the periodic table, a halogen compound, dispersed nanoparticles, and a combination thereof. 5. The optogenetic probe according to claim 4, wherein: the transition metal oxide is selected from the group consisting of ZnO, WO3, Nb2O5, Ti3O5, Zr3O5, Ag2O, AgNO3, V2O5, Cr2O3, Cu2O, and a combination thereof;the alkali metal oxide is selected from the group consisting of Li2O, Na2O, K2O, NaPO3, and a combination thereof;the pnictogen metal oxide is selected from the group consisting of P2O5, Sb2O3, Bi2O3, and a combination thereof;the oxide of the elements of group 13 of the periodic table is selected from the group consisting of In2O3, Ga2O3, and a combination thereof;the oxide of the elements of group 14 of the periodic table is selected from the group consisting of SiO2, GeO2, SnO2, and a combination thereof;the halogen compound is selected from the group consisting of AgI, PbF2, AgCl, and a combination thereof; an... | UNIVERSITE LAVAL | G02B006/44 A61B001/00 A61B005/0476 A61B001/015 C03B037/012 C03B037/026 C03B037/027 G02B006/42 G02B006/43 B82Y030/00 A61B005/00
US09808200 | B2 | 2017-11-07 | Medical imaging apparatus and method | 2014-05-05 | A medical device is used to image a body cavity using a plurality of axially and angularly spaced imaging sensors. Each imaging sensor generates an image that is distinct from one another due to distinct fields of vision. Each image includes an overlapping zone with commonalties that are used to extrapolate a greater calibrated image. | 1. A method comprising: imaging an interior portion of a balloon having distortable reference markings using a plurality of optical imaging signals that are received and digitized to produce an uncalibrated image in which the distortable reference markings are visible; andgenerating a calibrated image from the uncalibrated image by identifying the reference markings in the uncalibrated image and calibrating imaged distortions of the identified reference markings with known physical properties of the reference markings, wherein a correction factor is applied during generation to reduce errors in the calibrated image, wherein the balloon is of a catheter that is inserted into a body cavity. 2. The method of claim 1, further comprising: exposing the body cavity with imaging energy using the catheter, the imaging energy comprising one of white light, infrared light, sonar, ultrasonic, or LIDAR energy, and wherein the imaging sensor signals convert received imaging energy into electrical signal. 3. The method of claim 2, further comprising: using the catheter to place a stent in a portion of the body cavity corresponding to the calibrated image. 4. The method of claim 1, wherein generating the calibrated image further comprises identifying partially-overlapping regions of the plurality of images. 5. The method of claim 1, wherein the imaged distortions of the reference markings are used to extrapolate three-dimensional features of the body cavity that is in contact with the balloon. 6. The method of claim 1, wherein the balloon is pressurized within a portion of the body cavity having a first pliancy and a second pliancy, and wherein generating the calibrated image comprises calibrating imaged distortions of the reference markings over the first pliancy and imaged distortions of the reference markings over the second pliancy with known dimensions of the reference markings. 7. The method of claim 6, wherein the plurality of optical images are produced in accordance with changing balloon pressures or volumes, and wherein the imaged distortions of the reference markings over the first pliancy show a greater or lesser rate of distortional change, over the changing balloon pressures or volumes, than the imaged distortions of the reference markings over the second pliancy. 8. The method of claim 1, wherein the balloon has an expandable covering that promotes expansion of the balloon in a radial dimension and application of a uniform outward force against the body cavity. 9. A system comprising: an elongated catheter comprising: a balloon coupled to a distal portion of the elongated catheter; anda plurality of imaging sensors coupled to the catheter within the balloon, where the balloon has a plurality of distortable reference markings visible to one or more of the imaging sensors during an imaging operation of the plurality of imaging sensors; andan image processing device coupled to receive signals from the imaging sensors of the elongated catheter... | Headwater Partners II LLC | A61B008/14 A61B005/00 A61B001/00 A61B008/12 A61B008/00 A61F002/24 A61F002/958 A61B008/08 A61B017/32 A61M025/10 A61B090/00
US09791683 | B2 | 2017-10-17 | Microscope with multiple illumination channels for optogenetic stimulation and fluorescence imaging | 2016-01-08 | A miniaturized microscope provides the combined capability for simultaneous or sequential optogenetic stimulation of light sensitive ion channels with the capability for fluorescence imaging of fluorescent proteins for applications requiring simultaneous optical stimulation and monitoring of cell activity. The microscope includes a dual illumination output coupling for providing illumination to two different regions within the sample and/or two different illumination source inputs and a dual transmission band optical filter to clean two different bands of light that are separately available for optogenetic stimulation and fluorescence imaging. | 1. A miniature microscope for performing simultaneous or sequential optogenetic stimulation and fluorescence measurements on a sample having an attached optical cannula, the microscope comprising: a microscope body having a first connector for receiving at least one first optical illumination connection coupling light from at least two illumination sources having differing illumination wavelengths;a second connector of the microscope body adapted to mechanically connect to the cannula and optically align the cannula with the microscope body and having at least one optical interface for coupling light returning from the sample to the microscope and illumination from the at least two illumination sources to the sample; anda first optical filter having a dual band-pass characteristic for filtering the light coupled from the at least two illumination sources, wherein a first transmission band of the first optical filter passes light having a first wavelength corresponding to a first one of the illumination sources and the second transmission band of the first optical filter passes light having a differing second wavelength corresponding to a second one of the illumination sources. 2. The microscope of claim 1, further comprising an objective lens providing or coupled to the at least one optical interface of the second connector. 3. The microscope of claim 2, further comprising: an image sensor for receiving an image from the objective lens;a second optical filter having a dichroic characteristic optically coupled between the objective lens and the image sensor, whereby light coupled from the at least two illumination sources is directed to the at least one second optical connection, and whereby the light returning from the sample through the objective lens is directed to the image sensor. 4. The microscope of claim 1, wherein the first wavelength of the first illumination source is a wavelength for optogenetic stimulation of the sample, and wherein the second wavelength of the second illumination source is a wavelength for fluorescence imaging of the sample. 5. The microscope of claim 1, wherein the at least two illumination sources comprises a first illumination source and a second illumination source, and wherein the at least one optical interface comprises a first optical interface for coupling light returning from a first region within the sample to the microscope and coupling light from the first illumination source to the first region within the sample, and a second optical interface for coupling light from the second illumination source to a second region of the sample that is different from the first region. 6. The microscope of claim 5, wherein the first region and the second region are spacially separated areas within the sample. 7. The microscope of claim 5, wherein the second optical interface further couples light returning from the second region of the sample. 8. The microscope of claim 7, wherein the first optical filter filters... | OPTOMAK, INC. | G02B021/06 A61B005/00 A61N005/06 G02B021/08 G02B013/00 G02B021/02 F21V008/00 G02B027/14 G02B021/16 G02B021/00
US09826903 | B2 | 2017-11-28 | Multi modal method and system for transmitting information about a subject | 2015-02-27 | The present invention is directed to multimodal communications means for transmitting signals representing physiological, performance, and contextual information associated with a subject. In an exemplary embodiment, the multimodal communications means includes multiple radio subsystems (or modes) that enable connection to an external monitoring device to be acquired in a wide range of settings where a single radio mode would be ineffective. Additionally, combining the multimodal communications means of the invention with real-time data-processing allows the communications functionality to be engaged only when data determined to be relevant to the user is identified. | 1. A fitness monitoring system for monitoring physiological parameters of a subject engaged in a physical activity, comprising: a monitoring garment adapted to cover at least a portion of a subject's torso;a sensor subsystem incorporated into the garment comprising a first sensor and a second sensor, wherein the first and second sensors are configured to be proximate to the subject's chest region when the garment is worn by the subject and are configured to be responsive to changes in distance between the first and second sensors, and wherein the sensor subsystem is configured to generate and transmit a first signal based on a change in the distance between the first and second sensors and a second signal that represents a physiological parameter of the subject,wherein the first sensor comprises a first magnetometer, and the second sensor comprises a second magnetometer; anda data transmission subsystem configured to transmit the first and second signals to a data monitoring subsystem,wherein the data transmission subsystem has a first transmission mode and a second transmission mode for transmitting the first signal. 2. The fitness monitoring system of claim 1, further comprising a third magnetometer and a fourth magnetometer. 3. The fitness monitoring system of claim 2, wherein the third magnetometer is not horizontally aligned with the first magnetometer, and the fourth magnetometer is not horizontally aligned with the second magnetometer. 4. The fitness monitoring system of claim 2, wherein the third magnetometer is horizontally aligned with the first magnetometer and the fourth magnetometer is horizontally aligned with the second magnetometer. 5. The fitness monitoring system of claim 2, wherein the third magnetometer is located proximate to a base of a sternum of the subject, and the fourth magnetometer is horizontally aligned with the third magnetometer, but on a back of the subject. 6. The fitness monitoring system of claim 1, wherein the first magnetometer is located proximate to an umbilicus of the subject, and the second magnetometer is horizontally aligned with the first magnetometer, but on a back of the subject. 7. The fitness monitoring system of claim 1, further comprising a third magnetometer. 8. The fitness monitoring system of claim 1, wherein the monitoring garment comprises one of a shirt, a vest, and an accessory. 9. The fitness monitoring system of claim 1, wherein the monitoring garment further comprises an additional sensor that comprises one of an optical encoder, laser interferometry, accelerometer, gyroscope, and satellite positioning system. 10. The fitness monitoring system of claim 9, wherein the additional sensor is configured to produce a signal that represents an emotional parameter. 11. The fitness monitoring system of claim 1, wherein the monitoring garment further comprises a physiological sensor that comprises one of an EEG electrode, an ECG electrode, an EMG electrode, a SpO2 level sensor, a pulse ... | adidas AG | H04B001/00 H04B015/00 H04B001/38 H04B007/00 A61B005/00 A61B005/0205 A61B005/024 A61B005/0402 A61B005/0476 A61B005/0488 A61B005/113 G06F019/00 A61B005/01 A61B005/026 A61B005/145 A61B005/11 A61B005/1455 A61B007/00
US09795302 | B2 | 2017-10-24 | Tissue illumination system, device, and method | 2014-11-21 | Implementations of the tissue illumination systems, devices, and methods disclosed herein take advantage of the translucent nature of tissue to reveal properties by light transmission, for example, tissue type, tissue transition locations, underlying structures, and the like, that are not easily distinguished by reflected light. Illuminating a back-side of a translucent tissue permits a user to distinguish between different types of tissue, tissue transition locations, and/or structures that are difficult or impossible to discern under overhead or front-side illumination. Implementations include a light source that is positionable behind a tissue or disposable within a body cavity or duct, for example, within a heart ventricle. | 1. A method of visualizing translucent tissue structures, the method comprising, accessing a heart valve of a medical patient,positioning a portion of a body of an illumination device under an annulus of the heart valve on a ventricular side of the heart valve's leaflets, wherein the portion of the body positioned under the valve annulus includes a generally arc-shaped structure that carries a plurality of light emitters spaced apart along the arc of the structure,positioning the arc-shaped structure and the light emitters snugly around and under the valve annulus,causing light to be emitted from the light emitters and transmitted through tissue of the heart valve annulus, andviewing the light transmitted through the annulus from an atrial side of the annulus opposite the side of the annulus at which the arc-shaped structure is snugly positioned. | Edwards Lifesciences Corporation | A61B005/00 A61B090/30
US09814624 | B2 | 2017-11-14 | Device and method for delivering medicine into the tympanic cavity, with sliding assist | 2015-01-26 | A device insertable in the ear canal for substance delivery to and/or extraction from the tympanic cavity is described that includes a stationary body having a distal surface, a proximal surface; a movable body having a distal surface and a proximal surface, disposed within the stationary body and free to move relative to the stationary body in response to a pressure difference created in the ear canal; and at least one piercing element having a distal end and a proximal end, disposed in the movable body. | 1. A device for at least one of a substance delivery to and extraction from a tympanic cavity, comprising: a stationary body having a distal surface and a proximal surface;a movable body having a distal surface and a proximal surface, disposed within the stationary body and free to move relative to the stationary body;at least one piercing element, having a distal end and a proximal end, disposed within the movable body; anda means for moving the movable body relative to the stationary body;wherein the means for moving the movable body relative to the stationary body is air pressure difference at the proximal surface of the movable body and at the distal surface of the movable body. 2. The device for at least one of the substance delivery to and extraction from the tympanic cavity of claim 1, wherein the movable body further comprises: a large section of the movable body disposed in contact with the proximal surface of the movable body; anda small section of the movable body disposed in contact with the distal surface of the movable body. 3. The device for at least one of the substance delivery to and extraction from the tympanic cavity of claim 1, wherein the stationary body further comprises: a large section of the stationary body disposed in contact with the proximal surface of the stationary body; anda small section of the stationary body disposed in contact with the distal surface of the stationary body. 4. The device for at least one of the substance delivery to and extraction from the tympanic cavity of claim 3, wherein the small section of the stationary body further comprises: an outer compressible layer; andan inner rigid layer. 5. A method for at least one of a substance delivery to and extraction from a tympanic cavity, comprising: inserting the device of claim 1 into an ear canal;inducing movement of the movable body toward the tympanic membrane;penetrating the tympanic membrane with the at least one hollow piercing element of the device of claim 1; andperforming at least one of injecting a substance from and evacuating fluid into at least one chamber disposed in the movable body and connected to the at least one hollow piercing element of the device of claim 1. 6. A method for at least one of a substance delivery to and extraction from a tympanic cavity, comprising: creating a sufficiently airtight seal between the stationary body of the device in claim 1 and an ear canal;inducing flexing of the tympanic membrane toward the distal surface of the movable body of the device in claim 1;inducing movement of the movable body toward the tympanic membrane; and penetrating the tympanic membrane with the at least one piercing element of the device in claim 1. 7. The device for at least one of the substance delivery to and extraction from the tympanic cavity of claim 1, wherein the piercing element is hollow. 8. The device for at least one of the substance delivery to and extraction from the tympanic cavity of claim 7, further comprisi... | EntraTympanic, LLC | A61M037/00 A61F011/00 A61M031/00 A61B005/053 A61B005/12 A61B005/00
US09801668 | B1 | 2017-10-31 | Neurophysiological apparatus and procedures | 2015-09-15 | Neurophysiological instruments and techniques are improved through various enhancements. Stimulation of an instrument is possible while it is advancing into the spine or elsewhere, alerting the surgeon to the first sign the instrument or device (screw) may be too near a nerve. A directional probe helps surgeons determine the location of the hole in the pedicle. Electrically insulating sleeves prevent shunting into the soft tissues. According to a different improvement, the same probe to be used to stimulate different devices, such as screws and wires. Electrical impulses may be recorded from non-muscle regions of the body, including the spine and other portions of the central nervous system as opposed to just the extremities. | 1. An instrument for determining directionality during surgery, comprising: a rigid probe member having a distal end having a circumference and a longitudinal axis, said distal end circumferentially insulated and configured for advancement into a pedicle and rotatable in the pedicle, and an electrode region situated about only a portion of said circumference and extending along a portion of said longitudinal axis, said electrode region configured to transmit an electric stimulation signal radially from said longitudinal axis to a pedicle;a recording electrode; anda neurophysiology monitoring system communicatively linked to said probe and said recording electrode, and configured to record a stimulus received by the recording electrode caused by the electric stimulation signal of the probe member. 2. The instrument of claim 1, wherein said neurophysiology monitoring system determines the location of a pedicle breach by determining a threshold stimulation level at which muscles innervated by a nerve proximate said pedicle respond to said stimulation signal. 3. The instrument of claim 2, wherein said instrument indicates the location of said pedicle breach by determining the position of said electrode region about said longitudinal axis of said probe member when said stimulation threshold is lowest. 4. The instrument of claim 1, wherein a non-insulated area forms said electrode region. 5. The instrument of claim 1, wherein said probe member is insulated to prevent shunting of said stimulation signal along the length of said probe member. 6. The instrument of claim 1, wherein said probe member is electrically linked to said neurophysiology monitoring system via a connector attached to said probe member. 7. The instrument of claim 6, wherein said connector is a clip around at least a portion of said probe member. 8. The instrument of claim 1, wherein probe member further comprises a proximate end, and wherein the proximate end of the probe member and the distal end of the probe member are co-linear. 9. The instrument of claim 1, wherein the probe member comprises a handle having a circumference, and wherein the handle has a larger circumference than the distal end circumference. 10. The instrument of claim 1, wherein the electrode region has an electrode region distal end and an electrode region proximate end, and wherein the electrode region distal end is located at the probe member distal end. 11. The instrument of claim 1, wherein said electrode region is configured to transmit an electric stimulation signal perpendicularly from said longitudinal axis. 12. The instrument of claim 1, wherein said portion of said longitudinal axis has a constant diameter along the longitudinal axis. | NuVasive, Inc. | A61B017/70 A61B017/16 A61B005/00 A61B005/0488 A61B005/04 A61B017/00
US09782104 | B2 | 2017-10-10 | Systems, methods and devices for acquiring and processing physiological signals | 2015-03-26 | A system, method and device are described for biometric analysis using a wearable device. The device can be removeably securable to a user's skin surface and can include a plurality of electrodes positioned to acquire electrical signals from the skin surface. The device can also include an electromyography acquisition module and skin impedance acquisition module. A processing unit can use the electrodes to capture a plurality of electrical signals including at least one electromyography signal. The processing unit can generate calibration data based on a subset of the captured electrical signals and process the electromyography signals using the calibration data to determine a biometric for the user. | 1. A biometric analysis system comprising: a remote processing device; anda wearable device, the wearable device removeably securable to a user skin surface with a skin-facing surface of the wearable device positioned to make contact with the skin surface, the wearable device having:a plurality of electrodes positioned to acquire an electrical signal from the skin surface;a controller;a signal acquisition unit operatively coupled to the controller and to the plurality of electrodes, the signal acquisition unit having a skin impedance acquisition module and an electromyography acquisition module; anda wireless communication module operatively coupled to the controller to communicate with the remote processing device; andat least one of the controller and the remote processing device is configured to: cause the injection of at least one precursor signal into at least one of the plurality of electrodes;capture, via the signal acquisition unit and the plurality of electrodes, a plurality of electrical signals comprising at least one electromyography signal and at least one impedance signal based on the at least one precursor signal;generate calibration data based on the at least one impedance signal; andprocess the at least one electromyography signal using the calibration data to determine at least one biometric. 2. The system of claim 1, wherein: the at least one impedance signal comprises at least one skin impedance signal; andthe subset of the plurality of electrical signals used to generate the calibration data comprises the at least one skin impedance signal. 3. The system of claim 1, wherein the at least one biometric comprises at least one of muscle intensity, muscle coordination, muscle ratio, and fatigue. 4. The system of claim 1, wherein the wearable device further comprises at least one of an inertial measurement unit and a GNSS unit. 5. The system of claim 1, wherein each electrode is a passive electrode. 6. The system of claim 5, wherein the signal acquisition unit further comprises a bio-impedance acquisition module. 7. The system of claim 6, wherein the signal acquisition unit further comprises: a primary multiplexer module; andthe signal acquisition unit is coupled to the plurality of electrodes using the primary multiplexer module. 8. The system of claim 7, wherein the signal acquisition unit further comprises: the primary multiplexer module coupled to the plurality of electrodes; anda secondary multiplexer module coupled to the primary multiplexer module;wherein each of the skin impedance acquisition module and the bio-impedance acquisition module is coupled to the primary multiplexer module by the secondary multiplexer module; andthe electromyography acquisition module is coupled directly to the primary multiplexer module. 9. The system of claim 5, wherein the wearable device further comprises: an analog to digital converter coupled to the controller; andeach of the skin impedance acquisition module, the bio-impedance acqu... | GestureLogic Inc. | A61B005/0488 A61B005/053 A61B005/00 A61B005/11
US09805245 | B2 | 2017-10-31 | Image resolution recognition device, method for recognizing image resolution and image resolution recognition program | 2013-10-31 | To accurately show the size of a standard resolution on an image of a fingerprint. A device comprises: a starting point setting unit (12) which, according to a starting point specification instruction that is input on a scale included in a fingerprint image, sets a starting point on a pilot wire that is specified first by the starting point specification instruction out of pilot wires of divisions on the scale; a passage point setting unit (13) which, every time an input line that is drawn along the scale from the starting point passes a pilot wire, sets a passage point on the passed pilot wire; an inter-division distance calculation unit (14) which calculates an inter-division distance on the image using the position of the starting point and the position of the passage point; a standard resolution calculation unit (15) which calculates the size of a standard resolution that corresponds to the image using the inter-division distance and a unit length which is the length of one division in the scale; and a frame display unit (16) which displays a frame on the image so as to match the size in the standard resolution. | 1. An image resolution recognition device comprising: a starting point setting unit that, according to a starting point specification instruction that is input on a scale included in a photographed image of a fingerprint displayed on a display screen, sets a starting point on a pilot wire designated first by the starting point specification instruction from among the pilot wires to display divisions of the scale;a passing point setting unit that, every time an input line drawn along the scale from the starting point set by the starting point setting unit passes a pilot wire, sets a passing point on the passed pilot wire;an inter-division distance calculation unit that calculates inter-division distance on the image using a position of the starting point and positions of the passing points;a standard resolution calculation unit that calculates size of standard resolution corresponding to the image using the inter-division distance calculated by the inter-division distance calculation unit and unit length that is length of one division of the scale; anda frame display unit that displays a frame on the image in accordance with the size of the standard resolution calculated by the standard resolution calculation unit. 2. The image resolution recognition device according to claim 1, wherein the frame display unit displays a frame on the image according to a display position specification instruction that is input on the image, and changes size of the frame displayed on the image in accordance with the size of the standard resolution calculated by the standard resolution calculation unit. 3. The image resolution recognition device according to claim 2, wherein the display position specification instruction is an instruction to designate a range of an outline of the fingerprint. 4. The image resolution recognition device according to claim 2, wherein the display position specification instruction is an instruction to designate a center position of the frame to be displayed on the image. 5. The image resolution recognition device according to claim 1, wherein the inter-division distance calculation unit calculates the inter-division distance after correcting an inclination of the input line on the basis of an angle of the input line intersecting with each of pilot wires at a time of calculating the inter-division distance. 6. The image resolution recognition device according to claim 1, further comprising an input reception unit that receives input of the unit length,wherein the standard resolution calculation unit calculates the size of the standard resolution using the inter-division distance calculated by the inter-division distance calculation unit and the unit length received by the input reception unit. 7. The image resolution recognition device according to claim 1, wherein the starting point setting unit sets the starting point at a center position of a pilot wire;the passing point setting unit sets the passing points at center positions of... | NEC Corporation | G06K009/00 A61B005/1172
US09818177 | B2 | 2017-11-14 | Image processing apparatus, image processing method, and computer-readable recording medium | 2015-09-08 | The image processing apparatus is configured by including a palm luminance estimation unit and a surface reflected component enhancement unit that separate low-frequency components from a signal indicating luminance values of pixels of an input image, a denoising processing unit that executes a denoising process for the signal from which the low-frequency components have been separated, and a merging unit that generates an image by merging the signal after the denoising process and the low-frequency components. | 1. An image processing apparatus, comprising: a computer having hardware and units programmed to perform process functions, the units including,a separation unit that separates low-frequency components from a signal indicating luminance values of pixels of an input image;a demosaicking processing unit that executes a demosaicking process for the signal from which the low-frequency components have been separated; anda merging unit that generates an image by merging the signal after the demosaicking process and the low-frequency components; whereina noise level estimation unit estimates a noise level for the signal from which the low-frequency components have been separated;a correction unit corrects luminance values of pixels of the image generated by the merging unit so that unevenness of the luminance values of pixels of the image generated by the merging unit can be removed, andthe demosaicking processing unit executes the demosaicking process by using a result of an estimation of the noise level. 2. An image processing method, comprising: separating, by a computer, low-frequency components from a signal indicating luminance values of pixels of an input image;executing, by the computer, a demosaicking process for the signal from which the low-frequency components have been separated; andgenerating, by the computer, an image by merging the signal after the demosaicking process and the low-frequency components;the method further comprising:estimating, by the computer, a noise level for the signal from which the low-frequency components have been separated;executing, by the computer, a demosaicking process by using a result of an estimation of the noise level; andcorrecting, by the computer, luminance values of pixels of the generated image so that unevenness of the luminance values of pixels of the generated image can be removed. 3. A non-transitory computer-readable recording medium having stored there in a program for causing a computer to execute a digital signature process comprising: separating low-frequency components from a signal indicating luminance values of pixels of an input image;executing a demosaicking process for the signal from which the low-frequency components have been separated; andgenerating an image by merging the signal after the demosaicking process and the low-frequency components,the digital signature process further comprising:estimating a noise level for the signal from which the low-frequency components has been separated, and executing a demosaicking process by using a result of an estimation of the noise level; andcorrecting luminance values of pixels of the generated image so that unevenness of the luminance values of pixels of the generated image can be removed. | Fujitsu Frontech Limited | G06T005/00 G06T003/40 A61B005/1171 A61B005/117 G06K009/00 G06K009/46 G06T011/00 G06T011/60 A61B005/00
US09827368 | B2 | 2017-11-28 | Liquid injector with appropriate operating conditions set by selecting displayed image | 2014-03-25 | A contrast liquid injector has a controller that calculates a necessary volume of the contrast medium based on a read out necessary dose of effective component per unit weight, weight of the subject, and concentration of the effective component. The injection rate of a base-operation condition is changed based on the calculated necessary volume of the contrast medium, wherein the predetermined injection time is unchanged, for all volumes of the contrast medium, when making the injection pattern, whereby a timing of an optimum concentration of the contrast medium will be substantially unchanged for all volumes of the contrast medium. | 1. A liquid injector for injecting at least a contrast medium into a subject whose fluoroscopic image is to be captured by an imaging diagnostic apparatus, said liquid injector, comprising: a liquid injection mechanism for injecting at least a contrast medium into a subject;control means for controlling operation of said liquid injection mechanism under an operating condition;said control means is configured:(i) to read out a base-operation condition including data of a predetermined injection time for injection, a necessary dose of effective component of the contrast medium per unit weight of the subject, a concentration of the effective component, and an injection rate;(ii) to calculate a necessary volume of the contrast medium based on the read out necessary dose of effective component per unit weight, weight of the subject, and concentration of the effective component;(iii) to make an injection pattern based on the calculated necessary volume of the contrast medium and the base-operation condition, andwherein the injection rate of the base-operation condition is changed based on the calculated necessary volume of the contrast medium, and wherein the predetermined injection time is unchanged, for all volumes of the contrast medium, when making the injection pattern, whereby a timing of an optimum concentration of the contrast medium will be substantially unchanged for all volumes of the contrast medium by said liquid injection mechanism. 2. The liquid injector of claim 1, wherein the control means is configured to display icons for use in selecting body portions. 3. The liquid injector of claim 2, wherein the control means is further configured to accept a selection of one of the icons on said display. 4. A method of preparing an injection pattern, by using a computer, for injecting at least contrast medium into a subject, the method comprising steps of: (i) reading out a base-operation condition including data of a predetermined injection time for injection, a necessary dose of effective component of the contrast medium per unit weight of the subject, a concentration of the effective component, and an injection rate;(ii) calculating a necessary volume of the contrast medium based on the read out necessary dose of effective component per unit weight, weight of the subject, and concentration of the effective component;(iii) making an injection pattern based on the calculated necessary volume of the contrast medium and the base-operation condition, wherein the injection rate of the base-operation condition is changed based on the calculated necessary volume of the contrast medium, and wherein the predetermined injection time is unchanged, for all volumes of the contrast medium, whereby a timing of an optimum concentration of the contrast medium will be substantially unchanged for all volumes of the contrast medium,(iv) injecting the necessary volume of the contrast medium into a subject, according to the injection pattern. 5. The method of c... | Nemoto Kyorindo Co., Ltd. | A61M005/00 G06F019/00 A61B005/055 A61B006/03 A61B006/00 A61B008/08 A61M005/31 A61M005/145 A61M005/142
US09782561 | B2 | 2017-10-10 | Catheter tip | 2015-09-22 | Catheters and related methods for supporting a guidewire or delivering an agent through a vessel stenosis or other tortuous anatomy are disclosed. A catheter can comprise an elongate shaft body and a polymer tip member disposed at a distal end of the shaft body. The shaft body and the tip member can include a liner, one or both of a braid member or a coil member surrounding the liner, and a polymer cover surrounding the braid member or the coil member. The tip member can further include one or more filaments extending from a position overlapping, underlapping, or abutting a distal end of the braid member or the coil member on their proximal ends to a position distal to the braid member or the coil member on their distal ends. The one or more filaments can include a plurality of filaments arranged in a series of contacting helical windings about the liner, the braid member, or the coil member. | 1. A catheter, comprising: an elongate shaft body extending from a proximal end to a distal end and defining an inner lumen; and a polymer tip member disposed at the distal end of the shaft body, the shaft body and the tip member including a liner, both a braid member and a coil member surrounding the liner, and a polymer cover surrounding the braid member and the coil member, the tip member further including one or more filaments extending from a position overlapping a distal end of the braid member and abutting a distal end of the coil member on proximal ends of the filaments to a position distal to the braid member and the coil member on distal ends of the filaments. 2. The catheter of claim 1, wherein the one or more filaments include a plurality of polymer filaments. 3. The catheter of claim 2, wherein the plurality of polymer filaments includes at least 5 polymer filaments. 4. The catheter of claim 2, wherein the plurality of polymer filaments includes at least 10 polymer filaments. 5. The catheter of claim 2, wherein the plurality of polymer filaments is helically-wound around the braid member and abutting the coil member on the proximal ends of the filaments and helically-wound around the liner on the distal ends of the filaments. 6. The catheter of claim 5, wherein the polymer filaments are positioned in a contacting, side-by-side arrangement. 7. The catheter of claim 1, wherein the coil member includes a multi-layer coil having a first coil layer wound in a first direction and a second coil layer, surrounding the first coil layer, wound in a second direction opposing the first direction. 8. The catheter of claim 7, wherein the first and second coil layers each include a plurality of helically-wound wires having a fully round cross-section. 9. The catheter of claim 1, wherein the braid member is positioned between the liner and the coil member. 10. The catheter of claim 1, wherein the distal end of each of the braid member and the coil member extend beyond the distal end of the shaft body and into the Up member. 11. The catheter of claim 10, wherein the tip member includes a non-tapered proximal portion and a tapered distal portion, the liner extending into the non-tapered proximal portion and the tapered distal portion, the braid member and the coil member extending into the non-tapered proximal portion but not the tapered distal portion. 12. The catheter of claim 1, wherein a portion of the shaft body includes a first polymer cover and the tip member includes a second polymer cover different from the first polymer cover. 13. The catheter of claim 1, wherein the polymer cover extends inward through voids between successive windings of the coil or voids of the braid member. 14. The catheter of claim 1, further comprising one or more helical threads extending around an outer surface portion of the polymer cover. 15. The catheter of claim 14, wherein the one or more helical threads are positioned around a distal end portion o... | Vacular Solutions, Inc. | A61M025/00 A61M005/00 A61M025/09 A61B005/00
US09795321 | B2 | 2017-10-24 | Bed-leaving sensor and bed-leaving detection method | 2015-09-23 | A bed-leaving sensor for detecting bed-leaving by a user on a bed, including: a sitting position detection member that detects a sitting position of the user based on a detection value of a pressure sensor arranged on the bed; and a bed-leaving behavior detection member that detects bed-leaving behavior of the user with at least one of the following as a condition: (i) the sitting position being detected within a preset bed-leaving expectation region, (ii) a movement volume of a sitting position center of gravity which is a center of gravity of the sitting position exceeding a given threshold value within a given time, and (iii) the sitting position center of gravity moving toward a preset bed-leaving range. | 1. A bed-leaving sensor for detecting bed-leaving by a user on a bed, comprising: a sitting position detection member that detects a sitting position of the user based on a detection value of a pressure sensor arranged on the bed; anda bed-leaving behavior detection member that detects bed-leaving behavior of the user with the following as a condition: (i) the sitting position being detected within a preset bed-leaving expectation region, together with (ii) a movement volume of a sitting position center of gravity which is a center of gravity of the sitting position exceeding a given threshold value within a given time, and/or (iii) the sitting position center of gravity moving toward a preset bed-leaving range,wherein the sitting position detection member sets a determination area in a periphery of a pressure sensing center of a region in which pressure of a given contact threshold value or greater is detected within a preset sitting position determination region on the pressure sensor, and the sitting position is detected when a breadth of a region in which the pressure of the contact threshold value or greater is detected within the determination area is a designated value or greater. 2. The bed-leaving sensor according to claim 1, wherein the bed-leaving expectation region is set to be a foot side region on the bed. 3. The bed-leaving sensor according to claim 1, wherein the pressure sensor is arranged only at a bed-leaving possibility region on the bed. 4. The bed-leaving sensor according to claim 1, wherein the sitting position detection member detects the sitting position further if a ratio of a total pressure detected within the determination area occupying a total pressure detected within the sitting position determination region is a designated value or greater, and if the total pressure detected within the sitting position determination region is a designated value or greater. 5. The bed-leaving sensor according to claim 4, wherein the pressure sensing center is a center of gravity of the region in which the pressure of the contact threshold value or greater is detected with the pressure sensor. 6. The bed-leaving sensor according to claim 4, wherein the pressure sensing center is an area center of the region in which the pressure of the contact threshold value or greater is detected with the pressure sensor. 7. The bed-leaving sensor according to claim 1, further comprising a notification member for notifying that the bed-leaving behavior of the user has been detected. 8. A bed-leaving detection method for detecting bed-leaving by a user on a bed, comprising: a sitting position detection step for detecting a sitting position of the user based on a detection value of a pressure sensor arranged on the bed; anda bed-leaving behavior detection step for detecting bed-leaving behavior of the user with the following as a condition: (i) the sitting position being detected within a preset bed-leaving expectation region, together with (i... | SUMITOMO RIKO COMPANY LIMITED | G08B021/22 A61B005/11 G01L001/20 G08B029/18 A61B005/103
US09801560 | B2 | 2017-10-31 | Ophthalmic lens with a neural frequency detection system | 2013-08-27 | The present invention relates generally to an ophthalmic device capable of monitoring neural frequencies and correlating the measured frequencies them to specific brain activity/functions. In some embodiments, profiles specific to the user of the ophthalmic device can be pre-programmed to tailor a brain activity/function profiles according to a user. Based on the determined brain activity/function from the correlation, a signal may be generated to provide feedback to the user. The signal may be transmitted to the user in one or more form. For example, the signal may be outputted to a wireless device in wireless communication with the ophthalmic device, and/or through an audible signal projected by an acoustic element, and/or a visual signal projected using a photon emitter, both which may be included in the ophthalmic device. | 1. An ophthalmic device with neural frequency detection comprising: a media insert comprising a front and back arcuate surface forming a cavity therebetween that contains an energy source, a controller, a transmitter, and a neural frequency detection system,wherein the energy source, the controller, the transmitter, and the neural frequency detection system are located outside of an optical zone of the ophthalmic device,wherein the energy source energizes the controller,wherein the controller comprises a processor in digital communication with a digital media storage device,wherein the neural frequency detection system comprises one or more sensors capable of measuring neural frequencies of a wearer of the ophthalmic device;wherein the transmitter transmits the measured neural frequencies to the processor, andwherein the processor receives said measured neural frequencies and detects a brain function based on pre-recorded data stored in the digital media storage device. 2. The ophthalmic device of claim 1, additionally comprising: a radio frequency antenna capable of transmitting data from and to an external wireless device. 3. The ophthalmic device of claim 2, wherein the processor sends a signal to the external wireless device according to the detected brain function. 4. The ophthalmic device of claim 2, additionally comprising: a photon emitter capable of providing a visual signal to the wearer of the ophthalmic device. 5. The ophthalmic device of claim 4, wherein the processor receives a signal based on the brain function of a user of said external wireless device through the radio frequency antenna and provides said visual signal to the wearer of the ophthalmic device through the photon emitter. 6. The ophthalmic device of claim 5, wherein the visual signal is provided to the wearer of the ophthalmic device when the signal from the brain function of said user of said external wireless device is equivalent to the signal from the brain function of the wearer of the ophthalmic device receiving the visual signal. 7. The ophthalmic device of claim 6, wherein the external wireless device is another ophthalmic device being worn by said user. 8. The ophthalmic device of claim 2, additionally comprising: an electromechanical transducer capable of emitting an audible signal and providing said audible signal to the wearer of the ophthalmic device. 9. The ophthalmic device of claim 8, wherein the processor receives a signal based on the brain function of a user of said external wireless device through the radio frequency antenna and provides an audible signal to the wearer of the ophthalmic device through the electromechanical transducer. 10. The ophthalmic device of claim 9, wherein the audible signal is provided to the wearer of the ophthalmic device when the signal from the brain function of said user of said external wireless device is equivalent to the signal to the brain function of the wearer of the ophthalmic device receiving the audibl... | Johnson & Johnson Vision Care, Inc. | A61B005/04 A61B005/0478 A61B005/00 G02C007/04 G02C007/08
US09782125 | B2 | 2017-10-10 | Athletic or other performance sensing systems | 2014-09-19 | Articles of footwear and other devices include modules, e.g., for sensing physical and/or physiological characteristics associated with use of the footwear or other devices. Such devices may include an authentication or activation system for providing power or otherwise determining the connection of a sensor in the article of footwear. | 1. An article of footwear, comprising: an upper member at least partially defining a foot-receiving chamber;a sole structure engaged with the upper member;a housing for releasably receiving an electronic module, the housing engaged with at least one of the upper member or the sole structure, wherein the housing defines an electronic module receiving chamber formed in or on at least one of the upper member or sole structure;a system for releasably securing an electronic module with the electronic module receiving chamber; anda release mechanism for releasing an object mounted in the electronic module receiving chamber, wherein the release mechanism includes a spring member,wherein the spring member at least partially pushes the object out of the electronic module receiving chamber during a release operation. 2. An article of footwear according to claim 1, wherein the housing and the electronic module are asymmetrical in shape. 3. An article of footwear according to claim 1, wherein the spring member biases at least a portion of the system to a securing position. 4. An article of footwear according to claim 1, further comprising: a tool for activating the release mechanism. 5. An article of footwear according to claim 4, wherein the tool is engaged with the article of footwear. 6. An article of footwear according to claim 4, wherein the tool is engaged with the upper member. 7. An article of footwear according to claim 4, wherein the tool is engaged with the sole structure. 8. An article of footwear according to claim 4, wherein the tool includes an aglet for a shoe lace engaged with the article of footwear. 9. An article of footwear according to claim 1, wherein the housing is located in an impact-attenuating element of the article of footwear such that the electronic module receiving chamber faces the foot-receiving chamber. 10. An article of footwear according to claim 9, wherein the impact-attenuating element includes an opening defined therein that allows access to the release mechanism. 11. An article of footwear according to claim 9, further comprising: an insole member or a sock liner at least partially defining the foot-receiving chamber. 12. An article of footwear according to claim 11, wherein the insole member or sock liner includes an opening defined therein that allows access to the release mechanism. 13. An article of footwear according to claim 12, wherein the impact-attenuating element includes an opening defined therein that at least partially aligns with the opening in the insole member or sock liner to allow access to the release mechanism. 14. An article of footwear according to claim 13, further comprising: a tool for activating the release mechanism. 15. An article of footwear according to claim 14, wherein the tool includes an aglet for a shoe lace engaged with the article of footwear. 16. An article of footwear according to claim 1, further comprising: an electronic module releasably engaged at least partiall... | NIKE, Inc. | A43B003/00 A61B005/00 A43B013/00 G01C022/00 A61B005/11
US09819843 | B2 | 2017-11-14 | Head-mounted systems and methods for providing inspection, evaluation or assessment of an event or location | 2013-09-20 | Systems and methods for providing assessment of a local scene to a remote location are provided herein. The systems and methods facilitate collection, storage and transmission of data from the local scene to the remote location without limiting the mobility, dexterity, adaptability and interactive capability of an on-scene technician. For example, a head-mounted device according to an implementation discussed herein can include: a head-mounted frame that is configured to hold a transparent visor; an image capturing device attachable to at least one of the head-mounted frame and the transparent visor; and a micro-optic display system attachable to at least one of the head-mounted frame and the transparent visor. The micro-optic display system can be configured to render a heads up image on the transparent visor. In addition, the heads up image can define an outline of a field of view of the image capturing device. | 1. A head-mounted device, comprising: a head-mounted frame that is configured to hold a transparent visor;an image capturing device attachable to at least one of the head-mounted frame and the transparent visor;an interface configured to receive medical information, the interface having a wireless receiver configured to receive vital sign data for at least one patient from one or more medical instruments over a wireless communication link; anda micro-optic display system attachable to at least one of the head-mounted frame and the transparent visor, the micro-optic display system being configured to render a heads up image and the vital sign data on the transparent visor, wherein the heads up image comprises an outline of a field of view of the image capturing device and an outline of a field of view corresponding to an image as viewed at a remote work station, wherein the outline of the field of view corresponding to the image as viewed at the remote work station is dynamically displayed as a subset of the field of view of the image capturing device, the subset corresponding to a level of zoom on the image as viewed at the remote work station. 2. The head-mounted device of claim 1, wherein the outline of the field of view frames a scene captured by the image capturing device. 3. The head-mounted device of claim 1, further comprising an optical zoom control operable to control the level of zoom on the image as viewed at the remote work station, and resize the outline of the field of view corresponding to the image as viewed at the remote work station on the heads up image. 4. The head-mounted device of claim 1, wherein the micro-optic display system comprises: an image generator configured to project the heads up image on the transparent visor;a processor; anda memory communicatively connected to the processor having computer-executable instructions stored thereon that, when executed by the processor, cause the processor to: receive operating characteristics of the image capturing device; determine a field of view based on the operating characteristics; and determine dimensions of the outline of the field of view. 5. The head-mounted device of claim 4, wherein the operating characteristics of the image capturing device comprise at least one of an effective focal length and an amount of digital zoom. 6. The head-mounted device of claim 4 wherein the processor is operable to execute one or more instructions to receive the vital sign data for at least one patient from the one or more medical instruments and communicate the vital sign data to the micro-optic display system. 7. The head-mounted device of claim 6, wherein the wireless receiver comprises a Bluetooth receiver. 8. The head-mounted device of claim 1, wherein the heads up image further comprises a graphical information field. 9. The head-mounted device of claim 8, wherein the heads up image comprises a first region defined by the outline of the field of view and a second region incl... | Zeriscope Inc. | H04N005/222 G09G005/00 H04N005/225 G02B027/01 H04N005/232 G06F003/01 G06F003/0338 G06T011/00 G06F003/0362 G06T007/00 A61B005/00 G06F019/00
US09788722 | B2 | 2017-10-17 | Method for addressing medical conditions through a wearable health monitor with the aid of a digital computer | 2016-11-28 | Individuals who suffer from certain kinds of medical conditions, particularly conditions that only sporadically exhibit measurable symptoms, can feel helpless in their attempts to secure access to medical care because, at least in part, they are left to the mercy of their condition to present symptoms at the right time to allow diagnosis and treatment. Providing these individuals with ambulatory extended-wear health monitors that record ECG and physiology, preferably available over-the-counter and without health insurance preauthorization, is a first step towards addressing their needs. In addition, these individuals need a way to gain entry into the health care system once a medically-actionable medical condition has been identified. Here, the ECG and physiology is downloaded and evaluated post-monitoring against medical diagnostic criteria. Medical specialists are pre-identified and paired up with key diagnostic findings, such that an individual whose monitoring data indicates a medical concern will be automatically referred and treated. | 1. A method for addressing medical conditions through a wearable health monitor with the aid of a digital computer, comprising the steps of: providing a patient with a wearable health monitor comprising a disposable extended wear electrode patch and a monitor recorder, the electrode patch comprising a flexible backing and a plurality of electrocardiographic electrodes comprised on the contact surface of each end of the flexible backing, a non-conductive receptacle affixed to a non-contacting surface of the flexible backing, and a pair of flexible circuit traces affixed at each end of the flexible backing with each circuit trace connecting one of the electrocardiographic electrodes to an electro mechanical docking interface, the monitor recorder comprising a sealed housing adapted to be removably secured into the non-conductive receptacle and electronic circuitry comprising a memory and comprised within the sealed housing including an external interface configured to be removably connected to the electrocardiographic electrodes via the electro mechanical docking interface;sensing the patient's physiology through the disposable extended wear electrode patch over a monitoring period and recording the physiology throughout the monitoring period into the memory onboard the wearable health monitor;downloading the physiology into a computer upon completion of the monitoring period and comparing the physiology to a medical diagnostic criteria;generating with the computer a diagnostic overread that comprises one or more diagnostic findings;determining with the computer the severity of any medical condition indicated by the diagnostic findings; andreferring the patient to one or more pre-identified care providers via the computer for facilitated diagnosis as conditioned upon the severity of each such medical condition that exceeds a threshold tolerance level. 2. A method according to claim 1, wherein the physiology is selected from the group comprising heart rate, blood pressure, respiratory rate, temperature, oxygen saturation, and blood glucose. 3. A method according to claim 1, wherein the diagnostic criteria are selected from the group comprising lightheadedness, fainting, chest pain, hypoxia, syncope, palpitations, and congestive heart failure. 4. A method according to claim 1, further comprising at least one of the steps of: notifying a general practice physician upon a diagnostic finding comprising normal physiological data;notifying a cardiologist upon a diagnostic finding comprising one or more of atrial fibrillation, ectopy, palpitations, and supraventricular tachycardia; andnotifying an electrophysiologist upon a diagnostic finding comprising ventricular tachycardia, bradycardia, and heart blockage. 5. A method according to claim 1, further comprising the steps of: maintaining a database on the computer of the pre-identified care providers ordered by medical specialty and paired with one or more of the diagnostic criteria. 6. A method accor... | BARDY DIAGNOSTICS, INC. | A61B005/00 A61B005/01 A61B005/021 A61B005/0245 A61B005/0402 A61B005/046 A61B005/0464 A61B005/0468 A61B005/08 A61B005/145 A61B005/1477
US09775396 | B1 | 2017-10-03 | Data collecting head guard systems and methods thereof | 2016-04-13 | A head guard is provided. The head guard includes one or more sensors as part of an sensory input and communications system. The head guard wirelessly communicates data to remote computing devices for intelligent data collection. | 1. A data collection system, comprising a plurality of head guards, wherein each of the plurality of head guards comprises: a multi-layered sidewall, wherein the multi-layered sidewall forms a dome, wherein the dome defines a circular opening for a head of a wearer and a plurality of sensor ports, and wherein the multi-layered sidewall comprises: a stretchable fabric layer, anda padding layer; anda removable cover system selectably attachable to the multi-layered sidewall, wherein the removable cover system has an inner surface and an outer surface, wherein a portion of the inner surface of the removable cover system is attachable to a portion of the dome to maintain the position of the removable cover system relative to the dome; anda sensory input and communications system, wherein at least a portion of the sensory input and communications system is positioned between the cover system and the multi-layered sidewall, the sensory input and communications systems comprises: a microcontroller,a plurality of sensors that are each in communication with the microcontroller, wherein each of the plurality of sensors are removably received into a respective one of the plurality of sensors ports,a communications unit, anda local indicator, wherein the microcontroller is configured to activate the local indicator upon a detected impact in excess of a predefined magnitude, wherein the local indicator comprises any of a visual device and an audio device; andan activity monitoring computing system to wirelessly receive data collected by each of the plurality of head guards, the activity monitoring computing system comprising at least one processor and non-transitory computer readable medium having instructions stored thereon which when executed by a processor cause the processor to: for the data received from each of the plurality of head guards, determine athlete information associated with the data based on the source of the data;group data based on one or more factors selected from the group of factors comprising sport, age, and geography; andgenerate aggregated data reporting based on the grouped data. 2. The data collection system of claim 1, wherein the plurality of head guards comprises a first head guard and a second head guard, and wherein subsequent to an impact event detected by the activity monitoring computing system, the processor is to generate an impact report based on a first player associated with the first head guard and a second player associated with the second head guard. 3. The data collection system of claim 1, wherein each of the communications unit is configured to wirelessly transmit sensor data to an associated mobile communications device. 4. The data collection system of claim 3, wherein the activity monitoring computing system is to wirelessly receive sensor data from the associated mobile communications device. 5. The data collection system of claim 3, wherein the plurality of head guards are each associated with the mobile... | 2nd Skull, Inc. | G08B019/00 A42B003/04 H04B001/3827 H04L029/08 H04Q009/00 G08B021/04 A42B003/12 A63B024/00 A61B005/00 A61B005/024 A61B005/0402 A61B005/08 A61B005/145
US09801658 | B2 | 2017-10-31 | Removal device | 2015-03-27 | Method and system for treating a patient using an inflatable device. A removal device may be used to remove the inflatable device from the body. The removal device can have at least one manually-actuatable member and at least two opposing jaws at a distal end. At least one of the at least two jaws can be movable by actuation of the at least one manually-actuatable member. At least one of the at least two jaws can have a puncturing member. | 1. A removal device comprising: (a) at least one manually-actuatable member; and(b) at least two opposing jaws, at least one of the at least two opposing jaws being movable by actuation of the at least one manually-actuatable member, wherein at least one of the at least two opposing jaws comprises a gripping member and wherein at least one of the at least two opposing jaws comprises a puncturing member, wherein the puncturing member comprises a cannulated needle having a longitudinal axis and a lumen having a longitudinal axis, wherein the longitudinal axis of the lumen is aligned with the longitudinal axis of the cannulated needle. 2. The removal device of claim 1, wherein the jaw opposing the jaw with the puncturing member includes an opening configured to receive the cannulated needle when the at least two opposing jaws are in a closed position. 3. The removal device of claim 1, wherein each of the at least two opposing jaws comprises a puncturing member and an opening, the opening of each of the at least two jaws opposing being adapted to receive the puncturing member on the opposing jaw when the at least two opposing jaws are in a closed position. 4. The removal device of claim 3, wherein each puncturing member comprises a cannulated needle. 5. The removal device of claim 1, wherein the gripping member comprises a row of teeth. 6. The removal device of claim 1, wherein both of the at least two opposing jaws include gripping members. 7. The removal device of claim 6, wherein the gripping members comprise a row of teeth. 8. The removal device of claim 7, wherein in the closed position there is a gap between the rows of teeth of the at least two opposing jaws. 9. The removal device of claim 1, wherein the puncturing member is generally perpendicular to the jaw. | AttenueX Technologies, Inc. | A61B017/02 A61B017/34 A61F002/00 A61B001/00 A61B001/12 A61B001/307 A61B005/20 A61M039/22 A61M039/28 A61B017/295 A61M001/00 A61M029/02 A61B017/3205 A61F005/00 A61B017/00 A61M025/10 A61M039/00 A61B017/29 A61M025/00 A61M025/06 A61F002/04
US09802034 | B2 | 2017-10-31 | Medical equipment electrodes | 2015-10-07 | Cardiac electrodes and techniques for testing application of the electrodes to a victim are described herein. | 1. A medical device comprising: an external electrode assembly comprising a pair of electrodes and a camera included in the external electrode assembly, the camera configured to acquire image information indicative of a location of at least one of the pair of electrodes relative to external features of the chest of a patient;a processor and a memory;at least one output device; andan electronic interface, communicatively coupled to the external electrode assembly, the processor, the memory, and the at least one output device, and configured to obtain the image information from the camera and provide signals representing the image information to the processor,wherein the processor is configured to: receive the signals representing the image information,process the signals representing the image information,based on the processing of the signals representing the image information, determine electrode positioning feedback for a rescuer, andprovide the electrode positioning feedback to the at least one output device,and wherein the at least one output device is configured to provide the electrode positioning feedback as an instruction to move the at least one of the pair of electrodes. 2. The medical device of claim 1, wherein the processor is configured to determine a current location of the at least one of the pair of electrodes relative to the chest of the patient and to determine an x-translation distance and a y-translation distance between the current location of the at least one of the pair of electrodes and a desired location of the at least one of the pair of electrodes relative to the chest of the patient. 3. The medical device of claim 1, wherein the processor is configured to determine a current position of the at least one of the pair of electrodes relative to the chest of the patient based on the received image information. 4. The medical device of claim 3 wherein the electrode positioning feedback indicates the current position of the at least one of the pair of electrodes relative to the chest of the patient. 5. The medical device of claim 3, wherein the processor is configured to determine whether the current position of the at least one of the pair of electrodes corresponds to a recommended position for operation of the at least one of the pair of electrodes. 6. The medical device of claim 5 wherein the electrode positioning feedback for the rescuer indicates whether the current position of the at least one of the pair of electrodes corresponds to the recommended position for operation of the at least one of the pair of electrodes. 7. The medical device of claim 3, wherein the instruction to move the at least one of the pair of electrodes indicates a direction relative to the current position of the at least one of the pair of electrodes in which the rescuer should move the external electrode assembly. 8. The medical device of claim 7 wherein the at least one output device is a display and the electrode positioning feedback co... | ZOLL Medical Corporation | A61N001/39 A61N001/04 A61B005/06 A61B005/00
US09814807 | B2 | 2017-11-14 | Chest drainage systems and methods | 2015-03-30 | A chest drainage system including a collection device configured to receive fluid from the pleural cavity of a patient. A sensor is included to detect a pressure differential in the fluid. A display is configured to display a trend in occurrences of changes in pressure of the fluid over time in predetermined time increments based on a number of detections of pressure differentials that exceed a predetermined pressure differential during each of the predetermined time increments. The trend is correlative to the percentage of time that the patient is deemed to have an air leak in the pleural cavity in the predetermined time increments. The trend is derived from a ratio of the quantity of respiratory cycles of the patient for which the predetermined pressure differential is detected (QRCleak) in the predetermined time increments to the total quantity of respiratory cycles of the patient in respective predetermined time increments (QRCtotal). | 1. A chest drainage system comprising: a collection device configured to receive fluid from the pleural cavity of a patient via a fluid pathway extending between the patient and the collection device;a sensor for detecting a pressure differential between a first location in a proximal portion of the fluid pathway extending from the collection device toward the patient and a second location in a distal portion of the fluid pathway extending from the patient toward the collection device; anda display configured to display a trend in occurrences of changes in pressure of the fluid over time, the trend configured to correlate with the percentage of time that the patient is deemed to have an air leak from the patient's lung into the pleural cavity,wherein the trend is derived from a number of respiratory cycles of the patient for which it is detected that the pressure differential exceeds a predetermined pressure differential threshold. 2. The system of claim 1, wherein the display is configured to indicate when the collection device is disconnected from the chest drainage system. 3. The system of claim 1, wherein the display is configured to indicate a number of occurrences in which the detected fluid pressure exceeds a predetermined minimum magnitude. 4. The system of claim 1, wherein the display is configured to indicate a total amount of air leakage in a respiratory wave form of the patient. 5. The system of claim 1, wherein the display is generated from information including a duration of an air leak having at least a predetermined minimum magnitude. 6. The system of claim 1, wherein the trend is further derived from a ratio of the quantity of respiratory cycles of the patient for which it is detected that the pressure differential exceeds the predetermined pressure differential threshold during a time period, with respect to the total quantity of respiratory cycles of the patient during the time period. 7. The system of claim 1, wherein the trend is further derived from the number of respiratory cycles of the patient for which it is detected that the pressure differential exceeds the predetermined pressure differential threshold per unit time. 8. The system of claim 1, wherein the predetermined pressure differential threshold is +2 cmH2O. 9. The system of claim 1, wherein the trend is determined for each time increment in a plurality of predetermined time increments. 10. The system of claim 1, comprising a plurality of the predetermined pressure differential thresholds and a plurality of status indicators corresponding to each of the predetermined pressure differential thresholds, wherein the system is configured to activate each status indicator when that status indicator's corresponding pressure differential threshold is exceeded. 11. The system of claim 10, wherein the status indicators comprise different colored lights. | ATRIUM MEDICAL CORPORATION | A61M001/00 A61B005/08 A61M027/00 A61B090/00
US09788739 | B2 | 2017-10-17 | Management of cardiac data transmissions | 2013-04-29 | A medical alarm communications system comprises a pager like device to be kept near a patient. The pager like device communicates medical alerts regarding the patient to a remote central station, which can provide therapeutic and/or diagnostic assistance by communicating to the pager like device. When the pager like device determines that an alert should be sent to the central station, it attempts to establish communication with the central station according to a primary communication protocol. If this attempt is unsuccessful according to some predetermined criteria (e.g. too much time has elapsed before communication is established), then the pager like device generates a message to the patient indicative of the failure, and The attempts to establish communication with the central station according to a secondary communication protocol. | 1. A cardiac monitoring system having a patient monitoring system and external equipment located remotely, comprising: a. the patient monitoring system including a patient implanted device and an external device located near the patient; the implanted device configured to (1) record signal data from the heart, (2) to analyze the signal data to detect at least one type of cardiac event, said implanted device being configured to categorize said at least one cardiac event as an event which is selected from the group of urgent event and non-urgent event, and (3) to transmit data wirelessly to the external device;the external device configured to provide communication with the implanted device, a patient, and with the external equipment located remotely based upon the categorization transmitted by said implanted device, said external device being configured to transmit and receive data to and from the implanted medical device and further enabling patient interaction with the external device for performing event prioritization;b. the patient monitoring system being configured to provide event prioritization for categorizing the detected cardiac events into one of at least two priority levels, the external device being configured to transmit a priority level to the equipment at the remote location; and,c. the equipment at the remote location being configured to receive data communicated from the external device including the identified priority level of the detected cardiac event, the equipment further configured for operating contingently upon the priority level received, wherein operating contingently includes i) responding to a first priority level event types by substantially immediately issuing notification to a medical professional; and ii) responding to second priority level event types without providing substantially immediate notification thereof. 2. The cardiac monitoring system according to claim 1 wherein the priority level is transmitted from the external device as an event tag. 3. The cardiac monitoring system according to claim 1 wherein the priority level is transmitted from the external device as a message header. 4. The cardiac monitoring system according to claim 1 wherein the implanted device is configured to perform the event prioritization. 5. The cardiac monitoring system according to claim 1 wherein the external device is configured to perform the event prioritization. 6. The cardiac monitoring system according to claim 1 wherein the equipment at the remote location is configured to perform the event prioritization. 7. The cardiac monitoring system according to claim 1 wherein the external device is further configured for operating upon the signal data and generating at least one event tag. 8. The cardiac monitoring system according to claim 1 wherein the external device is further configured for allowing the patient to operate buttons of the external device to generate at least one event tag. 9. The cardiac monitoring sy... | Angel Medical Systems, Inc. | A61B005/0402 A61B005/04 A61B005/00 A61B005/02 G06F019/00 A61N001/08
US09801721 | B2 | 2017-10-31 | Sizing device and method of positioning a prosthetic heart valve | 2013-03-07 | A sizing device for a collapsible prosthetic heart valve includes a collapsible and expandable stent. A microelectromechanical sensor is coupled to the stent, the sensor being capable of collecting information related to the size and stiffness of tissue. | 1. A sizing device for use in implanting a collapsible prosthetic heart valve in a native valve annulus, the sizing device comprising: a collapsible and expandable stent having an annulus section at a first end with a first diameter and an aortic section at a second end with a second diameter, the second diameter being greater than the first diameter, and the first end and the second end being disposed at opposite sides of the stent; anda sensor coupled to the annulus section adjacent the first diameter of the stent, the sensor comprising at least one deflectable contacting member, and being capable of collecting information related to stiffness of the native valve annulus by contacting portions of the native valve annulus with the at least one deflectable contacting member. 2. The device of claim 1, wherein the stent is self-expandable. 3. The device of claim 2, wherein the stent comprises nitinol. 4. The device of claim 2, wherein the sensor is flexible. 5. The device of claim 1, wherein the information includes the diameter of the native valve annulus. 6. The device of claim 1, wherein the information includes data relating to the extent of calcification of tissue of the native valve annulus. 7. The device of claim 1, wherein the sensor includes at least one capacitor having variable capacitance, the capacitance corresponding to the information. 8. The device of claim 1, wherein the sensor includes at least one piezoelectric material. 9. The device of claim 1, wherein the sensor comprises a polymer. 10. The device of claim 9, wherein the polymer comprises polydimethylsiloxane. 11. The device of claim 1, wherein the sensor is a microelectromechanical sensor. 12. The device of claim 1, wherein the sensor comprises at least two electrodes mounted on a fabric. 13. The device of claim 1, further comprising a deployment device configured to expand the collapsible and expandable stent via a series of rotations. 14. A method for determining the fitment of a prosthetic heart valve within a native valve annulus, comprising: introducing a sizing device into the native valve annulus, the sizing device including (i) a collapsible and expandable stent having an annulus section at a first end with a first diameter and an aortic section at a second end with a second diameter, the second diameter being greater than the first diameter, and the first end and the second end being disposed at opposite sides of the stent and (ii) a sensor coupled to the annulus section of the stent, the sensor comprising at least one deflectable contacting member, and being capable of collecting information related to the native valve annulus by contacting portions of the native valve annulus with the at least one deflectable contacting member;expanding the diameter of the stent within the native valve annulus; andacquiring information related to the native valve annulus via the sensor by contacting portions of the native valve annulus with the at least one deflectable... | St. Jude Medical, Cardiology Division, Inc. | A61F002/24 A61B005/00 A61B005/107
US09807317 | B2 | 2017-10-31 | Multi-spectral imaging method for ultraweak photon emission and system thereof | 2012-04-23 | An ultra-weak light multispectral imaging method and an ultra-weak light multispectral imaging system, which can realize multispectral two-dimensional imaging of an ultra-weak light object by constituting a linear array from single-photon detectors of all response wavelengths and combining it with light-splitting technology. The ultra-weak light multispectral two-dimensional imaging system realizes high-resolution optical modulation by adopting the compressive sensing (CS) theory and the digital light processing (DLP) technology and using a linear array single-photon detector as a detection element; the ultra-weak light multispectral two-dimensional imaging system comprises a light filter, a first lens (1), a DMD control system, a second lens, a spectrophotometer, a linear array single-photon detector consisting of a plurality of single-photon detectors with different response wavelengths, and a central processing unit; and the sensitivity of the system can reach the single-photon level. The invention can be widely applied in the fields of biological self-illumination, medical diagnosis, nondestructive material analysis, astronomical observation, national defense and military, spectral measurement, quantum electronics and the like. | 1. A two-dimensional imaging method for high-resolution optical modulation by compressive sensing and digital light processing and using a linear array single-photon detector as a detection element, said method comprising: 1) compressive sampling by a digital micro-mirror device (DMD) control system, a first lens and a second lens, to convert two-dimensional image data into a one-dimensional data sequence to complete compressive sampling of signals to be measured, wherein the compressive sampling includesfiltering stray light by a light filter, imaging the light at the DMD control system through the first lens, controlling reflecting photons to the second lens by the DMD control system controlling focusing of the photons by the second lens; and imaging the light with different wavelengths by the linear array single-photon detector; and2) completing a sparse reconstruction by a combination of a central processing unit and the linear array single-photon detector to reconstruct the two-dimensional image data and obtain a two-dimensional photon density image, wherein i) the compressive sampling further includes mapping the signals to be measured from high-dimensional signals to low-dimensional signals where,where xεRn is data to be measured, yεRk is observation data, ΦεRk×n is a measurement matrix (k&lt;&lt;n) and e εRk is measurement noise, the compressive sample precess is y=Φx+e (1)ii) where a value of measurement times is k, a number of pixels in the two-dimensional image is n, a measurement matrix in formula (1) is Φ={Φ1, . . . , Φi, . . . , Φk wherein Φ1 is an ith row of Φ wherein columns of the two-dimensional images of size √{square root over (n)}×√{square root over (n)} are connected end to end to form an nx1 one-dimensional column vector, and corresponding to x in formula (1), each element of the nx1 one-dimensional column vector represents a photon density at a corresponding position; in the DMD control system, each micro-mirror has a same resolution and columns of the micro-mirrors of the DMD control system are connected end to end to form a 1xn one-dimensional row vector, which corresponds to a row in the measurement matrix Φ, wherein each element indicates whether the photon at a corresponding position is transmitted to the second lens;iii) where a measurement period is T and during this period the DMD control system is kept unchanged, light with certain wavelength directly projected onto the linear array single-photon detector with a corresponding wavelength; where a number of the photons detected by the linear array single-photon detector corresponding to a certain wavelength in the linear array single photon detector is N, then N/T is equivalent to an inner product value of the photon density image and measurement number array Φ on the DMD control system, corresponding to an elementyi=∑j=1n⁢⁢Φi,j⁢xj of an observation vector y in formula (1), wherein Φi,j and xj are jth elements of Φi and x respectively; according to the measurement ... | Center for Space Science and Applied Research, Chinese Academy of Sciences | H04N005/30 G01J003/28 G01J003/00 G02F001/01 G01N021/76 G01J001/42 A61B005/00 G01N021/64 G01J001/44
US09820664 | B2 | 2017-11-21 | Catheter with high density electrode spine array | 2014-11-20 | A catheter adapted or high density mapping and/or ablation of tissue surface has a distal electrode matrix having a plurality of spines arranged in parallel configuration on which a multitude of electrodes are carried in a grid formation for providing uniformity and predictability in electrode placement on the tissue surface. The matrix can be dragged against the tissue surface upon deflection (and/or release of the deflection) of the catheter. The spines generally maintain their parallel configuration and the multitude of electrodes generally maintain their predetermined relative spacing in the grid formation as the matrix is dragged across the tissue surface in providing very high density mapping signals. The spines may have free distal ends, or distal ends that are joined to form loops for maintaining the spines in parallel configuration. | 1. A catheter comprising: an elongated catheter body; anda distal electrode assembly, the distal electrode assembly comprising at least three spines extending in a common plane, each spine having at least one electrode, a proximal spine portion, a distal spine portion, and a free distal end,wherein the catheter body defines a longitudinal axis, and each proximal spine portion is configured to extend at an angle from the longitudinal axis of the catheter body and each distal spine portion is configured to extend generally parallel to the longitudinal axis of the catheter body and generally parallel to each other distal spine Portion. 2. The catheter of claim 1, further comprising an intermediate deflection section between the catheter body and the distal electrode assembly. 3. The catheter of claim 1, wherein the distal electrode assembly has an electrode density ranging between about 12 and 16 electrodes per square centimeter. 4. The catheter of claim 1, wherein the distal electrode assembly has an electrode density of about 32 electrodes per 2.2 cm2. 5. The catheter of claim 1, wherein each spine comprises a shape memory member and a nonconductive covering. 6. The catheter of claim 1, wherein the at least three spines ranges between about three and ten. 7. The catheter of claim 1, wherein the distal electrode assembly is configured to maintain the distal portions of the at least three spines in parallel arrangement upon exertion of a contact force. 8. The catheter of claim 1, wherein inner proximal spine portions extend at lesser angles from the longitudinal axis of the catheter body, and outer proximal spine portions extend at greater angles from the longitudinal axis of the catheter body. 9. A catheter comprising: an elongated catheter body; anda distal electrode assembly, the distal electrode assembly comprising a plurality of spines extending in a common plane, each spine having one or more electrodes, a proximal spine portion, a distal spine portion, and a free distal end,wherein the one or more electrodes are arranged in a predetermined pattern within the distal electrode assembly;wherein the catheter body defines a longitudinal axis, each proximal spine portion is configured to extend at an angle from the longitudinal axis of the catheter body, and each distal spine portion is configured to extend generally parallel to the longitudinal axis of the catheter body, andwherein inner proximal spine portions extend at lesser angles from the longitudinal axis of the catheter body, and outer proximal spine portions extend at greater angles from the longitudinal axis of the catheter body. 10. The catheter of claim 9, further comprising an intermediate deflection section between the catheter body and the distal electrode assembly. 11. The catheter of claim 9, wherein the distal electrode assembly has an electrode density ranging between about 12 and 16 electrodes per square centimeter. 12. The catheter of claim 9, wherein the distal elec... | Biosense Webster (Israel) Ltd. | A61B005/042 A61B018/14 A61B005/00 A61B018/00
US09788764 | B2 | 2017-10-17 | Method and system for non-invasive optical blood glucose detection utilizing spectral data analysis | 2012-09-11 | Systems and methods are disclosed for non-invasively measuring blood glucose levels in a biological sample based on spectral data. This includes at least one light source configured to strike a target area of a sample, at least one light detector positioned to receive light from the at least one light source and to generate an output signal, having a time dependent current, which is indicative of the power of light detected, and a processor configured to receive the output signal from the at least one light detector and based on the received output signal, calculate the attenuance attributable to blood in a sample present in the target area and eliminate effect of uncertainty caused by temperature dependent detector response of the at least one light detector, and based on the calculated attenuance, determine a blood glucose level associated with a sample present in the target area. | 1. A system for detecting glucose in a biological sample, comprising: at least one light beam generating structure generating one or more light beams having a wavelength in a range between 800 nm and 1600 nm to strike a target area of a biological sample;at least one photocurrent signal generating light detector containing one or more photodiodes positioned to receive detected blood glucose level information in the one or more light beams and to transform the detected blood glucose level information in the one or more light beams to the detected blood glucose level information in an output photocurrent signal, having a time dependent current, which is indicative of the power of light detected; anda blood glucose level caused light absorbance change determining algorithm implemented processor programmed to calculate a change in a light absorption caused by blood in the biological sample and configured to receive the output photocurrent signal from the at least one light detector and based on the received output photocurrent signal, calculate attenuance attributable to blood in a sample present in the target area and eliminate effect of uncertainty caused by temperature dependent detector response of the at least one light detector by calculating standard deviation of a logarithm of the time dependent output current generated by the light power from the same target area of the biological sample, and based on the calculated attenuance, determine a blood glucose level associated with a sample present in the target area. 2. The system for detecting glucose in a biological sample according to claim 1, wherein the processor is configured to eliminate the effect of uncertainty caused by temperature of the at least one light detector by calculating the following equation with the processor: σ[log ID(λi,t)]=σ[log P(λi,t], where the ith wavelength being represented by λi, based on the time dependent detector output current signal ID(λi,t) and the time dependent detector power signal P(λi,t), where log ID(λi,t) is the logarithm of the time dependent detector output current signal and log P(λi,t) is the logarithm of the time dependent light power received by the corresponding detector. 3. The system for detecting glucose in a biological sample according to claim 1, wherein the at least one light detector includes a preamplifier having a feedback resistor. 4. A method for detecting glucose in a biological sample, comprising: utilizing at least one light beam generating structure to generate one or more light beams having a wavelength in a range between 800 nm and 1600 nm to strike a target area of a biological sample;utilizing at least one photocurrent signal generating light detector positioned to receive detected blood glucose level information in the one or more light beams and to transform the detected blood glucose level information in the one or more light beams to the detected blood glucose level information in an output photocurrent signal, having a... | St. Louis Medical Devices, Inc. | A61B005/145 A61B005/1455 A61B005/00
US09808169 | B2 | 2017-11-07 | System for electroencephalographic detection of an inadequacy between the state of a patient placed under ventilatory support and the setting of the machine used for this support, and use of this detection to adjust the setting | 2011-07-06 | The invention relates to a system for electroencephalographic detection of an improper adjustment of a ventilator used on a mammal. The system includes an electroencephalograph, capable of measuring, as a function of time, an electroencephalographic signal representative of a breathing process; and an input for receiving a respiratory initiation signal, different from the electroencephalographic signal, capable of indicating a respiratory initiation time. The detection system further includes means for specifying a beta frequency band comprised between 15 and 30 Hz and with a width comprising between 5 and 10 Hz; means for processing the measured electroencephalographic signal, configured for processing the measured electroencephalographic signal as it is being acquired, in the sole specified beta frequency band; and means for identifying, for each breathing cycle, an improper adjustment of the ventilator from the electroencephalographic signals processed in the sole beta frequency band. | 1. A system for electroencephalographic detection of an improper adjustment of a ventilator used on a mammal, comprising: an electroencephalograph, capable of measuring, as a function of time, an electroencephalographic signal representative of a breathing process;an input receiving a respiratory initiation signal, different from the electroencephalographic signal, capable of indicating a respiratory initiation time;wherein the detection system further comprises:a beta band device specifying a sole beta frequency band comprised between 15 and 30 Hz and with a width comprised between 5 and 10 Hz;a processor configured to: process the measured electroencephalographic signal as it is being acquired, in the sole specified beta frequency band,calculate the instantaneous power of the electroencephalographic signal in the sole beta frequency band, andcalculate, for each breathing cycle, from the instantaneous power, an average power of the electroencephalographic signal in the sole beta frequency band over a first interval;memory, storing the calculated instantaneous power; andan identifier identifying, for each breathing cycle, an improper adjustment of the ventilator from the electroencephalographic signals processed in the sole beta frequency band: wherein the identifier comprises a comparator comparing the instantaneous power calculated in a second interval extending in advance of the respiratory initiation time with the average power calculated over the first interval, the first interval extending in advance of the second interval, andwherein the identifier further comprises a detector, capable of detecting, for each breathing cycle, a deviation between the average power calculated in the first interval and the instantaneous power calculated in the second interval, andwherein said deviation comprises a decrease of the instantaneous power calculated in the second interval compared to the average power calculated over the first interval and is representative of an improper adjustment of the ventilator. 2. The electroencephalographic detection system according to claim 1, wherein the identifier identifying the improper adjustment is configured to identify, for each breathing cycle, a possible desynchronization of the electroencephalographic signal in the specified beta frequency band, said desynchronization preceding the respiratory initiation time. 3. The electroencephalographic detection system according to claim 1, wherein the identifier further comprises a trigger capable of triggering an indicator if a deviation is detected for a predetermined number of successive breathing cycles. 4. The electroencephalographic detection system according to claim 1, further comprising, at an input of the beta band device, a beta band determiner determining the beta frequency band, wherein said beta band determiner is also configured to establish a time-frequency map of the electroencephalographic signal measured by the electroencephalograph, wherein said time... | UNIVERSITE PIERRE ET MARIE CURIE (PARIS 6) ASSISTANCE PUBLIQUE-HOPITAUX DE PARIS UNIVERSITE LILLE 2 DROIT ET SANTE CENTRE HOSPITALIER UNIVERSITAIRE DE LILLE | A61B005/04 A61B005/048 A61B005/08 A61M016/00 G06K009/00 A61B005/00
US09775554 | B2 | 2017-10-03 | Population cohort-linked avatar | 2007-12-31 | Avatars, computer program products, devices and systems are described that relate to at least one cohort-linked avatar comprising at least one avatar attribute that is at least partly based on an association of at least one physiologic activity of at least one member of a population cohort with at least one mental state. | 1. An avatar characteristic population cohort linking system, comprising: circuitry configured for receiving at least one characteristic for testing in conjunction with at least one avatar designated for use with at least one population cohort;circuitry configured for determining at least one preference of at least one member of the at least one population cohort responsive to at least one presentation of at least one avatar including the at least one characteristic to the at least one member of the at least one population cohort, including at least: circuitry configured for presenting the at least one avatar including the at least one characteristic to the at least one member of the at least one population cohort;circuitry configured for measuring at least one physiologic activity of the at least one member of the at least one population cohort, the measuring of the at least one physiologic activity proximate in time to presenting the at least one avatar including the at least one characteristic; andcircuitry configured for associating the at least one physiologic activity with at least one mental state indicative of at least one preference of the at least one member of the at least one population cohort with respect to the at least one characteristic; andcircuitry configured for linking the at least one characteristic to the at least one population cohort at least partly based on the at least one preference of the at least one member of the at least one population cohort with respect to the at least one characteristic. 2. The system of claim 1, wherein circuitry configured for presenting the at least one avatar including the at least one characteristic to the at least one member of the at least one population cohort comprises: circuitry configured for presenting the at least one avatar including the at least one characteristic to at least one virtual world participant, at least one social networking website participant, or at least one computer game participant representative of the at least one population cohort. 3. The system of claim 1, wherein circuitry configured for receiving at least one characteristic for testing in conjunction with at least one avatar designated for use with at least one population cohort comprises: circuitry configured for receiving at least one physical appearance characteristic for testing in conjunction with at least one avatar designated for use with at least one population cohort, the at least one physical appearance characteristic designated for linking with the at least one population cohort subsequent to a positive reaction to the at least one physical appearance characteristic by the at least one member of the population cohort as the at least one preference via the detection of the at least one physiologic activity and association with the at least one mental state. 4. The system of claim 3, wherein circuitry configured for receiving at least one physical appearance characteristic for testing in conjuncti... | Invention Science Fund I, LLC | G06F003/048 A61B005/16 A61B005/0484 A61B005/00 A61B005/04 A61B005/055 A61B005/1455
US09808166 | B2 | 2017-11-07 | Pulse photometry probe | 2009-06-24 | A pulse photometry probe includes: a light emitter having a first face from which light is emitted toward a living body; a light receiver having a second face which receives the light from the living body; a surface sheet which faces the first face of the light emitter and the second face of the light receiver; a cover sheet in which at least one slit is formed; and a lead wire which includes: a first lead wire connected to one of the light emitter and the light receiver; a second lead wire connected to the other one of the light emitter and the light receiver; and a basal portion at which the first lead wire and the second lead wire are bundled. The second lead wire includes at least one slack portion between the basal portion and the other one of the light emitter and the light receiver. | 1. A pulse photometry probe comprising: a light emitter having a first face from which light is emitted toward a living body;a light receiver having a second face which receives the light from the living body;a surface sheet which faces the first face of the light emitter and the second face of the light receiver;a cover sheet in which at least one slit is formed, the light emitter and the light receiver which are disposed between the surface sheet and the cover sheet; anda lead wire which includes: a first lead wire connected to one of the light emitter and the light receiver;a second lead wire connected to the other one of the light emitter and the light receiver; anda basal portion at which the first lead wire and the second lead wire are bundled, wherein the second lead wire includes at least one slack portion between the basal portion and the other one of the light emitter and the light receiver,wherein the pulse photometry probe is configured to circumferentially wrap around an axis extending from a proximal end to a distal end of an appendage and,wherein the at least one slit is configured to reduce aperture formation and thereby reduce interference from ambient light when the pulse photometry probe is wrapped around the axis extending from the proximal end to the distal end of the appendage. 2. The pulse photometry probe according to claim 1, further comprising a first sheet which covers the surface sheet and the cover sheet. 3. The pulse photometry probe according to claim 2, wherein the basal portion is covered by at least one of the cover sheet and the first sheet. 4. The pulse photometry probe according to claim 1, wherein the slit passes through the cover sheet. 5. The pulse photometry probe according to claim 1, wherein the slit has a bottom portion on the cover sheet without passing through the cover sheet. 6. The pulse photometry probe according to claim 1, wherein the second lead wire extends toward the other one of the light emitter and the light receiver at an angle, which is 30 degrees or more, from the basal portion, with respect to a direction in which the light emitter and the light receiver are arranged. 7. The pulse photometry probe according to claim 1, wherein a light intensity adjuster is disposed on a portion of the surface sheet, and the portion of the surface sheet faces the second face of the light receiver. 8. The pulse photometry probe according to claim 1, further comprising a second sheet to be in contact with the living body, wherein the surface sheet includes a third face which faces the light emitter and the light receiver and a fourth face opposite to the third face, and the second sheet faces the fourth face of the surface sheet. 9. The pulse photometry probe according to claim 8, wherein a light intensity adjuster is disposed in a portion of the second sheet, and the portion of the second sheet faces the second face of the light receiver. 10. The pulse photometry probe according to claim 1, where... | NIHON KOHDEN CORPORATION | A61B005/026 A61B005/1455 A61B005/00 A61B005/024
US09826926 | B2 | 2017-11-28 | Implantable sensor element | 2009-10-02 | The invention relates to a sensor element for detecting at least one analyte in a body fluid or in a body tissue, particularly for determining at least one metabolite concentration in a body fluid. The sensor element comprises an implantable, one-piece shaped body, which comprises a sensor end and a coupling end. The shaped body comprises, in the area of the sensor end, at least one sensor area, which comprises at least one sensor material. The sensor material changes at least one optically measurable property in the presence of the analyte. The shaped body also has at least one optically transparent coupling part, which is designed to transmit electromagnetic radiation in at least one spectral range between the sensor area and the coupling end. In the sensor area, the shaped body has at least one optically transparent matrix material, and the analyte can at least partially diffuse through the matrix material to the sensor material. The sensor material is embedded in the matrix material. The coupling part is formed at least partially by the matrix material. | 1. Sensor element for detecting at least one analyte in a body fluid or in a body tissue, particularly for determining at least one metabolite concentration in a body fluid, wherein the sensor element comprises an implantable, one-piece shaped body, wherein the shaped body comprises a sensor end and a coupling end, wherein the shaped body comprises, in the area of the sensor end, at least one sensor area, wherein the sensor area comprises at least one sensor material which changes at least one optically measurable property in the presence of the analyte, wherein the shaped body also has at least one optically transparent coupling part, wherein the coupling part is designed to transmit electromagnetic radiation in at least one spectral range between the sensor area and the coupling end, wherein the shaped body has, in the sensor area, at least one optically transparent matrix material, wherein the analyte can at least partially diffuse through the matrix material to the sensor material, wherein the sensor material is embedded in the matrix material, wherein the coupling part has a length between 1 millimeter and 8 millimeters and is formed at least partially by the matrix material, wherein the sensor material is at least partially embedded in the matrix material in microcapsules, wherein the matrix material comprises at least one curable plastic comprising a crosslinkable plastic. 2. Sensor element according to claim 1, wherein the coupling part is substantially free of sensor material and the crosslinkable plastic comprises a biocompatible crosslinkable plastic in cured form. 3. Sensor element according to claim 1, wherein the sensor material is immobilized in the matrix material. 4. Sensor element according to one of the claim 1, wherein the matrix material in the sensor area, in the implanted state of the sensor element, is configured to be in direct contact with the body fluid and/or the body tissue. 5. Sensor element according to claim 1, wherein the coupling part is an elongate coupling part with a substantially homogeneous refractive index. 6. Sensor element according to claim 1, wherein the crosslinkable plastic comprises at least one hydrogel and/or at least one random copolymer. 7. Sensor element according to claim 1, wherein the crosslinkable plastic is produced using at least one of the following materials: a nelfilcon polymer; a crosslinkable polyvinyl acetate or a crosslinkable derivative of the polyvinyl acetate; a crosslinkable polyvinyl alcohol or a crosslinkable derivative of the polyvinyl alcohol; a crosslinkable polymer based on polyethylene glycol, in particular based on at least one of the following polymers: bis(acryloyl) polyethylene glycol, bis(acrylamido) polyethylene glycol, a polyurethane based on polyethylene glycol, a bis- or tris-isocyanate, an acryloyl isocyanate; a crosslinkable polymer based on crosslinkable silicone hydrogel copolymers, in particular based on co-polycondenates of bis(aminodimethyl) siloxane... | EyeSense AG | A61B005/1455 A61B005/1459 A61B005/1473 B29C039/00 B29C039/02 B29C039/10 A61B005/145 A61B005/00 B29C039/24
US09808222 | B2 | 2017-11-07 | Intravascular ultrasound system for co-registered imaging | 2010-10-12 | An intravascular ultrasound imaging system with a catheter having an elongated body having a distal end and an imaging core arranged to be inserted into the elongated body. The imaging core is arranged to transmit ultrasonic energy pulses and to receive reflected ultrasonic energy pulses. The system further includes an imaging engine coupled to the imaging core and arranged to provide the imaging core with energy pulses to cause the imaging core to transmit the ultrasonic energy pulses. The energy pulses are arranged in repeated sequences and the energy pulses of each sequence have varying characteristics. The reflected pulses may be processed to provide a composite image of images resulting from each different characteristic. | 1. An intravascular ultrasound imaging system, comprising: a catheter having an elongated body having a distal end, the catheter having an imaging core configured to be inserted into the elongated body, the imaging core being configured to transmit ultrasonic energy pulses and to receive reflected ultrasonic energy pulses; andan imaging engine coupled to the imaging core and configured to provide the imaging core with energy pulses, the energy pulses being ultrasonic, to cause the imaging core to transmit the ultrasonic energy pulses, each energy pulse having a predefined energy pulse characteristic within a plurality of predefined energy pulse characteristics, the energy pulses being transmitted such that the associated predefined energy pulse characteristics are arranged in a repeating sequence, whereinthe imaging engine is further configured to process the reflected ultrasonic energy pulses into separate image frames, the separate image frames being acquired from the sequence of predefined energy pulse characteristics of the reflected ultrasonic energy pulses, each image frame corresponding to a predefined energy pulse characteristic, the imaging engine being further configured to identify prominent features of each separate image frame, the imaging engine being configured to provide display signals for simultaneously displaying the separate image frames as a composite image, the imaging engine is further configured to form the composite image by co-registering the separate image frames, such that features more prominent in a first image of the separate image frames are mapped to a second image of the separate image frames, and features more prominent in the second image are mapped to the first image, thereby including at least a portion of the prominent features of each of the separate image frames in the composite image. 2. The system of claim 1, wherein the imaging engine is further configured to provide the imaging core with repeated sequence of energy pulses such that each sequence of energy pulses includes at least two pulses. 3. The system of claim 1, wherein the imaging engine is further configured to provide the imaging core with repeated sequence of energy pulses, such that the energy pulses in each sequence having a different pulse energy. 4. The system of claim 1, wherein the imaging engine is further configured to provide the imaging core with repeated sequence of energy pulses, such that the energy pulses in each sequence having a different frequency. 5. The system of claim 1, wherein the imaging engine is further configured to provide the imaging core with repeated sequence of energy pulses, such that the energy pulses in each sequence having a different bandwidth. 6. The system of claim 1, wherein the imaging engine is further configured to provide the imaging core with repeated sequence of energy pulses such that each sequence of energy pulses includes three pulses. 7. The system of claim 6, wherein the imaging engine is... | ACIST Medical Systems, Inc. | A61B008/00 A61B008/12 G01S007/52 A61B008/08 G01S015/10 G01S015/89 A61B005/06
US09808163 | B2 | 2017-11-07 | Method and apparatus for heart rate monitoring | 2010-04-27 | A method and apparatus for monitoring heart rate. The method includes receiving a digital heart monitoring signal, determining the integrity of said signal with a “hand detect” signal that confirms electrical connection to the subject, dividing the digital heart monitoring signal into at least one frame, generating a folded correlation value for the center sample in the at least one frame, identify the number of peaks folded correlation values based on amplitude and distance parameters of the digital heart monitoring signal in the at least one frame, removing false peaks and collecting peaks corresponding to a length of time, determining the heart rate based on the identified peaks wherein the identified peaks relate to a minimum distance between valid heart beat peaks based on a heart rate estimate, computing the weighted average of heart rate value based on the number of previous heart rate values. | 1. A method for monitoring a heart rate, comprising: receiving ECG (electrocardiogram) signals;converting the ECG signals to a digital heart monitoring signal as a sequence of heart monitoring samples;receiving at a processor the digital heart monitoring signal;dividing the digital heart monitoring signal into successive overlapping frames, each with an odd number N of samples, and each frame encompassing multiple heart beats;setting amplitude and distance parameters, the distance parameter corresponding to a minimum distance between valid heart beat peaks based on a current estimate of heart rate;processing a current frame to determine a current heart rate using the amplitude and distance parameters by: generating folded correlation values based on a folded correlation of (N−1)/2 samples preceding a center sample, and time-reversed (N−1)/2 samples following the center sample;identifying a number of peaks using the folded correlation values, including removing false peaks using the amplitude parameter, andremoving false peaks using the distance parameter corresponding to a length of time between valid heart beat peaks;determining the current heart rate based on the identified peaks using at least the distance parameter;selectively updating at least the distance parameter in preparation for processing a next frame; andgenerating heart rate signals corresponding to the determined current heart rate. 2. The method of claim 1, further comprising: selectively updating the amplitude parameter in preparation for processing the next frame. 3. The method of claim 1, further comprising: computing a weighted average heart rate based on a predetermined number of previous heart rates, each computed from a respective frame. 4. The method of claim 1, further comprising: restricting a change in heart rate from the current frame to the next frame based on a predetermined timer. 5. An apparatus for monitoring a heart rate, comprising: means for receiving ECG (electrocardiogram) signals;means for converting the ECG signals to a digital heart monitoring signal as a sequence of heart monitoring samples;processing means for processing the digital heart monitoring signal, by: dividing the digital heart monitoring signal into successive overlapping frames, each with an odd number N of samples, and each frame encompassing multiple heart beats;setting amplitude and distance parameters, the distance parameter corresponding to a minimum distance between valid heart beat peaks based on a current estimate of heart rate;processing a current frame to determine a current heart rate using the amplitude and distance parameters, including generating folded correlation values based on a folded correlation of (N−1)/2 samples preceding a center sample, and time-reversed (N−1)/2 samples following the center sample;identifying a number of peaks using the folded correlation values, including removing false peaks using the amplitude parameter, andremoving false peaks using the distanc... | TEXAS INSTRUMENTS INCORPORATED | A61B005/04 A61B005/0245 A61B005/00
US09826905 | B2 | 2017-11-28 | Using colored probes in patient monitoring | 2009-05-07 | The present disclosure provides a sensor with color-coded indications that various patient physiological parameters are being monitored, such as blood oxygen saturation, blood pressure, respiration rate, and respiration effort. The sensor may sense a physical characteristic used to monitor the physiological parameter, and a visible light emitter emits visible light of a first color that is color-coded to the physiological parameter, but is not used to sense the physical characteristic. The visible light emitter may emit visibly flashing light in response to the sensor sensing a threshold value of the physical characteristic. The sensor may include a second light emitter that may sense the physical characteristic, and may emit light of a second color that is color-coded to a first or second physiological parameter. In some embodiments, the first and second colors may visibly mix. The first and second visible light emitters may emit light independently, including visibly flashing light. | 1. A sensor for sensing physical characteristics of a patient, the sensor comprising: a support structure; anda sensor component coupled to the support structure for sensing a physical characteristic of the patient, wherein the sensor component is configured to: generate a physiological signal indicative of the physical characteristic;transmit the physiological signal to a monitor external to the sensor, wherein the physiological signal is capable of being used by the external monitor to monitor first and second physiological parameters;receive a first drive signal from the external monitor indicating when the external monitor is using the physiological signal to monitor the first physiological parameter; andreceive a second drive signal from the external monitor indicating when the external monitor is using the physiological signal to monitor the second physiological parameter, the sensor component comprising:an infrared light emitter, wherein the infrared light is used to sense the physical characteristic;a detector for detecting the infrared light after it passes within the tissue of the patient;a first visible light emitter that emits visible light of a first color, wherein: the first visible light emitter is not used to sense the physical characteristic;the first color is color-coded to the first physiological parameter; andthe first visible light emitter emits visible light in response to the sensor component receiving the first drive signal indicating when the external monitor is using the physiological signal to monitor the first physiological parameter; anda second visible light emitter that emits visible light of a second color, wherein: the second visible light emitter is not used to sense the physical characteristic;the second color is color-coded to the second physiological parameter; andthe second visible light emitter emits visible light in response to the sensor component receiving the second drive signal indicating when the external monitor is using the physiological signal to monitor the second physiological parameter. 2. The sensor of claim 1 wherein the first physiological parameter comprises blood pressure, respiration rate, and/or respiration effort. 3. The sensor of claim 1 wherein the first visible light emitter emits visibly flashing light when the external monitor determines that the physical characteristic sensed by the sensor component for the first physiological parameter has reached a threshold. 4. The sensor of claim 1 wherein the first and second colors mix to appear as a third color that is color-coded to the first physiological parameter and the second physiological parameter. 5. The sensor of claim 4 wherein the third color appears to visibly flash when the external monitor determines that the physical characteristic sensed by the sensor component for the first physiological parameter and the second physiological parameter has reached a threshold. 6. The sensor of claim 1 wherein the second visible light em... | Nellcor Puritan Bennett Ireland | A61B005/00 A61B005/1455 A61B005/024 A61B005/08
US09801610 | B2 | 2017-10-31 | Method and device for measuring density of a bone | 2009-02-17 | An ultrasound device captures skeleton density information. This ultrasound device includes at least two transmitters 100, 102 having a distance of L1 between them for transmitting ultrasound signals into the skeleton, at least two receivers 104 having a distance of L2 between them for receiving ultrasound signals from the skeleton, as the calibrated known distance L is at least one of distances L1 and L2. The ultrasound device also includes a processing unit 101 for calculating a ultrasound velocity in the skeleton for forming skeleton density information on the basis of the ultrasound signals received by the receivers 104, 106 by dividing known distance by an average of a travel time difference for a ultrasound signal between the receivers 104, 106 receiving the ultrasound signal from the first transmitter 100 and the second transmitter 102. | 1. An ultrasound device for forming skeleton density information comprising: a measure head containing only two transmitters (100, 102) and only two receivers (104, 106), wherein the two transmitters (100, 102) are located between the two receivers (104, 106), and wherein the two transmitters (100, 102) and two receivers (104, 106) are located in the measure head (103), and electronics needed for processing measurement data is located in at least one location of a processor unit (101) and of the measure head (103);the two transmitters (100, 102) having a fixed known distance of L1 between common points on the transmitters for transmitting ultrasound signals of frequencies below 500 kHz into a skeleton;the two receivers (104, 106) for receiving from the skeleton the ultrasound signals of frequencies below 500 kHz transmitted by the transmitters (100, 102) which signals have travelled over said known distance L1 from opposite directions;the processor unit (101) for calculating an ultrasound velocity in the skeleton in order to form skeleton density information based on measured travel times of said ultrasound signals, which signals have travelled over said known distance L1 from opposite directions to eliminate soft tissue effects, by dividing the known distance L1 by an average of a travel time difference for an ultrasound signal between the receivers (104, 106) receiving said ultrasound signal from the first transmitter (100) and a travel time difference for an ultrasound signal between the receivers (104, 106) receiving said ultrasound signal from the second transmitter (102); andwherein the known distance L1 is the only known distance needed to calculate the ultrasound velocity. 2. The ultrasound device according to claim 1, wherein said transmitters (100, 102) transmit ultrasound signals into the skeleton at different moments in time. 3. The ultrasound device according to claim 1, further comprising, for checking quality of ultrasound signals, a transmitter (108) for transmitting ultrasound signals into the skeleton, wherein said transmitter (108) is located at a known distance L3 from the first transmitter (100) and at a known distance L4 from the second transmitter (102). 4. The ultrasound device according to claim 1, further comprising a calibration arrangement comprising a calibration program integrated programmatically to the ultrasound device, and a calibrating part (112) for calibrating the known distance L1, and in said calibrating part (112) the velocity of the ultrasound is known in different temperatures. 5. A method for forming skeleton density information, comprising: transmitting ultrasound signals of frequencies below 500 kHz into a skeleton by only two transmitters (100, 102) within a measure head (103) having a fixed known distance of L1 between common points on the transmitters;receiving ultrasound signals of frequencies below 500 kHz from the skeleton by only two receivers (104, 106) within the measure head (103), wherei... | Oscare Medical Oy | A61B008/08 G01N029/07 A61B005/00 A61B008/00
US09782122 | B1 | 2017-10-10 | Pain quantification and management system and device, and method of using | 2015-02-04 | The present invention relates to systems, devices and methods for acquiring, measuring, monitoring, processing and analyzing physiological signals. More particularly, the present invention relates to using physiological signals to determine a subject's response to various conditions, variables or constraints. Still more particularly, the present invention relates to monitoring the subject's external body motion and/or environmental factors and determining the amount of pain a subject is suffering as a result of the motion and factors. Still more particularly, the present invention relates to a system, device and methods of quantifying a subject's pain to provide an objective measurement of the subject's pain. The present invention further relates to establishing and improving pain management protocols and therapy or treatment for the subject's pain based on the quantified pain measurement. | 1. A portable pain measurement and quantification device or system comprising: at least one sensor having a signal, the sensor adapted to measure a subject's external body movement or motion or a physiological signal associated with external body movement or motion and at least one sensor adapted to measure a subject's speech;a processor with an algorithm adapted to receive the signal from the at least one sensor and calculate a quantified pain level based at least in part on the measured external body movement or motion and at least in part on speech of the subject; andan output device adapted to output the quantified level of pain. 2. The portable device or system of claim 1, further comprising an input device adapted to allow the subject to input personal pain data corresponding to the subject's experiences, habits, activities or limitations encountered during a monitoring period. 3. The portable device or system of claim 1, further comprising a pain treatment or therapy device adapted to provide pain treatment or therapy to the subject based on the quantified pain level. 4. The portable device or system of claim 1, wherein the output device is a visual display adapted to the quantified pain level visually as a representation of the pain index value. 5. The portable device or system of claim 2, wherein the output device is a visual display adapted to the quantified pain level visually as a representation of the pain index value, and the output device further visually displays a subject-inputted subject pain index value. 6. The portable device or system of claim 5, wherein the processor and algorithm are further adapted to output a two-dimensional representation of a three-dimensional data display comprising quantified pain data based on various conditions, variables, constraints or inputs corresponding to the subject's measured external body movement or motion. 7. A portable pain measurement and quantification device or system comprising: at least one sensor having a signal, the sensor adapted to measure a subject's speech;a processor with an algorithm adapted to receive the signal from the at least one sensor and calculate a quantified pain level based at least in part on the measured speech; andan output device adapted to output the quantified level of pain. 8. The portable device or system of claim 7, further comprising an input device adapted to allow the subject to input personal pain data corresponding to the subject's experiences, habits, activities or limitations encountered during a monitoring period. 9. The portable device or system of claim 7, further comprising a pain treatment or therapy device adapted to provide pain treatment or therapy to the subject based on the quantified pain level. 10. The portable device or system of claim 7, further comprising at least one sensor adapted to measure a subject's external body motion or movement and the quantified pain level is calculated based on both the measured external body mov... | Great Lakes NeuroTechnologies Inc | A61B005/00 A61B005/11 A61B005/16 A61M005/172 A61N001/36
US09795343 | B2 | 2017-10-24 | Method and apparatus for providing biometric information | 2015-10-22 | A method and apparatus for providing information are provided. The method includes measuring biometric information corresponding to a health condition of a body part of a user, determining a visual metaphor representation for characterizing a degree of the health condition based on mapping information between components of the visual metaphor representation and types of the measured biometric information, transforming, by a processor, the measured biometric information into the determined visual metaphor representation, and providing the transformed visual metaphor representation. | 1. A method of providing information, the method comprising: measuring biometric information corresponding to a health condition of a body part of a user;determining a visual metaphor representation for characterizing a degree of the health condition based on mapping information between components of the visual metaphor representation and types of the measured biometric information;transforming, by a processor, the measured biometric information into the determined visual metaphor representation; andproviding the transformed visual metaphor representation. 2. The method of claim 1, further comprising generating, by the processor, an image in which the determined visual metaphor representation is reflected on an object visualizing the body part to identify the health condition, wherein providing the transformed visual metaphor representation comprises providing the measured biometric information by displaying the generated image. 3. The method of claim 2, further comprising transmitting the generated image to an external device through a network. 4. The method of claim 1, further comprising processing, by the processor, the biometric information based on profile information of the user. 5. The method of claim 1, further comprising adjusting, by the processor, strength of a stimulus applied to the body part based on the biometric information. 6. The method of claim 1, further comprising selecting an image in which the biometric information is to be reflected from among a plurality of images obtained by photographing the body part. 7. The method of claim 1, further comprising receiving an image in which the biometric information is to be reflected, wherein the received image is obtained by photographing the body part from an external device connected through a network. 8. The method of claim 1, wherein the biometric information includes one of electrocardiogram information, electroencephalogram information, stress level information, bone mineral density information, body mass index information, calories burned information, and physical age information. 9. An apparatus for providing information, the apparatus comprising: a measuring unit configured to measure biometric information corresponding to a health condition of a body part of a user;a processor configured to determine a visual metaphor representation for characterizing a degree of the health condition based on mapping information between components of the visual metaphor representation and types of the measured biometric information, and transform the measured biometric information into the determined visual metaphor representation; anda display unit configured to provide the transformed visual metaphor representation. 10. The apparatus of claim 9, wherein the processor is further configured to generate an image in which the visual metaphor representation is reflected on an object visualizing the body part to identify the health condition, wherein the display unit is configured to pr... | Samsung Electronics Co., Ltd. | A61B005/00 A61B005/021 G06T007/00 G06T011/60 G06F019/00
US09814421 | B2 | 2017-11-14 | Apparatus for determining lifting capacity | 2015-11-30 | Apparatus for measuring lifting capacity are described. An apparatus for measuring lifting capacity has a container having an open top, opposed parallel side walls, a front wall, a back wall, and a base extending between the walls which defines an open space between the walls; and a handle removably attached to the base. The handle includes a bottom plate; a rod that extends vertically from the bottom plate, and a grip portion that is configured to connect to a locking means formed in a top portion of the rod. | 1. An apparatus for measuring lifting capacity, comprising: a container having an open top, opposed parallel side walls, a front wall, a back wall, and a base extending between the walls defining an open space therein; anda handle removably attached to the base, comprising:a bottom plate;a rod extending vertically from the bottom plate, a top portion of the rod having a locking means formed therein; anda grip portion configured to connect to the rod via the locking means;wherein the rod has an outer diameter and an inner diameter, a space being formed within the inner diameter for receiving a spring load assembly comprising:a screw having a threaded end and a cap opposite the threaded end;a washer and a spring positioned around the screw such that the washer is between the cap and the spring; anda spring load shaft having a threaded hole therein for receiving the screw; andwherein:in an initial position the spring is expanded thus biasing the washer toward the cap; andin a compressed position, the washer is biased away from the cap. 2. The apparatus of claim 1, wherein the handle grip portion comprises: a grip fastened to a grip base; andan extension member extending vertically from the grip base and having a pin inserted perpendicularly there through, wherein the pin engages with the locking means in the handle rod. 3. The apparatus of claim 2, wherein the bottom plate has a plurality of apertures formed therein which correspond to a plurality of apertures formed in the container base, and wherein the respective apertures are configured to receive attachment members for securing the bottom plate to the container base. 4. The apparatus of claim 3, wherein corresponding apertures are formed into the opposing sides walls of the container. 5. The apparatus of claim 4, wherein the handle is attached to the base such that it is closer to one of the front side or the back side of the container. 6. The apparatus of claim 5, wherein the rod receives weights for increasing the weight of the container, and wherein the back side of the container has an arced configuration generally corresponding to the shape of the weights. 7. An apparatus for measuring lifting capacity comprising: a container having an open top, opposed parallel side walls, a front wall, a back wall, and a base extending between the walls defining an open space therein; anda handle extending between the opposed side walls, the handle having a rod for receiving a spring load assembly, the spring load assembly comprising: a screw having a threaded end and a cap opposite the threaded end;a washer and a spring positioned around the screw such that the washer is between the cap and the spring; anda spring load shaft having a threaded hole therein for receiving the screw:wherein: a thickness of one of the back wall and the front wall is greater than the thickness of the other of the back wall and the front wall such that the open space is off-center; andthe handle is centered over the open... | Bardavon Health Innovations, LLC | A61B005/22
US09789206 | B2 | 2017-10-17 | Compounds, systems, and methods for monitoring and treating a surface of a subject | 2014-06-13 | Compounds, systems, and methods are provided for a non-invasive, analyte sensing dressing. The dressing includes a sensor and a matrix. The sensor is capable of detecting analytes such as molecular oxygen, carbon dioxide, nitric oxides, dissolved analytes in plasma, and hydrogen ions. The device emits a detectable signal when the sensor is excited in the presence of the analyte. The sensor may include a meso-unsubstituted metallated porphyrin that is sensitive towards oxygen. The porphyrin can be excited when illuminated at a first wavelength, followed by emission of phosphorescence at a second wavelength whose intensity is an indicator for oxygen concentration. An example metallated porphyrin has the Formula(I):;;wherein M is a metal, each R is independently an atom or a group of atoms, and at least one R is —OR1, wherein R1 is an atom or a group of atoms. | 1. A phosphorescent meso-unsubstituted porphyrin having the Formula (I):wherein M is a metal,wherein each R is independently an atom or a group of atoms, andwherein at least one R is —OR1, wherein R1 includes a moiety selected from the group consisting of alkynyl groups, triazolyl groups, alkylglutamate groups, and a pair of alkylglutamate groups. 2. The porphyrin of claim 1 wherein: R1 includes a triazolyl group. 3. The porphyrin of claim 1 wherein: R1 includes an alkylglutamate group. 4. The porphyrin of claim 1 wherein: R1 terminates in a pair of alkylglutamate groups. 5. The porphyrin of claim 2 wherein: R1 includes a triazolyl group, andR1 terminates in a pair of ethylglutamate groups. 6. The porphyrin of claim 5 wherein: every R is —OR1. 7. The porphyrin of claim 1 wherein: the porphyrin is encapsulated inside a second generation glutamic dendrimer. 8. The porphyrin of claim 1 wherein: the porphyrin is an oxygen-sensitive phosphor whose emission intensity is dependent on oxygen partial pressure. 9. The porphyrin of claim 1 wherein: the porphyrin can be excited when illuminated at a first wavelength in a range of 350-600 nanometers, followed by emission of phosphorescence at a second wavelength in a range of 600-700 nanometers. 10. A method for measuring oxygenation in tissue of a subject, the method comprising: (a) positioning a compound of claim 1 adjacent to the tissue of a subject;(b) causing the compound to phosphoresce; and(c) calculating an oxygen pressure based on phosphorescence intensity from the compound,wherein the compound is associated with a matrix incorporated in a dressing. 11. The porphyrin of claim 1 wherein: a plurality of R are —OR1. 12. The porphyrin of claim 1 wherein: every R is —OR1. 13. The porphyrin of claim 2 wherein: the triazolyl group is bonded to O via an alkyl chain. 14. The porphyrin of claim 1 wherein: M is platinum. 15. The porphyrin of claim 1 wherein: M is palladium. 16. The porphyrin of claim 9 wherein: the first wavelength is 532 nanometers, andthe second wavelength is 644 nanometers. | The General Hospital Corporation | C07D487/22 A61K049/00 A61B005/00 A61B005/1455
US09786040 | B2 | 2017-10-10 | Ultrasound diagnostic apparatus and ultrasound two-dimensional cross-section image generation method | 2013-09-12 | An ultrasound diagnostic apparatus includes a light source information setting unit setting light source data indicating a property of a light source irradiated to a cross section region of the object, an optical property setting unit setting a weight coefficient indicating an optical property of cross-section region data including intensity information on the cross section region to the light source, an illuminance slice data creation unit calculating illuminance at positions corresponding to coordinates of a plurality of the cross-section regions based on the light source data and the weight coefficient, and creating illuminance slice data of the plurality of the cross-section regions based on the calculated illuminance, and a composition unit compositing a two-dimensional cross-section image of the object from the plurality of illuminance slice data. | 1. An ultrasound diagnostic apparatus for use with tissue, the ultrasound diagnostic apparatus comprising: a light source configured to irradiate the tissue with light; anda processor programmed to: set light source data indicating a property of the light irradiated to a cross section region of the tissue;set a weight coefficient calculated based on a wavelength and illumination slice of the light source, the weight coefficient reflecting a behavior of the light irradiated in the tissue, the cross-section region data including intensity information on the cross-section region;calculate an illuminance at positions corresponding to coordinates of a plurality of cross-section regions based on the light source data and the weight coefficient, by calculating two-dimensional convolution of the light source data to generate two-dimensional convolution data;calculate illuminance slice data of the plurality of the cross-section regions based on the calculated illuminance by calculating weighted addition of the light source data and the two-dimensional convolution data based on the weight coefficient to create the illuminance slice data; andcomposite a two-dimensional cross-section image of the tissue from a plurality of the calculated illuminance slice data. 2. The ultrasound diagnostic apparatus according to claim claim 1, the processor further comprising: holding, as light source slice data, an initial value of the light source data and a calculated result of the weighted addition, andwhile switching the cross-section region data from an illuminance calculation starting slice to an illuminance calculation ending slice in the cross-section region data, the processor calculates two-dimensional convolution of the light source slice data to generate two-dimensional convolution data, and calculates weighted addition of the light source slice data and the two-dimensional convolution data based on the weight coefficient to calculate the illuminance slice data. 3. The ultrasound diagnostic apparatus according to claim 1, wherein the processor changes an order from the illuminance calculation starting slice to the illuminance calculation ending slice, based on a visual line direction. 4. The ultrasound diagnostic apparatus according to claim 1, wherein the processor sets the weight coefficient depending on intensity of the cross-section region data and the number of slices of the cross-section region data in which effective intensity continues in a prescribed direction. 5. The ultrasound diagnostic apparatus according to claim 1, wherein the processor includes at least one of: (i) averaging illuminance values of a plurality of the illuminance slice data, (ii) weighted addition of a plurality of the slice data, and (iii) rendering processing by an opacity table referred to by the intensity information of the cross-section region data, to composite the two-dimensional cross-section image. 6. The ultrasound diagnostic apparatus according to claim claim 1, furt... | HITACHI, LTD. | G06T007/00 A61B008/08 A61B005/00 G01S015/89
US09808159 | B2 | 2017-11-07 | Solid-state image sensor and imaging apparatus including the same | 2014-05-08 | An image sensor includes a first semiconductor chip including first and second surfaces; a second semiconductor chip including first and second surfaces; and a first adhesive layer between the second surface of the first semiconductor chip and the second surface of the second semiconductor chip, the first semiconductor chip being stacked on the second semiconductor chip via the first adhesive layer such that a footprint of the first semiconductor chip is larger than a footprint of the second semiconductor chip with respect to a plan view of the image sensor, the first semiconductor chip including an array of unit pixels configured to capture light corresponding to an image and to generate image signals based on the captured light, the second semiconductor chip including first peripheral circuits configured to control the array of unit pixels and receive the generated image signals. | 1. An image sensor comprising: a first semiconductor chip having a first surface and a second surface;a second semiconductor chip having a first surface and a second surface; anda first adhesive layer between the second surface of the first semiconductor chip and the second surface of the second semiconductor chip,the first semiconductor chip being stacked on the second semiconductor chip via the first adhesive layer such that the second surface of the first semiconductor chip faces the second surface of the second semiconductor chip, and a footprint of the first semiconductor chip is larger than a footprint of the second semiconductor chip with respect to a plan view of the image sensor,the first semiconductor chip including an image sensing area, the image sensing area including an array of unit pixels configured to capture light corresponding to an image and to generate image signals based on the captured light, the second surface of the first semiconductor chip including first contact pads connected to the unit pixels,the second semiconductor chip including first peripheral circuits configured to control the array of unit pixels and receive the generated image signals, the first peripheral circuits including a vertical scanning circuit, a horizontal scanning circuit, and a signal read-out circuit, the second surface of the second semiconductor chip including second contact pads connected to the first peripheral circuits, andthe first contact pads being electrically connected through the first adhesive layer to the second contact pads,wherein, the first peripheral circuits are located on a portion of the second semiconductor chip that falls inside a footprint of the image sensing area of the first semiconductor chip, with respect to the plan view of the image sensor,the horizontal scanning circuit and the vertical scanning circuit are each rectangular, and not square, in shape such that the horizontal scanning circuit and the vertical scanning circuit each have two parallel longer sides and two parallel shorter sides, andthe first peripheral circuits are arranged on the second surface of the second semiconductor chip such that the longer sides of the vertical scanning circuit are parallel to the longer sides of the horizontal scanning circuit. 2. The image sensor of claim 1, wherein the first semiconductor chip includes a silicon substrate having a thickness in a range of 0.5 to 20 microns. 3. The image sensor of claim 1 wherein, the first semiconductor chip includes first circuits, the first circuits being all circuits in the first semiconductor chip that include transistors, andeither, all the transistors included in the first circuits are NMOS type transistors and not PMOS type transistors, orall the transistors included in the first circuits are PMOS type transistors and not NMOS type transistors. 4. The image sensor of claim 1, wherein the first semiconductor chip includes a wiring layer and the wiring layer includes first wires, secon... | - | A61B006/00 A61B005/00 H01L027/146 G01T001/24
US09814531 | B2 | 2017-11-14 | System for diagnosing bloodflow characteristics, method thereof, and computer software program | 2012-08-27 | This system is a computer-based system for analyzing a blood flow at a target vascular site of a subject by means of a computer simulation, having a three-dimensional shape extraction unit, by a computer, for reading a captured image at the target vascular site and generating three-dimensional data representing a shape of a lumen of the target vascular site; a fluid dynamics analysis unit, by a computer, for determining state quantities (pressure and flow velocity) of blood flow at each position of the lumen of the target vascular site by means of computation by imposing boundary conditions relating to blood flow to the three-dimensional shape data; a blood flow characteristic determination unit for determining, from the state quantities of the blood flow determined by the fluid dynamics analysis unit, a wall shear stress vector at each position of the lumen wall surface of the target vascular site, determining relative relationship between a direction of the wall shear stress vector at a specific wall surface position and directions of wall shear stress vectors at wall surface positions surrounding the specific wall surface position, and from the morphology thereof, determining characteristics of the blood flow at the specific wall surface position and outputting the same as a determined result; and a display unit, by a computer, for displaying the determined result of the blood flow characteristic which is graphically superposed onto a three-dimensional shape model. | 1. A computer-based system for determining patient-specific vascular information, the system comprising at least one computer configured to: receive three dimensional data representing a shape of at least one vascular portion of the patient, the three dimensional data being suitable for simulations using computational fluid dynamics (CFD);determine information regarding blood flow properties including pressure and velocity at each computational mesh in the vascular portion of the patient using a CFD simulation on the three dimensional data;determine, based solely on the information regarding the blood flow properties, information relating to vascular wall vulnerability characteristics at each computational mesh located on an inner surface of the vascular portion of the patient,wherein the information relating to the vascular wall vulnerability characteristics includes a disturbance in blood flow exerted at each computational mesh and at surrounding computational meshes, located on the inner surface of the vascular portion of the patient;wherein the disturbance is determined based on at least a relative relationship among directions of wall shear stress vectors at each computational mesh and at surrounding computational meshes, located on the inner surface of the vascular portion of the patient; anddisplay the information relating to the vascular wall vulnerability characteristics graphically superposed on the three dimensional data. 2. The system of claim 1, wherein the computer is further configured to: determine if the relative relationship among the directions of the shear stress vectors at each computational mesh and at surrounding computational meshes, located on the inner surface of the vascular portion of the patient is “parallel”, “confluent”, “rotational”, or “divergent”, anddetermine the information relating to the vascular wall vulnerability characteristics to be benign (or non-malignant) if the relative relationship is “parallel”, otherwise malignant (or non-benign), anddisplay distinguishable symbols, each representing benign or malignant respectively, the symbols graphically superposed on the three dimensional data. 3. The system of claim 1, wherein the computer is further configured to: display a symbol representing potential wall-thinning at a computational mesh on the inner surface of the vascular portion of the patient, graphically superposed on the three dimensional data, if the relative relationship among the directions of the wall shear stress vectors at the mesh is “divergent”. 4. The system of claim 1, the computer is further configured to: receive an input from a user including a choice from a plurality of surgical treatments for a target site on the vascular portion, each of the surgical treatments being associated with a set of instructions to modify the three dimensional data for the surgical treatment,modify the three dimensional data, the three dimensional data being first three dimensional data, according to the i... | EBM CORPORATION | G06G007/50 A61B019/00 A61B005/055 A61B005/026 G06T019/00 G06F017/50 G06T017/00 A61B005/00 A61B017/08 A61B017/12 A61F002/82 A61F002/95 A61B034/00 A61B090/00 A61B006/00 A61B008/06 A61B005/02 G06F019/00 A61B034/10
US09815384 | B2 | 2017-11-14 | Wakefulness-maintenance apparatus | 2013-07-08 | A wakefulness-maintenance apparatus applies a stimulus, which is not perceived to be unpleasant to most seated persons, to a seated person, and effectively maintaining the wakefulness of the seated person. After a control device provided to the wakefulness-maintenance apparatus stimulates the seated person using a first stimulus at a timing close to a human heartbeat, if an index showing the wakefulness of the seated person departs within a predetermined time from a standard indicating that the wakefulness has been maintained, the control device drives a stimulus device to stimulate the seated person using a second stimulus having a timing that differs from the first stimulus. | 1. A wakefulness-maintenance apparatus comprising: a vehicle seat comprising a stimulus device for stimulating a seated person; anda control device for driving the stimulus device,wherein: the vehicle seat comprises heartbeat sensors for acquiring cardiac potential signals of the seated person, andthe control device:drives the stimulus device to stimulate the seated person using a first stimulus synchronous with a human heartbeat during the time when a QRS-wave occurs in an electrocardiographic waveform according to the cardiac potential signals, if an index showing wakefulness of the seated person departs from a standard value indicating that the wakefulness has been maintained, anddrives the stimulus device to stimulate the seated person using a second stimulus having a timing that differs from the first stimulus, if the index showing wakefulness of the seated person departs from the standard value within a predetermined time after driving the stimulus device to stimulate the seated person using the first stimulus. 2. The wakefulness-maintenance apparatus according to claim 1, wherein after the control device stimulates the seated person using the first stimulus, if the index departs from the standard value after the predetermined time passes, the control device drives the stimulus device to stimulate the seated person using the first stimulus. 3. The wakefulness-maintenance apparatus according to claim 1, wherein the second stimulus is a stimulus asynchronous with the heartbeat. 4. The wakefulness-maintenance apparatus according to claim 1, wherein the first stimulus is a stimulus defined from an average heartbeat information of a human and is a predefined stimulus. 5. The wakefulness-maintenance apparatus according to claim 1, wherein the control device drives the stimulus device to apply the second stimulus multiple times to the seated person. 6. The wakefulness-maintenance apparatus according to claim 1, wherein the control device drives the stimulus device to stimulate the seated person depending on a detection of an R-wave in the electrocardiographic waveform according to the cardiac potential signals. 7. The wakefulness-maintenance apparatus according to claim 1, wherein, in a case where the control device cannot acquire the cardiac potential signals from the heartbeat sensors, the control device drives the stimulus device based on the cardiac potential signals previously acquired. | TS Tech Co., Ltd. Public University Corporation Nagoya City University | B60N002/00 A61B005/0245 A61B005/0456 A61B005/18 A61M021/00 B60N002/44 B60K028/00 A61B005/00
US09813149 | B2 | 2017-11-07 | Method for communicating in a magnetic resonance apparatus and magnetic resonance apparatus operating therewith | 2016-01-08 | In a method for communicating between a first communication device of a magnetic resonance apparatus and a second communication device, in particular one that is mobile and on the patient side, of the magnetic resonance apparatus, a communication technology using visible light is used as the transmission medium for transmission of a useful signal from at least one of the communication devices to the other communication device, in particular from the first communication device to the second communication device. | 1. A method for communicating an information signal to an interior of a patient receptacle inside of a magnetic resonance (MR) data acquisition scanner of an MR apparatus, said MR data acquisition scanner having illumination lighting inside of said patient receptacle that emits light that illuminates the overall interior of said patient receptacle, said method comprising: providing an input information signal to a control circuit situated outside of said patient receptacle and, in said control circuit, generating an electrical control signal that represents information in said input information signal;communicating said control signal to said illumination lighting inside of said patient receptacle, and operating said illumination lighting with said control signal to embody said information in the light emitted by said illumination lighting;detecting said light embodying said information with a light detector circuit situated in the interior of said patient receptacle and, in said light detector circuit, recovering said information from the detected light and converting said information into an output information signal; andcommunicating said output information signal to an acoustic signal generator situated in the interior of said patient receptacle and, in said acoustic signal generator, converting said information signal into an acoustic signal, and emitting said acoustic signal from said acoustic signal generator in the interior of the patient receptacle. 2. A method as claimed in claim 1 comprising emitting visible light from said illuminating lighting inside of said patient receptacle. 3. A method as claimed in claim 2 comprising operating said illuminating lighting with said control signal to invisibly embody said information in said visible light. 4. A method as claimed in claim 2 wherein said visible light comprises a carrier signal, and wherein said method comprises, in said control circuit, embodying said information in said carrier signal of said visible light by modulating a brightness amplitude of said carrier signal. 5. A method as claimed in claim 4 comprising, in said control circuit, modulating said brightness amplitude using analog angle modulation of said carrier signal. 6. A method as claimed in claim 1 comprising operating said illumination lighting from said control circuit with a carrier signal that is phase-locked to a basic pulse signal of the magnetic resonance apparatus. 7. A method as claimed in claim 6 wherein said light detector circuit recovers said information from the detected light using an internal pulse signal, in said light detector circuit, that is phase-locked to said carrier signal. 8. A method as claimed in claim 1 comprising using a patient headset, adapted to be worn by a patient in said MR data acquisition scanner, as said acoustic signal generator. 9. A method as claimed in claim 1 comprising using a hand-held mobile communication device as said acoustic signal generator. 10. A magnetic resona... | Siemens Aktiegesellschaft | H04B010/00 H04B010/116 A61B005/055
US09788772 | B2 | 2017-10-17 | Wearable system and method for monitoring intoxication | 2016-12-12 | A system for transdermal alcohol sensing to be worn near a skin surface of a user, including: an alcohol sensor; a microporous membrane; a housing coupled to the alcohol sensor and the membrane, defining a volume between the alcohol sensor and a first membrane side, and fluidly isolating the volume from a second membrane side opposing the first membrane side; an electronics subsystem electrically coupled to the alcohol sensor, operable to power and receive signals from the alcohol sensor; and a fastener operable to position the second membrane side proximal the skin surface. | 1. A system for transdermal alcohol sensing to be worn near a skin surface of a user, the system comprising: an alcohol sensor comprising a fuel cell, the fuel cell comprising a sensing electrode and a counter electrode;an inlet providing a path for a sample from the skin surface of the user to the alcohol sensor;a housing coupled to the alcohol sensor and comprising a barrier defining an aperture proximal the inlet, the aperture and the inlet cooperatively controlling flow of the sample to the alcohol sensor;an electronics subsystem electrically coupled to the alcohol sensor by way of the housing, the electronics subsystem operable to receive signals from the alcohol sensor and power the alcohol sensor; anda fastener coupled to the housing and operable to position the aperture proximal the skin surface. 2. The system of claim 1, the fuel cell further comprising a reservoir wafer and a primary wafer, the primary wafer comprising a first catalytic coating and a second catalytic coating opposing the first catalytic coating, wherein: the counter electrode is electrically coupled to the first catalytic coating; andthe sensing electrode is electrically coupled to the second catalytic coating and to the reservoir wafer. 3. The system of claim 2, wherein the first catalytic coating and the second catalytic coating each comprise platinum. 4. The system of claim 1, further comprising a microporous membrane comprising a first membrane side and a second membrane side opposing the first membrane side wherein the microporous membrane is permeable to ethanol vapor. 5. The system of claim 4, wherein the microporous membrane is substantially impermeable to water. 6. The system of claim 5, wherein the fastener is further operable to position the second membrane side proximal the skin surface. 7. The system of claim 6, wherein the housing further defines a volume between the alcohol sensor and the first membrane side, wherein the volume is fluidly isolated from the second membrane side. 8. The system of claim 4, wherein the barrier comprises a first barrier side coupled to the microporous membrane, wherein the microporous membrane covers the aperture. 9. The system of claim 8, wherein: the housing further defines a volume between the alcohol sensor and the microporous membrane;the barrier further comprises a second barrier side opposing the first barrier side;the system further comprises a second microporous membrane coupled to the second barrier side and covering the aperture; andthe second microporous membrane is fluidly isolated from the volume. 10. The system of claim 1, wherein the fastener is further configured to retain the second membrane side against the skin surface. 11. The system of claim 1, wherein the fastener is operable to be repeatedly fastened and unfastened by the user. 12. The system of claim 11, wherein the fastener comprises a strap operable to encircle a forearm of the user, wherein the forearm comprises the skin surface. 13. The... | KHN Solutions, Inc. | G08B023/00 A61B005/1477 G01N027/416 A61B005/145 A61B005/00 A61B005/08 A61B005/18 G01N033/497 A61B010/00 A61B005/097
US09826889 | B2 | 2017-11-28 | Display device, medical device, display method and program | 2013-10-15 | A display device configured to display an endoscopic image obtained by an endoscope for observing an object is provided. The display device includes a projecting unit configured to project coordinate axes set to the object onto the endoscopic image, and a display unit configured to display thereon the endoscopic image and the coordinate axes projected onto the endoscopic image. | 1. A display device configured to display an endoscopic image obtained by an endoscope for observing an object, the display device comprising: a projecting unit configured to project a plurality of coordinate axes set to the object onto the endoscopic image such that an origin of the plurality of coordinate axes is aligned with a protruding direction of the endoscope;a display unit configured to display thereon the endoscopic image and the plurality of coordinate axes projected onto the endoscopic image; anda selecting unit configured to select a coordinate axis to be displayed on the display unit from among the plurality of coordinate axes, wherein the display unit is configured to display the coordinate axis selected by the selecting unit. 2. The display device according to claim 1, wherein the selecting unit is configured to select the coordinate axis on the basis of an angle formed between each of the plurality of coordinate axes and a view direction of the endoscope. 3. The display device according to claim 1, further comprising a first calculating unit configured to calculate an angle of a distal end of the endoscope. 4. The display device according to claim 3, wherein the first calculating unit is configured to calculate an elevation angle of the distal end of the endoscope. 5. The display device according to claim 3, wherein the first calculating unit is configured to calculate an azimuthal angle of the distal end of the endoscope. 6. The display device according to claim 1, further comprising a surgical instrument for performing a surgical procedure on a target for the object, which is provided at the distal end of the endoscope, wherein the surgical instrument is capable of protruding in the view direction of the endoscope. 7. The display device according to claim 6, wherein the display unit is configured to display the target for the object. 8. The display device according to claim 6, further comprising a second calculating unit configured to calculate a distance from the position of the distal end of the endoscope to the target for the object. 9. The display device according to claim 8, wherein the display unit is configured to display a distance from the position of the distal end of the endoscope to the target for the object. 10. The display device according to claim 6, wherein the endoscope is configured to be capable of setting a virtual protrusion distance of the surgical instrument when the surgical instrument is assumed to be protruded from the distal end of the endoscope. 11. The display device according to claim 10, wherein the display unit is configured to display the protrusion distance. 12. A medical apparatus comprising a display device configured to display an endoscopic image obtained by an endoscope for observing an object, the display device comprising: a projecting unit configured to project a plurality of coordinate axes set to the object onto the endoscopic image such that an origin of the plurality of c... | Shiga University of Medical Science GE Medical Systems Global Technology Company, LLC | A61B001/045 A61B001/00 G06T011/00 G01B011/25 H04N005/232 A61B005/055
US09782126 | B2 | 2017-10-10 | System for providing insightful lifestyle notifications | 2015-12-07 | A lifestyle service with data and routines storable in a memory of a wearable electronic device is discussed. The lifestyle service has at least a Sleep data routine and an Activity data routine. The lifestyle service can cooperate with an insight engine. The lifestyle service is aware of different types of information contained in one or more time synchronous applications. The system can compare tracked physical activity data with historical data in set increments in light of the information in the time synchronous applications and then communicate notifications to the user. | 1. An apparatus, comprising: a lifestyle service with data and routines storable in a memory in a wearable electronic device in an executable format, where the routines of the lifestyle service are configured to be executed by a processor in the wearable electronic device,where the routines of the lifestyle service include at least i) a Sleep data routine, andii) an Activity data routine configured to cooperate with an insight engine;where the Activity data routine is configured to collect and track data on a physical activity of a user of the wearable electronic device, via cooperation with any of an accelerometer, a magnetometer, a gyroscope, a barometer, and a light sensor, in the wearable electronic device,where the Activity data routine is configured to cooperate with a graphic user interface to display the tracked physical activity data on a display screen of the wearable electronic device;where the insight engine is configured to monitor, by pulling information from or receiving information from, one or more time synchronous applications,where the one or more time synchronous applications include any application that is configured to keep a time synchronized organization of at least two of i) personal communication information,ii) past and future event information, andiii) fact-based information, each of which are logged or scheduled in the one or more time synchronous applications,where the one or more time synchronous applications are resident on either the wearable electronic device or on a mobile computing device;where the insight engine is configured to cooperate and communicate with the wearable electronic device,where the insight engine is configured to also monitor tracked sleep data from the Sleep data routine and the tracked physical activity data from the Activity data routine,where the insight engine is configured to make correlations based on 1) the monitored tracked sleep data from the Sleep data routine or2) the monitored tracked physical activity data from the Activity data routine with3) any of the monitored personal communication information, the monitored past and future event information, or the monitored fact-based information logged or scheduled in the one or more time synchronous applications; andwhere the insight engine is configured to generate notifications with suggestions regarding i) sleep activity and/orii) physical activity to the user of the wearable electronic device based on the correlations,where the wearable electronic device is configured to communicate the notifications to the user of the wearable electronic device by cooperation with any combination of a group consisting of i) a speaker in the wearable electronic device to emit an audible notification,ii) a vibrator in the wearable electronic device to communicate a vibration notification, andiii) the display screen to display information on the display screen to communicate a visual notification,wherein the lifestyle service is further configured to ... | Pebble Technology, Corp. | A61B005/00 G06Q010/10
US09808173 | B2 | 2017-11-07 | Method for the spatially resolved determination of physcial, chemical and/or biological properties or state variables | 2004-04-15 | Method for the spatially resolved determination of physical, chemical and/or biological properties or state variables and/or the change therein in an examination area of an examination object by determining the change in the spatial distribution and/or the mobility, particularly the mobility in rotation, of magnetic particles in this examination area or in parts thereof as a function of the effect of physical, chemical and/or biological influencing variables on at least a part-area and/or in the physical, chemical and/or biological conditions in at least a part-area of the examination area, by means of the following steps: a) introducing covered and/or coated magnetic particles with at least one solid, viscous and/or liquid shell or coating into at least part of the examination area and/or introducing magnetic particles into at least part of the examination area and/or covering and/or coating at least some of these particles in the examination area, b) generating a magnetic field with a spatial profile of the magnetic field strength such that there is produced in the examination area a first part-area having a low magnetic field strength and a second part-area having a higher magnetic field strength, 15 c) changing the, in particular relative, spatial position of the two part-areas in the examination area or changing the magnetic field strength in the first part-area so that the magnetization of the particles is locally changed, d) detecting signals that depend on the magnetization in the examination area that is influenced by this change, and e) evaluating the signals so as to obtain information about the change in the spatial distribution and/or mobility of the magnetic particles in the examination area. The invention also relates to functionalized magnetic particle compositions and magnetic particle compositions suitable for use in the above method. The invention further also relates to an apparatus for the measurement of state variables in the examination area. | 1. A method for the spatially resolved determination of physical, chemical and/or biological properties or state variables, particularly substance concentrations, temperature, pH and/or physical fields, and/or the change in such physical, chemical and/or biological properties or state variables in an examination area of an examination object by determining a change in spatial distribution and/or mobility of magnetic particles in the examination area or in parts thereof as a function of the effect of physical, chemical and/or biological influencing variables on at least a part-area and/or in the physical, chemical and/or biological conditions in at least a part-area of the examination area, the method comprising: a) introducing coated magnetic particles having a coating into at least part of the examination area,b) generating a magnetic field with a spatial profile of the magnetic field strength such that there is produced in the examination area a first part-area having a low magnetic field strength and a second part-area having a higher magnetic field strength,c) changing the relative spatial position of the first and second part-areas in the examination area or changing the magnetic field strength in the first part-area so that the magnetization of the particles is locally changed,d) detecting signals that depend on magnetization in the examination area that is influenced by said changing, ande) evaluating the signals so as to obtain information about change in the spatial distribution and/or mobility of the magnetic particles in the examination area, wherein the coating is degradable and inhibits mobility of the particles. 2. A method as claimed in claim 1, wherein said changing takes place before said introducing, or said introducing and said changing are carried out essentially at the same time, and/or said changing, said detecting and said evaluating are repeated at least once. 3. A method as claimed in claim 1, wherein the examination object is a polymer material, a polymer melt, a microorganism, a plant, a plant part, a living thing or a part of a living thing. 4. A method as claimed in claim 3, wherein the polymer is a thermoplastic polymer or a polymer blend. 5. A method as claimed in claim 1, wherein a degree of mobility of the magnetic particles in the examination area is determined continuously or at intervals and is correlated with a state variable or property of the examination area, the state variable including a temperature, a concentration and/or a viscosity. 6. A method as claimed in claim 1, wherein a degree of mobility of the magnetic particles in a polymer melt that is forming or curing is determined continuously or at intervals and is correlated with a degree of curing or a degree of melting of a polymer material. 7. A method as claimed in claim 6, wherein the polymer material is a thermoplastic polymer. 8. A method as claimed in claim 1, wherein at least some of the magnetic particles have anisotropic properties. 9... | KONINKLIJKE PHILIPS N.V. | G01N033/48 A61B005/05 G06F011/07 A61B005/06 A61K049/18 G06G007/58
US09820679 | B2 | 2017-11-21 | Canine leg measurement device and method of use | 2007-04-16 | A measurement device and method for measuring a portion of an animal's leg to more accurately and fairly determine the height setting required for a jump apparatus to be jumped by the animal. The device, which is particularly useful for measuring dogs for Flyball competitions, comprises an elongated body member and a sliding mechanism slidably attached thereto. The preferred body member has a curved cross-section. A preferred sliding mechanism comprises a slot disposed in the body member and a slide member that slidably engages the slot. In use, the body member is placed against the ulna of the dog's front leg with the first end of the body member positioned at the intersection of the ulna and the accessory carpal bone so a contact plate of the slide member can be slid against the elbow. A color coded and labeled chart on the body member provides the jump height setting. | 1. A leg measurement device for measuring a leg of an animal, said leg measurement device comprising: a body member having a first end, a second end, an upper surface and a lower surface, said first end having a tip with an outwardly extending end that extends outwardly from said first end of said body member, said outwardly extending end configured to abut an accessory carpal bone of the leg;a sliding mechanism slidably attached to said body member, said sliding mechanism having a slide member configured to slide on said body member; andat least one measurement chart on said body member, said measurement chart configured to determine a jump height setting when a portion of the leg is disposed between said first end of said body member and said sliding mechanism. 2. The leg measurement device according to claim 1, wherein said slide member comprises a contact plate, a push plate and a base member interconnecting said contact plate and said push plate, said base member configured to slide in relation to said upper surface of said body member. 3. The leg measurement device according to claim 1, wherein said body member has a generally curved cross-section. 4. The leg measurement device according to claim 1, wherein said sliding mechanism further comprises a slot disposed in said body member, said slide member slidably engaging said slot. 5. The leg measurement device according to claim 4, wherein said sliding mechanism further comprises one or more fasteners configured to slidably attach said slide member to said body member through said slot. 6. The leg measurement device according to claim 4, wherein said slide member comprises a contact plate, a push plate and a base member interconnecting said contact plate and said push plate, said base member configured to slide in relation to said upper surface of said body member. 7. The leg measurement device according to claim 6, wherein said contact plate extends upwardly from said base member so as to contact the leg when said slide member is slid toward said first end of said body member. 8. The leg measurement device according to claim 1, wherein said measurement chart directly provides said jump height setting. 9. A leg measurement device for measuring a leg of an animal, said leg measurement device comprising: an elongated body member having a first end, a second end, an upper surface and a lower surface, said body member having a generally curved cross-section;a tip at said first end of said body member;a sliding mechanism slidably attached to said body member, said sliding mechanism having a slot disposed in said body member and a slide member slidably engaging said slot; andat least one measurement chart on said body member,wherein said leg measurement device measures the distance between an accessory carpal bone and an elbow of the leg to provide a jump height setting from said measurement chart when said upper surface of said body member is placed against an ulna of the leg, said tip en... | - | A61B005/107 A61B005/00
US09795334 | B2 | 2017-10-24 | Method and apparatus for penetrating tissue | 2007-07-09 | A tissue penetrating system has a housing member, a plurality of penetrating members positioned in the housing member and a plurality of sample chambers. Each sample chamber is associated with a penetrating member. A tissue stabilizing member has a tissue interface surface configured to be applied to a tissue surface and provide for spontaneous flow of blood for sample capture. The tissue stabilizing member is coupled to the housing. | 1. A tissue penetrating device, comprising: a housing member; an electronic penetrating member driver positioned in the housing member a plurality of penetrating members positioned in the housing member,each penetrating member coupled to and driven by the electronic penetrating member driver; a first processor coupled to the electronic penetrating member driver and to a feedback loop, the first processor storing penetrating member driver profiles in a memory, selecting a penetrating member driver profile from a set of alternative driver profiles based on a desired tissue penetration performance, and modifying an electronic penetrating member driver profile based on additional user input information, the first processor configured to modulate power from a power supply to the electronic penetrating member driver in response to a modification of the electronic penetrating member driver profile;a plurality of sample chambers positioned in the housing member, each sample chamber associated with a penetrating member;an interface coupled with the first processor coupled to a second processor, the second processor including a sleep mode and can run intermittently to conserve power, the second processor providing an alarm time set;a real time clock coupled to one or more alarms to provide a user with a reminder of a next penetrating event; andan interface surface that is coupled to the housing member. 2. The device of claim 1, wherein the interface surface includes a tissue stabilizing member to cause at least a portion of the tissue surface to press in an inward direction relative to the housing member. 3. The device of claim 2, wherein the interface surface is configured to enhance fluid flow from a wound created below the tissue surface. 4. The device of claim 2, wherein the interface surface is configured to stretch the tissue surface. 5. The device of claim 2, wherein the interface surface is configured to apply a force to the tissue surface. 6. The device of claim 2, wherein the interface surface is configured to be positioned to at least partially surround an impact location of the penetrating member on the tissue surface. 7. The device of claim 2, wherein the interface surface is a stimulating member. 8. The device of claim 1, further comprising: a vacuum source configured to assist creation of a low pressure environment in the interface surface or along a fluid path to a sample chamber. 9. The device of claim 1, wherein each sample chamber has a volume no greater than 1 μL. 10. The device of claim 1, wherein in a first direction each of the penetrating members moves toward the tissue surface at a speed that is different than a speed at which each of the penetrating members moves away from the tissue surface. 11. The device of claim 1, wherein in a first direction each penetrating member moves toward the tissue surface at a speed that is greater than a speed at which the penetrating member moves away from the tissue surface. 12. The dev... | SANOFI-AVENTIS DEUTSCHLAND GMBH | A61B005/151 A61B005/15 A61B005/155 A61M005/32 A61B005/157 A61B005/145 A61B005/1486 A61B005/00
US09804416 | B2 | 2017-10-31 | Energizable ophthalmic lens with an event-based coloration system | 2013-05-21 | The present invention provides a device for an energizable Ophthalmic Lens with an event coloration mechanism. The event coloration mechanism may color or change color based on some predefined event. For example, a predefined constituent or predefined condition of the tear fluid may be indicative of the predefined event, and the event coloration mechanisms may interact with the tear fluid, accordingly. The event coloration mechanism may provide the energizable functionality of the Ophthalmic Lens in some embodiments. In others, the event coloration mechanism may be passive but may interact and interface with the electrical components of the Ophthalmic Lens, such as, for example, those included within the Media Insert. Event coloration mechanisms may be combined with additional functionalities that may be included in an energizable Ophthalmic Lens. | 1. An ophthalmic lens, comprising: a soft lens portion comprising a polymerized reactive monomer mixture; anda media insert in contact with the soft lens portion, the media insert comprising an event coloration mechanism configured to provide a visual indication in the ophthalmic lens based on an occurrence of a predefined event, wherein the event coloration mechanism comprises: a power source;a coloration portion electrically connected to the power source by a conductive trace; anda tube portion comprising a body and an encapsulant layer capable of surrounding the body, the tube portion being in circuit with the conductive trace, the coloration portion, and the power source and being capable of limiting an activation of the coloration portion. 2. The ophthalmic lens of claim 1, wherein the event coloration mechanism is located within the media insert, and wherein the predefined event comprises a condition of a component located within the media insert. 3. The ophthalmic lens of claim 2, wherein the component comprises the power source and wherein the condition comprises a leeching of a caustic substance from a pre-designated area for the power source. 4. The ophthalmic lens of claim 2, wherein the predefined event comprises a predefined pH level within the media insert. 5. The ophthalmic lens of claim 2, wherein the media insert further comprises: a variable optic element capable of altering an optical power of the ophthalmic lens in electrical communication with a processor and the power source. 6. The ophthalmic lens of claim 1, wherein the coloration portion comprises: a liquid crystal, wherein the liquid crystal is capable of a resting orientation and a first energized orientation, and wherein the liquid crystal in the resting orientation comprises a first color and the liquid crystal in the first energized orientation comprises a second color. 7. The ophthalmic lens of claim 6, wherein the body comprises a reactive material capable of reacting with a specified tear fluid constituent, and wherein the specified tear fluid constituent is capable of permeating the encapsulant layer. 8. The ophthalmic lens of claim 1, wherein the predefined event comprises a pathogen level in tear fluid. 9. The ophthalmic lens of claim 1, wherein the predefined event comprises a biomarker level in tear fluid. 10. The ophthalmic lens of claim 1, the predefined event comprises an active agent level in tear fluid. 11. An ophthalmic lens, comprising: a soft lens portion comprising a polymerized reactive monomer mixture; anda media insert in contact with the soft lens portion, the media insert comprising an event coloration mechanism configured to provide a visual indication in the ophthalmic lens based on an occurrence of a predefined event;wherein the media insert comprises: a power source located within the media insert; anda processor in electrical communication with the power source, and wherein the event coloration mechanism comprises:an illumination ... | Johnson & Johnson Vision Care, Inc. | G02C007/04 A61B005/145 G02C007/08 G02C011/00 A61B003/10 G02C007/02 A61B005/00
US09776012 | B2 | 2017-10-03 | Variable frequency impedance measurement | 2013-03-08 | When a defibrillator selects a dosage of energy or current to be delivered to a patient, the defibrillator selects an excitation current frequency and applies the excitation current at the selected frequency to the patient. The frequency of the excitation current is selected as a function of the dosage to be delivered. The patient's response to the excitation current at the selected frequency will accurately reflect the impedance that the defibrillator will “see” when delivering the selected dosage of energy or current. | 1. A method comprising: applying a first excitation current at a first frequency to a patient;measuring a first impedance of the patient by measuring a response of the patient to application of the first excitation current;selecting a second frequency of a second excitation current as a function of both the first measured impedance and a dosage of at least one of energy and defibrillation current to be delivered to the patient with a defibrillator. 2. The method of claim 1, further comprising: applying the second excitation current to the patient at the selected second frequency; andmeasuring a second impedance of the patient by measuring a response of the patient to application of the second excitation current. 3. The method of claim 2, further comprising selecting an excitation current function as a function of the first measured impedance. 4. The method of claim 2, wherein the excitation current function comprises an equation that relates excitation current frequency to an energy or current dosage. 5. The method of claim 2, wherein the excitation current function comprises a lookup table that relates excitation current frequency to an energy or current dosage. 6. The method of claim 2, further comprising finding the difference between the first impedance and the second impedance. 7. The method of claim 1, wherein the defibrillator is an external defibrillator. 8. A defibrillator comprising: a current source configured to generate an excitation current at a selectable frequency;at least two electrodes configured to deliver the excitation current to a patient at a selected frequency;an impedance measuring circuit configured to measure the voltage across the electrodes and to measure an impedance as a function of the voltage and the excitation current; anda processor configured to select the selected frequency of the excitation current as a function of a dosage of at least one of energy and defibrillation current to be delivered to the patient. 9. The defibrillator of claim 8, wherein the current source is configured to generate a first excitation current at a first frequency and a second excitation current at a second frequency, and wherein the impedance measuring circuit is configured to measure a first impedance associated with the first excitation current at the first frequency and a second impedance associated with the second excitation current at the second frequency. 10. The defibrillator of claim 9, wherein the processor is configured to select the second selected frequency as a function of the first impedance. 11. The defibrillator of claim 8, wherein the processor is configured to select a dosage of at least one of energy and defibrillation current to be delivered to the patient. 12. The defibrillator of claim 8, wherein the defibrillator is an external defibrillator. | Physio-Control, Inc. | A61N001/39 A61B005/053
US09782082 | B2 | 2017-10-10 | Body temperature logging patch | 2013-06-25 | An actively-powered temperature data logger patch with wireless data communication includes a sealed, flexible battery comprising a printed electrochemical cell with an anode and a cathode, and a flexible circuit including a microprocessor, a temperature sensor configured to sense a temperature of a target subject, a wireless communication transceiver and an antenna. In one example, the flexible battery and flexible circuit are arranged in a covering, stacked arrangement between first and second substrate layers, and the patch configured to conform to a curved surface of the target subject. In another example, the patch is used in an actively-powered medical system for monitoring a body temperature of a patient, and includes an adhesive configured to be removably applied to skin of the patient. An external computing device is capable of two way communication with the wireless communication transceiver of the patch via an electromagnetic field. | 1. An actively-powered temperature data logger patch with wireless data communication, comprising: a first substrate layer;a sealed, flexible battery configured to provide continuous electrical power, comprising a printed electrochemical cell with an anode and a cathode, at least one of said anode and cathode being formed of a cured or dried ink, and first and second battery electrode contacts that are each electrically coupled to one of the anode and cathode;a flexible circuit comprising a microprocessor, a temperature sensor configured to sense a temperature of a target subject, a wireless communication transceiver and an antenna, the flexible circuit further comprising first and second battery contact pads that are each electrically coupled to one of the first and second battery electrode contacts to thereby electrically power the microprocessor and temperature sensor,wherein the microprocessor and the temperature sensor actively receive continuous electrical power from the flexible battery irrespective of whether the communication transceiver is powered,wherein the microprocessor further comprises a timer that actively receives electrical power from the flexible battery and is configured to enable the microprocessor to continuously obtain a plurality of temperature samples from the temperature sensor at a periodic time interval; anda second substrate layer comprising an adhesive configured to be removably applied to a surface of the target subject,wherein the flexible battery and flexible circuit together comprise an electronics inlay that is arranged in a covering, stacked arrangement between the first and second substrate layers, andwherein all of the first substrate layer, electronics inlay, and second substrate layer are sufficiently flexible so that the temperature data logger patch is configured to conform to a curved or variable surface of the target subject and is able to flex and move together with movement of the target subject without inadvertent removal of the temperature data logger patch from the target subject. 2. The patch of claim 1, wherein the wireless communication transceiver is configured to be passively powered by an electromagnetic field from an external computing device. 3. The patch of claim 2, wherein the wireless communication transceiver utilizes a standard NFC communications protocol. 4. The patch of claim 1, wherein the wireless communication transceiver utilizes a standard Bluetooth or Bluetooth low-energy communications protocol. 5. The patch of claim 1, wherein the microprocessor comprises a unique identification code (UID). 6. The patch of claim 5, wherein the microprocessor is configured to wirelessly transmit the plurality of temperature samples to an external computing device via the wireless communication transceiver and antenna. 7. The patch of claim 1, wherein the microprocessor further comprises a memory for storing the temperature samples together with an associated time-stamp for each temperat... | Blue Spark Technologies, Inc. | A61B005/00 A61B005/01 G01K001/02 G01K013/00
US09827431 | B2 | 2017-11-28 | Wearable defibrillator with no long-term ECG monitoring | 2013-05-30 | A wearable defibrillator system includes a support structure with one or more electrodes in an unbiased state. A monitoring device monitors, for the long term, a parameter of the person that is not the person's ECG; rather, the monitored parameter can be the person's motion, a physiological parameter, or both. When a value of the monitored parameter reaches a threshold, such as when the person is having an actionable episode, the electrode becomes mechanically biased against the person's body, for making good electrical contact. Then, if necessary, the person can be given electrical therapy, such as defibrillation. As such, the electrodes of the wearable defibrillator system can be worn loosely for the long term, without making good electrical contact. This can reduce the person's aversion to wearing the defibrillation system. | 1. A wearable defibrillator system configured to be worn by a person who may be moving and to defibrillate the person by delivering an electrical charge to the person's body, comprising: a support structure configured to be worn by the person;a monitoring device that includes a motion detector, the monitoring device configured to monitor at least one parameter of the person that is a motion of the person's body;a capacitor configured to store the electrical charge, for delivering to the person's body while the person is wearing the support structure;first electrodes coupled to the support structure, the coupling such that, while the support structure is worn by the person, at least a certain one of the first electrodes is at one of an unbiased state in which the certain electrode is moveable with respect to the person's body responsive to the person's moving, and a biased state in which the certain electrode is biased towards the person's body against the support structure so as to be less moveable with respect to the person's body than when in the unbiased state, the certain electrode including an adhesive material; anda first biasing mechanism configured to cause the certain electrode to transition from the unbiased state to the biased state responsive to a value of the monitored parameter reaching a first threshold, the first biasing mechanism causing the certain electrode to thus transition by exerting a biasing force, the first biasing mechanism in turn causing the certain electrode to adhere to the person's body due to the adhesive material and remain adhered even if the first biasing mechanism discontinues exerting the biasing force. 2. A wearable defibrillator system configured to be worn by a person who may be moving and to defibrillate the person by delivering an electrical charge to the person's body, comprising: a support structure configured to be worn by the person;a monitoring device configured to monitor at least one parameter of the person that is not an electrocardiogram (“ECG”) of the person;a capacitor configured to store the electrical charge, the stored electrical charge configured to be delivered to the person's body while the person is wearing the support structure;first electrodes coupled to the support structure, the coupling such that, while the support structure is worn by the person, at least a certain one of the first electrodes is at one of an unbiased state in which the certain electrode is moveable with respect to the person's body responsive to the person's moving, and a biased state in which the certain electrode is biased towards the person's body against the support structure so as to be less moveable with respect to the person's body than when in the unbiased state, the certain electrode including an adhesive material; anda first biasing mechanism configured to cause the certain electrode to transition from the unbiased state to the biased state responsive to a value of the monitored parameter reaching a fir... | WEST AFFUM HOLDINGS CORP. | A61N001/00 A61N001/39 A61B005/00 A61B005/0408 A61B007/00 A61B005/0205
US09795371 | B2 | 2017-10-24 | Surgical access system and related methods | 2016-03-01 | A surgical access system including a tissue distraction assembly and a tissue retraction assembly, both of which may be equipped with one or more electrodes for use in detecting the existence of (and optionally the distance and/or direction to) neural structures before, during, and after the establishment of an operative corridor to a surgical target site. | 1. A three-bladed retractor for use in forming a working corridor to a spine, the three-bladed retractor comprising: a mounting structure configured for connecting to three retractor blades;a first elongated retractor blade extending from a first proximal end coupled to the mounting structure to a first distal end that is spaced longitudinally from the first proximal end;a second elongated retractor blade extending from a second proximal end coupled to the mounting structure to a second distal end that is spaced longitudinally from the second proximal end; anda third elongated retractor blade extending from a third proximal end coupled to the mounting structure to a third distal end that is spaced longitudinally from the third proximal end, wherein the third elongated retractor blade includes a first edge extending between the third proximal end and the third distal end, a second edge extending between the third proximal end and the third distal end, a first inner face oriented toward an interior facing direction and defined by the third elongated retractor blade between the first and second edges, and a first outer face oriented toward an exterior facing direction and defined by the third elongated retractor blade between the first and second edges, wherein the first inner face of the third elongated retractor blade defines a first engagement groove dimensioned to couple to a first engagement portion of a first connectable element, wherein the first engagement groove is defined by the first inner face of the third elongated retractor blade between a first ridge extending along a first portion of the first engagement groove proximate the first edge of the third elongated retractor blade and a second ridge extending along a second portion of the first engagement groove proximate the second edge of the third elongated retractor blade, and wherein the first inner face of the third elongated retractor blade defines at least one first recess dimensioned to receive a first member of the first connectable element so as to hold the first connectable element in place with respect to the third elongated retractor blade. 2. The three-bladed retractor of claim 1, wherein the mounting structure is configured to position the first elongated retractor blade adjacent the second and third elongated retractor blades, the second elongated retractor blade adjacent the first and third elongated retractor blade, and the third elongated retractor blade adjacent the first and second elongated retractor blades to form a working corridor and wherein the first inner face of the third elongated retractor blade faces the first and second elongated retractor blades. 3. The three-bladed retractor of claim 1, wherein the three-bladed retractor is configured to form the working corridor along a lateral, trans-psoas path to a targeted intervertebral disc of a lumbar portion of the spine and wherein the third elongated retractor blade is a posterior-most retractor blade, the sec... | NuVasive, Inc. | A61B017/02 A61B001/32 A61B005/00 A61B017/00
US09792800 | B2 | 2017-10-17 | Wireless tag and home monitoring device for tracking individuals or objects with alcohol monitoring | 2016-12-16 | A methodology determining a response to exposure of a monitored person to a prohibited substance is provided. The methodology includes: first receiving movement data representing movement of the monitored person over time; second receiving an indication that a monitored person has been exposed to a quantity of a prohibited substance; first determining, from the received movement data, movement activity of the monitored person for a period of time before the exposure was detected; second determining a risk parameter based on at least (a) the amount of the quantity of the prohibited substance, and (b) the movement activity; issuing a notice to a supervising authority in substantially real time when the second determining results in a risk parameter that is above a first threshold; and issuing a notice to a supervising authority in non-real time when the second determining results in a risk parameter that is below the first threshold. | 1. A method performed on electronic computer hardware in conjunction with software for determining a response to exposure of a monitored person to a prohibited substance, comprising: first receiving an indication that a monitored person has been exposed to a prohibited substance;second receiving movement activity representing movement history of a monitored person;determining a risk presented by the monitored person's exposure to the prohibited substance based at least partially on the movement activity, comprising determining exposure when the movement activity indicates the monitored person was remote from a predetermined location during a period of time before exposure to be a greater risk relative to exposure to the prohibited substance when the movement activity indicates that the monitored person remained at the predetermined location during the period of time;issuing, in response to the determined risk being above a first threshold, a notice to a supervising authority in substantially real time; andissuing, in response to the determined risk being between a second and third threshold, a notice to a supervising authority in non-real time. 2. The method of claim 1, wherein the first and second thresholds are the same. 3. The method of claim 1, wherein the first and second thresholds are different. 4. The method of claim 1, wherein the third threshold is zero. 5. The method of claim 1, wherein the first receiving comprises receiving an indication of a type of the prohibited substance that a monitored person has been exposed to, and wherein the determining further comprising considering the type of the prohibited substance in determining the risk. 6. The method of claim 1, wherein the first receiving comprises receiving an indication of a quantity of how much of the prohibited substance that a monitored person has been exposed to, and wherein the determining further comprising considering the quantity of the prohibited substance in determining the risk. 7. The method of claim 6, wherein in response to quantity exceeding a fourth threshold, the determining sets the determined risk above the first threshold regardless of whether the monitored person was remote from or present at the predetermined location during the period of time. 8. The method of claim 6, wherein in response to the quantity being below a fourth threshold in combination with the monitored person being at the predetermined location, the method disregards the exposure and does not issue any corresponding alert to a supervising authority. 9. A method performed on electronic computer hardware in conjunction with software for determining a response to exposure of a monitored person to a prohibited substance, comprising: first receiving an indication that a monitored person has been exposed to a prohibited substance;second receiving movement activity representing movement history of a monitored person;determining a risk presented by the monitored person's exposure to the proh... | SATELLITE TRACKING OF PEOPLE LLC | G08B023/00 G08B021/02 A61B005/00 A61B005/11
US09775532 | B2 | 2017-10-03 | Remote control of implantable device through medical implant communication service band | 2007-09-10 | A system and method for communicating data and signals through the Medical Implant Communication Service Band using a repeater or base station in the proximity to an implantable device within a patient is disclosed. In a preferred embodiment, the device is capable for early detection and monitoring of congestive heart failure in a patient. Impedance measurements, or other health parameters depending on the type of implantable device or sensor used, are sent using a bi-directional low-power radio operating in the MICS band to a nearby base station which may provide signal processing and analysis. The base station may have an interface to one or more communications networks to connect to a remote location. The system and method of the present invention permits a healthcare professional to monitor an ambulatory patient's condition at a remote location and to program the implanted device. | 1. A device-implemented method of detecting and monitoring an event indicative of congestive heart failure in a patient, comprising: implanting the device on a lower left anterior lateral hemithorax of the patient;performing ongoing measurements of impedance of a portion of the patient's body generally occupied by the lungs, said measurements of impedance being indicative, at least in part, of said event indicative of congestive heart failure;transmitting transmissions including said measurements of impedance to a base station at a transmission time being based, at least in part, on an occurrence of the event indicative of congestive heart failure using a radio transmitter; andsending data relating to said measurements of impedance from the base station to a remote location via a telecommunications network. | Medtronic, Inc. | A61B005/05 A61B005/04 A61B005/00 A61B005/0205 A61B005/053 A61N001/362 A61B005/08 A61N001/365
US09781907 | B2 | 2017-10-10 | System to detect the level of stress/discomfort of aquatic animals | 2012-01-27 | System for detecting the level of discomfort of aquatic animals during experimental studies, with a water tank, a reference electrode placed at one of the sides of the tank, at least one recording electrode placed at another side of the tank, a bio amplifier for amplifying a received bio-signal at the recording electrode, a microprocessor for treating the signals and a low-pass filter for filtering the signals received. The system can provide a precise, quantifiable and specific indication of the level of stress/wellbeing of aquatic animals in normal living conditions as well as in experimental conditions without interfering with the animals' life. | 1. System for detecting the level of discomfort of aquatic animals during experimental studies, the system comprising: a water tank (1),a reference electrode (2) covering the entire surface of one of the sides of the tank in direct contact with the water, andan array of recording electrodes (3,3a,3b) in the form of printed circuit boards (PCB), covering the surface of another of the sides of the water tank, in direct contact with said side of the water tank,a bio amplifier (4) for amplifying a received bio-signal at the recording electrode,a micro-processor for treating the signals, anda low-pass filter for filtering the signals received, capable of filtering signals up to 20 Hz. 2. System according to claim 1 wherein the reference and signal electrodes (2, 3) are made of the same material. 3. System according to claim 2 wherein the electrodes are made of aluminum, copper or silver. 4. System according to claim 1, wherein the electrodes (2, 3a, 3b . . . ) are embedded in the walls of the tank (1). | FUNDACION AZTI | G01N027/00 A01K061/00 A01K029/00 A61B005/00 A61B005/04 A61B005/11 A61B005/16 G01N033/18 A01K063/00 A61D099/00 A01K061/10
US09775564 | B2 | 2017-10-03 | Anal fixation device | 2013-10-29 | Anorectal diagnostic procedures are oftentimes carried out with the patient laying in the left decubitus position, which leads to non-representative response to test maneuvers due to the unnatural position and patient anxiety. Devices spanning, or fastened on, the interglutial cleft of a patient allow for a more natural patient position and eliminate non-representative responses. | 1. An anal fixation device for use in anorectal diagnostic procedures, comprising: a housing configured to accept a probe and to maintain the probe in a fixed position relative to the subject's anus verge during an anorectal diagnostic procedure, said housing comprising two arms and a base which is a connection point of the two arms; anda probe support comprising a support arm and a retaining piece, wherein: a proximal end of said support arm is connected to the base such that the two arms can flex without moving said support arm,said support arm extending in-between the two arms,the retaining piece is connected to a distal end of said support arm, andthe retaining piece is configured to accept the probe for performing anorectal diagnostic procedure. 2. The anal fixation device according to claim 1, wherein the two arms of said housing are configured to span the interglutial cleft of a subject. 3. The device according to claim 1, wherein the housing has a first, uncompressed width in an uninstalled configuration, and has a second, compressed width in an installed configuration. 4. The device according to claim 1, wherein the two arms oppose each other and are positioned to create an open space therebetween, wherein said support arm extends into said open space. 5. The anal fixation device according to claim 1, wherein the housing includes a plastic material. 6. The anal fixation device according to claim 1, wherein the housing is configured to be fastened at the interglutial cleft of the subject. 7. The anal fixation device according to claim 1, wherein the two arms form an open object or a closed object. 8. The anal affixation device according to claim 1, wherein the probe support comprises a means for fastening or retaining the probe. 9. The device according to claim 8, wherein the fastening means is selected from the group consisting of a belt-like elastic band, a clamp-like structure, a lock piece, a plastic or rubber stopper-type mechanism and a tapered lock piece. 10. The anal fixation device according to claim 8, wherein the fastening means comprises a belt-like elastic band attached to one side of said probe support. 11. The anal fixation device according to claim 10, wherein said probe support has a said clamp-like structure with a side opening capable of receiving the probe. 12. The anal fixation device according to claim 11, wherein the belt-like elastic band is configured to be secured across the side opening of said clamp-like structure to thereby secure the probe, by positioning a slit formed in the middle of the belt-like elastic band over and onto a knob formed as part of said probe support. 13. The anal affixation device according to claim 1, wherein the housing comprises a semi-rigid plastic material. 14. The anal fixation device according to claim 1, wherein said housing has a width defined as a dimension of the housing spanning the two arms, and a height that is defined as a dimension of the housing perpendicular ... | GIVEN IMAGING LTD. | A61B001/32 A61B005/00 A61F005/00 A61M025/02 A61B005/11
US09808157 | B2 | 2017-11-07 | Hands-free interface | 2013-12-30 | A wearable device and external reader is provided for herein. In some embodiments of the present disclosure, the wearable device or the reader is configured to receive a level of radiant energy, detect a change in the received level of radiant energy, determine that the detected change in the received level of radiant energy is indicative of a predetermined pattern of received radiant energy, and responsively operate (or cause to be operated via the external reader) one or more external devices. | 1. A method comprising: a wearable device receiving from an external reader a level of radiant energy;the wearable device detecting a change in the received level of radiant energy;the wearable device determining that the detected change in the received level of radiant energy is indicative of a predetermined pattern of received radiant energy, wherein determining that the detected change in the received level of radiant energy is indicative of a predetermined pattern of received radiant energy comprises determining that the detected change in the received level of radiant energy is indicative of at least one body movement; andin response to the determining, the wearable device transmitting to the external reader an instruction that indicates that the wearable device has determined that the detected change in the received level of radiant energy is indicative of a predetermined pattern of received radiant energy. 2. The method of claim 1, wherein receiving a level of radiant energy comprises receiving a level of light incident on a light detector associated with the wearable device. 3. The method of claim 1, wherein receiving a level of radiant energy comprises receiving a level of radio frequency (RF) radiation from the external reader. 4. The method of claim 1, wherein determining that the detected change in the received level of radiant energy is indicative of a predetermined pattern of received radiant energy comprises determining that some combination of two or more of the following has occurred: the received level of radiant energy has, within a threshold period of time, changed from a first level to a second level and then returned to at or near the first level,the received level of radiant energy has changed from a first level to a second level and then remained at the second level for at least a threshold period of time, andthe received level of radiant energy has, over a period of time greater than or equal to a threshold period of time, changed from a first level to a second level and then returned to at or near the first level. 5. The method of claim 1, wherein the wearable device transmitting to an external reader an instruction causes the external reader to transmit an instruction to at least one external device. 6. The method of claim 1, wherein the at least one body movement comprises a blink, a nod, or a change of gaze. 7. The method of claim 6, wherein the at least one body movement comprises a blink, wherein determining that the detected change in the received level of radiant energy is indicative of a blink comprises: determining that the received level of radiant energy has, within a threshold period of time, changed from a first level to a second level and then returned to at or near the first level. 8. The method of claim 6, wherein the at least one body movement comprises a change of gaze, wherein determining that the detected change in the received level of radiant energy is indicative of a change of gaze comprises... | Verily Life Sciences LLC | G06F003/02 A61B005/00 A61B005/145 A61B005/1468 G06F003/01
US09795325 | B1 | 2017-10-24 | Auditory perceptual systems | 2014-03-14 | A new hearing profile assessment plays each of a plurality of frequency tones multiple times at a plurality of loudness levels. For each frequency, the patient indicates which tones that they can hear, establishing minimum-amplitude thresholds for each frequency. The assessment adapts loudness levels until the patient indicates consistent minimum-amplitude thresholds for each frequency. Also, a new system for training the hearing of a subject is disclosed. The system instructs the subject to vocalize various sound items and remember the sounds that they vocalize. The system plays either recorded samples of the subject's own voice, or comparison samples of a synthesized voice or other's vocalizations of the same sound items. The system prompts the subject to compare the played samples with the sounds they remembered vocalizing. The system uses the feedback to adjust sound processing parameters for a hearing aid, cochlear implant, sound output devices, or a training program. | 1. An audiometry method comprising: providing a computer game environment for an audiometric exercise, the game environment displaying a plurality of visual stimuli that represent objects in a natural environment or video game theme, wherein each of a first subset of the visual stimuli are associated with an audio stimulus, and each of a second subset of the visual stimuli is not associated with an audio stimulus, and a means is provided of selecting any of the plurality of visual stimuli;wherein each audio stimulus comprises a fixed frequency tone associated with a fixed loudness level, and the loudness levels of each audio stimulus is separated from each other audio stimulus by discrete decibel increments; andperforming the audiometric exercise by: (a) displaying the plurality of visual icons;(b) for each of at least two or more of the plurality of visual icons: receiving an initial selection of the visual icon to play audio stimulus associated with the visual icon;playing the audio stimulus, if any, associated with the visual icon; andif the user hears the audio stimulus associated with the visual icon, receiving the user's second selection of the visual icon to indicate that the user can hear the audio stimulus;(c) when a threshold loudness level is identifiable, identifying a threshold loudness level for the frequency tone of the audio stimuli and reducing a decibel increment between the loudness levels associated with the visual icons; and(d) when no threshold loudness level is identifiable because the user does not make any second selections, increasing the loudness levels associated with the plurality of visual icons. 2. The audiometry method of claim 1, further comprising challenging the user, in the computer game environment, to indicate as rapidly as possible whether they can hear the frequency tones, and providing sensory feedback on an accuracy and speed of the user's response. 3. The audiometry method of claim 2, further comprising: displaying to the user, in the computer game environment, a plurality of visual objects, each associated with one of the tones; andchallenging the user to hover a pointer over the objects to play the tones and to select the objects if they are able to hear the tones. 4. The audiometry method of claim 1, further comprising: selecting stimuli to use in a subsequent game that are a function of results obtained from the audiometry exercise. 5. The audiometry method of claim 4, wherein the subsequent game is an auditory perceptual training game, and the game stimuli comprise auditory signals. 6. The audiometry method of claim 1, further comprising for each frequency tone: identifying a minimum-amplitude threshold for the frequency tone as the quietest loudness level that the user indicated that they could hear;subsequent to identifying the minimum-amplitude threshold, playing the frequency tone at a plurality of loudness levels distributed around the minimum-amplitude threshold; andrepeating the precedin... | POSIT SCIENCE CORPORATION | A61B005/00 A61B005/12 H04R025/00 G09B019/04
US09805339 | B2 | 2017-10-31 | Method for monitoring and improving health and productivity of employees using a computer mouse system | 2013-09-24 | Provided are embodiments of systems, computer medium and computer-implemented methods for sensing health characteristics of a user using a computer mouse having a location sensor configured to sense movement of the computer mouse relative to a work surface on which the computer mouse is located, a temperature sensor, a blood condition sensor and a blood pressure sensor. A method including transmitting, to a computer workstation, location data corresponding to sensed movement of the computer mouse for use in determining movement of the mouse, receiving temperature data from the temperature sensor, blood condition data from the blood condition sensor and blood pressure data from the blood pressure sensor, and transmitting, to the computer workstation, health data corresponding to the temperature data, the blood condition data, and the blood pressure data for use in determining the body temperature, the blood saturation level, and the blood pressure of the user. | 1. A method for monitoring health characteristics of a user, the method comprising: sending, by a computer mouse controller of a computer mouse of a computer mouse system to a computer workstation, location data corresponding to sensed movement of the computer mouse for use in determining movement of the computer mouse, wherein the computer mouse system comprises: the computer mouse comprising: the computer mouse controller;a location sensor configured to sense movement of the computer mouse relative to a surface on which the computer mouse is located;a temperature sensor configured to sense a body temperature of the user;a blood condition sensor configured to sense a blood saturation level of the user; and a mouse wireless interface; anda mouse wireless interface; anda blood pressure cuff worn about the wrist of the user and configured to sense a blood pressure of the user, the blood pressure cuff comprising: a bladder;a pump;a pressure transducer;a blood pressure cuff controller;a blood pressure cuff wireless interface; anda battery configured to supply power to operate the pump, the pressure transducer, the blood pressure cuff controller and the blood pressure cuff wireless interface,wherein the bladder, the pump, the pressure transducer, the blood pressure cuff controller, the blood pressure cuff wireless interface and the battery are integrated with the blood pressure cuff, and wherein the blood pressure cuff is not physically tethered to the computer mouse;receiving, by the computer mouse controller from the temperature sensor, temperature data corresponding to a sensed body temperature of the user;receiving, by the computer mouse controller from the blood condition sensor, blood condition data corresponding to a sensed blood saturation level of the user;receiving, by the computer mouse controller from the computer workstation, a request for health data;in response to receiving the request for health data, sending, by the computer mouse controller to the blood pressure cuff controller via wireless communication between the mouse wireless interface and the blood pressure cuff wireless interface, a request for a blood pressure reading;in response to the request for the blood pressure reading, the blood pressure cuff controller performing the following operations: controlling operation of the pump of the blood pressure cuff to inflate the bladder;obtaining, from the pressure transducer of the blood pressure cuff, one or more pressure readings indicative of the blood pressure of the user; andsending, to the computer mouse controller via wireless communication between the mouse wireless interface and the blood pressure cuff wireless interface, blood pressure data corresponding to the one or more pressure readings;receiving, by the computer mouse controller from the blood pressure cuff controller via wireless communication between the mouse wireless interface and the blood pressure cuff wireless interface, the blood pressure data; andsending, by ... | Saudi Arabian Oil Company | G06Q010/00 G06Q010/10 G06F019/00 A61B005/00 A61B005/01 A61B005/0205 A61B005/021 A61B005/0476 A61B005/053 A61B005/08 A61B005/1455 A61B005/16 G06Q050/22 A61B003/113
US09814393 | B1 | 2017-11-14 | Bone oxygenation measurement | 2014-07-29 | A bone oximeter probe includes an elongated member and a sensor head at an end of the elongated member to make measurements for a bone. The measurements can indicate the viability or nonviability of the bone. In an implementation, the probe is advanced through an incision in soft tissue, towards the underlying bone, and positioned so that the sensor head faces the bone to be measured. Optical signals are sent from the sensor head and into the bone. The bone reflects some of the optical signals which are then detected so that measurements for the bone can be made. Some of these measurements include an oxygen saturation level value, and a total hemoglobin concentration value of the bone. | 1. A medical device comprising a probe for making an oxygen saturation measurement for a bone, the probe comprising: a rigid elongated member, wherein the rigid elongated member comprises a proximal end, a distal end opposite the proximal end, and a smooth outer surface therebetween;a sensor head, coupled to the distal end of the rigid elongated member, wherein the sensor head comprises:first and second openings; andfirst and second fiber optic cables, wherein the first fiber optic cable is coupled to the first opening, and the second fiber optic cable is coupled to the second opening, andwherein the sensor head is positioned so that light directed into the bone via the first fiber optic cable travels in a first direction from the proximal end towards the distal end and exits the first opening in the first direction, and light reflected from the bone enters the second opening in a second direction, and travels from the distal end towards the proximal end in the second direction and through the second fiber optic cable;a flexible tube, extending from the proximal end of the rigid elongated member; anda connector, coupled to the flexible tube. 2. The medical device of claim 1 comprising a beam combiner, external to the probe, wherein the beam combiner comprises: a first input coupled to a first radiation source of a first wavelength; anda second input coupled to a second radiation source of a second wavelength, different from the first wavelength, wherein the first fiber optic cable is coupled to the beam combiner and outputs light of the first wavelength and light of the second wavelength. 3. The medical device of claim 1 wherein a length of the rigid elongated member is from about 10 centimeters to about 50 centimeters. 4. The medical device of claim 1 wherein the rigid elongated member is a solid rod and the first and second fiber optic cables are coupled to a side surface of the solid rod which joins the distal and proximal ends. 5. The probe of claim 1 wherein the sensor head comprises a deformable material that can conform to a contour of the surface of the bone to be measured. 6. The probe of claim 1 wherein the sensor head comprises third and fourth openings, and in combination with light conducted via the first and second openings, light conducted via the third and fourth openings is used in making an oxygen saturation measurement for the bone. 7. The probe of claim 1 wherein to make an oxygen saturation measurement, the sensor head is adapted to be used on a surface of the bone to be measured. 8. A medical device probe for making a bone oxygen saturation measurement, the probe comprising: an elongated member, wherein the elongated member comprises a proximal end, and a distal end, opposite the proximal end;a sensor head, coupled to the elongated member closer to the distal end than the proximal end of the elongated member, wherein the sensor head comprises:first and second openings; andfirst and second fiber optic cables, wherein t... | ViOptix, Inc. | A61B005/00 A61B005/1455 A61B005/1459
US09804237 | B2 | 2017-10-31 | Body coil for magnetic resonance imaging | 2013-02-10 | A body coil for magnetic resonance imaging includes one or more coil elements incorporated in a shell material. The body coil has a rectangular basic shape. Two opposing edge sections of the body coil may be pivoted along one pivot axis, respectively, relative to a middle section. For this purpose, movement elements that effect pivoting are arranged on an edge section side. | 1. A body coil comprising: a middle section;two opposing edge sections connected to sides of the middle section, respectively;one or more coil elements incorporated in a shell material; andmovement elements arranged beneath the two opposing edge sections, a first subset of the movement elements and a second subset of the movement elements operable to expand, the first subset of expandable movement elements pressing against a first edge section of the two opposing edge sections from beneath the first edge section, pressing against a first rigid surface opposite the first edge section from above the first rigid surface, and pivoting the first edge section relative to the middle section when the first subset of the movement elements expands, the second subset of expandable movement elements pressing against a second edge section of the two opposing edge sections from beneath the second edge section, pressing against a second rigid surface opposite the second edge section from above the second rigid surface, and pivoting the second edge section relative to the middle section when the second subset of the movement elements expands, wherein the body coil has a rectangular basic shape, and wherein the two opposing edge sections are pivotable along one pivot axis, respectively, relative to the middle section. 2. The body coil as claimed in claim 1, wherein the body coil is for magnetic resonance imaging. 3. The body coil as claimed in claim 1, wherein the movement elements are constructed as inflatable cushions or balloons. 4. The body coil as claimed in claim 1, wherein each edge section of the two opposing edge sections is divided into a plurality of individual, separately pivotal section elements, to which one of the movement elements, respectively, is allocated. 5. The body coil as claimed in claim 1, wherein pocket-like receptacles are provided on the two opposing edge sections, the pocket-like receptacles being operable for detachably receiving one or more of the movement elements. 6. The body coil as claimed in claim 1, wherein a first top edge of the middle section is flush with a top edge of one edge section of the two opposing edge sections, and a second top edge of the middle section is flush with a top edge of the other edge section of the two opposing edge sections. 7. The body coil as claimed in claim 1, wherein the movement elements are elastically deformable. 8. The body coil as claimed in claim 7, wherein for pivoting, the movement elements are operable to change a shape of the movement elements, starting from a compressed shape to an expanded shape. 9. The body coil as claimed in claim 8, wherein the movement elements are made from a rubber-elastic plastics material that serves as a spring element. 10. The body coil as claimed in claim 7, wherein each edge section of the two opposing edge sections is flexible, and wherein one or more individual movement elements of the movement elements are provided along the respective edge se... | Siemens Aktiengesellschaft | A61B005/05 G01R033/341 A61B005/055 A61B005/00 G01R033/34 G01R033/3415
US09820811 | B2 | 2017-11-21 | System and method for mapping the functional nerves innervating the wall of arteries, 3-D mapping and catheters for same | 2015-12-01 | Disclosed herein are systems and methods for locating and identifying nerves innervating the wall of arteries such as the renal artery. The present invention identifies areas on vessel walls that are innervated with nerves; provides indication on whether energy is delivered accurately to a targeted nerve; and provides immediate post-procedural assessment of the effect of energy delivered to the nerve. The methods includes evaluating a change in physiological parameters after energy is delivered to an arterial wall; and determining the type of nerve that the energy was directed to (sympathetic or parasympathetic or none) based on the evaluated results. The system includes at least a device for delivering energy to the wall of blood vessel; sensors for detecting physiological signals from a subject; and indicators to display results obtained using said method. Also provided are catheters for performing the mapping and ablating functions. | 1. A proximal renal ablation method for ablation of sympathetic renal nerve of a subject, comprising the steps of: a) introducing a catheter into the lumen of a renal artery such that the tip of said catheter contacts a site on an inner wall of said renal artery, wherein said site is within a proximal region from an ostium of said renal artery, wherein said proximal region has a length equals to one-third of total length of said renal artery;b) applying electrical stimulation by introducing electrical current to said site via the catheter, wherein said electrical current is controlled to be sufficient to elicit changes in one or more physiological parameters when there is an underlying nerve at said site, said one or more physiological parameters are selected from the group consisting of systolic blood pressure, diastolic blood pressure, mean arterial pressure, and heart rate;c) measuring said one or more physiological parameters after said electrical stimulation, wherein an increase of said one or more physiological parameters as compared to measurements obtained before said electrical stimulation would indicate that a sympathetic renal nerve has been mapped at said site; andd) applying ablation energy through the catheter to the site identified in step (c) for ablation of said sympathetic renal nerve. 2. The method of claim 1, wherein said catheter is an ablative catheter currently in use to treat cardiac arrhythmias. 3. The method of claim 1, wherein said catheter is an ablative catheter designed specifically for mapping renal nerves for ablative procedures. 4. The method of claim 1, further comprises repeating the steps (b) to (c) on the site identified in (c), wherein a lack of change in said one or more physiological parameters confirms nerve ablation. 5. The method of claim 1, wherein said electrical current has one or more of the following parameters: (a) voltage between 2 and 30 volts;(b) resistance between 100 and 1000 ohms;(c) current between 5 and 40 milliamperes;(d) time of application between 0.1 and 20 milliseconds; and(e) total applied time between 1 to 5 minutes. 6. A proximal renal ablation method for ablation of renal nerve of a subject, comprising the steps of: a. determining the presence of a renal nerve by i. contacting a first site on an inner wall of a renal artery with one or more first electrodes, said first site is in a proximal region from an ostium of said renal artery, wherein said proximal region has a length equals to one-third of total length of said renal artery;ii. applying a first electrical stimulation by introducing electrical current to said first site via said first electrodes, wherein said electrical current is controlled to be sufficient to elicit changes in one or more physiological parameters when there is an underlying nerve at said first site, said one or more physiological parameters are selected from the group consisting of systolic blood pressure, diastolic blood pressure, mean arterial pressu... | SYMAP MEDICAL (SUZHOU), LTD | A61B018/00 A61B018/14 A61B005/00 A61B005/0205 A61B005/0215 A61B005/024 A61B005/04 A61B005/20 A61B018/18 A61N001/36
US09782073 | B2 | 2017-10-10 | Hand-held medical-data capture-device having a digital infrared sensor with no A/D converter and having interoperation with electronic medical record systems via an authenticated communication channel | 2014-12-30 | In one implementation, an apparatus estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point is described, estimates temperature from a digital infrared sensor and determines vital signs from a solid-state image transducer, or determines vital signs from a solid-state image transducer and estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point; after which the estimated and/or determined information is transmitted to an external database. | 1. A non-touch thermometer to measure a temperature, the non-touch thermometer comprising: a microprocessor;a wireless communication subsystem that is operably coupled to the microprocessor and that is operable to transmit a representation of the temperature;a battery operably coupled to the microprocessor;a single button operably coupled to the microprocessor;a digital infrared sensor operably coupled to the microprocessor with no analog-to-digital converter operably coupled between the digital infrared sensor and the microprocessor, the digital infrared sensor having only digital readout ports and comprising a low noise amplifier, an analog-to-digital converter, a digital signal processor, and a pulse width modulator (PWM), the digital readout ports operable to transmit to the microprocessor a digital signal that is representative of an infrared signal, acquired from a forehead of a user, that is detected by the digital infrared sensor, wherein the digital infrared sensor is configured with no need for recalibration after an initial factory calibration with an emissivity range from 0.1 to 1.0; anda display device operably coupled to the microprocessor,wherein the microprocessor is operable to receive from the digital readout ports the digital signal that is representative of the infrared signal detected by the digital infrared sensor and the microprocessor is operable to determine the temperature from the received digital signal by using the PWM,wherein when connection is established by the wireless communication subsystem to an external device thereafter data is pushed from the non-touch thermometer through the wireless communication subsystem, thereafter the external device controls all flow of the data between the non-touch thermometer and the external device, wherein the connection further comprises an authenticated communication channel. 2. The non-touch thermometer of claim 1, wherein the display device further comprises: a green traffic light operable to indicate that the temperature is good;an amber traffic light operable to indicate that the temperature is low; anda red traffic light operable to indicate that the temperature is high. 3. The non-touch thermometer of claim 1, the microprocessor further comprising: a pixel-examination-module configured to examine pixel values of the at least two images, a temporal-variation module to determine temporal variation of the pixel values between the at least two images being below a particular threshold, a signal processing module configured to amplify the temporal variation resulting in amplified temporal variation, and a visualizer to visualize a pattern of flow of blood in the amplified temporal variation in the at least two images. 4. The non-touch thermometer of claim 1 further comprising: the digital infrared sensor having no analog sensor readout ports. 5. The non-touch thermometer of claim 1 wherein the wireless communication subsystem transmits to a static Internet Protocol address... | ARC DEVICES, LTD | G08B023/00 A61B005/00 A61B005/01 G06F019/00 A61B005/021 A61B005/026 A61B005/08 A61B005/024 A61B005/0205 G06T003/40 G06T007/00 H04W048/16 H04W076/02 A61B005/0402 A61B005/03 G01J005/00 H04W080/04 G01J005/02 G01J005/10 G08C017/02 H04L029/08 H04W012/06 H04B007/26 H04L012/46 H04N005/33 H04W048/08 G08B005/36 G08B021/18 G01K013/00 G06K009/62 G06T007/174 G06T007/269 G06T007/90 A61B005/0245 H04W012/08
US09778736 | B2 | 2017-10-03 | Methods and systems for calibrating user devices | 2014-09-22 | Methods and systems are described herein for a media guidance application that improves the customization and calibration of user devices to a particular user. For example, in response to erroneously detecting (or failing to detect) a user input of a first type, the media guidance application may re-calibrate the user device based on subsequent corrective inputs issued using a second input type such that future user inputs of the first type will not be erroneously detected (or fail to be detected). | 1. A method of calibrating user devices, the method comprising: monitoring, with a user device, a first input type for a command;receiving a first user input comprising a first user attempt to issue the command using the first input type, wherein the first user input does not trigger the command, wherein the first user input is detected based on a detection of a first indicium of the first user input, and wherein the first user input does not trigger the command because a value corresponding to the first indicium of the first user input is below a threshold value;receiving a second user input comprising a second user attempt to issue the command using a second input type, wherein the second user attempt triggers the command; andin response to receiving the second user attempt, recalibrating the user device such that subsequent receipt of the first user input with the first input type triggers the command by reducing the threshold value based on the value corresponding to the first indicium of the first user input. 2. The method of claim 1, further comprising: searching data received while monitoring the first input type prior to the second user input for the first indicium of the first user input; andwherein reducing the threshold value comprises: reducing the threshold value to a value equal to or lower than the value corresponding to the first indicium of the first user input. 3. The method of claim 2, further comprising identifying the first indicium based on the temporal proximity of the first indicium to the second user input. 4. The method of claim 2, further comprising identifying the first indicium based on biometric activity associated with triggering the command. 5. The method of claim 1, further comprising updating a user profile to include the recalibration of the user device. 6. The method of claim 1, wherein monitoring the first input type for the command includes monitoring brain activity of a user. 7. The method of claim 6, wherein receiving the first user input further comprises: cataloging an instance of the brain activity;comparing the instance with a database listing instances of brain activity that correspond to the command to determine whether or not the instance corresponds to the command; andnot triggering the command in response to determining that the instance does not correspond to any of the instances of brain activity that correspond to the command. 8. The method of claim 7, wherein recalibrating the user device further comprises updating the database listing instances of brain activity that correspond to the command to include a listing corresponding to the instance. 9. The method of claim 6, wherein receiving the first user input further comprises: cataloging an instance of the brain activity;comparing a value corresponding to the instance to a threshold value; andnot triggering the command in response to determining that the value does not exceed the threshold value. 10. The method of claim 9, wherein reca... | Rovi Guides, Inc. | G06F003/00 G06F003/01 A61B005/00 H04N021/442 G06F017/30 A61B005/16 G06F001/16 H04L029/08 H04N005/445
US09808194 | B2 | 2017-11-07 | Cushioning support for monitoring user activity | 2015-09-09 | Multi-layer cushion supports are described which may generally comprise a first support having a first contact surface for contacting a portion of a body and a second surface opposite to the first surface, the first support defining a central chamber and a peripheral chamber surrounding the central chamber, wherein the first support is filled with a first gas or liquid and a second support attached to the first support along the first contact surface. The second support may be filled with a second gas or liquid which is relatively more viscous than the first gas or liquid. In particular, the first support may be filled with a volume of air and the second support may be filled with oil which is less than the volume of air. | 1. A monitoring system comprising: a support for supporting at least a portion of a user's body, wherein the support comprises a first layer and a second layer opposite the first layer, wherein the first and second layers define a central chamber and a peripheral chamber surrounding the central chamber, wherein the central chamber and the peripheral chamber are filled with a first gas or liquid, wherein the support further comprises a third layer, wherein the third layer is opposite both the first and second layers, wherein the second and third layers define an additional chamber, and wherein the additional chamber is filled with a second gas or liquid which is more viscous than the first gas or liquid;at least one pressure sensor in communication with the support; andan electronics unit in communication with the at least one pressure sensor,wherein the electronics unit is configured to transmit information relating to pressure sensed from the at least one pressure sensor to a remote device, and wherein the electronics unit and/or remote device is programmed to provide a warning or alert to inform the user that their posture is proper or improper when seated. 2. The system of claim 1, wherein the additional chamber is sized to contain a maximum volume which is less than a maximum volume of the central and peripheral chambers combined. 3. The system of claim 2 wherein the first gas or liquid comprises a volume of air. 4. The system of claim 2 wherein the second gas or liquid comprises a volume of oil. 5. The system of claim 2 further comprising a foam layer positioned adjacent to the first layer. 6. The system of claim 5 further comprising a cover configured to envelope the support and the foam layer. 7. The system of claim 1 wherein the electronics unit comprises a microprocessor. 8. The system of claim 1 wherein the electronics unit is further configured for wireless communication with the remote device. 9. The system of claim 1 wherein the remote device comprises a smart phone, tablet, PDA, laptop, or computer. 10. The system of claim 1 wherein the electronics unit and/or remote device is programmed to track a duration of time upon loading of the support. 11. The system of claim 10 wherein the electronics unit and/or remote device is programmed to provide a warning or alert to the user when the duration of time reaches or exceeds a predetermined level. 12. The system of claim 1 wherein the support comprises at least two fluid chambers separated from one another, each fluid chamber in communication with a corresponding pressure sensor which are in communication with the electronics unit. 13. The system of claim 12 wherein the electronics unit and/or remote device is programmed to determine a pressure differential between the fluid chambers and to provide a warning or alert if the pressure differential exceeds a predetermined level. 14. The system of claim 1 wherein the pressure sensor is further configured to detect and/or monitor o... | PRS Medical Technologies, Inc. | A61B005/00 A61B005/0205 A61B005/103 G01G019/44
US09807526 | B2 | 2017-10-31 | Apparatus and method for determining parameter using auditory model of hearing loss patient | 2014-03-26 | An apparatus and method are provided to determine a parameter using an auditory model of a hearing loss patient. The parameter determination apparatus determines a similarity between a neurogram of a normal subject and a neurogram of a hearing loss patient, and determines an optimal frequency band for the hearing loss patient based on the similarity. | 1. A parameter determination apparatus, comprising: a processor configured to determine a similarity index between a reference neurogram and a patient neurogram; anda frequency band determiner configured to determine an optimal frequency band based on the similarity index between a reference neurogram and a patient neurogram,wherein the optimal frequency band is based on a rate of change of the similarity index as a function of the frequency band, and the optimal frequency band is used to calibrate a hearing device for a hearing loss patient. 2. The parameter determination apparatus of claim 1, wherein the patient neurogram is determined according to an auditory model of the hearing loss patient. 3. The parameter determination apparatus of claim 2, wherein the auditory model of the hearing loss patient is determined based on at least one of a damage degree of a middle ear, a damage degree of an inner hair cell, and an outer hair cell of an inner ear. 4. The parameter determination apparatus of claim 1, wherein the patient neurogram is output when sound applying hearing loss compensation is applied to an auditory model of the hearing loss patient. 5. The parameter determination apparatus of claim 1, wherein the optimal frequency band is a calibrating parameter to control the hearing device of the hearing loss patient. 6. The parameter determination apparatus of claim 1, wherein the frequency band determiner is configured to determine the optimal frequency band based on whether the similarity index exceeds a threshold value or whether a gradient of the similarity index is less than or equal to a threshold gradient. 7. The parameter determination apparatus of claim 1, wherein the frequency band determiner wired or wirelessly transmits the optimal frequency band to optimize and reduce power consumption of the hearing device. 8. The parameter determination apparatus of claim 1, wherein the patient neurogram is varied according to an auditory model of a patient and the reference neurogram is a neurogram of a normal person. 9. An apparatus, comprising: a hearing device configured to receive an optimal frequency band for a hearing loss patient and automatically calibrate the hearing device to adapt to the hearing loss patient, wherein the optimal frequency band is based on an auditory model of the hearing loss patient and based on a similarity index between a reference neurogram and a patient neurogram of the hearing loss patient,wherein the optimal frequency band is based on a rate of change of the similarity index as a function of the frequency band. 10. A parameter determination method, comprising: determining a similarity index between a reference neurogram and a patient neurogram;determining an optimal frequency band between a reference neurogram and a patient neurogram based on the similarity index of the patient neurogram with respect to the reference neurogram; andapplying the optimal frequency band to calibrate a hearing device for a he... | Samsung Electronics Co., Ltd. | H04R025/00 A61B005/12 A61B005/0484
US09826931 | B2 | 2017-11-28 | Medical workstation with integrated support of process steps | 2015-01-20 | A device (13) is provided for use with a treatment device for treating a patient. The device (13) includes a detection device (17) for detecting at least one transition of the patient's treatment from a first phase of treatment to a second phase of treatment. A signal device (25) for sending at least one signal when the transition is detected. It proposes, furthermore, a workstation equipped herewith. | 1. A method for monitoring different phases of an anesthesia machine, the method comprising: providing an anesthesia device;providing a plurality of anesthesia phases of treatment with said anesthesia device, said plurality of anesthesia phases of treatment comprising an induction phase, a maintenance phase and an emergence phase;providing a detection means;providing a display means comprising a user interface;determining a current anesthesia phase of treatment of the patient with said detection means;displaying the current anesthesia phase of treatment on said display means;detecting at least one transition from said current anesthesia phase of treatment to a next anesthesia phase of treatment;displaying actions associated with one or more of said current anesthesia phase of treatment and said at least one transition from said current anesthesia phase of treatment to said next anesthesia phase of treatment on said display means, said actions corresponding to processes and operations associated with said current anesthesia phase of treatment and/or said at least one transition from said current anesthesia phase of treatment to said next anesthesia phase of treatment, wherein parameters of said processes and operations are displayed on said display means during one or more of said current anesthesia of treatment and said at least one transition from said current anesthesia phase of treatment;adjusting one or more of said parameters on said display means; andselecting at least one future anesthesia phase of treatment based on said actions displayed on said display means via said user interface, wherein said display means displays said selected at least one future anesthesia phase of treatment. 2. A method in accordance with claim 1, further comprising: providing a signal means;sending at least a signal when the at least one transition from said current anesthesia phase of treatment to said next anesthesia phase of treatment is detected via said signal means, wherein the at least one transition of one anesthesia phase to the next anesthesia phase is one of a transition from standby to preoxygenation, a transition from preoxygenation to intubation, a transition from intubation to a wash-in process, a transition from a wash-in process to instrumentalization, a transition from instrumentalization to maintenance, a transition from maintenance to heart-lung-machine-mode, a transition from heart-lung-machine-mode to maintenance, a transition from maintenance to a wash-out process, a transition from a wash-out process to extubation and a transition from extubation to standby, said parameters corresponding to one or more of gas concentration parameters, alarm parameters and ventilation parameters. 3. A method in accordance with claim 1, further comprising: providing a data bank means; andstoring work flows data with said data bank means. 4. A method in accordance with claim 3, further comprising: providing a control means; andcontrolling said work flows ... | DRÄGERWERK AG & CO. KGAA | A61B005/00 A61M016/00 A61M021/00 G06F019/00 A61B005/0205 A61B005/091
US09792787 | B2 | 2017-10-17 | Wearable electronic device | 2016-03-29 | An apparatus including an elongated apparatus structure configured to fit around a body part of a user and an actuator configured to change shape of the elongated apparatus structure. The actuator is configured to change tightness of the elongated apparatus structure around the body part of the user based on an action the apparatus is performing. | 1. An apparatus comprising: an elongated apparatus structure configured to fit around a body part of a user;an actuator configured to change shape of the elongated apparatus structure to tighten or loosen the elongated apparatus structure around the body part of the user; anda pressure sensor configured to detect compression force of the elongated apparatus structure against skin of the body part of the user;wherein the actuator is configured to automatically change tightness of the elongated apparatus structure around the body part, to achieve a predefined contact pressure of the apparatus structure against the skin of the body part of the user, in response to the pressure sensor detecting the compression force of the apparatus structure against the skin of the body part of the user. 2. The apparatus of claim 1, further comprising: a communication unit physically part of the elongated apparatus structure, the communication unit configured to receive wireless signals from a remote device, the wireless signals indicating activity in the remote device;wherein the actuator is configured to automatically change tightness of the elongated apparatus structure around the body part, in response to the communication unit receiving the wireless signals indicating activity in the remote device. 3. The apparatus of claim 1, further comprising: a processing unit physically separate from the elongated apparatus structure, the processing unit configured to cause transmission of wireless signals in response to activity in the processing unit;a communication unit physically part of the elongated apparatus structure, the communication unit configured to receive the wireless signals from the processing unit;wherein the actuator is configured to automatically change tightness of the elongated apparatus structure around the body part, in response to the communication unit receiving the wireless signals from the processing unit. 4. The apparatus of claim 1, further comprising: a communication unit configured to receive an indication of an incoming wireless signal;wherein the actuator is configured to automatically change tightness of the elongated apparatus structure around the body part, in response to the communication unit indicating receipt of an indication of an incoming wireless signal. 5. The apparatus of claim 1, further comprising: a communication unit configured to receive an incoming wireless signal;wherein the actuator is configured to automatically tighten the elongated apparatus structure around the body part of the user in response to initiating receipt of a wireless signal by the communication unit and automatically loosen the elongated apparatus structure around the body part of the user in response to ending the receipt of the wireless signal by the communication unit. 6. The apparatus of claim 1, further comprising: a second sensor configured to detect body signal data related to at least one of heart rate, skin moisture, or blood pressure of t... | Nokia Technologies Oy | H04B003/36 G08B006/00 H04M003/02 H04M003/42 A61B005/021 A61B005/00 A61B005/024 H04M001/725
US09788794 | B2 | 2017-10-17 | Method and apparatus for generating assessments using physical activity and biometric parameters | 2015-02-27 | The methods and apparatuses presented herein determine and/or improve the quality of one or more physiological assessment parameters, e.g., response-recovery rate, based on biometric signal(s) and/or motion signal(s) respectively output by one or more biometric and/or motion sensors. The disclosed methods and apparatuses also estimate a user's stride length based on a motion signal and a determined type of user motion, e.g., walking or running. The speed of the user may then be estimated based on the estimated stride length. | 1. An assessment generation system configured to assess a fitness level of a subject, the system comprising one or more processing circuits configured to: measure an average heart rate of the subject with a heart rate sensor worn by the subject over a period of time;measure an average cadence of the subject with a cadence sensor worn by the subject over the period of time; andcalculate a ratio of the average heart rate to the average cadence with a processor operatively coupled to the heart rate sensor and the cadence sensor;wherein the one or more processing circuits are further configured to determine a relative body temperature using the calculated ratio of the average heart rate to the average cadence, wherein the higher the calculated ratio the higher the relative body temperature. 2. The assessment generation system of claim 1 wherein the fitness level is inversely proportional to the ratio of the average heart rate to the average cadence. 3. An assessment generation system configured to evaluate an exercise activity comprising one or more sets of exercise repetitions performed by a user via at least one biometric sensor and at least one motion sensor comprised in a wearable monitoring device, the assessment generation system comprising: one or more processing circuits configured to: process a motion signal output by the at least one motion sensor and a biometric signal output by the at least one biometric sensor to identify a type of the exercise activity;process the motion signal using the identified exercise activity to determine an exercise cadence;process the biometric signal using the identified exercise activity to determine at least one biometric parameter;wherein the one or more processing circuits are further configured to estimate at least one of: a number of repetitions of the identified exercise activity using an integral of the exercise cadence; anda number of sets of the identified exercise activity using a time-dependent change of the at least one biometric parameter. 4. The assessment generation system of claim 3 wherein the one or more processing circuits are further configured to estimate an intensity of the identified exercise activity using the integral of the exercise cadence and an integral of the at least one biometric parameter. 5. An assessment generation system configured to determine a blood pressure characteristic of a user using a device comprising a photoplethysmography sensor and a motion sensor, wherein the photoplethysmography sensor is proximal to at least some skin of the user, the assessment generation system comprising one or more processing circuits configured to: screen data from the motion sensor to determine an integrity of data from the photoplethysmography sensor;buffer a plurality of pulses, at least some of the buffered pulses having integrity;generate an average pulse shape using the buffered plurality of pulses; anddetermine the blood pressure characteristic using a waveform representation... | VALENCELL, INC. | A61B005/00 A61B005/11 G06F019/00 A61B005/0205 A61B005/026 A61B005/021 A61B005/024 A61B005/0456 A61B005/08 A61B005/091
US09826912 | B2 | 2017-11-28 | Photoplethysmographic device for measuring a heart rhythm | 2016-10-19 | Photoplethysmographic device for measuring a heart rhythm, comprising a housing having a contacting surface and a control unit, a photoplethysmographic sensor mechanically attached to the contacting surface and connected to the control unit, wherein the control unit is arranged to receive a photoplethysmographic signal from the photoplethysmographic sensor, a fastening band having a first end and a second end, the first end and second end attached to the device for attaching the device to a subject with the contacting surface contacting the subject's skin, a pressure sensor attached to the contacting surface in close proximity to the photoplethysmographic sensor. The control unit is arranged determining whether the pressure from the pressure sensor is in a range suitable for photoplethysmographic measurement by the photoplethysmographic sensor. | 1. Photoplethysmographic device for measuring a heart rhythm, comprising a housing having a contacting surface and a control unit; a photoplethysmographic sensor mechanically mounted to the contacting surface and connected to the control unit, wherein the control unit is arranged to receive a photoplethysmographic signal from the photoplethysmographic sensor, wherein the control unit is arranged for determining a heart rhythm from the photoplethysmographic signal;a fastening band having a first end and a second end, the first end and second end attached to the device for attaching the device to a subject with the contacting surface contacting the subject's skin;a pressure sensor mounted to the contacting surface in close proximity to the photoplethysmographic sensor;wherein the control unit is arranged for determining whether the pressure from the pressure sensor is in a range suitable for photoplethysmographic measurement by the photoplethysmographic sensor, wherein the control unit is arranged for:determining a reliability factor for the PPG measurements from the pressure signal; anddetermining a heart rhythm using an actual photoplethysmographic measurement and a running history of measurements of the photoplethysmographic signal; andwherein a measurement of the photoplethysmographic signal is saved in the running history when the reliability factor is higher than a threshold. 2. Photoplethysmographic device according to claim 1, wherein the control unit is arranged for comparing a pressure signal strength from the pressure sensor with a first threshold value. 3. Photoplethysmographic device according to claim 1, wherein the control circuit is further arranged for comparing the pressure signal strength from the pressure sensor with a second threshold value, the second threshold value being higher than the first threshold value. 4. Photoplethysmographic device according to claim 3, further comprising an indicator connected to the control unit, wherein the control unit is arranged for activating the indicator when the pressure signal strength from the pressure sensor is equal to or exceeds the first threshold value and the pressure signal strength is equal to or less than the second threshold value. 5. Photoplethysmographic device according to claim 4, wherein the indicator is a sound alarm. 6. Photoplethysmographic device according to claim 1, wherein the photoplethysmographic sensor has a light source, and wherein the control unit is arranged for: comparing the pressure signal strength with a third threshold value; and wherein the control unit is arranged for deactivating the light source in the photoplethysmographic sensor when the pressure signal strength from the pressure sensor is equal to or less than the third threshold value. 7. Pressure sensor according to claim 6, wherein the control unit is arranged for deactivating the light source in the photoplethysmographic sensor when the pressure signal strength from the pressure sensor i... | LIVASSURED B.V. | A61B005/024 A61B005/00
US09826915 | B2 | 2017-11-28 | Systems and methods for assessing heart function | 2015-07-13 | Systems and methods can be used to provide an indication of heart function, such as an indication of mechanical function or hemodynamics of the heart, based on electrical data. For example, a method for assessing a function of the heart can include determining a time-based electrical characteristic for a plurality of points distributed across a spatial region of the heart. The plurality of points can be grouped into at least two subsets of points based on at least one of a spatial location for the plurality of points or the time-based electrical characteristics for the plurality of points. An indication of synchrony for the heart can be quantified based on relative analysis of the determined time-based electrical characteristic for each of the at least two subsets of points. | 1. A system, comprising: a computing device configured to store data and execute machine readable instructions, the data including electrical data representing measurements of electrical activity for a plurality of points across at least one spatial region of a patient's heart, wherein the plurality of points across the patient's heart is divided into a plurality of segments, each containing some of the plurality of points, the instructions to at least: compute at least one indication of synchrony based on analysis of the electrical data to provide an assessment of mechanical synchrony of the patient's heart based on electrical measurements by weighting each of the plurality of segments according to its respective contribution to mechanical heart function and/or hemodynamic performance;determine at least one therapy parameter based on the at least one indication of synchrony; anda therapy system to deliver therapy based on the at least one therapy parameter, the computing device to adjust the at least one therapy parameter in response to feedback corresponding to updated measurements of electrical activity. 2. The system of claim 1, wherein the instructions are further configured to determine which of a plurality of leads to use in delivering the therapy. 3. The system of claim 1, wherein the instructions are further configured to compute time-based electrical characteristics for the plurality of points based on the measurements of electrical activity during at least one time interval. 4. The system of claim 1, wherein the instructions of the computing device are further configured to determine the therapy parameter as including one or more lead placement locations to deliver the therapy. 5. The system of claim 1, wherein the instructions of the computing device are further configured to: compute the at least one indication of synchrony based on the electrical data that is acquired non-invasively via sensors distributed on the outer surface of the patient's body and invasively via one or more other sensors within the patient's body; andgenerate map data to display the at least one indication of synchrony. 6. A system, comprising: a computing device configured to store data and execute machine readable instructions, the data including electrical data representing measurements of electrical activity for a plurality of points across at least one spatial region of a patient's heart, the instructions to at least: compute at least one indication of synchrony based on analysis of the electrical data, the at least one indication of synchrony providing a quantitative assessment of at least one of heart electrical function, heart mechanical function or hemodynamic performance, the instructions to compute the at least one indication of synchrony further including instructions to; compute time-based electrical characteristics for the plurality of points based on the measurements of electrical activity during at least one time interval;calculate standard... | Cardioinsight Technologies, Inc. | A61B005/0402 A61B005/04 A61B005/00 A61N001/362 A61B005/0452 A61B005/02 A61B005/11 A61B005/042 A61N001/368
US09804243 | B2 | 2017-10-31 | Magnetic resonance imaging apparatus | 2014-04-23 | A magnetic resonance imaging apparatus of an embodiment includes: an imaging condition setting section configured to set an imaging condition for a main scan, a preparatory scan performing section configured to repeat a preparatory scan to be performed by the use of the imaging condition for a main scan and a changeably inputted center frequency so as to generate a preparatory image in real time, a display section on which the preparatory image can be displayed in real time, an optimum center frequency setting section configured to change the center frequency so as to change a situation in which a banding artifact appears and configured to set an optimum center frequency according to a user operation based on an observation of the preparatory image, and a main scan performing section configured to reconstruct an image for a diagnosis by using the optimum center frequency having been set. | 1. A magnetic resonance imaging (MRI) apparatus comprising: an MRI system including at least one processor configured toset an imaging condition for a main scan to obtain a diagnostic image;accept a user input center frequency changeably inputted in real time;repeat a preparatory scan performed by use of the set imaging condition and the center frequency changeably inputted in real time and generate preparatory images in real time visibly depicting a banding artifact using the preparatory scan data acquired in the repeated preparatory scans;display the preparatory images in real time;change the center frequency according to a user operation via a user interface so as to change where the banding artifact appears, the optimum center frequency being set according to a user operation based on observation of the preparatory images; andperform the main scan by using the set optimum center frequency and the set imaging condition for the main scan to reconstruct the diagnostic image. 2. The magnetic resonance imaging apparatus according to claim 1, wherein both the main scan and the preparatory scan are performed using an SSFP (Steady State Free Precession) method. 3. The magnetic resonance imaging apparatus according to claim 1, wherein the imaging condition set for the main scan includes information on an imaging cross section image which is to be captured in the main scan, andan imaging cross section which is same as a cross section to be used for the main scan is included in the preparatory scan. 4. The magnetic resonance imaging apparatus according to claim 3, wherein, when the imaging cross section in the main scan includes a plurality of slices, a middle one of the slices is included in the preparatory scan, and when the imaging cross section in the main scan is a slab having a certain thickness, one slice corresponding to a middle position of the slab is included in the preparatory scan. 5. The magnetic resonance imaging apparatus according to claim 1, wherein the preparatory scan is automatically started according to a user operation to start the main scan after the imaging condition is set, andthe main scan is automatically started when the optimum center frequency is set according to a user operation during the repeated preparatory scans. 6. The magnetic resonance imaging apparatus according to claim 1, wherein the set imaging condition for a main scan includes information on resolution in the main scan, andthe preparatory scan is performed in such a way that resolution is rendered the same as the resolution in the main scan. 7. The magnetic resonance imaging apparatus according to claim 6, wherein the preparatory scan acquires a signal only in a specific range centered at a phase encoding quantity of zero, and complements zero for a range excepting the specific range to perform reconstruction processing. 8. The magnetic resonance imaging apparatus according to claim 1, wherein information on an imaging cross section in a main scan to b... | TOSHIBA MEDICAL SYSTEMS CORPORATION | G01R033/56 A61B005/055 A61B005/00 G01R033/38
US09795338 | B2 | 2017-10-24 | Apparatus and method for detecting NIR fluorescence at sentinel lymph node | 2014-07-10 | The present invention relates to a device for observing a sentinel lymph node (SLN) in a human body. More particularly, the present invention relates to a device for observing an SLN by detecting near-infrared (NIR) fluorescence caused by a fluorescent material such as indocyanine green (ICG) at the SLN and a method for detecting NIR fluorescence at an SLN. Particularly, in the implementation of a composite image obtained by reproducing a fluorescent material such as ICG and NIR fluorescence emitted by excitation light together with a visible light image, it is possible to detect an SLN with high accuracy through a color contrast method and/or a temporal modulation method using an NIR fluorescence image signal and a visible reflection light image signal. | 1. A device for detecting near-infrared (NIR) fluorescence at a sentinel lymph node (SLN), the device comprising: a white light source configured to irradiate white light onto an object;a near-infrared (NIR) excitation light source configured to irradiate near-infrared (NIR) excitation light onto the object;an optical analyzing assembly configured to transmit white reflection light reflected off the object when the white light is irradiated on the object and near-infrared (NIR) fluorescence reflected off the object when the near-infrared (NIR) excitation light is irradiated on the object;a multispectral image processing unit configured to detect the white reflection light and the near-infrared (NIR) fluorescence, transmitted from the optical analyzing assembly, and process the white reflection light and the near-infrared (NIR) fluorescence as a visible (VIS) reflection light image signal and a near-infrared (NIR) fluorescence image signal, respectively; anda display unit configured to output a composite image obtained by combining the visible (VIS) reflection light image signal and the near-infrared (NIR) fluorescence image signal, processed by the multispectral image processing unit,wherein the multispectral image processing unit splits the visible (VIS) reflection light image signal into red (R), green (G) and blue (B) image signals, and performs image processing so that the visible (VIS) reflection light image signal is implemented with red (R) and green (G) in a pixel where the near-infrared (NIR) fluorescence image signal is not detected, and the near-infrared (NIR) fluorescent image signal is displayed with only blue (B) in a pixel where the near-infrared (NIR) fluorescent image signal is detected. 2. The device of claim 1, wherein the multispectral image processing unit includes a beam splitter configured to form a visible (VIS) channel and a near-infrared (NIR) channel by splitting a light path of the white reflection light and the near-infrared (NIR) fluorescence, transmitted from the optical analyzing assembly. 3. The device of claim 2, wherein the beam splitter is a beam splitting prism that selectively transmits or reflects light in a selected wavelength range. 4. The device of claim 3, wherein the beam splitter is a dichroic prism. 5. The device of claim 2, wherein the multispectral image processing unit is configured to have two image sensors respectively disposed on the light path split by the beam splitter. 6. The device of claim 5, wherein the two image sensors are a color image sensor disposed on the visible (VIS) channel and a monochrome image sensor disposed on the near-infrared (NIR) channel. 7. The device of claim 6, wherein a first optical filter is disposed between the monochrome image sensor and the beam splitter and a second optical filter is disposed between the color image sensor and the beam splitter. 8. The device of claim 5, wherein a gain amplifier and an analog/digital (A/D) converter are connected to each ... | INTHESMART CO. LTD. | A61B005/00 A61B001/00 A61B001/04 A61B001/06
US09802481 | B2 | 2017-10-31 | Vehicle driver monitoring system | 2016-08-26 | A vehicle driver monitoring system includes a supervisor computerized device and an operator computerized device. The system may also include a key container device. The devices of the vehicle driver monitoring system work in conjunction with each other to monitor various aspects of the user's operation of a corresponding vehicle during use. If a key container device is used, the devices also work to limit the operability of a vehicle key. For example, the supervisor computerized device is configured to remotely receive vehicle monitoring information, such as location information, speed information, and operator information, from the operator computerized device as the operator drives the vehicle. Further, the optional key container device is configured to selectively allow to the operability of a vehicle key based upon vehicle access info, identity and sobriety information as related to the user and as provided by the operator computerized device. | 1. A vehicle driver monitoring system, comprising: a proximity sensor;an operator computerized device comprising a processor; and a memory storing instructions that, when executed by the processor, cause the operator computerized device to perform the operation of:receiving a signal from the proximity sensor indicating that the operator is within a pre-defined proximity of the vehicle;providing an alert to an operator indicating the operator is permitted to operate a vehicle in response to receiving and testing operator identification information and information regarding the level of an ingestible substance; anda supervisor computerized device comprising a processor; and a memory storing instructions that, when executed by the processor, cause the supervisor computerized device to perform the operation of:receiving and saving vehicle monitoring information sent by the operator computerized device. 2. The vehicle driver monitoring system of claim 1, further comprising: a key disabling device;wherein the key disabling device is configured to selectively prevent use of a vehicle key at least partly contained therein, wherein a functional part of the key is accessible to the operator while use of the key is prevented;wherein the key disabling device is further configured to allow use of the vehicle key in response to receiving a signal from the operator computerized device, wherein the signal is based on the operator identification information and information regarding the level of an ingestible substance. 3. The system of claim 2, wherein the key comprises a wireless key. 4. The system of claim 1, wherein the ingestible substance is alcohol. 5. The system of claim 1, wherein the operator computerized device is configured to alert a supervisor when the level of the ingestible substance is outside of the pre-defined limits. 6. The system of claim 5, wherein the operator computerized device is configured to display to the operator information about a ride service when the level of the ingestible substance is outside the pre-defined limits. 7. The system of claim 1, wherein the operator computerized device is further configured to: compare time stamps of the substance test and the identification information test; andalert a supervisor and signal the operator to not drive when the time stamps are not reasonably contemporaneous. 8. The system of claim 1, wherein the operator identification information comprises an image of the operator's face. 9. The system of claim 1, wherein the operator computerized device is further configured to: receive a vehicle operation restriction, wherein the vehicle operation restriction comprises one or more of the following: a time period and a geographic area; andnotify the supervisor when the vehicle operation restriction is violated. 10. The system of claim 1, wherein the supervisor computerized device is further configured to: provide the vehicle monitoring information for use in calculating an insurance discou... | - | B60K028/06 A61B005/18 A61B005/00
US09788756 | B2 | 2017-10-17 | Systems, devices, methods, and computer-readable storage facilitating locating an implantable medical device within a body | 2015-04-20 | Techniques facilitating locating an implantable medical device (IMD) within a body of a patient are provided. An estimate of the location is determined based on strength information representative of strengths of communicative couplings between a communications head device and the IMD at various positions of the communications head device. The strength information can be updated periodically or based on specific events, such as changes to the body and/or an amount of time elapsed since strength information was previously obtained. Media representative of the estimate of the location and an image of the body can be output and can facilitate a patient or caregiver locating the IMD. In some embodiments, the media and image output can guide future placement of the communications head device on the patient to efficiently establish communication. Further, in some embodiments, numerous IMDs implanted within a single body can be identified and communication can commence. | 1. A device, comprising: a memory that stores executable modules;a processor coupled to the memory and configured to execute the following executable modules stored in the memory: a map composition module configured to generate information indicative of a first electronic map, wherein the first electronic map is representative of respective strengths of communicative couplings between a communications head device and an implantable medical device within a body at multiple positions of the communications head device relative to the body and during a first time period;a location prediction module configured to estimate an implant location of the implantable medical device within the body based on the information indicative of the first electronic map; anda media composition module configured to generate imaging information representative of an estimated implant location and at least a portion of the body. 2. The device of claim 1, further comprising: a display device configured to display indicia representative of at least a portion of the imaging information. 3. The device of claim 2, wherein the indicia comprises media including one or more of a video segment, a photograph, or schematic image. 4. The device of claim 3, wherein the display device is further configured to display a prompt to place the communications head device in a defined position relative to the body. 5. The device of claim 1, wherein the location prediction module is further configured to determine a location within the first electronic map at which one of the strengths of the communicative couplings is indicative of a strength intensity greater than a defined threshold. 6. The device of claim 1, wherein the location prediction module is further configured to predict a location within the first electronic map at which one of the strengths of the communicative couplings is indicative of maximal strength intensity. 7. The device of claim 1, wherein the map composition module is further configured to generate information indicative of a second electronic map, wherein the second electronic map is representative of respective strengths of communicative couplings between the communications head device and the implantable medical device within the body at multiple second locations of the communications head device relative to the body and during a second defined time period. 8. The device of claim 7, wherein the map composition module is further configured to generate the information indicative of the second electronic map based on a change in a size of the body, a time interval that has elapsed between commencement of the first defined time period and commencement of the second defined time period, or a defined change in a strength of a communicative coupling at a defined location of the body relative to an initial strength of an initial communicative coupling at the defined location of the body as indicated in the first electronic map. 9. The device of claim 7, wherein the l... | Medtronic, Inc. | A61N001/00 A61B005/06 A61N001/37 A61B005/00 A61N001/372
US09775682 | B2 | 2017-10-03 | Teleoperation system with visual indicator and method of use during surgical procedures | 2016-04-18 | Teleoperation systems and methods for use during a surgical procedure. The teleoperation system comprises a surgical tool and a master input device for being moved by a user. A slave device is coupled to the surgical tool and responsive to movement of the master input device. The slave device is configured to be operated in a first mode or a second mode. In the first mode, the surgical tool is placed in a desired pose with respect to a target region in response to a manual force applied to the slave device. In the second mode, the surgical tool is manipulated during the surgical procedure in response to movement of the master input device by the user. A visual indicator is coupled to the surgical tool or the slave device to emit light in the second mode to provide visual information to the user in the second mode. | 1. A teleoperation system for use during a surgical procedure, said teleoperation system comprising: a surgical tool;a master input device to be moved by a user;a slave device coupled to said surgical tool and responsive to movement of said master input device, said slave device configured to be operated in a first mode to place said surgical tool in a desired pose with respect to a target region and a second mode to manipulate said surgical tool during the surgical procedure; anda visual indicator coupled to and disposed on said slave device and configured to emit light in said second mode to provide visual information to the user in said second mode. 2. The teleoperation system of claim 1, wherein said visual indicator is configured to emit light in a plurality of different colors in said second mode to provide visual information to the user in said second mode. 3. The teleoperation system of claim 2, wherein said visual indicator comprises a multi-colored LED indicator. 4. The teleoperation system of claim 3, further comprising a display configured to display said surgical tool. 5. The teleoperation system of claim 1, further comprising a control system configured to lock said slave device in a desired pose. 6. The teleoperation system of claim 5, wherein said control system is configured to adjust said desired pose of said slave device in response to receiving a manual force on said slave device. 7. The teleoperation system of claim 5, wherein said control system is configured to lock only a portion of said slave device so that the user can adjust said slave device. 8. The teleoperation system of claim 1, further comprising a control system configured to constrain a trajectory of a tool guide passing through the target region. 9. The teleoperation system of claim 8, wherein said control system is configured to allow the user to pivot said tool guide about a pivot point while constraining said trajectory. 10. A teleoperation system for use during a surgical procedure, said teleoperation system comprising: a surgical tool;a master input device to be moved by a user;a slave device coupled to said surgical tool and responsive to movement of said master input device, said slave device configured to be operated in a first mode to place said surgical tool in a desired pose with respect to a target region and a second mode to manipulate said surgical tool during the surgical procedure; anda visual indicator coupled to and disposed on said surgical tool and configured to emit light in said second mode to provide visual information to the user in said second mode. 11. The teleoperation system of claim 10, wherein said visual indicator is configured to emit light in a plurality of different colors in said second mode to provide visual information to the user in said second mode. 12. The teleoperation system of claim 11, wherein said visual indicator comprises a multi-colored LED indicator. 13. The teleoperation system of claim 12, further co... | MAKO SURGICAL CORP. | G06F019/00 A61B034/30 A61F002/30 G06F003/01 A61B017/16 A61B034/37 A61B034/35 A61N001/05 A61N001/36 A61N001/372 A61F002/46 A61B005/11 A61B017/17 A61B017/00 A61F002/38 A61B005/00 A61B090/00 A61B034/10 A61B034/20 A61B034/00 A61B090/14
US09808179 | B2 | 2017-11-07 | Catheter and a system and a method for locating a catheter in a vessel | 2012-08-03 | A balloon catheter (60) includes an elongated catheter member (63) and an inflatable balloon (65) located adjacent a distal end (5) of the catheter member (63). Band type measuring electrodes (7) located on the catheter member (63) within the balloon (65) include a pair of outer stimulating electrodes (8) and sensing electrodes (9) for measuring the cross-sectional area and volume of the balloon (65) by impedance planimetry measuring. A light emitting diode (20) is located on each measuring electrode (7) for emitting light so that the location of the balloon catheter, and in particular, the measuring electrodes (7) can be detected in a vessel (91) with a translucent wall within the body of a human subject. By capturing an image of the vessel and the light from the light emitting diodes (20) emanating through a translucent wall of the vessel by an imaging device (92) located within the human body externally of the vessel (91), the location of the balloon catheter and the electrodes (7) within the vessel can be determined. | 1. A catheter comprising: an elongated catheter member,a plurality of spaced apart band electrodes located on the catheter member, andan indicating means comprising a light source adapted to output a detectable light signal, and a light outputting means provided by a light accommodating bore extending through at least one of the band electrodes for accommodating the detectable light signal therethrough for identifying the location of the corresponding band electrode. 2. A catheter as claimed in claim 1 in which the signal outputted by the indicating means is adapted to be receivable by an imaging device of an imaging system. 3. A catheter as claimed in claim 1 in which a corresponding light source is provided for each light accommodating bore. 4. A catheter as claimed in claim 1 further comprising an inflatable balloon mounted on the catheter member, and a communicating means being provided for communicating the balloon with an inflating medium source, at least one of the band electrodes being located on the balloon on an outer surface, thereof and having one of the light accommodating bores extending therethrough. 5. A catheter as claimed in claim 4 in which the balloon is located on the catheter member with the electrodes located on the catheter member within the balloon. 6. A catheter as claimed in claim 1 in which each electrode is provided with a corresponding light accommodating bore. 7. A catheter as claimed in claim 1 in which the light source comprises a light emitting diode. 8. A catheter as claimed in claim 1 in which a plurality of light sources are provided, and at least some of the light sources produce light of respective different wavelengths. 9. A system for locating a catheter in a vessel or lumen in a human or animal subject, the system comprising: a catheter locating in the vessel or lumen, the catheter comprising: an elongated catheter member,a plurality of spaced apart electrodes located on the catheter member, andan indicating means comprising a light source adapted to output a detectable light signal, and a light outputting means provided by a light accommodating bore extending through at least one of the electrodes for accommodating the detectable light signal therethrough for identifying the location of the corresponding electrode, andan imaging device adapted for locating in the body of the subject adjacent the vessel or lumen but exteriorly thereof, the imaging device being adapted for capturing an image of the vessel or lumen and the detectable light signal outputted through the light accommodating bore of the at least one electrode of the catheter member located within the vessel or lumen. 10. A system as claimed in claim 9 in which the imaging device comprises a charge coupled device (CCD) imaging chip. 11. A system as claimed in claim 9 in which the imaging device comprises a video camera. 12. A system as claimed in claim 9 in which the imaging device is adapted for use in a laparoscopic procedure. 13. ... | Flip Technologies Limited | A61B005/04 A61B005/06 A61B005/00 A61B001/05 A61B001/06 A61B001/07 A61B001/313 A61B005/053 A61B005/107 A61B005/042 A61M025/01 A61M025/10 A61B018/14 A61B018/00 A61B034/20
USRE046567 | E1 | 2017-10-10 | Magnetic resonance imaging apparatus | 2015-12-31 | Provided is an MRI apparatus. In the MRI apparatus, a data collection unit repetitively performs a tag mode of applying an RF wave to at least an upstream portion of an imaging area to perform fluid labeling of a fluid flown into the imaging area and, after a lapse of an inversion time from application of the RF wave, performing magnetic resonance data collection, while changing the inversion time. An image reconstruction unit reconstructs a plurality of tag images corresponding to a plurality of different inversion times based on the magnetic resonance data collected in the tag mode. A reference image generation unit generates a reference image based on the plurality of the tag images. A fluid image generation unit generates a subtraction image between each of the tag images and the reference image as a fluid image. | 1. A magnetic resonance imaging apparatus comprising: a memory; anda processor comprising processing circuitry, wherein the memory stores instructions, which when executed by the processor, cause the magnetic resonance imaging apparatus to perform operations comprising:a data collection unit that repetitively performsperforming a tag mode of applying an RF wave to at least an upstream portion of an imaging area to label a fluid flown into the imaging area and, after a lapse of an inversion time from application of the RF wave, performing magnetic resonance data collection, while changing the inversion time;an image reconstruction unit that reconstructsreconstructing a plurality of tag images corresponding to a plurality of different inversion times based on magnetic resonance data collected in the tag mode;a reference image generation unit that generatesgenerating a reference image based on the plurality of the tag images; anda fluid image generation unit that generatesgenerating a subtraction image between each of the tag images and the reference image as a fluid image. 2. The magnetic resonance imaging apparatus according to claim 1, wherein the reference image generation unit generates instructions, when executed by the processor, cause the magnetic resonance imaging apparatus to generate the reference image by using the tag image exhibiting a small difference in signal intensity between a blood flow portion and a background tissue among the plurality of the tag images. 3. The magnetic resonance imaging apparatus according to claim 1, further comprisingwherein the instructions, when executed by the processor, cause the magnetic resonance imaging apparatus to: a display unit that displaysdisplay a plurality of the fluid images generated by the fluid image generation unit as behavior expression in a cine display format or a parallel display format. 4. The magnetic resonance imaging apparatus according to claim 1, wherein at the time of performing the tag mode, the data collection unit appliesthe instructions, when executed by the processor, cause the magnetic resonance imaging apparatus to apply a non-selective inversion recovery pulse to a range including the imaging area and the fluid flown into the imaging area so that the magnetic resonance data collection is started at the time when a longitudinal magnetization of at least one type of tissue among tissues included in the imaging area becomes nearly zero. 5. The magnetic resonance imaging apparatus according to claim 1, wherein the data collection unit appliesinstructions, when executed by the processor, cause the magnetic resonance imaging apparatus to apply a saturation pulse to a range including the imaging area just before the magnetic resonance data collection is started. 6. The magnetic resonance imaging apparatus according to claim 1, wherein the reference image generation unit selectsinstructions, when executed by the processor, cause the magnetic resonance imaging apparatus to ... | Toshiba Medical Systems Corporation | G01V003/00 A61B005/055 G01R033/563 A61B005/026 G01R033/56
US09788724 | B2 | 2017-10-17 | Patient monitor screen aggregation | 2014-03-14 | A remote interface system according to some embodiments includes one or more patient monitoring devices having one or more sensors for patient monitoring, one or more first processors to receive information from the sensors and generate patient data based thereon, a web server, a first communication system, and a first screen to display a representation of at least a first portion of the patient data, and a remote interface device including a second communication system, a second screen, one or more second processors, an application executed by the one or more second processors and configured to detect a presence of the first communication system, establish a communication link between the first and second communication systems, establish a reliable connection channel (e.g. secure websocket connection) with the web server, receive the patient data via the connection, and display at least a second portion of the patient data on the second screen. | 1. A remote interface system for medical application, the remote interface system comprising: a patient monitoring device comprising: one or more sensors configured for physical attachment to a patient;a first processor configured to receive information from the one or more sensors and to generate patient data based on the information;a web server;a first communication system; and a first screen configured to display a representation of at least a first portion of the patient data; anda remote interface device comprising: a second communication system; a second screen;a second processor;an application executed by the second processor, the application configured to detect a presence of the first communication system, establish a communication link between the first and second communication systems, establish a websocket connection with the web server, send a first request to the web server for the patient data, receive a plurality of responses representing the patient data in response to the first request via the websocket connection, wherein the websocket connection is secure, receive the patient data via the websocket connection, and display at least a second portion of the patient data on the second screen;wherein the patient monitoring device is an external defibrillator,wherein the application is configured to receive a first alarm from the patient monitoring device and initiate a second alarm at the remote interface system based on the first alarm,wherein the application is further configured to accept a user command to suspend the second alarm for a period of time, and wherein the period of time is configurable by a user. 2. The remote interface system of claim 1, wherein the communication link is a direct communication link that does not travel over any wires outside of the patient monitoring device and the remote interface device. 3. The remote interface system of claim 1, wherein at least a portion of the communication link includes a wired network. 4. The remote interface system of claim 1, wherein the patient monitoring device is portable. 5. The remote interface system of claim 4, wherein the remote interface device is portable. 6. The remote interface system of claim 1, wherein the application is further configured to receive the patient data via the websocket connection in packets each representing 120 milliseconds of the patient data. 7. The remote interface system of claim 6, wherein the web server sends a new packet of the patient data to the remote interface device at least once every 120 milliseconds. 8. The remote interface system of claim 1, wherein the application is further configured to receive the patient data via the websocket connection in packets each representing between 40 and 200 milliseconds of the patient data. 9. The remote interface system of claim 1, wherein packets of the patient data sent by the web server to the remote interface device include ECG waveform data from the patient. 10. The remote inter... | ZOLL Medical Corporation | A61B005/00 G06F019/00 A61N001/39
US09804836 | B2 | 2017-10-31 | Medical device nucleus architecture | 2014-11-07 | A medical software package stored on a non-transitory, computer-readable storage medium is disclosed. The medical software package can include a plurality of software components programmed to provide one or more functionalities to one or more medical monitoring devices, where each software component is programmed to run independently as a daemon, where each software component is programmed to be modified without recompiling unchanged software components, where the software components are programmed to communicate to each other through at least one of the software components, and where the software package is hardware agnostic and operating system agnostic. A medical device including a tangible, non-transitory storage device storing a medical software package comprising a plurality of modular software components programmed to provide medical device functionality is also disclosed. | 1. A medical software package stored on a non-transitory, computer-readable storage medium, comprising: a plurality of software components programmed to provide one or more functionalities to one or more medical monitoring devices;wherein each of the plurality of software components is programmed to run independently as a daemon;wherein each of the plurality of software components is programmed to be modified without recompiling unchanged software components;wherein the plurality of software components is programmed to communicate to each other through at least one of the plurality of software components;wherein the medical software package is hardware agnostic and operating system agnostic;wherein the plurality of software components further comprise a plurality of reusable assets;wherein one of the plurality of reusable assets is a workflow manager programmed to store, add, remove, start, and stop manager programs resident in the non-transitory, computer-readable storage medium;wherein one of the plurality of reusable assets is a capability registry programmed to provide a bulletin board system where software components can post common object-oriented architecture messages and receive messages about changes to the posted common object-oriented architecture messages;wherein one of the plurality of reusable assets is a message manager programmed to provide an inter-process communications bus for sending common object-oriented architecture messages between the plurality of software components;wherein at least one of the posted common object-oriented architecture messages that is sent through the message manager includes a unique identifier and persists on the bulletin board system until another common object-oriented architecture message with the same unique identifier is posted to the bulletin board system;wherein one of the plurality of reusable assets is a configuration manager programmed to maintain persistent configuration of the one or more medical monitoring devices and provide a device configuration to a plurality of client software components;wherein one of the plurality of reusable assets is a component manager programmed to start up and shut down at least one of the plurality of software components;wherein the plurality of reusable assets form a single architecture that includes a set of portable and reusable software items designed to provide a base set of functionality for the one or more medical monitoring devices;wherein one of the plurality of reusable assets is a virtual operating system programmed to provide a fixed application programming interface between the single architecture and an underlying operating system;wherein one of the plurality of reusable assets is a virtual hardware programmed to provide a fixed application programming interface between the single architecture and underlying hardware and device drivers; andwherein the single architecture includes the workflow manager, the capability registry, the message manager... | Welch Allyn, Inc. | G06F009/44 A61B005/00 G06F019/00 G06F009/54 A61B005/01 A61B005/021 A61B005/024 A61B005/145
US09808590 | B2 | 2017-11-07 | Method and device for carrying out a signal-processing viewing of a measurement signal that is correlated to the respiratory activity of an individual | 2014-04-09 | The invention relates to a method and device for carrying out a signal-processing viewing of a measurement signal that is correlated with the respiratory activity of an individual, for example, of a measurement signal that is correlated with the respiratory gas. The aim of the invention is to provide solutions with which an improved electronic analysis of the signals that are representative with regard to respiratory activity can be achieved. To this end, the invention provides that viewing results are obtained within the scope of a signal-processing viewing of said measurement signal and make a differentiation between obstructive and central respiratory disorders possible. The viewing results are determined, in particular, while taking into account changes of selected breathing characteristics such as, for example, the change in the ratio of inhalation time to exhalation time. | 1. A method of controlling pressure provided to a patient by a pressure regulating device, the method comprising: providing a supply of respiratory gas at a first pressure level above ambient pressure level;measuring a signal indicative of respiratory activity;determining, using a signal processing device, an inspiration phase and an expiration phase of each breath from the signal;detecting, using the signal processing device, an inspiration time and an expiration time of each breath for successive breaths;determining, using the signal processing device, a ratio of the inspiration time to the expiration time for each said breath;detecting, using the signal processing device, a variation in the ratio;determining, using the signal processing device, presence of obstruction of the upper respiratory tract based on the in the ratio;and adjusting the first pressure level based on the determined presence of obstruction. 2. The method according to claim 1, wherein the first pressure level is increased based on the determined presence of obstruction of the upper respiratory tract. 3. The method according to claim 1, wherein the presence of obstruction of the upper respiratory tract is determined based on comparative analysis of successively occurring changes in the ratio of the inspiration time to the expiration time. 4. The method according to claim 1, wherein the signal indicative of respiratory activity is a flow signal. 5. The method of claim 4, further comprising determining a first derivative of flow from the flow signal. 6. The method of claim 5, wherein the first derivative of flow is determined for a region of phase change in each breath. 7. The method of claim 6, wherein the determining of the presence of obstruction of the upper respiratory tract includes determining changes in the first derivative of flow over successively occurring breaths. 8. The method according to claim 6, wherein the region of phase change is directed to at least one of the beginning of the inspiration phase and the end of the inspiration phase. 9. The method of claim 5, further comprising analyzing a curve shape during the inspiration phase. 10. The method of claim 9, further comprising calculating an average gradient in simple form for intervals that extend over a predetermined percentage of a time duration of the inspiration phase. 11. The method of claim 10, further comprising variably calculating a gradient at a phase change within a window over the inspiration phase. 12. The method according to claim 10, wherein the predetermined percentage is at least 50%. 13. The method according to claim 10, wherein the predetermined percentage is 80%. 14. The method according to claim 13 wherein the curve shape during the inspiration phase is analysed in an interval between 10% after the beginning of the inspiration phase and 10% before the end of the inspiration phase. 15. The method according to claim 9, wherein the analysis of the curve shape includes determinin... | RESMED R&D GERMANY GMBH | A61B005/08 A61M016/00 A61B005/087 A61F005/56 A61B005/085 A61B005/00
US09782092 | B2 | 2017-10-10 | Patient signal analysis based on vector analysis | 2014-11-14 | Disclosed herein is a framework for facilitating patient signal analysis based on vector analysis. In accordance with one aspect, a set of vectors is generated from a patient signal data waveform. The vectors may be directed from a common center to points of interest on the patient signal data waveform. The framework may further extract one or more vector parameters from the set of vectors, and determine one or more vector ratios based on the vector parameters to monitor changes in the patient signal data waveform. | 1. A computer-implemented method of patient signal analysis, comprising: receiving, from a patient monitor, a patient signal data waveform;generating, by a processor, a set of vectors directed from a common center to points of interest on the patient signal data waveform;extracting, by the processor, one or more vector parameters from the set of vectors; anddetermining, by the processor, one or more vector ratios based on the one or more vector parameters to monitor changes in the patient signal data waveform, wherein determining the one or more vector ratios comprises determining one or more unilateral ratios from the vectors or determining one or more bilateral ratios from the vectors. 2. The method of claim 1 wherein the patient signal data waveform comprises an electrocardiogram waveform. 3. The method of claim 1 wherein generating the vectors comprises generating the vectors directed from an R wave peak center point on the patient signal data waveform. 4. The method of claim 1 wherein generating the vectors comprises generating the vectors directed from the common center point to at least one of P, Q, S, T or U wave peak points on the patient signal data waveform. 5. The method of claim 1 wherein generating the vectors comprises generating the vectors directed from the common center point to an H point defined at a half cycle time point. 6. The method of claim 1 wherein extracting the one or more vector parameters comprises extracting one or more time durations measured from a half cycle time stamp to time stamps of the points of interest. 7. The method of claim 6 wherein determining the one or more vector ratios comprises determining at least one ratio of distances associated with two vectors selected from the set of vectors. 8. The method of claim 6 wherein determining the one or more vector ratios comprises determining at least one ratio of time durations associated with two vectors selected from the set of vectors. 9. The method of claim 6 wherein determining the one or more vector ratios comprises determining at least one ratio comparing a vector amplitude associated with a first vector to a time duration associated with a second vector, wherein the first and second vectors are selected from the set of vectors. 10. The method of claim 9 wherein the first and second vectors are different. 11. The method of claim 9 wherein the first and second vectors are same. 12. The method of claim 1 wherein extracting the one or more vector parameters comprises extracting vector circles, wherein each of the vector circles is centered at the common center and has a radius defined by one of the vectors. 13. The method of claim 12 wherein determining the one or more vector ratios comprises determining at least one ratio of areas associated with two of the vector circles. 14. The method of claim 12 wherein determining the one or more vector ratios comprises determining at least one ratio that compares an area associated with a vector circle o... | SIEMENS MEDICAL SOLUTIONS USA, INC. | A61N001/00 A61B005/04 A61B005/00 A61B005/0456 A61B005/046 A61B005/0464 A61N001/365
US09826935 | B2 | 2017-11-28 | Interchangeable cases for biometric monitoring devices | 2016-10-14 | A biometric monitoring device and multiple carrying cases for same are provided. In some implementations, the case may be made from a flexible viscoelastic material and the biometric monitoring device may be slipped into a receptacle in the case through an opening in the case; the opening may become distended during the insertion of the biometric monitoring device. In some implementations, the case may feature a display window that, in combination with materials of the biometric monitoring device, may mask a display of the biometric monitoring device from view when the display is off and may allow the display to be seen when the display is displaying content. | 1. An apparatus comprising: a housing;a plurality of biometric sensors, the plurality of biometric sensors including a magnetic sensor; andprocessing circuitry, wherein: the plurality of biometric sensors and the processing circuitry are located within the housing, andthe processing circuitry is configured to: determine one or more activity-related metrics based on data from the plurality of biometric sensors, andswitch from a first operating mode to a second operating mode when data from the magnetic sensor corresponds with a first magnetic signature associated with a first mounting fixture, wherein the first mounting fixture includes a magnetic source that produces the first magnetic signature. 2. The apparatus of claim 1, wherein the magnetic sensor is a Hall effect sensor. 3. The apparatus of claim 1, wherein the magnetic sensor is a magnetometer. 4. The apparatus of claim 1, wherein the processing circuitry is further configured to use at least one algorithm to process data from the plurality of biometric sensors while in the second operating mode that is different from one or more algorithms that the processing circuitry uses to process data from the plurality of biometric sensors while in the first operating mode. 5. The apparatus of claim 4, wherein the first mounting fixture is configured for use with a bicycle and the processing circuitry is further configured to process data from the plurality of biometric sensors to determine bicycle-related activity metrics when in the second operating mode and the bicycle-related activity metrics include at least one activity-related metric that is different from the activity-related metrics determined by the processing circuitry in the first operating mode. 6. The apparatus of claim 4, wherein the first mounting fixture is part of a strap that allows the housing to be worn on a person's arm and the processing circuitry is further configured to process data from the plurality of biometric sensors to determine sleep-related activity metrics when in the second operating mode and the sleep-related activity metrics include at least one activity-related metric that is different from the activity-related metrics determined by the processing circuitry in the first operating mode. 7. The apparatus of claim 1, further comprising a display, wherein the processing circuitry is further configured to change the orientation of content shown on the display in the second operating mode as compared with the orientation of content shown on the display in the first operating mode. 8. The apparatus of claim 7, wherein the first mounting fixture is part of a wristband case that includes a magnet. 9. The apparatus of claim 8, wherein the wristband case includes injection-molded portions including: a molded body;a first molded strap extending away from a first end of the molded body;a second molded strap extending away from a second end of the molded body opposite the first end, wherein the molded body, the first ... | Fitbit, Inc. | G06F003/00 A61B005/00 A45C011/00 A61B005/11 A61B007/00 A61B005/0205 G06F015/00 B65D025/22 G01C005/06 B42D015/00 G09F003/00 G09F003/14 G09F003/04 A45C013/42 A61B005/05 A61B005/0408 A61B005/021 A61B005/024 A61B005/08 A61B005/145 G01C022/00
US09808604 | B2 | 2017-11-07 | Guide wire | 2015-09-24 | A guide wire includes a wire main body that has a plate-like portion in a distal end portion of the wire main body and a reinforcement portion that is provided on at least one surface of the plate-like portion and reinforces the plate-like portion. The reinforcement portion is made up of two wire rods, and both end portions of the reinforcement portion are fixed to the plate-like portion. The plate-like portion is formed in a ribbon shape, and the reinforcement portions are diagonally provided on the same plane of the plate-like portion, and intersect each other. The guide wire has a coil that covers the plate-like portion and the reinforcement portion. | 1. A guide wire comprising: a wire main body including a plate-like portion in a distal end portion of the wire main body, the plate-like portion having a distal-most end and a proximal-most end; anda reinforcement portion provided on at least one surface of the plate-like portion, and which reinforces the plate-like portion;wherein the reinforcement portion extends from the distal-most end to the proximal-most end of the plate-like portion;wherein the reinforcement portion includes two wire rods;wherein the plate-like portion defines a first planar surface and a second planar surface; andwherein each of the two wire rods is provided on the first planar surface of the plate-like portion, and the two wire rods intersect each other at a center portion of the plate-like portion. 2. The guide wire according to claim 1, wherein the reinforcement portion includes opposing end portions and the reinforcement portion is fixed to the plate-like portion at only the opposing end portions. 3. The guide wire according to claim 1, wherein the plate-like portion is formed in a ribbon shape, andwherein the reinforcement portion is provided on a diagonal line of the plate-like portion. 4. The guide wire according to claim 1, wherein rigidity of the reinforcement portion is the same as that of the plate-like portion. 5. The guide wire according to claim 1, wherein rigidity of the reinforcement portion is different from that of the plate-like portion. 6. The guide wire according to claim 1, wherein the reinforcement portion is configured as a member that is separate from the plate-like portion. 7. The guide wire according to claim 1, wherein the reinforcement portion is formed integrally with the plate-like portion. 8. The guide wire according to claim 1, further comprising: a coil that covers the plate-like portion and the reinforcement portion. 9. A guide wire comprising: a wire main body including a first wire disposed on a distal end side of the wire main body and a second wire disposed on a proximal end side of the wire main body, the first wire being joined to the second wire, and the first wire including a planar portion having a first planar surface and a second planar surface; anda reinforcement portion provided on at least one of the first planar surface and the second planar surface of the planar portion;wherein the reinforcement portion includes a first wire rod and a second wire rod;wherein the first wire rod and second wire rod each include a first terminal end portion, a second terminal end portion and a rod body portion extending therebetween,wherein the first wire rod and second wire rod are secured to the at least one of the first planar surface and the second planar surface of the planar portion only at the first terminal end portion and the second terminal end portion such that the rod body portion is not secured to the at least one of the first planar surface and the second planar surface of the planar portion, the rod body portion thus ... | Terumo Kabushiki Kaisha | A61B005/00 A61M025/00 A61M025/09
US09805577 | B2 | 2017-10-31 | Motion sensing necklace system | 2013-11-05 | A motion sensing necklace comprising includes a pendant attached to lower ends of two lanyards. The pendant includes a help button for a user to generate a signal for help, and a battery to power the motion sensing necklace. A neck strap is attached to upper ends of the first and second lanyards. Two motion sensors are located within opposing ends of the neck strap. A microprocessor located within the neck strap processes signals from the motion sensors to determine whether motion has occurred. A transmitter or transceiver located within the neck strap transmits motion detection signals and alerts generated by the microprocessor. | 1. A motion sensing necklace system comprising: at least two motion sensors coupled to a necklace, wherein when the necklace is worn by a user, the motion sensors are configured to be located on opposed and offset sides of a neck of the user, and the motion sensors are configured to sense respective different left and right side motion information about a particular motion of the user over a first duration and, in response, generate respective different time-varying left and right side motion profiles, wherein each of the motion profiles includes motion information corresponding to multiple different instances over the first duration, and wherein the particular motion of the user includes a weight shift from one side of the user's body to the other;a microprocessor configured to receive the respective different time-varying left and right side motion profiles, process the received motion profiles together to determine a comprehensive movement profile associated with the user wherein the comprehensive movement profile corresponds to the same particular motion of the user over the first duration that includes the weight shift from one side of the user's body to the other, and compare subsequent sensor information with the comprehensive movement profile to provide a comparison result, wherein the subsequent sensor information includes signals later received from the same motion sensors and includes motion information corresponding to multiple different instances over a different second duration; anda transmitter configured to transmit signals or alerts generated by the microprocessor and based on the comparison result. 2. The motion sensing necklace system of claim 1, further comprising a pendant attached to lower ends of first and second lanyards, the pendant including a help button and a battery; and a neck strap attached to upper ends of the first and second lanyards, wherein the at least two motion sensors are located at opposing ends of the neck strap. 3. The motion sensing necklace system of claim 2, wherein the microprocessor and the transmitter are located within the neck strap between the at least two motion sensors. 4. The motion sensing necklace system of claim 2, further comprising conductive wiring extending through the first lanyard between the pendant and the neck strap, such that the help button and battery are conductively connected to one or more of the motion sensor, microprocessor and transmitter. 5. The motion sensing necklace system of claim 4, further comprising a breakaway safety clasp between the second lanyard and the pendant, wherein the breakaway safety clasp triggers an alarm when the clasp is disconnected. 6. The motion sensing necklace system of claim 1, wherein the motion sensors are selected from a group consisting of accelerometers, gyroscopes, magnetometers and altimeters. 7. The motion sensing necklace system of claim 1, further comprising biometric sensors that are coupled to the necklace. 8. The motion se... | Nortek Security & Control, LLC | G08B021/04 A61B005/00 A61B005/11
US09788754 | B2 | 2017-10-17 | Method of managing a brain stroke | 2014-05-16 | The present invention relates to a medical method of managing a cerebrovascular insult, the method comprising the following steps in this order: a) placing a patient on a patient support unit;b) positioning a mobile tomographic imaging system in a predetermined position relative to the patient support unit with the patient placed on the patient support unit;c) imaging at least part of the patient's brain using an imaging unit of the tomographic imaging system, the imaging comprising in particular generating an image describing the functioning of the patient's blood vessel system;d) determining, in dependence on the result of the imaging, whether the patient support unit should be rotated relative to the tomographic imaging system with the patient placed on the patient support unit so that the patient is free of the imaging unit in order to conduct a medical intervention on the patient's blood vessel system, orwhether the patient support unit may remain in its position relative to the tomographic imaging system with the patient placed on the patient support unit. | 1. A medical method of managing a cerebrovascular insult, the method comprising: a) placing a patient on a patient support unit;b) positioning a mobile tomographic imaging system in a predetermined position relative to the patient support unit with the patient placed on the patient support unit;c) imaging at least part of the patient's brain using an imaging unit of the tomographic imaging system, the imaging comprising generating an image describing functioning of a blood vessel system of the patient; andd) determining, in dependence on the result of the imaging,whether the patient support unit should be rotated relative to the tomographic imaging system with the patient placed on the patient support unit so that the patient is free of the imaging unit in order to conduct a medical intervention on the patient's blood vessel system, or whether the patient support unit may remain in its position relative to the tomographic imaging system with the patient placed on the patient support unit,wherein positioning the tomographic imaging system in a predetermined position relative to the patient support unit comprises fastening the patient support unit with the patient placed on the patient support unit to a pedestal of the tomographic imaging system, the pedestal having a predetermined position relative to the imaging unit of the tomographic imaging system and being rotatable around a vertical axis running through the pedestal and patient support unit and rotatable relative to the imaging unit, wherein the pedestal is an integral part of the mobile tomographic system. 2. The method of claim 1, further comprising: determining, as a result of the imaging, whether the patient suffers from an ischemic stroke or a hemorrhagic stroke. 3. The method of claim 1, wherein the predetermined position is a fixed position. 4. The method of claim 3, wherein the patient support unit is detachably fastened to the pedestal. 5. The method of claim 1, wherein the imaging unit is a scanner configured to generate a magnetic resonance tomography or a computed x-ray tomography. 6. The method of claim 5, wherein the magnetic resonance tomography is a magnetic resonance-based angiography. 7. The method of claim 5, wherein the computed x-ray tomography is an x-ray-based angiography. 8. The method of claim 7, wherein the x-ray-based angiography is a computed tomography angiography. 9. The method of claim 1, wherein the imaging includes generating an angiography of at least part of the patient's brain. 10. The method of claim 1, wherein the patient support unit is a bed or a flat table. 11. The method of claim 1, wherein the patient support unit is translucent for x-rays. 12. The method of claim 1, wherein positioning the tomographic imaging system in a predetermined position relative to the patient support unit comprises leaving the patient support unit stationary and moving the imaging unit towards the patient support unit. 13. The method of claim 1, wherein if deter... | Brainlab AG | A61B005/05 A61B005/055 A61B006/04 A61B006/00 A61B005/00 A61B006/03
US09775526 | B2 | 2017-10-03 | Blood pressure meter and pump driving system | 2014-10-17 | In the present invention, a pump driving circuit includes a step-up unit that steps up a first DC voltage from a power supply and outputs it as a second DC voltage, and an H bridge unit that has first and second series circuits that each include two switching elements connected in series between a high potential corresponding to the second DC voltage and a reference potential. According to a control signal from the control unit, the two switching elements of the first series circuit and the two switching elements of the second series circuit are switched on and off. A voltage generated between a first contact point between the two switching elements of the first series circuit, and a second contact point between the two switching elements of the second series circuit is used as a driving voltage for driving a pump. | 1. A blood pressure meter, comprising: a piezoelectric pump configured to send a fluid to a blood pressure measurement cuff;a pump driving circuit for driving the piezoelectric pump; anda processor configured to control the pump driving circuit in order to measure blood pressure, whereinthe pump driving circuit includes: a step-up unit that steps up a first DC voltage from a power supply and outputs it as a second DC voltage according to a step-up control signal from the processor, andan H bridge unit including first and second series circuits that each include two switching elements that are connected in series between a high potential corresponding to the second DC voltage and a reference potential that is lower than the high potential,the two switching elements of the first series circuit are switched on and off in a complementary manner according to bridge control signals from the processor, and the two switching elements of the second series circuit are switched on and off in a complementary manner with a phase that is an inverse of that of the switching on and off of the two switching elements of the first series circuit,a voltage that is generated between a first contact point between the two switching elements of the first series circuit and a second contact point between the two switching elements of the second series circuit is used as a driving voltage for driving the piezoelectric pump, andthe processor is configured to set a rest period in which all switching elements are off between an on period of a switching element in each of the first and second series circuits and an on period of another switching element that follows the on period. 2. The blood pressure meter according to claim 1, wherein the step-up unit is a step-up regulator that varies and outputs the second DC voltage according to the step-up control signal from the processor. 3. The blood pressure meter according to claim 2, wherein the step-up control signal from the processor is a pulse width modulation signal, andthe step-up regulator serving as the step-up unit varies and outputs the second DC voltage according to the pulse width of the pulse width modulation signal. 4. A blood pressure meter, comprising: a piezoelectric pump configured to send a fluid to a blood pressure measurement cuff;a pump driving circuit for driving the piezoelectric pump; anda processor configured to control the pump driving circuit in order to measure blood pressure, whereinthe pump driving circuit includes: a step-up unit that steps up a first DC voltage from a power supply and outputs it as a second DC voltage according to a step-up control signal from the processor, andan H bridge unit including first and second series circuits that each include two switching elements that are connected in series between a high potential corresponding to the second DC voltage and a reference potential that is lower than the high potential,the two switching elements of the first series circuit are switc... | OMRON HEALTHCARE CO., LTD. MURATA MANUFACTURING CO., LTD. | A61B005/021 F04B045/047 A61B005/0225 A61B005/024 A61B005/022
US09802081 | B2 | 2017-10-31 | Bike system for use in rehabilitation of a patient | 2015-12-11 | A system for use in rehabilitation of a target patient is provided. The system includes at least two bicycle devices for use by the target patient and a second operator other than the target patient. The at least two bicycle devices each include pedals. At least one of the pedals may have at least one sensor mounted thereon for monitoring operation of the first bicycle device and the target's condition. A servomotor is coupled to the pedals for providing gear-like resistance or pedal assistance for the at least two bicycle devices. A controller is programmed to electrically couple the at least two bicycle devices to each other. | 1. A system for use in rehabilitation of a target patient, the system comprising: at least two bicycle devices for use by the target patient and a second operator other than the target patient, the at least two bicycle devices each including:pedals, indicating operation of the first bicycle device and the target's condition;a servomotor coupled to the pedals for providing resistance or assistance for the at least one bicycle device; anda controller programmed to electrically couple the at least two bicycle devices to each other, said controller including a human in the loop control system programmed to:react and adapt to changes initiated by at least one of the target patient and the second operator;detect and accommodate inappropriate interaction by at least one of the target patient and the second operator;accommodate different skill level and competency levels between the target patient and the second operator;detect and respond to wear and fault conditions of the system; and/orprotect the at least one of the target patient and the second operator and the system. 2. The system of claim 1, wherein the controller is further programmed to acquire performance data related to each of the at least two bicycle devices. 3. The system of claim 1, wherein the controller is further programmed to: collect performance data from each of the target patient and the second operator; andgenerate a mathematical model of an interaction between the target patient and the second operator from the performance data. 4. The system of claim 3, wherein the mathematical model provides the electrical coupling between the at least two bicycle devices. 5. The system of claim 3, wherein the controller is further programmed to: generate a statistical model developed from patient data of other patients; andadjust operation of the system responsive the mathematical model and the statistical model. 6. The system of claim 5, wherein the controller adjusts operation of the system by being further programmed to: dynamically alter a cadence and a torque experienced by each of the target patient and the second operator through a real-time rehabilitation management control algorithm. 7. The system of claim 6, wherein the controller is further programmed to dynamically alter a cadence and a torque experienced by each of the target patient and the second operator through a machine-learning algorithm. 8. The system of claim 1, wherein the system operates with the target patient operating one of the at least two bicycle devices and a second user different from the target patient operating another of the at least two bicycle devices. 9. The system of claim 1, wherein the system operates with the target patient operating one of the at least two bicycle devices and the controller operating another of the at least two bicycle devices. 10. The system of claim 1, wherein the human in the loop control system includes: a plant component that includes actual output data of the target pati... | KENT STATE UNIVERSITY CASE WESTERN RESERVE UNIVERSITY | A63B024/00 A61B005/00 A63B021/28 B62K003/14 G06F019/00 A63B022/06 A63B021/005 A63B022/00 A63B022/02
US09795314 | B2 | 2017-10-24 | Medical devices for mapping cardiac tissue | 2016-08-08 | Medical devices and methods for making and using medical devices are disclosed. An example method may include a method of identifying an activation time in a cardiac electrical signal. The method may include sensing a cardiac electrical signal, generating an approximation signal based at least in part on one or more parameters of the cardiac electrical signal, identifying a fiducial point on the approximation signal and determining, based at least in part on a timing of the fiducial point in the approximation signal, an activation time in the cardiac electrical signal. | 1. A system for mapping the electrical activity of the heart, the system comprising: a catheter comprising a plurality of electrodes positioned on a distal end of the catheter, the plurality of electrodes configured to sense a plurality of signals, wherein each of the plurality of signals is a function of time;a processor configured to: receive the plurality of signals sensed by the plurality of electrodes;generate a plurality of alternate signals, wherein each alternate signal is a function of frequency and wherein each alternate signal of the plurality of alternate signals corresponds to a respective signal of the plurality signals;generate a composite signal of the plurality of alternate signals;determine a characteristic frequency from the generated composite signal by identifying a frequency of the composite signal having a maximum power; andoutputting to a display device the determined characteristic frequency. 2. The system of claim 1, wherein to generate a composite signal of the plurality of signals, the processor is configured to calculate: a mean, a median or a mode at each frequency of the plurality of alternate signals. 3. The system of claim 1, wherein the processor is configured to: determine a plurality of phase values that correlates each alternate signal to its respective signal of the plurality of signals;phase shift each of the plurality of alternate signals using a respective phase value of the plurality of phase values; andgenerate the composite signal using the phase shifted alternate signals. 4. The system of claim 1, wherein the processor is configured to determine at least one fiducial of the composite signal. 5. The system of claim 4, wherein to determine at least one fiducial, the processor is configured to determine at least one of: a peak amplitude, change in slope and deflection of the correlated sinusoidal signal. 6. The system of claim 4, wherein to determine at least one fiducial, the processor is configured to determine a window for the composite signal and determine the at least one fiducial within the determined window. 7. The system of claim 4, wherein the at least one fiducial corresponds to an activation time. 8. The system of claim 1, further comprising a band-pass filter configured to filter the sensed signals of the plurality of signals. 9. The system of claim 8, wherein the band-pass filter filters signals outside of the 3 to 7 Hz band. 10. The system of claim 1, wherein to generate a plurality of alternate signals, the processor is configured to: transform the sensed plurality of signals using a Fourier Transform or a Fast Fourier Transform, Wavelet Transform or Wigner-Ville Transform. 11. A method for mapping the electrical activity of the heart, the method comprising: receiving a plurality of signals sensed by a plurality of electrodes positioned within the heart, wherein each of the plurality of signals is a function of time;generating a plurality of sinusoidal signals, wherein each sinuso... | BOSTON SCIENTIFIC SCIMED, INC. | A61B005/04 A61B005/044 A61B005/00 A61B005/042 A61B005/046 A61N001/02 A61B005/0452
US09829327 | B2 | 2017-11-28 | Method for planning a route for physical training purposes | 2014-08-07 | Method of route planning for physical training, wherein the heart rate of the subject should be kept within a permitted range. Prior to route planning, map data including starting and target points and altitude data are stored. A load profile is created by measuring the heart rate of the subject along routes of different steepness and storing the data in an electronic table, taking into account the required effort and whether the subject could have partially regenerated. The route planning includes searching and selecting a possible route, dividing the route into segments of steepness, and determining the costs based on the steepness data. The cost represents the load when the subject takes that section. Next, expected heart rate is calculated and it is determined whether it will remain in the permitted range along the whole first route. If the heart rate will be in the permitted range, the route is permitted, if not, the steps are repeated for subsequent possible routes where the heart rate in the associated route will remain in the permitted range. | 1. Method of route planning for physical training purposes, wherein during movement along a planned route the heart rate of the training subject should be kept within a permitted range; comprising the steps of: as preparatory steps prior to the actual route planning: storing in a navigation device map data of an area concerned that includes required starting and target points, the map data include also altitude data;determining a load profile of the subject including:measuring the heart rate of the subject along routes of different steepness;storing the so obtained data in an electronic table stored in the device;the actual route planning includes the further steps of: searching possible routes between said starting and target points;selecting a first one of the possible routes anddividing the first possible route into segments that have respective constant steepness;determining the respective costs typical of each section based on the steepness data of the section using the table of said stored load profile, wherein the cost represents the load when the subject takes that section;calculating from the determined cost values in the course of the movement of the subject proceeding along said first possible route his expected heart rate;examining whether the expected heart rate will remain in the permitted range along the whole first route;accepting the first route as planned route if the examination shows that the heart rate will be in the permitted range; andrepeating the planning steps for subsequent possible routes as long as the examination at a route shows that the heart rate in the associated route will remain in the permitted range, and accepting that route as the planned routed, characterized in, that during the step of determining the load profile the time elapsed since the beginning of the movement is also stored in said table, and said determination step is carried out repeatedly and the extent of regeneration of the subject is determined as a result of a larger load followed by a smaller load; and at said calculation step of the expected heart rate the cost of any section depends also on the time elapsed since the beginning of the movement and the total load during said elapsed time; and in said calculating step examining at each section whether at the previous section or sections there was a decrease in the required level of effort in which the subject could have partially regenerated and if such is experienced the previously determined regeneration is also taken into account in determining the cost associated with the concerned section. 2. The method as claimed in claim 1, wherein navigating the subject with the navigation device along said accepted planned route. 3. The method as claimed in claim 1, wherein said starting and target points are the same, said route is planned on the basis of a predetermined training time and/or so that the heart rate of the subject remains in the permitted range. 4. The method as claimed in claim 2,... | GPS TUNER KFT. | A63B024/00 G01C021/34 G06F019/00 A61B005/024 A61B005/11 A61B005/18 A61B005/00 G01C021/20
US09775681 | B2 | 2017-10-03 | Haptic guidance system and method | 2011-08-01 | A surgical apparatus includes a surgical device, configured to be manipulated by a user to perform a procedure on a patient, and a computer system. The computer system is programmed to implement control parameters for controlling the surgical device to provide at least one of haptic guidance to the user and a limit on user manipulation of the surgical device, based on a relationship between an anatomy of the patient and at least one of a position, an orientation, a velocity, and an acceleration of a portion of the surgical device, and to adjust the control parameters in response to movement of the anatomy during the procedure. | 1. A method of operating a surgical system comprising: manually manipulating a surgical instrument;tracking a position of the surgical instrument and a position of an anatomy of a patient using a detection device, wherein the positions of the surgical instrument and the anatomy are tracked relative to a reference point within a coordinate frame of reference of the detection device;associating a virtual object with an anatomy of a patient, wherein the virtual object is representative of a cutting boundary;wherein the surgical instrument and the anatomy move independently of the other and relative to the reference point;determining, based on the tracked position of the anatomy, that the anatomy has moved relative to the reference point;maintaining registration between the anatomy and the virtual object when the anatomy has moved relative to the reference point by automatically adjusting the cutting boundary relative to the reference point in response to the determined movement of the anatomy;providing haptic feedback to a user manually manipulating the surgical instrument to constrain the surgical instrument from penetrating the adjusted cutting boundary; andproviding visual information to the user, wherein the visual information comprises a representation of a desired portion of material to be removed by the surgical instrument in a first color that is updated to represent a portion of the desired material actually removed by the surgical instrument in a second color, based on the tracked position of the surgical instrument, and wherein the visual information further comprises a representation of a portion of material actually removed that exceeds the desired portion of material to be removed in a third color. 2. The method of claim 1, wherein the surgical instrument is a cutting tool. 3. The method of claim 2, wherein the cutting boundary defines a cutting surface relative to the anatomy. 4. The method of claim 1, wherein the haptic feedback is provided based on a relationship between the anatomy and at least one of a position, an orientation, a velocity and an acceleration of at least a portion of the surgical instrument. 5. The method of claim 1, wherein the step of associating a virtual object with the anatomy includes: creating a virtual representation of the anatomy;registering the virtual representation of the anatomy to the anatomy; andassociating the virtual object with the virtual representation of the anatomy. 6. The method of claim 1, wherein the step of associating the virtual object with the anatomy of the patient includes associating the virtual object with a bone of the patient. 7. The method of claim 6, wherein the bone is a femur. 8. The method of claim 6, wherein the bone is a tibia. 9. The method of claim 1, wherein at least a portion of the cutting boundary includes material to be removed from the anatomy. 10. The method of claim 1, wherein the step of providing haptic feedback to the user includes controlling an actu... | MAKO Surgical Corp. | A61B034/30 A61B090/00 A61F002/30 G06F003/01 A61B017/16 A61B034/37 A61B034/35 A61N001/05 A61N001/36 A61N001/372 A61F002/46 A61B005/11 A61B017/17 A61B017/00 A61F002/38 A61B005/00 A61B034/10 A61B034/20 A61B034/00 A61B090/14
US09813078 | B2 | 2017-11-07 | Systems and methods for automated electrical element addressing | 2016-10-21 | The present disclosure relates generally to automating the task of assignment of labels to identify electrical elements (e.g., electrode contacts, electrodes including a plurality of electrode contacts, and/or non-addressable electrical elements, like wires). A system that can automate the task of assignment of labels can include an electrical element, a microelectronic circuit associated with the electrical element, and an acquisition system. The microelectronic circuit can transmit a sequence comprising a label corresponding to the electrical element. The acquisition system can assign the label corresponding to the electrical element to a recording channel after decoding the sequence. | 1. A method comprising: encoding, by a microcircuit associated with an electrical element, a sequence with a label corresponding to the electrical element;transmitting, by the microcircuit, the sequence to an acquisition system;decoding, by the acquisition system, the sequence to determine the label corresponding to the electrical element; andassigning, by the acquisition system, the label corresponding to the electrical element to a recording channel,wherein the electrical element comprises an addressable electrode contact, an addressable electrode comprising a plurality of electrode contacts, or a non-addressable electrical element. 2. The method of claim 1 further comprising placing an adaptor comprising the microcircuit serially inline with the electrical element. 3. The method of claim 1, wherein the sequence comprises a start, a label identifier, and a stop, and wherein the start sequence, the address sequence, and the stop sequence are identifiable due to at least one of interpulse intervals, amplitudes, and frequencies. 4. The method of claim 3, wherein the label identifier comprises a binary sequence of positive and negative values according to an encoding scheme. 5. The method of claim 3, wherein at least one of the start and the stop comprises a series of pulses with small inter-pulse intervals. 6. The method of claim 5, wherein the label identifier comprises a series of pulses with larger inter-pulse intervals or pulse heights than the at least one of the start and the stop, and wherein the series of pulses of the label identifier describes information relative to a specific electrode. 7. The method of claim 1, further comprising permitting, by the acquisition system, the label to be edited. 8. The method of claim 7, wherein the label is edited by: receiving, via a computer interface, an input comprising an edit to the label;sending, by the acquisition system, the edited label to the microcircuit; andreplacing, by the microcircuit, the label with the edited label,wherein the edited label replaces the label stored on the microcircuit. 9. The method of claim 1, wherein the microcircuit physically contacts the electrical element or the microcircuit physically contacts a connector between the acquisition system and the electrical element. 10. A system comprising: a non-addressable electrical element associated with an adaptor insert comprising addressing circuitry and a controller;a microelectronic circuit associated with the non-addressable electrical element to transmit a sequence comprising a label corresponding to the non-addressable electrical element; andan acquisition system to assign the label to a recording channel after decoding the sequence. 11. The system of claim 10, wherein the acquisition system is designed for a vendor application of the non-addressable electrical element. 12. The system of claim 10, wherein the addressing circuitry comprises: a first relay to allow a line signal to pass;a signal generator to pr... | The Cleveland Clinic Foundation | A61B005/04 H03M007/30 G06K019/06 A61B005/00 A61B005/0408
US09801526 | B2 | 2017-10-31 | Minimally invasive surgical system | 2012-05-07 | A surgical instrument is inserted through a guide tube. A telemanipulation system moves the distal end of the surgical instrument in all six Cartesian degrees of freedom independently of any guide tube movements. | 1. A method comprising: advancing a distal end of a straight first rigid surgical instrument body tube through a first channel of a guide tube, the distal end of the first rigid surgical instrument body tube being coupled to an end component, wherein as the distal end of the first rigid surgical instrument body tube advances through the guide tube, a portion of the first rigid surgical instrument body tube proximal to and outside the guide tube is resiliently bent; andadvancing a second straight rigid surgical instrument body tube of a second surgical instrument through a second channel of the guide tube, a portion of the second rigid surgical instrument body tube proximal to and outside the guide tube being resiliently bent, the guide tube being coupled to a platform, each of the first and second surgical instruments being coupled to the platform by a different teleoperated manipulator assembly, and the platform being coupled to a robotic manipulator arm of a patient side support system. 2. The method of claim 1 further comprising: rotating the guide tube around a remote center point. 3. The method of claim 2: wherein a guide tube longitudinal axis is defined through proximal and distal ends of the guide tube; andwherein the rotating the guide tube comprises rotating the guide tube around an axis perpendicular to the guide tube longitudinal axis. 4. The method of claim 2: wherein a guide tube longitudinal axis is defined through proximal and distal ends of the guide tube; andwherein the rotating the guide tube comprises rotating the guide tube around the guide tube longitudinal axis. 5. A method of operating a patient side support system comprising: advancing a first distal portion of a first straight rigid surgical instrument body tube of a first surgical instrument through a first channel of a guide tube, with a first proximal portion of the first rigid surgical instrument body tube being resiliently bent between a first transmission mechanism of the first surgical instrument and a proximal end of the guide tube as the first distal portion of the first rigid surgical instrument body tube advances through the first channel of the guide tube;wherein the first transmission mechanism is coupled to the first proximal portion of the first rigid surgical instrument body tube;wherein a first surgical end effector is coupled to the first distal portion of the first rigid surgical instrument body tube and is coupled to the first transmission mechanism;wherein the guide tube is coupled to a platform;wherein the first transmission mechanism is coupled to a first teleoperated manipulator assembly of a plurality of teleoperated manipulator assemblies; andwherein the plurality of teleoperated manipulator assemblies is coupled to the platform; andwherein the platform is coupled to a robotic manipulator arm of the patient side support system. 6. The method of claim 5, further comprising: moving the first surgical end effector in at least one of the six Ca... | Intuitive Surgical Operations, Inc. | A61B001/00 A61B001/04 A61B017/00 A61B005/05 A61B019/00 A61B034/37 A61B034/30 A61B001/018 A61B001/002 A61B001/06 A61B005/00 A61B008/12 A61B001/05 A61B090/10 A61B034/00 A61B034/20 A61B001/005 A61B017/34 A61B090/00
US09826955 | B2 | 2017-11-28 | Air conduction sensor and a system and a method for monitoring a health condition | 2011-03-31 | According to embodiments of the present invention, an air conduction sensor for detecting a sound from a user is provided. The air conduction sensor includes a housing comprising an opening, wherein a rim of the opening is configured to at least substantially attach to a skin or a clothing of the user; a microphone coupled to the housing such that there is an air gap between the microphone and the skin or the clothing, and wherein the microphone is configured to detect the sound. A system and a method for monitoring a health condition of a user are also provided. | 1. A system for monitoring a health condition of a user, comprising: an air conduction sensor for detecting a sound from the user, comprising: a housing comprising an opening, wherein a rim of the opening is configured to be disposed on a skin or a clothing of the user; anda microphone coupled to the housing such that there is an air gap between the microphone and the skin or the clothing, wherein the microphone is configured to detect the sound,a filter and an amplifier in electrical communication with the microphone, the filter and the amplifier configured to process the sound;an analogue-to-digital converter configured to convert the sound detected into a digital signal;a selection circuit configured to select a predetermined number of samples from the digital signal;a transformation circuit configured to transform the selected predetermined number of samples into a plurality of frequency components;an entropy determination circuit configured to determine at least one entropy-based parameter value based on the plurality of frequency components; andan entropy comparator circuit configured to compare the at least one entropy-based parameter value determined with a predetermined threshold value to detect a health signal for monitoring the health condition. 2. The system as claimed in claim 1, wherein the microphone is configured to detect the sound in a frequency ranging from 20 Hz to 20 kHz. 3. The system as claimed in claim 1, wherein the microphone is configured to convert the sound into an electrical signal. 4. The system as claimed in claim 1, wherein the air gap is in a range from 0.1 cm to 3 cm. 5. The system as claimed in claim 1, wherein the rim of the housing is configured to cover an area on the skin or the clothing in a range from 0.19 cm2 to 30 cm2. 6. The system as claimed in claim 1, wherein the housing has a shape of a hemisphere. 7. The system as claimed in claim 1, wherein the microphone is configured to detect the sound from at least one of a lung and a heart of the user. 8. A method for monitoring a health condition of a user, comprising: detecting a sound from the user using an air conduction sensor the air conduction sensor comprising: a housing comprising an opening, wherein a rim of the opening is configured to be disposed on a skin or a clothing of the user; anda microphone coupled to the housing such that there is an air gap between the microphone and the skin or the clothing, wherein the microphone is configured to detect the sound;processing the sound using a filter and an amplifier in electrical communication with the microphone;converting the sound detected into a digital signal;selecting a predetermined number of samples from the digital signal;transforming the selected predetermined number of samples into a plurality of frequency components;determining at least one entropy-based parameter value based on the plurality of frequency components; andcomparing the at least one entropy-based parameter value determi... | Nanyang Technological University | A61B005/02 A61B007/00 A61B007/04 H04R001/46 A61B005/0205 H04R001/40
US09826938 | B2 | 2017-11-28 | Motion compensation for optical heart rate sensors | 2015-09-14 | A method of optical heart rate sensing includes receiving a motion frequency from a motion sensor and receiving an optical signal from an optical sensor. The motion frequency is then filtered from the optical signal, and an estimated heart rate frequency is determined based on the filtered optical signal. A heart rate is reported that is based on the motion frequency, but not based on the estimated heart rate frequency, when a local neighbor magnitude ratio of the estimated heart rate frequency is below a confidence threshold. | 1. A method of optical heart rate sensing, comprising: receiving a motion frequency via a motion sensor;receiving an optical signal from an optical sensor;filtering the motion frequency from the optical signal;determining an estimated heart rate frequency based on the filtered optical signal;reporting a heart rate based on the motion frequency, but not based on the estimated heart rate frequency responsive to a local neighbor magnitude ratio of the estimated heart rate frequency being less than a confidence threshold; andreporting a heart rate based on the estimated heart rate frequency responsive to the local neighbor magnitude ratio of the estimated heart rate frequency being greater than the confidence threshold. 2. The method of claim 1, further comprising: dynamically adjusting the confidence threshold based on the estimated heart rate frequency. 3. The method of claim 1, wherein the local neighbor magnitude ratio of the estimated heart rate frequency comprises a ratio of a magnitude of the estimated heart rate frequency to a sum of magnitudes of heart rate frequencies within a threshold of the estimated heart rate frequency. 4. The method of claim 3, wherein the magnitudes of heart rate frequencies within a threshold of the estimated heart rate frequency include magnitudes of heart rate frequencies within 30 beats per minute of the estimated heart rate frequency. 5. The method of claim 3, further comprising: transforming a recent window of the filtered optical signal from a time domain to a frequency domain;determining the magnitude of the estimated heart rate frequency based on the transformed filtered optical signal; anddetermining magnitudes of heart rate frequencies within a threshold of the estimated heart rate frequency based on the transformed filtered optical signal. 6. The method of claim 5, wherein the recent window of the filtered optical signal is a recent five second window. 7. The method of claim 1, wherein an amount of motion signal filtering remains constant with changing motion frequency. 8. The method of claim 7, wherein filtering the optical signal based on the motion frequency includes applying a comb filter to the optical signal. 9. The method of claim 8, wherein filtering the optical signal based on the motion frequency further includes: adjusting one or more parameters of the comb filter based on a recent motion frequency while maintaining the amount of motion signal filtering. 10. An optical heart rate sensor, comprising: an optical source configured to illuminate one or more blood vessels through a user's skin;an optical sensor configured to measure illumination from the one or more blood vessels;one or more motion sensors; anda storage machine holding instructions executable by a logic machine to: receive an optical signal from the optical sensor;receive a motion signal from the one or more motion sensors;determine a motion frequency based on the motion signal;filter the motion frequency from the optical s... | MICROSOFT TECHNOLOGY LICENSING, LLC | A61B005/02 A61B005/00 A61B005/024 A61B005/01
US09782087 | B2 | 2017-10-10 | Heart rate estimating apparatus and method | 2013-06-12 | A heart rate estimating apparatus may include a transmitter to transmit microwave with respect to a subject, a receiver to receive and detect reflected wave from the subject irradiated with the microwave, and acquire a detected result, a sensor to sense a movement of the subject, and acquire a sensed result, and an estimating unit. The estimating unit may estimate a heart rate based on feature points remaining after excluding, from heart rate candidates, feature points that are obtained by frequency analysis of the detected result and are located in a vicinity of frequencies at which feature points are obtained by frequency analysis of the sensed result. | 1. A heart rate estimating apparatus comprising: a transmitter configured to transmit microwave with respect to a subject;a receiver configured to receive and detect reflected wave from the subject irradiated with the microwave, and acquire a detected result;a sensor configured to sense a movement of the subject, and acquire a sensed result;a storage unit configured to store a program; anda processor configured to execute the program stored in the storage unit and perform a process including performing a first frequency analysis of the detected result to obtain a first result of the first frequency analysis versus frequency, wherein the first result includes feature points that are formed by peaks having a first amplitude or greater;performing a second frequency analysis of the sensed result to obtain a second result of the second frequency analysis versus frequency, wherein the second result includes feature points that are formed by peaks having a second amplitude or greater; andestimating a heart rate based on feature points remaining after excluding the feature points included in the first result and located at frequencies within a predetermined frequency range of frequencies of the peaks forming the feature points included in the second result. 2. The heart rate estimating apparatus as claimed in claim 1, wherein the sensor includes at least one of a microphone, an acceleration sensor, and an angular velocity sensor. 3. The heart rate estimating apparatus as claimed in claim 1, wherein the predetermined frequency range is 0.5 Hz or higher and 4.0 Hz or lower. 4. The heart rate estimating apparatus as claimed in claim 1, wherein the estimating performs a filter process with respect to the detected result, using a frequency of one of the peaks forming the feature points remaining after excluding the feature points included in the first result and located at the frequencies within the predetermined frequency range, as a center of the filter process. 5. The heart rate estimating apparatus as claimed in claim 4, wherein the processor performs the process further including performing a first order differentiation on output data of the filter process to acquire a heart beat interval; andcomputing the heart rate based on the heart beat interval. 6. A heart rate estimating method comprising: acquiring a detected result from a receiver that receives and detects reflected wave from a subject irradiated with microwave;acquiring a sensed result from a sensor that senses a movement of the subject;performing, by a processor, a first frequency analysis of the detected result to obtain a first result of the first frequency analysis versus frequency, wherein the first result includes feature points that are formed by peaks having a first amplitude or greater;performing, by the processor, a second frequency analysis of the sensed result to obtain a second result of the second frequency analysis versus frequency, wherein the second result includes feature ... | FUJITSU LIMITED | A61B005/05 A61B005/024 A61B005/00
US09821003 | B2 | 2017-11-21 | Methods and devices for treating conditions associated with autonomic dysfunction | 2016-03-21 | Methods and devices treating an autonomic nervous system associated disease condition in a subject are provided. Aspects of the invention include inducing one or more physiological response selected from the group consisting of sweating, gastric emptying, enhanced heart rate variability and enhanced quantitative sensory test responsiveness in a manner sufficient to modify the autonomic nervous system so as to treat the subject for the disease condition. The methods and devices find use in a variety of applications, e.g. in the treatment of subjects suffering from conditions arising from disorders of the autonomic nervous system. | 1. A method of treating a subject having a condition arising from autonomic dysfunction, the method comprising: enhancing one or more physiological responses selected from the group consisting of: sweating,gastric emptying,heart rate variability,ability to rapidly transition from prone to suspended orientation, andquantitative sensory test responsiveness,in a manner sufficient to modify the autonomic nervous system so as to treat the subject for the disease condition by applying a pharmacological stimulus, an electrical stimulus, a thermal stimulus, or a combination of stimuli thereof, to induce the one or more physiological responses multiple times during a course of treatment. 2. The method according to claim 1, wherein the physiological response is sweating. 3. The method according to claim 2, wherein method comprises whole body induction of sweating. 4. The method according to claim 2, wherein the method comprises localized induction of sweating. 5. The method according to claim 1, wherein the physiological response gastric emptying. 6. The method according to claim 5, wherein the method comprises electrically stimulating gastric emptying. 7. The method according to claim 6, wherein the method comprises pharmacologically stimulating gastric emptying. 8. The method according to claim 1, wherein the physiological response is heart rate variability. 9. The method according to claim 8, wherein enhancing comprises applying an electrical stimulus to induce heart rate variability multiple times during the course of treatment. 10. The method according to claim 8, wherein the enhancing comprises applying a pharmacological stimulus to induce heart rate variability multiple times during the course of treatment. 11. The method according to claim 1, wherein the physiological response is ability to rapidly transition from prone to suspended orientation and the method comprises rapidly transitioning the subject from a prone to suspended orientation. 12. The method according to claim 1, wherein the physiological response is quantitative sensory test responsiveness. 13. The method according to claim 1, wherein the subject is a mammalian subject. 14. The method according to claim 13, wherein the mammalian subject is a human. 15. The method according to claim 1, wherein the method further comprises diagnosing the subject as having an autonomic nervous system condition. 16. A device comprising: a physiological response stimulator; anda controller configured to operate the device in a manner sufficient to perform the method according to claim 1. 17. The device according to claim 16, wherein the device is an implantable device. 18. The implantable device according to claim 17, wherein the implantable device is an implantable pulse generator. 19. The implantable device according to claim 17, wherein the implantable device is an implantable active agent delivery device. 20. The device according to claim 16, wherein the device is configured to be s... | Palo Alto Investors | A61K031/70 A61K031/7048 A61M005/00 A61B005/145 A61B005/00 A61F007/00 A61H023/00 A61K031/137 A61K031/166 A61K031/21 A61K031/4468 A61K031/454 A61N001/30 A61N001/36 A61N001/362 A61B005/024 A61B005/11
US09820760 | B2 | 2017-11-21 | Aiming device for distal locking of intramedullary nails and methods of use | 2015-04-24 | An aiming device system that is configured to align at least one bone anchor guide to a first distal hole of an intramedullary nail can include an aiming arm that is configured to be operatively coupled to the intramedullary nail and an aiming guide. The aiming arm can define a sliding support that is elongate along a first direction. The aiming guide can include a guide body that defines at least a first bone anchor guide. The aiming guide is attached to the aiming arm such that the aiming guide is selectably movable along the sliding support, and selectably rotatable relative to the aiming arm. When the first bone anchor guide is retaining the guide sleeve, at least one of the selectable movement and selectable rotation of the guide body at least partially aligns the first bone anchor guide with the first distal hole of the intramedullary nail. | 1. An aiming device system configured to align at least one guide sleeve to a first distal hole of an intramedullary nail, the aiming device system comprising: an aiming arm configured to be operatively coupled to the intramedullary nail, the aiming arm defining a sliding support that is elongate along a first direction;an aiming guide including a guide body that defines at least a first bone anchor guide that is configured to retain a first guide sleeve, the aiming guide further including an attachment mechanism that is configured to mate with the sliding support to thereby attach the guide body to the aiming arm such that the guide body is (i) selectably movable along the sliding support along the first direction, and (ii) selectably rotatable relative to the aiming arm about a guide body axis that is perpendicular to the first direction,wherein when the first bone anchor guide retains the first guide sleeve, at least one of the selectable movement and selectable rotation of the guide body at least partially aligns the first guide sleeve with the first distal hole of the intramedullary nail,wherein (i) the sliding support is a slot that is elongate along the first direction, (ii) the attachment mechanism includes a pin that extends out from the guide body such that the pin defines the guide body axis, and (iii) the pin is configured to be received by the slot such that the pin is slidable within the slot,wherein (i) the guide body defines a curved slot and (ii) the attachment mechanism further includes first and second aiming guide attachment members, the first aiming guide attachment member being configured to extend through the slot of the aiming arm and through the curved slot of the guide body, and the second aiming guide attachment member being configured to mate with the first aiming guide attachment member to thereby attach the guide body to the aiming arm,wherein (i) the guide body defines a second bone anchor guide that is configured to retain a second guide sleeve that is configured to align with a second distal hole of the intramedullary nail and (ii) the aiming device system further comprising the first and second guide sleeves, each of the first and second guide sleeves including a respective bore that is configured to receive a bone anchor and guide the bone anchor to a respective one of the first and second distal holes; anda magnetic probe and a sensor unit, the magnetic probe having a magnet that is configured to extend into the intramedullary nail, the sensor unit being configured to be retained by the second bone anchor guide prior to the bone anchor guide retaining the second guide sleeve, wherein alignment of the sensor unit and magnet aligns the first bone anchor guide with the first distal hole of the intramedullary nail. 2. The aiming device system of claim 1, wherein the curved slot is curved such that the pin is at a center of the curve. 3. The aiming device system of claim 1, wherein (i) the first aiming guide attac... | DePuy Synthes Products, Inc. | A61B017/88 A61B019/00 A61B017/72 A61B017/17 A61B017/16 A61B005/06 A61B017/68 A61B017/90
US09788788 | B2 | 2017-10-17 | Three dimensional imaging of veins | 2015-05-14 | An apparatus and method for creating a three dimensional imaging system is disclosed. There is a first source of laser light and a second source of laser light having a wavelength different from the wavelength of the laser light of the first source. The laser light from the first and second sources are combined, and the combined laser light is transmitted to a scanner. The scanner further transmits the combined light to a surface to be imaged. | 1. A method of creating a layered vein image of subcutaneous veins beneath a target skin surface of a patient, and for projecting the image onto the surface, said method comprising: (a) emitting a first beam of light at a first wavelength;(b) emitting a second beam of light at a second wavelength being different than said first wavelength;(c) combining said first beam of light and said second beam of light into a single combined beam of light;(d) scanning said single combined beam of light in a pattern, and onto the target skin surface, using a scanner;(e) receiving a contrasted vein image formed of said first wavelength of light, by a photodetector responsive to said first wavelength, after being differentially absorbed and reflected by veins and surrounding tissue of the target skin surface;(f) converting said contrasted vein image into a digital image, using an analog-to-digital converter;(g) storing said digital vein image in a memory;(h) increasing an intensity of said first wavelength of light, by one or more increments, and repeating steps (a) through (g) for each said one or more incrementally increased light intensities, and storing, in said memory, a corresponding one or more digital vein images being at one or more incrementally increased depths beneath the target skin surface;(i) creating a processed vein image by layering of each vein in said stored vein images, by depicting veins being at a greater depth as appearing beneath veins at a shallower depth;(j) converting said processed vein image from a digital image into an analog image, using a digital-to-analog converter; and(k) scanning said analog processed vein image onto the target skin surface using said second wavelength of light. 2. The method according to claim 1 further comprising replacing each vein of said processed vein image with a unique pattern. 3. The method according to claim 1 further comprising replacing each vein of said processed vein image with a unique color of light. 4. The method according to claim 1 further comprising concurrently displaying said processed vein image on computer screen. 5. The method according to claim 1 further comprising incrementally increasing said intensity of said first wavelength of light for at least two of said incremental increases, and incrementally increasing said intensity in linear increments, with each said incremental increase being the same amount of increase in intensity. 6. The method according to claim 1 further comprising scanning of said pattern of said single combined beam of light in a raster pattern. 7. The method according to claim 1 wherein said scanning of said pattern of said single combined beam of light onto the target skin surface comprises scanning a field of view of said scan pattern to correspond with a top of the target skin surface of the patient. 8. The method according to claim 7 further comprising correcting a size of said veins in each of said stored vein images prior to said layering and scanni... | AccuVein, Inc | A61B006/00 A61B005/00 A61M005/42 A61M005/46 A61B017/34 A61B019/00 A61B090/00
US09810628 | B2 | 2017-11-07 | System and method for detection of materials using orbital angular momentum signatures | 2017-01-05 | An apparatus for detecting a presence of a predetermined material within a sample uses signal generation circuitry for generating a first signal having a first distinct signature including a first eccentricity of a mode intensity, a first shift in a center of the mode intensity and a first rotation of an ellipsoidal intensity output of the mode intensity and applying the first signal to the sample. A detector receives the first signal after the first signal passes through the sample and detects a second distinct signature including a second eccentricity of the mode intensity, a second shift in the center of the mode intensity and a second rotation of the ellipsoidal intensity output of the mode intensity. The detector also determines the presence of the predetermined material within the sample based on the detected second distinct signature within the first signal received from the sample and provides an output of an indication of the presence of the predetermined material responsive to the determination. | 1. An apparatus for detecting a presence of a predetermined material within a sample, comprising: signal generation circuitry for generating a first signal having a first distinct signature including a first eccentricity of a mode intensity, a first shift in a center of the mode intensity and a first rotation of an ellipsoidal intensity output of the mode intensity and applying the first signal to the sample;a detector for receiving the first signal after the first signal passes through the sample, for detecting a second distinct signature including a second eccentricity of the mode intensity, a second shift in the center of the mode intensity and a second rotation of the ellipsoidal intensity output of the mode intensity and for determining the presence of the predetermined material within the sample based on the detected second distinct signature within the first signal received from the sample, the detector providing an output of an indication of the presence of the predetermined material responsive to the determination. 2. The apparatus of claim 1, wherein the signal generation circuitry applies the first distinct signature using a spatial light modulator. 3. The apparatus of claim 1, wherein the signal generation circuitry applies the first distinct signature using a digital light processor. 4. The apparatus of claim 1, wherein the signal generation circuitry applies the first distinct signature using a hologram. 5. The apparatus of claim 1, wherein the signal generation circuitry applies the first distinct signature using an amplitude mask. 6. The apparatus of claim 1, wherein the signal generation circuitry applies the first distinct signature using a phase mask. 7. The apparatus of claim 1, wherein the detector further implements Photoshop to identify the second eccentricity of the mode intensity. 8. The apparatus of claim 1, wherein the detector implements Matlab to identify the second distinct signature. 9. The apparatus of claim 1, wherein the detector negates effects of at least one of sample distribution due to gravity, angle of camera recording the mode intensity to the sample, a container holding the sample, an angle of incidence of the first signal to the sample and an angle exit of the first signal from the sample to detect a change in an eccentricity of the mode intensity. 10. The apparatus of claim 1, wherein the detector further includes circuitry for determining a phase of the first signal after it passes through the sample, wherein the circuitry determines the phase by interfering the first signal having the second distinct signature therein with the first signal having plane waves therein. 11. A method for detecting a presence of a predetermined material within a sample based upon a unique orbital angular momentum signature, comprising: generating a first signal having a first distinct signature including a first eccentricity of a mode intensity, a first shift in a center of the mode intensity and a first rotation... | NxGen Partners IP, LLC | G01N021/59 G01N021/17 G01N033/483 G01N033/487 G01N021/21 A61B005/145 A61B005/1455 G01N024/00
US09826400 | B2 | 2017-11-21 | Method and apparatus that facilitates a wearable identity manager | 2014-07-28 | Various aspects directed towards a wearable identity manager system are disclosed. In a first aspect, an association status between a user and a wearable identity manager device is ascertained based on whether the wearable identity manager device is worn by the user, and motion data associated with a movement of the wearable identity manager device is monitored. Authentication data, which includes the motion data, is then transmitted based on the association status. In another aspect, an association status between a user and a wearable identity manager device is again determined based on whether the wearable identity manager device is worn by the user. Here, however, the wearable identity manager device is paired with a pairing device, and authentication data is transmitted to the pairing device based on the association status to facilitate a user authentication via the pairing device. | 1. A wearable identity manager device, comprising: a memory; anda processor communicatively coupled to the memory, the processor configured to: determine an association status between a user and the wearable identity manager device based in part on whether the wearable identity manager device is worn;pair the wearable identity manager device with a pairing device, wherein the pairing device is configured as a relay device between the wearable identity manager device and an external device based on whether the wearable identity manager device is continuously worn, and wherein the wearable identity manager device is configured to unpair from the pairing device in response to a discontinuity of the wearable identity manager device being worn; andtransmit authentication data from the wearable identity manager device to the external device via the pairing device based on whether the wearable identity manager device is worn, wherein the authentication data is data used to determine a user authentication. 2. The wearable identity manager device of claim 1, the processor further configured to: monitor and collect motion data associated with a movement of the wearable identity manager device, wherein the authentication data includes the collected motion data, and wherein the collected motion data is configured to facilitate a pairing of the wearable identity manager device with the pairing device. 3. The wearable identity manager device of claim 2, wherein the motion data comprises a path traversed by the wearable identity manager device. 4. The wearable identity manager device of claim 1, wherein the processor is configured to associate the wearable identity manager device with the user via an association procedure, wherein the association status is based on an outcome of the association procedure. 5. The wearable identity manager device of claim 1, wherein the processor is further configured to retrieve sensor data from at least one of a gyro, a global positioning system (GPS) device, a touch-sensitive sensor, or a microphone, and wherein the authentication data further comprises the sensor data. 6. The wearable identity manager device of claim 1, wherein the processor is further configured to infer whether the wearable identity manager device is worn based on data retrieved from at least one of a clasp sensor, a pressure sensor, a temperature sensor, a pulse sensor, a motion sensor, or a stretch sensor. 7. The wearable identity manager device of claim 1, wherein the processor is further configured to: receive an authentication request;provide a credential; andtransmit the credential based on the authentication request. 8. The wearable identity manager device of claim 1, wherein the processor is further configured to receive an authentication request;ascertain a security level associated with the authentication request; andtransmit the authentication data based on the security level. 9. A non-transitory machine-readable storage medium having one... | QUALCOMM Incorporated | G06F021/00 H04W012/06 G06F021/35 H04L029/06 A61B005/00 G06F001/16 G06F021/32 G06F003/0346 G06F021/34 A61B005/024 A61B005/11
US09795692 | B2 | 2017-10-24 | Ocular detection of amyloid proteins | 2014-06-05 | Described are methods for the detection, in the eye of an individual, of protein aggregates or other misfolded proteins associated with disease using peptide or peptide mimic probes that preferentially associate with the protein aggregates or misfolded proteins, which can be accomplished non-invasively. | 1. A method for detecting Aβ protein aggregates in the eye of an individual, comprising: (A) administering to the individual a peptide probe that consists of from 10 to 50 amino acid residues and comprises an amino acid sequence that is at least 60% identical to SEQ ID NO:1, wherein said probe binds to Aβ protein aggregates and(B) detecting probe associated with any Aβ protein aggregates present in the eye. 2. The method of claim 1, wherein said probe is labeled with a detectable label. 3. The method of claim 2, wherein said probe is labeled at separate sites with a first detectable label and a second detectable label, and wherein said labels generate a signal when said probe binds to said Aβ protein aggregates. 4. The method of claim 3, wherein said first and second labels comprise pyrene. 5. The method of claim 1, wherein said probe comprises an amino acid sequence that is at least 85% identical to SEQ ID NO:1. 6. The method of claim 1, wherein the probe comprises an amino acid sequence selected from the group consisting of SEQ ID NO:56 and SEQ ID NO:62. 7. The method of claim 1, wherein said administering is by systemic injection, direct injection, contacting the probe with the conjunctiva or cornea of the eye, intranasal administration or inhalation administration. 8. The method of claim 1, wherein said probe binds to Aβ protein aggregates localized in a region of the eye selected from the group consisting of the optic nerve and the retina. 9. The method of claim 1, wherein said Aβ protein aggregates are detected in a region of the eye selected from the group consisting of the optic nerve, lens, ciliary body, vitreous body, the retina and ocular blood vessels. 10. The method of claim 1, wherein said detecting is conducted using a retinal imaging device, a fundus camera, a slit-lamp operated with a fluoroscopic device, or by direct inspection using regular or laser light. 11. The method of claim 1, wherein said Aβ protein aggregates are associated with an amyloidogenic disease selected from the group consisting of Alzheimer's Disease and adult macular degeneration. 12. The method of claim 11, wherein the adult macular degeneration is a wet or dry variant. 13. A method for detecting Aβ protein aggregates in the eye of an individual, comprising: (A) administering to the individual a peptide probe that consists of from 10 to 50 amino acid residues and comprises an amino acid sequence that is at least 60% identical to SEQ ID NO:1, wherein said probe (i) binds to Aβ protein aggregates and (ii) generates a detectable signal when said probe binds to said Aβ protein aggregates; and(B) detecting any detectable signal resulting from said probe binding to any protein aggregates present in the eye, wherein the presence of a detectable signal indicates the presence of a condition selected from Alzheimer's Disease, age related macular degeneration, diabetes, scrapie, BSE, CJD, chronic wasting disease (CWD) and transmissible spongiform encephalop... | SYSTEM OF SYSTEMS ANALYTICS, INC. | A61K051/00 A61M036/14 A61K049/00 G01N033/68 A61B005/00 A61B003/12
US09831709 | B2 | 2017-11-28 | Method and apparatus for wireless charging of a bioelectronic device implanted in a laboratory animal or in a human being | 2013-06-18 | The invention relates to an apparatus for charging a remote feedable circuit bioelectronic implanted in a patient or in a laboratory animal, said apparatus comprising a composable container configured to define a closed environment suitable to receive a patient or a laboratory animal, said container comprising a plurality of composable walls made of a nonmagnetic material and connected to each other so as to define said closed environment, said container comprising at least one first winding whose axis is arranged in a first direction (Z) and at least one second winding whose axis is arranged in a second direction (Y) perpendicular to said first direction (Z). The apparatus further comprises a system for powering and driving the windings of the composable container, said system comprising a switching power driver for each winding, a plurality of phase locked loop circuits respectively connected to each switching power driver and connected to a programmable logic circuit of the powering and driving system, said programmable logic circuit being configured to perform a phase comparison, the programmable logic circuit being in turn connected to a microprocessor of the powering and driving system, said microprocessor being configured to provide driving signals to the windings for generating inside the container a rotating magnetic field. | 1. An apparatus for charging a remote feedable circuit bioelectronic implanted in a patient or in a laboratory animal, said apparatus comprising: a) a composable container configured to define a closed environment suitable to receive a patient or a laboratory animal, said container comprising a plurality of composable walls made of a nonmagnetic material and connected to each other so as to define said closed environment, said container comprising at least one first winding whose axis is arranged in a first direction (Z) and at least one second winding whose axis is arranged in a second direction (Y) perpendicular to said first direction (Z), said windings being installed in or on the walls and configured so as to: be not accessible to the patient and/or laboratory animal confined in the container,have electrical terminals accessible only from the outside of the container when it is mounted,radiate an electromagnetic field, when supplied with alternating current, towards the inside of said closed environment along said first direction (Z) and said second direction (Y); andb) a system for powering and driving the windings of the composable container, comprising: a switching power driver for each winding,a plurality of phase locked loop circuits respectively connected to each switching power driver and to a programmable logic circuit of the powering and driving system, said programmable logic circuit being configured to perform a phase comparison, the programmable logic circuit being in turn connected to a microprocessor of the powering and driving system, said microprocessor being configured to provide driving signals to the windings for generating inside the container a rotating magnetic field. 2. The charging apparatus according to claim 1, wherein at least one capacitor is connected at the terminals of each winding forming therewith a resonant circuit connected to a respective switching power driver, and wherein the values of the capacitance of the capacitors and of the inductance of the windings of each resonant circuit are dimensioned so that the resonant circuit resonates at a frequency corresponding to the frequency of a driving signal generated by the respective switching power driver. 3. The charging apparatus according to claim 1, comprising a pair of first identical windings arranged coaxially to each other and mutually spaced in said first direction (Z) of the container, as well as a pair of second identical windings arranged coaxially to each other and mutually spaced in the second direction (Y) of the container, perpendicular to the first direction (Z), wherein the windings of each pair are configured and arranged so as to form Helmholtz coils and said powering and driving system of the windings is configured to power the windings of each pair so as to generate magnetic fields having the same direction and to provide control signals suitable to generate inside the container a rotating magnetic field in a plane identified by the fir... | AB MEDICA HOLDING S.P.A. | H02J007/00 H02J007/02 A01K001/03 A01K011/00 A61B005/00 H01F038/14 H02J007/04 H02J005/00 H02J017/00
US09824444 | B2 | 2017-11-21 | Image processing apparatus | 2016-10-14 | An image processing apparatus includes an input section, a specification section, an extraction section, and a calculation section. The input section receives fluorescence image information obtained by picking up an image of fluorescence based on application of excitation light to a subject provided with a fluorescent substance with a specific effect on a living tissue and therapeutic light position image information including an application position of therapeutic light. The specification section specifies an application region of the therapeutic light. The extraction section extracts a luminance value corresponding to the application region of the therapeutic light and a luminance value corresponding to a region other than the application region of the therapeutic light. The calculation section calculates and outputs a ratio of the extracted luminance value corresponding to the application region to the luminance value corresponding to the region other than the application region. | 1. An image processing apparatus comprising: an input section configured to receive fluorescence image information obtained by picking up an image of fluorescence based on application of excitation light to a subject provided with a fluorescent substance with a specific effect on a living tissue and therapeutic light position image information including an application position of therapeutic light for causing the fluorescent substance to make an effect on a living body;a specification section configured to specify an application region of the therapeutic light in the fluorescence image information based on the therapeutic light position image information inputted to the input section;an extraction section configured to extract a luminance value corresponding to the application region of the therapeutic light specified by the specification section in the fluorescence image information inputted to the input section and a luminance value corresponding to a region other than the application region of the therapeutic light in the fluorescence image information; anda calculation section configured to calculate and output a ratio of the luminance value corresponding to the application region extracted by the extraction section to the luminance value corresponding to the region other than the application region. 2. The image processing apparatus according to claim 1, further comprising a memory configured to output information regarding a progression of the effect on the living body based on the ratio calculated by the calculation section. 3. The image processing apparatus according to claim 2, wherein the memory generates change information indicating a temporal change in the ratio after start of the application of the therapeutic light and displays the change information on a display apparatus configured to display the fluorescence image. 4. The image processing apparatus according to claim 2, further comprising: a set value input circuit configured to input a set value for stopping the application of the therapeutic light; and a controller configured to perform control for stopping the application of the therapeutic light when the ratio calculated by the calculation section or a value calculated based on the ratio becomes a value corresponding to the set value. 5. The image processing apparatus according to claim 2, wherein the memory uses a plurality of times after a time of a start of the application of the therapeutic light and a plurality of values of the ratio respectively calculated by the calculation section in the plurality of times to calculate prediction values of the values of the ratio in time periods after a current time and outputs the prediction values to the display apparatus. 6. The image processing apparatus according to claim 1, wherein the extraction section extracts a region with a luminance value lower than the luminance value corresponding to the application region of the therapeutic light by a predetermined value or more, ... | Olympus Corporation | G06T007/00 A61B001/00 A61N005/06 G02B023/24 A61B001/04 A61B001/06 A61B005/00
US09820678 | B2 | 2017-11-21 | Kinetic assessment and alignment of the muscular-skeletal system and method therefor | 2013-09-13 | A system is disclosed herein for providing a kinetic assessment and preparation of a prosthetic joint comprising one or more prosthetic components. The system comprises a prosthetic component including sensors and circuitry configured to measure load, position of load, and joint alignment. The system further includes a remote system for receiving, processing, and displaying quantitative measurements from the sensors. The kinetic assessment measures joint alignment under loading that will be similar to that of a final joint installation. The kinetic assessment can use trial or permanent prosthetic components. Furthermore, adjustments can be made to the applied load magnitude, position of load, and joint alignment by various means to fine-tune an installation. The kinetic assessment increases both performance and reliability of the installed joint by reducing error that is introduced by elements that load or modify the joint dynamics not taken into account by prior assessment methods. | 1. A system to support kinetic assessment, joint modification, and installation of a final prosthetic joint comprising: a device configured to couple to a joint of a first bone and a second bone to measure load, position of load, and joint alignment, wherein the device includes one or more sensors;a computer configured to receive measurement data from the one or more sensors of the device, wherein kinetic assessment results in at least one change performed under musculoskeletal loading that affects alignment, whereby a final joint installation has similar loading, position of load, or alignment; anda display coupled to the computer, wherein the computer is configured to calculate alignment from the measurement data, and wherein the display is configured to display one of a bone offset relative to a mechanical axis or a total offset relative to the mechanical axis generated from the measurement data. 2. The system of claim 1, wherein the prosthetic component comprises: at least one articular surface;a plurality of load sensors underlying the at least one articular surface configured to measure load and position of load; andat least one inertial sensor configured to measure a movement of the muscular-skeletal system under loading, wherein the measurement data corresponds to and is converted to an alignment measurement. 3. The system of claim 2, wherein the computer is configured to calculate an offset of the first bone relative to the mechanical axis from the measurement data generated by a first movement of the joint. 4. The system of claim 2, wherein the computer is configured to calculate an offset of the second bone relative to the mechanical axis from the measurement data generated by a second movement of the joint. 5. The system of claim 2, wherein the display is configured to display a portion of a surgical field in an operating room. 6. The system of claim 5, wherein the computer and the display are configured to indicate if the total offset is outside a predetermined range. 7. The system of claim 5, wherein the display is configured to display a contact point on the device from the measurement data. 8. The system of claim 2, wherein the display is configured to display contact point rotation using the measurement data. 9. The system of claim 1, wherein the display is configured to display at least a portion of the first bone relative to a portion of the second bone. 10. The system of claim 9, wherein the display is configured to display a flexion angle corresponding to an angle formed by the first bone relative to the second bone. 11. The system of claim 9, wherein the computer and the display is configured to display a anterior-posterior slope of a bone cut on the first bone or the second bone measured by the one or more sensors. 12. The system of claim 1, wherein the computer and the display are configured to provide real-time feedback corresponding to the at least one change as it is being made using the measurement data. 13... | ORTHOSENSOR INC | A61B005/103 A61B005/00 A61B017/15 A61B034/20 A61B034/00 A61B090/00 A61B005/11 A61B017/17 A61F002/38 A61B005/107 G06F003/0481 A61F002/46 A61B034/10
US09775586 | B2 | 2017-10-03 | Urine hat flush water capture device | 2015-08-12 | The present invention relates to a urine specimen capture device comprising a capture bowl to capture a urine specimen, a mounting frame connected to the capture bowl to support the capture bowl within a toilet bowl and one or more flush water capture flange for capturing flush water emitted from a toilet bowl rim during a toilet flush cycle and diverting the flush water to the capture bowl to rinse the capture bowl of the urine specimen. | 1. A urine specimen capture device for analyzing a urine specimen, comprising: a capture bowl to capture a urine specimen;an opening at a bottom of the capture bowl;a mounting frame connected to the capture bowl to support the capture bowl within a toilet bowl;a flush valve within the opening at the bottom of the capture bowl, andone or more flush water capture flange fluidly connected to and extending from the capture bowl to capture flush water emitted from a toilet bowl rim during a toilet flush cycle and divert the flush water to the capture bowl to thereby rinse the capture bowl of the urine specimen,wherein the flush valve is operably connected to a toilet flush mechanism and configured to open upon actuation of the toilet flush mechanism. 2. The capture device according to claim 1, wherein the flush valve is operably connected to a user sensor configured to close the flush valve when the presence of a user on a toilet seat is detected and open the flush valve when the presence of a user on the toilet seat is not detected. 3. The capture device according to claim 2, wherein the user sensor comprises a toilet seatload sensor. 4. The capture device according to claim 2, wherein the user sensor comprises a user proximity sensor. 5. A urine specimen capture device for analyzing a urine specimen, comprising: a capture bowl to capture a urine specimen;a mounting frame connected to the capture bowl to support the capture bowl within a toilet bowl; andone or more flush water capture flange fluidly connected to and extending from the capture bowl to capture flush water emitted from a toilet bowl rim during a toilet flush cycle and divert the flush water to the capture bowl to thereby rinse the capture bowl of the urine specimen, wherein at least one of the one or more capture flange is flexible. 6. A urine specimen capture device for analyzing a urine specimen, comprising: a capture bowl to capture a urine specimen;a mounting frame connected to the capture bowl to support the capture bowl within a toilet bowl;a fluid level sensor;one or more flush water capture flange fluidly connected to and extending from the capture bowl to capture flush water emitted from a toilet bowl rim during a toilet flush cycle and divert the flush water to the capture bowl to thereby rinse the capture bowl of the urine specimen; anda flush valve operably connected to the fluid level sensor and configured to selectively close and open to maintain the urine specimen at a predetermined volume in the capture bowl. 7. The capture device according to claim 6, further comprising a flush valve operably connected to the fluid level sensor and configured to selectively close and open to prevent the urine specimen from exceeding a predetermined volume in the capture bowl. 8. A urine specimen capture device for analyzing a urine specimen, comprising: a capture bowl to capture a urine specimen; a mounting frame connected to the capture bowl to support the capture bowl within a ... | - | A47K003/20 A47K004/00 E03D009/00 A61B010/00 A61B005/00
US09808158 | B2 | 2017-11-07 | Device and method for the non-invasive measurement of state of tension, biomechanical and viscoelastic properties of surfaces of soft biological tissues | 2013-04-12 | A device and method for the non-invasive measurement of state of tension, biomechanical and viscoelastic properties of surfaces of soft biological tissues includes electronic, digital and mechanical elements and sensors, and a testing end. The elements within the housing can subject the testing end to a mechanical impulse force, and can sense the movement of the testing end. The testing end has a contact surface which can be adhered to the surface of the soft biological tissue, and subjected to a mechanical force in a direction parallel to the surface. | 1. A device for measurement of biomechanical and viscoelastic properties of skin of a soft biological tissue with elastic properties, the device comprising: a housing including an actuator shaft provided with an actuator, and an accelerometer;a testing end with a contact surface configured for adhering to a surface of the skin of the soft biological tissue, the testing end configured to be subjected to a mechanical force along an axis of movement of the testing end in a direction parallel to the surface of the skin by actuation of the actuator, the accelerometer being mounted to the actuator shaft so as to move along with the testing end, and arranged to record acceleration of the testing end, the testing end being shaped such that the contact surface extends in a plane that is parallel to the axis of movement of the testing end,wherein the actuator is arranged to subject the testing end to a force such that the skin is subjected to a pre-set tension parallel to the axis of movement of the testing end, and then to apply a brief mechanical impulse parallel to the axis of movement of the testing end, while continuing to apply the pre-tension force, and wherein after the application of the brief mechanical impulse the resulting damped natural oscillation is recorded, as an acceleration curve, with data from the accelerometer; anda signal processing unit for performing signal processing on the data from the accelerometer during the damped natural oscillation to obtain values of the biomechanical and viscoelastic properties of the skin. 2. A device as claimed in claim 1 wherein the contact surface has an area from 20 mm2 to 113 mm2. 3. A device as claimed in claim 1 wherein the contact surface is circular or oval. 4. A device as claimed in claim 1, further comprising a gravitational sensor and a feedback circuit for ensuring that the preset pre-tension and the brief mechanical impulse, in operation of the device, are unaffected by orientation of the housing relative to a gravitational field. 5. A device as claimed in claim 1 wherein the testing end is L-shaped. 6. A device as claimed in claim 1, further comprising a control circuit, wherein the accelerometer is arranged to provide a signal indicative of orientation an inclination of the housing from a horizontal axis to the control circuit to compensate for the inclination from the horizontal axis. 7. A method of measuring biomechanical and viscoelastic properties of skin of a soft biological tissue with elastic properties with a device comprising: a housing including an actuator shaft provided with an actuator, and an accelerometer; and a testing end with a contact surface configured for adhering to a surface of the skin of the soft biological tissue, the testing end being movable along an axis of movement of the testing end in a direction parallel to the surface of the skin by actuation of the actuator, the testing end being shaped such that the contact surface extends in a plane that is paral... | MYOTON AS | A61B005/00 B06B001/04 A61B009/00 G01N003/317 A61B090/00
US09826920 | B2 | 2017-11-28 | Disposable dual tipped ear curette incorporating depth measurement system | 2012-05-25 | A disposable dual ended ear curette incorporates specialized tip configurations and a depth measurement system. The ear curette may comprise a central handle portion; a first tip extending from one end of the handle portion and including an annular array of projections; and a second tip extending from an opposed end of the handle portion and including a blade member having a first lateral scraping edge on one edge of the blade member positioned adjacent a concave surface of the blade member, and a second lateral scraping edge on an opposed edge of the blade member. The ear curette may further include a sequence of spaced, visible measurement indicia on each tip, configured to provide an operator with a visual indication of the depth of operation of the associated tip in a patient's ear. | 1. An ear curette comprising: a central handle portion;a first tip extending from one end of the handle portion and including a base having a surface the outer periphery of which is directly coupled to the handle portion and an annular array of at least three projections extending from the base wherein a proximal end of each projection extends from a common base plane on the surface of the base, and wherein a shape of the array conforms to a shape of at least a portion of the outer periphery of the base; anda second tip extending from an opposed end of the handle portion and including a blade member having a first lateral scraping edge on one edge of the blade member positioned adjacent a concave surface of the blade member, and a second lateral scraping edge on an opposed edge of the blade member. 2. The ear curette of claim 1 wherein the first tip includes having the base formed as a ring member extending along the common base plane and wherein the array of projections extends from the ring member substantially perpendicular to the common base plane. 3. The ear curette of claim 2 wherein the array of projections extend from one side of the ring member, and wherein a side of the ring member opposed from the array of projections includes an annular scraping edge. 4. The ear curette of claim 2 wherein the array of projections extend from one side of the ring member and wherein each projection is formed of one of i) an arc segment projection generally conforming to a radii of a portion of the ring member from which the arc segment projection extends, and ii) a frusto-conical member. 5. The ear curette of claim 4 wherein at least one tip further includes a sequence of spaced, visible measurement indicia configured to provide an operator with a visual indication of a depth of operation of the associated tip in a patient's ear wherein the spaced, visible measurement indicia includes at least one of i) evenly spaced unit gradation indicia and evenly spaced subunit gradations between the unit gradation indicia, ii) alpha-numerical indicia indicative of the depth of operation of the associated tip in the patient's ear, and iii) a spectrum of colors that varies in accordance with the depth of operation of the associated tip in the patient's ear. 6. The ear curette of claim 1 further including a sequence of spaced, visible measurement indicia on at least one tip, configured to provide an operator with a visual indication of a depth of operation of the associated tip in a patient's ear. 7. The ear curette of claim 6 wherein the sequence of spaced, visible measurement indicia includes at least one of i) evenly spaced unit gradation indicia and evenly spaced subunit gradations between the unit gradation indicia, ii) a spectrum of colors that varies in accordance with the depth of operation of the associated tip in the patient's ear, and iii) alpha-numerical indicia indicative of the depth of operation of the associated tip in the patient's ear. 8. The ... | - | A61B005/107 A61F011/00 A61B017/3207 A61B090/00
US09814894 | B2 | 2017-11-14 | Systems and methods for detecting health disorders | 2013-05-31 | According to some embodiments, a wearable medical device capable of treating a patient presenting with syncope is provided. The wearable medical device includes a memory storing event profile information, a battery, at least one treatment electrode coupled to the battery, at least one processor coupled to the memory and the at least one treatment electrode, and an event manager executed by the at least one processor. The event manager is configured to detect an event associated with syncope; store, in the memory, data descriptive of the event in association with an indication that the data includes data descriptive of a syncopal event; and address the event. | 1. A method of treating a patient presenting with syncope, the method comprising: monitoring, by a plurality of physiological sensors of an external wearable defibrillator worn by the patient and configured to provide a therapeutic treatment to the patient via one or more external treatment electrodes disposed in the external wearable defibrillator, a plurality of physiological parameters, the plurality of physiological parameters including at least an electrocardiogram signal of the patient;detecting, by the external wearable defibrillator, an event associated with syncope based on the plurality of physiological parameters and data received from at least one motion sensor disposed on the patient;storing, by the external wearable defibrillator and responsive to detecting the event, data descriptive of the event in association with an indication that the data includes data descriptive of a syncopal event; andaddressing the event. 2. The method according to claim 1, wherein detecting the event includes detecting a cardiac arrhythmia. 3. The method according to claim 2, wherein detecting the cardiac arrhythmia includes detecting at least one of ventricular tachycardia, ventricular fibrillation, and bradycardia. 4. The method according to claim 3, wherein providing the therapeutic treatment includes issuing at least one of a pacing pulse and a defibrillation shock. 5. The method according to claim 1, wherein addressing the event further includes issuing a notification to an external entity. 6. The method according to claim 1, wherein detecting the event further includes identifying a pattern of data in the data received from the plurality of physiological sensors and the at least one motion sensor. 7. The method according to claim 6, wherein identifying the pattern of data includes identifying a pattern of data indicative of a level of dyskinesia. 8. The method according to claim 7, wherein identifying the pattern of data includes identifying a pattern of data indicative of a location of the patient within a space. 9. An external wearable defibrillator capable of treating a patient presenting with syncope, the external wearable defibrillator comprising: a memory storing event profile information;a battery;a plurality of physiological sensors disposed on the patient;at least one motion sensor disposed on the patient;at least one external treatment electrode disposed on the patient and operationally coupled to the battery;at least one processor coupled to the memory and the at least one external treatment electrode; andan event manager executed by the at least one processor and configured to: monitor a plurality of physiological parameters received from the plurality of physiological sensors, the plurality of physiological parameters including at least an electrocardiogram signal of the patient;detect an event associated with syncope based on the plurality of physiological parameters and data received from the at least one motion sensor dispose... | ZOLL MEDICAL CORPORATION | A61N001/39 A61B005/00 A61B005/0402 A61B005/11
US09814429 | B2 | 2017-11-14 | System and method for discrimination of central and obstructive disordered breathing events | 2008-12-22 | Disordered breathing events may be classified as central, obstructive or a combination of central an obstructive in origin based on patient motion associated with respiratory effort. Central disordered breathing is associated with disrupted respiration with reduced respiratory effort. Obstructive disordered breathing is associated with disrupted respiration accompanied by respiratory effort. A disordered breathing classification system includes a disordered breathing detector and a respiratory effort motion sensor. Components of the disordered breathing classification system may be fully or partially implantable. | 1. A method of operating an implantable device to classify disordered breathing events of a patient, comprising: sensing respiration with a respiration sensor and generating a respiration signal indicating patient respiration cycles;sensing chest wall motion with a motion sensor and generating a motion signal indicating chest wall motion associated with inspiratory effort;detecting disordered breathing based on the respiration signal using a disordered breathing detector;time correlating the motion signal with the respiration signal using a processor of the implantable device;classifying the disordered breathing event as a central disordered breathing event in response to the motion signal remaining below a motion threshold during a time period in which the respiration signal remains below an inspiration threshold; orclassifying the disordered breathing event as an obstructive disordered breathing event in response to the motion signal rising above the motion threshold during a time period in which the respiration signal remains below the inspiration threshold; anddelivering a therapy to treat disordered breathing based on the classification of the disordered breathing event with a therapy delivery unit coupled to the processor. 2. The method of claim 1, wherein: the sensing respiration and generating the respiration signal comprises sensing and generating using a transthoracic impedance sensor; andthe sensing chest wall motion and generating the motion signal comprises sensing and generating using an accelerometer. 3. The method of claim 1, wherein the sensing respiration, the sensing chest wall motion, and the detecting disordered breathing are each performed at least in part implantably by the implantable device. 4. The method of claim 3, wherein the classifying is also performed at least in part implantably by the implantable device. 5. The method of claim 1, wherein the sensing respiration is performed using a sensor selected from the group of a transthoracic impedance sensor, a microphone configured to detect snoring sounds, and airflow sensor, and a blood gas sensor. 6. The method of claim 1, wherein sensing chest wall motion includes distinguishing chest wall motion associated with respiratory effort from chest wall motion not associated with respiratory effort. 7. A system for classifying disordered breathing of a patient, comprising: a respiration sensor configured to provide a respiration signal indicative of patient respiration cycles;a motion sensor configured to provide a motion signal indicative of chest wall motion associated with inspiratory effort;a disordered breathing detector coupled to the respiration sensor and configured to detect a disordered breathing event based on the respiration signal;a disordered breathing classification processor of an implantable device coupled to the motion sensor and the disordered breathing detector, the disordered breathing classification processor configured to time correlate the motion... | Cardiac Pacemakers, Inc. | A61B005/08 A61B005/00 A61B005/0205 A61B005/113 A61B005/0476 A61B005/0478 A61B005/0488 A61B005/0496 A61B005/11
US09782229 | B2 | 2017-10-10 | Surgical robot platform | 2013-06-21 | A medical robot system, including a robot coupled to an effectuator element with the robot configured for controlled movement and positioning. The system may include a transmitter configured to emit one or more signals, and the transmitter is coupled to an instrument coupled to the effectuator element. The system may further include a motor assembly coupled to the robot and a plurality of receivers configured to receive the one or more signals emitted by the transmitter. A control unit is coupled to the motor assembly and the plurality of receivers, and the control unit is configured to supply one or more instruction signals to the motor assembly. The instruction signals can be configured to cause the motor assembly to selectively move the effectuator element and is further configured to (i) calculate a position of the at least one transmitter by analysis of the signals received by the plurality of receivers; (ii) display the position of the at least one transmitter with respect to the body of the patient; and (iii) selectively control actuation of the motor assembly in response to the signals received by the plurality of receivers. | 1. A medical robot system, comprising: a robot coupled to an effectuator element, the robot configured for controlled movement and positioning, wherein the robot is configured to receive a three-dimensional (3D) image of a target anatomy of a patient;at least one radiofrequency (RF) transmitter configured to emit one or more RF signals, the at least one RF transmitter being coupled to an instrument coupled to the effectuator element;a motor assembly coupled to the robot, the motor assembly being configured to move the effectuator element along one or more of a first x-axis, a first y-axis, and a first z-axis such that movement of the effectuator element along one of the first x-, y-, or z-axes occurs independently of movement of the effectuator element along the other axes of the first x-, y-, and z-axes;a plurality of RF receivers configured to receive the one or more RF signals emitted by the at least one RF transmitter; anda control unit coupled to the motor assembly and the plurality of RF receivers, the control unit configured to supply one or more instruction signals to the motor assembly, the instruction signals configured to cause the motor assembly to selectively move the effectuator element along at least one of the first x-, y-, and z-axes, the control unit being further configured to(i) calculate a position, relative to the target anatomy of the patient, of the at least one RF transmitter coupled to the instrument by analyzing a time of flight of the one or more signals to the plurality of RF receivers in relation to a known distance between each of the plurality of RF receivers and the 3D image of the target anatomy of the patient;(ii) display the position of the at least one RF transmitter on the 3D image of the target anatomy of the patient; and(iii) selectively control actuation of the motor assembly in response to the one or more RF signals received by the plurality of RF receivers to drive the instrument along a preprogrammed trajectory. 2. The system of claim 1, wherein the effectuator element comprises a tube having a pitch axis, a roll axis and a tube axis defining an axis of rotation of the tube. 3. The system of claim 1, wherein the effectuator element comprises a needle. 4. The system of claim 1, wherein the plurality of RF receivers comprises at least three RF receivers. 5. The system of claim 1, wherein the control unit is configured to receive the one or more RF signals from the plurality of RF receivers in an iterative fashion. 6. The system of claim 1, wherein the control unit is configured to receive the one or more RF signals from the plurality of RF receivers in a dynamic fashion. 7. The system of claim 1, wherein the control unit is configured to allow a user to pick the desired location within the patient's body and to dynamically provide information that allows the motor assembly to move the effectuator element to the desired location. 8. The system of claim 1, wherein the at least one RF transmitter is o... | GLOBUS MEDICAL, INC. | A61B034/30 A61B005/06 A61N001/05 A61B017/02 A61B034/32 A61B010/02 A61B017/16 A61B034/20 A61B090/00 A61B034/00 A61B090/14 A61B034/10
US09795296 | B2 | 2017-10-24 | Mouthpiece for measurement of biometric data of a diver and underwater communication | 2012-04-26 | Embodiments of the invention provide a system, apparatus and methods for measurement of biometric data of a diver. In many embodiments, the system includes a mouthpiece having a sensor device comprising a light emitter and detector configured to emit and detect light at a wavelength having an absorbance correlated with a level of a blood gas saturation e.g., oxygen, nitrogen, CO2. The emitter is positioned to emit light onto oral tissue of the diver and the detector positioned to detect light which is received from the oral tissue either by transmittance of light through the oral tissue or by reflection of light from the tissue. The target oral tissue can include one or both of gum or buccal tissue. Such embodiments allow data to be collected without having to wear additional sensors or measurement devices and allow for measurement of blood gas levels as the diver breaths through their mouthpiece. | 1. A mouthpiece apparatus for outputting biometric data of a diver, the mouthpiece comprising: a flexible mouthpiece configured to be worn in the mouth of the diver, comprising:a processor configured to control an intensity of emitted light;a light emitter positioned to emit light onto an oral tissue surface completely inside the mouth of the diver, the light being at a wavelength having an absorbance correlated with a level of a blood gas saturation, the processor configured to account for reflectance caused by an enamel surface in the mouth of the diver by controlling the intensity of emitted light and the wavelength of emitted light; anda light detector positioned to detect light which is received from the oral tissue and to generate an output signal correlated to the detected light, wherein the processor includes logic configured to modulate a characteristic of the emitted light in response to a diving condition to maintain an accuracy of a blood gas saturation measurement based on the output signal. 2. The mouthpiece apparatus of claim 1, wherein the blood gas saturation is oxygen saturation. 3. The mouthpiece apparatus of claim 2, wherein the wavelength includes at least a first and second wavelength. 4. The mouthpiece apparatus of claim 1, wherein the light emitter and the light detector are positioned to detect light transmitted through the oral tissue. 5. The mouthpiece apparatus of claim 1, wherein the light emitter and the light detector are positioned to detect light transmitted through a lingual frenulum portion of the tongue. 6. The mouthpiece apparatus of claim 1, wherein at least one of the light emitter or the light detector is positioned on a surface of the mouthpiece. 7. The mouthpiece apparatus of claim 1, wherein at least one of the light emitter or the light detector is recessed below a surface of the mouthpiece. 8. The mouthpiece apparatus of claim 1, wherein at least one of the light emitter or the light is embedded below a surface of the mouthpiece. 9. The mouthpiece apparatus of claim 1, wherein the mouthpiece is configured to conform to a contour of the diver's oral tissue to maintain substantial optical contact between the diver's oral tissue and the light emitter and the light detector during movement of the diver's mouth, cheek, jaw or teeth. 10. The mouthpiece apparatus of claim 9, wherein the mouthpiece is configured to maintain substantial optical contact between the diver's oral tissue and the light emitter and the light detector during said movement, without impeding the diver's respiration. 11. The mouthpiece apparatus of claim 1, wherein the processor is coupled to the light emitter and light detector, and configured for controlling a characteristic of the emitted light other than the intensity, and for receiving and converting an output signal from the detector into a blood gas saturation level. 12. The mouthpiece apparatus of claim 11, wherein the characteristic is at least one of a duration or wa... | InCube Labs, LLC | A61B005/1455 A61B005/00 A61B005/08 B63C011/26 B63C011/18
US09788736 | B2 | 2017-10-17 | Blood pressure information measurement device cuff and blood pressure information measurement device provided therewith | 2012-08-23 | A cuff includes a fluid bladder, an outer cover that contains the fluid bladder toward one end portion, a surface fastener, and another surface fastener. The outer cover includes, in an area of the outer cover that includes a region in which the other surface fastener is provided, a first wide portion in which the width of the outer cover is greater than a width of an area of the outer cover that includes a region in which the surface fastener is provided. The other surface fastener includes a second wide portion in which the width of the other surface fastener is greater than the width of the surface fastener. The cuff can be fitted to a measurement area with certainty. | 1. A blood pressure information measurement device cuff comprising: a fluid bladder that applies pressure to a body;an outer cover, comprising a first main surface and a second main surface, that contains the fluid bladder toward a first end portion;a first fastening member provided toward the first end portion on the first main surface; anda second fastening member provided toward a second end portion on the first main surface or toward the second end portion on the second main surface, wherein the outer cover is adapted to be wrapped around the body in a ring shape;the first fastening member and the second fastening member are adapted to hold the wrapped outer cover on the body in a secured state;the outer cover comprises, in an area of the outer cover that comprises a region in which the second fastening member is provided, a first wide portion in which the width of the outer cover is greater than a width of an area of the outer cover that comprises a region in which the first fastening member is provided; andthe second fastening member comprises a second wide portion in which the width of the second fastening member is greater than a width of the first fastening member,wherein the second wide portion of the second fastening member is on the first wide portion of the outer cover,wherein at least one side edge of the outer cover in the first wide portion is curved outwardly from an elongated center line of the outer cover,wherein at least one side edge of the second fastening member in the second wide portion is curved outwardly from the elongated center line of the outer cover,wherein the second end portion of the outer cover is wider than the width of the first end portion of the outer cover, andwherein the second end portion of the second fastening member is wider than the width of the first end portion of the outer cover. 2. The blood pressure information measurement device cuff according to claim 1, wherein the first wide portion in the outer cover and the second wide portion in the second fastening member have approximately arc shapes in which the widths gradually increase outward in the width direction as the outer cover progresses from the first end portion toward the second end portion. 3. The blood pressure information measurement device cuff according to claim 1, wherein the first wide portion in the outer cover and the second wide portion in the second fastening member have shapes in which the widths increase outward on both sides in the width direction as the outer cover progresses from the first end portion toward the second end portion. 4. A blood pressure information measurement device comprising: the blood pressure information measurement device cuff according to claim 1;an inflation/deflation mechanism that inflates/deflates the fluid bladder; anda blood pressure information obtainment unit that obtains blood pressure information. 5. The blood pressure information measurement device cuff according to claim 1, wherein the w... | OMRON HEALTHCARE Co., Ltd. | A61B005/02 A61B005/022
US09805164 | B2 | 2017-10-31 | Method and apparatus for providing contraction information during labour | 2007-03-09 | A system and apparatus for implementing a user interface for displaying uterine contraction information is provided. The graphical user interface displays first information conveying a rate of uterine contractions, the first information being derived at least in part on the basis of at least a portion of a contraction signal. The graphical user interface also displays, concurrently with the first information, second information conveying a threshold rate of uterine contractions. In specific examples of implementation, the graphical user interface is adapted for selectively causing an alarm event based at least in part on a rate of uterine contractions conveyed by the first information and the threshold rate of uterine contractions. | 1. A non-transitory computer readable storage medium storing a program element suitable for execution by a processor, said program element implementing a process for monitoring an obstetrics patient during labour, said processor when executing said program element being configured for: (a) receiving a contraction signal associated with the obstetrics patient, the contraction signal conveying information related to uterine contractions over time;(b) processing the contraction signal to derive first information conveying rates of uterine contractions over time;(c) causing said first information to be displayed in a viewing window on a display device;(d) deriving information conveying if an anomalous contraction rate has persisted for a prolonged period of time, said information being derived at least in part by determining if the rates of uterine contractions conveyed by said first information have fallen outside a limit set by a threshold rate of uterine contractions for a time duration exceeding a predetermined time duration;(e) receiving contraction medication information conveying whether contraction inducing medication was administered to the obstetrics patient;(f) selectively causing an alarm event at least in part based on: (i) the information conveying if the anomalous contraction rate has persisted for the prolonged period of time; and(ii) said contraction medication information. 2. A non-transitory computer readable storage medium as defined in claim 1, said processor when executing said program element being configured for: (a) receiving fetal heart rate information;(b) selectively causing the alarm event based at least in part: (i) on the information conveying if the anomalous contraction rate has persisted for the prolonged period of time;(ii) on said contraction medication information; and(iii) on said fetal heart rate information. 3. A non-transitory computer readable storage medium as defined in claim 1, wherein said first information includes a first tracing conveying the rates of uterine contractions over time; and wherein said processor when executing said program element is operative for causing said first tracing to be displayed in the viewing window concurrently with a second tracing, said second tracing conveying the threshold rate of uterine contractions. 4. A non-transitory computer readable storage medium as described in claim 1, wherein said alarm event includes displaying a visual indicator. 5. A non-transitory computer readable storage medium as described in claim 1, wherein said alarm event includes causing an audio signal to be issued. 6. A non-transitory computer readable storage medium as described in claim 1, wherein said alarm event includes causing a message signal to be transmitted to a remote device. 7. A non-transitory computer readable storage medium as described in claim 6, wherein the remote device is a device selected from the set consisting of a PDA, telephone, pager and computing terminal. 8. A no... | PERIGEN, INC. | G06F019/00 A61B005/00 A61B005/024 A61B005/0444 A61B008/02 A61B008/08
US09792721 | B2 | 2017-10-17 | Method and device for displaying an object | 2012-06-28 | The invention related to a method and a device for displaying an object, in particular biological tissue. Said method having the following steps: a) generating a first image of at least one sub-region of the object using a first device; b) generating a second image of at least one sub-region of the object using a second device; c) ascertaining first coordinates of at least some image points of the second image in a first coordinate system; d) ascertaining second coordinates of the image points of the second image by projecting the first coordinates in a second coordinate system which is different from the first coordinate system and which is assigned to the first device; and e) generating a combined image of the object from the first and the second image using the ascertained second coordinates of the image points of the second image. | 1. A method for determining image points while displaying an object, in particular biological tissue, comprising the steps of: a) generating a first image of at least one sub-region of the object by means of a endoscopy device;b) generating a second image of at least one sub-region of the object by means of a second device;c) ascertaining first coordinates of at least some image points of the second image in a first coordinate system;d) ascertaining second coordinates of at least some of the image points of the second image by projecting the first ascertained coordinates in a second coordinate system which is different from the first coordinate system and which is assigned to the first device; ande) generating a combined image of the object from the first and the second image using the ascertained second coordinates of the image points of the second image, wherein determining image points is achieved by using a sub-region of the object of interest which is marked successively in a marking process: with a multitude of light spots as supporting points, wherein the endoscopy device comprises a navigated light source, orby marking a sub-region of the object of interest with a plurality of supporting points with a separate position determining device,so that generation of the combined image is achieved following the marking process by replacing or overlaying image points of the first image within the selected sub-region of interest with image points of the second image. 2. The method according to claim 1, wherein ascertaining of the second coordinates comprises determining the spatial position of the second device with a position determining device and projecting the first coordinates of the image points in spatial coordinates in a spatial coordinate system assigned to the position determining device. 3. The method according to claim 2, wherein a time delay of the image triggering of the first device, the second device and the position determination of the positioning determining unit is ascertained and is provided for the time coordination by a central processing unit for displaying the combined image. 4. The method according to claim 2, wherein a common spatial reference coordinate system of the first and the second device is not defined by the position determining device, but by at least one marking element attached relative to the object which can be captured by the position determining device such that a movement or the object can be compensated. 5. The method according to claim 2, wherein the sub-region of interest is selected by means of a reconstruction of a three-dimensional area of the object, wherein the reconstruction is generated by using a plurality of images generated by the second device. 6. The method according to claim 5, wherein the reconstruction of three-dimensional region of the object comprises an interpolation between two time successive images of the second device depending on their respective spatial positions ascertaine... | Scopis GmbH | G06T015/00 G06T015/08 A61B001/04 A61B001/00 A61B008/00 A61B005/06 G06T011/60 G06T015/50 G06T007/33 A61B006/03 A61B006/00 A61B008/08 A61B034/20 A61B090/00
US09814425 | B2 | 2017-11-14 | Health monitoring appliance | 2014-12-11 | A monitoring system includes a wearable appliance; and a processor coupled to the wearable appliance to analyze vital data or wellness data. | 1. A home control and monitoring system for providing health and security maintenance, wherein the system comprises a plurality of detection, data storage and control elements connected in a mesh network configuration; and wherein: a) said detection elements comprise a plurality of health monitoring devices for tracking the physiological status of a user, each health monitoring device comprising a health parameter sensor;b) said data storage elements comprise at least one memory for storing data detected by said detection elements;c) said control elements comprise at least one control device for controlling at least one non-health related appliance coupled to said mesh network;d) said detection, data storage and control elements are connected to each other as a mesh network which further comprises a location determination system and said mesh network is connected to at least one base station outside of the mesh network configuration. 2. The system of claim 1 wherein said base station receives the data provided by at least one of the health parameter sensors and makes said data accessible to an authorized third party who can provide an appropriate response. 3. The system of claim 2 which further comprises means for verbal and visual communication between said user and said authorized third party. 4. The system of claim 3, wherein said detection elements further comprise a motion detecting device for detecting the motion of said user, a processor and a health application executed by said processor which processes data from at least one of said detection elements in accordance with the health application to provide analysis of the physical condition of said user to said base station. 5. The system of claim 4, wherein said analysis provides the identification of an activity of said user and provides an alert to said base station if the analysis detects that said user is need of help. 6. The system of claim 4 wherein said at least one health monitoring device is wearable by said user and wherein if said wearable health monitoring device moves beyond the range of said mesh network, the wearable health monitoring device can utilize a cell phone network to establish contact with said base station. 7. The system of claim 4 wherein said health application utilizes data obtained from the at least one health monitoring device to detect the blood pressure of the user. 8. The system of claim 4 wherein said health application utilizes data obtained from the at least one health monitoring device to detect if the user is having a stroke. 9. The system of claim 4 wherein said health application utilizes data obtained from the at least one health monitoring device to detect if the user is having a cardiac problem. 10. The system of claim 1 wherein said at least one health monitoring device is worn on the user's wrist. 11. The system of claim 1 wherein said at least one non-health related device is a kitchen appliance. 12. The system of claim 1 wherein sa... | KONINKLIJKE PHILIPS N.V. | A61B005/00 A61B005/0205 A61B007/04 A61B005/11 A61B005/021 A61B008/00 A61B005/0402
US09820329 | B2 | 2017-11-14 | Method for managing a wireless link between a first device and a second device | 2014-06-06 | The invention relates to a method for managing a wireless link between a first device and a second device. The method includes the steps of polling an activity of a first wireless interface of the first device during a first predetermined lapse of time, suspending the wireless link and polling an activity of a first body-coupled communication interface of the first device during a second predetermined lapse of time when no activity is detected on the first wireless interface of the first device during the first predetermined lapse of time, and resuming the wireless link when at least one polling packet comprising a resume request is detected by the first body-coupled communication interface of the first device during the second predetermined lapse of time. | 1. A method for managing a wireless link between a first device and a second device, said method comprising the steps of: polling an activity of a first wireless interface of the first device (101) during a first predetermined lapse of time (T1);suspending the wireless link and polling an activity of a first body-coupled communication interface of the first device (102) during a second predetermined lapse of time (T2) when no activity is detected on the first wireless interface of the first device during the first predetermined lapse of time;resuming the wireless link (103) when at least one polling packet comprising a resume request is detected by the first body-coupled communication interface of the first device during the second predetermined lapse of time (T2). 2. The method according to claim 1, comprising a step of polling for a link suspend on a second wireless interface of the second device (105) during the first predetermined lapse of time (T1). 3. The method according to claim 1, comprising a step of generating polling packets by a second body-coupled communication interface of the second device (106) when the link is suspended on the second wireless interface of the second device during a third predetermined lapse of time (T3). 4. The method according to claim 3 comprising a step of generating at least a polling packet comprising a resume request by the second body-coupled communication interface of the second device (107) when the second device has data to send to the first device through the wireless link. 5. The method according to claims 3 comprising a step of polling for a link resume on the second wireless interface of the second device (108) during a fourth predetermined lapse of time (T4) after the generation of the polling packet. 6. The method according to claim 3, comprising a step of polling an activity of the second BCC interface of the second device (109) during the fourth (T4) predetermined lapse of time after the generation of the polling packet. 7. The method according to claim 1, comprising a preliminary step of pairing (200) the first device and the second device in order to establish the wireless link between the first device and the second device and a step of iterating the method according to claim 1 in order to maintain said wireless link. 8. The method according to claim 1 comprising a step of terminating the communication between the two devices when: no activity is detected on the first body-coupled communication interface of the first device during the second predetermined lapse of time (T2) or no activity is detected on the second wireless interface or the second BCC interface during a third and fourth predetermined lapse of time (T3/T4). 9. The method according to claim 1 wherein the first device is a contactless device being configured to contain a user's personal information in order for said user to authenticate itself to a secure system and wherein the second device is a proxy between the first ... | GEMALTO SA | H04B007/00 H04W076/04 H04B013/00 H04W052/02 H04W084/18 A61B005/00
US09824440 | B2 | 2017-11-21 | Methods and systems for three-dimensional real-time intraoperative surgical margin evaluation of tumor tissues | 2013-11-20 | In one aspect, the present invention relates to a system evaluating a surgical margin of tumor tissues of a living subject. In one embodiment, the system includes a light source configured to emit a source light; at least one optical probe; a scanner; a spectrometer; and a controller coupled with the scanner and the spectrometer for operably controlling the scanner and the spectrometer. In operation, a working end of the optical probe is positioned proximate to a surface of a specimen of the tumor tissues. A source channel of the optical probe deliver the source light emitted by the light source from the working end to the surface of the specimen, and a plurality of collection channels collect from the working end diffused/reflected light generated from interaction of the source light with the specimen. The spectrometer receives the collected diffused/reflected light to evaluate a margin status of the specimen. | 1. A system for evaluating a surgical margin of tumor tissues of a living subject, comprising: a light source configured to emit a source light;at least one optical probe coupled with the light source, each of the at least one optical probe has a working end, a source channel and a plurality of collection channels, wherein when the working end is positioned proximate to a surface of a specimen of the tumor tissues, the source channel is configured to deliver the source light emitted by the light source from the working end to the surface of the specimen, and the collection channels are configured to collect from the working end diffused/reflected light generated from interaction of the source light with the specimen;a scanner coupled with the at least one optical probe, configured to cooperatively move the specimen and the at least one optical probe so as to probe a plurality of specimen locations, wherein the specimen locations cover the surface of the specimen;a detector coupled with the at least one optical probe, configured to receive the collected diffused/reflected light to evaluate a margin status of the specimen; anda controller coupled with the scanner and the detector, configured to operably control the scanner and the detector,wherein each optical probe comprises a plurality of optical fibers spatially arranged in a fiber array;wherein for each optical probe, the optical fibers comprise a source optical fiber for delivering the source light emitted from the light source to the surface of the specimen, and a plurality of collection optical fibers for collecting diffused/reflected light generated from the interaction of the source light with the specimen;wherein the source optical fiber is positioned in a center of the fiber array, and the plurality of collection optical fibers is positioned in one or more rings having a center at the source optical fiber such that each collection optical fiber is offset from the source optical fiber;wherein the plurality of collection optical fibers is divided into N groups, such that a first group of the collection optical fibers each is positioned away along a radial direction from the source optical fiber with a first spatial offset R1, and a N-th group of the collection optical fibers each is positioned away along the radial direction from the source optical fiber with a N-th spatial offset RN, wherein N is an positive integer, and R1≦R2≦ . . . ≦RN-1≦RN; andwherein the tumor tissues include breast cancer tissues, N=3, R1=1.57 mm, R2=2.68 mm, and R3=3.50 mm. 2. The system of claim 1, wherein the detector comprises a spectrometer. 3. The system of claim 1, wherein the detector comprises a charge-coupled device (CCD). 4. The system of claim 1, wherein the detector is configured to: generate optical data from the received collected diffused/reflected light;obtain a raw spectrum from the optical data;generate a fluorescence spectrum (FS) and a Raman spectrum (RS) from the raw spectrum; anddetermine th... | VANDERBILT UNIVERSITY | G06T007/00 G01N021/64 G01N021/65 A61B005/00
US09789315 | B2 | 2017-10-17 | Method and system for modulating neural activity | 2016-06-06 | Methods and related systems for modulating neural activity by cyclically modulating neural activity in peripheral neural structures are disclosed. Neural activity may be modulated cyclically by stimuli delivered via various types of stimulus sources. In an aspect, activity of a sensory nerve is modulated. Neural modulation may be used, for example, to modulate an undesired sensation, such as pain, or an immune or inflammatory response or process. Delivery of stimuli for modulating neural activity may be controlled in part in response to an input from a user input device. | 1. A neural modulation system comprising: a first signal input structure adapted to receive a user input signal indicative of an instruction from a user regarding operation of a neural modulation system to deliver a stimulus configured to reversibly modulate neural activity in at least a portion of a peripheral neural structure of a subject, the user input signal transmitted from a user input device adapted to receive a user input indicative of the instruction from the user;a second signal input structure adapted to receive an activity signal indicative of an activity state of at least a portion of a body of the subject innervated by the peripheral neural structure from a sensor operatively connected to the second signal input structure; anda signal processing portion including circuitry for modifying operation of the neural modulation system responsive to receipt of the user input signal;circuitry for distinguishing a first activity state of the at least a portion of the body of the subject innervated by the peripheral neural structure from a second activity state of the at least a portion of the body of the subject innervated by the peripheral neural structure from the activity signal received at the second signal input structure; andcircuitry for generating a stimulus control signal for driving production of the stimulus by a stimulus source, the stimulus configured to modulate neural activity in at least a portion of a peripheral neural structure of the subject during at least a portion of the first activity state. 2. The neural modulation system of claim 1, wherein the circuitry for generating a stimulus control signal includes circuitry for generating a blocking stimulus control signal for driving production of a blocking stimulus by the stimulus source. 3. The neural modulation system of claim 1, wherein the circuitry for generating a stimulus control signal includes circuitry for generating a stimulus control signal for driving production of the stimulus by a stimulus source implanted in the body of the subject. 4. The neural modulation system of claim 1, wherein the circuitry for generating the stimulus control signal includes circuitry for generating the stimulus control signal for driving production of the stimulus by a stimulus source positioned on or substantially adjacent the body of the subject. 5. The neural modulation system of claim 1, wherein the circuitry for generating a stimulus control signal includes circuitry for generating the stimulus control signal for driving production of an electrical stimulus, a magnetic stimulus, an ultrasonic stimulus, an optical stimulus, a thermal stimulus, or a chemical stimulus. 6. The neural modulation system of claim 1, wherein the first signal input structure is configured to receive the user input signal from a cell phone. 7. The neural modulation system of claim 1, wherein the first signal input structure is configured to receiving the user input signal from a remote controller. 8... | Gearbox, LLC | A61N001/18 A61M035/00 A61N001/36 A61B005/11 A61B005/00 A61F007/00 A61M005/142 A61M005/172 A61M037/00 A61N001/04 A61N007/00 A61B005/0205 A61B005/021 A61B005/024 A61B005/0402 A61B005/0476 A61B005/0488 A61B005/145 A61M005/168 A61N001/30 A61N001/32 A61N001/372 A61N002/00
US09801577 | B2 | 2017-10-31 | Sensitivity calibration of in vivo sensors used to measure analyte concentration | 2017-06-07 | Methods, devices, systems, and computer program products are provided to improve sensitivity calibration of an in vivo analyte sensor and usability of an associated analyte monitoring system. In certain embodiments, methods are provided that improve the user experience of using an analyte monitoring system. Certain embodiments of the present disclosure include features that reduce the amount of calibration or re-calibration performed by the analyte monitoring system. More specifically methods of using a suspect calibration attempt to avoid having to recalibrate by adjusting the calibration or mitigating effects of sensor signal attenuation that caused the calibration attempt to be suspect are provided. Additional features are disclosed. | 1. A method, comprising: determining a sample lot sensitivity associated with sensors in a manufacturing lot;selecting a sensor in the lot;for the selected sensor, determining one or more sensor-specific correction factors;determining a sensor-specific sensitivity for the selected sensor based on the one or more sensor-specific correction factors;updating a sensor code based on the sensor-specific sensitivity; andcalibrating analyte level data obtained by the selected sensor based on the updated sensor code. 2. The method of claim 1, wherein the sample lot sensitivity is a median sensitivity for the sensors of the lot. 3. The method of claim 1, wherein the one or more sensor-specific correction factors are predetermined. 4. The method of claim 3, wherein the one or more sensor-specific correction factors are predetermined based on prior sensor use. 5. The method of claim 1, wherein the sample lot sensitivity corresponds to a nominal sensor code sensitivity. 6. The method of claim 1, wherein the one or more sensor-specific correction factors include a correction factor dependent on a physical characteristic of the sensor. 7. The method of claim 6, wherein the correction factor dependent on the physical characteristic of the sensor is further dependent on a period of time of sensor wear. 8. The method of claim 1, further comprising displaying an analyte concentration value corresponding to the calibrated analyte level data. 9. The method of claim 1, wherein the selected sensor is one of a glucose sensor or a lactate sensor. 10. The method of claim 1, wherein the analyte level data is calibrated in real time. 11. The method of claim 1, further comprising transmitting the calibrated analyte level data to a remote device. 12. The method of claim 11, wherein the calibrated analyte level data is transmitted via a radio frequency (RF) communication link. 13. The method of claim 11, wherein the calibrated analyte level data is transmitted in one or more batches in response to a request from the remote device. 14. The method of claim 1, further comprising determining a risk distribution for the selected sensor. 15. The method of claim 1, further comprising delivering medicament with a medicament delivery device based on the calibrated analyte level data. 16. The method of claim 1, further comprising modifying a stored user profile based on the calibrated analyte level data. | Abbott Diabetes Care Inc. | A61B005/00 A61B005/1495 A61B005/145
US09801542 | B2 | 2017-10-31 | Health monitoring appliance | 2015-11-13 | A heart monitoring system for a person includes one or more wireless nodes; and wearable appliance in communication with the one or more wireless nodes, the appliance monitoring vital signs. | 1. A wearable device for use in a wireless body area network which comprises one or more additional devices, wherein said network is configured as a mesh network and said wearable device and said one or more additional devices are placed on one or more parts of a user's body and each of said wearable device and said one or more additional devices each comprise a sensor for monitoring a vital parameter related to the health of the user, and wherein at least one of said sensors is woven into a garment worn by the user and coupled to the network using optical fibers. 2. The wearable device of claim 1 wherein said body area network operates on an assigned medical frequency band and one or more of said sensors can detect a vital health parameter indicating the user's need for assistance with respect to the vital health parameter. 3. The wearable device of claim 2 further comprising a wireless transceiver for communicating data indicative of the user's need for assistance to a receiver located outside of the wireless body area network. 4. The wearable device of claim 3 wherein said wireless transceiver communicates over cell phone networks. 5. The wearable device of claim 3 wherein said wireless transceiver communicates over a WLAN. 6. The device of claim 2 wherein the vital parameter is cerebrovascular in nature. 7. The device of claim 2 wherein the vital parameter is cardiovascular in nature. 8. The device of claim 2 wherein the vital parameter is indicative of SIDS (sudden infant death syndrome). 9. The device of claim 2 wherein a vital parameter is glucose level. 10. The device of claim 2 wherein a vital parameter is arterial blood pressure. 11. The device of claim 2 wherein a vital parameter is a body motion of the user. 12. The device of claim 2 wherein a vital parameter is a sudden change in activity. 13. A method for maintaining wellness in a user comprising the following steps: a. using one or more sensors physically coupled to the user and connected to a network, to monitor one or more vital parameters, providing data representative of the user's physical condition;b. analyzing the one or more vital parameters using a statistical analyzer, trained with training data representing physiological conditions determined to be undesirable for the user to analyze the vital parameters to determine if the physiological conditions are undesirable; andc. generating a warning indication when the user's physical condition is undesirable. 14. The method of claim 13 wherein a vital parameter is arterial blood pressure. 15. The method of claim 13 wherein a vital parameter is body motion patterns of the user. 16. The method of claim 13 wherein a vital parameter is a sudden change in activity. 17. The system of claim 16 further comprising a microphone and speaker which are easily accessible to the user and one of the transceivers. 18. The system of claim 16 wherein one or more of the sensors detects bioelectrical impedance. 19. The system of c... | KONINKLIJKE PHILIPS N.V. | A61B005/00 A61B005/021 A61B005/0402 A61B005/0476 A61B005/11 A61B007/04 G08B021/02 A61B005/0205 A61B005/01 A61B005/03 A61B005/1455 A61B005/20 A61B005/0488 A61B005/0492 A61B005/145 A61B005/08 A61B005/0215 A61B005/022 A61B005/024
US09808183 | B2 | 2017-11-07 | Device and method for real-time measurement of parameters of mechanical stress state and biomechanical properties of soft biological tissue | 2011-07-07 | A device and a method for simultaneous recording, in real time, parameters characterising the mechanical tension, elasticity, dynamical stiffness, creepability and mechanical stress of soft biological tissue are provided. By means of the myometer, a constant external pre-pressure is created, independently of the device's position, between the tissue and the testing end of the device. Next, the tissue is subjected to a short-term external dynamic influence. A mechanical change in the shape of the tissue and its mechanical response are registered as a graph of the tissue's oscillations. For calculating the parameters, a time span on the graph is used which involves an oscillation period from the beginning to the end of the effect on the tissue plus its subsequent first 1.5 self-oscillation period. This enables recording and data-processing to be carried out simultaneously as well as statistically significant estimates to be made in real time. | 1. A device for simultaneous measurement of the parameters characterising the mechanical stress, elasticity, dynamic stiffness, creep and mechanical stress relaxation time of soft biological tissue, comprising: a body, a processor and controller for managing the measuring process and calculating the parameters, a testing end having a movement axis and being movable along its movement axis, a drive of the testing end comprising a movable frame and a solenoid, one end of the testing end being fixed to the movable frame, position sensors for the movable frame, and an accelerometer arranged to sense movement of the testing end,wherein the drive of the testing end is operable in a translational motion having the same direction as the movement axis of the testing end, the movable frame being supported by elastic plate elements so the drive of the testing end is operable with negligible mechanical friction,and the device also comprises light or sound signals arranged to indicate, in response to data from the position sensors, when the elastic plate elements are stress free, when an impulse of current may be transmitted to the solenoid to exert an external impact on the biological tissue; andwherein the controller is arranged such that in use an impulse of current is transmitted to the solenoid to exert the external impact on the biological tissue, after which the biological tissue along with the testing end, the movable frame and the accelerometer undergoes a series of damped oscillations, from which the processor is arranged to calculate the parameters characterising the biological tissue. 2. A device as claimed in claim 1, wherein the moving frame contains a sleeve with permanent magnets placed in the centre of the solenoid, the magnets being oriented with south poles face to face or oriented with north poles face to face. 3. A device as claimed in claim 2, wherein the one end of the testing end that is fixed to the movable frame includes an electrical steel cone-shaped tip that is placed in the pull zone of the permanent magnet which is located closer to the testing end, so said cone-shaped tip is fixed selfrigidly vis-à-vis the movable frame. 4. A device as claimed in claim 3, wherein said device is equipped with a friction-free element for carrying signals from the movements sensor from the movable frame to the processor and controller, the friction-free element being a flexible flat cable. 5. A device as claimed in claim 1 wherein the light signals or sounds signals are placed around an aperture in the body through which the testing end may project. 6. A device as claimed in claim 5, also comprising an arrester system comprising a drive, an actuating screw, a slider moving relative to the body, wherein the slider is equipped with a shutter, and wherein the device comprises stoppers at the extreme positions of the movable frame, together with position sensors mounted on the body. 7. A device as claimed in claim 1 wherein the elastic plates ha... | MYOTON AS | A61B005/00 A61B005/103 A61B009/00 H01F007/122 H01F007/16 G01D005/20
US09795718 | B1 | 2017-10-24 | Biocompatible devices with dissolvable substrates and methods of forming the same | 2016-08-05 | The present disclosure provides a method of forming a biocompatible structure that includes forming biodissolvable substrate comprising a flexible network of peptides, and a biocompatible structure having a biodissolvable substrate and, optionally, an electronic device on a surface thereof for use in implantable electronics. | 1. A method of forming a biocompatible structure, said method comprising: forming a solution comprising a plurality of peptides;depositing said solution on a surface of a wafer, andforming a biodissolvable substrate on said surface of said wafer. 2. The method of claim 1, further comprising forming an electronic device on an exposed surface of said biodissolvable substrate. 3. The method of claim 2, wherein said electronic device comprises at least one semiconductor component. 4. The method of claim 1, wherein forming said solution comprises mixing 9-fluorenylmethoxycarbonyl (Fmoc) or pyrene (Pyr) and a plurality of peptide amphiphiles in water. 5. The method of claim 4, wherein said solution comprises 0.1 mg/ml of peptide to 10 mg/ml of peptide. 6. The method of claim 4, wherein a first subset of said plurality of peptide amphiphiles are two amino acids in length. 7. The method of claim 6, wherein a second subset of said plurality of peptide amphiphiles are at least 3 amino acids in length. 8. The method of claim 4, wherein said peptide amphiphiles of said plurality of peptide amphiphiles are either two amino acids in length or three amino acids in length. 9. The method of claim 1, wherein said deposition comprises spin coating said solution on an exposed surface of said wafer at a spin rate between 500 rpm and 3000 rpm. 10. The method of claim 9, wherein said forming said biodissolvable substrate comprises heating said solution to a temperature of at least 30° C. and subsequently cooling said solution to a temperature of about 20° C. 11. A biocompatible structure comprising: a biodissolvable substrate, wherein said biodissolvable substrate comprises a flexible network of peptides; andan electronic device having at least one semiconductor component on said substrate. 12. The structure of claim 11, wherein said flexible network of peptides comprises a mixture of 9-fluorenylmethoxycarbonyl (Fmoc) or pyrene (Pyr) and a plurality of peptide amphiphiles. 13. The structure of claim 11, wherein said biodissolvable substrate has a thickness from 100 nanometers to 500 microns. 14. The structure of claim 11, wherein at least one of said semiconductor components is a flexible semiconductor component or a stretchable semiconductor component. 15. The structure of claim 11, wherein a bottommost surface of said electronic device is in direct physical contact with a topmost surface of said biodissolvable substrate. 16. A biocompatible structure comprising: a biodissolvable substrate, wherein said biocompatible substrate has a Young's modulus selected from the range of 0.5 MPa and 10 GPa; andan electronic device having at least one semiconductor component on said substrate. 17. The structure of claim 16, wherein at least one of said semiconductor components is a flexible semiconductor component or a stretchable semiconductor component. 18. The structure of claim 16, wherein a bottommost surface of said electronic device is in direct physical con... | INTERNATIONAL BUSINESS MACHINES CORPORATION | H01L051/00 G01N031/22 A61L031/14 A61L031/00 A61B005/11 H01L023/498
US09788737 | B2 | 2017-10-17 | Vital signs measurement system, detecting method of the vital signs measurement system, and vital signs measurement earphone | 2015-01-10 | A vital signs measurement system includes a plurality of light sources emitting into a subject's skin. A plurality of photo sensors receives lights reflected from the subject's skin and converts the lights to a plurality of signals. A processing module receives the plurality of signals and transforms the plurality of signals to a PPG signal by analyzing a correlation coefficient between every two ones of the plurality of signals. The vital signs measurement system improves the measurement accuracy of the physiological information of the participant by the correlation coefficient. | 1. A vital signs measurement system to monitor heartbeat of a subject, comprising: a plurality of light sources adapted to emit light into a subject's skin;a plurality of photo sensors adapted to receive light reflected from the subject's skin and converting the received light to a plurality of signals;a processing module receiving the plurality of signals and transforming the plurality of signals to a PPG (Photoplethysmography) signal by calculating a correlation coefficient between every two ones of the plurality of signals, the correlation coefficient being calculated as follows:Rxy=Σi⁡[(xi-x_)⁢(yi-y_)]Σi⁡(xi-x_)2·Σi⁡(yi-y_)2where Rxy is the correlation coefficient between the every two ones of the plurality of signals, x and y are two signal values corresponding to the every two ones of the plurality of signals respectively;wherein if the correlation coefficient is within an allowable value, then sum all of the plurality of signals to obtain the PPG signal; if not, use a following formula to get a correlation signal between the every two ones of the plurality of signals, the correlation signal is expressed in Sxy: S=(1−Rxy)(x+y)then sum all of the correlation signal Sxy to obtain the PPG signal to monitor heartbeat of the subject. 2. The vital signs measurement system as claimed in claim 1, wherein the PPG signal is transformed from a time domain into a frequency domain by fast Fourier transformation to monitor heartbeat. 3. The vital signs measurement system as claimed in claim 1, wherein the plurality of photo sensors includes a first photo sensor, a second photo sensor and a third photo sensor which convert the light reflected from the subject's skin to a first signal S1, a second signal S2 and a third signal S3 respectively, the processing module processes the first signal S1, the second signal S2 and the third signal S3 to eliminate noise, then use the correlation coefficient (Rxy) formula to get three correlation coefficients R12, R23 and R31, if all of the three correlation coefficients are within the allowable value, use the following formula to obtain the PPG signal: PPG signal=S1+S2+S3if not, use the following formula to obtain the PPG signal: PPG signal=(1−R12)(S1+S2)+(1−R23)(S2+S3)+(1−R31)(S1+S3)then the PPG signal is transformed from a time domain into a frequency domain by fast Fourier transformation to monitor heartbeat. 4. The vital signs measurement system as claimed in claim 3, further comprising a gravity sensing element which is configured to sense a vibration and then convert the vibration to a motion signal, the processing module compares an amplitude of the PPG signal with an amplitude of the motion signal. 5. The vital signs measurement system as claimed in claim 1, wherein the photo sensors transmit the plurality of signals into the processing module through an analog front-end sensor. 6. The vital signs measurement system as claimed in claim 1, further comprising a flash memory and a wireless communication modul... | CHENG UEI PRECISION INDUSTRY CO., LTD. | A61B005/02 A61B005/024 A61B005/00
US09827065 | B2 | 2017-11-28 | Tomography apparatus with integrated lighting | 2013-10-24 | A tomography apparatus has a gantry with an exterior surface, and a tunnel-shaped opening proceeding through the gantry that defines an examination region, from which tomographic data are acquired from a patient in the opening. An illuminant is integrated into the gantry. The illuminant has a light exit window with a smooth transition to the exterior surface of the gantry. | 1. A tomography apparatus comprising: a gantry having a tunnel-shaped opening proceeding therethrough in an axial direction and defining an examination region adapted to receive a patient therein, said gantry having an exterior surface said tunnel-shaped opening having a constant diameter along said axial direction;a plurality of electronic components configured to rotate around said tunnel-shaped opening in order to acquire tomographic data from a portion of the patient situated in said examination region of said gantry;said gantry comprising a funnel-shaped taper leading to an entrance of said tunnel-shaped opening, said funnel-shaped taper having a smallest diameter, adjacent to said tunnel-shaped opening, and a largest diameter at an end of said funnel-shaped opening farthest from said tunnel-shaped opening; andan illuminant integrated into said funnel-shaped taper of said gantry, said illuminant comprising a light exit window through which light proceeds from said funnel-shaped taper so as to be oriented to illuminate at least a portion of said examination region in said tunnel-shaped opening, said light exit window forming a portion of said exterior surface of said gantry and having a smooth transition to a remainder of said exterior surface of said gantry. 2. A tomography apparatus as claimed in claim 1 wherein said illuminant comprises a plurality of individual lighting elements, and wherein said electronic components include a lighting control component configured to control the individual lighting elements separately from each other. 3. A tomography apparatus as claimed in claim 1 wherein said illuminant emits said light along a propagation direction, and wherein said illuminant is configured to allow said direction of propagation of said light to be selectively adjusted. 4. A tomography apparatus as claimed in claim 1 wherein said illuminant is configured to simultaneously illuminate at least said portion of said examination region from multiple directions, thereby reducing a shadow cast by said illuminant. 5. A tomography apparatus as claimed in claim 1 wherein said illuminant proceeds radially along said tunnel-shaped opening. 6. A tomography apparatus as claimed in claim 1 wherein said illuminant proceeds axially along said tunnel-shaped opening. 7. A tomography apparatus as claimed in claim 1 wherein said illuminant is configured to direct light emitted thereby onto a localized region of said examination region. 8. A tomography apparatus as claimed in claim 1 wherein said illuminant is comprised of at least one light-emitting diode. 9. A tomography apparatus as claimed in claim 1 wherein said illuminant is comprised of at least one halogen lamp. 10. A tomography apparatus as claimed in claim 1 wherein said plurality of components are configured to obtain computed tomography data from said portion of said patient in said examination region, as said tomographic data. 11. A tomography apparatus as claimed in claim 1 wherein ... | Siemens Aktiengesellschaft | A61B090/30 A61B006/08 A61B006/00 A61B005/055 A61B006/03 G01R033/30 G01R033/34 G01R033/48
US09795340 | B2 | 2017-10-24 | Vacuum-pump sucker | 2010-12-13 | This invention relates to a vacuum-pump sucker for high-resolution microscopy comprising a sucker body and a transparent plate. The vacuum-pump sucker is designed as a stabilizer for sample stabilization in a clinical optical virtual biopsy system with sub-micron resolution. The sucker is connected with a vacuum pump. As the vacuum pump pumps out air or gas, tissues around the imaged area will be sucked by the sucker and the distance between objective lens and the imaged tissue can be stabilized. In this way, the stability and resolution of the clinical biopsy system can be greatly improved. | 1. A vacuum-pump sucker to stabilize a tissue of a subject for an imaging system of a microscopy device having a focusing lens which is movable for changing observation direction and depth of images, comprising a sucker body, wherein(1) the sucker body allows the focusing lens to be removably inserted into the sucker body in order for the focusing lens to be independently moved within the sucker body for changing observation direction and depth of images,(2) one end of the sucker body is adapted to fix the vacuum-pump sucker onto a main body of the imaging system of the microscopy device,(3) another end of the sucker body provides a contact surface facing or contacting the tissue of the subject, and(4) the sucker body comprises a rigid material;a single vacuum pump;an imaging window positioned at the center of the contact surface of the sucker body to define an imaging window zone which is not connected with the single vacuum pump, wherein the imaging window allows an imaging light to pass through the focusing lens onto the tissue and the tissue to be observed via the imaging window; anda ring-shaped channel positioned at the circumference of the imaging window to define a channel zone, wherein the ring-shaped channel is connected with the single vacuum pump to create a suction only in the channel zone and to pump gas or liquid out of the channel zone;wherein the channel zone and the imaging window zone are dislocated and separated by a barrier to allow the suction to be created by the vacuum pump only in the channel zone, not in the imaging window zone so that only gas or liquid in the channel zone is pumped out and only the tissue of the subject in the channel zone is sucked, thereby stabilizing the tissue in the imaging window zone. 2. The vacuum-pump sucker of claim 1, further comprising a transparent plate covering the imaging window for protecting the focusing lens from contamination. 3. The vacuum-pump sucker of claim 2, wherein the transparent plate is a glass plate or a plastic plate. 4. The vacuum-pump sucker of claim 1, wherein the sucker body further comprises a media injection spot for injecting oil media or water-media. 5. The vacuum-pump sucker of claim 1, wherein the sucker body comprises metal or alloy. | NATIONAL TAIWAN UNIVERSITY | A61B005/00
US09775531 | B2 | 2017-10-03 | Sensor device, processing device, and measurement system for acquiring a biopotential | 2013-07-19 | A measurement system for acquiring biopotentials is provided, which comprises a sensor device and a processing device, which are electrically connected to cooperate during use. The measurement system may be capable of simultaneous measurement of biopotentials by the sensor device and wireless transmission of amplified biopotentials by the processing device, and may be provided with means for preventing the measurements from being disturbed by the transmissions. Also disclosed are (i) a sensor device for measuring a biopotential having at least one sensor with a measurement electrode for measuring the biopotential at a measurement portion of a body and (ii) a processing device for processing biopotentials that has at least one sensor input for receiving the amplified biopotentials measured by a sensor of the sensor device. | 1. A device for measuring a biopotential Vp, comprising at least one sensor comprising: (i) a measurement electrode for measuring the biopotential Vp at a measurement portion of a body,(ii) a measurement connection for a lead,(iii) a first reference connection for receiving a reference potential Vref, and(iv) a differential amplifier connected to the measurement connection and arranged for generating an amplified biopotential VP at a differential amplifier output by amplification of an electrical potential difference between the measured biopotential Vp and the reference potential Vref, wherein the differential amplifier comprisesa first operational amplifier with an input connected to the reference potential Vref,a second operational amplifier with an input connected to the reference potential Vref and a further input connected to an output of the first operational amplifier,wherein a second operational amplifier output is connected in parallel with the measured biopotential Vp, forming a feedback loop with the first operational amplifier,and wherein the differential amplifier further comprises a low-pass filter arranged to filter the biopotential Vp prior to entering the first operational amplifier, and a high-pass filter in combination with the second operational amplifier. 2. The device according to claim 1, further comprising a reference electrode with a second reference connection, arranged for supplying the reference potential Vref to a reference portion of the body. 3. The device according to claim 1, wherein the at least one sensor further comprises a power connection for receiving a power signal for feeding the differential amplifier. 4. The device according to claim 3, wherein the measurement connection, the first reference connection, and the power connection are integrally formed and combined with the lead. 5. The device according to claim 1, wherein the differential amplifier comprises at least one RF-radiation hardened operational amplifier arranged for shielding the at least one sensor from external EM RF fields. 6. A processing device for processing biopotentials, comprising: (i) at least one sensor input for receiving an analog amplified biopotential VP measured by a sensor,(ii) a reference voltage source for generating a reference potential Vref,(iii) an amplifier power source for generating a power signal, and(iv) at least one reference output and at least one power output, which are arranged for outputting the reference potential Vref to a differential amplifier and the power signal to the sensor, the differential amplifier comprises a first operational amplifier with an input connected to the reference potential Vref,a second operational amplifier with an input connected to the reference potential Vref and a further input connected to an output of the first operational amplifier,wherein a second operational amplifier output is connected in parallel with the measured biopotential Vp, forming a feedback loop with the firs... | Techmedic Development International B.V. | A61B005/0428 A61B005/04 A61B005/00
US09782123 | B2 | 2017-10-10 | Integrated resuscitation | 2007-10-30 | A resuscitation system for use by a rescuer for resuscitating a patient, comprising at least two high-voltage defibrillation electrodes, a first electrical unit comprising circuitry for providing resuscitation prompts to the rescuer, a second electrical unit separate from the first unit and comprising circuitry for providing defibrillation pulses to the electrodes, and circuitry for providing at least one electrical connection between the first and second units. In another aspect, at least two electrical therapy electrodes adapted to be worn by the patient for extended periods of time, circuitry for monitoring the ECG of the patient, an activity sensor adapted to be worn by the patient and capable of providing an output from which the patient's current activity can be estimated, and at least one processor configured for estimating the patient's current activity by analyzing the output of the activity sensor, analyzing the ECG of the patient, and determining whether electrical therapy should be delivered to the electrodes. | 1. A resuscitation system for use by a rescuer for resuscitating a patient, comprising: at least two high-voltage defibrillation electrodes;a first electrical unit comprising circuitry for monitoring at least an ECG signal;a second electrical unit mechanically separate from the first unit and comprising circuitry for providing defibrillation pulses to the electrodes; andat least one detachable electrical connection between the electrodes and the second unit for transmitting the defibrillation pulses to the electrodes;wherein the two electrodes and the first unit are configured to be worn by the patient while the patient is moving;wherein the second unit is configured to be electrically attached to the electrodes during some periods of time, and electrically detached during other periods; andwherein the at least two high-voltage defibrillation electrodes are adapted to be worn by the patient for extended periods of time. 2. The resuscitation system of claim 1 wherein the second unit is configured to be mechanically detachable from the first unit when the second unit is electrically detached from the electrodes. 3. The resuscitation system of claim 1 wherein the first unit is separate from the two electrodes, and connected to the two electrodes by one or more cables. 4. The resuscitation system of claim 1 wherein the first unit is capable of providing resuscitation prompts without being electrically connected to the second unit. 5. The resuscitation system of claim 4 wherein the first unit comprises a source of electrical power and a processor. 6. The resuscitation system of claim 4 wherein the first unit comprises a speaker for providing the resuscitation prompts. 7. The resuscitation system of claim 4 wherein the resuscitation prompts comprise spoken and visual prompts. 8. The resuscitation system of claim 1 wherein the resuscitation prompts comprise CPR prompts. 9. The resuscitation system of claim 1 wherein the circuitry for providing at least one detachable electrical connection between the electrodes and second unit comprises at least one cable providing a direct electrical connection to the electrodes. 10. The resuscitation system of claim 1 wherein the at least one detachable electrical connection comprises one or more cables that carry the defibrillation pulses to the electrodes. 11. The resuscitation system of claim 1 wherein the first unit comprises a microphone and circuitry for storing sounds recorded during use of the unit. 12. The resuscitation system of claim 1 wherein the second unit is configured to be supported on the patient when electrically attached. 13. The resuscitation system of claim 1 wherein the second unit delivers defibrillation pulses to the electrodes via an electrical connection from the second unit to the first unit, and the first unit delivers the defibrillation pulses to the electrodes via an electrical connection from the first unit to the electrodes, and wherein the connection from the second unit t... | ZOLL Medical Corporation | A61N001/39 A61N001/04 A61B005/00 A61H031/00 A61B005/11 A61B005/024 A61B005/0402 A61B005/1455 A61B007/02 A61N001/362 A61B005/0245 A61B005/04
US09789208 | B2 | 2017-10-17 | Synthesis and composition of amino acid linking groups conjugated to compounds used for the targeted imaging of tumors | 2016-09-08 | The present disclosure relates to compounds that are useful as near-infrared fluorescence probes, wherein the compounds include i) a pteroyl ligand that binds to a target receptor protein, ii) a dye molecule, and iii) a linker molecule that comprises an amino acid or derivative thereof. The disclosure further describes methods and compositions for incorporating the compounds as used for the targeted imaging of tumors. Conjugation of the amino acid linking groups increase specificity and detection of the compound. Methods and compositions for use thereof in diagnostic imaging are contemplated. | 1. A method of imaging a carcinoma that expresses a folate receptor comprising the steps of: a. administering to a patient in need of an effective amount of a compound capable of binding to a cancerous cell having the formula:or a pharmaceutically acceptable salt or isotope thereof, wherein: X is an single amino acid or a single amino acid derivative thereof, wherein the single amino acid or single amino acid derivative contains an —OH, —NH2, or —SH functional group, and Y is a dye that has a fluorescence excitation and emission spectra in the near infra-red range, wherein Y is represented by the formula:wherein, R1 is independently selected from the group consisting of O, S, N and C, and R2 is independently selected from the group consisting of CH2 and CH2CH2, and the compound enhances the fluorescence of the dye, Y; and b. fluorescent imaging of an area of concern in the patient's body where the compound has been bound to a cancerous cellwherein the carcinoma is selected from the group consisting of cervical cancer, ovarian cancer, breast cancer, leukemia cancer, and lung cancer. 2. The method of claim 1 wherein the single amino acid of the compound is selected from the group consisting of tyrosine, cysteine, serine, and lysine. 3. The method of claim 1 wherein the single amino acid of the compound is tyrosine. 4. The method of claim 1 wherein the single amino acid derivative is selected from the group consisting of a derivative of tyrosine, a derivative of cysteine, a derivative of serine and a derivative of lysine. 5. The method of claim 1 wherein the patient is a human. | Purdue Research Foundation | A61B005/00 A61B008/00 A61B010/00 A61K049/00
US09826886 | B2 | 2017-11-28 | Two-way photodynamic therapy stereo colposcope for diagnosing and treating diseases of the female genital tract | 2011-04-15 | A stereo colposcope having variable linearity filter systems for both the excitation step and the suppression step, and can be used universally with any fluorescent compound or drug, as is the case of photodynamic diagnosis (PDD). The colposcope is a two-way colposcope because the treatment can be administered by an optical system or by a light-producing radio-frequency electrical current with a diathermic effect which facilitates photodynamic treatment. The colposcope produces ozone, which has an antiseptic effect when applied to the genital tract. A monitor provides for three-dimensional viewing through the use of two video cameras with the DLP (Digital Light Processing) and HDTV (High Definition Television) systems with the use of active lenses. | 1. A system for diagnosis and treatment of female genital tract diseases comprising: a stereocolposcope for three-dimensional observation including a head with a lighting system, the lighting system including: a first light source;collecting lenses for concentrating light beam;a variable linear type excitation filter system configured to choose any light frequency range to generate a light beam to send to any compound located at the female genital tract covering a range from 400 nanometers to 1080 nanometers, a light beam band being selectable every 20 nanometers;two variable linear suppressors or barrier filters, one for each observation optical axis for adjusting emission beams issued from a compound to be reacted in a dyed female genital tract with a light spectrum from 400 to 700 nanometers; andtwo diaphragms, one on each optical axis with a manual opening controller to contrast an image and watching a depth of field; anda light source housing with a focusable optical system and directed to an optical fiber including: a second light source,collecting converging lenses; anda variable linearity filter configured to choose any light frequency range to generate a light beam to send to any compound located at the female genital tract covering a range from 400 nanometers to 700 nanometers, the light beam band being selectable every 20 nanometers,wherein the stereocolposcope is configured to perform two way photodynamic therapy by manual selection of one light source that is activated based on a desired treatment. 2. The system according to claim 1, wherein the optical fiber is a medical grade optical fiber or an optical fiber made of ethylene and propylene fluorinate (FEP) resin. 3. The system according to claim 1, further comprising a third light source housing with d'Arsonval system or violet ray, consisting of a Ruhmkorff induction coil, Leiden bottles or capacitors, lighters, a high frequency d'Arsonval coil, Oudin resonator which radio frequency emission end cable thereof is connected to an applicator's electrode, which is Strong Finch vacuum glass tube; and wherein the vacuum glass tube is of vaginal type with emission in different colors spanning a visible light spectrum. 4. The system according to claim 1, wherein the lighting system consists of episcopic illumination, based on Kohler principle. 5. The system according to claim 1, wherein the three-dimensional observation is by direct observation through an eyepiece or by two high resolution video cameras placed on said head, each camera being placed in each visual axis. 6. The system according to claim 1, wherein the three-dimensional observation is by a three-dimensional video vision system including DLP (digital light processing) HDTV (high definition television) with use of three-dimensional active lenses. 7. The system according to claim 1, further comprising a high-amperage power line isolator transformer. 8. A system comprising: a stereocolposcope for three-dimensional observa... | - | A61N005/06 A61B005/00 A61B001/00 A61B001/303 G01N021/64 A61B001/05 A61B001/06 A61B001/07 A61N001/00 A61B001/04
US09820673 | B2 | 2017-11-21 | Method and apparatus for monitoring an organ of a patient | 2014-03-14 | An implanted device for an organ of a patient including a housing. The device includes a detector having electrodes that have a varying distance over time between them which produces a detector signal based on electrical signals derived from the organ. The device includes a signal processor disposed in the housing in communication with the detector which determines admittance from the detector signal based on the varying distance over time between the electrodes. The device includes a drive circuit disposed in the housing to cause the electrodes to generate emitted electrical signals. A method for monitoring a patient's organ. | 1. An implanted device for an organ of a patient comprising: a housing;a detector having electrodes that have a varying distance over time between them which produces a detector signal based on electrical signals derived from the organ;a signal processor disposed in the housing in communication with the detector which determines admittance from the detector signal based on the varying distance over time between the electrodes; anda drive circuit disposed in the housing to cause the electrodes to generate emitted electrical signals, the detector that generates the detector signal based on electrical signals derived from tissue and fluid from the organ; and the signal processor in communication with the detector which subtracts in real time a tissue component from the detector signal and produces a fluid volume signal, the electrodes produce a combined signal that has a tissue component and a fluid component. 2. The implanted device as described in claim 1 wherein the fluid is blood, the tissue is myocardium and the processor determines LV, RV, LA, or RV length, area, and/or volume of the heart in 1, 2 or 3 dimensions. 3. The implanted device as described in claim 2 wherein the myocardial component Rm is determined byRm=-Im⁢{Z_}×(1+(ωɛmσm)2)ωɛmσm,where Z is the complex impedance; ω=2πf where f is the frequency; ∈m is permittivity of muscle; and σm is the conductivity of muscle. 4. The implanted device as described in claim 3 wherein the fluid is blood and blood component Rb is determined byRb=Re⁢{Z_}-Rm1+(ω⁢⁢Rm⁢Cm)2,where Z is the complex impedance; ω=2πf where f is the frequency; Rm is the resistance of muscle, and Cm is the capacitance of muscle. 5. The implanted device as described in claim 4 wherein the signal processor subtracts in real time the myocardial component from the detector signal and produces the left and right ventricle, and left and right atrial blood volume signal. 6. An implanted device for an organ of a patient comprising: a housing;a detector having electrodes that have a varying distance over time between them which produces a detector signal based on electrical signals derived from the organ;a signal processor disposed in the housing in communication with the detector which determines admittance from the detector signal based on the varying distance over time between the electrodes; anda drive circuit disposed in the housing to cause the electrodes to generate emitted electrical signals, the detector that generates the detector signal based on electrical signals derived from tissue and fluid from the organ, the signal processor in communication with the detector which subtracts in real time a tissue component from the detector signal and produces a fluid volume signal, the electrodes produce a combined signal that has a tissue component and a fluid component, the fluid is blood, the tissue is myocardium and the processor determines LV, RV, LA, or RV length, area, and/or volume of the heart in 1, 2 or 3 dimensions, the m... | Board of Regents, The University of Texas System | A61N001/05 A61B005/053 A61B005/107 A61B005/20 A61B005/0452 A61B005/0476 A61B005/0488 A61B005/08 A61B005/029 A61B005/0295 A61B005/00 A61B005/0215 A61N001/365 A61N001/368 A61N001/39
US09827420 | B2 | 2017-11-28 | Transcutaneous electrical nerve stimulator with user gesture detector and electrode-skin contact detector, with transient motion detector for increasing the accuracy of the same | 2014-05-05 | Apparatus for transcutaneous electrical nerve stimulation in a user, the apparatus comprising: a housing; stimulation means carried by the housing for electrically stimulating at least one nerve; a pair of electrodes releasably mounted to the housing and connectable to the stimulation means for electrical stimulation of the at least one nerve; monitoring means for monitoring user gesture, electrode-skin contact integrity and transient motion; analysis means for analyzing the output of the monitoring means for determining user gesture, electrode-skin contact integrity and transient motion; and control means for controlling the output of the stimulation means in response to the determined user gesture, electrode-skin contact integrity and transient motion. | 1. Apparatus for transcutaneous electrical nerve stimulation in a user, the apparatus comprising: a housing;stimulation means for electrically stimulating at least one nerve;a pair of electrodes connectable to the stimulation means for electrical stimulation of the at least one nerve;monitoring means for monitoring at least one selected from the group consisting of user gesture, electrode-skin contact integrity and transient motion of the user;analysis means for analyzing at least one of (i) the user gesture and the transient motion monitored by the monitoring means for determining whether an intentional user gesture has occurred, and (ii) the electrode-skin contact integrity and the transient motion monitored by the monitoring means for determining whether a non-transient degradation of the electrode-skin contact integrity has occurred; andcontrol means for automatically modifying operation of the stimulation means in response to at least one of the determination of an intentional user gesture and a non-transient degradation of the electrode-skin contact integrity. 2. Apparatus according to claim 1 wherein the determined user gesture is used to control output of the stimulation means. 3. Apparatus according to claim 1 wherein said monitoring of the user gesture and transient motion is accomplished using at least one accelerometer. 4. Apparatus according to claim 3 wherein the accelerometer provides an acceleration signal associated with acceleration along at least one axis of three dimensional space. 5. Apparatus according to claim 4 wherein the acceleration signal is high-pass filtered to remove the static earth gravity element in the acceleration signal. 6. Apparatus according to claim 1 wherein monitoring the electrode-skin contact integrity is accomplished with impedance measurement. 7. Apparatus according to claim 6 wherein impedance is determined by dividing the voltage difference across the electrodes by the stimulation current provided by the stimulating means. 8. Apparatus according to claim 4 wherein said analysis means analyzes the acceleration signal to recognize pulse events. 9. Apparatus according to claim 8 wherein the acceleration signal is from one accelerometer axis. 10. Apparatus according to claim 8 wherein the acceleration signal is derived from acceleration data associated with acceleration along all three axes of three dimensional space. 11. Apparatus according to claim 8 wherein a pulse event is characterized by a segment of the acceleration signal exceeding a pulse threshold for a duration no longer than a pulse duration threshold. 12. Apparatus according to claim 11 wherein exceeding a pulse threshold occurs when the acceleration signal value is greater than a positive threshold value. 13. Apparatus according to claim 11 wherein exceeding a pulse threshold occurs when the acceleration signal value is smaller than a negative threshold value. 14. Apparatus according to claim 11 wherein exceeding a pulse thres... | Neurometrix, Inc. | A61N001/00 A61N001/36 A61N001/08 G06F003/01 A61B005/11 A61N001/04
US09808715 | B2 | 2017-11-07 | Interactive gaming system | 2012-05-30 | An interactive gaming system for a user suffering tinnitus. The system comprises a game machine configured to provide a game scenario in which the user's performance in the game scenario is based on the user's response to audio cue arrangements comprising a combination of sounds, and where at least one of the sounds has one or more sound characteristics, that are configured based on one or more sound characteristics of the user's perceived tinnitus. The system further comprises an audio delivery device or devices that are driven by the game machine to present the audio cue arrangements to the user in accordance with the game scenario and a user interface in signal communication with the game machine and which is operable by a user to interact with the game scenario by responding to the audio cue arrangements. | 1. An interactive auditory training rehabilitation gaming system for a user suffering tinnitus comprising: an electronic game machine having a processor configured to provide a game scenario in which the user's performance in the game scenario is based on the user's response to audio cue arrangements comprising a combination of sounds, and where at least one of the sounds has one or more sound characteristics that are configured based on one or more sound characteristics of the user's perceived tinnitus, and wherein the sound characteristics of the user's perceived tinnitus comprises spatial attributes indicative of the spatial location in 3D auditory space of the sound source location of the tinnitus as perceived by the user;an audio delivery device or devices that are driven by the game machine to present the audio cue arrangements to the user in accordance with the game scenario; anda user interface in signal communication with the game machine and which is operable by a user to interact with the game scenario by responding to the audio cue arrangements. 2. An interactive gaming system according to claim 1 wherein the sound characteristics of the user's perceived tinnitus comprise any one or more of the following sound attributes: frequency, patch, bandwidth, loudness, temporal properties, or type of stimuli. 3. An interactive gaming system according to claim 1 wherein at least one audio cue arrangement comprises a sequential playback of sounds in accordance with a playback timeline. 4. An interactive gaming system according to claim 1 wherein the game machine is configured to present background noise or ambient sound constantly or intermittently over the audio delivery device(s) in combination with the audio cue arrangements. 5. An interactive gaming system according to claim 1 further comprising a visual display screen that is driven by the game machine to present visual information relating to the game scenario, and wherein the game machine is configured to display information indicative of the user's performance in the game scenario on the visual display screen. 6. An interactive gaming system according to claim 1 wherein the game machine is configured to customize the audio cue arrangements of the game scenario to the user based on a user tinnitus profile that defines one or more assessed tinnitus sound characteristics of the user. 7. An interactive gaming system according to claim 6 wherein the user tinnitus profile is in the form of a stored electronic data file that is retrieved from an associated data storage medium by the game machine. 8. An interactive gaming system according to claim 6 wherein the game machine is configured to run an initial user calibration function prior to commencing game play of a game scenario to re-assess the user's perceived tinnitus sound characteristic relative to those stored in the user tinnitus profile and the user calibration function generating calibration data indicative of the user's current ... | Auckland Uniservices Ltd. | A63F013/00 A63F013/424 A63F009/00 A63F009/24 A61B005/12 A61B005/00 A61F011/00
US09829480 | B2 | 2017-11-28 | Remote breath alcohol monitor | 2015-03-16 | A portable handheld wireless breath alcohol monitoring device (RBAM) utilizes facial recognition from an enrollment image or gallery of images accumulated over time, and automatic retesting if an initial test is positive for alcohol or of an initial facial match is negative. A location fix is captured with each breath test taken by an offender. After each breath test, the breath alcohol content (BrAC), date and time of the breath test, facial image data, and location fix are uploaded through a built-in cellular phone module in the RBAM to a monitoring station. The monitoring station evaluates each breath test and determines if immediate notification to a supervising agency is needed. If so, an email, text message, or page is sent to the supervising agency. RBAM enables the monitoring of lower-risk offenders or offenders who have earned the privilege of a less intrusive alcohol testing and monitoring program. | 1. A method for remote breath alcohol monitoring comprising the steps of: (a) analyzing by a remote breath alcohol monitor a first breath sample for a first breath alcohol content;(b) capturing a first digital image of at least a portion of a face of a person providing the first breath sample;(c) performing a first facial match on a first template extracted from the first digital image of the person compared to a gallery template derived from a gallery of images of the person; and(d) sending by the remote breath alcohol monitor at least one of the first breath alcohol content and a result of the first facial match to a monitor network. 2. The method for remote breath alcohol monitoring according to claim 1 further comprising the step of: receiving the first breath sample in the remote breath alcohol monitor. 3. The method for remote breath alcohol monitoring according to claim 1 wherein capturing step (b) further comprise the step of: capturing the first digital image with a camera in the remote breath alcohol monitor. 4. The method for remote breath alcohol monitoring according to claim 1 wherein performing step (c) further comprise the step of: performing the first facial match by at least one of the remote breath alcohol monitor and the monitor network. 5. The method for remote breath alcohol monitoring according to claim 1 wherein the first digital image is substantially all of the face of the person providing the first breath sample. 6. The method for remote breath alcohol monitoring according to claim 1 further comprising the steps of: when the first breath alcohol content is greater than or equal to a predefined value, performing the following steps before sending step (d):(c1) generating by the remote breath alcohol monitor an output requesting a retest;(c2) receiving a second breath sample in the remote breath alcohol monitor;(c3) analyzing the second breath sample for a second breath alcohol content;(c4) capturing a second digital image of the face of the person providing the second breath sample;(c5) performing a second facial match on a second template extracted from the second digital image compared to the gallery template derived from the gallery of images of the person; and(c6) sending by the remote breath alcohol monitor at least one of the second breath alcohol content and the result of the second facial match along with the at least one of the first breath alcohol content and the result of the first facial match to the monitor network. 7. The method for remote breath alcohol monitoring according to claim 6 wherein performing steps (c) and (c5) further comprise the steps of: comparing the first template to the gallery template derived from the gallery of images, and determining a first match score;comparing the second template to the gallery template, and determining a second match score;determining a first facial match for the first digital image if the first match score is greater than a threshold value; or, determining a... | ALCOHOL MONITORING SYSTEMS, INC. | G08B023/00 G08B013/14 G08C015/06 G01N033/497 G06K009/00 A61B005/08 G01N033/00
US09782648 | B2 | 2017-10-10 | Athletic training, data collection, dynamic, and personified sporting method, apparatus, system, and computer program product | 2017-02-20 | A computer implemented personified sporting apparatus, system, method and/or computer program product may provide an electronically and programmably controlled personification attributes of a sport, and can include: a sporting device, comprising a processor; a memory coupled to the processor; at least one personified feature, attribute, or movement; and at least one user interface coupled to sporting device, wherein said user interface comprises at least one of: a display device, at least one input device, at least one output device, a keyboard, or a touchscreen, and wherein said at least one sporting device is configured to at least one of: enable at least one user to interact with said at least one sporting device; receive a selection of at least one sporting routine; or receive instructions to control said at least one sporting device. The device can include a robot, display(s), and/or optional speech bubbles. | 1. A personified sporting robotic device apparatus comprising: at least one robot comprising: a plurality of joint axes, wherein each joint axis of said plurality of joint axes comprises: at least one motor;at least one structural member coupled to at least one of said each of said plurality of joint axes motors,wherein each of said at least one motor is configured to move an associated coupled at least one of said at least one structural member;at least one sporting device coupled to said at least one robot, wherein said at least one sporting device comprises at least one personified feature, attribute, or movement; andat least one computer processor, wherein said at least one computer processor is coupled to said plurality of joint axes motors;at least one memory coupled to said at least one computer processor;at least one user interface coupled to said at least one robot, wherein said at least one user interface is configured to interact with a user to receive a selection of at least one sporting routine,wherein said at least one user interface comprises: at least one input device coupled to said at least one computer processor; andat least one output device coupled to said at least one computer processor;wherein the personified sporting robotic device apparatus comprises at least one of: a) wherein said user interface comprises at least one of: a display device,said at least one input device,said at least one output device,a keyboard, ora touchscreen;b) wherein said at least one computer processor is configured to save or retrieve said at least one sporting routine from said at least one memory;c) wherein said at least one computer processor is configured to randomize at least one challenge by said at least one sporting device; ord) wherein said computer processor is configured to at least one of: combine a plurality of previously saved of said at least one sporting routine, orshuffle a plurality of previously saved of said at least one sporting routine; andfurther comprisingat least one of:a) wherein said at least one sporting device comprises being configured to at least one of: pitch or throw;catch or hit;automatically adjust;automatically move;move, adjust or modify position;move in at least two directions;move or adjust with respect to player position;move with at least one of a velocity, or an acceleration;automatically move in at least one of a horizontal plane, or a vertical plane;adjust or move at least one goal;adjust or move at least one target;adjust or move in five axes of freedom relative to a support;adjust or move with at least one fixed dimension;adjust or move within a precise vertical or horizontal dimensional range;provide at least one sporting goal for a game;provide a plurality of sporting goals for a game;provide an adjustable or moving goal;provide at least one goal automatically moveable based on a received input; orprovide at least one goal stationary for at least a predetermined interval;b) wherein said at least one... | Christopher DeCarlo | A63B063/08 A63B069/00 A63B024/00 A63B063/00 A63B071/06 F41J001/00 A63B069/38 A63B069/36 G06N003/08 A61B005/00 F41J001/01 F41J001/10 F41J003/00 F41J005/02 F41J005/052 G06F019/00 G06N005/02 G09B019/00 A63B102/14 A63B069/10 A61B005/11 G06Q050/00
US09795339 | B2 | 2017-10-24 | Sensor for measuring biometric information and item of clothing including the same | 2015-10-09 | A sensor for measuring biometric information includes a biometric electrode sensing biometric information in a state in which the biometric electrode is in contact with a surface of the skin; and a signal line transmitting the sensed biometric information, wherein the biometric electrode is printed on an inner surface of an item of clothing that comes into contact with a human body. | 1. A sensor for measuring biometric information, the sensor comprising: a biometric electrode sensing biometric information in a state in which the biometric electrode is in contact with a surface of the skin; anda signal line transmitting the sensed biometric information,wherein the biometric electrode was formed by printing a mixture solution, obtained by mixing a carbon nano-tube (CNT) dispersion with textile paint, on an inner surface of the item of clothing that comes into contact with a human body, andwherein the CNT dispersion was prepared by mixing 3.0wt. % of CNT and an isopropyl alcohol (IPA) solvent, by being stirred by a stirrer for 0.5 to 4 hours, by being aged for 24 hours, and a supernatant was subsequently removed using a centrifugal separator. 2. The sensor of claim 1, wherein the signal line is formed by printing the mixture solution on an inner surface of the item of clothing. 3. The sensor of claim 2, wherein the textile paint is printed on a surface of the signal line. 4. The sensor of claim 1, further comprising: a snap button provided at an end portion of the signal line; anda signal processing device having a detachable recess formed to be attached to or detached from the snap button, and receiving biometric information transmitted through the signal line. 5. The sensor of claim 4, wherein the signal processing device includes: an amplifier amplifying the received biometric information;a filter filtering the amplified biometric information; anda communications module wirelessly transmitting the filtered biometric information outwardly. 6. An item of clothing comprising: a sensor for measuring biometric information including a biometric electrode sensing biometric information in a state in which the biometric electrode is in contact with a surface of the skin and a signal line transmitting the sensed biometric information,wherein the biometric electrode was formed by printing a mixture solution, obtained by mixing a carbon nano-tube (CNT) dispersion with textile paint, on an inner surface of the item of clothing that comes into contact with a human body, andwherein the CNT dispersion was prepared by mixing 3.0wt. % of CNT and an isopropyl alcohol (IPA) solvent, by being stirred by stirrer for 0.5 to 4 hours, by being aged for 24 hours, and a supernatant was subsequently removed using a centrifudal separator. 7. The item of clothing of claim 6, wherein the signal line is formed by printing the mixture solution on an inner surface of the item of clothing. 8. The item of clothing of claim 7, wherein the textile paint is printed on a surface of the signal line. 9. The item of clothing of claim 6, wherein the sensor further includes: a snap button provided at an end portion of the signal line; anda signal processing device having a detachable recess formed to be attached to or detached from the snap button, and receiving biometric information transmitted through the signal line. 10. The item of clothing of claim 9, wher... | Industry-Academic Cooperation Foundation, Chosun University | H04Q009/00 G08C017/00 G11C008/10 A61B005/00 A61B005/0488 A61B005/024
US09801578 | B2 | 2017-10-31 | Reduced size optical coupler for fluorescence detection | 2014-01-21 | Optical systems are disclosed for use in identifying an analyte, such as glucose in blood or interstitial fluid (ISF), using a biomaterial, such as glucose binding protein (GBP), that is brought into contact with the analyte. An optical system includes a first filter adapted to reflect light emitted from a light-emitting diode to illuminate a fluorescent body, and further adapted to transmit light emitted from the fluorescent body, and a second filter adapted to separate light transmitted by the first filter into signal band light and reference band light. | 1. An optical system comprising: a first filter adapted to reflect light emitted from a light-emitting diode in a first frequency band to illuminate a fluorescent body, and further adapted to transmit light emitted from the fluorescent body in a second frequency band;a second filter adapted to separate light transmitted by the first filter by reflecting a signal band light and transmitting a reference band light; anda single optical connector comprising at least two integral lenses and an inclined surface upon which said first filter is mounted and said second filter is positioned to direct the reflected signal band light through the inclined surface. 2. The optical system of claim 1, further comprising a photodiode adapted to receive at least one selected from the set consisting of signal band light and reference band light transmitted by the second filter. 3. The optical system of claim 1, wherein the first filter comprises a first glass filter comprising a first dichroic filter coating on a first surface of the first glass filter, wherein the first dichroic filter is adapted to reflect light emitted by the light-emitting diode and is further adapted to transmit light emitted from the fluorescent body. 4. The optical system of claim 1, wherein the first filter is further adapted to block wavelengths outside of a signal band wavelength range and a reference band wavelength range. 5. The optical system of claim 1, wherein the first filter comprises a glass filter coated by a dichroic filter coating. 6. The optical system of claim 1, wherein the second filter comprises a second glass filter comprising a second dichroic filter coating on a first surface of the second glass filter, wherein the second dichroic filter is adapted to reflect light of wavelengths associated with a signal band, and is further adapted to transmit light of wavelengths associated with a reference band. 7. The optical system of claim 1, wherein wavelengths associated with the signal band are shorter than wavelengths associated with the reference band. 8. The optical system of claim 1, wherein the second filter comprises a diffraction grating layer adapted to spread light in accordance with wavelength of light. 9. The optical system of claim 1, wherein the first filter and the second filter are formed on opposing sides of a filter element. 10. An optical system comprising: a first filter adapted to reflect light emitted from a light-emitting diode in a first frequency band to illuminate a fluorescent body, and further adapted to transmit light emitted from the fluorescent body in a second frequency band;a second filter adapted to separate light transmitted by the first filter by reflecting a first beam of light in a signal band and transmitting a second beam of light in a reference band; anda single optical connector comprising at least two integral lenses and an inclined surface upon which said first filter is mounted and said second filter is positioned to direct th... | Becton, Dickinson and Company | A61B005/1455 A61B005/154 G01N021/64 A61B005/145 A61B005/00 A61B005/1459
US09778382 | B2 | 2017-10-03 | Motion-corrected pet images | 2013-10-01 | A method is disclosed for generating a motion-corrected PET image of an examination area in a combined MR-PET system. In an embodiment, the method includes recording PET events from the examination area in a first recording time frame; recording a number of MR images of the examination area in at least the first recording time frame; calculating an at least two-dimensional movement information of the examination area on the basis of the number of MR images, wherein the movement information describes the movement information of the examination area during the first recording time frame, and determining the motion-corrected PET image from the PET events using the calculated movement information. | 1. A method for generating a motion-corrected PET image of an examination area in a combined MR-PET system, the method comprising: recording PET events from the examination area in a first recording time frame;recording a number of MR images of the examination area in at least the first recording time frame;calculating an at least two-dimensional movement information of the examination area on the basis of the recorded number of MR images, wherein the at least two-dimensional movement information describes the movement of the examination area during the first recording time frame, the movement of the examination area being a cyclical movement and the number of MR images are recorded in different time segments of the cyclical movement; anddetermining the motion-corrected PET image from the recorded PET events using the calculated at least two-dimensional movement information, the determining including, generating time segment PET images for the different times segments of the cyclical movement, a respective time segment PET image of the time segment PET images is generated on the basis of the PET events occurring in the associated time segment, the generated time segment PET images being associated with movement segment information of the at least two-dimensional movement information, andcorrecting the time segment PET images with the associated movement segment information such that a corrected time segment PET image appears for each time segment of the cyclical movement, in which the movement occurring in the associated time segment is corrected. 2. The method of claim 1, wherein at least two-dimensional movement segment information of the examination area is determined for each of the time segments, which describes the movement of the examination area in the associated time segment. 3. The method of claim 1, wherein the determining includes, summing the corrected time segment PET images to generate the motion-corrected PET image. 4. The method of claim 3, wherein the summing is based on sizes of the movement occurring in the time segments, and is the time segment PET images are weighted relatively lower in the summation if the movement was relatively larger in the associated time segment than another time segment PET image, in which the movement was relatively lower in the associated time segment. 5. The method of claim 1, further comprising: determining an associated trajectory followed by an annihilation radiation is for each PET event, and the determining the motion-corrected PET image includes, calculating corrected trajectories for the PET events with the aid of the two-dimensional movement information, anddetermining the motion-corrected PET image from the corrected trajectories. 6. The method of claim 5, wherein the determining the associated trajectory determines deformed trajectories for the PET events on the basis of the at least two-dimensional movement information of the examination area, andthe determining the motion-corrected... | SIEMENS AKTIENGESELLSCHAFT | A61B005/05 G01T001/29 A61B005/00 A61B005/055 A61B005/113 A61B006/03 A61B006/00 G01R033/48
US09788780 | B2 | 2017-10-17 | Allergy skin test applicator, and related testing tray, testing kit and testing method | 2014-07-11 | An allergy testing kit contains a plurality of allergy testing applicators, an allergy testing tray and a plurality of allergen bottles each containing an allergen. Each of the applicators contains an elongated handle, a plurality of arms extending from the elongated handle and disposed in an asymmetrical configuration, and a plurality of legs with tines extending from each of the arms. The allergy testing tray contains a main body having an underside and a top surface, a cover for locking with the main body and a plurality of reservoirs extending from the underside of the main body. The reservoirs each have a chamber with an opening extending from the top surface. The reservoirs are disposed in different groups and each group has an asymmetrical configuration matching that of the applicator. | 1. An allergy skin testing applicator, comprising: an elongated handle;a plurality of arms extending from said elongated handle, said plurality of arms having an asymmetrical configuration wherein said arms are disposed on opposite sides of said elongated handle except for an odd numbered arm being disposed on a center line and offset from said arms situated on said opposite sides of said elongated handle; anda plurality of legs each having tines and a respective one of said legs extending from a respective one of said arms, said respective leg extending from said odd numbered arm having a top and said elongated handle extending a distance into said top of said odd numbered arm, said legs having a main body region and an upper tapered region extending from said main body region and being wider than said main body region, all of said upper tapered region being distal from said tines. 2. The allergy testing applicator according to claim 1, wherein a number of said arms or said legs is an odd number selected from the group consisting of 3, 5, 7, 9 and 11. 3. The allergy testing applicator according to claim 1, wherein said elongated handle has a first end and a second end defined along a longitudinal extent of said elongated handle, and one of said arms extends directly from a distal end of one of said first or second ends. 4. The allergy testing applicator according to claim 1, wherein: said elongated handle has long sides and shorter sides both extending in a vertical direction;a shoulder extending perpendicularly from both of said long sides of said elongated handle in a horizontal direction and away from said elongated handle, said shoulder extending over a majority of each of said long sides. 5. The allergy testing applicator according to claim 1, wherein said elongated handle has a bottom side and a shoulder extends out from said bottom side of said elongated handle. 6. The allergy testing applicator according to claim 1, wherein said upper tapered region is conical or parabolic in shape. 7. The allergy testing applicator according to claim 1, further comprising raised indicators, a respective one of said raised indicators is disposed on a top side of a respective one of said arms. 8. The allergy testing applicator according to claim 1, wherein said elongated handle has a plurality of ribs formed at a given angle to said elongated handle for assisting a hand of a user to grip said elongated handle. 9. The allergy testing applicator according to claim 1, further comprising reinforcements extending along a top of said arms. | - | A61B005/00 A61B050/22 A61B050/20 A61B050/30 A61B050/33 A61B090/94 A61B010/00 A61B050/00
US09823138 | B2 | 2017-11-21 | Non-invasive automatic monitoring of pet animal's core temperature | 2013-08-21 | A system, device and method monitoring whether a core temperature of a warm-blooded pet animal is within a normal range for the pet animal comprises a sensing assembly including (i) a skin temperature sensor positioned such that a sensing surface of the skin temperature sensor faces the animal, the skin temperature sensor configured to produce a skin temperature output, (ii) an ambient temperature sensor spaced away from the animal and configured to produce an ambient temperature output, and (iii) an accelerometer for sensing an acceleration of the pet animal and producing an acceleration output; and a processor for receiving the outputs, calculating an activity level from the acceleration data and determining whether the core temperature of the pet animal is within the normal range based on a pre-defined function relating the skin temperature TS, the ambient temperature TA, and the activity level of the pet animal. | 1. A system for monitoring a health of a warm-blooded pet animal including non-invasively monitoring and providing early detection of whether a core temperature of the warm-blooded pet animal is within a normal range for the pet animal, wherein the pet animal is a dog or cat, the system comprising: a non-invasive collar, including an elastic band, for positioning on a neck of the pet animal, the collar having an outer surface and an inner surface, the collar configured to output signals for ambient temperature, skin temperature, at least one of heart rate and respiration and at least one of posture and movement of the pet animal, while being positioned on a single integrally connected body part of the pet animal,the collar including a sensing assembly having sensors positioned at different points along the collar including (i) a skin temperature sensor positioned such that a sensing surface of the skin temperature sensor faces a body of the pet animal, the skin temperature sensor configured to produce a skin temperature output, (ii) an ambient temperature sensor positioned on the outer surface and configured to produce an ambient temperature output, (iii) an accelerometer for sensing an acceleration level of the pet animal and producing an acceleration level output including posture and movement of the pet animal and (iv) an acoustic sensor configured to measure the at least one of heart rate and respiration and output at least one signal representing vital sign data for the at least one of heart rate and respiration, the acoustic sensor having multiple elastic projections of a substantially same length that face the neck and that are configured in shape to penetrate a fur of the dog or cat and to absorb noise from friction due to movement of a head of the dog or cat, the absorption of the noise increasing a signal to noise ratio of the at least one signal conducted to the acoustic sensor, the length running radially from the collar to the neck; andone or more processors configured to receive the skin temperature output, TS, the ambient temperature output TA, and the acceleration level output and to determine an activity level, A, of the pet animal from the acceleration level output,the one or more processors also configured to(i) use TA to select a zone from one of three zones including (i) a thermoneutral zone, (ii) a zone below lower critical temperature (Tlc) and (iii) a zone above upper critical temperature (Tuc);(ii) use a pre-defined function TMODEL=a*TA+b to calculate from TA a baseline model skin temperature, wherein TA is a variable of the ambient temperature output and a and b are coefficients that are set depending upon an animal size and kind and the selected zone of TA, wherein, in each of the three zones, the baseline model skin temperature varies across varying values of TA;(iii) use c1*A+d1 to calculate from activity A and coefficients (ci, di) an adjustment of the baseline model skin temperature yielding a lower margin of a tole... | PETPACE LTD | G01K013/00 A61B005/01 A61B005/11 A01K029/00 G01K007/42
US09778131 | B2 | 2017-10-03 | Pressure data acquisition assembly | 2014-05-21 | A method and assembly for acquiring pressure data. A pressure sensor is applied to a target surface on an individual. A calibrator and a processing element are in communication with the pressure sensor. Processing element receives pressure data and provides an integrated pressure value over a measurement time period. The integrated pressure value is compared to an alert value and to a change condition value. Where an alert value is exceeded, an alert is transmitted to an output device for display. Where a change condition value is exceeded, a measurement parameter of the pressure sensor is changed, or the calibrator is applied to the pressure sensor to recalibrate the pressure sensor to a recalibrated pressure range. | 1. A pressure data acquisition assembly comprising: a support web for applying to a target surface of a body of an individual;a first pressure sensor connected to the support web for sensing at a first frequency a first pressure applied to the target surface;a calibrator in operative communication with the first pressure sensor for calibrating the first pressure sensor to a first pressure range within which the first pressure sensor senses the first pressure;a processing element in operative communication with the first pressure sensor for receiving first pressure signals corresponding to values of the first pressure from the first pressure sensor at the first frequency, and with the calibrator for recalibrating the first pressure sensor, the processing element comprising: a computer readable memory for storing measurement parameters comprising a first threshold value corresponding to a first threshold pressure, a first measurement time period, a first alert value, and a first change condition; anda processor in operative communication with the computer readable memory for accessing the measurement parameters, comparing the first pressure to the first threshold pressure to provide a first comparative pressure, integrating the first comparative pressure over the first time measurement period to provide a first integrated pressure, comparing the first integrated pressure with the first alert value to determine a first alert status, comparing the first integrated pressure with the first change condition to determine a first change status, and changing at least one measurement parameter or recalibrating the first pressure sensor to a first recalibrated pressure range, according to the first change status; andan output device in communication with the processing element for receiving signals comprising data from the processing element and displaying the data. 2. The assembly of claim 1 further comprising a second pressure sensor for sensing at a second frequency a second pressure applied to the target surface; wherein the calibrator is further in operative communication with the second pressure sensor for calibrating the second pressure sensor;the processing element is in operative communication with the second pressure sensor for receiving second pressure signals corresponding to values of the second pressure from the second pressure sensor at the second frequency, and with the calibrator for recalibrating the second pressure sensor;the measurement parameters further comprise a second threshold value corresponding to a second threshold pressure, a second measurement time period, a second alert value, and a second change condition;the processing element is further for comparing the second normalized pressure to a second threshold pressure value to determine a second comparative pressure;the processor is in operative communication with the computer readable memory for accessing the measurement parameters, comparing the second pressure to the second th... | ORPYX MEDICAL TECHNOLOGIES INC. | G01L025/00 A61B005/00 G01D003/08 A61B005/103 A61G007/057 G01L019/12 G01L027/00
US09808140 | B2 | 2017-11-07 | Steerable segmented endoscope and method of insertion | 2014-09-15 | A system for determining the shape of a surgical instrument includes an elongate instrument body having a selectively steerable distal tip, a pair of automatically controlled segments proximal to the selectively steerable distal tip, and a joint that couples the pair of adjacent automatically controlled segments together. An actuator changes an angle of the joint and a position encoder provides information associated with the joint angle. A controller receives the information associated with the angle of the joint and generates a three dimensional model of a shape of the instrument. A method for determining the three dimensional shape of an instrument includes providing axial position data to a controller, providing angular position data to the controller from a respective segment controller, and generating a three-dimensional model of a shape of said instrument using the axial position data and the angular position data. | 1. A method for determining the three dimensional shape of an instrument during or after insertion of said instrument into a body cavity, wherein said instrument comprises an elongate body comprising a selectively steerable distal tip, and a plurality of pivotably coupled and automatically controlled segments proximal to said selectively steerable distal tip, wherein adjacent segments of the plurality of segments are pivotable about two intersecting, independent axes, said method comprising: advancing said instrument into said body cavity;providing axial position data to a controller, wherein said axial position data is relative to a datum;providing angular position data to said controller from a respective segment controller associated with each of said plurality of automatically controlled segments, the angular position data comprising relative angular positions of the adjacent segments of the plurality of segments; andgenerating in electronic memory of said controller a three-dimensional model of a shape of said instrument using the provided axial position data and the angular position data. 2. The method for determining the three dimensional shape of an instrument according to claim 1, wherein the providing angular position data further comprises: providing angular position data to said controller from each segment controller using an optical fiber communications channel between said controller and each segment controller. 3. The method for determining the three dimensional shape of an instrument according to claim 1, wherein the providing angular position data further comprises: providing angular position data to said controller from each segment controller using an electrically conductive communication channel between said controller and each segment controller. 4. The method for determining the three dimensional shape of an instrument according to claim 1, wherein the providing angular position data further comprises: wirelessly providing angular position data to said controller from each segment. 5. The method of claim 1, further comprising the controller commanding each segment controller to control the geometry of the segment with which each segment controller is associated during the advancing. 6. The method of claim 1, further comprising the controller controlling an orientation of the two independent axes between adjacent segments in the plurality of automatically controlled segments during the advancing. 7. The method of claim 1, further comprising displaying the three-dimensional model on a display. 8. The method of claim 1, further comprising combining the three-dimensional model with a three dimensional image of said body cavity to form a combined three dimensional image. 9. The method of claim 8, further comprising displaying corresponding movement of the instrument within the body cavity based on the combined three dimensional image. 10. The method of claim 1, further comprising the controller commanding each segment c... | INTUITIVE SURGICAL OPERATIONS, INC. | A61B001/005 A61B005/06 A61B001/00 A61B001/008 A61B001/31 A61B001/015 A61B001/018 A61B001/05 A61B034/00 A61B034/30
US09788789 | B2 | 2017-10-17 | Systems, articles, and methods for stretchable printed circuit boards | 2014-08-28 | Improved stretchable printed circuit boards, and fabrication methods thereof, are described. The improved stretchable printed circuit boards include a serpentine conductive trace enclosed by stretchable dielectric material. The stretchable dielectric material has a serpentine shape itself, realized by crenulated edges. The crenulated edges reduce torsional strain on the conductive trace and are formed, for example, by cutting away sections of the stretchable dielectric material proximate segments of the serpentine conductive trace where the serpentine conductive trace changes direction. | 1. A method of fabricating a stretchable printed circuit board, the method comprising: fabricating a flexible printed circuit board, wherein fabricating a flexible printed circuit board comprises: depositing a layer of flexible dielectric material;depositing a layer of conductive metal on top of the flexible dielectric material; andetching a circuit pattern into the layer of conductive metal, wherein the circuit pattern includes at least one serpentine conductive trace that includes a plurality of changes in direction,removing at least some portions of the flexible dielectric material that are not covered by the at least one serpentine conductive trace; andenclosing at least a portion of the flexible printed circuit board in a stretchable dielectric material, wherein the stretchable dielectric material forms a crenulated shape that includes a plurality of crenulations across a length thereof, wherein each crenulation is positioned proximate and corresponds to a respective change in direction in the at least one serpentine conductive trace, and wherein enclosing at least a portion of the flexible printed circuit board in a stretchable dielectric material comprises:placing the flexible printed circuit board into a mold that provides a crenulated shape for the stretchable dielectric material;injecting the stretchable dielectric material, in liquid form, into the mold;solidifying the stretchable dielectric material; andremoving the mold. 2. The method of claim 1 wherein removing at least some portions of the flexible dielectric material that are not covered by the at least one serpentine conductive trace includes cutting away the at least some portions of the flexible dielectric material that are not covered by the at least one serpentine conductive trace. | THALMIC LABS INC. | H05K001/02 H05K001/14 A61B005/00 A61B005/0492
US09788808 | B2 | 2017-10-17 | Method of displaying virtual ruler on separate image or medical image of object, medical image obtaining apparatus, and method and apparatus for displaying separate image or medical image with virtual ruler | 2015-02-02 | A medical image displaying apparatus includes: an image obtainer that obtains a medical image of an object without placing a physical ruler beside the object; a controller that generates a virtual ruler based on a location of the image obtainer when the medical image is obtained; and a display that displays the virtual ruler on the medical image. | 1. A medical imaging apparatus comprising: an X-ray emitter configured to emit first X-rays and second X-rays toward an object;an X-ray detector configured to detect the first X-rays and the second X-rays that have been emitted to obtain a first X-ray image and a second X-ray image, corresponding to a plurality of areas of the object;a processor, which is configured to obtain information about a first detector location for obtaining the first X-ray image and information about a second detector location, wherein the X-ray detector moves from the first detector location, for obtaining the second X-ray image, is configured to generate a first virtual ruler based on a first distance from a reference point to the first detector location and a second virtual ruler based on a second distance from the reference point to the second detector location, and is configured to obtain a composite image by combining the first X-ray image and the second X-ray image based on the first virtual ruler and the second virtual ruler; anda display configured to display the composite image on which the first virtual ruler and the second virtual ruler are displayed,wherein, when a user input for moving the first X-ray image of the composite image along a first direction is received, the processor moves the first X-ray image along the first direction without moving the first virtual ruler, andthe display is configured to display the first virtual ruler and the second virtual ruler on the composite image by aligning the first virtual ruler and the second virtual ruler as a one-dimensional ruler in correspondence to the first image and the second image. 2. The medical imaging apparatus of claim 1, wherein the X-ray detector is configured to detect the first X-rays and the second X-rays that have been emitted to obtain the first X-ray image and the second X-ray image without placing a physical ruler beside the object between the X-ray emitter and the X-ray detector. 3. The medical imaging apparatus of claim 1, wherein the processor obtains the first distance from the reference point to the X-ray detector when the first X-ray image is obtained, and the second distance from the reference point to the X-ray detector when the second X-ray image is obtained, and the first virtual ruler is associated with the first distance and the second virtual ruler is associated with the second distance. 4. The medical imaging apparatus of claim 1, wherein the processor generates the first virtual ruler associated with the first distance corresponding to the first detector location when the first X-ray image is obtained and the second virtual ruler associated with the second distance corresponding to the second detector location when the second X-ray image is obtained. 5. The medical imaging apparatus of claim 4, wherein the processor obtains a first value for a first side of the first X-ray image and a second value for a second side of the first X-ray image, by using the first distance, gene... | SAMSUNG ELECTRONICS CO., LTD. | G06K009/00 A61B006/00 G06T007/00 A61B005/00 A61B005/107 G06T011/20 G06T011/60
US09782080 | B2 | 2017-10-10 | Object information acquiring apparatus and control method for the object information acquiring apparatus | 2013-11-07 | An object information acquiring apparatus comprises an irradiating unit which irradiates light on an object; a holding member which includes a light-absorbing member whose optical characteristics with respect to irradiated light are known and which holds the object; a detector which detects acoustic waves generated by tissue in the object and the light-absorbing member due to light irradiation; and a computing unit which calculates optical characteristics of the tissue in the object and the light-absorbing member by using the detected acoustic waves and which corrects the optical characteristics of the tissue in the object by using the calculated optical characteristics of the light-absorbing member and the known optical characteristics. | 1. An apparatus comprising: an irradiating unit configured to irradiate an object with light;a holding member configured to hold the object, wherein the holding member includes at least two light-absorbing members and optical characteristics of the at least two light-absorbing members are known;a detector configured to detect acoustic waves generated from in the object and the at least two light-absorbing members due to light irradiation and output a detection signal; anda computing unit configured to: select at least one pair of light absorbing members from among the at least two light-absorbing members,calculate optical characteristics of the tissue in the object and the selected at least one pair of light-absorbing members by using the detection signal,calculate an error between the calculated optical characteristics and the known optical characteristics with respect to the selected at least one pair of light-absorbing members, andcorrect the calculated optical characteristics of a region of the object that is positioned between the selected at least one pair of light-absorbing members on a basis of the error. 2. The apparatus according to claim 1, wherein said computing unit is configured to correct the calculated optical characteristics by any one of methods of linear interpolation, polynomial interpolation, spline interpolation, and a least-squares method on the basis of the error. 3. The apparatus according to claim 1, wherein the optical characteristics of the at least two light-absorbing members and the object are absorption coefficients. 4. The apparatus according to claim 1, wherein the optical characteristics of the at least two light-absorbing members and the object are oxygen saturation,said irradiating unit is configured to irradiate the object and the at least two light-absorbing members with light of each of a plurality of wavelengths, andsaid computing unit is configured to calculate the oxygen saturation based on an absorption coefficient calculated for each of the plurality of wavelengths. 5. The apparatus according to claim 1, wherein said holding member and the at least two light-absorbing members are made of a polyurethane gel. 6. The apparatus according to claim 1, wherein the at least two light-absorbing members are spherical members arranged in a lattice pattern on a contact surface in contact with the object in said holding member. 7. The apparatus according to claim 6, wherein the at least two light-absorbing members are arranged at a plurality of depths from the contact surface in contact with the object in said holding member. 8. The apparatus according to claim 1, wherein the at least two light-absorbing members are cylindrical members arranged in a lattice pattern on a contact surface in contact with the object in said holding member. 9. The apparatus according to claim 1, wherein the region is a straight line connecting the selected at least one pair of light-absorbing members. 10. The apparatus according... | CANON KABUSHIKI KAISHA | A61B005/00 A61B005/145 A61B005/1455
US09789267 | B2 | 2017-10-17 | Vein scanner with user interface | 2015-05-21 | A portable vein viewer apparatus may be battery powered and hand-held to reveal patient vasculature information to aid in venipuncture processes. The apparatus comprises a first laser diode emitting infrared light, and a second laser diode emitting only visible wavelengths, wherein vasculature absorbs a portion of the infrared light causing reflection of a contrasted infrared image. A pair of silicon PIN photodiodes, responsive to the contrasted infrared image, causes transmission of a corresponding signal. The signal is processed through circuitry to amplify, sum, and filter the outputted signals, and with the use of an image processing algorithm, the contrasted image is projected onto the patient's skin surface using the second laser diode. Revealed information may comprise vein location, depth, diameter, and degree of certainty of vein locations. Projection of vein images may be a positive or a negative image. Venipuncture needles may be coated to provide visibility in projected images. | 1. A vein imaging apparatus, for use in imaging veins of a patient and for locating the veins for a medical practitioner by projecting the image onto the patient's skin surface, said vein imaging apparatus comprising: a first laser configured to emit at least a first wavelength of light;one or more photodiodes, said one or more photodiodes each configured to be responsive to said at least a first wavelength of light, to thereby receive a contrasted vein image formed by reflected light at said first wavelength, and to convert said image into a signal;a second laser configured to emit at least a second wavelength of light, being different than said at least a first wavelength of light;a user interface board comprising: electronic circuitry; a display screen; one or more push buttons; a memory; and firmware, said firmware configured to display a scrollable menu screen list of two or more configuration settings, and the parameters for each of said two or more configuration settings, on said display screen, said firmware further configured to permit actuation of said one or more buttons to select and to modify any of said two or more configuration settings to be the default setting for said apparatus, and to permit said selected default settings to be stored in said memory, said two or more configuration settings comprising a plurality of vein visualization settings; andwherein said electronic circuitry is configured to receive said signal from said one or more photodiodes and to output said signal to said second laser, said electronic circuitry configured to cause said second laser to use said signal to emit said at least a second wavelength of light to project said vein image in accordance with said selected vein visualization setting, being optimized for the skin surface. 2. The vein imaging apparatus according to claim 1, wherein said electronic circuitry comprises a microcontroller. 3. The vein imaging apparatus according to claim 2, further comprising a scanner configured to scan a beam of said first wavelength of light emitted from said first laser and a beam of said second wavelength of light emitted from said second laser in a pattern, and onto the skin surface. 4. The vein imaging apparatus according to claim 3, further comprising a dichroic element configured to align said beams of light from said first and second lasers to be on the same optical axis, to be seamed simultaneously by said scanner. 5. The vein imaging apparatus according to claim 4, said two or more configuration settings comprising setting a vein display time out interval. 6. The vein imaging apparatus according to claim 5, said two or more configuration settings comprising setting a backlight intensity of said display screen. 7. The vein imaging apparatus according to claim 6, said two or more configuration settings comprising setting a name for said apparatus. 8. The vein imaging apparatus according to claim 7, further comprising a speaker, said speaker configured t... | AccuVein, Inc. | A61B006/00 A61M005/42 A61B005/00 H02J007/00
US09798385 | B1 | 2017-10-24 | User physical attribute based device and content management system | 2016-05-31 | Systems and methods for device and content management include determining a user is viewing virtual reality content from a user device, and receiving, through a network from the user device, body information associated with the user while the user is viewing the virtual reality content. A user device management configuration associated with the body information is determined. A user device management action is retrieved using the user device management configuration. A notification associated with the user device management action that causes the user device to perform the user device management action and modify the virtual reality content being viewed by the user is sent through the network to the user device. | 1. A system, comprising: a non-transitory memory; andone or more hardware processors coupled to the non-transitory memory and configured to read instructions from the non-transitory memory to cause the system to perform operations comprising: determining a user is viewing virtual reality content from a user device;receiving, through a network from the user device, body information associated with the user while the user is viewing the virtual reality content;determining a user device management configuration associated with the body information;retrieving a user device management action using the user device management configuration; andsending, through the network to the user device, a notification associated with the user device management action that causes the user device to perform the user device management action and modify the virtual reality content being viewed by the user. 2. The system of claim 1, wherein the body information includes facial information associated with eyes of the user, wherein the facial information indicates eye strain. 3. The system of claim 2, wherein the facial information includes one selected from the group consisting of eye blinking rate information, eye redness information, and eye dryness information. 4. The system of claim 3, wherein the eye dryness information is using thermal images of tear films of the user, wherein the thermal images are generated by an infrared sensor of the user device. 5. The system of claim 1, wherein the body information includes gesture information associated with a gesture, wherein the gesture indicates head pain. 6. The system of claim 1, wherein the user device management action is an action to display a break reminder on the user device to remind the user to take a break from using the user device. 7. The system of claim 1, wherein the user device management action is selected from the group consisting of switching off a display of the user device, changing brightness with one or more images of the virtual reality content, changing sizes of texts and images of the virtual reality content, and changing a speed of image transitions of the virtual reality content. 8. A method, comprising: receiving, by a system provider device through a network from a user device associated with a user, body information associated with the user;determining, by the system provider device, a user device management configuration associated with the body information;retrieving, by the system provider device, a user device management action using the user device management configuration; andsending, by the system provider device through the network to the user device, a notification associated with the user device management action that causes the user device to perform the user device management action. 9. The method of claim 8, further comprising: receiving, through the network from a medical history provider device associated with a medical history provider, medical history information asso... | PAYPAL, INC. | G09G005/00 G06F003/01 H04L029/08 G06T019/00 G06K009/00 G02B027/01 A61B005/11 A61B003/10 A61B005/00 A61B003/14 A61B003/113
US09832370 | B2 | 2017-11-28 | Cognitive sensor and method of operating of the same | 2015-07-17 | A cognitive sensor and a method of operating the same. The cognitive sensor includes a sensor unit, which generates electric signals in response to outside stimulations; a signal processing unit, which generates sensed data regarding the outside stimulations by processing electric signals generated by the sensor unit; a cognitive circuit unit, which specifies an area of interest in the sensed data processed by the signal processing unit; and an output unit, which outputs the sensed data generated by the signal processing unit, wherein at least a portion of the sensed data output by the output unit is sensed data regarding the area of interest. | 1. A recognition apparatus comprising: a first cognitive sensor comprising:a sensor unit configured to generate electric signals in response to outside stimulations;a signal processing unit configured to generate sensed data regarding the outside stimulations by processing electric signals generated by the sensor unit;a cognitive circuit unit of the first cognitive sensor configured to continuously specify an area of interest in the sensed data processed by the signal processing unit when a target is signaled in the sensed data and periodically specify the area of interest when the target is not signaled in the sensed area; andan output unit configured to output the sensed data generated by the signal processing unit,wherein at least a portion of the sensed data output by the output unit corresponds to the area of interest; anda second sensor configured to obtain initial information for specifying an area of interest,wherein the cognitive circuit unit of the first cognitive sensor is configured to specify the area of interest based on information obtained via the second sensor. 2. The recognition apparatus of claim 1, wherein the sensor unit, the signal processing unit, the cognitive circuit unit, and the output unit are integrated on a single chip. 3. The recognition apparatus of claim 1, wherein the signal processing unit is configured to replace data corresponding to areas other than the area of interest in the sensed data with existing data, null data, dummy data, or low-resolution data. 4. The recognition apparatus of claim 1, wherein the sensor unit comprises a sensing element array including a plurality of sensing elements, each of which being configured to convert outside stimulations into electric signals. 5. The recognition apparatus of claim 4, wherein the signal processing unit is configured to generate sensed data regarding outside stimulations corresponding to the area of interest by processing electric signals obtained from the sensing elements corresponding to the area of interest based on the area of interest specified by the cognitive circuit unit. 6. The recognition apparatus of claim 4, further comprising a driving circuit unit configured to drive the sensing element array, wherein the driving circuit unit is configured to selectively drive sensing elements corresponding to the area of interest based on information regarding the area of interest received from the cognitive circuit unit. 7. The recognition apparatus of claim 6, wherein, while sensed data regarding outside stimulations corresponding to the area of interest is being generated, the driving circuit unit is configured to drive only sensing elements corresponding to the area of interest and to not drive the other sensing elements. 8. The recognition apparatus of claim 4, wherein one or more of the plurality of sensing elements comprises a photoelectric conversion element configured to convert a light into electric signals. 9. The recognition apparatus of clai... | Samsung Electronics Co., Ltd. | H04N005/228 H04N005/232 A61B005/00 H04N005/345 A61B003/113 A61B008/08 A61B005/11
US09782112 | B2 | 2017-10-10 | Connectors for making connections between analyte sensors and other devices | 2016-02-18 | Analyte sensor connectors that connect analyte sensors, e.g., conductive members of analyte sensors, to other devices such as sensor electronics units, e.g., sensor control units, are provided. Also provided are systems that include analyte sensors, analyte sensor connectors, and analyte sensor electronics units, as well as methods of establishing and maintaining connections between analyte sensors and analyte sensor electronics units, and methods of analyte monitoring/detection. Also provided are methods of making analyte sensor connectors and systems that include analyte sensor connectors. | 1. An electrode, arrangement, comprising: a base substrate;a first conductive layer having a width (W1) and positioned on the base substrate;a first dielectric layer positioned over at least a portion of the first conductive layer;a second conductive layer having a with (W2) and positioned on the first dielectric layer;a conductive strip positioned in contact with the second conductive layer; andat least two fiducial; marks on the base substrate adjacent to the first conductive layer;wherein the fiducial marks are positioned to indicate the width and location of the first conductive layer on the base substrate;wherein the conductive strip comprises a material that is different from the second conductive layer. 2. The electrode arrangement of claim 1, wherein the fiducial marks extend beyond the first dielectric layer on the base substrate. 3. The electrode arrangement of claim 1, wherein W1 is less than W2 or W2 is less than W1. 4. The electrode arrangement of claim 3, wherein when W1 is less than W2 the first conductive layer is spaced away from the first and second side edges of the dielectric has substrate. 5. The electrode arrangement of claim 4, wherein the second conductive layer terminates at a distal end which is spaced hack from the distal end opt first conductive layer. 6. The electrode arrangement of claim 3, wherein when W2 is less than W1 the second conductive layer is spaced away from the first and second side edges of the first dielectric layer. 7. The electrode arrangement of claim 6, wherein the second conductive layer terminates at a distal end which is spaced back from the distal end of the first conductive layer. 8. The electrode arrangement of claim 1, wherein the first and second conductive layers comprise carbon or gold. 9. The electrode arrangement of claim 1, wherein the first conductive layer has a length (L1) and the second conductive layer has a length (L2). 10. The electrode arrangement of claim 9, wherein the conductive strip has a length that is less than L1 and L2. 11. The electrode arrangement of claim 1, wherein the conductive strip is positioned orthogonal to the second conductive layer. 12. The electrode arrangement of claim 1, wherein the second conductive comprises carbon and the conductive strip comprises Ag/AgCl. 13. The electrode arrangement of claim 1, wherein the first conductive layer comprises one or more sensing elements, the sensing elements comprising an analyte responsive enzyme and a mediator. 14. The electrode arrangement of claim 1, wherein first dielectric layer has a thickness which is less than that of the base substrate. 15. The electrode arrangement of claim 14, further comprising a second dielectric layer positioned to cover at least a portion of the second conductive layer, wherein the second dielectric layer has a second thickness which is less than that of the dielectric base substrate. 16. The electrode arrangement of claim 15, wherein W1 is less than the width of the di... | Abbott Diabetes Care Inc. | H05K005/00 A61B005/1486 A61B005/00 A61B005/145 A61B005/1468 A61B005/1473 H05K003/32 H01L023/00 H05K003/36
US09820819 | B2 | 2017-11-21 | Suspension system for remote catheter guidance | 2015-01-02 | A robotic catheter system includes an arch-shaped suspension system comprising a first and second vertical span, a horizontal span between the pair of vertical spans, and at least one robotic catheter head coupled to the horizontal span. The arch-shaped suspension system can include a pair of linear guide blocks that can be attached to a patient bed and in some embodiments the linear guide blocks can be configured to slidably move along a rail. The arch-shaped suspension system can be moved to allow for proper placement of the arch-shaped suspension system for use on a patient or for placing the system away from a patient or for storage. In some embodiments the robotic catheter system can include a suspension system cart for storing and moving the arch-shaped suspension system. | 1. An arch-shaped suspension system comprising: a first vertical span, a second vertical span, a horizontal span, and a robotic catheter head;wherein the horizontal span comprises a first end and a second end,wherein a first end of the horizontal span is coupled to the first vertical span and a second end of the horizontal span is coupled to the second vertical span, wherein the robotic catheter head is operably connected to the horizontal span, wherein the horizontal span is configured to rotate around a longitudinal axis of the horizontal span such that the robotic catheter head can be placed in a vertical position, and wherein the first vertical span and the second vertical span are configured to releaseably couple to a patient bed. 2. The arch-shaped suspension system according to claim 1 further comprising a first linear guide block coupled to the first vertical span and a second linear guide block coupled to the second vertical span. 3. The arch-shaped suspension system according to claim 2 wherein the first and second linear guide blocks further comprise a guide housing and at least one roller. 4. The arch-shaped suspension system according to claim 3 wherein the guide housing is configured to couple to a linear guide track and the at least one roller is configured for moving the arch-shaped suspension system along the guide track. 5. The arch-shaped suspension system according to claim 1 wherein the robotic catheter head further comprises a catheter cartridge and a sheath cartridge. 6. The arch-shaped suspension system according to claim 1 wherein the robotic catheter head is a first robotic catheter head and the arch-shaped suspension system further comprises a second robotic catheter head. 7. The arch-shaped suspension system according to claim 1 further comprising a robotic mount coupled to the horizontal span and coupled to the robotic catheter head. 8. The arch-shaped suspension system according to claim 7 wherein the robotic mount is configured to be slidably movable along the horizontal span and the robotic catheter head is configured to be slidably movable along the robotic mount. 9. The arch-shaped suspension system according to claim 7 wherein the robotic mount further comprises a coupler extension that is coupled to the robotic mount and to the robotic catheter head. 10. The arch-shaped suspension system according to claim 9 wherein the coupler extension is configured to adjust at least one of the extension, pan, rotation, pitch, and yaw of the robotic catheter head. 11. The arch-shaped suspension system according to claim 1 wherein one of the first vertical span and the second vertical span further comprises an arch clamp configured for securing the arch-shaped suspension system to a patient table. 12. A robotic catheter system comprising: an arch-shaped suspension system comprising a first vertical span, a second vertical span, a horizontal span, and a robotic catheter head; anda suspension system cart comprising a... | St. Jude Medical, Cardiology Division, Inc. | A61B019/00 A61B005/00 A61M025/01 A61B018/14 A61B090/57 A61B090/50 A61B034/37 A61B018/00 A61B034/30
US09805463 | B2 | 2017-10-31 | Systems and methods for predicting location, onset, and/or change of coronary lesions | 2013-08-27 | Systems and methods are disclosed for predicting the location, onset, or change of coronary lesions from factors like vessel geometry, physiology, and hemodynamics. One method includes: acquiring, for each of a plurality of individuals, a geometric model, blood flow characteristics, and plaque information for part of the individual's vascular system; training a machine learning algorithm based on the geometric models and blood flow characteristics for each of the plurality of individuals, and features predictive of the presence of plaque within the geometric models and blood flow characteristics of the plurality of individuals; acquiring, for a patient, a geometric model and blood flow characteristics for part of the patient's vascular system; and executing the machine learning algorithm on the patient's geometric model and blood flow characteristics to determine, based on the predictive features, plaque information of the patient for at least one point in the patient's geometric model. | 1. A method for predicting information relating to a coronary lesion, the method comprising: acquiring, for each of a plurality of individuals, a geometric model, blood flow characteristics, and plaque information for at least part of each individual's vascular system;identifying, for each of a plurality of locations in the geometric models of the plurality of individuals, features associated with a probability of a presence of plaque within the geometric models, the probability of the presence of plaque being determined by an analysis of the intensity distribution of individual-specific image data, by an image processor;forming feature vectors comprising blood flow characteristics at each of the plurality of locations and information related to each of the plurality of locations in the geometric models, for each of the plurality of individuals;associating the feature vectors with the identified features associated with the probability of the presence of plaque within the geometric models, for each of the plurality of individuals;training a machine learning algorithm to predict the probability of the presence of plaque within a geometric model, from feature vectors comprising blood flow characteristics at a plurality of locations on a geometric model and information related to each of the plurality of locations, using the associated feature vectors;acquiring, for a patient, non-invasively obtained patient-specific image data of structures of the patient's anatomy and blood flow characteristics for each of a plurality of locations for at least part of the patient's vascular system;segmenting the received patient-specific image based on regions with known blood flow characteristics;generating a geometric model of the patient for at least part of the patient's vascular system using the segmented patient-specific image data, the geometric model including the plurality of locations and blood flow characteristics at each of the plurality of points;forming feature vectors comprising blood flow characteristics at each of the plurality of locations and information related to each of the plurality of locations in the geometric model of the patient;executing the trained machine learning algorithm on the patient's geometric model and blood flow characteristics to predict the probability of the presence of plaque within the geometric model of the patient, using the formed feature vectors; anddetermining, non-invasively, a patient-specific prediction or a patient-specific probability of a development of artery disease for at least one location in the patient's geometric model using the predicted probability of the presence of plaque within the geometric model of the patient. 2. The method of claim 1, wherein the plaque information includes one or more of: a presence or absence of plaque, a growth or shrinkage of plaque, a plaque composition, a change in plaque composition, a plaque shape, and a change in plaque shape. 3. The method of claim 1, further compri... | HEARTFLOW, INC. | G06F019/10 G06T007/00 A61B005/02 A61B005/00 G06N099/00 G06F019/00 G06K009/46 G06K009/66 G06N007/00
US09814417 | B2 | 2017-11-14 | Noninvasive measurement of flavonoid compounds in biological tissue | 2009-01-13 | Methods and apparatus are disclosed which facilitate the rapid, noninvasive and quantitative measurement of the concentration of flavonoid compounds, as well as their isomers and metabolites, in biological tissue such as human skin. Low-intensity, visible-light illumination of intact tissue provides for high spatial resolution, and allows for precise quantification of the flavonoid levels in the tissue. The preferred embodiments make use of a previously unknown, low-oscillator strength, optical absorption transition of flavonoids. This makes it possible to optically excite flavonoids in living human tissue outside the absorption range of other, potentially confounding skin chromophores. A system constructed in accordance with the invention includes a source of light for illuminating a localized region of tissue with light that overlaps the absorption bands of a flavonoid compound; a device for detecting the fluorescence emitted by the flavonoid compound resulting from the illumination; and a processor for determining the concentration level of the flavonoid compound based upon the detected fluorescence. | 1. A noninvasive method of measuring flavonoid levels in biological tissue, comprising the steps of: illuminating a localized region of tissue with light having wavelength from about 500 to 700 nm to overlap the long-wavelength absorption band tail of a flavonoid compound;detecting the fluorescence emitted by the flavonoid compound resulting from the illumination; anddetermining the concentration level of the flavonoid compound based upon the detected fluorescence. 2. The method of claim 1, including the step of using the concentration level to assess the antioxidant status of the tissue. 3. The method of claim 1, including the step of comparing the concentration level to levels of normal biological tissue to assess the risk or presence of a malignancy or other disease. 4. The method of claim 1, wherein the step of detecting the fluorescence emitted by the flavonoid compound includes the use of fluorescence spectroscopy. 5. The method of claim 1, wherein the step of detecting the fluorescence emitted by the flavonoid compound includes the use of an optical detector. 6. The method of claim 5, wherein the light is in the red region of the spectrum. 7. The method of claim 5, wherein the localized region of tissue is substantially melanin-free. 8. The method of claim 1, wherein the tissue is human skin. 9. The method of claim 8, wherein the skin is on a fingertip or other portion of a hand. 10. The method of claim 1, wherein the illumination is outside the absorption range of potentially confounding skin chromophores. 11. The method of claim 1, wherein the illumination excites the tissue flavonoids in their long-wavelength absorption tail outside the absorption range of other skin chromophores. 12. The method of claim 10, wherein the chromophores include carotenoids, blood, elastin, and collagen. 13. The method of claim 11, wherein the chromophores include carotenoids, blood, elastin, and collagen. 14. A system for measuring flavonoid levels in biological tissue noninvasively, comprising: a source of light for illuminating a localized region of tissue with light that overlaps the absorption bands of a flavonoid compound;a device for detecting the fluorescence emitted by the flavonoid compound resulting from the illumination; anda processor for determining the concentration level of the flavonoid compound based upon the detected fluorescence. 15. The system of claim 14, wherein the processor is operative to assess the antioxidant status of the tissue based upon the concentration level. 16. The system of claim 14, further including a memory for storing flavonoid concentration levels associated with normal biological tissue; and wherein the processor is operative to compare the determined concentration level to the stored levels to assess the risk or presence of a malignancy or other disease. 17. The system of claim 14, wherein the device for detecting the fluorescence emitted by the flavonoid compound includes a fluorescence spectrograp... | Longevity Link Corporation | A61B005/1455 A61B005/145 G01N021/64
US09804246 | B2 | 2017-10-31 | Magnetic resonance imaging apparatus | 2011-01-20 | According to one embodiment, an image processing apparatus includes a storage unit configured to store data of a series of slice images associated with a region including a target region of an object, a first rest period specifying unit configured to specify a first rest period based on a change between images of the series of slice images, and a second rest period specifying unit configured to specify a second rest period shorter than the first rest period by tracking the target region on a plurality of slice images corresponding to the specified first rest period or a rest period enlarged from the first rest period. | 1. A magnetic resonance (MR) imaging apparatus comprising: gradient magnetic field generators, at least one radio frequency (RF) coil coupled to an imaging volume, transmitting and receiving units for the at least one RF coil, and a computer system connected to control said gradient magnetic field generators and said RF coil to acquire and process MR signals from patient tissue disposed in said imaging volume, said computer system being configured torepeat an MR imaging scan on an overall heart of a patient,detect displacement of a patient diaphragm due to respiratory movement of the patient by executing an MR probe scan immediately before each of the MR imaging scans,displace an imaging range for each of the MR imaging scans based on the detected displacement of the diaphragm,before each said MR probe scan, generate a series of slice images by executing a preliminary MR scan which repeatedly images a region including the heart using electrocardiogram (ECG) gating,specify, for the series of slice images, a first rest period in which a variation in position of a coronary artery falls within a predetermined range in a cardiac cycle, based on a change in the MR image of the overall heart,specify a second rest period of the coronary artery in the cardiac cycle, after specifying the first rest period, by tracking movement of the coronary artery only within a local range including the coronary artery for a plurality of slice images of the series of slice images, wherein the plurality of slice images corresponds to the first rest period or a rest period enlarged from the first rest period, and the second rest period is included in the first rest period or the rest period enlarged from the first rest period, andreconstruct an image based on MR data acquired using said MR imaging scan in the second rest period. 2. The apparatus according to claim 1, wherein the computer system is further configured to remove a still portion from the series of slice images. 3. The apparatus according to claim 2, wherein the computer system is further configured to Fourier-transform a temporal change in pixel value of each pixel of the series of slice images and remove a pixel exhibiting a frequency not more than a predetermined frequency as the still portion. 4. The apparatus according to claim 2, wherein the computer system is further configured to determine, based on a spatial change in difference between a maximum pixel value and a minimum pixel value in a predetermined period, whether each pixel of the series of slice images is the still portion. 5. The apparatus according to claim 1, wherein the computer system is further configured to specify the first rest period based on a difference result between adjacent frames of the series of slice images. 6. The apparatus according to claim 1, wherein the computer system is further configured to specify the first rest period based on a correlation coefficient between adjacent frames of the series of slice images. 7. The... | TOSHIBA MEDICAL SYSTEMS CORPORATION | A61B005/055 G01R033/565 G06T007/20 G01R033/563 G01R033/567
US09802129 | B2 | 2017-10-31 | Internet sports computer cellular device | 2007-09-18 | A system that wirelessly integrates actual sports equipment with a computer and the internet to allow players remotely located from one another to play a competitive simulated sports game. An individual player may opt to play solo or practice to improve basic techniques. The system includes motion sensors connected to the player and a motion sensing device, all containing circuits and contact or motion sensors coupled with signal processing and radio frequency transmitter circuitry, thereby wirelessly communicate game performance information to a remote receiver-computer. The computer displays player information and visually simulates and controls a game between two players via the internet, having similar equipment and remotely located from each other. Standard sports equipment may be retrofitted with the sensors and associated circuitry to convert such equipment into “smart equipment” for use with the system. The system employs specially developed computer software to process player performance data, control game play, communicate game information between players, generate and control visual simulations and display player performance information. | 1. A system including a first sports device, a second sports device, and a third sports device, each sports device connected to a network of sports devices, the first sports device comprising a first sports apparatus, a first display screen, a first digital camera, a first processor, a first set of one or more motion sensors, a first set of one or more physiological sensors, a first GPS apparatus, a first speaker, a first microphone, a first power supply, the first sports device being operated by a first user in a first location;the second sports device comprising a second sports apparatus, a second display screen, a second digital camera, a second processor, a second set of one or more motion sensors, a second set of one or more physiological sensors, a second GPS apparatus, a second speaker, a second microphone, a second power supply, the second sports device being operated by a second user in a second location;the third sports device comprising a third sports apparatus, a third display screen, a third digital camera, a third processor, a third set of one or more motion sensors, a third set of one or more physiological sensors, a third GPS apparatus, a third speaker, a third microphone, a third power supply, the third sports device being operated by a third user in a third location;the first set of one or more motion sensors configured to be worn by the first user, the second set of one or more motions sensors configured to be attached to the first sports apparatus, the third set of one or more motion sensors configured to be worn by the third user or attached to the third sports apparatus;the first, second, and third sets of one or more motion sensors configured to capture velocity data in a non-visual format;the first processor in wireless communication with the first set of one or more motion sensors and the first GPS apparatus, and programmed to receive a first set of motion data from the first GPS apparatus and the first set of one or more motion sensors, the first processor configured to calculate a first numerical representation using the first set of motion data, transmit the first set of motion data to the second sports device and the third sports device, and receive a second set of motion data from the second sports device and a third set of motion data from the third sports device;the third processor in wireless communication with the third set of one or more motion sensors and the third GPS apparatus, and programmed to receive the third set of motion data from the third GPS apparatus and the third set of one or more motion sensors, the third processor configured to transmit the third set of motion data to the first sports device and the second sports device, and receive the first set of motion data from the first sports device and the second set of motion data from the second sports device;the second processor in wireless communication with the second set of one or more motion sensors and the second GPS apparatus, and programmed to ... | - | A63B069/36 A63F013/92 A63F013/34 A63B069/00 A63F013/00 A63F013/332 A63F013/327 A63F013/816 A63F013/216 A63F013/87 A61B005/00 A63B024/00 A63B069/26 A63B069/38 A63B071/06 A63B102/32 A63B102/02 A63B102/18
US09782068 | B2 | 2017-10-10 | System for diagnosis and therapy of gaze stability | 2010-02-02 | A system to improve vestibulo-ocular reflex (VOR) includes a display device; a head motion detector, the head motion detector being configured to detect at least one of a first angular rotational rate of a subject's head or a first acceleration of said subject's head in a first direction; a signal conditioner in communication with the head motion detector; a first gain control in communication with the signal conditioner; a second gain control in communication with the signal conditioner; a motion controller in communication with the display device and the first gain control and the second gain control; and a head motion-inducing device configured to induce a head motion on the subject. The head motion-inducing device includes a motor configured to rotate a head of the subject based on an instruction produced by a data processor. | 1. A system to improve vestibulo-ocular reflex (VOR), comprising a display device;a head motion detector, said head motion detector configured to detect at least one of a first angular rotational rate of a subject's head or a first acceleration of said subject's head in a first direction;a signal conditioner in communication with said head motion detector;a first gain control in communication with said signal conditioner;a second gain control in communication with said signal conditioner;a motion controller in communication with said display device and said first gain control and said second gain control; anda head motion-inducing device configured to induce a head motion on said subject, wherein the head motion-inducing device includes a motor configured to rotate a head of the subject based on an instruction produced by a data processor, wherein said signal conditioner is adapted to: remove interfering signals from head motion data,transmit a positive rectified signal to said first gain control, andtransmit a negative rectified signal to said second gain control, wherein the first gain control and the second gain control are individually adjusted,said motion controller is adapted to receive information encoding the subject's head motion from said first gain control and said second gain control, andsaid motion controller is further adapted to control the display device to display a visual target to said subject such that said visual target moves in a second direction opposite to said first direction in correspondence with the subject's head motion. 2. The system of claim 1, wherein said display device is at least one of a cathode ray tube (CRT) monitor, a liquid crystal display (LCD) monitor, a digital light projection (DLP) monitor, a projector display, a laser projector, a plasma screen, an organic light emitting diode (OLED). 3. The system of claim 1, wherein said motion controller comprises a galvanometer. 4. The system of claim 1, wherein said motion controller is further adapted to adjust a velocity of said visual target displayed by said display device. 5. The system of claim 1, wherein said display device, head motion detector, and motion controller are adapted to be worn by the subject. 6. The system of claim 1, wherein the correspondence is a correspondence between a head motion velocity and a visual target velocity. | THE JOHNS HOPKINS UNIVERSITY | A61B013/00 A61B003/113 A61B005/11 A61B005/00
US09814405 | B2 | 2017-11-14 | R-R interval measurement using multi-rate ECG processing | 2012-09-07 | A method and system for R-R interval measurement of a user are disclosed. In a first aspect, the method comprises detecting an electrocardiogram (ECG) signal of the user. The method includes performing QRS peak detection on the ECG signal to obtain a low resolution peak and searching near the low resolution peak for a high resolution peak. The method includes calculating the R-R interval measurement based upon the high resolution peak. In a second aspect, a wireless sensor device comprises a processor and a memory device coupled to the processor, wherein the memory device includes an application that, when executed by the processor, causes the processor to carry out the steps of the method. | 1. A device-implemented method for R-R interval measurement of a user, the method comprising: detecting, using the device, an electrocardiogram (ECG) signal of the user;performing, using the device, QRS peak detection on the ECG signal to obtain a low resolution peak;searching, using the device, proximate the low resolution peak for a high resolution peak by scanning a range of ECG sample points that span to the left and to the right of the low resolution peak; andcalculating, using the device, the R-R interval measurement based upon the high resolution peak, wherein the device comprises a wireless sensor device comprising a processor and a memory device coupled to the processor, and wherein the memory device comprises executable instruction. 2. The device-implemented method of claim 1, wherein the detecting is at a high sampling rate, further comprising: filtering, using the device, the ECG signal using an anti-aliasing filter;converting, using the device, the ECG signal from an analog domain to a digital domain using an analog-to-digital converter (ADC);filtering, using the device, the digital ECG signal using a band-pass filter;down-sampling, using the device, the digital ECG signal at a low sampling rate;processing, using the device, at least one ECG sample of the digital ECG signal using a digital signal processing (DSP) algorithm to detect a peak of the down-sampled ECG signal; andsearching, using the device, near the peak of the down-sampled ECG signal at the high sampling rate for the high resolution peak. 3. The device-implemented method of claim 2, wherein the processing further comprises: processing, using the device, every fourth ECG sample of the digital ECG signal using a robust digital signal processing (DSP) algorithm to detect a peak of the down-sampled ECG signal. 4. The device-implemented method of claim 2, wherein the searching further comprises: searching, using the device, a range of ECG sample points spanning left and right sides of the low resolution peak to determine the high resolution peak. 5. The device-implemented method of claim 4, wherein the range of ECG sample points spans from 3 points to the left to 3 points to the right of the low resolution peak. 6. The device-implemented method of claim 1, wherein the detecting is at a low sampling rate, further comprising: filtering, using the device, the ECG signal using an anti-aliasing filter;converting, using the device, the ECG signal from an analog domain to a digital domain using an analog-to-digital converter (ADC);filtering, using the device, the digital ECG signal using a band-pass filter;processing, using the device, the digital ECG signal using a digital signal processing (DSP) algorithm to detect the low resolution peak; andsearching, using the device, near the low resolution peak for the high resolution peak. 7. The device-implemented method of claim 6, wherein the searching further comprises: selecting, using the device, a range of ECG sample points span... | Vital Connect, Inc. | A61B005/0456 A61B005/00 A61B005/04
US09814412 | B2 | 2017-11-14 | Method and device for determining a blood constituent | 2012-11-26 | A method and device for determining the concentration of blood constituents, in particular haemoglobin, in a hose line of an extracorporeal blood circuit of an extracorporeal blood treatment apparatus, and an extracorporeal blood treatment apparatus with a device for determining the concentration of a blood constituent, are based on the correction of the influence of the blood flow rate of the blood flowing through the hose line on the determination of the concentration of the blood constituent. The device comprises a computing and evaluation unit configured such that a correction factor is ascertained for the influence of the blood flow rate on the determination of the concentration of the blood constituent. The concentration of the blood constituent is then determined based on a relationship describing the dependence of the concentration of the blood constituent on the intensity of the decoupled electromagnetic radiation, taking account of the correction factor. | 1. A method for determining a concentration of a blood constituent during an extracorporeal blood treatment performed by an extracorporeal blood treatment apparatus including a central control and computing unit that performs control and computing operations of the extracorporeal blood treatment apparatus, wherein blood flows through an arterial hose line into a blood chamber and via a venous hose line out of the blood chamber of a dialyser divided by a semipermeable membrane into the blood chamber and a dialysing fluid chamber, the method comprising: calibrating, prior to use in the extracorporeal blood treatment, a measuring arrangement using a calibration blood flow rate;storing the calibration blood flow rate in a memory of a computing and evaluation unit;determining, via the central control and computing unit, a current blood flow rate;storing the current blood flow rate in the memory of the computing and evaluation unit;setting a preset threshold value for the difference between the current blood flow rate and the calibration blood flow rate;storing the preset threshold value in the memory of the computing and evaluation unit;coupling, via an emitter device of the measuring arrangement, electromagnetic radiation through the arterial hose line or the venous hose line into the blood;measuring, via a detector of the measuring arrangement, the electromagnetic radiation decoupled through the arterial hose line or the venous hose line from the blood;ascertaining, via the computing and evaluation unit, a correction factor to correct for an influence of the blood flow rate on a determination of the concentration of the blood constituent wherein the correction factor is ascertained only if a deviation between the current blood flow rate and the calibration blood flow rate is greater than the preset threshold value as determined by the computing and evaluation unit;determining, via a determining device, the concentration of the blood constituent based on (a) a relationship describing a dependence of the concentration of the blood constituent on a measured intensity of the decoupled electromagnetic radiation, and (b) the correction factor; andtransmitting the concentration of the blood constituent from the determining device to the central control and computing unit of the extracorporeal blood treatment apparatus. 2. The method according to claim 1, wherein the relationship describing the dependence of the concentration of the blood constituent on the measured intensity of the decoupled electromagnetic radiation is described by the following equation: Hb=a·S(Qbk)−b wherein a and b are constants and S(Qbk) is a measurement signal correlating with the measured intensity of the decoupled electromagnetic radiation. 3. The method according to claim 1, wherein, if the deviation between the current blood flow rate and the calibration blood flow rate is greater than the preset threshold value as determined by the computing and evaluation unit, in order to t... | Fresenius Medical Care Deutschland GmbH | A61B005/1455 A61B005/145 G01N021/85 G01N021/53 A61M001/36
US09808148 | B2 | 2017-11-07 | Spatial 3D sterioscopic intraoral camera system | 2013-03-14 | Spatial 3D stereoscopic video is generated with a dual camera subassembly in a hand held device. The 3D Left-Right Multiplexed Spatial Stereoscopic video stream is wirelessly coupled to a processing system for simultaneous visualization comparison with intraoral cavity digital dental impression 3D CAD formatted files, Cone Beam Panoramic Computed Tomography images, Optical Coherence Tomography planar images and Panoramic X-ray image digital files for Stereoscopic visualization on stereoscopic displays and occluding viewing glasses. | 1. A spatial 3D stereoscopic intraoral camera system comprising: processor;memory;wireless component;dual side-by-side camera CCD subassemblies;the sensor subassemblies each having an optical path through a lens disposed to direct light to a CCD;each lens comprising an optical path through an optical lens or pin hole disposed to direct light on CCD;the duel side-by-side subassemblies each having a CCD imaging chip converging optical paths of subassemblies by tilt, tilt range of between 1.degree. to 5.degree. off the surface mount plane, or shift, shift gap symmetrical with the duel assembly centerline of between 0.5 mm and 5.0 mm, for focus control to achieve the stereoscopic convergence;the CCDs operatively connected and mounted on daughterboard's with a lens focal gap of between 0.01 and 3.0 mm each;each subassembly surrounded by an array of LEDs;each subassembly surface mounted on a motherboard supporting electronic components for processing the CCD image signals and LEDs power and control,software instructions in the form of device logic stored in memory for enabling the device, under control of the processors to format the signal images for wireless transport to remote processing system for storage and display in 3D Line Sequential Left-Right Multiplexed Spatial Stereoscopic format; andeach subassembly with programmable logic for frequency sweep through an array of LEDs with source light spectrum spanning from UV to Near Infrared; the LED array source light disposed to enter a beamsplitter in the subassembly adjacent to the lens;each beamsplitter refracted incoming light from the LEDs to a reference leg optical path normal to the incoming light and with a mirror on the reference path end reflecting reference beam back to beamsplitter and bending toward the subassembly lens;each LED between the beamsplitter mirror and the lens centerline positioned in a symmetric pattern LED array emitting Near Infrared light from a transparent surface mounted LED package and each beamsplitter having a straight through optical path for illumination and reflected image light recombining with reference light on route to the subassembly lens;whereby spatial 3D intraoral images are converged and captured by the duel subassemblies for processing and display in stereoscopic rendering and viewing. 2. A spatial 3D stereoscopic intraoral camera system as in claim 1 wherein each LED array comprise high-efficiency very bright white light, LED's interspersed between an RGB based LEDs that can be tuned to a specific color temperature and tunable discrete selectable blue and 400 nm to 450 nm wavelength UV illumination. 3. A spatial 3D stereoscopic intraoral camera system as in claim 1 wherein LEDs are symmetrically arranged encircling each the two camera subassemblies so as to minimize shadows or hot-spots on the target subject. 4. A spatial 3D stereoscopic intraoral camera system as in claim 1 wherein the white light and ultraviolet LED's are arranged in an alternating... | Jan Erich Sommers | A61B001/06 A61B001/00 A61B001/05 A61B001/24 A61B005/00 A61B006/00 A61B006/14
US09788748 | B2 | 2017-10-17 | Method of using human physiological responses as inputs to hydrocarbon management decisions | 2010-05-12 | A method of analyzing hydrocarbon-related data is disclosed. Data representative of a hydrocarbon entity is presented. A physiological response of a viewer of the data is sensed. The physiological response is associated with the data. The data and a representation of the associated physiological response is outputted. | 1. A method comprising: presenting, with a processor, a subsurface image of a hydrocarbon entity, wherein the hydrocarbon entity includes hydrocarbons that consist of hydrogen and carbon;sensing, with a sensor, a physiological response of a viewer of the subsurface image of the hydrocarbon entity;associating, with the processor, the physiological response with a localized region, which includes the hydrocarbons, of the subsurface image of the hydrocarbon entity, the localized region being less than an entirety of the subsurface image of the hydrocarbon entity, and the localized region being responsible for the physiological response;outputting, with the processor, the subsurface image of the hydrocarbon entity and a representation of the associated physiological response to the localized region, which includes the hydrocarbons; andusing the processor and the representation of the associated physiological response to identify a location for a new well or to identify a previously uninterpreted geological feature in the localized region of the subsurface image. 2. The method of claim 1, wherein presenting the subsurface image of the hydrocarbon entity comprises the processor causing the subsurface image of the hydrocarbon entity to be displayed on a display device. 3. The method of claim 2, wherein the representation of the associated physiological response is displayed in a graphical form. 4. The method of claim 2, wherein the representation of the associated physiological response is superimposed upon the subsurface image of the hydrocarbon entity. 5. The method of claim 1, wherein the physiological response comprises at least one of brainwave activity, movement of an eye, position of an eye, gaze, muscle movement, body temperature, heart rate, pulmonary performance, and change in tone of voice. 6. The method of claim 1, wherein the physiological response comprises at least one of a rate of use of an input device, and a position of an input device relative to the presented data representative of the hydrocarbon entity. 7. The method of claim 1, wherein outputting the subsurface image of the hydrocarbon entity and the associated representation of the physiological response comprises the processor causing the image and the representation to be stored in a memory. 8. The method of claim 1, wherein outputting the subsurface image of the hydrocarbon entity and the representation of the associated physiological response comprises the processor causing the data and a graphical representation of the physiological response to be displayed on a display device. 9. The method of claim 1, further comprising interpreting the physiological response based on information regarding the viewer. 10. The method of claim 1, wherein the viewer of the data is a first viewer of the subsurface image of the hydrocarbon entity, and further comprising: sensing a physiological response of a second viewer of the subsurface image of the hydrocarbon entity;associating th... | ExxonMobil Upstream Research Company | G01N033/48 G01N033/50 A61B005/0484 G06F019/12 G06F019/26 G06F019/28 A61B005/0205 A61B005/16
US09814431 | B2 | 2017-11-14 | Methods and systems for retrospective internal gating | 2008-05-05 | The present invention, in one form, is a method for deriving respiratory gated PET image reconstruction from raw PET data. In reconstructing the respiratory gated images in accordance with the present invention, respiratory motion information derived from individual voxel signal fluctuations, is used in combination to create usable respiratory phase information. Employing this method allows the respiratory gated PET images to be reconstructed from PET data without the use of external hardware, and in a fully automated manner. | 1. A method for retrospective internal gating comprising: acquiring a series of images at times t1 . . . tn;extracting time-activity information for individual voxels;prioritizing voxels for phase analysis, and assigning weighting factors;applying frequency filter to voxel time-activity curves;using prioritization, combining voxel time-activity information into a time varying object motion function;using the time varying object motion function for the mapping of image data to corresponding motion phases; andusing the mapping of image data to corresponding motion phases to generate at least one motion corrected image. 2. The method of claim 1, wherein time-activity information is defined as a voxel's value over times t1 . . . tn. 3. The method of claim 1, wherein the weighting factors include voxel specific weighting factors based on the mean voxel activity over times t1 . . . tn. 4. The method of claim 1, wherein the weighting factors include voxel specific weighting factors based upon the voxel's proximity to greater spatial signal gradients on an ungated image. 5. The method of claim 4, herein the ungated image is comprised of combined image data from times t1 . . . tn. 6. The method of claim 1, wherein the weighting factors include voxel specific weighting factors assigned based on magnitude of signal variation over times t1 . . . tn, for each voxel. 7. The method of claim 1, wherein at least some voxels are deemed unimportant and weighted at zero prioritization value. 8. The method of claim 1, wherein time-activity signals are filtered in frequency space for windows encompassing expected valid periodicity of the motion. 9. The method of claim 8, wherein the frequency window used may be adjusted to be patient or data specific. 10. The method of claim 8, wherein the frequency window may be adjusted over times t1 . . . tn. 11. The method of claim 1, wherein voxel time-activity information is processed serially in order of prioritization to yield a time varying object motion function. 12. The method of claim 11, wherein the time varying object motion function spans t1 . . . tn. 13. The method of claim 1, wherein the initial time varying object motion function is assigned to be the time-activity curve of the voxel with the highest priority. 14. The method of claim 1, wherein individual voxel information is combined with the evolving time varying object motion function in three possible scenarios: the first scenario is leaving the current time varying object motion function unaltered, the second scenario is adding the individual voxel time-activity information to the current time varying object motion function, and the third scenario is subtracting the individual voxel time-activity information from the current time varying object motion function, to account for possible phase mismatch; of these three scenarios, the one with the most significant improvement is chosen as the new time varying object motion function, to be used in evaluat... | Memorial Sloan Kettering Cancer Center | A61B006/00 A61B006/03 A61B005/00
US09814395 | B2 | 2017-11-14 | Method and apparatus for determination of physiological parameters using cervical impedance | 2012-08-07 | Implanted electrodes can be used to deliver electrical stimulation signals to areas near blood vessels, nerves, or other internal body locations. In an example, an electrode can be implanted in a cervical location and can be used to measure dimensional changes in an artery using impedance plethysmography. Measured artery dimensional changes can be used to determine one or more physiological parameters associated with a patient's health status, such as pulse transit time, relative pulse pressure, or aterial compliance, among others. These parameters can be used to monitor a patient health status or to modulate a patient's therapy, among other uses. In some examples, an electrode configured to deliver an electrostimulation signal to nerve tissue can be used to provide non-neurostimulating electrical stimulation plethysmography signals near a blood vessel. | 1. An apparatus, comprising: an electrode configured to be extravascularly implanted, the extravascularly-implanted electrode adapted for encircling extravascular nerve tissue near a blood vessel without encircling the blood vessel;an electrical energy delivery circuit configured to be coupled to the extravascularly-implanted electrode and configured to: provide an electrical test signal directed from the extravascularly-implanted electrode to a location internal to a patient body and near the blood vessel when the electrode encircles the extravascular nerve tissue without encircling the blood vessel; andprovide an electrostimulation signal directed from the extravascularly-implanted electrode toward the extravascular nerve tissue when the electrode encircles the extravascular nerve tissue without encircling the blood vessel;a detector circuit configured to receive an electrical response signal in response to the electrical test signal; anda signal processor circuit, coupled to the electrical energy delivery circuit and the detector circuit, the signal processor circuit configured to: determine a vessel dimension of the blood vessel near the extravascularly-implanted electrode using the received electrical response signal;compare the determined vessel dimension to a reference vessel dimension to determine a vessel dimension change; andoutput an indication of the determined vessel dimension change. 2. The apparatus of claim 1, wherein the electrical energy delivery circuit is configured to provide the electrical test signal to the extravascular nerve tissue, wherein the extravascular nerve tissue includes a vagus nerve in a cervical region near the blood vessel including a carotid artery. 3. The apparatus of claim 2, wherein the electrical energy delivery circuit is configured to provide the electrical test signal as a non-neurostimulating electrical test signal. 4. The apparatus of claim 2, wherein the electrical energy delivery circuit is configured to provide a series of non-neurostimulating electrical test signals; wherein the detector circuit is configured to receive a corresponding series of electrical response signals in response to the series of non-neurostimulating electrical test signals; andwherein the signal processor circuit is configured to determine the vessel dimension change of the blood vessel near the extravascularly-implanted electrode using the series of received electrical response signals. 5. The apparatus of claim 1, further comprising an electrode; wherein the electrical energy delivery circuit is configured to provide a neurostimulating therapy signal to the electrode; andwherein the electrical energy delivery circuit is further configured to provide a non-neurostimulating test signal to the same electrode. 6. The apparatus of claim 1, wherein the electrical energy delivery circuit is configured to provide a neurostimulating therapy signal directed from the extravascularly-implanted electrode toward a vagus nerve; an... | Cardiac Pacemakers, Inc. | A61B005/02 A61B005/00 A61B005/021 A61B005/0215 A61B005/0295 A61B005/053 A61B005/107
US09826198 | B2 | 2017-11-21 | Biological imaging device and biological imaging method | 2011-05-19 | A biological imaging device comprising: an irradiation unit which emits parallel light onto a first part in a living body; and an imaging unit which takes images of the first part and a second part that is connected to the first site. | 1. A biological imaging device, comprising: a light emitting device configured to emit light to a fingerprint part of a finger in a spot manner;an imaging device; anda processor configured to control the imaging device to image a fingerprint of the fingerprint part and a blood vessel of a finger part other than the fingerprint part of the finger connected to the fingerprint part at a same time, and to control the imaging device to output an image of a fingerprint, the image of the fingerprint captured reflected light from the fingerprint part, the reflected light has been emitted by the light emitting device and has been reflected on the surface of the fingerprint part, and an image of a blood vessel, the image of the blood vessel captured radiated light from the finger part connected to the fingerprint part of the finger, the radiated light has been emitted by the light emitting unit, has been propagated inside the finger part and has been radiated. 2. The biological imaging device according to claim 1: wherein the fingerprint part is a fingerprint part at least between a distal interphalangeal joint and a fingertip; andwherein the finger part is a part between at least a proximal interphalangeal joint and the distal interphalangeal joint. 3. The biological imaging device according to claim 1, wherein the fingerprint part and the finger part are kept at a position in which they are not contacted with any other part of the biological imaging device. 4. The biological imaging device according to claim 1, comprising a second light emitting device configured to emit light to a third part that is connected to the finger part and that is in a side opposite to that of the fingerprint part. 5. A biological imaging method, comprising: emitting light to a fingerprint part of a finger in a spot manner by a light emitting device;imaging a fingerprint of the fingerprint part, the fingerprint part with reflected light that has been emitted by the light emitting device and has been reflected on the surface of the fingerprint part, and a blood vessel of a finger part other than the fingerprint part of the finger connected to said fingerprint part at a same time, the finger part of the living body with radiated light that has been emitted by said light emitting unit, has been propagated inside the finger part and has been radiated from the finger part; andcontrolling output an image of the fingerprint, the image of the fingerprint imaged reflected light from the fingerprint part, and an image of a blood vessel, the image of the blood vessel imaged radiated light from the finger part connected to the fingerprint part. 6. The biological imaging method according to claim 5: wherein the fingerprint part is a fingerprint part at least between a distal interphalangeal joint and a fingertip; andwherein the finger part is a part at least between a proximal interphalangeal joint and the distal interphalangeal joint. 7. The biological imaging method according to cla... | NEC CORPORATION | G06K009/74 H04N007/18 A61B005/1171 A61B005/117 A61B005/1172 G06K009/00 A61B005/00 G01B011/00
US09800882 | B2 | 2017-10-24 | Dynamic digital image compression based on digital image characteristics | 2016-11-15 | Systems and techniques are disclosed for dynamically and automatically selecting an appropriate compression technique and/or compression parameters for digital images in order to reduce or prevent loss of significant information that may negatively impact the utility or usefulness of the digital images. For example, based on various image characteristics associated with a digital image, the system may dynamically compress the image using particular compression techniques and/or by adjusting compression parameters, to maintain significant information of the image. The system may select compression techniques and/or compression parameters based on one or more compression rules, which may be associated with image characteristics, patient characteristics, medical history, etc. Further, the system may, based on the one or more compression rules, compress the image to a maximum degree of compression while maintaining the significant information of the image. | 1. A method of digital image compression, the method comprising: causing execution of software instructions by one or more hardware computing devices in order to: receive an uncompressed digital image;determine a characteristic associated with the uncompressed digital image;access a compression rule associated with the characteristic;compress, based on the compression rule, the uncompressed digital image using each of a plurality of compression levels to generate a set of compressed digital images that are each compressed at different compression levels;for each of the compressed digital images of the set, determine a respective amount of errors based on comparison of the respective compressed digital images to the uncompressed digital image; anddetermine a first compressed digital image of the set that: is associated with an amount of errors that satisfies a threshold defined by the compression rule, andis compressed at a highest compression level of the plurality of compression levels that has an amount of errors that satisfies the threshold defined by the compression rule. 2. The method of claim 1, wherein the characteristic associated with the uncompressed digital image includes at least one of: an imaging modality, an anatomical feature, or an acquisition type. 3. The method of claim 1, wherein the plurality of compression levels each includes a compression algorithm and a set of respective quality factors, and the set of compressed digital images is generated based on the respective compression algorithms and quality factors. 4. The method of claim 1, wherein determining an amount of errors by comparison of a compressed digital image to the uncompressed digital image comprises: causing execution of software instructions by one or more hardware computing devices in order to: determine a difference between the uncompressed digital image and the compressed digital image to generate difference image data; anddetermine the amount of errors by at least one of: determining a number of pixels in the difference image data having a value indicative of an error, ordetermining a degree of error in one or more pixels of the difference image data. 5. The method of claim 1, wherein determining an amount of errors by comparison of a compressed digital image to the uncompressed digital image comprises: causing execution of software instructions by one or more hardware computing devices in order to: identify one or more pixels of the uncompressed digital image having an intensity value satisfying a threshold;for each pixel of the one or more pixels, determine a difference between the pixel of the uncompressed digital image and the pixel of the compressed digital image to generate difference image data;determine the amount of errors by at least one of: determining a number of pixels in the difference image data having a value indicative of an error, ordetermining a degree of error in one or more pixels of the difference image data. 6. The method of claim... | D.R. SYSTEMS, INC. | G06T007/00 H04N019/426 G06F019/00 H04N019/103 H04N019/119 H04N019/122 H04N019/157 H04N019/67 H04N019/895 H04N019/182 H04N019/154 H04N019/176 G06T009/00 H04N019/136 H04N019/65 A61B005/00 A61B005/055 H04N019/30 H04N019/124
US09782324 | B2 | 2017-10-10 | System and method for treating skin and underlying tissues for improved health, function and/or appearance | 2014-03-11 | Systems and methods for treating facial soft tissues of a patient in a healthcare or cosmetic treatment environment such as, for example, a spa, clinic, or a medical practitioner's office are presented herein. The facial treatment system includes a facial stimulator instrument and/or an acoustic oscillator to deliver therapeutic stimulation to the facial tissues of a patient including percussive massage, electrical stimulation, and acoustic stimulation. A facial treatment application executing on a processor of a computing device identifies a treatment protocol and operates the facial stimulator instrument and/or acoustic oscillator to implement the facial tissue treatment. | 1. A system for cosmetically treating tissue of a patient for at least one of improving skin appearance, reducing skin wrinkles, improving skin tone, or improving tissue function, the system comprising: a display configured to display information associated with the treatment of the tissue;an input in electrical communication with the display and including a key board and touch screen, the input configured to receive information associated with the treatment of the tissue;a central processing unit (CPU) in electrical communication with the input;a memory in electrical communication with the CPU and including treatment parameters associated with the treatment of the tissue;a first radio frequency (RF) head capable of being placed in electrical communication with the CPU and comprising an array of piezoelectric transducers, the array configured to generate RF over a range of frequencies not possible via a single piezoelectric transducer; anda RF receiver antenna capable of being placed in electrical communication with the CPU and configured to detect RF energy transmitted through the tissue from the first RF head;a plurality of second RF heads, each second RF head having a piezoelectric transducer tuned to a unique frequency and capable of being placed in electrical communication with the CPU; andan EMG sensor capable of being placed in electrical communication with the CPU and configured to detect electromyogram in the tissue;wherein, when the first RF head and RF receiver antenna are applied to the tissue, the system is configured to: a) cause the first RF head to administer RF energy to the tissue over a range of RF frequencies; b) cause the RF receiver antenna to sense the administered RF energy transmitted through the tissue; c) identify which RF frequency of the range of RF frequencies administered to the tissue has the most transmissibility through the tissue; and d) recommend a second RF head of the plurality of second RF heads that is capable of providing the identified RF frequency,wherein, when the recommended second RF head and EMG sensor are applied to the tissue, the system is configured to: a) cause the recommended second RF head to administer RF energy at the identified RF frequency to the tissue over a range of pulse frequencies; b) cause the EMG sensor to detect electromyogram in the tissue arising due to the RF energy administered to the tissue over the range of pulse frequencies; c) identify which pulse frequency of the range of pulse frequencies administered to the tissue causes the highest electromyogram readings in the tissue; and d) treat the tissue with the recommended second RF head at the identified RF frequency at the identified pulse frequency. 2. The system of claim 1, wherein the array is configured to generate RF over a range of between 500 KHz and 1.5 MHz. 3. The system of claim 2, wherein the array is configured to generate RF over a range of between 500 KHz and 1.5 MHz at steps of between 50 KHz and 200 KHz. 4.... | Sigma Instruments Holdings, LLC | A61H023/02 A61N001/32 A61B005/053 A61B005/05 A61H023/00 A61B005/0488
US09791281 | B2 | 2017-10-17 | Cognitive tour guide system | 2016-01-25 | Methods, computer program products, and systems are presented. The methods include, for instance: providing a cognitive tour guide service to a group of participants for a tour with an initial route planned by participants registration information and environment information along the initial route. During the tour, real time sensory data on the participants health characteristics, change in environment are collected by a cognitive agent accompanying the group to lead the tour are relayed to a cognitive tour guide engine, and real time multi-objective optimization is modeled and performed. The participants are regrouped responsive to their respective enjoyment and circumstances of the environment and, among a set of optimal solutions, a new route for the tour is selected based on participant preference modeled and trained by participant feedback. During the tour, the cognitive tour guide engine continuously optimizes route responsive to incoming real time sensory data. | 1. A computer implemented method for providing a cognitive tour guide service to a group of participants, comprising: collecting, by a cognitive tour guide engine running on a computer, sensory data on the participants and on environments of a tour in real time captured by two or more cognitive agents accompanying and leading the group for the tour with a present route such that the cognitive tour guide engine learns location of the group, respective pace of travel for the participants, and respective level of interest of the participants for a landmark in the present route;optimizing, by the cognitive tour guide engine, the present route by use of at least two objectives of the tour in real time during the tour, a first objective of the tour from the at least two objectives being to minimize expenses for the tour, and a second objective of the tour from the at least two objectives being to maximize enjoyment of the participants, based on the collected real time sensory data; andnotifying, by the cognitive tour guide engine, the participants and the two or more cognitive agents of two or more new routes generated from the optimizing in real time, respective new routes of the two or more new routes corresponding to each subgroup of two or more subgroups formed from the group of participants based on each objective of the at least two objectives of the tour, such that the participants of each subgroup may continue the tour according to the respective new route corresponding to respective subgroups. 2. The method of claim 1, further comprising: registering a first participant among the participants for the tour with an initial route prior to the tour, the initial route being calculated from individual participant registration data as provided by the first participant, the participant registration data comprising face image, body measurements, and data for health characteristics. 3. The method of claim 1, wherein the two or more cognitive agents are respective unmanned aerial vehicles capable of real time sensory data collection and real time data communication,wherein the sensory data comprises visual data, auditory data, and/or gestural data of the participants captured by respective devices installed on each cognitive agent,wherein each cognitive agent of the two or more cognitive agents communicates in real time with the cognitive tour guide engine, other cognitive agents, and a mobile communication terminal of respective participants in the group, andwherein said each cognitive agent relays data stream from the mobile communication terminal to the cognitive tour guide engine, the data stream being generated by an applet installed on the mobile communication terminal. 4. The method of claim 1, wherein the sensory data on the environment of the tour comprises topographical scan of the present route, obstacles on the present route, weather, delays of the tour caused by traffic, concentration of visitors in a landmark in the present route and/or ... | INTERNATIONAL BUSINESS MACHINES CORPORATION | G01C021/20 B64C039/02 G06Q050/14 A61B005/00 G06N005/02 A61B005/11 A61B005/01 A61B003/113 A61B007/04
US09788758 | B2 | 2017-10-17 | Method for determining abnormal gait | 2014-04-15 | The invention relates to a method for determining abnormal gait comprising the following steps: (a) measuring ground reaction force generated during walking by a plurality of sensors arranged on the left-foot and the right-foot respectively; (b) applying measurements from each of the plurality of sensors to a predetermined fuzzy membership function to transform the measurements into the first fuzzy values; (c) applying the first fuzzy values to a predetermined fuzzy logic to generate the second fuzzy values for a plurality of gait phases; and (d) comparing the second fuzzy values with pre-stored data of normal gait to determine whether it is abnormal gait or not. Therefore, it is possible to determine abnormal gait more accurately even with fewer sensors. | 1. A method for determining abnormal gait comprising: measuring, by a plurality of sensors, ground reaction force generated during walking, wherein the plurality of sensors are arranged on a left foot and a right foot of a walker;applying, by a fuzzy calculator, measurement values measured by each of the plurality of sensors to a predetermined fuzzy membership function to transform the measurement values into first fuzzy values;applying, by the fuzzy calculator, the first fuzzy values to a predetermined fuzzy logic to generate second fuzzy values corresponding to a plurality of gait phases; andcomparing, by a gait-state determiner, the second fuzzy values with pre-registered data about normal gait; anddiagnosing an abnormal gait or a normal gait of the walker based on the comparison of the second fuzzy values with the pre-registered data and reporting the diagnosis to a practitioner,wherein the fuzzy membership function is defined asfLarge⁡(x)=12⁢tanh⁡(s⁡(x-x0))+1∈[0,1]fSmall⁡(x)=1-fLarge⁡(x)∈[0,1]and, wherein fLarge(x) and fSmall(x) are the first fuzzy values, x is a measurement value, x0 is a predetermined reference value, and s is a sensibility coefficient. 2. The method according to claim 1, wherein in the applying the first fuzzy values, Perry's Gait Phase is applied for the plurality of gait phases and wherein the second fuzzy values corresponding to each Perry's Gait Phase are obtained by logic operation of the first fuzzy values which are the measurement values from at least two sensors of the plurality of sensors, by the fuzzy calculator. 3. The method according to claim 2, wherein in the comparing, the second fuzzy values are scaled by formulasf⁡(k)=1∑μPhase,i⁡(k)(where, μPhase,i(k) is the second fuzzy value for each gait phase at time k) and the scaled values are summed up for each gait phase, by a gait-state determiner. 4. The method according to claim 1, wherein the plurality of sensors are five sensors per foot. 5. The method according to claim 4, wherein the five sensors of the plurality of sensors are positioned in the plantar region of each foot. 6. The method according to claim 4, wherein the five sensors of the plurality of sensors are pressure sensors. 7. The method according to claim 4, wherein the five sensors of the plurality of sensors are foot sensor resistor. | KOREA UNIVERSITY RESEARCH AND BUSINESS FOUNDATION | A61B005/103 G06N005/04 A61B005/00
US09814841 | B2 | 2017-11-14 | Medical device with sliding frontal attachment and retractable needle | 2013-09-06 | A medical device having a frontal attachment and a connector housing having or attachable to an associated medical apparatus, the frontal attachment slidably engaging the connector housing and having a forwardly projecting, rearwardly biased needle and a needle retraction assembly, and the connector housing having a needle retraction cavity laterally offset from the needle in a first position, the needle retraction cavity being selectively movable relative to the frontal attachment following use to reposition the needle retraction cavity into alignment with the needle to permit retraction. | 1. A medical device comprising: a connector housing further comprising a fluid pathway and a needle retraction chamber that is laterally spaced apart from the fluid pathway;a frontal attachment disposed in sliding engagement with the connector housing, the frontal attachment further comprising a retractable needle that is forwardly projecting prior to use of the medical device and a needle retraction assembly with a needle holder that is rearwardly biased prior to use of the medical device; anda fluid seal disposed between the frontal attachment and the connector housing;wherein the connector housing and the frontal attachment comprise cooperatively configured elements forming a sliding interface that is transverse to a longitudinal axis through the needle, and allows sliding lateral movement between the frontal attachment and the connector housing between a first position where the fluid pathway is coaxially aligned with the needle retraction assembly and a second position where the needle retraction chamber is aligned with the needle retraction assembly. 2. The medical device of claim 1 further comprising a removable needle cover. 3. The medical device of claim 2 wherein the removable needle cover is configured to limit relative sliding movement between the frontal attachment and the connector housing prior to removal of the needle cover. 4. The medical device of claim 2 wherein the removable needle cover further comprises a locking member that limits relative sliding lateral movement between the connector housing and the frontal attachment. 5. The medical device of claim 1 in combination with an associated medical apparatus selected from the group consisting of syringes, vascular catheter insertion devices, fluid infusion sets, and fluid collection devices. 6. The medical device of claim 5 wherein the associated medical apparatus is a syringe. 7. The medical device of claim 1 wherein the connector housing further comprises a rearwardly facing connector that is configured to engage an associated medical apparatus in fluid communication with the retractable needle. 8. The medical device of claim 7 wherein the associated medical apparatus is selected from the group consisting of syringes, vascular catheter insertion devices, fluid infusion sets, and fluid collection devices. 9. The medical device of claim 7 wherein the associated medical apparatus is a syringe. 10. The medical device of claim 1 wherein the connector housing further comprises a luer lock or luer slip connector. 11. The medical device of claim 1 wherein the fluid seal is an annular seal seated in a recess disposed around a portion of the fluid pathway in a forwardly facing opening in the connector housing and blocks fluid leakage both prior to and following needle retraction. | Retractable Technologies, Inc. | A61M005/32 A61B005/15 A61B005/153 A61M025/06
US09791527 | B2 | 2017-10-17 | Extended detuning in local coils | 2014-06-30 | A local coil for an MRI imaging system includes an antenna containing a first detuning circuit and a second detuning circuit, and a connection connected to the antenna between a first connection point on the antenna and a second connection point on the antenna. The connection is configured to be short-circuited by at least one diode. The first connection point and the second connection point are situated spatially between a first partial region and a second partial region of the antenna. | 1. A local coil for an MRI imaging system, the local coil comprising: an antenna comprising a first detuning circuit and a second detuning circuit; anda connection connected to the antenna between a first connection point on the antenna and a second connection point on the antenna;wherein the connection is configured to be short-circuited by at least one diode,wherein the first connection point and the second connection point are situated spatially between a first partial region and a second partial region of the antenna, andwherein the first partial region and the second partial region of the antenna are substantially equal in size. 2. The local coil of claim 1, further comprising a connection between the first partial region and the second partial region of the antenna, wherein the connection between the first partial region and the second partial region of the antenna is configured for optional inhibition and activation by a single diode and a current source. 3. The local coil of claim 1, further comprising a connection between the first partial region and the second partial region of the antenna, wherein the connection between the first partial region and the second partial region of the antenna is configured for optional inhibition and activation by a first diode and a second diode, andwherein the first diode and the second diode are connected back-to-back. 4. The local coil of claim 1, wherein the first partial region and the second partial region form a conductor of the antenna. 5. The local coil of claim 1, wherein the first partial region and the second partial region are spatially symmetrical with respect to one another, wherein the at least one diode is connected to the antenna between the first connection point and the second connection point in a region between the first partial region and the second partial region. 6. The local coil of claim 1, wherein the first partial region and the second partial region are spatially symmetrical with respect to one another, wherein the at least one diode is connected on the antenna between the first partial region and the second partial region and between the first connection point and the second connection point, andwherein the first connection point and the second connection point are opposite one another in a region of an axis of mirror symmetry, on the antenna, or in the region of the axis of mirror symmetry and on the antenna. 7. The local coil of claim 1, wherein the first partial region and the second partial region are spatially symmetrical with respect to one another, wherein the at least one diode is connected to the antenna between the first partial region and the second partial region and between the first connection point and the second connection point,wherein each of the first connection point and the second connection point lies in a region of an axis of point symmetry of the antenna and in a region of overlapping crossover of the antenna, the antenna having a butterfly-type... | Siemens Aktiengesellschaft | G01V003/00 G01R033/385 G01R033/36 A61B005/055 G01R033/341
US09788783 | B2 | 2017-10-17 | Methods and apparatus for the planning and delivery of radiation treatments | 2016-12-16 | Methods and apparatus are provided for planning and delivering radiation treatments by modalities which involve moving a radiation source along a trajectory relative to a subject while delivering radiation to the subject. In some embodiments the radiation source is moved continuously along the trajectory while in some embodiments the radiation source is moved intermittently. Some embodiments involve the optimization of the radiation delivery plan to meet various optimization goals while meeting a number of constraints. For each of a number of control points along a trajectory, a radiation delivery plan may comprise: a set of motion axes parameters, a set of beam shape parameters and a beam intensity. | 1. A method for delivering radiation dose to a target area within a subject, the method comprising: defining a trajectory comprising relative movement between a treatment radiation source and the subject and determining a radiation delivery plan;sensing a position of the subject;effecting relative movement between the treatment radiation source and the subject along the trajectory and, during the relative movement between the treatment radiation source and the subject along the trajectory, delivering a treatment radiation beam from the treatment radiation source to the subject according to the radiation delivery plan to impart a dose distribution on the subject, wherein delivering the treatment radiation beam from the treatment radiation source to the subject comprises: varying an intensity of the treatment radiation beam over at least a portion of the trajectory;sensing that the position is outside of an acceptable range and deactivating delivery of the treatment radiation beam based on the sensing that the position is outside of an acceptable range; andsensing that the position is within the acceptable range and reactivating delivery of the treatment radiation beam based on the sensing that the position is within the acceptable range; andwherein varying the intensity of the treatment radiation beam over at least a portion of the trajectory comprises varying a radiation output rate of the treatment radiation source in accordance with the radiation delivery plan during the relative movement between the treatment radiation source and the subject along the trajectory. 2. A method according to claim 1 wherein the subject is a living subject. 3. A method according to claim 1 wherein defining the trajectory comprises defining an initial trajectory and wherein determining the radiation delivery plan comprises modifying the initial trajectory to provide the trajectory. 4. A method according to claim 1 wherein the position of the subject comprises a location of the target area and wherein sensing the position of the subject comprises sensing a location of the target area. 5. A method according to claim 1 wherein delivering the treatment radiation beam from the treatment radiation source to the subject comprises varying a shape of the treatment radiation beam over at least a portion of the trajectory. 6. A method according to claim 1 wherein sensing the position comprises, during the relative movement between the treatment radiation source and the subject along the trajectory, obtaining two-dimensional projection images of the target area at a plurality of locations along the trajectory. 7. A method according to claim 1 wherein effecting relative movement between the treatment radiation source and the subject along the trajectory comprises effecting continuous relative movement between the treatment radiation source and the subject. 8. A method according to claim 1 wherein effecting relative movement between the treatment radiation source and the s... | Varian Medical Systems International AG | A61B005/00 A61N005/10 A61B005/08 A61B005/113
US09808198 | B2 | 2017-11-07 | Biological information measuring device | 2015-12-02 | A biological information measuring device includes a case section having, in sectional view, a trapezoidal shape including an upper base and a lower base shorter than the upper base, a first leg substantially orthogonal to the upper base and the lower base, and a second leg that is the opposite side of the first leg, a display section disposed on the second leg side, a main circuit board housed in the case section, a flexible board configured to electrically connect the main circuit board and the display section, and a pulse-wave sensor section disposed on the first leg side and configured to detect a pulse wave signal of a user. The flexible board is disposed on the lower base side. | 1. A biological information measuring device comprising: a case section having, in sectional view, a trapezoidal shape including an upper base and a lower base shorter than the upper base, a first leg crossing the upper base and the lower base, and a second leg that is an opposite side of the first leg;a display section disposed on the second leg side;a circuit board housed in the case section;a flexible board configured to electrically connect the circuit board and the display section;a pulse wave sensor disposed on the first leg side and configured to detect a pulse wave signal of a user; anda vibration motor housed on the first leg side in the case section in sectional view,wherein the vibration motor is disposed further on the upper base side than a center of a surface on the first leg side. 2. The biological information measuring device according to claim 1, wherein the first leg and the upper and lower bases cross substantially orthogonally. 3. The biological information measuring device according to claim 1, wherein the upper base and the lower base are disposed in parallel. 4. The biological information measuring device according to claim 1, wherein the pulse wave sensor is disposed in a region including a center of gravity of the case section on a surface of the case section on the first leg side. 5. The biological information measuring device according to claim 1, wherein the pulse wave sensor is provided on a surface of the case section on the first leg side and disposed in a sensor projecting section projecting to a body side of the user. 6. The biological information measuring device according to claim 1, further comprising at least one battery housed in the case section, whereinthe battery is disposed with a center of gravity of the battery located further on the upper base side than a center of a surface on the first leg side. 7. The biological information measuring device according to claim 1, further comprising a band section configured to fix the case section to an organism, whereinthe band section includes a first band connected to the upper base side, a second band connected to the lower base side, and a connecting section that connects the first band and the second band. 8. The biological information measuring device according to claim 7, wherein surfaces of the first band and the second band on the organism side and a surface of the case section on the first leg side form a continuous surface. 9. The biological information measuring device according to claim 1, wherein a surface of the case section on the first leg side has a curved surface. 10. The biological information measuring device according to claim 1, wherein an atmospheric pressure sensor configured to detect an atmospheric pressure is housed in the case section, andthe atmospheric pressure sensor is disposed not to overlap the pulse wave sensor in plan view. 11. The biological information measuring device according to claim 10, wherein the atmospheric pr... | Seiko Epson Corporation | A61B005/0205 A61B005/00 A61B005/021
US09795884 | B2 | 2017-10-24 | Interactive gaming analysis systems and methods | 2016-01-11 | An interactive gaming system is disclosed. The system comprises at least one sensor that conveys information to the system about the physical, intellectual, mental, emotional, psychological or other type of ability of a user. The system uses the information to assess the existence and extent of a disability, and then implements a change to an aspect of the gaming environment, thus optimizing the gaming experience for the game player by accounting for the game player's disabilities. | 1. An interactive system for optimizing a gaming experience, the system comprising: a gaming analysis engine programmed to: receive sensor data from at least one sensor, wherein the sensor data comprises data captured by the at least one sensor and associated with a game player's interactions with a gaming environment while playing a computer game within the gaming environment;determine at least one disability status of the game player as a function of the sensor data, each of the at least one disability status being indicative of a respective disability; andcause a modification of a motion input corresponding to an input command of the computer game within the gaming environment as a function of the disability status, wherein the motion input is associated with a first user motion performed by the user detected by a motion-detecting user interface to cause the computer game to execute the input command, the modification comprising replacing the association of the motion input with the first user motion with an association of the motion input to a second user motion performed by the user detected by the motion-detecting user interface to cause the computer game to execute the input command, the second user motion corresponding to a motion performed according to a therapy program associated with the disability status. 2. The system of claim 1, wherein the step of causing the modification further comprises implementing the modification while the game player is playing the computer game. 3. The system of claim 1, wherein the disability status includes at least one indicator of at least one disability and of a severity level of the disability and wherein the modification is based at least in part on the severity level. 4. The system of claim 3, wherein the gaming analysis engine is further programmed to: receive updated sensor data from the at least one sensor, wherein the updated sensor data comprises data captured by the at least one sensor and associated with the game player's interactions with the gaming environment while playing the computer game within the gaming environment at a second point in time;update the at least one disability status of the game player as a function of the updated sensor data; andadjust the modification as a function of the updated at least one disability status. 5. The system of claim 1, wherein the at least one disability status comprises at least one disability indicator derived from the sensor data and from at least one reference data set. 6. The system of claim 5, wherein the at least one reference data set is generated based on aggregated sensor data corresponding to a plurality of game players known to have no disabilities. 7. The system of claim 1, wherein the at least one disability status comprises at least one of a physical disability, a sensory disability, a motor disability, a cognitive disability, a developmental disability, a social disorder, an intellectual disability, an emotional disability, a le... | DATADNA, INC. | A63F013/00 A63F013/79 A61B005/16 A63F013/45 A63F013/212 G06F003/048 A63F013/215 A63F013/218 A63F013/211 A61B005/00
US09814540 | B2 | 2017-11-14 | Method, apparatus and article for detection of transponder tagged objects, for example during surgery | 2016-11-10 | The presence or absence of objects is determined by interrogating or exciting transponders coupled to the objects using pulsed wide band frequency signals. Interrogation is broken down into a number of subsample scan cycles each having interrogation cycles a start time forward in time by a fraction of a period of an expected transponder response signal. Ambient or background noise is evaluated and a threshold adjusted based on the level of noise. Adjustment may be based on multiple noise measurements or samples. Noise detection may be limited, with emphasis placed on interrogation to increase the signal to noise ratio. Matched filtering may be employed. Presence/absence determination may take into account frequency and/or Q value to limit false detections. Appropriate acts may be taken if detected noise is out of defined limits of operation, for example shutting down interrogation and/or providing an appropriate indication. | 1. A transponder detection device, comprising: an antenna;a processor operatively coupled to the antenna; anda nontransitory processor-readable medium communicatively coupled to the processor and that stores at least one of instructions or data executable by the processor, which cause the processor to: execute a number of successive interrogation cycles, the number of successive interrogation cycles greater than one, wherein, during each of the number of successive interrogation cycles the transponder detection device, emits a periodic electromagnetic interrogation signal during a transmit portion of the interrogation cycle;receives electromagnetic signals during a receive response portion of the interrogation cycle that follows the transmit portion of the interrogation cycle; andwaits a respective delay period before a respective start of each of the successive interrogation cycles of the number of interrogation cycles, each respective delay period for the number of interrogation cycles equal to a respective fraction of a nominal tag response signal period, and, for each one of the successive interrogation cycles of the number of interrogation cycles, the respective delay period is different than at least one other respective delay period of an interrogation cycle of the number of successive interrogation cycles by a duration equal to the nominal tag response signal period divided by the number of successive interrogation cycles. 2. A transponder detection device, comprising: an antenna;a processor operatively coupled to the antenna; anda nontransitory processor-readable medium communicatively coupled to the processor and that stores at least one of instructions or data executable by the processor, which cause the processor to: obtain a number of subsamples, each of the subsamples associated with a respective delay period equal to a respective fraction of a nominal tag response signal period, the respective fractions different from one another, each of the number of subsamples obtained by: during each of a plurality of successive interrogation cycles associated with each of the number of subsamples, the transponder detection device, emits via the antenna a periodic electromagnetic interrogation signal during a transmit portion of the interrogation cycle;receives via the antenna electromagnetic signals during a receive response portion of the interrogation cycle that follows the transmit portion of the interrogation cycle; andwaits a delay period of the one of the number of subsamples to which the interrogation cycle is associated before a start of each of the successive interrogation cycles of the plurality of interrogation cycles, wherein waiting a delay period before a start of each of the successive interrogation cycles of the plurality of interrogation cycles comprises waiting a delay period before a start of each of the successive interrogation cycles of the plurality of interrogation cycles and, the respective delay period associated with... | Covidien LP | A61B090/98 G06K007/10 G01V015/00 A61B005/06 A61B090/90
US09821169 | B2 | 2017-11-21 | Systems and methods for providing a magnetic resonance treatment to a subject | 2017-06-26 | Disclosed here in are a system and method for providing a magnetic resonance treatment to a subject. In one described embodiment, a system for providing a magnetic resonance treatment to a subject is provided, the system including components that provide the system with an ability to treat pain and relieve symptoms of the subject. In another embodiment, a method of providing a magnetic resonance treatment to a subject is provided, wherein the system includes components that provide the system with an ability to treat pain, relieve symptoms, provide relaxation, and improve the overall comfort and well-being of the subject. | 1. A method of providing a magnetic resonance treatment to a subject, comprising: querying a database of a plurality of magnetic resonance treatment protocols and parameters;generating a magnetic resonance treatment protocol based at least in part on a result of the query;providing a first magnetic resonance treatment to the subject using a magnetic resonance device comprising a magnetic assembly;controlling, by a magnetic resonance driver, the magnetic resonance device according to the magnetic resonance treatment protocol, wherein the magnetic resonance driver comprises an electronic waveform generator;compensating using a compensation network comprising a resistor and capacitor network configured to match an impedance of the magnetic resonance driver, wherein the magnetic assembly is coupled in parallel with the compensation network;determining a magnetic field strength output by the magnetic resonance device; andmodifying the magnetic resonance treatment based in part on a biometric measurement and the determined magnetic field strength. 2. The method of claim 1, wherein the biometric measurement comprises a measurement received from one or more of: a blood pressure sensor;a perspiration sensor;a body weight sensor;a body temperature sensor;a haptics glove sensor;an electroencephalograph; oran electrocardiograph. 3. The method of claim 1, wherein querying the database comprises at least one of: searching the database for prior treatment data of the subject;searching the database for a treatment of a like-situated subject with a medical history similar to the subject; orsearching the database for a subject population trend. 4. The method of claim 1, wherein modifying the magnetic resonance treatment comprises applying a second magnetic resonance treatment to the subject at a different magnetic field strength than the first magnetic resonance treatment. 5. The method of claim 4, wherein modifying the magnetic resonance treatment further comprises: monitoring the subject during the treatment;providing subject data to a processor; anddetermining whether a treatment adjustment is needed, and if it is determined that the treatment adjustment is needed, identifying the needed adjustment. 6. The method of claim 5, wherein determining whether a treatment adjustment is needed comprises comparing an intended result of the treatment to an actual result of the treatment. 7. The method of claim 1, further comprising ending treatment, wherein ending the treatment of the subject comprises: commanding a processor to end the treatment;using a processor to collect and store subject data; andproviding the subject data to a database. 8. The method of claim 1, further comprising obtaining access to the magnetic resonance treatment protocol. 9. The method of claim 8, wherein obtaining access to the magnetic resonance treatment protocol comprises determining whether an operator has a subscription to the magnetic resonance treatment protocol. 10. The method ... | Applied Magnetics, LLC | G06F015/18 A61N002/02 A61B005/00 G06F019/00 A61B005/0205 A61B005/021 A61B005/01 A61B005/0402 A61B005/0476
US09791535 | B2 | 2017-10-17 | Apparatus and method for mapping and measurement of spatially and temporally distributed induced small phase deviations in magnetic resonance utilizing deconvolution | 2013-03-15 | Phase variations of the transverse magnetization in magnetic resonance induced by superimposed physical phenomenae or by intrinsic deviations of the main magnetic B0 field are separated from Feature Space set by demodulation and deconvolution, either by electrical circuits or by equivalent computational methods, permitting mapping and measurement of these induced phase variations independent of Feature Space. | 1. A system for identifying spatially and temporally distributed phase deviations within a sample under evaluation, comprising: a magnetic resonance machine including at least one quadrature coil set and a magnetic resonance machine radio frequency generator, wherein the magnetic resonance machine applies a main magnetic field to the sample under evaluation within the magnetic resonance machine;an induction component configured to induce spatially and temporally ordered phase deviations of the transverse magnetization within the sample under evaluation;a computer system connected to the magnetic resonance machine, wherein the computer system is configured to: receive a first voltage from the at least one quadrature coil set;receive a second voltage from the at least one quadrature coil set;receive an output of the magnetic resonance machine radio frequency generator and phase-shift the output to a first and second phase-shifted voltage;multiply the first voltage by the first phase-shifted voltage to generate a first result;multiply the second voltage by the second phase-shifted voltage and invert the result to generate a second result;sum the first and second results to generate a spatial Fourier transform of a Feature Space representing an original native data set of the sample under evaluation produced by the magnetic resonance machine without the induction of the spatially and temporally ordered phase deviations;apply an inverse Fourier Transform to the spatial Fourier transform of the Feature Space representing the original native data set to generate a Feature Space which is the original native data set;multiply the first voltage by the second phase-shifted voltage to generate a third result;multiply the second voltage by the first phase-shifted voltage to generate a fourth result; andsum the third and fourth results to generate a spatial Fourier transform of an Aberrational Feature Space representing the product everywhere of the original native data set and the induced spatially and temporally ordered phase deviations;apply a deconvolution algorithm to the spatial Fourier transform of the Aberrational Feature Space and the spatial Fourier transform of the Feature Space to generate a fifth result; andapply an inverse Fourier Transform to the fifth result to generate a Phase Aberration Space which is the data set of the spatially and temporally ordered phase deviations. 2. The system of claim 1, wherein the computer system is further configured to generate at least one image using the Feature Space which is the original native data set. 3. The system of claim 1, wherein the computer system is further configured to generate at least one image using the Phase Aberration Space. 4. A method of identifying spatially and temporally distributed phase deviations within a sample under evaluation by a magnetic resonance machine with at least one quadrature coil set and a magnetic resonance machine radio frequency generator, comprising: applying a m... | Wollin Ventures, Inc. | G01R033/565 A61B005/055 G01R033/36 G01R033/24 G01R033/34 G01R033/56
US09814896 | B2 | 2017-11-14 | Intra cardiac device, system and methods | 2015-09-25 | An intra cardiac device is disclosed. The device comprises means for transforming kinetic energy from heart tissue movement into electrical energy in use, from which electrical energy information in respect of heart function is obtainable. Furthermore, a system is disclosed, comprising one such intra cardiac device and at least one receiver, wherein the intra cardiac device comprises means of communication, through which said at least one device communicates with the receiver(s) wirelessly. In this way energy from heart movement provides self contained intra cardiac devices for conveniently monitoring or stimulating a patient's heart. | 1. A method for monitoring heart function and operating an intra cardiac device, said heart function being a heart muscle condition, said method comprising: implanting said intra cardiac device within a body of a patient such that said device has at least one mechanical connection to the patient's heart;transforming kinetic energy from heart tissue movement into electrical energy via said device;powering said device with said electrical energy;wirelessly communicating information related to said electrical energy to a receiver for interpretation and analyzing said information to determine a characteristic of said electrical energy that is indicative of said heart function;wherein implanting said intra cardiac device comprises fixing a screw of said intra cardiac device in ventricular heart tissue; and,wherein transforming said kinetic energy comprises transforming with an energy conversion unit that comprises a piezo-electric element configured to convert the kinetic energy from ventricular heart tissue movement into an electrical signal. 2. The method according to claim 1, wherein said receiver is located in an external data assembly equipment outside of the patient body. 3. The method according to claim 2, wherein said external data assembly equipment is a mobile terminal such as a mobile telephone, or a fix-net telephone line, an intensive care monitor, an infusion pump or a transceiver of a telemetry system. 4. The method according to claim 1, further comprising recording said electrical signal and analyzing said recording to obtain a maximum amplitude of the electrical signal, which is indicative of how fast the heart is accelerating. 5. The method according to claim 1, further comprising recording said electrical signal and analyzing said recording to measure an amount of energy produced. 6. The method according to claim 1, further comprising using information obtained in said observing step in an external data assembly equipment for monitoring heart movement. 7. The method according to claim 1, further comprising recording said electrical signal and analyzing said recording to obtain electrical energy information indicative of said heart function for monitoring said heart muscle condition. 8. The method according to claim 7, wherein said analyzing said recording to obtain electrical energy information indicative of said heart function comprises obtaining information in respect of heart movement amplitude, or heart movement acceleration from said electrical signal. | Tor Peters | A61N001/378 A61N001/39 H02N002/18 A61B005/0245 A61N001/365 A61B005/11 A61B005/00 A61N001/362 A61N001/372
US09808176 | B2 | 2017-11-07 | Method and magnetic resonance system for magnetic resonance thermometry | 2014-03-20 | In magnetic resonance (MR) thermometry, first and second magnetic resonance data that are acquired sequentially, from which a phase shift is determined between a measurement phase and a reference phase. In a first region of an examined person, the measurement phase is determined from the second magnetic resonance data and the reference phase is determined from the first magnetic resonance data. In a second region of the examined person, both the reference phase and measurement phase are determined from the second magnetic resonance data. Techniques for reference-based magnetic resonance thermometry and reference-free magnetic resonance thermometry can be applied. | 1. A method for magnetic resonance (MR) thermometry of a patient, comprising: operating an MR data acquisition unit, in which the patient is situated, to acquire first MR data from the patient at a first point in time;operating said MR data acquisition unit in which said patient is situated to acquire second MR data from the patient at a second point in time that occurs after said first point in time;providing said first MR data and said second MR data to a computerized processor and, in said processor, automatically determining a temperature of a first region of the patient by determining a phase shift between said second MR data in said first region, as a measurement phase, and said first MR data in said first region, as a reference phase;in said processor, automatically determining a temperature in a second region of the patient, which is different from said first region, by determining a phase shift between said second MR data in said second region, as a measurement phase, and said second MR data in a third region of the patient that is different from said first region, and that is situated adjacent to said second region, as a reference phase; andmaking the temperature determined in at least one of said first region or said second region available at an output of said processor in electronic form. 2. A method as claimed in claim 1 comprising operating said MR data acquisition unit with said patient situated therein, prior to said first point in time, to acquire a plurality of sets of additional MR data from the patient; and providing said sets of additional MR data to said processor and, in said processor, establishing said first region and said second region and said third region from said sets of additional MR data. 3. A method as claimed in claim 2 comprising, in said processor: determining a spatially resolved first temperature curve for said sets of additional MR data by determining a phase shift between each individual set of additional MR data, as a measurement phase, and defined data among said additional MR data, as a reference phase;determining a spatially resolved second temperature curve for said sets of additional MR data by determining a phase shift between a same region in each individual additional MR data set, as a measurement phase, and a region adjacent to said same region in each individual set of additional MR data, as a reference phase; andestablishing said first region and said second region and said third region dependent on a spatially resolved evaluation of said first temperature curve and said second temperature curve. 4. A method as claimed in claim 3 comprising: allowing no introduction of heat into said patient during acquisition of said sets of additional MR data; anddetermining said first region as at least one region of the patient for which said first temperature curve is temperature-stable; anddetermining said second region as at least one region of the patient for which said second temperature curve is t... | Siemens Aktiengesellschaft | G01R033/48 A61B005/055 A61B005/01
US09801592 | B2 | 2017-10-31 | Caster system for mobile apparatus | 2014-03-14 | A mobile imaging device includes a base having at least one caster, a first drive mechanism that moves the system in a transport mode and translates an imaging component relative to the base in a scan mode, and a second drive mechanism that extends caster relative to the base to raise the base off the ground in the transport mode, and retracts the caster relative to the base to lower the base to the ground in the scan mode. A caster system for a mobile apparatus includes a base containing a housing and a caster, attached to the base, the caster having a wheel defining a wheel axis and a swivel joint defining a swivel axis and a pivot point defining a pivot axis, wherein the caster pivots on the pivot axis as the caster is retracted into the housing and extended out of the housing. | 1. A mobile imaging system, comprising: a base having at least one caster;a first drive mechanism that moves the entire mobile imaging system in a transport mode and translates at least one imaging component relative to the base in a scan mode; anda second drive mechanism that extends the at least one caster relative to the base to raise the base off the ground in the transport mode, and retracts the at least one caster relative to the base to lower the base to the ground in the scan mode, wherein the base comprises at least one housing for the at least one caster, and the at least one caster retracts into the at least one housing and extends out of the at least one housing and the at least one caster comprises a wheel defining a wheel axis and a swivel joint defining a swivel axis and a pivot point defining a pivot axis, wherein the at least one caster pivots on the pivot axis as the at least one caster is retracted into the at least one housing and extended out of the at least one housing, the at least one caster has a first offset distance between the wheel axis and the swivel axis when the at least one caster is extended and a second offset distance between the wheel axis and the swivel axis when the at least one caster is retracted, and the second offset axis is less than the first offset axis, and the at least one housing for the caster has at least one dimension such that the at least one housing is not large enough to receive the at least one caster when the at least one caster has the first offset distance. 2. The mobile imaging system of claim 1, wherein the at least one imaging component is raised away from the ground in response to the second drive mechanism extending the at least one caster relative to the base and the at least one imaging component is lowered towards the ground in response to the second drive mechanism retracting the at least one caster relative to the base. 3. The system of claim 2, wherein the second drive mechanism generates a motive force in a direction substantially along a length of the base, and a linkage apparatus translates the motive force to a substantially vertical force that extends or retracts the at least one caster. 4. The system of claim 2, further comprising a cable management system in the base that couples power and/or data between the base and at least one component that translates relative to the base. 5. The system of claim 4, wherein the cable management system comprises a cable chain having a first end connected to the base and a second end connected to the first drive mechanism. 6. The system of claim 5, wherein the base comprises a housing and the cable chain is located in the housing and extends in a loop between the first and the second end, and the first drive mechanism translates relative to the base in the scan mode and a leading edge of the loop travels at a lower speed than a speed at which the first drive mechanism translates relative to the base. 7. The system of claim 6, wh... | MOBIUS IMAGING, LLC | B60B033/06 A61B006/03 A61B005/055 A61B006/00
US09801693 | B1 | 2017-10-31 | Method and system for correlating anatomy using an electronic mobile device transparent display screen | 2017-04-03 | Method and system form a correlated view of human or other animal anatomy using at least one transparent display screen associated with an electronic mobile device. The view relates an optical view with other electromagnetic spectrum images with a non-optical electromagnetic image of selected portions of human or other animal anatomy. At least three visible position markers associate with selected positions of a predetermined portion of human or other animal anatomy. The disclosure forms a correlated view of the predetermined portion of human or other animal anatomy by relating said at least one non-optical electromagnetic image of the at least three visible position markers with a visual image of said at least three visible position markers. The view correlates the size and dimensions of the optical view and non-optical electromagnetic image of the predetermined portion of human or other animal anatomy. | 1. A method for forming a correlated view of human or other animal anatomy using at least one transparent display screen associated with an electronic mobile device, said view relating an optical view with other electromagnetic spectrum images with a non-optical electromagnetic image of selected portions of human or other animal anatomy, the method comprising the steps of: associating at least three visible position markers with selected positions of a predetermined portion of human or other animal anatomy, said at least three visible position markers comprising a predetermined measure of opacity for selected non-optical electromagnetic frequencies;imaging said predetermined portion of human or other animal anatomy using at least a subset of said selected non-optical electromagnetic frequencies using an electromagnetic imaging device for forming at least one non-optical electromagnetic image of said predetermined portion of human or other animal anatomy, including said at least three visible position markers and further wherein said at least three visible position markers exhibit at least partial opacity for at least one of said subset of selected non-optical electromagnetic frequencies;forming an optical view of said predetermined portion of human or other animal anatomy through at least a portion of a transparent display screen associated with said electronic mobile device;forming a correlated view of said predetermined portion of human or other animal anatomy by relating said at least one non-optical electromagnetic image of said at least three visible position markers with a visual image of said at least three visible position markers, said correlated view correlating the size and dimensions of said optical view and said at least one non-optical electromagnetic image of said predetermined portion of human or other animal anatomy. 2. The method of claim 1, wherein said non-optical electromagnetic frequencies comprise X-ray frequencies and said at least one non-optical electromagnetic image comprises an X-ray image. 3. The method of claim 1, wherein said non-optical electromagnetic frequencies comprise magnetic resonance frequencies and said at least one non-optical electromagnetic image comprises a magnetic resonance image. 4. The method of claim 1, wherein said predetermined portion of human or other animal anatomy comprises flesh tissue. 5. The method of claim 1, wherein said predetermined portion of human or other animal anatomy comprises skeletal anatomy. 6. The method of claim 1, wherein said at least three visible position markers exhibit at least partial opacity to the non-optical electromagnetic frequencies. 7. The method of claim 1, wherein said at least three visible position markers further comprise electromagnetic transmitters for transmitting position signals from said selected positions of said predetermined portion of said human or other animal anatomy. | SK Commercial Construction, Inc. | A61B090/00 A61B005/00 A61B005/055 A61B006/00
US09786048 | B2 | 2017-10-10 | System and method using precious-metal nanoparticle contrast agent for microwave medical imaging | 2014-02-07 | A system and method of imaging tissue includes administering a contrast agent having charged gold nanoparticles in suspension into a vessel of the subject, such that the nanoparticles are carried into the tissue; and performing microwave imaging of the tissue after administering the contrast agent. In embodiments, the nanoparticles have a tissue-selective protein tag. In embodiments, images are taken prior to administering the contrast agent, and further images may be taken during an agent—washout period after imaging with contrast agent. The contrast agent is injectable, with the nanoparticles suspended as a colloid in a biocompatible, isotonic, carrier. In particular embodiments, the nanoparticles have median diameter of less than fifty nanometers, or less than five nanometers, and may have a tissue-selective protein tag. A microwave imaging system has injection apparatus with the gold-nanoparticle agent, and is configured to take, and difference, pre and post contrast images as well as washout images. | 1. A system for performing microwave imaging of biological tissue comprising: apparatus adapted to administer a contrast agent, the apparatus containing a contrast agent comprising a suspension of gold nanoparticles in a biocompatible, isotonic, carrier; wherein the gold nanoparticles are charged nanoparticles having median diameter less than fifty nanometers;at least one microwave radio frequency source;an antenna array having a plurality of antennae disposed about a cavity;at least one microwave receiver;switching apparatus configured to couple the at least one radio frequency source to particular antennae of the antenna array under control of a processor, and to couple particular antennae of the antenna to receivers of the at least one microwave receiver under control of the processor;a data acquisition subsystem adapted to provide the processor with measurements derived from the at least one microwave receiver; andwhere the processor is coupled to a memory containing microwave imaging routines for constructing a model of tissue and fitting permittivity parameters of the model to the measurements derived from the at least one microwave receiver, and for constructing images representing distribution of the permittivity parameters. 2. The system of claim 1 wherein the nanoparticles have median diameter of less than five nanometers. 3. The system of claim 2, the memory further containing routines comprising machine readable instructions that are executable by the processor for obtaining a pre-contrast image derived from at least permittivity parameters fitted to measurements derived from the at least one microwave receiver prior to operation of the apparatus for administering the contrast agent, and for obtaining a post-contrast image derived from at least permittivity parameters fitted to measurements derived from the at least one microwave receiver after operation of the apparatus for administering the contrast agent. 4. The system of claim 3 wherein the cavity is adapted to permit positioning of a human breast within the cavity. 5. The system of claim 3, the memory further containing differencing routines comprising machine readable instruction executable by the processor to prepare a differential image from the pre-contrast and post-contrast images. 6. The system of claim 5 wherein the memory further contains routines comprising machine readable instruction executable by the processor for obtaining a washout image derived from at least permittivity parameters fitted to measurements derived from the at least one microwave receiver a predetermined time after operation of the apparatus for administering the contrast agent, and for obtaining a differential washout image by differencing the washout image from the post-contrast image. 7. The system of claim 1 wherein the nanoparticles have a tissue-selective protein tag. 8. A method of imaging tissue in a subject comprising: administering a contrast agent comprising charged gold nanoparticle... | THE TRUSTEES OF DARTMOUTH COLLEGE | G06K009/00 G06T007/00 A61K049/00 A61B005/05 A61B005/00 A61K049/18 G06T001/00
US09792531 | B2 | 2017-10-17 | Intelligent multi-scale medical image landmark detection | 2017-01-03 | Intelligent multi-scale image parsing determines the optimal size of each observation by an artificial agent at a given point in time while searching for the anatomical landmark. The artificial agent begins searching image data with a coarse field-of-view and iteratively decreases the field-of-view to locate the anatomical landmark. After searching at a coarse field-of view, the artificial agent increases resolution to a finer field-of-view to analyze context and appearance factors to converge on the anatomical landmark. The artificial agent determines applicable context and appearance factors at each effective scale. | 1. A method for intelligent multi-scale image parsing, the method comprising: specifying, by a processor, a state space of an artificial agent for discrete portions of a training image, the state space specified by a parametric space and a scale space for the discrete portions of the training image;determining, by the processor, a set of actions, the set of actions comprising parametric actions specifying a possible change in the parametric space with respect to the training image and scale actions specifying a possible change in the scale space with respect to the training image;establishing, by the processor, a reward system based on applying each action of the set of actions and based on at least one target location of the training image; andlearning, by the processor, an optimal action-value function approximator of the artificial agent specifying the behavior of the artificial agent to maximize a cumulative future reward value of the reward system, wherein the behavior of the artificial agent is a sequence of actions moving the agent towards the at least one target location of the training image, the sequence of actions comprising at least one scale action. 2. The method of claim 1, further comprising: applying, by the processor, the learned artificial agent on a test image to automatically parse image content for a landmark location, the applying comprising: evaluating the optimal action-value function approximator for a current state space;simultaneously obtaining the optimal action-value function approximator for all possible actions at each current state space; andapplying a reward policy of the optimal action-value function approximator; anddisplaying the identified landmark location in the test image. 3. The method of claim 2, wherein applying the reward policy of the optimal action-value function approximator comprises determining a next action of the artificial agent based on balancing maximization of the cumulative future reward value by actions changing the parametric space and by actions changing the scale space. 4. The method of claim 1, wherein the at least one target location is an anatomical landmark location, and wherein the target location is defined by the position parameters of the anatomical landmark, and wherein a reward value is indicative of a proximity of the state space to the at least one target state. 5. The method of claim 1, wherein learning the optimal action-value function approximator further comprises: generating an experience memory database including a predefined number of last evaluated state spaces for the training image;sampling the experience memory database; andupdating parameters of the optimal action-value function approximator based on the experience memory. 6. The method of claim 1, wherein learning the optimal action-value function approximator further comprises: parameterizing the action-value function approximator using a deep neural network; andoptimizing the action-value function using an... | Siemens Healthcare GmbH | G06K009/62 G06N003/08 G06T007/00 G06K009/20 G06K009/66 G06T007/70 A61B008/08 A61B005/055 A61B005/00 A61B008/00 G06F019/00
US09820660 | B2 | 2017-11-21 | Devices, systems, and methods for assessing a vessel with optimized proximal and distal pressure measurements obtained without the use of a hyperemic agent | 2014-10-20 | Embodiments of the present disclosure are configured to assess the severity of a blockage in a vessel and, in particular, a stenosis in a blood vessel. In some particular embodiments, the devices, systems, and methods of the present disclosure are configured to assess the severity of a stenosis in the coronary arteries without the administration of a hyperemic agent. In some embodiments, the devices, systems, and methods of the present disclosure are configured to optimize proximal and distal pressure measurements utilized to assess the vessel. | 1. A method of evaluating a vessel of a patient, comprising: obtaining proximal pressure measurements associated with blood flow measured at a position proximal of a stenosis of the vessel for a plurality of heartbeat cycles;obtaining distal pressure measurements associated with the blood flow measured at a position distal of the stenosis of the vessel for the plurality of heartbeat cycles;evaluating the obtained proximal and distal pressure measurements to identify any irregular heartbeat cycles within the plurality of heartbeat cycles;optimizing the obtained proximal and distal pressure measurements by removing pressure measurements associated with the irregular heartbeat cycles;calculating a pressure ratio between the optimized distal pressure measurements and the optimized proximal pressure measurements; andoutputting the calculated pressure ratio to a display. 2. The method of claim 1, wherein evaluating the obtained proximal and distal pressure measurements to identify any irregular heartbeat cycles within the plurality of heartbeat cycles includes comparing the obtained proximal and distal pressure measurements to a library of pressure measurements associated with normal heartbeat cycles. 3. The method of claim 1, wherein evaluating the obtained proximal and distal pressure measurements to identify any irregular heartbeat cycles within the plurality of heartbeat cycles includes comparing the obtained proximal and distal pressure measurements for one heartbeat cycle of the plurality of heartbeat cycles to the obtained proximal and distal pressure measurements for other heartbeat cycles of the plurality of heartbeat cycles. 4. The method of claim 1, further comprising: obtaining ECG signals for the plurality of heartbeat cycles. 5. The method of claim 4, further comprising: evaluating the obtained ECG signals to identify any irregular heartbeat cycles within the plurality of heartbeat cycles. 6. The method of claim 5, wherein evaluating the obtained ECG signals to identify any irregular heartbeat cycles within the plurality of heartbeat cycles includes comparing the obtained ECG signals to a library of ECG signals associated with normal heartbeat cycles. 7. The method of claim 5, wherein evaluating the obtained ECG signals to identify any irregular heartbeat cycles within the plurality of heartbeat cycles includes comparing the obtained ECG signals for one heartbeat cycle of the plurality of heartbeat cycles to the obtained ECG signals for other heartbeat cycles of the plurality of heartbeat cycles. 8. The method of claim 1, wherein the proximal pressure measurements are obtained from a pressure sensing catheter. 9. The method of claim 8, wherein the distal pressure measurements are obtained from a pressure sensing guidewire. 10. The method of claim 1, wherein the calculated pressure ratio is output to the display in real time. 11. A system comprising: a processing unit in communication with first and second pressure sensing device... | VOLCANO CORPORATION | A61B005/0215 A61B005/02 A61B005/0452
US09788769 | B2 | 2017-10-17 | Handheld multi-parameter measuring instrument | 2015-01-16 | A handheld multi-parameter measuring instrument includes: a housing having a sidewall and opposed first and second surfaces, the sidewall, the first surface and the second surface cooperatively forming a cavity; a cover attached to the sidewall and rotatably connected to the housing; the blood oxygen transducer includes a light emitting unit and a light receiving unit; the light emitting unit is configured to emit red light and infrared light alternately; the light receiving unit is configured to receive the red light and the infrared light emitted from the light emitting unit; the signal processing unit is received in the cavity of the housing and is configured to control the light emitting unit to emit the red light and the infrared light alternately; the signal processing unit is further configured to calculate oxygen saturation based on a signal received by the light receiving unit; the light emitting unit is located at a side of the cover attached to the housing, and the light receiving unit is located at the sidewall. When measuring blood oxygen saturation, one hand holds the first surface and the second surface of the housing, and one finger of the hand is inserted between the cover and the sidewall of the housing attached to the cover. This handheld multi-parameter measuring instrument is easy to use and has a high measuring accuracy. | 1. A handheld multi-parameter measuring instrument, comprising: a housing having a sidewall and opposed first and second surfaces, the sidewall, the first surface and the second surface cooperatively forming a cavity;a cover attached to the sidewall and rotatably connected to the housing, a finger being capable of being received between the cover and the sidewall of the housing attached to the cover; anda blood oxygen measuring device comprising a blood oxygen transducer and a signal processing unit, the blood oxygen transducer comprising a light emitting unit and a light receiving unit; the light emitting unit being configured to emit red light and infrared light alternately; the light receiving unit being configured to receive the red light and the infrared light emitted from the light emitting unit; wherein the signal processing unit is received in the cavity of the housing and controls the light emitting unit to emit the red light and the infrared light alternately; the signal processing unit is further configured to calculate oxygen saturation based on a signal received by the light receiving unit;an electrocardiogram (ECG) sensor, comprising: a first ECG electrode located on the first surface and adapted to engage a surface of a thumb of a right hand while the right hand is supporting the instrument and the finger is received between the cover and the sidewall, and a second ECG sensor located on the sidewall and adapted to engage a palm of a left hand while the right hand is supporting the instrument;wherein (a) the light emitting unit is located at a side of the cover attached to the housing, and the light receiving unit is located at the sidewall of the housing attached to the cover; or (b) the light emitting unit is located at the sidewall facing the cover, and the light receiving unit is located at a side of the cover facing the housing; or (c) the light emitting unit and the light receiving unit are both located at the sidewall facing the cover;wherein the cover has an open and a closed position, the closed position being adapted to prevent entry of a finger into the instrument, and the open position being adapted to receive the finger such that the forward end of the finger is completely accessible to an environment external to the instrument while the blood oxygen measuring device measures blood oxygen. 2. The handheld multi-parameter measuring instrument according to claim 1, wherein the blood oxygen measuring device further comprises: an optoelectronic interface unit, by which the light emitting unit and the light receiving unit are electrically coupled to the signal processing unit;a memory unit received in the cavity and electrically coupled to the signal processing unit, the memory unit being configured to store the oxygen saturation calculated by the signal processing unit; anda display unit located at the housing, the display unit being electrically coupled to the signal processing unit and configured to display the oxygen sa... | SHENZHEN VIATOM TECHNOLOGY CO., LTD. | A61B005/1455 A61B005/024 A61B005/026 A61B005/00 A61B005/0404
US09832305 | B2 | 2017-11-28 | Configure smartphone based on user sleep status | 2014-12-12 | Various systems and methods for configuring a smartphone based on a user's sleep status are described herein. A compute device includes a determination module to determine a physiological state of a person and a configuration module to configure a quiet mode of the compute device based on the physiological state of the person. | 1. A compute device comprising: a determination module to determine a physiological state of a person; anda configuration module to configure a quiet mode of the compute device based on the physiological state of the person, wherein to configure the quiet mode of the compute device, the configuration module is to: determine a distance from the person to the compute device;configure an audio volume of the compute device as a function of the distance; andconfigure an application executing on the compute device to selectively use an audible notification when the compute device is in the quiet mode. 2. The compute device of claim 1, wherein to determine the physiological state of the person, the determination module is to: use contextual information to strengthen an inference of the physiological state of the person. 3. The compute device of claim 1, wherein to determine the physiological state of the person, the determination module is to: access a sensor to obtain sensor data; anduse the sensor data to determine the physiological state. 4. The compute device of claim 3, wherein the sensor is separate from the compute device. 5. The compute device of claim 3, wherein the sensor is incorporated into the compute device. 6. The compute device of claim 5, wherein the sensor is one of a camera, a heart rate sensor, a skin temperature sensor, or an accelerometer. 7. The compute device of claim 3, wherein the sensor includes a heart rate sensor, and wherein the sensor data includes a heart rate of the person, and wherein to use the sensor data to determine the physiological state, the determination module is to: determine whether the heart rate is slower than a threshold rate; andwhen the heart rate is slower than the threshold rate, declare that the physiological state of the person is resting. 8. The compute device of claim 3, wherein the sensor includes a posture sensor, and wherein the sensor data includes an indication that the person is prone or supine, and to use the sensor data to determine the physiological state, the determination module is to: determine a duration the person has been prone or supine for longer than a threshold period; andwhen the duration is longer than the threshold period, declare that the physiological state of the person is resting. 9. The compute device of claim 3, wherein to access the sensor, the determination module is to: poll the sensor; andobtain the sensor data in response to the polling. 10. The compute device of claim 3, wherein to obtain the sensor data, the determination module is to: listen for sensor data; andreceive the sensor data from the sensor. 11. The compute device of claim 1, wherein to configure the quiet mode of the compute device, the configuration module is to mute a selection of applications executing on the compute device. 12. The compute device of claim 1, wherein the application executing on the compute device is a telephone application, and wherein to selectively use the audible noti... | Intel Corporation | H04M001/00 H04M001/725 G06F001/16 G06F003/01 A61B005/00 A61B005/01 A61B005/024 G06F003/16 H04M003/42 H04M001/02
US09824503 | B2 | 2017-11-21 | System and method for image composition | 2016-12-30 | A system and method for obtaining a composite image by combining multiple sub-images are provided. In some embodiments, the method may include retrieving overlapping images corresponding to sub-images including 3D volume data, generating two-dimensional (2D) projection images and pixel maps based on the overlapping images, performing one or more registrations based on the 2D projection images and the pixel maps, calibrating the sub-images based on the results of the registration(s), and fusing the sub-images to produce a composite image. In some embodiments, the method may include setting a plurality of parameters relating to an X-radiation source or a radiation detector based on a preliminary number of exposures and a preliminary exposure region, controlling, based on at least one of the plurality of parameters, a motion of the X-radiation source or a motion of the radiation detector to capture a plurality of sub-images, and combining the plurality of sub-images. | 1. A method implemented on at least one device each of which has at least one processor and storage, the method comprising: retrieving a first sub-image and a second sub-image; the first sub-image corresponding to a first set of 3D volume data, and the second sub-image corresponding to a second set of 3D volume data;retrieving a first overlapping image from the first sub-image;retrieving a second overlapping image from the second sub-image;generating a first 2D projection image and a first pixel map based on maximum intensity projection of the first overlapping image onto a plane;generating a second 2D projection image and a second pixel map based on maximum intensity projection of the second overlapping image onto the plane;performing 2D registration based on the first 2D projection image, the first pixel map, the second 2D projection image, and the second pixel map;performing 3D registration based on the 2D registration, the first pixel map, or the second pixel map;identifying a correlation between of the first sub-image including the first overlapping image and the second sub-image including the second overlapping image based on the 2D registration or the 3D registration; andfusing the first overlapping image and the second overlapping image based on the correlation to provide a composite image. 2. The method of claim 1, the performing 3D registration comprising calculating a plurality of differences between first locations of maximum intensity of the first overlapping image in the direction perpendicular to the plane and second locations of maximum intensity of the second overlapping image in the direction perpendicular to the plane, each one of the plurality of differences corresponding to a pixel within the plane;comparing the plurality of differences to obtain a probability of the differences; anddesignating an offset in the direction perpendicular to the plane based on the probability. 3. The method of claim 1, the first sub-image or the second sub-image comprising a coronal image, a sagittal image, or a transverse image. 4. The method of claim 1, the plane comprising a coronal plane, a sagittal plane, or a transverse plane. 5. The method of claim 1, retrieving the first sub-image or retrieving the second sub-image comprising using DSA (digital subtraction angiography), CT (computed tomography), CTA (computed tomography angiography), PET (positron emission tomography), X-ray, MRI (magnetic resonance imaging), MRA (magnetic resonance angiography), SPECT (single-photon emission computerized tomography), or US (ultrasound scanning). 6. The method of claim 1, retrieving the first overlapping image or retrieving the second overlapping image comprising using Digital Imaging and Communication in Medicine. 7. The method of claim 1, performing 2D registration comprising calculating an offset in a direction within the plane, and the offset comprising an X offset, a Y offset, a Z offset, a coronal offset, a sagittal offset, or a transverse off... | SHANGHAI UNITED IMAGING HEALTHCARE CO., LTD. | G06K009/00 G06T019/20 G06T005/50 A61B006/00 G06F019/00 G06T007/30 G06T007/32 A61B005/055 A61B006/03 A61B008/08 G06K009/62 H04N005/232 H04N005/32
US09782102 | B2 | 2017-10-10 | Electrocardiography to differentiate acute myocardial infarction from bundle branch block or left ventricular hypertrophy | 2015-03-24 | Acute myocardial infarction (AMI) is diagnosed if: (1) the ECG traces satisfy an ST Elevation Myocardial Infarction (STEMI) criterion and the ECG traces do not indicate the subject has a confounding cardiac condition, or (2) the ECG traces satisfy the STEMI criterion and the ECG traces also indicate the subject has the confounding cardiac condition and a three-dimensional vector cardiograph (3D-VCG) signal generated from the ECG traces includes an ST vector in the ST segment of the 3D-VCG signal and a terminal QRS vector of maximum magnitude in a terminal portion of the QRS complex of the 3D-VCG signal for which the angle between the ST vector and the terminal QRS vector is less than a threshold angle, e.g. in the range [130°, 170°] inclusive. The confounding cardiac condition may be bundle branch block (BBB), left ventricular hypertrophy (LVH), or interventricular conduction delay (IVCD). | 1. A cardiac monitoring system comprising: an electrocardiograph (ECG) monitor configured for operative connection with a plurality of ECG electrodes, the ECG monitor including a display device and an electronic data processing component configured to perform ECG analysis of ECG traces for a plurality of ECG leads acquired by the ECG monitor including the operations of:determining whether the ECG traces satisfy an ST Elevation Myocardial Infarction (STEMI) criterion;analyzing the ECG traces to detect a confounding cardiac condition that is not acute myocardial infarction (AMI) and that produces ST deviation;conditional on the STEMI criterion being satisfied and not detecting the confounding cardiac condition, displaying an AMI alarm on the display device of the ECG monitor;conditional on the STEMI criterion being satisfied and also detecting the confounding cardiac condition, performing the further operations of: converting the ECG traces to a three-dimensional vector cardiograph (3D-VCG) signal;computing an angle between (1) an ST vector characterizing the ST segment of the 3D-VCG signal and (2) a terminal QRS vector characterizing a terminal portion of the QRS complex of the 3D-VCG signal; andconditional on the computed angle being less than a threshold angle, displaying the AMI alarm on the display device of the ECG monitor. 2. The cardiac monitoring system of claim 1 wherein the converting operation comprises: converting the ECG traces to said 3D-VCG signal comprising a vector sum of lead vectors in a three-dimensional space wherein each lead vector has magnitude corresponding to a lead voltage and lies along an anatomical direction corresponding to a spatial orientation of the lead voltage. 3. The cardiac monitoring system of claim 2 wherein the ECG traces for the plurality of ECG leads acquired by the ECG monitor are 12-lead ECG traces and the converting operation further comprises: calculating the lead voltages comprising Frank lead voltages from the 12-lead ECG traces. 4. The cardiac monitoring system of claim 1 wherein the confounding cardiac condition includes one or more of: left bundle branch block (LBBB), right bundle branch block (RBBB), left ventricular hypertrophy (LVH), and Interventricular Conduction Delay (IVCD). 5. The cardiac monitoring system of claim 1 wherein the operation of analyzing the ECG traces to detect the confounding cardiac condition comprises: detecting the confounding cardiac condition if the QRS complex exceeds a threshold duration. 6. The cardiac monitoring system of claim 1 wherein: the operation of analyzing the ECG traces to detect the confounding cardiac condition includes analyzing the ECG traces to detect a bundle branch block (BBB) condition and analyzing the ECG traces to detect a left ventricular hypertrophy (LVH) condition; andthe AMI alarm is displayed on the display device of the ECG monitor conditional on the computed angle being less than (i) a threshold angle ThBBB if BBB is detected or (i... | KONINKLIJKE PHILIPS N.V. | A61B005/04 A61B005/0472 A61B005/00 A61B005/0432 A61B005/0408
US09775543 | B2 | 2017-10-03 | Transcutaneous analyte sensor | 2013-12-30 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A continuous analyte monitoring assembly adapted to assist in determining an analyte level of a fluid, the continuous analyte monitoring assembly comprising: a housing having a bottom, the bottom forming a recess;electronics located within the housing, the electronics being adapted to assist in determining an analyte level of a fluid sample;an implantable sensor comprising a sensor body and at least one electrode disposed on the sensor body, wherein the implantable sensor is adapted to move from a retracted position to an inserted position through the recess, wherein in the retracted position the implantable sensor is at least partially curved and stowed at least partially inside the housing, wherein in the retracted position the at least one electrode resides above a host's skin, wherein in the inserted position the implantable sensor extends at least partially out of the housing, wherein in the inserted position the at least one electrode is adapted to at least partially reside under the host's skin; anda cannula being adapted to assist in placing the implantable sensor. 2. The assembly of claim 1, wherein the electronics is a printed circuit. 3. The assembly of claim 1, further including a removable adhesive liner, the adhesive liner adapted to attach to skin. 4. The assembly of claim 1, wherein the housing includes a disposable housing and a reusable housing, the reusable housing being adapted to contain the electronics. 5. The assembly of claim 4, wherein the disposable housing includes the implantable sensor, the cannula and an adhesive liner. 6. The assembly of claim 1, wherein the implantable sensor is a wire that includes a reagent. 7. The assembly of claim 1, wherein the implantable sensor is an electrochemical sensor. | DexCom, Inc. | A61B005/00 A61B005/145 A61B005/1486 A61B005/1495
US09814552 | B2 | 2017-11-14 | Gum detection using an optical detector in a dental hygiene detection apparatus | 2013-12-20 | A detection apparatus enables detection of a substance that may be present on the surface based on measurement of a stream probe signal correlating a substance at least partially obstructing the passage of fluid through a stream probe includes a distal optical gum detector transmission portion and a distal optical gum detector reception portion in a position to transmit and to receive, respectively, an optical signal controlled by a controller, enables the controller to determine if an open port of a distal tip of the stream probe is in contact with a substance at least partially obstructing the passage of fluid through the open port and not in contact with the gums of a subject or of a user of the detection apparatus to override false positive signals. | 1. A dental hygiene detection apparatus for detecting the presence of a substance on a surface, the detection apparatus comprising: a distal oral insertion portion defining a proximal end and a distal end, the distal oral insertion portion including:a stream probe having distal probe portion, the distal probe portion of the stream probe configured to be immersed in a first fluid, the distal probe portion defining a distal tip having an open port to enable passage of a gas therethrough, the open port of the distal tip having a size and shape configured to detect a substance that may be present on a surface,a distal optical gum detector transmission portion defining a proximal end and a distal tip, the distal tip of the distal optical gum detector transmission portion extending to the vicinity of the distal end of the distal oral insertion portion; anda distal optical gum detector reception portion defining a proximal end and a distal tip, the distal optical gum detector reception portion extending to the vicinity of the distal end of the distal oral insertion portion,the apparatus configured such that passage of the gas through the distal tip of the distal probe portion of the stream probe enables detection of a substance that may be present on the surface based on measurement of a stream probe signal correlating a substance at least partially obstructing the passage of fluid through the open port of the distal tip of the distal probe portion of the stream probe, and configured such that the distal optical gum detector transmission portion and the distal optical gum detector reception portion are in a position to transmit and to receive, respectively, an optical signal that upon transmission of the optical signal and reception of the optical signal by a controller, the controller configured to determine, based on the stream probe signal and the optical signal, if the open port of the distal tip of the distal probe portion of the stream probe is in contact with a substance at least partially obstructing the passage of fluid through the open port and not in contact with the gums,wherein the optical signal is obtained using at least two wavelengths, with a first wavelength below 600 nm and a second wavelength above 600 nm. 2. The dental hygiene detection apparatus according to claim 1, wherein the distal optical gum detector transmission portion includes a first distal transmitting optical fiber defining a proximal end and a distal tip extending to the vicinity of the distal end of the distal oral insertion portion and wherein the distal optical gum detector transmission portion further includes a second distal transmitting optical fiber defining a proximal end and a distal tip, the distal tip of the second distal transmitting optical fiber extending to the vicinity of the distal end of the distal oral insertion portion. 3. The dental hygiene detection apparatus according to claim 1, wherein the distal optical detector transmitting portion comprise... | KONINKLIJKE PHILIPS N.V. | A61C003/00 A61C001/00 A61C019/04 A61C017/22 G01N007/00 A61B005/00 A46B015/00
US09795307 | B2 | 2017-10-24 | Pressure sensing guidewires | 2015-12-04 | Medical devices and methods for making and using medical devices are disclosed. An example medical device may include a medical device for measuring blood pressure. The medical device may include an elongated shaft having a proximal region and a distal region. An optical fiber may extend along the proximal region. The optical fiber may be secured to an inner surface of the shaft. An optical pressure sensor may be coupled to the optical fiber. The optical pressure sensor may be disposed along the distal region. A sealing member may be attached to the optical fiber and may have a surface engaged with the inner surface of the shaft. | 1. A medical device for measuring blood pressure, comprising: an elongated shaft having a proximal region and a distal region;an optical fiber extending along the proximal region;wherein the optical fiber is secured to an inner surface of the shaft;an optical pressure sensor coupled to the optical fiber, the optical pressure sensor being disposed along the distal region;a centering member coupled to the optical fiber and positioned adjacent to the optical pressure sensor; anda sealing member attached to the centering member and extending proximally therefrom, the sealing member having a surface engaged with the inner surface of the shaft. 2. The medical device of claim 1, wherein the sealing member is secured to a proximal end of the centering member. 3. The medical device of claim 1, wherein the sealing member includes a conical portion. 4. The medical device of claim 3, wherein the conical portion is wedged against the inner surface of the shaft. 5. The medical device of claim 1, wherein the sealing member includes a cylindrical section. 6. The medical device of claim 5, wherein the cylindrical section contacts the inner surface of the shaft. 7. The medical device of claim 1, wherein the shaft includes an outer coating. 8. The medical device of claim 7, wherein the sealing member is designed to limit distal migration of the coating within the shaft. 9. The medical device of claim 1, wherein the shaft has a plurality of slots formed therein. 10. The medical device of claim 1, wherein the distal region of the shaft has a distal inner diameter, wherein the proximal region of the shaft has a proximal inner diameter, and wherein the distal inner diameter is larger than the proximal inner diameter. 11. The medical device of claim 10, wherein the sealing member is positioned at a transition region between the proximal region and the distal region. 12. The medical device of claim 1, wherein the medical device is a guidewire for measuring fractional flow reserve. 13. A medical device for measuring blood pressure, comprising: a tubular member having a proximal region, a distal region, and a sensor housing region;a coating disposed along an outer surface of the tubular member;wherein the tubular member has a plurality of slots formed therein;an optical fiber extending along the proximal region;an optical pressure sensor coupled to the optical fiber, the optical pressure sensor being disposed within the sensor housing region; anda sealing member attached to the optical fiber, the sealing member being disposed within the sensor housing region and having a surface extending into and engaged with the inner surface of the tubular member along the distal region so as to limit distal migration of the coating within the tubular member. 14. The medical device of claim 13, wherein the sealing member includes a conical portion that is wedged against the inner surface of the tubular member. 15. The medical device of claim 13, wherein the sealing member ... | BOSTON SCIENTIFIC SCIMED, INC. | A61B005/0215 A61B005/026 A61B005/00 A61B090/30 A61B005/02
US09814378 | B2 | 2017-11-14 | Full spectrum LED illuminator having a mechanical enclosure and heatsink | 2016-08-25 | An apparatus for providing a light output to an optical guide for illumination of an imaged object including a plurality of solid state light-emitting sources each of which are independently powered and independently controlled, each light-emitting source emitting light at a wavelength which is different from the wavelength emitted by the other light-emitting sources. The apparatus also includes a heat sink configured to thermally couple the plurality of solid state light-emitting sources and provide conduction of heat generated by the plurality of solid state light-emitting sources. The apparatus further includes an optical elements to collect, collimate, and combine the emissions from the plurality of solid state light-emitting sources into a combined beam of light to be optically coupled to the light guide. | 1. An apparatus for providing a light output to an optical guide for illumination of an object to be imaged, the apparatus comprising: a plurality of solid state light-emitting sources each of which are independently powered and independently controlled, each light-emitting source emitting light at a wavelength that is different from a wavelength emitted by the other light-emitting sources;a heat sink configured to thermally couple the plurality of solid state light-emitting sources and provide conduction of heat generated by the plurality of solid state light-emitting sources, wherein the heat sink comprises a heat spreader plate having a planar surface and each of the solid-state light-emitting sources is mounted to the planar surface to orient each of the solid-state light-emitting sources along a common optical plane;optical elements configured to collect, collimate, and combine the emissions from the plurality of solid state light-emitting sources into a combined beam of light to be optically coupled to a light guide at an output of the apparatus; anda mechanical enclosure comprising a linear array of input ports matching a linear pattern of the light-emitting sources on the planar surface, wherein at least some of the optical elements are not directly attached to the light-emitting sources and the at least some of the optical elements not directly attached to the light-emitting sources are mounted in the mechanical enclosure and the heat spreader plate is configured to be assembled to the enclosure so as to seal the enclosure,wherein light emitted from each of the light-emitting sources travels an optical path length from the respective light-emitting source to the output, the optical path lengths from the light-emitting sources to the output varying based on the wavelength of the light emitted from the respective light-emitting source. 2. The apparatus of claim 1, wherein the heat sink comprises a passive cooling system. 3. The apparatus of claim 2, wherein the passive cooling system is a finned heat sink or a heat pipe. 4. The apparatus of claim 1, wherein the heat sink comprises an active cooling system. 5. The apparatus of claim 4, wherein the active cooling system is a thermoelectric cooler or a liquid cooler. 6. The apparatus of claim 1, wherein the solid state light-emitting sources comprise light emitting diodes and diode lasers. 7. The apparatus of claim 1, wherein the optical elements comprise a field lens and an aspheric lens configured to collect and collimate the emission from each of the plurality of solid state light-emitting sources. 8. The apparatus of claim 1, further comprising a dichroic filter configured to couple the collimated emission from each of the plurality of solid state light-emitting sources into the combined beam of light directed along a common path to an output port. 9. The apparatus of claim 1, wherein the optical elements are arranged such that the optical path length of each of the plurality of s... | Novadaq Technologies Inc. | F21V013/02 F21V009/00 A61B001/12 A61B001/06 F21V029/70 F21V005/04 F21V009/08 F21K009/60 F21V029/89 A61B005/00 F21V015/01 F21V023/02 F21V008/00 G02B023/24 F21Y101/00 F21Y115/30 F21Y115/10 F21Y113/13 A61B090/30 F21Y103/10
US09776005 | B2 | 2017-10-03 | Transvascular diaphragm pacing systems and methods of use | 2015-08-28 | Transvascular diaphragm pacing systems (TDPS) and methods are disclosed for providing respiratory therapy to a patient. The TDPS can provide rapid insertion and deployment of endovascular pacing electrodes in critically ill patients who require intubation and invasive PPMV in order to support the physiological requirements of the human ventilatory system. The systems and methods make best use of the contractile properties of the diaphragm muscle and prevent muscle disuse and muscle atrophy. This can be carried out by engaging the phrenic nerves using patterned functional electrical stimulation applied to endovascular electrodes that are temporarily and reversibly inserted in central veins of the patient, such as the left subclavian vein and the superior vena cava. The TDPS can be designed to seamlessly interface with any commercially available positive-pressure ventilatory assistance/support equipment such as is commonly in use in hospital intensive care units (ICU) for treating critically ill patients with breathing insufficiencies, pain, trauma, sepsis or neurological diseases or deficits. | 1. A method of stimulating a diaphragm to provide a diaphragm assist level, comprising: (a) obtaining an attribute of a breath cycle of a patient, wherein the patient receives respiratory assistance from a ventilator during the breath cycle;(b) based on a stimulation parameter, the attribute of the breath cycle, and a desired value of a breath parameter, transmitting a first stimulation signal to a first electrode positioned within a blood vessel of the patient such that the electrode recruits a phrenic nerve of the patient;(c) after transmitting the first stimulation signal, determining an actual value of the breath parameter from a breath;(d) comparing the actual value of the breath parameter to the desired value of the breath parameter; and(e) either (1) maintaining the stimulation parameter if the actual value of the breath parameter is within a selected range of the desired value of the breath parameter; or (2) modifying the stimulation parameter if the actual value of the breath parameter is not within the selected range of the desired value of the breath parameter; and(f) repeating steps (b)-(e),wherein the method further includes transmitting a second stimulation signal to a second electrode to recruit the phrenic nerve, and wherein transmitting the first stimulation signal to the first electrode occurs at a different time as transmitting the second stimulation signal to the second electrode. 2. The method of claim 1, wherein the first stimulation signal includes at least one pulse, and the stimulation parameter includes at least one of a rate, a duration, a pulse width, a frequency, or an amplitude. 3. The method of claim 1, wherein the blood vessel is at least one of a left jugular vein or a left subclavian vein, the phrenic nerve is a left phrenic nerve, and the method further includes transmitting a stimulation signal to a third electrode positioned in a superior vena cava such that the third electrode recruits a right phrenic nerve of the patient. 4. The method of claim 1, wherein the blood vessel is at least one of a left jugular vein or a left subclavian vein, the phrenic nerve is a left phrenic nerve, and the method further includes transmitting stimulation signals to a third electrode and a fourth electrode positioned in the superior vena cava such that the third and fourth electrodes recruit a right phrenic nerve of the patient. 5. The method of claim 4, wherein transmitting the stimulation signal to the third electrode occurs at a different time as transmitting the stimulation signal to the fourth electrode. 6. The method of claim 1, wherein each of (a) the step of obtaining the attribute of the breath cycle of the patient, and (b) the step of determining the actual value of the breath parameter from the prior breath includes detecting with a sensor at least one of airflow, volume, pressure, motion, or electrical signals. 7. The method of claim 1, wherein the step of comparing the actual value of the breath parameter to t... | Lungpacer Medical Inc. | A61N001/36 A61M016/00 A61N001/05 A61B005/08
US09795357 | B2 | 2017-10-24 | Positioning distance control for X-ray imaging systems | 2012-11-07 | The present invention relates to providing guiding information for operating an X-ray imaging system. In order to provide an improved positioning distance control facilitating the work flow, an X-ray imaging system (10) is provided, comprising a moveable X-ray source (12) and/or a moveable X-ray detector (14). The system further comprises a plurality of object-surface detecting sensors (16), a positioning detection arrangement (18), a processing unit (20), and a display unit (22). The sensors are arranged such to detect object data of objects located between the X-ray source and the X-ray detector. The positioning detection arrangement is provided to detect the current position of the X-ray source and/or the X-ray detector and the position of the sensors. The processing unit is configured to compute a situation-map (26) of the current spatial situation between the X-ray source and the X-ray detector based on the object data provided by the sensors and the current position, the situation-map distinguishing at least between empty spaces and spaces occupied by rigid objects. The situation-map comprises a representation (25) of the X-ray source and the X-ray detector in relation to the spatial situation. Further, the display unit is configured to display the situation-map to a user operating the X-ray imaging system. | 1. An X-ray imaging system, comprising: a moveable X-ray source and a moveable X-ray detector;a plurality of object-surface detecting sensors;a positioning detection arrangement;a processing unit; anda display unit;wherein the sensors are arranged such to detect object data of objects located between the X-ray source and the X-ray detector;wherein the positioning detection arrangement is provided to detect the current position of the X-ray source and the X-ray detector and the position of the sensors;wherein the processing unit is configured to compute a situation-map of the current spatial situation of a positional arrangement exclusively between the X-ray source and the X-ray detector based on the object data provided by the sensors and the current position, the situation-map distinguishing at least between empty spaces and spaces occupied by rigid objects; and wherein the situation-map comprises a representation of the X-ray source and the X-ray detector in relation to the positioning detection arrangement;wherein the display unit is configured to display the situation-map to a user operating the X-ray imaging system;wherein the X-ray imaging system is a C-arm imaging system comprising a C-arm structure, wherein the X-ray source and the X-ray detector are mounted to opposing ends of a C-arm;wherein the sensors are mounted to the C-arm structure in a distributed manner along an inner side of the arc; andwherein the positioning detection arrangement comprises angulation detectors for providing information about the movement of the C-arm. 2. The X-ray imaging system according to claim 1, wherein the object-surface detecting sensors are provided as ultrasound sensors. 3. The X-ray imaging system according to claim 1, wherein the sensors provide surface data of a number of spatial segments; and wherein the processing unit is configured to combine the surface data of the spatial segments to form a topogram of the current spatial situation. 4. The X-ray imaging system according to one of the preceding claims, wherein the situation-map comprises a visual representation of the X-ray source and the X-ray detector in their current position. 5. The X-ray imaging system according to claim 1, wherein the current spatial situation in the situation-map is provided in relation to a spatial frame of reference. 6. The X-ray imaging system according to claim 1, wherein the processing unit is configured to indicate, in the situation-map, a geometrical form of an X-ray beam to be radiated from the X-ray source to the X-ray detector; wherein the situation-map is a positioning guide for the operator of the X-ray imaging system. 7. The X-ray imaging system according to claim 1, wherein the processing unit is configured to compute a simulation of at least one movement of the X-ray source and the X-ray detector; and wherein the display unit is configured to display at least one resulting simulated spatial situation. 8. The X-ray imaging system according to claim ... | Koninklijke Philips N.V. | A61B006/00 A61B006/10 A61B005/11
US09782131 | B2 | 2017-10-10 | Method and system for self-managed sound enhancement | 2016-10-17 | A system and method are provided for capturing hearing characteristics from self-administered hearing tests, including loudness tolerance levels at different sound frequencies, as an individualized shaped auditory profile for automatically enhancing audio to complement and address an individual's hearing deficits experienced via a particular signal pathway including the audio device used to administer the test, the sound environment, and the user's hearing capabilities. The user may self-administer a hearing test on a smartphone or other device. The shaped auditory profile may be used to produce enhanced sound subsequently transmitted to the individual via the same signal pathway. In some aspects, the invention is useful for any individual seeking an enhanced hearing experience, whether having hearing within normal range or hearing that is impaired. Manual gain adjustments applied by a user to the enhanced sound thereby produced may result in changes to a stored shaped auditory profile or the generation of a new profile. | 1. A method for providing enhanced audio to an individual, tailored to audiological characteristics of the individual, comprising: capturing the individual's audio hearing characteristics, including determining a threshold uncomfortable loudness level of the user for at least one frequency and the individual's minimum audible level for at least one audiometric frequency, at least in part by a self-administered test;analyzing the individual's hearing characteristics to generate an individualized audiological profile;storing the individualized audiological profile in a storage means;using the individualized audiological profile to determine a first frequency-specific gain adjustment for an input audio signal to compensate for frequency-dependent loudness deficits in the individual's hearing;determining a modified audio signal produced by applying the first gain adjustment to an input signal;providing the modified audio signal to the individual using a personal electronic device that reproduces sound;applying a second frequency-specific gain adjustment, input manually by the individual after the modified audio signal is provided to the individual, to convert the modified audio signal to an enhanced audio signal; andproviding the enhanced audio signal to the individual using the personal electronic device. 2. The method of claim 1, further comprising generating an environment profile including at least an environmental sound measurement;using the environment profile and the individualized audiological profile to determine the first frequency-specific gain adjustment;providing the modified audio signal to the individual using the personal electronic device in the profiled environment; andproviding the enhanced audio signal to the individual using the personal electronic device in the profiled environment. 3. The method of claim 2, further comprising periodically updating the environment profile during provision of the enhanced audio to the individual. 4. The method of claim 1, further comprising said capturing the individual's audio hearing characteristics comprising capturing the audio hearing characteristics of the individual for a test signal pathway including at least one specific component selected from a circuit of the personal electronic device, a transducer delivering the audio from the personal electronic device to the individual's ear, whether the transducer fits onto or into the individual's ear, the anatomy of the user's ear canal, and an environmental sound measurement;said individualized audiological profile being a shaped auditory profile;providing the modified audio signal to the individual using the personal electronic device through a listening signal pathway similar to the test signal pathway; andproviding the enhanced audio signal to the individual using the personal electronic device through a listening signal pathway similar to the test signal pathway. 5. The method of claim 1, further comprising generating a customized indiv... | ACE Communications Limited (HK) | H04R025/00 A61B005/00 A61B005/12 H04R005/04
US09781965 | B2 | 2017-10-10 | Wireless monitoring of safety helmets | 2016-07-21 | Remote monitoring of a subject wearing a sports helmet is enabled. In one aspect, an example system includes a safety helmet and a sensor integrated with the helmet for continuously gathering head acceleration force data, the head acceleration force data associated with the head movements of a subject. The system can also include a wireless transceiver coupled to the sensor for transmitting the head acceleration force data and a mobile device for receiving the head acceleration force data from the wireless transceiver. The system can further include a database engine for displaying the head acceleration force data to a user. | 1. A system, comprising: a processor; anda memory that stores executable instructions that, when executed by the processor, facilitate performance of operations, comprising: receiving, from a wireless transceiver of a helmet corresponding to a head of a person, first pressure data representing a first pressure that has been applied to the head during a first type of activity determined to be performed by the person; anddisplaying the first pressure data representing the first pressure that has been applied to the head during the first type of activity with second pressure data representing a second pressure that has been applied to the head during a second type of activity determined to be performed by the person that is different from the first type of activity. 2. The system of claim 1, wherein the operations further comprise: determining whether head acceleration force data associated with movements of the head during the first type of activity satisfies a defined condition related to a head injury. 3. The system of claim 2, wherein the operations further comprise: in response to determining that the head acceleration force data satisfies the defined condition, sending the first pressure data directed to a device. 4. The system of claim 1, wherein the operations further comprise: correlating head acceleration force data associated with movements of the head with the first type of activity and the second type of activity. 5. The system of claim 4, wherein the operations further comprises: in response to the correlating, displaying respective axes of movement of respective uses associated with the first type of activity and the second type of activity. 6. The system of claim 1, wherein the operations further comprise: determining an acceleration of the head; anddetermining a period of time since the acceleration. 7. The system of claim 6, wherein the operations further comprise: displaying a symbol corresponding to the period of time. 8. The system of claim 1, further comprising an array of sensors, wherein the receiving further comprises: obtaining the first pressure data via the wireless transceiver from the array of sensors. 9. The system of claim 8, wherein the array of sensors comprises a microelectromechanical system accelerometer. 10. The system of claim 1, wherein the receiving comprises: receiving the first pressure data from the wireless transceiver utilizing a wireless communication protocol based on an Institute for Electronics Engineers 802.15.4 physical radio specification. 11. A method, comprising: receiving, by a system comprising a processor, pressure data from a transceiver coupled to a pressure sensitive sensor of a helmet, wherein the pressure data represents a first pressure applied to a portion of a head engaged in a first type of activity; anddisplaying the first pressure data representing the first pressure applied to the portion of the head with second pressure data representing a second pressure applied to the... | AT&T INTELLECTUAL PROPERTY I, L.P. | G08B023/00 A42B003/04 G08B021/04 A61B005/00 A61B005/11
US09782133 | B2 | 2017-10-10 | Respiratory apparatus with improved flow-flattening detection | 2014-03-14 | In a respiratory apparatus for treatment of sleep apnea and other disorders associated with an obstruction of a patient's airway and which uses an airflow signal, an obstruction index is generated which detects the flattening of the inspiratory portion of the airflow. The obstruction index is used to differentiate normal and obstructed breathing. The obstruction index is based upon different weighting factors applied to sections of the airflow signal thereby improving sensitivity to various types of respiration obstructions. | 1. A method for determining the presence of a respiratory obstruction with a processor, the method comprising: accessing flow data samples from a signal generated by a sensor configured to detect flow, the flow data samples representing breath data; andgenerating an obstruction signal in an obstruction detector of the processor when at least one breathing cycle in the flow data samples has, during an inspiratory portion, two high positive lobes and a relatively flat zone between the lobes. 2. The method of claim 1, wherein at least some of the data samples are successive. 3. The method of claim 1, wherein at least some of the data samples are in regular intervals. 4. A method for determining a response for a respiratory obstruction with a processor, the method comprising: accessing data samples from a signal generated by a sensor configured to detect flow, the data samples representing breath data;detecting at least one breathing cycle in the data samples that has, during an inspiratory portion, two high positive lobes and a relatively flat zone between the lobes,calculating an obstruction index by assigning different weighting factors to inspiration flow samples depending on a magnitude of each sample with respect to a mean inspiration flow and a time-wise position of each sample with respect to a mid-inspiration point; anddetermining a command signal for activating or altering an operation of a respiratory apparatus, based on the calculated obstruction index. 5. An apparatus for determining the presence of a respiratory obstruction, the apparatus comprising: a processor configured to:access flow data samples from a signal generated by a sensor configured to detect flow, the flow data samples representing breath data; andgenerate an obstruction signal in an obstruction detector of the processor when at least one breathing cycle in the flow data samples has, during an inspiratory portion, two high positive lobes and a relatively flat zone between the lobes. 6. The apparatus of claim 5 wherein at least some of the data samples are successive. 7. The apparatus of claim 5 wherein at least some of the data samples are in regular intervals. 8. An apparatus for determining a response for a respiratory obstruction, the apparatus comprising: a processor configured to:access data samples from a signal generated by a sensor configured to detect flow, the data samples representing breath data;detect at least one breathing cycle in the data samples that has, during an inspiratory portion, two high positive lobes and a relatively flat zone between the lobes,calculate an obstruction index by assigning different weighting factors to inspiration flow samples depending on a magnitude of each sample with respect to a mean inspiration flow and a time-wise position of each sample with respect to a mid-inspiration point; anddetermine a command signal for activating or altering an operation of a respiratory apparatus, based on the calculated obstruction index. | ResMed Limited | A61B005/08 A61M016/00 A61B005/00 A61B005/085 A61B005/087 A61M016/06
US09808156 | B2 | 2017-11-07 | Systems, computer medium and computer-implemented methods for monitoring and improving biomechanical health of employees | 2014-02-14 | Provided are embodiments of systems, computer medium and computer-implemented methods for monitoring the health of an employee while the employee is located in an employee workstation. A method including collecting, from a plurality of biomechanical sensors disposed throughout the employee workstation, biomechanical health data indicative of biomechanical characteristics of the employee, determining an employee biomechanical health profile based at least in part on the collected biomechanical health data. The employee biomechanical health profile including biomechanical health characteristics, biomechanical health conditions and/or biomechanical health risks for the employee. The method including generating a health plan for the employee based at least in part on the biomechanical health profile, and displaying the employee biomechanical health profile and the health plan to the employee. | 1. A system for monitoring employee health, the system comprising: an employee workstation comprising: an employee computer comprising a display;a computer mouse;a chair disposed on a floor of the workstation, the chair comprising a seat portion and a back portion configured to support an employee seated in the chair;a floor mat disposed on the floor of the workstation and configured to support feet of the employee seated in the chair; anda set of biomechanical health sensors configured to collect health data via a plurality of points of contact with the employee seated in the chair and to output biomechanical sensor data corresponding to the health data collected,the set of biomechanical health sensors in communication with a communications network and comprising temperature transducers, force transducers and a camera,the biomechanical sensor data comprising temperature data output by the temperature transducers that is indicative of a body position of the employee, force data output by the force transducers that is indicative of the body position of the employee, and image data output by the camera comprising three-dimensional images indicative of the body position and eye movement of the employee;the floor mat comprising floor mat biomechanical sensors for detecting biomechanical characteristics of the employee via contact with feet of the employee defining a first point of contact with the employee, the floor mat biomechanical sensors comprising: a floor mat temperature sensor comprising one of the temperature transducers configured to sense body temperature of the employee; anda floor mat force sensor comprising one of the force transducers configured to sense a force applied to the floor mat,the chair comprising chair biomechanical sensors for detecting biomechanical characteristics of the employee via contact with portions of the employee seated in the chair defining a second point of contact with the employee, the chair biomechanical sensors comprising: a chair temperature sensor comprising one of the temperature transducers configured to sense body temperature of the employee; andchair force sensors comprising: a chair seat force transducer located in the seat portion of the chair and configured to sense a force applied to the seat portion of the chair; anda chair back force transducer located in the back portion of the chair and configured to sense a force applied to the back portion of the chair, andthe camera positioned to capture images of the employee seated in the chair defining a third point of contact with the with the employee, the images of the employee seated in the chair comprising the three-dimensional images indicative of the body position and eye movement of the employee;a database in communication with the communications network and storing employee health information associated with one or more employees; anda computer server in communication with the communication network, the computer server configured to: conduct a he... | Saudi Arabian Oil Company | A61B005/00 A61B005/01 A61B005/11 A61B005/021 G06F019/00 A61B005/0205 A61B005/053 A61B005/08 A61B005/145
US09791417 | B2 | 2017-10-17 | Acoustic wave detection probe and photoacoustic measurement apparatus provided with the same | 2014-08-12 | An acoustic wave detection probe includes a light projection section that emits measuring light to be projected onto a subject, an acoustic wave transducer disposed adjacent to the light projection section and capable of detecting an acoustic wave, an acoustic lens provided on a detection side of the acoustic wave transducer, and a housing accommodating the light projection section, the acoustic wave transducer, and the acoustic lens, in which the acoustic lens and a surface portion of the housing adjacent to the acoustic lens are formed to have an optical property in which the average diffuse reflection factor is 85% or more and the average absorption factor is 10% or less in a wavelength range of the measuring light. | 1. An acoustic wave detection probe, comprising: a light projection section that emits measuring light to be projected onto a subject, wherein a wavelength range of the measuring light is a near infrared wavelength range;an acoustic wave transducer disposed adjacent to the light projection section and capable of detecting an acoustic wave;an acoustic lens provided on a detection side of the acoustic wave transducer; anda housing accommodating the light projection section, the acoustic wave transducer, and the acoustic lens,wherein the acoustic lens and a surface portion of the housing adjacent to the acoustic lens have an optical property in which the average diffuse reflection factor is 85% or more and the average absorption factor is 10% or less in the wavelength range of the measuring light,wherein the optical property of the acoustic lens and/or the surface portion of the housing further includes an optical property in which the average diffuse reflection factor in a blue to green wavelength range is 70% or less. 2. The probe as claimed in claim 1, wherein the optical property of the surface portion of the housing is given by a coating material containing a first inorganic pigment. 3. The probe as claimed in claim 2, wherein the first inorganic pigment is at least one type of oxide particles selected from the group consisting of titanium oxide, zirconium oxide, ferric oxide, and cerium oxide. 4. The probe as claimed in claim 2, wherein the particle size of the first inorganic pigment is 0.05 to 0.35 μm. 5. The probe as claimed in claim 2, wherein the additive amount of the first inorganic pigment is 2 to 65 wt %. 6. The probe as claimed in claim 1, wherein the optical property of the surface portion of the housing is given by a diffuse reflection sheet. 7. The probe as claimed in claim 1, wherein the optical property of the surface portion of the housing is given by forming the surface portion with a high reflective material. 8. The probe as claimed in claim 7, wherein the high reflective material is a resin which includes an inorganic pigment and at least one type of resin selected from the group consisting of polyester, polyethylene, polycarbonate, polytetrafluoroethylene (PTFE), perfluoroalkoxy fluororesin (PFA), tetrafluoroethylene-hexafluoropropylene copolymer (FEP), ethylene-tetrafluoroethylene copolymer (ETFE), and ethylene-chlorotrifluoroethylene copolymer (ECTFE). 9. The probe as claimed in claim 1, wherein the optical property of the acoustic lens is given by forming the acoustic lens with a material containing a second inorganic pigment. 10. The probe as claimed in claim 9, wherein the second inorganic pigment is at least one type of oxide particles selected from the group consisting of titanium oxide, zirconium oxide, ferric oxide, and cerium oxide. 11. The probe as claimed in claim 9, wherein the particle size of the second inorganic pigment is 0.05 to 0.35 μm. 12. The probe as claimed in claim 9, wherein the additive ... | FUJIFILM Corporation | G01N029/24 A61B008/00 G01N029/22 A61B005/00 G01N029/04 B06B001/06
US09820658 | B2 | 2017-11-21 | Systems and methods for providing interoperability among healthcare devices | 2006-08-30 | A medical system includes a network; one or more medical data collection appliances coupled to the network, each appliance transmitting data conforming to an interoperable format; and a server coupled to the network to store data for each individual in accordance with the interoperable format. | 1. An interoperable health-care system, comprising: a network;a first medical appliance including a first sensor to capture a first vital information including weight and coupled to the network, the first medical appliance transmitting the first vital information conforming to an interoperable format to a base station; wherein the interoperable format is used by a plurality of vendors, anda second medical appliance including a second sensor to capture a second vital information including blood pressure and coupled to the network, each of the first and second medical appliances converting the vital information in accordance with the interoperable format and processing the first and second vital information,wherein the interoperable format enables different appliances monitoring medical conditions to share medical data, wherein the interoperable format includes a protocol for the first and second medical appliances to locate each other's data and detect patient health based on a combination of the different first and second medical condition data captured through the medical appliances, wherein the network translates data into an interoperable format accessible by a third medical appliance to receive at least patient weight and blood pressure generated by the first and second medical appliances over the network. 2. The system of claim 1, wherein the first medical condition is represented by weight data and the second medical condition is represented by blood pressure data and wherein the first or second appliances communicate data conforming to the interoperable format over a wired or wireless connection. 3. The system of claim 2, wherein the first or second appliances communicates over two or more protocols selected from a group consisting of: cellular protocol, ZigBee protocol, Bluetooth protocol, WiFi protocol, WiMAX protocol, USB protocol, ultrawideband protocol. 4. The system of claim 2, wherein the first medical appliance transmits the first vital information over a first protocol to a computer, wherein the computer transmits the first vital information to the second medical appliance over a second protocol. 5. The system of claim 1, wherein the first or second appliance comprises a statistical analyzer to process the vital information, wherein the statistical analyzer is selected from a group consisting of rule-based inference, fuzzy logic, if-then logic, Hidden Markov Models (HMMs), dynamic time warping, and template matching. 6. The system of claim 1, wherein the protocol supports an automatic configuration of the second appliance with the base station when the second appliance announces its presence to the base station and the base station queries the device for its identity, wherein the second device requests registration with base station and subsequently requests application data from base station, wherein the base station searches for a responsive device from a registration list and forwards the request to a matching device and u... | - | G06Q010/00 G06Q050/00 A61B005/0205 A61B005/00 A61B005/0404 A61B005/0492 G06F019/00 G06Q050/22 A61B005/021 A61B005/024 A61B005/0476 A61B005/0496 A61B005/053 A61B005/08 A61B005/11 A61B005/113 A61B005/145 A61B007/00 A61B008/06 A61B008/00 A61B008/08
US09820688 | B2 | 2017-11-21 | Sinus illumination lightwire device | 2008-05-19 | An illuminating wire medical device may include an elongate flexible housing and an illuminating fiber and a core wire extending through at least part of the housing, the core wire providing desired pushability and torquability. The illuminating device may further include a connector assembly cooperating with the illuminating fiber to accommodate changes in length and to absorb forces applied during use of the device. | 1. An illuminating guidewire device for use at a site within a patient's body, the guidewire device comprising: an elongate flexible housing having a proximal end, a distal end, and a lumen extending through the housing from the proximal end to the distal end, wherein the distal end is configured to be passed into a nostril of the patient and through an opening into a paranasal sinus of the patient, and wherein the elongate flexible housing is configured to have a length sufficient to extend from the distal end disposed in the paranasal sinus through the nostril to the proximal end disposed outside of the patient;at least one illuminating fiber extending through the lumen of the housing;a connector having a channel, wherein the connector is sized and configured to slidably receive the elongate flexible housing within the channel such that the connector is coupled with the proximal end of the elongate housing and a proximal end of the at least one illuminating fiber, the connector including a sliding member attached to the elongate housing so that the proximal end of the housing can move longitudinally within the connector; anda core wire disposed in the housing, wherein the core wire longitudinally terminates at a proximal end and at a distal end;wherein the proximal end of the at least one illuminating fiber is fixedly attached to the connector and extends into the proximal end of the housing such that the housing can slide over the at least one illuminating fiber as the housing moves longitudinally within the connector and the at least one illuminating fiber is stationary with respect to the connector; andwherein the core wire is directly attached to the housing at two attachment points, wherein one of the attachment points is located at or near the housing distal end and the other attachment point is located at or near a junction between the proximal and distal ends of the housing, wherein the distal end of the core wire is secured at the attachment point at or near the housing distal end and the proximal end of the core wire is secured at the attachment point located at or near the junction. 2. The device of claim 1, wherein elongate housing comprises: a proximal portion; anda distal portion that is more flexible than the proximal portion. 3. The device of claim 2, wherein the proximal portion comprises a cable tube and the distal portion comprises a coil, and wherein a distal end of the cable tube is attached to a proximal end of the coil. 4. The device of claim 2, wherein the at least one illuminating fiber comprises at least two illuminating fibers. 5. The device of claim 1, wherein the core wire extends proximally from the junction but does not extend all the way to the proximal end of the housing, and wherein a proximal end of the core wire is not attached to the housing. 6. The device of claim 1, wherein the sliding member comprises a cylindrical member configured to rotate within the connector, thus allowing the housing to rotate ... | Acclarent, Inc. | A61B005/00 A61M025/00 A61B001/00 A61B017/24 A61B001/233 A61M025/01
US09782075 | B2 | 2017-10-10 | Electronic delivery of information in personalized medicine | 2013-03-15 | According to an aspect of an embodiment, a method of delivering information-enabled personalized healthcare in a clinical, non-research setting may include capturing one or more data streams where each of the data streams relates to health care of a patient. The method may further include integrating the data streams to generate integrated diagnostic data and analyzing the integrated diagnostic data to generate analyzed diagnostic data. The method may further include curating the analyzed diagnostic data and generating an integrated report for presentation to a physician of the patient based on the curated analyzed diagnostic data. | 1. A system for electronic delivery of information in personalized health care, the system comprising: one or more servers configured to host a healthcare social network particular to providing personalized health care for an individual patient, the healthcare social network including one or more private portions and a public portion, the one or more servers configured to: accept the individual patient as a new patient in the healthcare social network, the individual patient being associated with a current participant in the healthcare social network,communicate with testing equipment coupled with the healthcare social network that are associated with the individual patient within the healthcare social network, the testing equipment including one or more of: a genome sequencing device; a magnetic resonance imaging device, and a positron emission tomography scanner, wherein the testing equipment provides signals that are generated by diagnostic testing performed on the individual patient using the testing equipment;transforming the signals that are based on one or more physical measurements into signal data in a first data structure,add content data for the individual patient in the healthcare social network, the content data including the signal data,invite a new participant to the healthcare social network to participate in the personalized health care of the individual patient in the healthcare social network, the new participant being invited to a first private portion of the healthcare social network based on an association between the current participant and the new participant outside of the healthcare social network,specify recipients within the healthcare social network particular to the individual patient for delivery of the content data, the recipients including: the new participant, the current participant, and one or more other participants of the healthcare social network that is not part of the first private portion of the healthcare social network, andbased on the specified recipients, provide a secure interface for the new participant, the current participant, and the one or more other participants to exchange bidirectional secure correspondence within the healthcare social network regarding the personalized health care of the individual patient and transfer the content data there between, the correspondence including communications regarding a particular treatment for the individual patient, wherein the secure interface is particular to the individual patient such that the secure interface limits participation to participants in the healthcare social network invited to participate in the personalized health care of the individual patient;one or more quantitation devices coupled with the one or more servers and configured to perform a quantitative analysis on the first data structure to generate quantitation data, the quantitation data added to the content data of the individual patient;a data analytics unit networked to the one o... | i2Dx, Inc. | A61B005/00 G06Q050/24
US09814892 | B2 | 2017-11-14 | Leadless pacemaker with tripolar electrode | 2016-08-25 | A leadless implantable medical device comprises a first electrode configured to deliver electrical pacing energy, a second electrode configured to sense intrinsic electrical cardiac activity, and a third electrode configurable to both deliver electrical pacing energy and sense intrinsic electrical cardiac activity. The first and third electrodes are used for delivering electrical pacing energy and the second and third electrodes are used to sense intrinsic electrical cardiac activity. None of the first, second and third electrodes are incorporated into a lead. | 1. A leadless implantable medical device comprising: a housing having an outer surface;a first electrode having a first electrode surface area;a second electrode positioned along the outer surface of the housing, the second electrode having a second electrode surface area;a third electrode positioned along the outer surface of the housing, the third electrode having a third electrode surface area;the first electrode surface area is less than one-tenth the second electrode surface area and less than one-tenth the third electrode surface area;wherein the leadless implantable medical device is configured to use both the second electrode and the third electrode as a pacing anode in combination with the first electrode as a pacing cathode for delivering electrical pacing energy; andwherein the leadless implantable medical device is configured to use at least one of the second electrode and the third electrode to sense intrinsic electrical cardiac activity. 2. The leadless implantable medical device of claim 1, wherein: the second electrode is spaced a first distance from the first electrode;the third electrode is spaced a second distance from the first electrode;wherein the first distance is the same as the second distance. 3. The leadless implantable medical device of claim 2, wherein: the second electrode is spaced a third distance from the third electrode, and the third distance is greater than the first distance and the second distance. 4. The leadless implantable medical device of claim 3, wherein the leadless implantable medical device is configured to use both the second electrode and the third electrode to sense intrinsic electrical cardiac activity. 5. The leadless implantable medical device of claim 1, wherein the housing has a first end, a second end and a side wall extending between the first end and the second end, and wherein the first electrode is incorporated into a pin that extends substantially orthogonal from the first end of the housing. 6. The leadless implantable medical device of claim 5, wherein the second electrode and the third electrode are positioned along the side wall of the housing. 7. The leadless implantable medical device of claim 5, wherein the second electrode and the third electrode are positioned along the first end of the housing. 8. A leadless implantable medical device comprising: a housing;a first electrode exposed to outside of the housing and having a first electrode surface area;a second electrode exposed to outside of the housing and having a second electrode surface area;a third electrode exposed to outside of the housing and having a third electrode surface area;wherein the second electrode surface area is at least ten times the first electrode surface area;a control circuit configured to select the first electrode as a pacing cathode and both the second electrode and the third electrode as a pacing anode for delivering electrical pacing energy; andthe control circuit further configured to select ... | Cardiac Pacemakers, Inc. | A61N001/375 A61N001/362 A61B005/042 A61N001/372 A61N001/05 A61N001/368
US09795789 | B2 | 2017-10-24 | Systems and methods to optimize anti-tachycardial pacing (ATP) | 2015-07-28 | Apparatus, systems and methods are provided for prevention and/or remediation of cardiac arrhythmias, e.g. optimizing anti-tachycardia pacing (ATP) algorithms. More particularly, implantable devices are provided that measure and treat cardiac arrhythmias. By monitoring the ATP attempt from additional electrodes, far-field morphology analyses, and/or measuring the return interval from a failed ATP attempt; the devices may estimate when entrainment has occurred, the amount of delay within the reentrant tachycardia, and/or tachycardia termination/acceleration. These variables and occurrences can be used to optimize the first and/or subsequent ATP attempts. Furthermore, other exemplary embodiments describe methods to integrate electrical restitution properties into the design of ATP pacing algorithms to facilitate tachycardia termination. | 1. A method for delivering therapy using an implant including a processor and a plurality of electrodes implanted in a patient's body adjacent a heart to terminate a ventricular tachycardia event, comprising: sensing cardiac signals from the heart via one or more of the plurality of electrodes, the processor coupled to the one or more of the plurality of electrodes to identify a ventricular tachycardia in response to the sensed cardiac signals having a tachycardia cycle length;delivering a plurality of overdrive pacing pulses to the heart via one or more pacing electrodes, the overdrive pacing pulses having overdrive pacing cycle lengths sufficient to advance or terminate the tachycardia;monitoring cardiac signals via one or more sensing electrodes spaced apart from the one or more pacing electrodes and the processor while delivering the overdrive pacing pulses to identify local depolarization of the heart after each overdrive pacing pulse;measuring time intervals between a time that the overdrive pacing pulses are transmitted by the one or more pacing electrodes and a time at which subsequent local depolarizations are identified by the one or more sensing electrodes; andadjusting delivery of the overdrive pacing pulses based at least in part on the time intervals between the overdrive pacing pulses and subsequent local depolarizations. 2. The method of claim 1, wherein monitoring cardiac signals comprises a combination of monitoring the cardiac signals using the one or more sensing electrodes and far-field morphology analyses. 3. The method of claim 2, further comprising the processor recording one or more time intervals from the one or more pacing electrodes to the one or more sensing electrodes when not in a ventricular tachycardia to assist in monitoring the ventricular tachycardia. 4. The method of claim 3, wherein if the measured time intervals indicate there has been conduction block within the ventricular tachycardia, adjusting delivery comprises stopping delivery of the overdrive pacing pulses. 5. The method of claim 1, wherein adjusting delivery of the overdrive pacing pulses comprises continuing to deliver overdrive pacing pulses at the overdrive pacing cycle lengths until the measured time intervals are unchanged for two consecutive pulses within a specified threshold. 6. The method of claim 5, wherein adjusting delivery of the pacing pulses includes delivering one or more priming pulses via the one or more pacing electrodes, the priming pulses having priming pulse lengths longer than the overdrive pacing cycle lengths used to advance the tachycardia via the one or more pacing electrodes; and thereafter, delivering at least one accelerating pulse via the one or more pacing electrodes, the at least one accelerating pulse having an accelerating pulse length shorter than the priming pulse lengths. 7. The method of claim 6, wherein the priming pulse cycle lengths are determined based at least in part on the measured time intervals a... | CARDIOFLOW TECHNOLOGIES, LLC | A61N001/36 A61N001/362 A61N001/365 A61B005/00 A61B005/0464
US09816956 | B2 | 2017-11-14 | Portable analytical device | 2015-08-19 | A portable analytical device includes a housing, a measuring facility and a processor in the housing, and an electrical connector coupled to the housing and having one end electrically connected to the processor and an opposite end extending outwardly from the housing. The measuring facility is configured to analyze an optical or electrochemical change in the test element having received a liquid sample of the body fluid and to generate measuring values resulting from the analysis. A memory has stored therein software at least a portion of which is executable by the processor to process the measuring values to produce analytical measuring data taking into account calibration values. The electrical connector may be a male universal serial bus (USB) electrical connector, a Type-A USB electrical connector or a mini-USB electrical connector having electrical connections for electrical power, ground, data and ID. | 1. A portable, patient-operable analytical device for analysis of a medically significant component of a body fluid for self-monitoring by a patient, comprising: a device housing defining an opening therein,a measuring facility arranged inside the device housing configured to receive a test element via the housing opening, the measuring facility configured to analyze one of an optical and electrochemical change in the test element having received a liquid sample of the body fluid and to generate measuring values resulting from the analysis,a processor arranged inside the device housing and including a memory having stored therein software at least a portion of which is executable by the processor to process the measuring values to produce analytical measuring data taking into account calibration values, andan electrical connector selected from the group consisting of: a male universal serial bus (USB) electrical connector carried by the housing and having one end electrically connected to the processor and an opposite end extending outwardly from the device housing, the device housing, measuring facility, processor and male USB electrical connector together forming a USB device,a Type-A universal serial bus (USB) electrical connector coupled to the housing and having one end electrically connected to the processor and an opposite end extending outwardly from the device housing, anda mini-universal serial bus (USB) electrical connector coupled to the housing and having one end electrically connected to the processor and an opposite end extending outwardly from the device housing, the mini-USB electrical connector having electrical connections for electrical power, ground, data and ID. 2. The analytical device of claim 1 wherein the opposite end of the electrical connector is configured to operatively connect to an external electronic device and to receive electrical power to operate the processor and the measuring facility from the external electronic device. 3. The analytical device of claim 1 wherein the opposite end of the electrical connector is configured to operatively connect to an external electronic device and includes a voltage supply input configured to transmit an external supply of electrical power to the analytical device from the external electronic device. 4. The analytical device of claim 1 wherein the opposite end of the electrical connector is configured to operatively connect to an external electronic device, and wherein the analytical device is configured to provide one or more portions of the software through the electrical connector to the external electronic device when the electrical connector is operatively connected to the external electronic device. 5. The analytical device of claim 1 wherein the opposite end of the electrical connector is configured to operatively connect to an external electronic device, and wherein one of operation of the analytical device and action of the analytical device to analyze the one of... | Roche Diabetes Care, Inc. | G01N027/327 A61B005/145 G01N033/487 G01N021/78 G01N033/66 G01N021/84 A61B005/00 G06F019/00
US09795796 | B2 | 2017-10-24 | Implantable electronic device employing fused thermoplastic-conductor subassembly and method of manufacturing the subassembly and the device | 2016-03-21 | Disclosed herein is an implantable electronic device. The device has a housing and a header connector assembly coupled to the housing. The header connector assembly has a connector assembly and a header enclosing the connector assembly. The connector assembly has a subassembly including an electrically conductive component at least partially residing within a first material that is provided about the electrically conductive component. The header has a second material that is provided about the connector assembly and the subassembly subsequent to the first material setting up about the electrically conductive component. | 1. A method of manufacturing an implantable electronic device, the method comprising: forming a fused thermoplastic-conductor subassembly by causing an electrically conductive component of a connector assembly of a header connector assembly to fuse to a thermoplastic; andforming a header of the header connector assembly about the fused thermoplastic-conductor subassembly;wherein the thermoplastic comprises at least two opposed portions and the electrically conductive component is sandwiched between the two opposed portions prior to the electrically conductive component being fused to the thermoplastic. 2. The method of claim 1, wherein the fusing is brought about by melting the thermoplastic via at least one of inductive heating, Ohmic heating or laser heating of the electrically conductive component. 3. The method of claim 1, wherein the thermoplastic comprises at least one of Tecothane, Pellethane, Bionate or Elasthane. 4. The method of claim 1, wherein the electrically conductive component comprises at least one of a wire, ribbon, antenna, block, or ring. 5. The method of claim 1, wherein the header is formed about the thermoplastic-conductor subassembly via at least one of a backfilling process, an overmolding process, an injection molding process, or a casting process. 6. The method of claim 1, wherein the sandwiched opposed portions and electrically conductive component are compressed between pressure platens when the electrically conductive component is being fused to the thermoplastic. 7. The method of claim 1, wherein the resulting implantable electronic device comprises an implantable pulse generator for administering electrotherapy via an implantable medical lead configured to couple with the implantable pulse generator. 8. The method of claim 1, wherein the resulting implantable electronic device comprises an implantable cardiac monitor. | PACESETTER, INC. | A61N001/00 A61N001/375 A61B005/0402 A61B005/00
US09797708 | B2 | 2017-10-24 | Apparatus and method for profiling a depth of a surface of a target object | 2013-05-03 | An apparatus (1) for profiling the depth of a surface of a target object (30), having a two-dimensional array of lasers (5), an optical device (15) for projecting a two-dimensional illumination pattern (31) onto an area of the surface of the target object, an image capture device (10) arranged to capture an image of the two-dimensional illumination pattern projected onto the area of the surface of the target object and a processor (25) configured to process the captured image in order to reconstruct a depth profile of the two-dimensional area of the surface of the target object from the image captured by the image capture device. | 1. A hand-held depth profiling apparatus for profiling a depth of a surface of a target object, the apparatus comprising: a first light source comprising a two-dimensional array of laser, wherein the first light source has an opening angle θp substantially between 5 and 45 degrees, and wherein the opening angle θp is inversely correlated to a focal length of the first light source;an optical device for projecting a two-dimensional multiline illumination pattern onto an area of the surface of the target object, wherein the optical device has a focal length on the order of several millimeters and wherein the two-dimensional multiline illumination pattern is distorted by a depth profile of the surface of the target object;an image capture device arranged to capture an image of the two-dimensional illumination pattern projected onto the area of the surface of the target object, wherein the image capture device is oriented at an angle θd relative to the first light source, and wherein the image capture device has a field of view of θs, wherein the angle θd depends on a depth range and area of the surface of the target object;a processor configured to process the captured image in order to reconstruct the distorted depth profile of a two-dimensional area of the surface of the target object from the image captured by the image capture device, andmeans for determining distance between the apparatus and the surface of the target object,wherein the two dimensional array comprises a plurality of rows, wherein at least one row is laterally offset with respect to an adjacent row. 2. An apparatus according to claim 1, wherein the two-dimensional array of lasers comprises at least two lasers having different shapes. 3. An apparatus according to claim 1, wherein the two-dimensional array is configured to be switched between a homogeneous illumination mode and a structured illumination mode. 4. An apparatus according to claim 1, wherein the means for determining distance between the two-dimensional array and the surface of a target object comprises a second light source inclined at a predetermined angle relative to the first light source, wherein the second light source is configured to project a second illumination pattern onto the surface of the target object and the image capture device is configured to capture an image of the second illumination pattern, andwherein the processor is further configured to process the captured image of the second illumination pattern and determine a distance between the first light source and the surface of the target object based on the predetermined angle and the captured image of the second illumination pattern. 5. An apparatus according to claim 1, wherein the two-dimensional array of lasers comprises vertical cavity surface emitting lasers. 6. A method of profiling a depth of a surface of a target object using a hand-held apparatus, the method comprising: projecting a two-dimensional multiline illumination pattern from... | KONINKLIJKE PHILIPS N.V. | G01B011/00 G01B011/02 A61B005/107 A61B005/00 G01B011/22 G01B011/25
US09826941 | B1 | 2017-11-28 | Pilot health monitoring and hypoxia prevention system and method | 2015-06-26 | A vehicular system includes a wearable device and a computing device. The wearable device is configured to be worn by an operator of a vehicle. The wearable device includes a sensor configured to measure a physiological state of the operator. The wearable device is configured to output physiological data indicative of the physiological state of the operator or operator health alert data indicative of a determined health problem associated with the operator. The computing device includes a processor configured to receive the physiological data associated with the operator or the operator health alert data from the wearable device and to determine whether the operator is experiencing a health problem based on the physiological data or the operator health alert data. | 1. An aviation system, comprising: a wearable device configured to be worn by a pilot of an aircraft, including: at least one sensor configured to measure at least one physiological state of the pilot;at least one wearable device processor coupled to the at least one sensor; anda wearable device display coupled to the at least one wearable device processor;wherein the wearable device is configured to output at least one of physiological data indicative of the at least one physiological state of the pilot or pilot health alert data indicative of a determined health problem associated with the pilot; andat least one computing device implemented in the aircraft, including: at least one processor configured to: receive at least one of the physiological data associated with the pilot or the pilot health alert data from the wearable device;determine whether the pilot is experiencing a health problem based on at least one of the physiological data or the pilot health alert data;route at least a portion of at least one of the physiological data or the pilot health alert data to at least one control station;receive an instruction from the at least one control station to engage an autopilot system of the aircraft after routing at least the portion of at least one of the physiological data or the pilot health alert data to the at least one control station; andin response to receipt of the instruction from the at least one control station to engage the autopilot system of the aircraft, engage the autopilot system of the aircraft. 2. The aviation system of claim 1, wherein the at least one sensor includes an oximeter, and the physiological data includes oxygen level data. 3. The aviation system of claim 1, wherein the at least one wearable device processor is configured to determine whether the pilot is experiencing the health problem based on the physiological data. 4. The aviation system of claim 3, wherein the at least one sensor includes an oximeter configured to measure arterial oxygen saturation levels, the physiological data includes oxygen level data indicative of measured arterial oxygen saturation levels, and the at least one wearable device processor is configured to determine whether the pilot is experiencing hypoxia based on the oxygen level data. 5. The aviation system of claim 4, wherein the at least one sensor further includes a sensor configured to measure heart rate, the physiological data further includes heart rate data indicative of a measured heart rate, and the at least one wearable device processor is configured to determine whether the pilot is experiencing the health problem based at least on the heart rate data and the oxygen level data. 6. The aviation system of claim 3, wherein the at least one wearable device processor is configured to output the pilot health alert data to be routed to at least one output component onboard the aircraft to be issued as at least one of a visual alert, an aural alert, or a vibratory alert. 7. ... | Rockwell Collins, Inc. | A61B005/00 A61B005/1455
US09788793 | B2 | 2017-10-17 | Method and system to identify motion artifacts and improve reliability of measurements and alarms in photoplethysmographic measurements | 2013-07-22 | A system (10) and method to identify motion artifacts. Measurements of a physiological parameter of an associated patient are received from a probe (12) positioned on or proximate to the associated patient. Further, measurements of acceleration are received from an accelerometer (26) positioned on, proximate to, or integrated with the probe (12). Measurements of the physiological parameter are labeled based on the measured acceleration, such as being with or without motion. | 1. A system for identifying motion artifacts, said system comprising: a probe configured to sense a physiological parameter of an associated patient, the probe configured to be positioned on or proximate to the associated patient, the probe being configured to generate one or more physiological signals indicative of the sensed physiological parameter;an accelerometer configured to be positioned on, proximate to, or integrated with the probe, the accelerometer configured to generate one or more acceleration signals measuring acceleration;a first signal processing unit configured to process the one or more physiological signals from the probe to determine measurements of the physiological parameter;a second signal processing unit configured to process one or more acceleration signals from the accelerometer to determine motion characteristics, the processing of the second signal processing unit being performed in parallel, and independent of, the processing of the first signal processing unit; anda labeling unit configured to label measurement of the physiological parameter with temporally corresponding motion characteristics based on the determined motion characteristics; anda signal quality index (SQI) threshold unit configured to discard or down weight measurements of the physiological parameter with signal quality indices (SQI)s below a SQI threshold and passing a remaining or down weighted measurements of the physiological parameter to the labeling unit; and wherein the SQI threshold unit is configured to determine the SQI threshold by: determining whether there is motion based on the measured acceleration;in response to determining there is no motion, setting the SQI threshold to a predetermined value;in response to determining there is motion, determining whether the motion is correlated with a measured motion artifact of the one or more physiological signals that is generated by, and received from, the probe;in response to determining the motion is correlated, setting the SQI threshold to a first value using a correlated lookup function of a motion strength value; andin response to determining the motion is uncorrelated, setting the SQI threshold to a second value using an uncorrelated lookup function of the motion strength value. 2. The system according to claim 1, wherein the second signal processing unit is further configured to generate a motion information signal (MIS) from the determined motion characteristics, an amplitude of the MIS corresponding to the measured acceleration. 3. The system according to claim 2, wherein the second signal processing unit is further configured to generate the MIS to include at least one of: digital messages indicating at least one of movement speed, acceleration, direction, and duration; anddigital messages that each identify a start and end time for a period of movement. 4. The system according to claim 1, further including: an alarm unit configured to generate alarms based on the measurements of th... | KONINKLIJKE PHILIPS N.V. | A61B005/00 A61B005/024 A61B005/11 A61B005/0205 A61B005/145 G01P015/00 A61B005/1455 A61B005/021
US09792814 | B2 | 2017-10-17 | Traffic disruption detection using passive monitoring of vehicle occupant frustration level | 2015-12-29 | Aspects of the present disclosure include a navigation system and computer-implemented methods for detecting traffic disruption events based on an analysis of input component data obtained from navigation-enabled devices of vehicles near a particular location. Traffic disruption events are events such as accidents, construction road closures, police and speed traps, or road hazards that cause a decrease in the flow of traffic along a particular route and thus, added time delays for occupants of vehicles traveling along those routes. The navigation system scores the input component data associated with each vehicle and aggregates the scored input component data to obtain a frustration score associated with the vehicle. The navigation system may detect traffic disruption events based on a number of vehicles near a particular area having associated frustration scores above a certain threshold. | 1. A system comprising: one or more processors of a machine; andone or more machine-readable mediums storing instructions to configure the one or more processors to perform operations comprising: detecting a traffic disruption event at a particular location, the detecting of the traffic disruption event including: identifying a plurality of navigation-enabled devices within a predefined distance of the particular location based on location data received from geolocation components of the plurality of navigation-enabled devices;determining a frustration score associated with each of the plurality of navigation-enabled devices based on input component data received from one or more input components coupled to or embedded in the plurality of navigation-enabled devices, the determining of the frustration score associated with each of the plurality of navigation-enabled devices including: scoring respective input component data based on input component type, andaggregating respective input component data scores to determine the frustration score, anddetermining that at least a portion of the plurality of the navigation-enabled devices within the predefined distance of the particular location have an associated frustration score above a threshold frustration score, the detecting of the traffic disruption event being in response to the determining that at least the portion of the plurality of the navigation-enabled devices within the predefined distance of the particular location have an associated frustration score above the threshold frustration score;in response to detecting the traffic disruption event, generating traffic disruption report data for the traffic disruption event; andstoring the traffic disruption report data at a database coupled to a navigation server. 2. The system of claim 1, wherein the operations further comprise transmitting a notification of the traffic disruption event to one or more of the plurality of navigation-enabled devices. 3. The system of claim 1, wherein the operations further comprise causing display of a map within a user interface, the map including an indicator of the traffic disruption event. 4. The system of claim 1, wherein the operations further comprise verifying a previous traffic disruption report maintained in the networked database using the traffic disruption report data. 5. The system of claim 4, wherein the verifying the previous traffic disruption report includes comparing location information included in the traffic disruption report data with location information included in previous traffic disruption data stored in the networked database. 6. The system of claim 1, wherein the scoring respective input component data includes comparing the input component data to baseline data stored in respective user account records maintained in the networked database. 7. The system of claim 1, wherein the input component data includes audio data obtained from a microphone, and wherein the determining of t... | eBay Inc. | G08G001/09 A61B005/0205 A61B005/024 G08G001/01 G08G001/0969 A61B005/00 A61B005/11
US09805922 | B2 | 2017-10-31 | System and method for rapid evaporative ionization of liquid phase samples | 2016-02-22 | According to some embodiments, systems and methods for rapid evaporation of liquid phase samples are provided. The method includes directing liquid samples to a thermal evaporation ionizing device, thermally evaporating the liquid samples to create gaseous molecular ions, and directing the gaseous molecular ions to an ion analyzer to analyze and provide information regarding the chemical composition of the liquid samples. | 1. A method for analyzing liquid phase samples, comprising generating a sample aerosol from a liquid sample;transporting the sample aerosol to an analyzer device;introducing the sample aerosol into a sub-atmospheric pressure regime of the ion analyzer device;desolvating the sample aerosol in the sub-atmospheric pressure regime of the ion analyzer device to produce sample ions. 2. The method of claim 1, wherein generating a sample aerosol comprises thermally evaporating the liquid sample. 3. The method of claim 2, wherein thermally evaporating the liquid sample comprises applying an electric current to the liquid sample through a gap between a pair of electrodes. 4. The method of claim 2, wherein thermally evaporating the liquid sample comprises heating the sample using at least one electromagnetic radiation beam. 5. The method of claim 4, wherein the at least one electromagnetic radiation beam comprises a laser. 6. The method of claim 1, wherein thermally evaporating the liquid sample comprises applying an electric current to the liquid sample between a pair of electrodes. 7. The method of claim 6, wherein the pair of electrodes is at least partially submerged in the liquid sample. 8. The method of claim 6, wherein applying an electric current comprises applying an alternating current. 9. The method of claim 6, wherein applying an electric current comprises applying a direct current. 10. The method of claim 1, wherein the ion analyzer device is a mass spectrometer or ion mobility spectrometer. | MICROMASS UK LIMITED | H01J049/16 G01N027/62 G01N033/493 H01J049/04 G01N030/72 H01J049/00 A61B005/20 A61B005/00
US09795177 | B1 | 2017-10-24 | Head-mounted impact sensing and warning device | 2015-03-01 | A head-mounted impact sensing and warning device includes a band adapted to be worn around at least a portion of the head or helmeted area of a user; an impact sensing device configured to sense an impact sustained by the head or the helmeted area of the user; at least one visual indicating device operatively coupled to the impact sensing device, the at least one visual indicating device configured to produce a visual indicator when the impact sensing device senses an impact that exceeds a predetermined impact level, thereby providing a visual alert that the impact sustained by the head or the helmeted area of the user exceeded the predetermined impact level; and a unit housing containing the impact sensing device and the at least one visual indicating device, the unit housing being coupled to the band or the helmeted area. | 1. A head-mounted impact sensing and warning device comprising, in combination: a band adapted to be worn around at least a portion of the head or helmeted area of a user;an impact sensing device configured to sense an impact sustained by said head or said helmeted area of said user, said impact sensing device being capable of 360 degree impact measurement, thereby allowing full range measurement of g-forces on and around said head or said helmeted area of said user;at least one visual indicating device operatively coupled to said impact sensing device, said at least one visual indicating device configured to produce a visual indicator when said impact sensing device senses an impact that exceeds a predetermined impact level, thereby providing a visual alert that said impact sustained by said head or said helmeted area of said user exceeded said predetermined impact level;a microcontroller operatively coupled to said impact sensing device and said at least one visual indicating device, said microcontroller configured to process impact data sensed by said impact sensing device, and to selectively activate said at least one visual indicating device when an impact sensed by said impact sensing device exceeds said predetermined impact level;a motion sensing device, said motion sensing device being a component that is separate from said impact sensing device, said motion sensing device configured to sense a motion imparted on said head-mounted impact sensing and warning device by one or more external forces and to output a signal to said microcontroller, whereby, upon receiving said signal from said motion sensing device, said microcontroller is configured to activate said impact sensing device and to change said head-mounted impact sensing and warning device from an inactive mode to an active mode for a predetermined period of time; anda unit housing containing said impact sensing device, said at least one visual indicating device, said microcontroller, and said motion sensing device, said unit housing being coupled to said band or said helmeted area;wherein said at least one visual indicating device is disposed within said unit housing, and wherein said unit housing is substantially transparent such that said visual indicator produced by said at least one visual indicating device is visible through said unit housing. 2. The head-mounted impact sensing and warning device according to claim 1, wherein said band comprises one of the following: (i) a headband; (ii) a bandana; and (iii) a strap for securing glasses or goggles around said head or said helmeted area of said user. 3. The head-mounted impact sensing and warning device according to claim 1, wherein said band comprises one of the following materials: (i) a moisture-wicking material; (ii) polyester; (iii) a moisture management antimicrobial performance fabric; (iv) spandex; (v) cotton; (vi) a cotton blend; (vii) a polyester/spandex mix; (viii) elastic; (ix) leather; (x) nylon; (xi) silicone; ... | - | A61B005/00 G08B023/00 A42B003/04
US09782078 | B2 | 2017-10-10 | Device, system and method for blood vessel imaging and marking | 2014-02-13 | The present invention relates to a system and a method for blood vessel imaging, detection, and marking and in particular, to such a system and method in which optical enhancers are utilized to increase the accuracy and quality of the process for detecting a blood vessel. | 1. An optical blood vessel enhancing system for identify a blood vessel within a target area, the system comprising: a. a light source configured to transmit at least one wavelength incident light within the infrared spectrum onto said target area, said light source comprising at least one wavelength;b. an optical sensor responsive to said infrared spectrum light, wherein said optical sensor is configured to receive light reflected from said target area therein forming a contrasted image based on said reflected light and at least a selective portion of said incident light;c. a controllable optical enhancer in the form of a controllable polarizer provided for controlling the optical properties of at least one of said incident light and/or said reflected light; andd. a processor functionally coupled with said optical sensor for enhancing said contrasted image to identify a blood vessel with said target area by distinguishing between a potential blood vessel from the surroundings target area; and for controlling said controllable optical enhancer; wherein enhancing said contrasted image is provided for with said processor and includes: i) obtaining a plurality of images of the target area, characterized in that each image is obtained at a different polarity rotation; andii) said plurality of images are evaluated relative to one another by way of superimposing said plurality of images to define said enhanced image. 2. The blood vessel enhancing system of claim 1 comprising at least two controllable optical enhancers wherein each of said optical enhancer is a polarizer; a first controllable optical enhancer disposed about said light source for controlling the optical properties of said incident light;a second controllable optical enhancer disposed about said optical sensor for controlling the optical properties of said detected reflected light;and wherein said processor provides for controlling said first optical enhancer and said second optical enhancer with respect to one another. 3. The blood vessel enhancing system of claim 2 wherein said polarizers are selected from S rotation polarizer or P rotation polarizer. 4. The blood vessel enhancing system of claim 2 further comprising at least two polarizers disposed about said optical sensor and at least two polarizers disposed about said light source. 5. The blood vessel enhancing system of claim 1 wherein said optical enhancer is provided in the form selected from: a motorized rotating polarizer, a liquid crystal cell (LCD), pi-cell liquid crystal, TN (twisted nematic) liquid crystal cell, liquid crystal polarization rotator, liquid crystal polarization retarder, and any combination thereof. 6. The blood vessel enhancing system of claim 1 further comprising: a. a printer to mark the blood vessel within said target area to reveal the blood vessel locations therein; andb. a fixation module to minimize or prevent relative movement between the imaging module and the printer head to ensure that the bl... | - | A61B005/00
US09775530 | B2 | 2017-10-03 | Foley catheter with ring electrodes | 2014-11-03 | An apparatus for monitoring a nerve includes a Foley type catheter having an exterior surface, and first and second pairs of ring electrodes formed on the exterior surface of the Foley catheter. Each of the first and second pairs of ring electrodes is configured to perform at least one of stimulating the nerve and recording nerve activity. | 1. An apparatus for monitoring a nerve, comprising: a Foley type catheter having an exterior surface;first and second pairs of ring electrodes formed on the exterior surface of the Foley catheter, wherein the second pair of ring electrodes delivers bipolar nerve stimulation, and the first pair of ring electrodes provides bipolar recording of nerve activity;a processing apparatus to process the recorded nerve activity; anda first conductor coupled to the processing apparatus and the first pair of ring electrodes to carry signals representing the recorded nerve activity to the processing apparatus. 2. The apparatus of claim 1, and further comprising: a second conductor coupled to the processing apparatus and the second pair of ring electrodes and configured to carry signals between the second pair of ring electrodes and the processing apparatus. 3. The apparatus of claim 1, wherein the first and second pairs of ring electrodes each comprise an anode ring electrode and a cathode ring electrode. 4. The apparatus of claim 1, and further comprising: a third pair of ring electrodes formed on the exterior surface of the Foley catheter, wherein the third pair of ring electrodes is configured to perform at least one of stimulating the nerve and recording nerve activity. 5. The apparatus of claim 4, wherein the first pair of ring electrodes includes first and second ring electrodes that are separated by a distance of about 5 mm to 10 mm, and wherein the second pair of ring electrodes includes third and fourth ring electrodes that are separated by a distance of about 5 mm to 10 mm, and wherein the third pair of ring electrodes includes fifth and sixth ring electrodes that are separated by a distance of about 5 mm to 10 mm. 6. The apparatus of claim 5, wherein a center of the first pair of ring electrodes is separated from a center of the second pair of ring electrodes by a distance of about 18 mm to about 22 mm, and wherein a center of the second pair of ring electrodes is separated from a center of the third pair of ring electrodes by a distance of about 18 mm to about 22 mm. 7. The apparatus of claim 6, wherein the third pair of ring electrodes is most distally positioned on the Foley catheter, and wherein the third pair of ring electrodes is positioned proximally from a proximal end of a balloon of the Foley catheter by a distance of about 35-40 mm. 8. The apparatus of claim 1, wherein the first pair of ring electrodes includes first and second ring electrodes that are separated by a distance of about 5 mm to 10 mm, and wherein the second pair of ring electrodes includes third and fourth ring electrodes that are separated by a distance of about 5 mm to 10 mm. 9. The apparatus of claim 8, wherein a center of the first pair of ring electrodes is separated from a center of the second pair of ring electrodes by a distance of about 8 mm to about 12 mm. 10. The apparatus of claim 9, wherein the second pair of ring electrodes is most distally positioned... | Medtronic Xomed, Inc. | A61B005/04 A61N001/05 A61B005/00 A61M025/00 A61M025/10 A61N001/36
US09801574 | B2 | 2017-10-31 | Techniques to improve polyurethane membranes for implantable glucose sensors | 2015-10-05 | The invention provides an implantable membrane for regulating the transport of analytes therethrough that includes a matrix including a first polymer; and a second polymer dispersed throughout the matrix, wherein the second polymer forms a network of microdomains which when hydrated are not observable using photomicroscopy at 400× magnification or less. In one aspect, the homogeneous membrane of the present invention has hydrophilic domains dispersed substantially throughout a hydrophobic matrix to provide an optimum balance between oxygen and glucose transport to an electrochemical glucose sensor. | 1. An implantable device for measuring an analyte in a biological fluid, comprising: an implantable sensor comprising an electrochemically reactive surface; anda membrane disposed over at least a portion of the electrochemically reactive surface, wherein the membrane comprises a matrix comprising a first polymer and a second polymer dispersed throughout the matrix, wherein the first polymer is polyurethane urea, wherein the second polymer comprises at least one hydrophilic component, wherein the hydrophilic component of the second polymer forms in the membrane hydrophilic segments along which the analyte diffuses. 2. The implantable device of claim 1, wherein the second polymer comprises polyethylene oxide. 3. The implantable device of claim 2, wherein polyethylene oxide comprises about 20% by weight of the second polymer. 4. The implantable device of claim 2, wherein the polyethylene oxide has an average molecular weight from 200 to 3,000 Daltons. 5. The implantable device of claim 1, wherein the hydrophilic segments comprise a network of microdomains. 6. The implantable device of claim 5, wherein the microdomains are not photomicroscopically observable when hydrated at 400× magnification or less. 7. The implantable device of claim 5, wherein the network of are substantially dispersed throughout the matrix. 8. The implantable device of claim 1, wherein the membrane has a thickness from 5 to 100 microns. 9. The implantable device of claim 1, wherein the membrane is configured to yield calibration factors that are stable over the life of the sensor. 10. The implantable device of claim 1, wherein the analyte is glucose. | DexCom, Inc. | A61B005/05 A61B005/1486 C12Q001/00 A61L031/12 B01D069/14 B01D071/54 B01D071/80 A61B005/145 A61B005/1473 B05D003/00 B05D005/00
US09814394 | B2 | 2017-11-14 | Noise suppression in an optoacoustic system | 2013-03-15 | An optoacoustic imaging system includes an ultrasound transducer array, first and second light sources and a switching power supply that generates power for the light sources. The switching power supply includes an input for impeding its switching operation. A data acquisition unit samples the ultrasound transducer array for a first predetermined period of time after a pulse of light from the first light source and for a second predetermined period of time after a pulse of light from the second light source and stores the sampled data. A master processor utilizes the input to impede operation of the switching operation of the switching power supply during the first predetermined period of time after a pulse of light from the first light source, and during the second predetermined period of time after a pulse of light from the second light source. | 1. An optoacoustic imaging system comprising: a handheld probe having a distal end and a proximal end, the handheld probe comprising an ultrasound transducer array having an active end located at the distal end of the handheld probe, the ultrasound transducer array capable of receiving an optoacoustic return signal during an acquisition period and generating data signals corresponding to the optoacoustic return signal, the handheld probe further comprising an optical window located at the distal end of the handheld probe;one or more light sources having a common output, the one or more light sources configured to generate pulses of light to produce first light events at a first predominant wavelength and to produce second light events at a second predominant wavelength;one or more sync detectors configured to generate a first sync signal upon detection of a pulse of light at the first predominant wavelength and to generate a second sync signal upon detection of a pulse of light at the second predominant wavelength;a switching power supply generating power for the one or more light sources, the switching power supply configured to impede power to the one or more light sources upon receipt of a blanking input;a light path configured to direct light from the common output to the optical window of the handheld probe;a data acquisition unit configured to sample the data signals during a first data acquisition period and to store the sampled data signals, wherein the first data acquisition period is less than about one percent (1%) of the first light event period; anda master processor configured to provide the blanking input to the switching power input to impede power to the one or more light sources during a data acquisition period upon receipt of the first sync signal or the second sync signal to cause the optoacoustic imaging system to produce a plurality of light event sequences, each of the light event sequences comprising (i) a first light event at the first predominant wavelength, (ii) a first light event period comprising a first data acquisition period following the first light event, and (iii) a second light event at the second predominant wavelength. 2. The optoacoustic imaging system of claim 1, wherein the data acquisition unit is further configured to sample the data signals during a second data acquisition period and to store the sampled data signals, wherein the second data acquisition period is less than about one percent (1%) of a second light event period, andwherein the master processor is further configured to provide the blanking input to the switching power supply to impede power to a first light source and a second light source during the second data acquisition period so as to produce a light event sequence further comprising a second light event period comprising the second data acquisition period following the second light event. 3. The optoacoustic imaging system of claim 2, wherein the first and second data acquisition ... | Seno Medical Instruments, Inc. | A61B005/05 A61B005/00 A61B005/145 A61B008/00
US09775553 | B2 | 2017-10-03 | Method and apparatus for a fluid sampling device | 2008-10-01 | A blood analyzer device has a housing with a top section coupled to a bottom section, a driver and a plurality of penetrating members housed in a disposable positionable in the housing. A gripper engages each penetrating member. A manually actuated button advances the disposable to move penetrating members into launch positions. A power is source coupled to the driver. A display is at the housing. | 1. A blood analyzer device, comprising: a housing; a penetrating member driver positioned in the housing; a plurality of penetrating members housed in a disposable positionable in the housing;a gripper assembly that includes a gripper that engages a penetrating member of the plurality of penetrating members with the driver prior to launch of a penetrating member during a lancing event, an open channel that is positioned over the gripper in the gripper assembly, a gripper shaft assembly including a flag, a non-polymeric rod, a polymeric gripper shaft, a magnet and a stepped gripper portion that captures and retains a penetrating member;an actuator assembly that includes a gripper track coupled to the gripper assembly, a sensor, and a gripper insertion arm;a sweeper arm coupled to the gripper assembly, the sweeper arm configured to physically retain the gripper assembly and prevent it from accidentally exiting the housing; anda power source coupled to a drive assembly and the penetrating member. 2. The device of claim 1, further comprising: a sterility barrier that maintains the plurality of penetrating members in a sterile environment. 3. The device of claim 1, further comprising: a mechanical assembly that includes a cam. 4. The device of claim 1, further comprising: a sterility barrier opener that moves a sterility barrier out of the way of a launched penetrating member. 5. The device of claim 4, wherein a sterility barrier opener is a punch that moves at least a portion of the sterility barrier out of a way of a launched penetrating member. 6. The device of claim 5, wherein the punch lifts up at least a portion of the sterility barrier. 7. The device of claim 1, wherein the power source is a battery. 8. The device of claim 5, further comprising: battery contacts. 9. The device of claim 1, further comprising: a door assembly. 10. The device of claim 9, further comprising: a door latch. 11. The device of claim 10, further comprising: a door latch retainer. 12. The device of claim 1, wherein the top section of the housing includes a window. 13. The device of claim 12, further comprising: a removable window protection film. 14. The device of claim 13, further comprising: a display. 15. The device of claim 14, wherein the display is a liquid crystal display. 16. The device of claim 15, further comprising: a liquid crystal display gasket positioned between the housing and the liquid crystal display. 17. The device of claim 1, wherein the housing includes a finger interface section. 18. The device of claim 17, wherein the finger interface section is indented. 19. The device of claim 18, wherein the finger interface section is conformable to an anatomy of a fingertip. 20. The device of claim 1, further comprising: a fire button that launches penetrating members. 21. The device of claim 1, further comprising: a slider handle. 22. The device of claim 1, further comprising: a sterility barrier detect spring. 23. The device of claim 1... | SANOFI-AVENTIS DEUTSCHLAND GMBH | A61B005/00 B65D081/00 A61B017/14 A61B017/32 A61B005/157 A61B005/151 A61B005/15
US09820695 | B2 | 2017-11-21 | Method for detecting contact with the wall of a region of interest | 2010-03-29 | A method for detecting when a tool contacts a wall of a region of interest includes the steps of: tracking the motion of the tool and creating a tool motion profile; calculating a correlation value representative of the correlation between the tool motion profile and a wall motion profile indicative of the motion of the wall; and determining when the correlation value exceeds a threshold indicative of contact of the tool with the wall of the region of interest. | 1. A method for tracking a tool relative to a wall of a region of interest of a patient, the method comprising: tracking the motion of the tool, wherein the tool includes a positioning sensor at a distal end thereof configured to (i) sense one or more characteristics of a magnetic field produced by a medical positioning system and (ii) generate a signal indicative of a sensor position of the positioning sensor, and creating a tool motion profile including generating, using the medical positioning system in response to said signal, the tool position;calculating a correlation value representative of the correlation between the tool motion profile and a wall motion profile indicative of the motion of the wall, wherein calculating the correlation value comprises comparing at least one of the (1) frequency-based components, (2) time-based components and (3) spatial-based components of the tool motion profile with the corresponding at least one components of the wall motion profile, wherein the at least one component is the frequency-based component, and wherein said comparing includes determining whether a predetermined frequency component present in the wall motion is present in the tool motion;determining when the correlation value exceeds a threshold, the correlation value exceeding the threshold being indicative of contact of the tool with the wall of the region of interest; andproviding a contact indicator for signaling when the threshold level has been exceeded, therein indicating the tool is in contact with the wall of the region of interest. 2. The method of claim 1 wherein said step of tracking and creating the tool motion profile further includes: generating time-based position data representative of the motion of the tool using the positioning sensor, said time-based position data comprising a plurality of tool positions at respective times. 3. The method of claim 1 further including: creating the wall motion profile based on a predetermined motion pattern data for the region of interest. 4. The method of claim 3 wherein creating the wall motion profile includes: obtaining motion pattern data from patients other than said patient. 5. The method of claim 1 further including: selecting a motion measurement strategy for the region of interest from the group comprising image analysis, electrical parameter measurement including electro-cardiogram analysis, and positioning sensor-based measurements;obtaining measured region of interest motion data using the selection measurement strategy; andcreating the wall motion profile based on at least the measured region of interest motion data. 6. The method of claim 5 wherein creating the wall motion profile is further performed based on one of (i) predetermined motion pattern data for the region of interest and (ii) a combination of the predetermined motion pattern data and the measured region of interest motion data. 7. The method of claim 5 wherein the selected motion measurement strategy compri... | St. Jude Medical International Holding S.àr.l. | A61B005/00 A61B005/06
US09801548 | B2 | 2017-10-31 | Apparatus and method for fluorescence imaging and tomography using spatially structured illumination | 2013-08-05 | Apparatus and method for performing depth sectioned fluorescence imaging of a turbid sample including a fluorescent turbid medium, uses an apparatus for quantitative modulated fluorescence imaging, the apparatus including projection optics with a first optical axis, to expose the turbid sample to a periodic pattern of excitation radiation to provide depth-resolved discrimination of fluorescent structures within the turbid medium; an image capture module, including a second optical axis and a detection beam path, to receive a data image from the sample; and a signal processor to transform the data image from the sample, spatially filter the transformed data image from the sample, and reconstruct the filtered, transformed data image from the sample. | 1. An apparatus to perform depth sectioned fluorescence imaging of a turbid sample including a fluorescent turbid medium, the apparatus comprising: projection optics, including a first optical axis, to expose the turbid sample to a periodic pattern of excitation radiation to provide depth-resolved discrimination of fluorescent structures within the turbid medium;an image capture module, including a second optical axis and a detection beam path, to receive a data image from the sample, whereby the first optical axis is inclined relative to the second optical axis;a signal processor to transform the data image from the sample, spatially filter the transformed data image from the sample, and reconstruct the filtered, transformed data image from the sample;the projection optics including an object plane and an image plane;the periodic pattern of excitation radiation having periodicity in a direction perpendicular to a direction of a projection of the first optical axis onto the image plane, so that the phase of the periodic pattern of excitation radiation does not change with increasing depth into an image space; andthe projection optics having an angle of inclination relative to an image plane of the apparatus, the angle of inclination being selected such that the component of excitation radiation incident upon the sample that is not absorbed by the sample is scattered by the sample. 2. The apparatus of claim 1, wherein the projection optics comprise one or more non-telecentric Scheimpflug lens systems. 3. The apparatus of claim 2, wherein the projection optics comprise two non-telecentric Scheimpflug lens systems that are mirror-symmetric. 4. The apparatus of claim 2, wherein the one or more non-telecentric Scheimpflug lens systems are zoomable. 5. The apparatus of claim 2, wherein the one or more non-telecentric Scheimpflug lens systems include a plurality of fixed-focal lens systems, each providing a different magnification. 6. The apparatus of claim 1, wherein the projection optics comprise one or more Scheimpflug lens systems providing object space telecentricity. 7. The apparatus of claim 6, wherein the projection optics comprise two Scheimpflug lens systems providing object space telecentricity that are mirror-symmetric. 8. The apparatus of claim 6, wherein the one or more Scheimpflug lens systems providing object space telecentricity are zoomable. 9. The apparatus of claim 6, wherein the one or more Scheimpflug lens systems providing object space telecentricity include a plurality of fixed-focal lens systems each providing a different magnification. 10. The apparatus of claim 1, wherein the projection optics comprise one or more Scheimpflug lens systems providing image space telecentricity. 11. The apparatus of claim 10, wherein the projection optics comprise two Scheimpflug lens systems providing image space telecentricity that are mirror-symmetric. 12. The apparatus of claim 10, wherein the Scheimpflug lens systems providing ima... | Bruker Biospin Corporation | A61B005/00 G01N021/64 G01N021/17
US09801549 | B2 | 2017-10-31 | Apparatus for finding a functional tissue area in a tissue region | 2015-04-14 | The invention is directed to an apparatus for finding a functional tissue area in a tissue region. The apparatus has a measurement illuminating device suitable for emitting measurement illumination to the tissue region and a camera which can capture light reflected by the tissue region. The camera has a green channel and/or a blue channel wherein there is a change in an optical property of the light reflected by the tissue region during the stimulation thereof which is undertaken at least intermittently. An evaluation unit captures the change in the optical property only by a signal of the green channel and/or of the blue channel of the camera. A display unit can display an output signal of the evaluation unit for the functional tissue area in the tissue region. | 1. An apparatus for finding a functional tissue area in a tissue region, the apparatus comprising: a measurement illuminating device configured to emit a measurement illumination on to the tissue region;a camera for capturing light reflected from the tissue region;said camera being configured to have at least one of a green channel and a blue channel wherein there is a change in at least one optical property of the light reflected by the tissue region during stimulation of the tissue region undertaken at least intermittently;said camera being further configured to emit a camera signal exclusively of at least one of said green channel and said blue channel;an evaluation unit for detecting the change in the at least one optical property only via said camera signal and supplying an output signal; and,a display unit for displaying said output signal of said evaluation unit for the functional tissue area of the tissue region. 2. The apparatus of claim 1, wherein said at least one optical property is an intensity or a wavelength of the reflected light. 3. The apparatus of claim 1, wherein the change of the at least one optical property is pulsed. 4. The apparatus of claim 1, wherein said measurement illuminating device is configured to emit a pulsed measurement illumination. 5. A method of finding a functional tissue area in a tissue region, the method comprising the steps of: providing a camera exclusively having a green channel and/or a blue channel;illuminating the tissue region with a measurement illuminating device;at least intermittently stimulating the tissue region to cause a change of at least one optical property of the light reflected by the tissue region;capturing the change of the at least one optical property by an evaluation unit only via a signal of the green channel and/or blue channel of the camera; and,displaying an output signal of the evaluation unit on a display unit. 6. A surgical microscope comprising: an apparatus for finding a functional tissue area in a tissue region, the apparatus including:a measurement illuminating device configured to emit a measurement illumination on to the tissue region;a camera for capturing light reflected from the tissue region;said camera being configured to have at least one of a green channel and a blue channel wherein there is a change in at least one optical property of the light reflected by the tissue region during stimulation of the tissue region undertaken at least intermittently;said camera being further configured to emit a camera signal exclusively of at least one of said green channel and said blue channel;an evaluation unit for detecting the change in the at least one optical property only via said camera signal and supplying an output signal; and,a display unit for displaying said output signal of said evaluation unit for the functional tissue area of the tissue region. 7. The surgical microscope of claim 6, wherein said at least one optical property is an intensity or a wavelen... | Carl Zeiss Meditec AG Technische Universitaet Dresden Institut fuer Biomedizinische Technik Technische Universitaet Dresden Klinik und Poliklinik fuer Neurochirurgie | A61B005/00 A61B005/026 A61B005/1455 A61B090/20
US09795451 | B2 | 2017-10-24 | Surgical monitoring system and related methods for spinal instrument angular relation | 2015-08-31 | The present invention relates to a system and methods generally aimed at monitoring the angular orientation between two locations within a fluoroscopic image and especially for monitoring the angular orientation between two locations within a fluoroscopic image and a predetermined target angle. | 1. A system for use during spinal surgery, comprising: a control unit comprising a computer,a fluoroscope in communication with the control unit;a reticle comprising a laser, and one or more radiopaque markers, said reticle attached to an image intensifier on the fluoroscope;a spinal surgical instrument with a tilt sensor; anda display communicatively linked to said control unit;wherein said control unit is configured to receive a fluoroscopic image from the fluoroscope, obtain a reading of a gravitational angular orientation from the tilt sensor of the spinal surgical tool, and determine an angular orientation between at least two image locations within the fluoroscopic image;and wherein, said display is configured to display at least one of a first alpha-numeric, graphic, and color indicia related to the angular orientation between the at least two image locations within the fluoroscopic image, wherein the angular orientation between the at least two image locations is based on a vertical reference line relative to gravity and one of a medial-lateral trajectory line and a cranial-caudal trajectory line of a spine. 2. The system of claim 1, wherein the medial-lateral trajectory line is a line from a central position in a pedicle to an anterior point of a vertebral body. 3. The system of claim 1, wherein the cranial-caudal trajectory line is a line from a central position in a pedicle to an anterior point of a vertebral body. 4. The system of claim 1, wherein the fluoroscopic image is a lateral image in relation to a direction of gravity. 5. The system of claim 1, wherein the control unit is further configured to monitor neurophysiologic changes in and adjacent the spine. 6. The system of claim 1 further comprising, the display being further configured to display at least one of a second alpha-numeric, graphic, and color indicia indicative of a relationship between said angular orientation and a target. 7. The system of claim 6, wherein said second color indicia includes at least one of a green color indicative of an optimal relationship between said angular orientation and the target, a red color indicative of an unacceptable relationship between said angular orientation and the target, and a yellow color indicative of an acceptable yet not optimal relationship between said angular orientation and the target. 8. The system of claim 6, wherein the target is a cranial-caudal angle. 9. The system of claim 1, wherein the tilt sensor is in communication with the control unit. 10. The system of claim 1, wherein the reticle further comprises an integrated tilt sensor. 11. The system of claim 1, wherein the radiopaque marker is a cross-hair. 12. A non-transitory computer readable medium with computer executable instructions stored thereon, that when executed on a processor, perform the steps of: receive a fluoroscopic image from a C-arm fluoroscopy device;determine from a lateral fluoroscopic image a first spinal angle relative to a direction... | NuVasive, Inc. | A61B019/00 A61B006/00 A61B006/12 A61F002/46 A61B017/17 A61B005/04 A61N001/05 A61B034/20 A61B090/00
US09814850 | B2 | 2017-11-14 | Patient sleep therapy self management tool | 2013-06-10 | A tool for use in conjunction with a pressure support system that is structured to provide therapy to a patient to treat a condition of the patient by delivering a flow of breathing gas to the patient. The tool may be implemented on a portable electronic device or a PC and is configured to, among other things, provide customized/personalized education and feedback to the patient based, at least in part, on data that is measured by the pressure support system during the provision of therapy to the patient. The tool utilizes certain patient/therapy metrics, wherein each patient/therapy metric comprises raw data that was measured by the pressure support system and that has been processed (e.g., summarized and/or otherwise manipulated) to form the patient/therapy metric. | 1. A method of providing information to a patient, wherein the patient uses a pressure support system structured to provide therapy to the patient to treat a condition of the patient by delivering a flow of breathing gas to the patient, comprising: receiving in a computing device first information identifying one or more particular components of the pressure support system;receiving in the computing device second information identifying a data connection methodology by which data measured by the pressure support system during a provision of therapy by the pressure support system is to be transferred from the pressure support system to a location where the data is processed; anddisplaying third information on a display of the computing device, the third information including: (i) information relating to use of or care for the pressure support system, (ii) information relating to transferring the data from the pressure support system, and (iii) information relating to one or more solutions for problems relating to use of the pressure support system, the third information being selected based on either or both of the first information and the second information. 2. The method according to claim 1, further comprising obtaining in the computing device patient/therapy metric data, wherein the patient/therapy metric data is formed by processing the data measured by the pressure support system during the provision of therapy to the patient, and displaying the patient/therapy metric data on the display of the computing device. 3. The method according to claim 2, wherein the patient/therapy metric data is displayed in a metrics area of a first screen on the display and at least a portion of the third information is simultaneously displayed in a dynamic messaging area of the first screen. 4. The method according to claim 2, wherein the patient/therapy metric data is at least one of: (i) a first metric relating to a degree of mask leak, (ii) a second metric indicating a length of use of the pressure support device, and (iii) a third metric indicating a severity or state of the condition of the patient. 5. The method according to claim 1, wherein the one or more particular components of the pressure support system include a particular mask or a particular pressure generating device, wherein the first information includes information identifying the particular mask or information identifying the particular pressure generating device. 6. The method according to claim 5, wherein the third information includes one of (i) a device guide relating to operation and use of the particular pressure generating device, the device guide including at least one of text and one or more links to a video, and (ii) a mask guide relating to use of the particular mask, the mask guide including at least one of a text and one or more links to a video. 7. The method according to claim 2, wherein the third information is based on the first information and the patient/therapy met... | Koninklijke Philips N.V. | A61M016/00 G06F019/00 A61B005/00 A61M016/06
US09801557 | B2 | 2017-10-31 | Catheter or guidewire device including flow sensing and use thereof | 2016-02-18 | Devices and methods are provided for performing procedure on tissue with flow monitoring using flow sensors. The devices include an elongated member, and at least one flow sensor disposed on the elongated member. The flow sensor includes at least one temperature sensor and at least one heating element having a cavity. At least a portion of the at least one temperature sensor is housed in the cavity. A temperature measurement of the temperature sensor provides an indication of the flow rate of a fluid proximate to the flow sensor. | 1. A device for determining a flow rate of a fluid proximate to a portion of a tissue, comprising: an elongated member having a proximal portion and a distal portion;a flow sensor disposed proximate to the distal portion of the elongated member, the flow sensor comprising: at least one temperature sensor; andat least one heating element to heat an area proximate to the elongated member, at least a portion of the at least one heating element forming a cavity,wherein: at least a portion of the at least one temperature sensor is housed in a portion of the cavity; anda temperature measurement of the temperature sensor provides a first indication of a flow rate of the fluid proximate to the flow sensor. 2. The device of claim 1, further comprising an inflatable and/or expandable body coupled to a portion of the elongated member and having a proximal portion and a distal portion, wherein the distal portion of the inflatable and/or expandable body is disposed proximate to the flow sensor. 3. The device of claim 2, further comprising an electronic circuit coupled with the inflatable and/or expandable body, wherein the electronic circuit comprises at least one stretchable interconnect, and wherein the electronic circuit is stretchable and conformable such that the electronic circuit accommodates an expansion of the inflatable and/or expandable body. 4. The device of claim 3, wherein the electronic circuit further comprises at least one passive electronic component and/or at least one active electronic component, and wherein the at least one stretchable interconnect electrically couples at least two electronic components of the electronic circuit. 5. The device of claim 3, wherein the electronic circuit further comprises a plurality of electrodes, and wherein at least one electrode of the plurality of electrodes is a radiofrequency electrode to deliver a radiofrequency energy to a surface proximate to the radiofrequency electrode. 6. The device of claim 3, wherein the electronic circuit comprises at least one ablative element. 7. The device of claim 1, wherein the device is adapted for performing a procedure on the portion of tissue, and wherein the procedure is a denervation procedure or a nerve stimulation procedure. 8. The device of claim 7, wherein the procedure is a carotid sinus denervation, a carotid body disruption, a vagus nerve stimulation, a pulmonary artery denervation, a celiac ganglion disruption, a bladder trigone ablation, or a renal denervation. 9. The device of claim 1, wherein the temperature sensor is a thermistor or a thermocouple. 10. The device of claim 1, wherein the elongated member is a catheter or a guide wire. 11. The device of claim 1, further comprising a reference temperature sensor disposed on a proximal portion of the elongated member. 12. The device of claim 11, wherein a comparison of a temperature measurement of the reference temperature sensor to a measurement of the flow sensor provides a second indication o... | MC10, Inc. | A61B018/04 A61B005/027 A61N001/36 A61B034/35 A61B090/70 A61B018/14 A61N001/05 A61B090/30 A61B005/00 A61B018/00 A61B034/00 A61B090/00 A61N007/02 A61B018/02 A61B018/22 A61B018/18
US09798796 | B2 | 2017-10-24 | Ranking device, ranking method, and program | 2012-10-12 | A device, a system, a method, and a program are realized which are capable of predicting the order of the feeling and the preference of a subject to products which are the research objects, for marketing research and others. The images of a plurality of research objects are presented a plurality of times as visual stimuli in order to measure the brain wave. A brain wave data of an event related electrical potential for the research objects which the subject has selected as a target immediately after the stimulus presentation, is processed by a linear discriminant analysis to quantitatively represent with a single index, in order to rank the research objects. | 1. A ranking device for ranking a plurality of research objects based on differences in a degree of a rising attention of a subject to each of the plurality of research objects, configured to: classify target and non-target for brain wave data of event related electrical potentials, wherein the brain wave data is measured by a brain wave monitor with electrodes attached on a head of the subject;calculate a weighting factor of discriminant equation based on data point of time series data and the number of a channel according to a number of measured positions where the brain wave data of the even related electrical potential is acquired when each of the plurality of research objects is the target;calculate a discriminant point for each product as the target based on the discriminant equation; andrank the research objects based on the discriminant point calculated for each product. 2. A ranking system for ranking a plurality of research objects based on difference in a degree of a rising attention of a subject to each of the plurality or research objects, comprising: a stimulus presentation device configured to present stimuli;a brain wave monitor; anda processing device configured to process brain wave data from the brain wave monitor,wherein the stimulus presentation device presents the plurality of research objects as the stimuli a plurality of times,wherein the brain wave monitor measures brain waves immediately after the stimuli are presented by the stimulus presentation device by electrodes attached on the head of the subject,wherein the processing device classifies target and non-target for brain wave data of event related electrical potentials, calculates a weighting factor of discriminant equation based on data point of time series data and the number of a channel according to a number of measured positions where the brain wave data of the event related electrical potential is acquired when each of the plurality of research objects is the target, and calculates a discriminant point for each product as the target based on the discriminant equation, to rank the research objects based on the discriminant point calculated for each product. 3. A ranking method for ranking a plurality of research objects based on differences in the degree of the rising attention of a subject to each research object, comprising: classifying of target and non-target for brain wave data of event related electrical potentials, wherein the brain wave data is measured by a brain wave monitor with electrodes attached on the head of the subject;calculating weighting factor of discriminant equation based on data point of time-series data and the number of the channel according to the number of the measurement positions where the brain wave data of the event related electrical potential is acquired when each research object is the target; andcalculating a discriminant point for each product as the target, to rank the research objects. | National Institute of Advanced Industrial Science and Technology | G06F017/30 A61B005/0476 G06Q030/02 A61B005/0484 G06F019/00
US09782099 | B2 | 2017-10-10 | Basket catheter with improved spine flexibility | 2014-12-31 | A catheter with basket-shaped electrode assembly with spines configured for hyper-flexing in a predetermined, predictable manner when a compressive force acts on the assembly from either its distal end or its proximal end. At least one spine has at least one region of greater (or hyper) flexibility that allows the electrode assembly to deform, for example, compress, for absorbing and dampening excessive force that may otherwise cause damage or injury to tissue wall in contact with the assembly, without compromising the structure and stiffness of the remaining regions of the spine, including its distal and proximal regions. The one or more regions of greater flexibility in the spine allow the spine to flex into a generally V-shape configuration or a generally U-shape configuration. | 1. A catheter comprising: an elongated catheter body having proximal and distal ends and at least one lumen therethrough;a basket-shaped electrode assembly at the distal end of the catheter body, the basket-shaped electrode assembly having proximal and distal ends and comprising a plurality of spines, at least one spine having an equatorial region of greater flexibility and proximal and distal regions of lesser flexibility, the basket-shaped electrode assembly being movable between a collapsed configuration, an expanded configuration in which the spines bow radially outward relative to the catheter body to form a first expanded shape, and a hyper-expanded configuration in which the spines bend at the equatorial region of greater flexibility to thereby deform the first expanded shape, shorten a length of the basket-shaped electrode assembly and assume a V shape with an acute bend in the equatorial region of greater flexibility. 2. The catheter of claim 1, wherein the at least one spine has a thickness dimension and the one equatorial region of greater flexibility has a lesser thickness. 3. The catheter of claim 1, wherein the at least one spine has a width dimension and the equatorial region of greater flexibility has a lesser width. 4. The catheter of claim 1, wherein the equatorial region of greater flexibility includes a notch. 5. The catheter of claim 4, wherein the notch is lateral. 6. The catheter of claim 4, wherein the notch is inner or outer. 7. The catheter of claim 4, wherein the notch has a generally smooth contour. 8. The catheter of claim 4, wherein the notch has a stepped contour. 9. The catheter of claim 1, further comprising at least one tensile member elastically connecting adjacent spines to each other. 10. The catheter of claim 9, wherein the at least one tensile member is affixed to equatorial regions of the adjacent spines. 11. A catheter comprising: an elongated catheter body having proximal and distal ends and at least one lumen therethrough;a basket-shaped electrode assembly at the distal end of the catheter body, the basket-shaped electrode assembly having proximal and distal ends and comprising a plurality of spines, at least one spine having proximal and distal regions of greater flexibility and a mid region of lesser flexibility between the proximal and distal regions of greater flexibility, the basket-shaped electrode assembly being movable between a collapsed configuration, an expanded configuration in which the spines bow radially outward relative to the catheter body to form a first expanded shape, and a hyper-expanded configuration in which the spines bend at the proximal and distal regions of greater flexibility to thereby deform the first expanded shape and assume a U shape with first and second acute bends in the proximal and distal regions of greater flexibility. 12. The catheter of claim 11, wherein the at least one spine has a thickness dimension and each of the proximal and distal regions of grea... | Biosense Webster (Israel) Ltd. | A61B005/042 A61B018/14 A61B005/00 A61B018/00
US09782132 | B2 | 2017-10-10 | Health monitoring systems and methods | 2014-12-09 | Systems, methods and devices for reducing noise in health monitoring including monitoring systems, methods and/or devices receiving a health signal and/or having at least one electrode or sensor for health monitoring. | 1. A device for monitoring a physiological parameter, the device being configured to be adhered to the skin of a subject for the physiological parameter monitoring; the device comprising: a substrate;a conductive sensor connected to the substrate, anda double-sided composite adhesive structure having: at least first and second layers; the first layer having: at least one conductive adhesive portion, andat least one non-conductive adhesive portion, the at least one conductive portion being disposed as at least one island within the at least one non-conductive portion and being surrounded thereby;the second layer being non-conductive and being disposed on the first layer and covering a part of the at least one non-conductive adhesive portion adjacent the at least one conductive adhesive portion and covering some but not all of the at least one conductive adhesive portion;the double-sided composite adhesive structure being attached to the substrate and the conductive sensor; and,the at least one conductive adhesive portion being disposed in conductive communicative contact and connected with the conductive sensor to substantially remove movement of the sensor relative to the conductive adhesive portion, andbeing configured to be securely conductively connected to the skin by being adhered to the skin of the subject to substantially remove movement of the conductive sensor relative to the skin to one or both reduce noise or provide a clean signal and to maintain the sensor substantially fixed relative to the skin and thereby one or both reduces or eliminates possible sensor movement relative to the skin for conductive signal communication from the subject to the conductive sensor. 2. A device according to claim 1 wherein the physiological parameter is acquired by the conductive sensor and includes one or more signals for one or more or all of electrocardiography, photoplethysmography, pulse oximetry, temperature, or subject acceleration. 3. A device according to claim 2 wherein the conductive sensor is an electrode for electrocardiography. 4. A device according to claim 3 wherein the electrode is a proxy driven-right-leg electrode. 5. A device according to claim 1 comprising a plurality of conductive sensors connected to the substrate,the composite adhesive structure having a plurality of conductive adhesive portions,the plurality of conductive adhesive portions being disposed in conductive communicative contact with respective ones of the plurality of conductive sensors, and being adapted to be conductively adhered to the skin of the subject for conductive signal communication from the subject to the conductive sensor. 6. A device according to claim 5 wherein the plurality of conductive sensors include at least one, two or three electrodes for electrocardiography; whereby when one of which is for electrocardiography, the other one or more of the plurality are for one or more other functionalities. 7. A device according to claim 6 wherein one o... | RHYTHM DIAGNOSTIC SYSTEMS, INC. | A61B005/00 A61B005/0205 A61B005/026 A61B005/0295 A61B005/0408 A61B005/11 A61B005/0432 A61B005/0402 A61B005/08 A61B005/021 A61B005/1455 A61B005/024 A61B005/0456
US09820670 | B2 | 2017-11-21 | Methods and apparatus for electrode placement and tracking | 2016-12-21 | An electrode carrier system includes one or more electrode assemblies having an electrode body. One or more tubular members extend from the electrode body and define a lumen terminating in a distal opening. The electrode assemblies carry a reservoir containing a conductive fluid or gel. The reservoir is in fluid communication with the lumens in the tubular members, and the electrode assemblies are typically supported on a backing which may optionally be configured as a headband. Systems are for tracking patient movement may be used in combination with the electrode carrier system. | 1. An electrode assembly, comprising: an electrode body;one or more tubular members extending from the electrode body and having a distal tip, at least some of the one or more tubular members having a lumen with a distal opening in the distal tip; andmeans on the electrode body for dispensing a conductive fluid or gel from a reservoir through the lumen(s) and out of the distal opening(s) of the tubular member(s), wherein the one or more tubular members each have a distribution channel extending across a tissue-contacting surface thereof for receiving the conductive fluid or gel from the lumen so that the conductive fluid or gel may flow through the channel over the patient tissue to form an electrically conductive path, wherein the fluid or gel enters each distribution channel directly from the lumen. 2. The assembly of claim 1 comprising at least two tubular members. 3. The assembly of claim 2 wherein the at least two tubular members are straight. 4. The assembly of claim 3 wherein the at least two tubular members are configured to extend from a bottom of the electrode body. 5. The assembly of claim 4 wherein the at least two tubular members are configured to extend from the generally planar bottom of the electrode body at a generally perpendicular angle in the range relative to the plane. 6. The assembly of claim 1 wherein the distal tip of at least some of the tubular members define a skin preparation surface. 7. The assembly of claim 6 wherein the skin preparation surface comprises at least one of (a) an abrasive and (b) surface features on the distal tip. 8. The assembly of claim 1 wherein the tubular members are formed entirely from an electrically conductive material. 9. The assembly of claim 1 wherein the tubular members are formed at least partly from an electrically non-conductive material. 10. The assembly of claim 1 wherein the tubular member(s) comprise rigid prong(s) formed on a lower base of the electrode body. 11. The assembly of claim 10 wherein the the one or more tubular members each have a single distribution channel. 12. The assembly of claim 11 wherein the lower surface of the prong has at least one skin preparation surface region adjacent to the distal opening. 13. The assembly of claim 11 wherein the lower surface of the prong has a crescent shape and the distribution channel has an arcuate shape. 14. The assembly of claim 13 wherein the crescent-shaped lower surface of the prong has at least one skin preparation surface region adjacent to the arcuate distribution channel. 15. The assembly of claim 11 wherein the lower surface of the prong has a raised portion and a recessed portion, wherein the distal opening is in the recessed portion. 16. The assembly of claim 1 wherein the reservoir is disposed in the electrode body, said reservoir containing a conductive fluid or gel. 17. The assembly of claim 16 wherein the reservoir comprises a sealed dispensing container which is constrained within a chamber in the ... | Ceribell, Inc. | A61B005/0478 A61B005/00 A61B005/0402
US09785748 | B2 | 2017-10-10 | Systems and methods for estimating hemodynamic forces acting on plaque and monitoring patient risk | 2016-06-30 | Computer-implemented methods are disclosed for estimating values of hemodynamic forces acting on plaque or lesions. One method includes: receiving one or more patient-specific parameters of at least a portion of a patient's vasculature that is prone to plaque progression, rupture, or erosion; constructing a patient-specific geometric model of at least a portion of a patient's vasculature that is prone to plaque progression, rupture, or erosion, using the received one or more patient-specific parameters; estimating, using one or more processors, the values of hemodynamic forces at one or more points on the patient-specific geometric model, using the patient-specific parameters and geometric model by measuring, deriving, or obtaining one or more of a pressure gradient and a radius gradient; and outputting the estimated values of hemodynamic forces to an electronic storage medium. Systems and computer readable media for executing these methods are also disclosed. | 1. A computer implemented method for estimating values of hemodynamic forces acting on plaque or lesions, the method comprising: receiving one or more patient-specific parameters of at least a portion of a patient's vasculature that is prone to plaque progression, rupture, or erosion;constructing a patient-specific geometric model of at least a portion of a patient's vasculature that is prone to plaque progression, rupture, or erosion, using the received one or more patient-specific parameters;measuring, deriving, or obtaining a pressure gradient along one or more points of the patient-specific geometric model;identifying one or more plaque progressions or erosions of the patient-specific geometric model;receiving a maximum allowable value of a hemodynamic force on the identified plaque progressions or erosions;estimating, using one or more processors, a value of the hemodynamic force acting on the identified plaque progressions or erosions, using the pressure gradient;assessing a risk posed by each identified plaque progression or erosion, based on a comparison of the maximum allowable value of the hemodynamic force and the estimated value of the hemodynamic force acting on the identified plaque progressions or erosions; andoutputting, the assessed risk to an electronic storage medium or display. 2. The computer implemented method of claim 1, wherein constructing a patient-specific geometric model includes: receiving one or more images of a patient anatomy comprising at least a portion of a vasculature that may be prone to a plaque progression, rupture, or erosion, using, one or more of, computerized tomography,optical coherence tomography,intravascular ultrasound imaging,intravascular pressure-wire,motorized pull-back system,angiogram, ora combination thereof;receiving or updating the one or more patient-specific parameters, using the received one or more images regarding patient anatomy;deriving a patient-specific geometric model, by identifying one or more portions of a patient's vasculature on the received one or more images; andupdating the patient specific geometric model based on the received one or more patient-specific parameters. 3. The computer implemented method of claim 1, wherein estimating a value of a hemodynamic force includes, one or more of: measuring, deriving, or obtaining the pressure gradient from, one or more of, the patient-specific parameters,computerized tomography,optical coherence tomography,intravascular ultrasound imaging,intravascular pressure-wire, orflow simulations of the geometric model; andmeasuring, deriving, or obtaining the axial plaque stress using the pressure gradient. 4. The computer implemented method of claim 1, wherein the patient-specific parameters may include one or more measured, derived, or obtained physiological characteristics, including, one or more of, blood pressure, heart rate, hematocrit level, or a combination thereof. 5. The computer implemented method of claim 1, wherein the patie... | HeartFlow, Inc. | G06G007/50 G06F019/00 A61B005/026 A61B005/02 A61B005/0215
US09808189 | B2 | 2017-11-07 | Concentration-measurement device and concentration-measurement method | 2012-04-06 | A concentration measurement apparatus measures a temporal relative change amount (ΔcHb, ΔO2Hb) of either or both of total hemoglobin concentration and oxygenated hemoglobin concentration in the head that vary due to repetition of chest compression, and includes a light incidence section making measurement light incident on the head, a light detection section detecting the measurement light propagated through the interior of the head and generating a detection signal in accordance with the intensity of the measurement light, and a CPU determining, based on the detection signal, the relative change amount (ΔcHb, ΔO2Hb) and performing a filtering process of removing frequency components less than a predetermined frequency from frequency components contained in the relative change amount (ΔcHb, ΔO2Hb). | 1. A concentration measurement apparatus measuring a temporal relative change amount of at least one of total hemoglobin concentration, oxygenated hemoglobin concentration, and deoxygenated hemoglobin concentration in the head that vary due to repetition of chest compression, comprising: a light source that makes measurement light having at least a first wavelength λ1 and a second wavelength λ2 incident on the head;a light detector that detects the measurement light that has propagated through the interior of the head and generating a first detection signal corresponding to the first wavelength λ1 and a second detection signal corresponding to the second wavelength λ2 in accordance with an intensity of the measurement light; anda processor electrically connected to the light detector and programmed to determine, based on the first detection signal and the second detection signal, the temporal relative change amount of at least one of the total hethoglobin concentration (ΔcHb), the oxygenated hemoglobin concentration (ΔO2Hb), and the deoxygenated hemoglobin concentration (ΔHHb),wherein the processor is programmed to perform a filtering process that removes frequency components less than a predetermined frequency of not more than 1.66 Hz to extract information on the concentration change due to chest compression, and calculate information related to a variation frequency of the relative change amount corresponding to a current frequency of chest compression, andwherein a calculation cycle of the relative change amount is not more than 0.2 seconds. 2. The concentration measurement apparatus according to claim 1, further comprising a display that displays the calculation result obtained by the processor, and wherein the processor is programmed to determine the temporal relative change amount of the total hemoglobin concentration and the temporal relative change amount of the oxygenated hemoglobin concentration, and determines first information related to a ratio (A2/A1) of an amplitude (A1) of the temporal variation of the temporal relative change amount of the total hemoglobin concentration and an amplitude (A2) of the temporal variation of the temporal relative change amount of the oxygenated hemoglobin concentration, after the filtering process, andthe display displays the first information. 3. The concentration measurement apparatus according to claim 1, further comprising a display that displays the calculation result obtained by the processor, and wherein the processor is programmed to determine the temporal relative change amount of the total hemoglobin concentration and the temporal relative change amount of the oxygenated hemoglobin concentration, and determines at least one of an integrated value (I1) of the temporal variation of the temporal relative change amount of the total hemoglobin concentration and an integrated value (I2) of the temporal variation of the temporal relative change amount of the oxygenated hemoglobin concentration, ... | St. Marianna University School of Medicine HAMAMATSU PHOTONICS K.K. | A61B005/1455 A61B005/145 A61B005/00 A61B005/026 A61H031/00 A61B005/0295
US09820666 | B2 | 2017-11-21 | Integrated analysis of electrophysiological data | 2015-02-04 | A method can include analyzing non-invasive electrical data for a region of interest (ROI) of a patient's anatomical structure to identify one or more zones within the ROI that contain at least one mechanism of distinct arrhythmogenic electrical activity. The method also includes analyzing invasive electrical data for a plurality of signals of interest at different spatial sites within each of the identified zones to determine intracardiac signal characteristics for the plurality of sites within each respective zone. The method also includes generating an output that integrates the at least one mechanism of distinct arrhythmogenic electrical activity for the one or more zones with intracardiac signal characteristics for the plurality of sites within each respective zone. | 1. A method comprising: receiving, by a system comprising a processor, non-invasive electrical data from a plurality of electrodes on an external surface of a patient's body, wherein the plurality of electrodes define a region of interest (ROI) of the patient's anatomical structure;analyzing, by the system, the non-invasive electrical data to identify one or more zones within the ROI that contain at least one mechanism of distinct arrhythmogenic electrical activity;analyzing, by the system, invasive electrical data for a plurality of signals of interest at different spatial sites within each of the identified zones to determine intracardiac signal characteristics for the plurality of sites within each respective zone;determining, by the system, a number of mechanisms of distinct arrhythmogenic electrical activity that occur within each respective zone during at least one time interval based on the intracardiac signal characteristics for the plurality of sites within each respective zone, wherein the number of mechanisms of distinct arrhythmogenic electrical activity within each respective zone comprises at least one of a number of rotors that occur within the at least one time interval or a number of foci that occur within the at least one time interval, and a number of fast bursting cycle lengths that occur within the at least one time interval;generating, by the system, an output map that integrates the at least one mechanism of distinct arrhythmogenic electrical activity for the one or more zones and the intracardiac signal characteristics for the plurality of sites within each respective zone; anddisplaying, by a display device of the system, a hybrid graphical map comprising the output map superimposed on a model of the cardiac surface. 2. The method of claim 1, wherein the one or more zones are a plurality of zones, the method further comprises: determining, by the system, a zonal priority among the plurality of zones based on a relative number of the mechanisms of distinct arrhythmogenic electrical activity that occur within in each respective zone during the at least one time interval. 3. The method of claim 2, the method further comprising evaluating, by the system, a plurality of different intracardiac signal characteristics determined for the plurality of sites within each zone to determine a local priority among the plurality of sites within each respective zone. 4. The method of claim 3, wherein the plurality of different intracardiac signal characteristics comprises at least two of a cycle length, cycle length variation, percentage of continuous activation and fractionation determined for at least one measurement interval. 5. The method of claim 4, the method further comprising assigning, by the system, different priority weighting to at least two of the different intracardiac signal characteristics for the plurality of sites within a given zone according to the mechanisms of distinct arrhythmogenic electrical activity identifie... | CARDIOINSIGHT TECHNOLOGIES, INC. | A61B005/044 A61B005/00 A61B005/0402 A61B005/06 A61B005/042
US09808202 | B2 | 2017-11-07 | Mattress for measuring physiological parameters of heart | 2015-11-04 | The present invention provides a mattress for measuring physiological parameters of a heart, comprising a mattress body, a first electrode, a second electrode and an electromagnetic shield, wherein the first electrode and the second electrode are located on the upper surface of the mattress body and the first electrode and the second electrode have a space therebetween; at least a portion of the electromagnetic shield is located on the lower surface of the mattress body and is insulated from the first electrode and the second electrode. In the above mattress provided by the present invention, the first electrode and the second electrode are shielded from ambient electromagnetic radiation effectively. The electromagnetic shield, by blocking the electromagnetic wave propagation path, prevents the electrodes from being polluted by electromagnetic radiation, achieving more accurate measurement. | 1. A mattress for measuring physiological parameters of a heart, comprising a mattress body, a first electrode, a second electrode and an electromagnetic shield, wherein the first electrode and the second electrode are located on an upper surface of the mattress body, and the first electrode and the second electrode have a space therebetween, the first electrode is configured to be placed on an upper position of the heart and the second electrode is configured to be placed on a lower position of the heart, wherein the first and second electrodes have a width between 9-12 cm and a length from 60-85 cm and a distance between the first and second electrodes is between 13-20 cm;the electromagnetic shield, the first electrode and the second electrode comprise a conductive fabric, having a thickness of 1-1.8 mm;at least a portion of the electromagnetic shield is located on a lower surface of the mattress body and is insulated from the first electrode and the second electrode. 2. The mattress according to claim 1, wherein the electromagnetic shield has a width equal to that of the mattress body. 3. The mattress according to claim 2, wherein a portion of the electromagnetic shield located on the lower surface of the mattress body covers the entire lower surface of the mattress body. 4. The mattress according to claim 1, wherein the electromagnetic shield extends from the lower surface of the mattress body along a tail of the mattress body to the upper surface of the mattress body in the width of the electromagnetic shield. 5. The mattress according to claim 4, wherein, on the upper surface of the mattress body, a distance between the electromagnetic shield and the second electrode is from 20 cm to 40 cm. 6. The mattress according to claim 1, wherein the electromagnetic shield extends from the lower surface of the mattress body along a head of the mattress body to the upper surface of the mattress body in the width of the electromagnetic shield. 7. The mattress according to claim 6, wherein, on the upper surface of the mattress body, a distance between the electromagnetic shield and the first electrode is from 12 cm to 25 cm. 8. The mattress according to claim 1, wherein the electromagnetic shield extends from the lower surface of the mattress body along a head and a tail of the mattress body to the upper surface of the mattress body in the width of the electromagnetic shield. 9. The mattress according to claim 8, wherein, on the upper surface of the mattress body, a distance d1 between the electromagnetic shield and the first electrode and a distance d2 between the electromagnetic shield and the second electrode meet the following condition: d2=a*d1, aε[1,2]. 10. The mattress according to claim 8, wherein the electromagnetic shield has an end edge on the upper surface of the mattress body that is parallel to the head of the mattress. 11. The mattress according to claim 8, wherein the electromagnetic shield has a width equal to that of the mattr... | Shenzhen Novocare Medical Devices Co, INC | A61B005/0408 A61B005/00
US09821165 | B2 | 2017-11-21 | Method for symmetry-based implant control | 2016-09-20 | Apparatus comprising (1) a breathing sensor, configured to detect a breathing-related factor of a subject; (2) at least a first electrode configured to be placed in a vicinity of a respective first hypoglossal nerve, and to be driven, in response to the detected breathing-related factor, to apply a first electrical current to the first hypoglossal nerve; (3) at least a second electrode configured to be placed in a vicinity of a respective second hypoglossal nerve, and to be driven, in response to the detected breathing-related factor, to apply a second electrical current to the second hypoglossal nerve; and (4) circuitry configured to, in response to a detected symmetry-related factor indicative of a degree of symmetry of the subject, configure at least one current selected from the group consisting of: the first current and the second current. Other embodiments are also described. | 1. A method for use with a body of a subject, the method comprising: detecting a breathing-related factor of the subject;in response to the detected breathing-related factor: applying a first electrical current to a first hypoglossal nerve of the subject, andapplying a second electrical current to a second hypoglossal nerve of the subject;detecting a symmetry of a response of the subject to the applied first and second currents; andin response to the detected symmetry, configuring at least one current selected from the group consisting of: the first current and the second current. 2. The method according to claim 1, wherein detecting the breathing-related factor comprises extracorporeally detecting the breathing-related factor. 3. The method according to claim 1, wherein: detecting the breathing-related parameter comprises detecting the breathing-related parameter using a breathing sensor;applying the first electrical current comprises applying the first electrical current using a first electrode disposed in a vicinity of the first hypoglossal nerve;applying the second electrical current comprises applying the second electrical current using a second electrode disposed in a vicinity of the second hypoglossal nerve; andconfiguring the at least one selected current comprises configuring the at least one selected current using circuitry configured to configure the at least one selected current in response to the detected symmetry. 4. The method according to claim 3, wherein detecting the symmetry comprises detecting the symmetry using a symmetry sensor. 5. The method according to claim 4, wherein detecting the symmetry comprises detecting a mechanical symmetry of the response of the subject. 6. The method according to claim 4, wherein detecting the symmetry comprises detecting the symmetry using a symmetry sensor that includes an accelerometer. 7. The method according to claim 4, wherein detecting the symmetry comprises detecting an electrical symmetry of the response of the subject. 8. The method according to claim 4, wherein detecting the symmetry comprises detecting the symmetry using a symmetry sensor that includes an electrode. 9. The method according to claim 8, wherein detecting the symmetry comprises detecting the symmetry using a symmetry sensor that includes an electromyographic electrode. 10. The method according to claim 4, wherein: detecting the symmetry comprises detecting the symmetry-related factor using an extracorporeal control unit that includes the symmetry sensor, andthe method further includes transmitting a wireless signal at least in part responsively to the detected symmetry. 11. The method according to claim 3, wherein: the circuitry includes a first circuitry,a first implant includes the first circuitry and the first electrode,applying the first current comprises applying the first current using the first electrode of the first implant,a second implant includes the second electrode and second circuitry, andapplyi... | BLUEWIND MEDICAL LTD. | A61N001/05 A61N001/36 A61N001/375 A61N001/378 A61B005/08 A61N001/372 A61B005/0488 A61B005/07 A61B005/00 A61B005/113
US09814440 | B2 | 2017-11-14 | Measurement apparatus and a method for temperature dependent frequency change measurement in a body part via ultrasonic measurements | 2015-02-13 | To measure the content percentage of fat in a liver without increasing hardware size, provided is a measurement apparatus that supplies an ultrasonic signal to a body part and receives a response signal that has passed through an inner portion of the body part, the measurement apparatus including an ultrasonic signal generating section that generates the ultrasonic signal according to a control signal; a loop control section that supplies the control signal to the ultrasonic signal generating section in response to receiving the response signal that has passed through the body part; and a frequency measuring section that measures a repeating frequency of the control signal repeatedly supplied by the loop control section. | 1. A measurement apparatus comprising: an ultrasonic signal generator that generates an ultrasonic signal according to a control signal and supplies the ultrasonic signal to a body part;a loop controller that supplies the control signal to the ultrasonic signal generator in response to receiving a response signal that has passed through the body part;a frequency counter that measures a repeating frequency of the control signal repeatedly supplied by the loop controller; anda processor configured to detect change in the frequency measured by the frequency counter;whereinthe ultrasonic signal generator supplies the ultrasonic signal to a portion containing a liver, which is the body part,the ultrasonic signal generator, the body part, and the loop controller form a loop path through which a signal based on the ultrasonic signal is sequentially transferred,the frequency counter measures a loop frequency at which the signal travels around the loop path, andthe processor is further configured todetect physical matter contained in the body part based on the detected change in the frequency,detect a content percentage of fat in the liver based on a result of the detection,detect the physical matter contained in the body part based on change in the frequency corresponding to temperature change of the inner portion of the body part anddetect the content percentage of fat in the body part based on a slope of the change in the frequency corresponding to the temperature change. 2. The measurement apparatus according to claim 1, wherein the processor is further configured to identify the liver as a fatty liver when the content percentage of fat in the liver is greater than or equal to a predetermined value. 3. The measurement apparatus according to claim 1, comprising: a response signal passing section that is provided between the ultrasonic signal generator and the loop controller, and includes an amplification circuit that amplifies or attenuates a signal input thereto. 4. The measurement apparatus according to claim 1, wherein the response signal is a reflected signal that has been reflected from the inner portion of the body part. 5. A measurement apparatus comprising: an ultrasonic signal generator that generates an ultrasonic signal according to a control signal and supplies the ultrasonic signal to a body part;a loop controller that supplies the control signal to the ultrasonic signal generator in response to receiving a response signal that has passed through the body part;a frequency counter that measures a repeating frequency of the control signal repeatedly supplied by the loop controller;a processor configured to detect change in the frequency measured by the frequency counter; anda response signal passing section that is provided between the ultrasonic signal generator and the loop controller, passes the response signal to the loop controller, and attenuates the ultrasonic signal,whereinthe ultrasonic signal generator, the body part, and the ... | ADVANTEST CORPORATION | A61B008/08 A61B005/00
US09826954 | B2 | 2017-11-28 | Diagnostic for in situ deformation and strain measurements applicable to traumatic internal injury investigation and prevention | 2015-02-04 | A diagnostic gage (12) that can be implemented into a tissue-simulating headform (17) or other anthropomorphic surrogate test device (11) as a means of determining the internal strain within the test surrogate. One embodiment of the gage consists of a matrix or substrate embedded with x-ray contrast agents (14) and a series of holes within the substrate (15) that provide contrasting markers in an x-ray image and a means of closely coupling the gage to the test specimen. The relative motion of these contrasting markers can be monitored using x-ray fluoroscopy equipment (e.g., source (10) and detector (13)). This gage provides a means of determining the internal strain within a headform surrogate model for the purpose of evaluating the performance of helmets in terms of reducing the occurrence of concussion among other biomechanical injuries from trauma. | 1. A diagnostic gage device involving a plurality of radiopaque markers enabling a full-field, time-resolved measurement of dynamic deformation and strain within an internal plane of a specimen using x-ray visualization equipment for the purpose of predicting, evaluating or preventing a traumatic biomechanical injury. 2. The see device according to claim 1, further comprising an x-ray contrast agent that is deposited onto or within a substrate or matrix material layer that is mechanically similar, having density and/or elastic Young's modulus not to exceed double the value of these same properties in material constituents of said specimen, in a manner that creates a plurality of high contrast markers as seen in an x-ray image, whereby said markers and substrate are mechanically coupled to said specimen. 3. The gage device according to claim 2, wherein the x-ray contrast agent further comprises any single or multitude of materials or coating comprising, at a minimum, one chemical element with an atomic number higher than the atomic number of each of the chemical elements within the materials used to create said specimen. 4. The gage device according to claim 1, wherein the specimen further comprises an anthropomorphic surrogate model of a biological system, body or body part that are used to predict, evaluate or prevent a biomechanical injury. 5. The gage device according to claim 1, wherein the specimen further comprises any one or combination of a multitude of protective barriers, such as a helmet or external padding that are designed to protect the specimen from trauma. 6. The gage device according to claim 5, wherein the trauma further comprises a direct impact to the specimen, both in a protected or unprotected embodiment, or the result of shock wave, blast wave or other dynamic or static pressure loading of the specimen. 7. The gage device according to claim 5, wherein the trauma further comprises a rapid acceleration or deceleration of the specimen. 8. The gage device according to claim 1, wherein the x-ray visualization equipment further comprises a single or multitude of x-ray sources that are in continuous, discrete or multi-anode source configurations as well as a single or multitude of x-ray digital or analog detectors, fluoroscopes, or film that are used to record the relative movement of the x-ray contras agent markers. 9. The gage device according to claim 1, wherein the x-ray visualization equipment further comprises an arrangement in a single line of sight or stereoscopic orientation that would enable the visualization of the x-ray contrast agents in a planar or in complex curved planes and other three-dimensional orientations respectively. 10. The gage device according to claim 2, wherein the x-ray contrast agent is introduced into said specimen as a series of discrete particles or powder grains, wherein the contrast agent is placed onto or within a substrate or matrix layer that are integrated into the specimen, onto or... | - | A61B005/107 A61B006/12 A61B006/00 A61B005/11 A42B003/06 G01B015/06 G01N023/04 G09B023/30
US09808260 | B2 | 2017-11-07 | Noninvasive protection from emboli | 2015-05-04 | A system for diverting emboli within a patient can include a device that detects a presence of emboli in a first blood vessel of a patient. A compression member can be aligned with a second blood vessel of the patient when a collar supporting the compression member is positioned at least partially around a portion of the patient. In response to the detection device, a controller can actuate the compression member, when the presence of emboli is detected, from an unactuated state to an actuated state in which at least a portion of the compression member (i) is closer to the second blood vessel than while in the unactuated state and (ii) compresses and limits blood flow through the second blood vessel. | 1. A method for diverting emboli within a patient, comprising: detecting a presence of emboli in a first vascular location of a patient;when the presence of emboli is detected, actuating a compression member aligned with a second vascular location of the patient when a collar supporting the compression member is positioned at least partially around a portion of the patient, from an unactuated state to an actuated state in which at least a portion of the compression member (i) is closer to the second vascular location than while in the unactuated state and (ii) compresses and limits blood flow through the second vascular location. 2. The method of claim 1, wherein the second vascular location is downstream of the first vascular location. 3. The method of claim 1, further comprising: determining whether the detected emboli exceed a threshold amount;wherein the compression member is actuated when the detected emboli exceed the threshold amount. 4. The method of claim 1, further comprising: determining whether the detected emboli do not exceed a threshold amount;transitioning the compression member to the unactuated state when the detected emboli do not exceed the threshold amount. 5. The method of claim 1, further comprising transitioning the compression member to the unactuated state after the compression member has been actuated for a predetermined time limit. 6. The method of claim 1, wherein the detecting is by a Doppler ultrasound device, a Doppler probe device, an oscillotonometry device, an electroencephalography device, a transcranial Doppler device, or a cerebral oximetry device. 7. The method of claim 1, wherein the first vascular location is a heart valve, an aorta, a calf vein, a femoral vein, a popliteal vein, an iliofemoral vein, an iliac vein, an inferior vena cava, or a peripheral vein of the patient. 8. The method of claim 1, wherein the second vascular location is a carotid artery or a vertebral artery of the patient. 9. A method of diverting emboli from cerebral circulation of a patient, comprising: while a device is positioned around a neck of a patient, expanding a first carotid compression member of the device toward a first carotid artery of the patient to compress and limit blood flow through the first carotid artery;while the device is positioned around the neck, expanding a first vertebral compression member of the device toward a first vertebral artery of the patient to compress and limit blood flow through the first vertebral artery. 10. The method of claim 9, wherein expanding the first carotid compression member comprises radially advancing a portion of the first carotid compression member toward a cervical spine of the patient. 11. The method of claim 9, wherein expanding the first vertebral compression member comprises radially advancing a portion of the first vertebral compression member toward a cervical spine of the patient. 12. The method of claim 9, further comprising: while the device is positioned ar... | ZHADKEVICH MEDICAL, INC. | A61B017/00 A61B017/135 A61B017/132 A61B005/02 A61B005/026 A61B005/022 A61H009/00 A61H001/00 A61B005/00
US09775544 | B2 | 2017-10-03 | Analyte sensors having a membrane with low temperature sensitivity | 2016-11-09 | Embodiments of the present disclosure relate to analyte determining methods and devices (e.g., electrochemical analyte monitoring systems) that have a membrane with low temperature sensitivity. The sensing layer is disposed on a working electrode of in vivo and/or in vitro analyte sensors, e.g., continuous and/or automatic in vivo monitoring using analyte sensors and/or test strips. Also provided are systems and methods of using the, for example electrochemical, analyte sensors in analyte monitoring. | 1. A system comprising: an analyte sensor comprising: a counter electrode;a working electrode comprising a sensing layer thereon, the sensing layer comprising an analyte responsive enzyme; anda diffusion limiting polymeric membrane disposed on the sensing layer, wherein the polymeric membrane comprises a poly(4-vinylpyridine)-co-polystyrene polymer, an ethylene oxide substituent and a propylene oxide substituent,wherein diffusion of an analyte through the diffusion limiting membrane to the sensing layer does not change more than 5% per ° C. over a range of physiological temperatures; andan electrochemical sensor control unit comprising: a housing; anda transmitter for transmitting sensor signals from the electrochemical sensor. 2. The system of claim 1, wherein the sensor control unit further comprises a data storage unit for storing data from the electrochemical sensor. 3. The system of claim 1, wherein the sensor control unit further comprises a reference voltage generator. 4. The system of claim 1, wherein the sensing layer comprises a glucose-responsive enzyme. 5. The system of claim 4, wherein the glucose-responsive enzyme comprises glucose oxidase. 6. The system of claim 1, wherein the sensing layer further comprises a redox mediator. 7. The system of claim 6, wherein the redox mediator comprises a ruthenium-containing complex or an osmium-containing complex. 8. The system of claim 6, wherein one or more of the redox mediator and the enzyme is covalently bound to a polymer. 9. The system of claim 8, wherein the analyte sensor is an in vivo glucose sensor. 10. The system of claim 8, wherein the poly(4-vinylpyridine)-co-polystyrene polymer comprises a compound of the formula:wherein x and y are each positive integers and the ratio of x to y is 1:1. 11. The system of claim 1, wherein the ethylene oxide substituent and a propylene oxide substituent comprises a formula:whereinm, n and o are each independently positive integers,R′ and R″ are each a linking group, andEO is ethylene oxide and PO is propylene oxide. 12. The system of claim 1, wherein the polymeric membrane comprises a compound of the formula:whereinx and y are each positive integers,m, n and o are each independently positive integers,R′ and R″ are each a linking group, andEO is ethylene oxide and PO is propylene oxide. | Abbott Diabetes Care Inc. | A61B005/05 A61B005/145 A61B005/1486 A61M005/172 A61L031/04 A61L031/06
US09827039 | B2 | 2017-11-28 | Methods and apparatuses for remodeling tissue of or adjacent to a body passage | 2014-03-11 | Medical devices and methods for making and using the same are disclosed. An medical device may include a medical device for renal nerve ablation. The medical device may include a catheter shaft having a distal region. An expandable balloon may be coupled to the distal region. An electrode assembly may be coupled to the balloon. The electrode assembly may include a first electrode pad including one or more electrodes. The first electrode pad may have a first lead-in edge, a first protruding edge, and a first transition region with a continuously changing curvature disposed between the first lead-in edge and the first protruding edge. | 1. A medical device for renal nerve ablation, comprising: a catheter shaft having a distal region;an expandable balloon coupled to the distal region;an electrode assembly coupled to the balloon, the electrode assembly including a first electrode pad including one or more electrodes and a first tail positioned proximal to the first electrode pad;wherein the first electrode pad comprises a base layer and a sensor is positioned within an opening in the base layer;wherein the electrode is positioned over the base layer;wherein the electrode assembly has a first edge comprising a first lead-in edge on the first tail, a first protruding edge on the first electrode pad, and a first transition region disposed between the first lead-in edge and the first protruding edge; andwherein there is a first continuously changing curvature along said first edge within the first transition region. 2. The medical device of claim 1, wherein the electrode assembly includes a second electrode pad positioned proximal to the first tail and second tail positioned proximal to the second electrode pad. 3. The medical device of claim 2, wherein the second electrode pad is circumferentially offset from the first electrode pad. 4. The medical device of claim 3, wherein the electrode assembly comprise a central axis, wherein the first electrode pad is asymmetrically disposed to one side of the central axis, and wherein the second electrode pad is asymmetrically disposed to an opposite side of the central axis. 5. The medical device of claim 2, wherein the electrode assembly has a second edge comprising a second lead-in edge on the second tail, a second protruding edge on the second electrode pad, and a second transition region disposed between the second lead-in edge and the second protruding edge; andwherein there is a second continuously changing curvature along said second edge within the second transition region. 6. The medical device of claim 1, wherein the first electrode pad includes one or more pairs of bipolar electrodes. 7. The medical device of claim 1, further comprising one or more additional electrode assemblies coupled to the balloon. 8. The medical device of claim 1, wherein the first continuously changing curvature has a second derivative that is constant along the first transition region. 9. The medical device of claim 1, wherein the first transition region has a first curvature transition point where the first continuously changing curvature changes directions, and wherein the first continuously changing curvature has a second derivative that is a mathematically continuous function at the first curvature transition point. 10. The medical device of claim 1, wherein the first continuously changing curvature is defined by a mathematically continuous function. 11. The medical device of claim 10, wherein the mathematically continuous function includes a parabolic function. 12. A medical device for renal nerve ablation, comprising: a catheter shaft having... | Boston Scientific Scimed, Inc. | A61B018/12 A61B018/14 A61M025/10 A61N001/36 A61B018/00 A61B005/00 A61N001/05 A61B090/00
US09823312 | B2 | 2017-11-21 | Apparatus and method for indirectly cooling superconducting quantum interference device | 2015-04-21 | An apparatus and a method for indirectly cooling a superconducting quantum interference device (SQUID) are provided. The apparatus includes an outer container extending in a vertical direction; a metallic inner container inserted into the outer container to store a liquid coolant, the metal inner container including a top plate; a SQUID sensor module disposed between a bottom surface of the outer container and a bottom surface of the inner container; a heat transfer pillar adapted to cool the SQUID sensor module, the heat transfer pillar having one end connected to the bottom surface of the inner container and the other end directly or indirectly connected to the SQUID sensor module; a magnetic shield part formed of a superconductor covering a top surface of the SQUID sensor module; and a heat conduction plate being in thermal contact with the other end of the heat transfer pillar. | 1. An apparatus for indirectly cooling a superconducting quantum interference device (SQUID), comprising: an outer container extending in a vertical direction;a metallic inner container inserted into the outer container to store a liquid coolant, the metal inner container including a top plate;a SQUID sensor module disposed between a bottom surface of the outer container and a bottom surface of the inner container;a heat transfer pillar adapted to cool the SQUID sensor module, the heat transfer pillar having one end connected to the bottom surface of the inner container and the other end directly or indirectly connected to the SQUID sensor module;a magnetic shield part formed of a superconductor covering a top surface of the SQUID sensor module; anda heat conduction plate being in thermal contact with the other end of the heat transfer pillar, the heat conduction plate being stacked on a top surface of the magnetic shield part,wherein the inner container has a space spaced apart from the outer container and a space between the inner container and the outer container is in a vacuum state. 2. The apparatus as set forth in claim 1, further comprising: an auxiliary heat transfer layer disposed to a bottom surface of the SQUID sensor module. 3. The apparatus as set forth in claim 2, wherein the auxiliary heat transfer layer is formed of a copper mesh, and wherein wefts and warps of the copper mesh are electrically insulated from each other. 4. The apparatus as set forth in claim 2, wherein the auxiliary heat transfer layer comprises: an epoxy layer; andcopper conductor lines stacked on the epoxy layer and extending parallel with each other at regular intervals. 5. The apparatus as set forth in claim 2, further comprising: a fixing ring formed of a conductor to fix the auxiliary heat transfer layer to the bottom surface of the inner container. 6. The apparatus as set forth in claim 1, further comprising: a fixing block formed of a copper material and fixed to the bottom surface of the inner container to transfer heat to the heat transfer pillar. 7. The apparatus as set forth in claim 1, wherein the outer container comprises: an upper outer cylinder;a washer-shaped outer middle plate connected to a bottom surface of the upper outer cylinder;a lower outer cylinder being in contact with an inner side surface of the outer middle plate to vertically extend; andan outer base plate disposed on a bottom surface of the lower outer cylinder. 8. The apparatus as set forth in claim 1, wherein the inner container comprises: a neck portion connected to an inner side surface of a through-hole formed in the center of a top plate, the neck portion extending vertically;a washer-shaped inner middle plate connected to a bottom surface of the neck portion, the inner middle plate having a through-hole in its center;a cylindrical inner body portion extending vertically from an outer side surface of the inner middle plate; andan inner base plate disposed on a bottom su... | Korea Research Institute of Standards and Science | F28F007/00 G01R033/00 A61B005/05 G01R033/035 G01R001/18 H01L023/367 H01L023/473 H01L039/02 A61B005/04
US09782585 | B2 | 2017-10-10 | Method and system for providing electrical stimulation to a user | 2016-03-02 | A method for providing electrical stimulation to a user, the method comprising: providing an electrical stimulation device, in communication with a controller, at a head region of the user; with the electrical stimulation device, providing a stimulation treatment having a waveform configured for neuromodulation in the user; with the controller, performing an adjustment to the stimulation treatment, wherein performing the adjustment includes: generating a transformed waveform with application of a transfer function to the waveform, wherein the transfer function scales the waveform and selectively attenuates extreme waveform values while maintaining a frequency characteristic of the waveform in the transformed waveform; and applying the transformed waveform with the electrical stimulation device, thereby modulating the stimulation treatment. | 1. A method for providing electrical stimulation to a user, the method comprising: providing an electrical stimulation device, in communication with a controller, at a head region of the user;with the electrical stimulation device, providing a stimulation treatment having a waveform configured for neuromodulation in the user; wherein the waveform comprises a superposed waveform including superposition of two or more of: delta-band, theta-band, alpha-band, gamma-band, and beta-band components;at the controller, receiving an adjustment to the stimulation treatment, by the user;in near-real time and in response to the adjustment, generating a transformed waveform with application of a transfer function to the waveform, wherein the transfer function scales the waveform and selectively attenuates extreme waveform values while maintaining a characteristic of the waveform in the transformed waveform; andapplying the transformed waveform with the electrical stimulation device, thereby modulating the stimulation treatment in near-real time. 2. The method of claim 1, wherein the waveform comprises a band-limited random noise waveform. 3. The method of claim 2, wherein the band-limited random noise waveform has a frequency band within the range from: 50-700 Hz. 4. The method of claim 1, wherein the waveform comprises a superposed waveform including superposition of theta-band and gamma-band frequency components. 5. The method of claim 1, wherein the waveform includes a low-amplitude portion having a first duration and a high-amplitude portion having a second duration. 6. The method of claim 5, wherein the low-amplitude portion is a segment of a first sinusoidal waveform and the high-amplitude portion is a segment of a second sinusoidal waveform, wherein a P-P value of the second sinusoidal waveform is greater than an RMS value of the first sinusoidal waveform. 7. The method of claim 1, wherein application of the transfer function includes application of a sigmoidal transfer function to the waveform that scales the waveform and selectively attenuates extreme values and transients in the waveform. 8. The method of claim 7, wherein the sigmoidal transfer function includes one or more of: an error function, a hyperbolic tangent function, a Gudermannian function, an inverse tangent function, and a logarithmic function. 9. The method of claim 1, wherein application of the transfer function includes application of a compound transfer function that scales the waveform, applies a sigmoidal transformation to the waveform, and then rescales extreme values of the waveform. 10. The method of claim 9, wherein the compound transfer function has the form Y=C1·erf(C2·X), where Y is an output, C1 is a current limit of the electrical stimulation device, erf is an error function, C2 is an initial scaling factor, and X is an output of the transfer function. 11. The method of claim 1, wherein receiving the adjustment includes receiving an input indicative of a desire t... | Halo Neuro, Inc. | A61N001/05 A61N001/36 A61B005/16 A61B005/0205 A61B005/0476 A61B005/0488 A61B005/0496 A61B005/00 A61N001/04
US09792411 | B2 | 2017-10-17 | Poroelastic dynamic mechanical analyzer for determining mechanical properties of biological materials | 2013-03-14 | A system for determining parameters of porous media or material, which in an embodiment is biological tissue, includes an actuator and a displacement monitor. The actuator is adapted to apply a displacement to tissue at a particular frequency selected from a range of frequencies, and the force monitor adapted to monitor a mechanical response of tissue. The system also has a processor coupled to drive the actuator and to read the mechanical response, the processor coupled to execute from memory a poroelastic model of mechanical properties of the material, and a convergence procedure for determining parameters for the poroelastic model such that the model predicts mechanical response of the tissue to within limits. | 1. A system for determining poroelastic mechanical properties of porous materials comprising: a dynamic mechanical analyzer comprising an actuator and a force measuring device, the actuator adapted to apply a displacement in a first axis to the porous materials over a range of frequencies, and the force measuring device is adapted to measure a mechanical response of the porous materials, wherein the actuator is configured to apply the displacement to the porous materials through a first nonslip surface, the system configured to hold the porous materials on a second nonslip surface; at least one processor coupled to drive the actuator to apply the displacement to the porous materials and to read the mechanical response of the porous materials, the at least one processor also being configured to execute machine-readable instructions of a poroelastic material model recorded in a memory; the memory also containing limits;the memory further containing machine-readable instructions which are implemented on the at least one processor, to cause the at least one processor to perform a convergence procedure configured to determine parameters for the poroelastic material model such that the porous material model, when executed using the determined parameters to simulate a mechanical response, produces an estimated porous materials response that matches the measured mechanical response of the porous materials to within the limits. 2. The system of claim 1 further comprising a magnetic resonance elastography (MRE) system, and wherein the determined parameters are used to validate material property values acquired by the MRE system. 3. The system of claim 1 wherein the dynamic mechanical analyzer is configured to measure the response of the porous material over a plurality of frequencies, and wherein the machine-readable instructions of the convergence procedure use the response over that plurality of frequencies to determine the model parameters. 4. The system of claim 1 wherein the determined parameters are provided to a computer model of displacement of tumor tissue during neurosurgery. 5. A system for modeling mechanical responses of tissue, comprising the system of claim 4, the processor configured to execute a second mechanical model of in-vivo tissue incorporating a modeled location of a tumor, the processor configured to execute the second mechanical model, and the second mechanical model is configured with the determined parameters. 6. The system of claim 1 wherein the poroelastic model is derived from Biot's theory of consolidation implementing the equations:∇·μ⁡(∇u_+∇u_T)+∇(λ⁢∇·u_)-(1-β)⁢∇p_=-ω2⁡(ρ-βρf)⁢u_i⁢⁢ω⁡(∇·u_)-∇·q_=0where⁢⁢β=ωϕ2⁢ρf⁢κi⁢⁢ϕ2+k⁢⁢ω⁡(ρa+ϕρf)⁢⁢and⁢⁢q_=-κ⁢⁢i⁢⁢ϕ2⁡(∇p-ω2⁢ρf⁢u_)i⁢⁢ϕ2+κω⁡(ρa+ϕρf)and where, u is displacement, p is pore-pressure, μ is shear modulus, λ is compressional modulus, φ is porosity, κ is a hydraulic conductivity, ω is a vibration frequency, ρf is a fluid density, and ρa is an apparent mass density, β is a com... | THE TRUSTEES OF DARTMOUTH COLLEGE | A61B005/103 A61B005/117 G06F019/00 G01N003/32 A61B005/00 A61B005/055 G01R033/563 A61B005/11
US09814401 | B2 | 2017-11-14 | Angiosome-based perfusion monitoring system | 2013-06-27 | A compression device includes at least one pressurizable bladder to substantially occlude blood flow into skin capillary beds adjacent to the at least one pressurizable bladder, and a plurality of perfusion sensors. In operation a first-angiosome sensor detects the perfusion parameter of a skin capillary bed in a first angiosome of the limb, and a second-angiosome sensor detects the perfusion parameter of a skin capillary bed in a second angiosome of the limb that is different from the first angiosome. A control circuit maps sensor signals from the first-angiosome sensor to the first angiosome or a first artery of the limb, and maps sensor signals from the second-angiosome sensor to the second angiosome or a second artery of the limb different from the first artery of the limb. For each perfusion sensor, the control circuit determines whether the received sensor signals are indicative of peripheral artery disease. | 1. An angiosome-based perfusion monitoring system for peripheral artery disease, the monitoring system comprising: a compression device sized and shaped for placement on a limb of a subject, the compression device including: a plurality of pressurizable bladders configured to exert a suitable compressive force on the limb of the subject when pressurized to substantially occlude blood flow into skin capillary beds adjacent to the plurality of pressurizable bladders, anda plurality of spaced apart perfusion sensors, wherein each sensor of the plurality of perfusion sensors is located generally adjacent at least one bladder of the plurality of pressurizable bladders, each perfusion sensor configured to detect a perfusion parameter of skin capillary beds adjacent the perfusion sensor for quantifying skin capillary bed perfusion and generate a signal indicative of the perfusion parameter,wherein at least one sensor of the plurality of perfusion sensors is a first-angiosome sensor for detecting the perfusion parameter of a skin capillary bed in a first angiosome of the limb and at least one sensor of the plurality of perfusion sensors is a second-angiosome sensor for detecting the perfusion parameter of a skin capillary bed in a second angiosome of the limb that is different from the first angiosome, the first-angiosome sensor located generally adjacent a first bladder of the plurality of pressurizable bladders and the second-angiosome sensor located generally adjacent a second bladder of the plurality of pressurizable bladders that is different from the first bladder. 2. The monitoring system of claim 1, wherein the first-angiosome sensor comprises a plurality of first-angiosome sensors, and wherein the second-angiosome sensor comprises a plurality of second-angiosome sensors. 3. The monitoring system of claim 2, wherein the first-angiosome sensors are arranged in a first row extending across the first bladder, and wherein the second-angiosome sensors are arranged in a second row, different from the first row, extending across the second bladder. 4. The monitoring system of claim 3, wherein the first and second rows are spaced apart along a length of the compression device. 5. The monitoring system of claim 2, wherein the first-angiosome sensors and the second-angiosome sensors are arranged in a same row extending across the first bladder and the second bladder. 6. The monitoring system of claim 1, wherein the plurality of pressurizable bladders are spaced apart along a length of the compression device, and wherein the compression device is sized and shaped for placement on a leg of the subject. 7. The monitoring system of claim 1, wherein the perfusion sensors comprise pulse oximetry sensors. 8. The monitoring system of claim 1, in combination with a control device, the control device comprising: a source of pressurized fluid for introducing into the compression device that is placed on the limb of the subject; anda control circuit configured ... | Covidien LP | A61B005/026 A61B005/0295 A61B005/0205 A61B005/02 A61B005/1455 A61B005/00
US09788820 | B2 | 2017-10-17 | Methods and devices for cervical cell and tissue sampling | 2012-11-06 | The present invention comprises methods and devices comprising a cutting element used to contact circumferentially and longitudinally the surface walls of the endocervical canal to provide for broad and complete contact of the intended surface with the device, resulting in attainment of a sufficient volume and comprehensive tissue sample for analysis as an endocervical curettage or screening pap smear. The device may provide for a reservoir for the obtained sample to be contained when removing from the cervical canal and vagina. The device may be comprised of a detachable means or be of a material that allows the user to cut the sampling head of the device for placement in collection means to maximize the amount of tissue or cells being sent to the laboratory for analysis. | 1. A method of obtaining an endocervical sample, comprising: a) introducing into the endocervical canal of a female animal at least a portion of a helical slit in a distal portion of a cervical sampling apparatus comprising, i. a handle having a distal end and a longitudinal axis;ii. an elongate wand extending outwardly from the distal end of the handle substantially along the longitudinal axis, wherein the wand has an exterior surface, a front end that defines a conical tip having a proximal and a distal end, and a back end, wherein the back end is fixedly mounted to the distal end of the handle;iii. a selectively moveable actuator coupled to the handle and to a proximal end of a sheath; wherein the sheath encapsulates a portion of the wand and is fixedly mounted to the proximal end of the tip, and to a portion of an actuator, wherein the slit is located in a distal portion of the sheath and is selectively movable between a closed position, in which opposing edges of the helical slit substantially adjoin, and an open position, in which the opposing edges of the slit are spaced from each other; and wherein the actuator is selectively movable between a first position, in which the slit is positioned in the closed position, and a second position, in which the helical slit is positioned in the open position;b) moving the actuator to the second position to selectively move the slit to the open position;c) selectively urging the open positioned slit against tissue in the endocervical cavity, while simultaneously rotating the handle of the cervical sampling apparatus to affect corresponding rotation of the open positioned slit, to collect tissue and cells into a sample collection cavity;d) moving the actuator to the first position to selectively move the slit to the closed position to selectively close the sample collection cavity; ande) withdrawing the cervical sampling apparatus from the subject. 2. The method of claim 1, wherein the step of rotating the handle of the cervical sampling apparatus includes rotating the handle at least one complete revolution in a clockwise or counterclockwise direction. 3. The method of claim 1, wherein the method is performed to obtain cells for diagnosing cervical lesions. 4. The method of claim 1, wherein the method is performed to monitor the post-surgical status of the endocervix and the effectiveness of surgical removal of cervical lesions. 5. The method of claim 1, wherein the method is performed to diagnose cervical sources of undiagnosed bleeding. 6. A cervical sampling apparatus, comprising: a) a handle having a distal end and a longitudinal axis;b) an elongate wand extending outwardly from the distal end of the handle substantially along the longitudinal axis, wherein the wand has an exterior surface, a front end that defines a conical tip having a proximal and a distal end, and a back end, and wherein the back end is fixedly mounted to the distal end of the handle;c) a selectively moveable actuator cou... | FEMASYS INC | A61B010/04 A61B005/00 A61B010/02 A61B017/3207
US09775525 | B2 | 2017-10-03 | Concentration presence/absence determining device and content evaluation apparatus | 2012-12-17 | A concentration presence/absence determining device includes: a brain blood flow volume obtaining unit configured to obtain a brain blood flow volume of the user; a varying threshold obtaining unit configured to obtain a varying threshold that is a value smaller than the brain blood flow volume of the user at a base time; and an determining unit configured to determine that the user concentrates, when the brain blood flow volume obtained by the brain blood flow volume obtaining unit falls below the varying threshold obtained by the varying threshold obtaining unit. | 1. A concentration presence/absence determining device that determines whether or not a user is concentrating, the concentration presence/absence determining device comprising: a near-infrared spectroscopy sensor which senses a plurality of values of a hemoglobin state in brain blood of the user;a processor; anda non-transitory computer-readable recording medium having stored thereon executable instructions, which when executed by the processor, cause the concentration presence/absence determining device to:obtain the plurality of values of the hemoglobin state in the brain blood of the user from the near-infrared spectroscopy sensor;calculate a varying threshold;obtain a value of the varying threshold at a base time;determine that the user is concentrating, when a first value of the plurality of values of the hemoglobin state in the brain blood of the user falls below the value of the varying threshold;evaluate, based on the determination that the user is concentrating, a degree of concentration of the user; andpresent a result of the evaluated degree of concentration of the user, whereinthe varying threshold is calculated using repeated loop processing, by increasing a reference time between a start time and an end time,the obtained value of the varying threshold is a negative maximum value of the calculated varying threshold, based on a result of the repeated loop processing, andthe result of the evaluated degree of concentration of the user is based on a comparison of a characteristic of the plurality of values of the hemoglobin state in the brain blood of the user with a characteristic of a task executed by the user. 2. The concentration presence/absence determining device according to claim 1, wherein the executable instructions, when executed by the processor, cause the concentration presence/absence determining device to further:calculate, as a decrease start time, a time at which the first value of the hemoglobin state in the brain blood of the user falls below the value of the varying threshold;calculate, as a recovery time, a time at which a second value of the plurality of values of the hemoglobin state in the brain blood of the user exceeds the value of the varying threshold after the calculated decrease start time;calculate, as an increase start time, a time at which a group of values of the plurality of values of the hemoglobin state in the brain blood of the user start to increase between the calculated decrease start time and the calculated recovery time; andcalculate, as a concentration section, a period in which the user is concentrating, the period ranging from the calculated decrease start time to the calculated increase start time. 3. The concentration presence/absence determining device according to claim 2, wherein the executable instructions, when executed by the processor, cause the concentration presence/absence determining device to further:calculate, as the degree of concentration of the user, a ratio of the concent... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | A61B005/1455 A61B005/0205 A61B005/026 A61B005/00 A61B005/16
US09775534 | B2 | 2017-10-03 | Biological signal management | 2013-05-29 | Systems and techniques for managing biological signals. In one implementation, a method includes receiving a cardiac biological signal that includes information describing events, determining a merit of each event based on one or more of a severity of a cardiac condition associated with the event and a quality of the event, and handling a subset of the events that meet a merit criterion. The subset can be handled for medical purposes. | 1. A system comprising: a memory comprising: a plurality of data structures, each data structure comprising: a span identification information field which identifies a time span, the time span having a duration, andan allocation information field describing a cardiac event allocated to the data structure, wherein the described cardiac event occurred within the time span identified in the span identification field; andcircuitry configured to perform a method comprising: receiving a cardiac biological signal that includes information describing a first cardiac event relevant to a medical purpose, wherein the first cardiac event occurred within a first time span,determining a first measure of medical relevance of the information describing the first cardiac event for the medical purpose,determining whether the first measure of medical relevance is greater than a threshold level of medical relevance, wherein the threshold level of medical relevance is a level of medical relevance of a second cardiac event currently allocated to a first data structure of the plurality of data structures, wherein a span identification field of the first data structure identifies the first time span, andbased on determining that the first measure of medical relevance is greater than the threshold level of medical relevance, allocating the first cardiac event to the first data structure. 2. The system of claim 1, wherein each data structure describes a single cardiac event. 3. The system of claim 1, wherein each of the plurality of data structures identifies a unique time span. 4. The system of claim 1, wherein the method performed by the circuitry further comprises: based on determining that the first measure of medical relevance is less than the threshold level of medical relevance, discarding the first cardiac event. 5. The system of claim 1, wherein the first measure of medical relevance embodies a severity of a cardiac condition associated with the first cardiac event. 6. The system of claim 5, wherein determining the first measure of medical relevance comprises determining a noise metric for the first cardiac event. 7. The system of claim 1, wherein the method performed by the circuitry further comprises identifying the first cardiac event, wherein the first cardiac event is one or more of an asystole event, a tachycardia event, a bradycardia event, and an atrial fibrillation/flutter event. 8. The system of claim 1, wherein a duration of the first time span is greater than a duration of the first cardiac event. 9. The system of claim 1, wherein the time span of each of the plurality of data structures has a duration including a fixed number of heart beats. 10. The system of claim 1, wherein the second cardiac event is allocated to the first data structure by a second pointer, and wherein allocating the first cardiac event to the first data structure comprises associating the first event with the first data structure using a first pointer. 11. A method com... | Braemar Manufacturing, LLC | A61B005/04 A61B005/046 A61B005/0464 A61B005/0468
US09820722 | B1 | 2017-11-21 | Simultaneous multiple method out-patient uterus biopsy device and method | 2016-05-27 | A biopsy device, comprising a flexible coaxial structure, comprising an obturator within a sheath, the obturator being adapted to be displaced with respect to the sheath along the coaxial axis by a force applied at a proximal end; a disruptor, at a distal end of the obturator, adapted to disrupt a tissue surface to free cells therefrom, having a first position covered within the sheath and a second position freely extending beyond the sheath; an element, having a fixed position on the obturator, and a close clearance with respect to an inner surface of the sheath, such that a retraction of the obturator within the sheath causes an associated change in pressure in a space between the obturator distal to the element and the sheath, to thereby draw fluid proximate to an end of the sheath within the sheath. | 1. A tissue sampling device, comprising: a coaxial structure comprising: a flexible sheath having at least a distal portion configured to maintain an internal vacuum; anda displaceable structure provided within the sheath, having a first end extending from a proximal end of the sheath configured to be manipulated by an axial movement, and second end configured to, in a first state of axial movement, extend from a distal end of the sheath, and in a second state of axial movement, to be retracted into the distal end of the sheath, the second end of the displaceable structure having a cellular sampling structure, preceded by a sampling region and sealed piston;the coaxial structure being configured such that a tension on the first end of the displaceable structure at the proximal end of the sheath results in the axial movement to cause a retraction of the displaceable structure, sealed piston, and cellular sampling structure from the first state of axial movement to the second state of axial movement,the sealed piston, displaceable structure, and sheath together being configured to create suction to generate the internal vacuum in the sheath distal to the sealed piston by the retraction, to cause a displacement of media external to the distal end of the sheath into the sheath distal to the sealed piston in the sampling region,wherein the cellular sampling structure distal to the sealed piston. 2. The tissue sampling device according to claim 1, wherein the displaceable structure terminates at the second end in an atraumatic bulb. 3. The tissue sampling device according to claim 1, wherein the cellular sampling structure comprises a brush. 4. The tissue sampling device according to claim 3, wherein the brush comprises a plurality of bristles extending radially from the displaceable structure. 5. The tissue sampling device according to claim 3, wherein the brush has a cross section which tapers with respect distance from the second end. 6. The tissue sampling device according to claim 3, wherein the brush has a helical cross sectional profile. 7. The tissue sampling device according to claim 1, wherein the coaxial structure is configured for insertion into the cervical os of a uterus of a human, to retrieve an endometrial biopsy sample, and to be withdrawn from the cervical os of the uterus. 8. The tissue sampling device according to claim 7, wherein the coaxial structure is further configured to be: inserted into the cervical os with the displaceable structure in the second state;extended into the first state with the cellular sampling structure within the uterus;manipulated by a user by movement of the first end of the displaceable structure to dislodge cells within the uterus;retracted into the second state within the uterus, to cause the internal vacuum to draw the media external to the distal end of the sheath, comprising a liquid sample surrounding the cellular sampling structure, into the distal end of the sheath; andretracted from the c... | - | A61B005/00 A61B010/02 A61M025/09
US09801547 | B2 | 2017-10-31 | Portable monitoring devices for processing applications and processing analysis of physiological conditions of a user associated with the portable monitoring device | 2014-03-20 | An activity monitoring device and associated methods for using and interfacing with the activity monitoring device are provided. In one example, the activity monitoring device includes a housing configured for attachment to a body part of a user. Also included ins a display screen attached to the housing and a sensor for capturing physiological conditions of the user. The sensor is disposed along a surface of the housing so that the sensor is proximate to the body part, the body part having at least some exposed skin. The activity monitoring device also includes memory for storing the captured physiological conditions, and a processor for examining the captured physiological conditions. At a particular time, the processor automatically selects an application to execute from a plurality of applications based on characteristics of the captured physiological conditions. | 1. An activity monitoring device configured to be worn on a wrist of a user, comprising, a housing configured for attachment to the wrist of the user, the housing having an upper surface and a generally planar lower surface, the lower surface configured to be closer to the skin of the user than the upper surface when the activity monitoring device is worn by the user, the housing further comprising: a protrusion formed on the lower surface of the housing and extending away from the lower housing of the housing, the protrusion comprising an optically opaque material and covering only a portion of the lower surface of the housing, the protrusion having (i) a lower surface formed on the protrusion, and (ii) arcuate sides formed on the protrusion to provide contact between the lower surface of the protrusion and the skin of the user but not between the lower surface of the housing and the skin of the user; andan optically transparent layer covering at least a portion of the lower surface of the housing;a display screen attached to the housing;a first sensor for capturing motion data of the user, the motion data being representative of one or more activities;a second sensor for capturing physiological conditions of the user, the second sensor being directed toward a surface of the housing so that the second sensor is proximate to the wrist when the activity monitoring device is worn by the user, the wrist having at least some exposed skin, the second sensor including: a pair of first light pipes formed in the protrusion and having a geometry defined by the optically opaque material of the protrusion, each one of the pair of first light pipes forming an elongate cylindrical structure having an upper end and a lower end, the upper end disposed adjacent the lower surface of the housing and the lower end disposed at the lower surface of the protrusion, each one of the pair of first light pipes having a light emitting device disposed at the lower surface of the housing and at the upper end of the first light pipe, wherein each of the pair of first light pipes is configured to transmit light generated by the light emitting device through the protrusion and along a direction of light travel defined by the geometry of the first light pipe; anda second light pipe formed in the protrusion and disposed between the pair of first light pipes and having a geometry defined by the optically opaque material of the protrusion, the second light pipe having an upper end and a lower end, the upper end of the second light pipe disposed adjacent the lower surface of the housing and the lower end of the second light pipe disposed at the lower surface of the protrusion, the second light pipe having a detector disposed at the lower surface of the housing and at the upper end of the second light pipe, wherein the second light pipe is configured to transmit reflected light along a direction of light travel defined by the geometry of the second light pipe toward the detector, whe... | Fitbit, Inc. | G08B021/04 A61B005/00 A61B005/11 A61B005/01 A61B005/024 A61B005/0205 A61B005/021 A61B005/22 G01C022/00 G06F019/00
US09808178 | B2 | 2017-11-07 | Tracking system using field mapping | 2013-09-23 | In some aspects, a method includes (i) securing multiple sets of current injecting electrodes to an organ in a patient's body, (ii) causing current to flow among the multiple sets of current injecting electrodes to generate a field in the organ, (iii) in response to current flow caused by the multiple sets of current injecting electrodes, measuring the field at each of one or more additional electrodes, (iv) determining expected signal measurements of the field inside the organ using a pre-determined model of the field, and (v) determining a position of each of the one or more additional electrodes in the organ based on the measurements made by the additional electrodes and the determined expected signal measurements of the field. | 1. A method comprising: determining expected signal measurements of a field inside an organ of a body using a noninterpolation based calculation, the expected signal measurements being based on signals generated by current injecting electrodes that are mounted on a first catheter placed inside the body or body surface electrodes and measured by electrodes on a second catheter; anddetermining a position of each of one or more additional electrodes on a third catheter in the organ based on measurements made by the additional electrodes and the determined expected signal measurements of the field;the method further comprising:causing current to flow among multiple sets of current injecting electrodes to generate a field in the organ in the body, the multiple sets of current injecting electrodes being secured to one or more stable locations relative to the organ in the body;obtaining positions within the organ of one or more measuring electrodes used for measuring the field generated by the multiple sets of current injecting electrodes;in response to the current flow, measuring the field at multiple locations in the organ using the one or more measuring electrodes;modeling the field using the measurements of the field from the one or more measuring electrodes and the positions within the organ of the one or more measuring electrodes; anddetermining the expected signal measurements of the field at additional locations within the organ using the model of the field. 2. The method of claim 1, wherein modeling the field comprises modeling the field based on physical characteristics using Laplace's equation. 3. The method of claim 1, wherein modeling the field comprises modeling the field based on physical characteristics using Poisson's equation. 4. The method of claim 1, wherein modeling of the field further comprises using a function that correlates field measurements with position coordinates. 5. The method of claim 4, wherein the function is a differentiable function. 6. The method of claim 1, wherein the multiple sets of current injecting electrodes are mounted on one or more catheters that are placed inside the body. 7. The method of claim 1, wherein the multiple sets of current injecting electrodes comprise one or more body-surface electrodes. 8. The method of claim 1, wherein the multiple sets of current injecting electrodes comprise both electrodes mounted on one or more catheters that are placed inside the body and body surface electrodes. 9. The method of claim 8, wherein the electrodes mounted on one or more catheters comprise one or more electrodes secured to the organ. 10. The method of claim 8, wherein the electrodes mounted on the one or more catheters comprise one or more electrodes that can be moved and positioned at multiple locations in an organ. 11. The method of claim 1, wherein obtaining the position of the one or more measuring electrodes comprises using a tracking system to track the position of the one or more measuring... | Boston Scientific Scimed Inc. | A61B005/06 A61B005/053 A61B034/20 A61B005/00 A61B018/14
US09808186 | B2 | 2017-11-07 | Method and system for providing an integrated analyte sensor insertion device and data processing unit | 2014-09-26 | Method and apparatus for providing an integrated analyte sensor and data processing unit assembly is provided. | 1. An apparatus, comprising: a housing;a data processing unit disposed in the housing;an introducer disposed within the housing and including a portion having a sharp distal end configured for piercing a body of a subject, wherein the portion of the introducer is removable from the housing after piercing the body of the subject; andan analyte sensor including a first portion and a second portion, the first portion of the analyte sensor at least partially disposed, prior to the removal of the introducer, within the portion of the introducer having the sharp distal end, and the second portion of the analyte sensor coupled to the data processing unit;wherein the first portion of the analyte sensor is configured for positioning in fluid contact with a biological fluid in the subject, and the second portion of the analyte sensor is in electrical contact with the data processing unit prior to positioning of the first portion of the analyte sensor. 2. The apparatus of claim 1, further including an adhesive layer provided on a lower surface of the housing, wherein the adhesive layer is configured to removably attach the housing to the body of the subject. 3. The apparatus of claim 2, wherein the adhesive layer includes a removable protective layer. 4. The apparatus of claim 1, wherein the apparatus is configured such that when the sharp distal end pierces the body of the subject, the first portion of the analyte sensor is positioned in fluid contact with the biological fluid in the subject. 5. The apparatus of claim 1, wherein the introducer is configured to decouple from the data processing unit after the first portion of the analyte sensor is positioned in the body of the subject. 6. The apparatus of claim 1, wherein the data processing unit includes a data communication unit configured to receive one or more signals associated with a monitored analyte level from the analyte sensor. 7. The apparatus of claim 6, wherein the data communication unit is configured to wirelessly communicate data associated with the one or more signals received from the analyte sensor to a remote location. 8. The apparatus of claim 1, wherein the introducer is positioned through an aperture of the housing. 9. The apparatus of claim 1, wherein the second portion of the analyte sensor extends substantially transverse to the first portion of the analyte sensor. 10. An assembly, comprising: a housing;a data processing unit disposed within the housing;an introducer disposed within the housing and including a portion having a sharp distal end configured for piercing a body of a subject, wherein the portion of the introducer is removable from the housing after piercing the body of the subject; andan analyte sensor coupled to the housing, a first portion of the analyte sensor configured for positioning in fluid contact with a biological fluid in the subject, and a second portion of the analyte sensor disposed within the housing and in electrical communication with the data ... | ABBOTT DIABETES CARE INC. | A61B005/00 A61B005/145 A61B005/15 A61M005/172 A61M005/142
US09830576 | B2 | 2017-11-28 | Computer mouse for monitoring and improving health and productivity of employees | 2013-09-24 | Provided are embodiments of a computer mouse for sensing health characteristics of a user. The computer mouse including a location sensor adapted to sense movement of the computer mouse, a temperature sensor adapted to sense a body temperature of the user, a blood condition sensor adapted to sense a blood saturation level of the user, and a blood pressure sensor adapted to sense a blood pressure of the user. The computer mouse including a controller to transmit, to a computer workstation, location data corresponding to the sensed movement of the computer mouse, temperature data corresponding to the sensed body temperature, blood condition data corresponding to the sensed blood saturation level, and blood pressure data corresponding to the sensed blood pressure, for determining movement of the mouse and body temperature, blood saturation level, and the blood pressure of the user. | 1. A computer mouse system for monitoring health characteristics of a user, the computer mouse system comprising: a computer mouse comprising: a computer mouse controller;a computer mouse body;a mouse location sensor configured to sense movement of the computer mouse relative to a work surface on which the computer mouse is located;a mouse temperature sensor configured to sense a body temperature of the user;a mouse blood condition sensor configured to sense a blood saturation level of the user;a mouse wireless interface;a first mouse body portion; anda mouse lid coupled to the first mouse body portion via a hinge, the mouse lid configured to rotate about the hinge between an open position and a closed position, wherein rotating the mouse lid to the open position is configured to expose an opening between the first mouse body portion and the mouse lid to provide access to the opening, and wherein rotating the mouse lid to the closed position is configured to enclose the opening between the first mouse body portion and the mouse lid to inhibit access to the opening,wherein the mouse blood condition sensor is housed in the opening such that rotating the mouse lid to the open position is configured to provide access to the mouse blood condition sensor and rotating the mouse lid to the closed position is configured to inhibit access to the mouse blood condition sensor; anda blood pressure sensor configured to sense a blood pressure of the user, the blood pressure sensor comprising: a bladder;a pump;a pressure transducer;a blood pressure sensor controller; anda blood pressure sensor wireless interface,wherein the bladder, the pump, the pressure transducer, the blood pressure sensor controller, and the blood pressure sensor wireless interface are integrated with the blood pressure sensor, and wherein the blood pressure sensor is not physically tethered to the computer mouse; the computer mouse controller configured to;send, to the blood pressure sensor controller via wireless communication between the mouse wireless interface and the blood pressure sensor wireless interface, a request for a blood pressure reading,the blood pressure sensor controller configured to, in response to the request for the blood pressure reading, perform the following operations: control operation of the pump to inflate the bladder;obtain, from the pressure transducer, one or more pressure readings indicative of a blood pressure of the user; andsend, to the computer mouse controller via wireless communication between the mouse wireless interface and the blood pressure sensor wireless interface, blood pressure datacorresponding to the one or more pressure readings;the computer mouse controller being further configured to: receive, from the mouse location sensor, movement data corresponding to sensed movement of the computer mouse for use in determining movement of the computer mouse;receive, from the mouse temperature sensor, temperature data corresponding to a sensed body temp... | Saudi Arabian Oil Company | G06Q010/00 G06Q010/10 G06F019/00 A61B005/00 A61B005/01 A61B005/0205 A61B005/021 A61B005/0476 A61B005/053 A61B005/08 A61B005/1455 A61B005/16 G06Q050/22 A61B003/113
US09821079 | B2 | 2017-11-21 | Fluorescent conjugates | 2016-10-19 | The invention provides a conjugate with fluorochrome moiety F coupled to a targeting moiety T. Moiety F is described by formula (I), and the targeting moiety T is characterized in that it has affinity to a tumor marker, such as an antibody or antibody fragment. The conjugate may be used for tumor diagnostics including photodetection of tumor nodules during resection surgery. | 1. A fluorescent conjugate comprising a fluorochrome moiety F coupled to a targeting moiety T, wherein moiety F has the formula (I):whereinY is independently for each occurrence selected from SO3H, SO3—, and SO3M,wherein M is a monovalent cation;x, z and y are independently selected from an integer of 1 to 8;and wherein moiety T can bind or associate with a tumour marker, wherein the tumour marker is a tumour antigen. 2. The conjugate of claim 1, wherein Y is selected from SO3— and SO3Na. 3. The conjugate of claim 2, wherein moiety F has the formula (II): 4. The conjugate of claim 1, wherein the tumour antigen is carcinoembryonic antigen (CEA). 5. The conjugate of claim 1, wherein moiety T is selected from an antibody, an antibody fragment, and a fusion protein comprising at least one variable region of an antibody; and wherein the antibody is optionally a chimeric monoclonal antibody (mAb). 6. The conjugate of claim 1, wherein the moiety T is a chimeric monoclonal antibody that binds to one or more epitopes of the CEA antigen. 7. The conjugate of claim 6, wherein moiety F has the formula (II), and wherein the chimeric monoclonal antibody is directed against the GOLD-2 epitope of CEA, comprises heavy chains of the G1m3 allotype and light chains of the km3 allotype, each heavy chain and each light chain comprising at least one mouse IgG1 variable domain and at least one human constant domain. 8. The conjugate of claim 7, wherein the chimeric monoclonal antibody comprises a variable light chain region of SEQ ID NO: 4 and a variable heavy chain of SEQ ID NO: 3. 9. The conjugate of claim 6, wherein there are multiple F moieties, nF, coupled to one T moiety, and wherein n is an integer selected from 1 to 4. 10. A composition comprising the conjugate of claim 1 in combination with one or more inactive ingredients comprising water for injection. 11. The composition of claim 10, wherein the average degree of conjugation of the conjugate in the composition is about 0.5 to 3. 12. The conjugate of claim 3, wherein the moiety T is a chimeric monoclonal antibody that binds to one or more epitopes of the CEA antigen. 13. The conjugate of claim 3, wherein the conjugate comprises nF moieties of Formula (II), coupled to one T moiety, wherein n is an integer selected from 1 to 4. 14. A composition according to claim 10, wherein the composition (a) is sterile and comprises an aqueous buffer and L-arginine, and/or(b) further comprises a compound comprising a T moiety which is not coupled to an F moiety. | SURGIMAB S.A.S. | A61B010/00 A61B005/00 A61B008/00 A61K049/00 A61K039/395 C07K016/30 C07K016/28 G01N033/574
US09814411 | B2 | 2017-11-14 | Robotic knee testing (RKT) device having decoupled drive capability and systems and methods providing the same | 2013-09-17 | Various limb manipulation and evaluation devices are provided. The devices generally include three drives, namely a first drive configured to manipulate a first bone relative to a second bone in a first direction, a second drive configured to manipulate the first bone relative to the second bone in a second direction, a third drive configured to manipulate the first bone relative to the second bone in a second direction. The three directions are different relative to each other and in some embodiments represent three distinct axes. The devices are further configured such that at least one of the drives is mutually decoupled relative to at least one other drive, such that operation of the one drive does not affect the position or movement of the another drive. One or multiple of the drives may be decoupled. A corresponding method of operating such decoupled drives is also provided. | 1. A limb manipulation and evaluation device comprising: a frame;a first drive supported by the frame and configured to articulate a first bone relative to a second bone about a first axis; a second drive supported by the frame and configured to articulate said first bone relative to said second bone about a second axis; and a third drive supported by the frame and configured to articulate said first bone relative to said second bone about a third axis;wherein: said first, second, and third axes are different relative to each other, andat least one drive of the first, second, and third drives is mutually decoupled relative to another drive of the first, second, and third drives, such that operation of said one drive does not affect position of said another drive relative to the frame, and such that operation of said another drive does not affect position of said one drive. 2. The limb manipulation and evaluation device of claim 1, further comprising a device for recording rotation by or torque encountered upon a respective drive of the first, second, and third drives. 3. The limb manipulation and evaluation device of claim 1, wherein: said first, second, and third axes are substantially mutually orthogonal. 4. The limb manipulation and evaluation device of claim 1, wherein: at least two of the first, second, and third axes lie in a same plane. 5. The limb manipulation and evaluation device of claim 1, wherein: the first, second, and third axes lie in respective, different planes. 6. A limb manipulation and evaluation device comprising: a first drive configured to articulate a first bone relative to a second bone about a first axis;a second drive configured to articulate said first bone relative to said second bone about a second axis; anda third drive configured to articulate said first bone relative to said second bone about a third axis,wherein: said first, second, and third axes are each oriented at an angle relative to each other, andat least one drive of the first, second, and third drives is mutually decoupled relative to another of the first, second, and third drives, such that operation of said one drive does not affect a rotational axis of said another drive, and such that operation of said another drive does not affect a rotational axis of said one drive. 7. The limb manipulation and evaluation device of claim 6, further comprising a device for recording rotation by or torque encountered upon a respective drive of the first, second, and third drives. 8. The limb manipulation and evaluation device of claim 6, wherein said first, second, and third axes are substantially orthogonal to one another. 9. The limb manipulation and evaluation device of claim 6, wherein at least two of the first, second, and third axes lie in a same plane. 10. The limb manipulation and evaluation device of claim 6, wherein the first, second, and third axes lie in respective, different planes. 11. A limb manipulation and evaluation device comprising: a f... | EMRI, Inc. | A61H001/02 A61B005/11 A61B005/00
US09808591 | B2 | 2017-11-07 | Methods and systems for breath delivery synchronization | 2014-08-15 | Systems and methods for triggering inspiration and/or cycling exhalation with a ventilator are described herein. In particular, systems and methods for synchronizing ventilator breath delivery with patient breath demand utilizing a digital sample counting trigger mode are described herein. The digital sample counting triggering mode characterizes digital samples taken from a measured or estimated parameter during the patient inhalation/exhalation period to synchronize breath delivery with patient breath demand. | 1. A method for ventilating a patient with a ventilator, comprising: monitoring a physiological parameter of the patient based on one or more received sensor measurements;calculating a first derivative of the physiological parameter for a first sample period;selecting a first trigger count threshold;selecting a first level for the first derivative;updating a sample count to form a first updated sample count for the first sample period based on first comparison results from at least one of the following: comparing the first trigger count threshold to a previous trigger count threshold selected in a previous sample period,comparing the first derivative to a previous first derivative calculated for the previous sample period, andcomparing the first derivative to the first level,comparing the first updated sample count to the first trigger count threshold; andtriggering inspiration based on a first result of the comparing of the first updated sample count to the first trigger count threshold. 2. The method of claim 1, wherein the first result is that the first updated sample count is equal to the first trigger count threshold. 3. The method of claim 1, wherein the comparing of the first updated sample count to the first trigger count threshold has a second result, wherein the second result is that the first updated sample count is not equal to the first trigger count threshold. 4. The method of claim 3, further comprising: calculating a next first derivative of the physiological parameter for a second sample period;selecting a second trigger count threshold and a second level for the first derivative;updating the first updated sample count to form a second updated sample count for the second sample period based on second comparison results from at least one of the following: comparing the second trigger count threshold to the first trigger count threshold;comparing the next first derivative to the first derivative for the first sample period, andcomparing the first derivative to the second level;comparing the second updated sample count to the second trigger count threshold;determining that the second updated sample count is equal to the second trigger count threshold based on the comparison of the second updated sample count to the second trigger count threshold; andtriggering inspiration based on the determining that the second updated sample count is equal to the second trigger count threshold. 5. The method of claim 1, wherein the first derivative is amplified. 6. The method of claim 1, wherein the first trigger count threshold is selected from a count of 2 or 4 passing samples. 7. The method of claim 1, wherein the first level is selected from one of the following groups: −3 cm of H2O per second and −8 cm of H2O per second;−4 cm of H2O per second and −9 cm of H2O per second;−5 cm of H2O per second and −10 cm of H2O per second;−10 cm of H2O per second and −15 cm of H2O per second; and−15 cm of H2O per second and −20 cm of H2O per second. 8... | COVIDIEN LP | A61M016/00 A61B005/00 A61B005/08 A61B005/113
US09791529 | B2 | 2017-10-17 | Method and device for indicating differentiation between tissues | 2014-12-31 | The present invention provides a method for indicating differentiation between tissues. For each tissue amongst multiple tissues, a magnetization vector corresponding to each tissue is generated on the basis of a random scan sequence; on the basis of the magnetization vector corresponding to each of the multiple tissues, a differentiation-indicating value between each pair of the multiple tissues is calculated. Thus, a physician can select a suitable random scan sequence, according to the method provided in the present invention, and generate a magnetic resonance image comprising multiple brightness curves corresponding to multiple tissues, such that the trends of the brightness curves corresponding to the multiple tissues in the magnetic resonance image differ significantly. The present invention also provides a device for indicating differentiation between tissues. | 1. A method for indicating differentiation between tissues of a subject, comprising: in a computer, identifying multiple tissues in a subject and, for each tissue among said multiple tissues, executing a simulation comprising a random scan sequence, wherein a magnetic resonance scan of each tissue is simulated with a selected parameter of the scan being randomly varied so as to have different parameter values and in which, as a result of the different parameter values, the respective tissue is given a magnetization vector in the simulation;in said computer, dependent on the respective magnetization vectors for each of the multiple tissue, calculating a differentiation-indicating value between each pair of tissues among said multiple tissues;in said computer, simulating a pair of brightness curves for each said pair of tissues among said multiple tissues and using the respective differentiation-indicating values to differentiate between each pair of brightness curves, and thereby obtaining a differentiation for each pair of brightness curves; andin said computer, based on said differentiation, selecting a magnetic resonance fingerprinting (MRF) protocol for scanning the multiple tissues, and making said MRF protocol available from the computer in electronic form with a format configured to operate a magnetic resonance imaging apparatus in order to execute the selected MRF protocol. 2. The method as claimed in claim 1, comprising, in said simulation, generating, for each tissue among said multiple tissues, said magnetization vector corresponding to the tissue based on said random scan sequence, by: generating a radio frequency (RF) pulse rotation matrix RRF of each tissue on the basis of the random scan sequence;generating a relaxation matrix Rrelax of each tissue on the basis of the random scan sequence;determining whether the multiple tissues include a tissue having an off-resonance property df (Hz);if so, then generating an off-resonance rotation matrix Roff of each tissue having an off-resonance property on the basis of the random scan sequence, applying the RF pulse rotation matrix RRF, the relaxation matrix Rrelax and the off-resonance rotation matrix Roff to an initial magnetization vector, and generating a magnetization vector of each tissue by iterative calculation; andotherwise, applying the RF pulse rotation matrix RRF and the relaxation matrix Rrelax to an initial magnetization vector, and generating the magnetization vector of each tissue by iterative calculation. 3. The method as claimed in claim 2, comprising selecting the random scan sequence as at least one random scan sequence from the group consisting of a flip angle random scan sequence, and a repetition time random scan sequence. 4. The method as claimed in claim 3, comprising generating said RF pulse rotation matrix RRF of the tissue on the basis of the random scan sequence by: determining whether the random scan sequence comprises the flip angle random scan sequence;if the... | Siemens Aktiengesellschaft | G01N033/48 G01N033/50 G01R033/50 G06F017/10 G06F019/18 G06F017/11 G06F019/12 A61B005/055 A61B005/00 G01R033/56
US09779751 | B2 | 2017-10-03 | Respiratory biofeedback devices, systems, and methods | 2010-02-19 | Respiratory-based biofeedback devices, systems, and methods are provided. A respiratory biofeedback method includes producing a respiratory signal in response to a user's respiratory activity, generating an audio output signal that includes a modified version of the respiratory signal, and converting the audio output signal into sound waves output to the user to provide biofeedback. The sound waves can be output to the user in real time response to the user's respiratory activity. A microphone can be used to generate the respiratory signal. The generated audio output signal can includes the respiratory signal modified to increase a volume level of a portion of the respiratory signal where the volume level exceeds a specified volume level. | 1. A biofeedback method, comprising: producing a respiratory signal in response to a user's respiratory activity;generating an audio output signal comprising a modified version of the respiratory signal, wherein a modification to the respiratory signal changes based on an output of a state machine, wherein the state machine is implemented on a digital signal processor and wherein the state machine controls an audio processor on the digital signal processor and is operable to: perform calibration during a calibrate state of the state machine using the respiratory signal, wherein the calibration comprises setting a low threshold and a high threshold; andresponsive to results from the calibration, detecting breaths of the user in a breath detection state of the state machine using the low threshold and the high threshold established during the calibrate state, wherein the detecting comprises detecting a plurality of troughs and peaks associated with a respective breath of the breaths; andconverting the audio output signal into sound waves output to the user to provide biofeedback using a biofeedback device in real time responsive to a mode of operation selected by the user and user respiratory activity as measured by the detected breaths. 2. The method of claim 1, wherein the respiratory signal is modified based on a characteristic other than intensity. 3. The method of claim 1, wherein the respiratory signal is modified based on a numerical calculation. 4. The method of claim 1, wherein the respiratory signal is modified to at least one of enhance, accentuate, or amplify a portion of the respiratory signal. 5. The method of claim 1, wherein the audio output signal is generated by using a first device, and the method further comprises transmitting at least one of the respiratory signal or the audio output signal to a second device for at least one of processing, analysis, or storage in memory. 6. The method of claim 5, wherein the transmission occurs over a communication network. 7. The method of claim 1, wherein the biofeedback provides at least one of education, behavioral modification, stress reduction, snoring reduction, or training to the user. 8. The method of claim 7, wherein the biofeedback provides fitness training to the user. 9. The method of claim 1, wherein the respiratory signal is modified to accentuate a portion of the respiratory signal having an intensity level above a specified intensity level. 10. The method of claim 1, further comprising storing data in a memory device, the data comprising at least one of: the respiratory signal;data derived from the respiratory signal;the audio output signal; ordata derived from the audio output signal. 11. The method of claim 1, wherein producing a respiratory signal comprises converting sound of the user's respiratory activity into the respiratory signal. 12. The method of claim 11, wherein the generated audio output signal comprises the respiratory signal modified to increase the ... | BREATH RESEARCH, INC. | A61B005/08 G10L021/00 A61B005/11 A61B005/00 A61B007/00 A61F005/56 A61B005/01 A61B005/0205 A61B005/021 A61B005/024 A61B005/0476 A61B005/145 G06F019/00
US09775538 | B2 | 2017-10-03 | System and method for determining the position of the tip of a medical catheter within the body of a patient | 2009-12-03 | Method and system for determining the current position of a selected portion of a medical catheter inserted into a tubular organ (118) of the body of a patient, the method comprising the procedures of inserting a medical positioning system (MPS) (102) catheter into the tubular organ (118), acquiring a plurality of mapping positions (120) within the tubular organ (118), displaying a mapping position (120) representation of the mapping positions (120), constructing a mapping path (122) according to the mapping positions (120), inserting the medical catheter into the tubular organ (118) until the selected portion reaches the initial position, displaying an operational image of the tubular organ (118), a path representation of the mapping path (122), and an initial position representation of the initial position superimposed on the operational image, registering the selected portion with the initial position, measuring a traveled length of the medical catheter within the tubular organ (118) from the initial position, and estimating the current position. | 1. A method for determining the current position of a selected portion of a medical catheter inserted into a tubular organ, the method comprising: acquiring a plurality of mapping positions of a Medical Positioning System (MPS) catheter within the tubular organ according to output from a sensor included on the MPS catheter, the sensor electrically coupled with an MPS;displaying a mapping position representation of the mapping positions, superimposed on an image of the tubular organ;constructing a mapping path according to the mapping positions, a selected one of the mapping positions being defined as an initial position of the mapping path;displaying an operational image of the tubular organ, a path representation of the mapping path, and an initial position representation of the initial position superimposed on the operational image, and a marker image of the selected catheter position of the medical catheter,registering a selected catheter portion with the initial position when the selected catheter portion is disposed at the initial position;measuring a traveled length of the catheter within the tubular organ from the initial position; andestimating a current three-dimensional position of the selected catheter portion according to the traveled length, according to the mapping positions, and according to a plurality of calculated distances between each of the mapping positions, and the initial position, along the mapping path. 2. The method according to claim 1, further comprising acquiring a pre operational image of the tubular organ, by an imager. 3. The method according to claim 2, further comprising registering a three-dimensional coordinate system associated with the MPS with a two-dimensional coordinate system associated with the pre operational image. 4. The method according to claim 1, further comprising acquiring an operational image of the tubular organ, after constructing the mapping path, to enable display of a representation of the at least one mapping path, on the operational image. 5. The method according to claim 4, further comprising superimposing the representation of the at least one mapping path on the operational image. 6. The method according to claim 5, further comprising displaying a superimposed operational image of the representation of the at least one mapping path on the operational image. 7. The method according to claim 1, further comprising registering the tip of the medical catheter with the initial position. 8. The method according to claim 1, further comprising acquiring an organ timing signal of an organ. 9. The method according to claim 8, further comprising acquiring a plurality of pre operational images of the organ, according to the organ timing signal. 10. The method according to claim 9, further comprising registering each of the mapping positions with respective two-dimensional coordinates of a respective pre operational image. 11. The method according to claim 8, further comprising grouping t... | MediGuide Ltd. | A61B005/05 A61B006/00 A61B005/06 A61B005/00 A61B008/12 A61B008/00 A61B008/08 A61B034/20 A61M025/01 A61B006/12
US09788771 | B2 | 2017-10-17 | Variable speed sensor insertion devices and methods of use | 2006-10-23 | An automatic sensor inserter is disclosed for placing a transcutaneous sensor into the skin of a living body. According to aspects of the invention, characteristics of the insertion such as sensor insertion speed may be varied by a user. In some embodiments, insertion speed may be varied by changing an amount of drive spring compression. The amount of spring compression may be selected from a continuous range of settings and/or it may be selected from a finite number of discrete settings. Methods associated with the use of the automatic inserter are also covered. | 1. A sensor insertion device, comprising: a shuttle movable between a first position and a second position, the shuttle configured to retain a transcutaneous sensor while in the first position and release the sensor while in the second position;a drive mechanism coupled to the shuttle to drive the shuttle from the first position to the second position based on a first driving force, wherein the drive mechanism includes a first compression spring configured to apply the first driving force;a retraction mechanism coupled to the shuttle to drive the shuttle from the second position to a neutral position based on a second driving force, wherein the retraction mechanism includes a second compression spring configured to apply the second driving force, and wherein the neutral position is between the first position and the second position; anda control mechanism operatively coupled to the drive mechanism, wherein the control mechanism is manually adjustable by a user and configured to vary the first driving force of the first compression spring by varying a load on the first compression spring while the shuttle is stationary. 2. The insertion device of claim 1, wherein the manually adjustable control mechanism is configured to vary the first driving force over a continuous range of settings. 3. The insertion device of claim 1, wherein the manually adjustable control mechanism is configured to alternately set the first driving force to one of a plurality of discrete settings. 4. The insertion device of claim 1, wherein the manually adjustable control mechanism is solely mechanical. 5. The insertion device of claim 4, wherein the manually adjustable control mechanism includes a knob configured to be twisted by the user of the insertion device such that twisting of the knob varies the first driving force of the first compression spring. 6. The insertion device of claim 4, wherein the manually adjustable control mechanism includes a thumbwheel configured to be turned by the user of the insertion device such that turning of the thumbwheel varies the first driving force of the first compression spring. 7. The insertion device of claim 4, wherein the manually adjustable control mechanism includes a threaded rod. 8. The insertion device of claim 1, wherein the manually adjustable control mechanism includes a component configured to be set to one of a plurality of discrete, alternate shuttle orientations by the user of the insertion device, wherein each shuttle orientation varies the first driving force of the first compression spring. 9. The insertion device of claim 1, wherein the manually adjustable control mechanism is configured to allow the user of the insertion device to manually set a sensor insertion velocity. 10. The insertion device of claim 1, wherein the manually adjustable control mechanism is configured to allow an amount of compression of the first compression spring of the drive mechanism to be manually changed by the user. 11. The ins... | ABBOTT DIABETES CARE INC. | A61B005/1468 A61B005/1486 A61B005/145
US09811935 | B2 | 2017-11-07 | Emotive advisory system and method | 2008-04-28 | Information about a device may be emotively conveyed to a user of the device. Input indicative of an operating state of the device may be received. The input may be transformed into data representing a simulated emotional state. Data representing an avatar that expresses the simulated emotional state may be generated and displayed. A query from the user regarding the simulated emotional state expressed by the avatar may be received. The query may be responded to. | 1. An emotive advisory system for use by an occupant of an automotive vehicle, the system comprising: a computer configured toreceive input indicative of an operating state of the vehicle and input indicative of an emotional state of the occupant,generate data representing an avatar having an appearance and data representing a spoken statement for the avatar having at least one speech characteristic, the appearance and the at least one speech characteristic (i) depending on at least one of the operating state of the vehicle and the emotional state of the occupant and (ii) conveying a simulated emotional state of the avatar to the occupant to prompt the occupant to pose a query regarding a cause of the avatar's simulated emotional state,receive input indicative of a query from the occupant regarding the avatar's simulated emotional state,in response to the query, generate data representing another spoken statement for the avatar and including information about the cause of the avatar's simulated emotional state, andoutput the data representing the avatar for visual display and the data representing the statements for the avatar for audio play. 2. The system of claim 1 wherein the computer is further configured to determine whether a predetermined event has occurred and to query the occupant via spoken dialog if the predetermined event has occurred. 3. The system of claim 2 wherein the computer is further configured to receive a spoken response from the occupant to the computer's query. 4. The system of claim 3 wherein the computer is further configured to alter an operating parameter of the vehicle based on the spoken response to the computer's query. 5. The system of claim 1 wherein the input indicative of an emotional state of the occupant is based on at least one of audio and visual characteristics of the occupant. 6. The system of claim 1 wherein the computer is further configured to determine whether a predetermined event has occurred, to access information from sources remote from the vehicle if the predetermined event has occurred and to provide at least a portion of the accessed information to the occupant visually or audibly. 7. An emotive information system for a user of an automotive vehicle, the system comprising: at least one sensor configured to detect at least one of audio and visual characteristics of the user;a computer configured to receive the detected information,receive information about an operating state of the vehicle, andgenerate data representing a simulated emotional state based on at least a portion of the received information; andan interface configured to convey the simulated emotional state to the user to prompt the user to pose a query regarding a cause of the simulated emotional state, wherein the computer is further configured to receive a query from the user regarding the simulated emotional state and generate data representing a response to the query including information about the cause of the simulated e... | Ford Global Technologies, LLC | G05D001/00 G05D003/00 G06F007/00 G06F017/00 G06T013/40 B60W050/10 G01C021/36 G06F003/01 G06N003/00 G10L017/26 G06T013/20 A61B005/16 A61B005/18 B60W040/08 A61B005/00
US09808181 | B2 | 2017-11-07 | System and method for in-rumen monitoring | 2012-06-15 | A system for monitoring digestive efficiency within the rumen of one or more ruminant animals comprises rumen boluses shaped and sized to be retained within the rumen dorsal sac and each comprising temperature, pH sensor, and redox sensors, and a wireless transmitter. A processor is arranged to derive from the sensor data one or more parameters indicative of animal digestive efficiency including any combination of one or more of hydrogen scale (rH), partial pressure of hydrogen (pp[H2]), oxygen fugacity (f(O2)) and free energy of the system (ΔG). A method and bolus are also claimed. | 1. A system for monitoring the process of fermentation within the rumen of one or more ruminant animals comprising: one or more rumen boluses each having a body comprising at or towards one end of the body opposed lateral extensions or tabs which are flexible toward the body but have memory towards a position extending outwardly of the body to retain the bolus within the dorsal sac of the rumen of the associated animal, the bolus having buoyancy within the dorsal sac against falling to the bottom of the rumen, and each bolus comprising within the body: a temperature sensor to provide data indicative of temperature within the rumen,a pH sensor to provide data indicative of pH within the rumen, anda redox sensor to provide data indicative of the oxidation-reduction potential within the rumen, anda wireless transmitter for transmitting data or derived parameters to a remote station, andin the or each bolus or at a remote station a processor arranged to derive from the sensor data one or more parameters indicative of the process of fermentation. 2. A system according to claim 1 wherein each bolus comprises stored in memory in the bolus an identification code unique to the individual bolus. 3. A system according to claim 1 arranged to transmit data at least once per week, at least once per day, or at least twice per day. 4. A system according to claim 3 arranged to transmit data for at least one week, for at least four weeks, for at least three months, or indefinitely. 5. A system according to claim 1 wherein the processor is located within the remote station. 6. A system according to claim 1 wherein the processor is located within the bolus. 7. A system according to claim 1 arranged to store locally in the bolus the sensor data and/or derived parameters and transmit same to the remote station when the animal comes into proximity of a receiver. 8. A system according to claim 7 wherein the receiver is located proximal to an animal milking, watering or handling station. 9. A system according to claim 1 wherein the one or more parameters includes any combination of one or more of hydrogen scale (rH), partial pressure of hydrogen (pp[H2]), oxygen fugacity (f(O2)) and free energy of the system (ΔG). 10. A system according to claim 9 wherein the equation stored in memory for obtaining rH is: rH=(Eh*2*F/(2.303*R*T))+2*pHwhere: Eh=hydrogen potential (V) derived from the redox sensor dataF=faraday's constant (96.485 kj·V−1·mol−1)R=gas constant (0.008314472 kJ·mol−1·K−1)T=temperature (K) derived from the temperature sensor data,rH=−log (H2), andpH=−log (H+) derived from the pH sensor data. 11. A system according to claim 10 wherein rH is indicative of the reductive or oxidative intensity of the rumen environment and the processor is arranged to output a value ranging from 0to 42, wherein 0 represents a saturate atmosphere of hydrogen. 12. A system according to claim 9 wherein the equation stored memory for obtaining pp[H2] is: pp[H2]=10−rH (atm)where... | Kahne Limited | A61B005/00 A61B005/07 A61B005/01 A61B005/145 A61B005/03 A61B005/053
US09808197 | B2 | 2017-11-07 | Electronic apparatus and control method | 2015-02-27 | According to one embodiment, a wearable electronic apparatus detects movement of a human body wearing the apparatus. The apparatus causes a first number of light emitters to emit light when a magnitude of the detected movement is less than a first value. The first number of light emitters is among a plurality of light emitters which emit light having different wavelengths, in which the light are received by a light receiver to obtain physiological information. The apparatus causes a second number greater than the first number of light emitters among the plurality of light emitters to emit light when the magnitude of the detected movement is greater than or equal to the first value. | 1. A wearable electronic apparatus, comprising: a plurality of light emitters including a first LED, a second LED, a third LED and a fourth LED, the first LED emitting red light, the second LED emitting infrared light, the third LED and the fourth LED emitting light different from the light emitted by the first LED and the second LED, and having wavelengths different from each other;a light receiver to receive reflected light or transmitted light emitted from the light emitters, the reflected light from a human body wearing the apparatus, the transmitted light passing through the human body;a sensor to detect movement of the human body wearing the apparatus; anda CPU configured to execute various programs including a physiological information acquisition program,wherein the CPU is configured to:two light emitters including the first LED which emits red light and the second LED which emits infrared light to emit light when the magnitude of the movement detected by the sensor is less than a threshold value, and acquire physiological information of the human body based on an output signal from the light receiver which receives the reflected light or the transmitted light; and;cause four light emitters including the first LED which emits red light, the second LED which emits infrared light, and the third LED and the fourth LED which emit light different from the light emitted by the first LED and the second LED, and having wavelengths different from each other to emit light when the magnitude of the movement detected by the sensor is greater than or equal to the threshold value, and acquire the physiological information of the human body based on an output signal from the light receiver which receives the reflected light or the transmitted light. 2. The apparatus of claim 1, wherein the first LED, the second LED, the third LED and the fourth LED are arranged such that paths of the light emitted from the plurality of light emitters to the light receiver are equidistant. 3. The apparatus of claim 2, wherein the first LED, the second LED, the third LED and the fourth LED are arranged on a circumference of the light receiver. 4. The apparatus of claim 1, wherein the third LED emits light having a wavelength which deoxyhemoglobin [Hb] absorbs more strongly than oxyhemoglobin [HbO2], andthe fourth LED emits light having a wavelength which oxyhemoglobin [HbO2] absorbs more strongly than deoxyhemoglobin [Hb]. 5. The apparatus of claim 1, wherein the light receiver is configured to output the output signal in accordance with a synchronous signal indicative of light-emission timing of the plurality of light emitters. 6. A method of controlling a wearable electronic apparatus, the method comprising: detecting, by a sensor, movement of a human body wearing the apparatus;causing two light emitters including a first LED which emits red light and a second LED which emits infrared light to emit light when the magnitude of the movement detected by the sensor is ... | Kabushiki Kaisha Toshiba | A61B005/1455 A61B005/00 A61B005/11 A61B005/024
US09775573 | B2 | 2017-10-03 | Peroperative sensing head adapted to be coupled to an ablation tool | 2014-12-23 | A peroperative probe for guiding a manual excision tool. The probe includes a detection head, an optical fiber for the reception and guidance of a signal emitted by radioactive tracers and fluorescent molecules in a tissue area, a photo-detector for converting the emitted signal into an electrical signal, a transmitter for transmitting information carried by the electrical signal to an analysis equipment, and a fastener for attaching the probe onto the manual excision tool, so that the excision tool can be used to remove a portion of tissue from the tissue area emitting the signal. | 1. A preoperative probe for use with an analysis equipment for guiding a manual excision tool, comprising: a single-use detection head comprising a first body having a first optical fiber for the reception and guidance of a signal representative of both beta and gamma particles emitted by radioactive tracers in a tissue area, a second optical fiber for reception and guidance of a signal representative of gamma radiation only, and a first planar connection surface having a connection stud, the first and second optical fibers each having an end disposed flush with the first planar connection surface; anda reusable detachable part comprising a second body having a second planar connection surface and a connection orifice, a first photo-detector for converting the signal guided by the first optical fiber into a first electrical signal, a second photo-detector for converting the signal guided by the second optical fiber into a second electrical signal, and an integrated circuit coupled to the first and second photo-detectors to output information for processing to the analysis equipment, each of first and second photo-detectors disposed flush with the second planar connection surface;a transmitter for transmitting the information output by the integrated circuit to the analysis equipment; andlocking component for removably coupling the second body to the first body so that the first and second planar surfaces mate, the first optical fiber is disposed adjacent to the first photo-detector, the second optical fiber is disposed adjacent to the second photo-detector and the connection stud and the connection orifice interengage. 2. The probe according to claim 1, wherein the transmitter comprises a wireless communication module. 3. The probe according to claim 1, wherein the transmitter comprises a transmission cable coupled to the reusable detachable part. 4. The probe according to claim 1, wherein at least one of the first and second photo-detectors comprises a silicon photomultiplier. 5. The probe according to claim 1, wherein the detection head comprises a handpiece configured to be manipulated by a clinician. 6. The probe according to claim 5, wherein handpiece is configured to receive an excision tool. 7. The probe according to claim 1, further comprising a temperature sensor for sensing a temperature of one of the first and second photo-detectors. 8. The probe according to claim 7, further comprising a battery for supplying a power signal to the first and second photo-detectors and the integrated circuit is configured to adjust the power signal based on the temperature sensed by the temperature sensor. 9. The probe according to claim 1, wherein the transmitter is arranged to transmit information carried by the first electrical signal and the second electrical signal to the analysis equipment. 10. The probe according to claim 1, further comprising a third photo-detector, wherein the first photo-detector comprises a surface in contact with an... | Centre National De La Recherche Scientifique | A61B006/00 A61B005/00 A61B010/02 A61B090/00 A61B001/00 A61B017/22 A61B017/00 A61B017/32
US09795297 | B2 | 2017-10-24 | Hand-held medical-data capture-device having detection of body core temperature by a microprocessor from a signal from a digital infrared sensor on a separate circuit board with no A/D converter and having interoperation with electronic medical record systems without specific discovery protocols or domain name service | 2014-12-29 | In one implementation, an apparatus estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point is described, estimates temperature from a digital infrared sensor and determines vital signs from a solid-state image transducer, or determines vital signs from a solid-state image transducer and estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point; after which the estimated and/or determined information is transmitted to an external database. | 1. A device comprising: a first circuit board including:a microprocessor;a wireless communication subsystem that is operably coupled to the microprocessor and that is operable to transmit a representation of a temperature; anda first digital interface operably coupled to the microprocessor;a second circuit board including:a second digital interface, the second digital interface being operably coupled to the first digital interface; anda digital infrared sensor operably coupled to the second digital interface, the digital infrared sensor having ports that provide only digital readout signals, the digital infrared sensor further comprising a low noise amplifier, an analog-to-digital converter, a digital signal processor, and a pulse width modulator (PWM) in a single industry standard package; wherein the digital infrared sensor is configured with no need for recalibration after an initial factory calibration with an emissivity range from 0.1 to 1.0;wherein the microprocessor is operable to receive from the ports of the digital infrared sensor the digital readout signals that are representative of an infrared signal, the infrared signal acquired from a forehead of a user and generated by the digital infrared sensor, and the microprocessor is further operable to determine the temperature from the received digital readout signals by using the PWM, andwherein no analog-to-digital converter is operably coupled between the digital infrared sensor and the microprocessor. 2. The device of claim 1 wherein a display device that is operably coupled to the microprocessor further comprises: a green traffic light operable to indicate that the temperature is good;an amber traffic light operable to indicate that the temperature is low; anda red traffic light operable to indicate that the temperature is high. 3. The device of claim 1 further comprising: the digital infrared sensor having no analog sensor readout ports. 4. The device of claim 1 further comprising: a camera operably coupled to the microprocessor and providing at least two images to the microprocessor; andthe microprocessor including a cropper that is operable to receive the at least two images and operable to crop the images to exclude a border area of the images, generating at least two cropped images, the microprocessor also including a temporal-variation-amplifier of at least two cropped images that is operable to generate a temporal variation, the microprocessor also including a vital-sign generator that is operably coupled to the temporal-variation-amplifier that is operable to generate at least one vital sign from the temporal variation. 5. The device of claim 4, wherein the temporal-variation-amplifier further comprises: skin-pixel-identifier. 6. The device of claim 4, wherein the temporal-variation-amplifier further comprises: a regional facial clusterial module. 7. The device of claim 4, wherein the temporal-variation-amplifier further comprises: a first frequency filter. 8. The devic... | ARC DEVICES, LTD | G08B023/00 A61B005/00 A61B005/01 G06F019/00 A61B005/021 A61B005/026 A61B005/08 A61B005/024 A61B005/0205 G06T003/40 G06T007/00 H04W048/16 H04W076/02 A61B005/0402 A61B005/03 G01J005/00 H04W080/04 G01J005/02 G01J005/10 G08C017/02 H04L029/08 H04W012/06 H04B007/26 H04L012/46 H04N005/33 H04W048/08 G08B005/36 G08B021/18 G01K013/00 G06K009/62 G06T007/174 G06T007/269 G06T007/90 A61B005/0245 H04W012/08
US09778489 | B2 | 2017-10-03 | Electro-active spectacles | 2016-08-24 | A spectacle lens is disclosed. The disclosed lens provides a vision correcting area for the correction of a wearer's refractive error. The viewing correction area provides correction for non-conventional refractive error to provide at least a part of the wearer's vision correction. The lens has a prescription based on a wave front analysis of the wearer's eye and the lens can further be modified to fit within an eyeglass frame. | 1. Electro-active spectacles comprising: spectacle frames;a first lens held by the spectacle frames, the first lens comprising a first plurality of stacked electro-active layers having refractive indices that vary in response to voltage, the first plurality of stacked electro-active layers covering a majority of the first lens;a second lens held by the spectacle frames, the second lens comprising a second plurality of stacked electro-active layers having refractive indices that vary in response to voltage, the second plurality of stacked electro-active layers covering a majority of the second lens;a power supply that is in electrical communication with the first plurality of stacked electro-active layers and the second plurality of stacked electro-active layers, and is configured to operatively supply a voltage to the first plurality of stacked electro-active layers and the second plurality of stacked electro-active layers;a plurality of photodetectors that is mechanically coupled to the spectacle frames and is configured to operatively sense positions of edges of pupils of the user's eyes when the user is wearing the frames; anda controller that is coupled to the plurality of photodetectors, the first plurality of stacked electro-active layers, the second plurality of stacked electro-active layers, and the power supply, is configured to operatively convert the positions of the edges of the pupils of the user's eyes to an inter-pupillary distance, and is configured to operatively regulate the voltage supplied from the power supply to the first electro-active element and the second electro-active element based on the inter-pupillary distance so as to vary the refractive indices of the first plurality of stacked electro-active layers and the second plurality of stacked electro-active layers. 2. Electro-active spectacles comprising: spectacle frames;a first lens held by the spectacle frames, the first lens comprising a first electro-active element covering a majority of the first lens;a second lens held by the spectacle frames, the second lens comprising a second electro-active element covering a majority of the second lens;a power supply that is in electrical communication with the first electro-active element and the second electro-active element and is configured operatively to supply a voltage to the first electro-active elementand the electro-active layer;a plurality of photodetectors that is mechanically coupled to the spectacle frames and is configured to operatively sense positions of edges of pupils of the user's eyes when the user is wearing the frames; anda controller that is coupled to the plurality of photodetectors, the first electro-active element, the second electro-active element, and the power supply, and is configured to operatively regulate the voltage supplied from the power supply to the first electro-active element and the second electro-active element based on the positions of edges of pupils of the user's eyes sensed by the... | e-Vision Smart Optics, Inc. | G02C007/06 G02C007/08 G02B007/28 G02C007/02 G02C007/10 G02F001/1335 A61B003/028 A61B003/04 G02C011/00 A61B005/11 G02F001/29
US09808182 | B2 | 2017-11-07 | Breathing detection apparatus | 2015-03-20 | An apparatus including a tubular channel having a pathway extending between a proximal end and a distal end of the tubular channel. The proximal end is configured to connect directly or indirectly to an end tidal CO2 monitor. The distal end has an opening. The apparatus further includes a detection member including a chamber disposed in fluid communication with the pathway of the tubular channel such that gas entering the tubular channel via the opening on the distal end passes into the chamber. A detection element is disposed within the chamber and includes a component that is sensitive to one or more systemic biomarkers such that, upon exposure to a predetermined concentration level of the one or more systemic biomarkers contained in the gas, a state of the detection element experiences a permanent alteration and the detection member indicates that the predetermined concentration level of the one or more systemic biomarkers is present in the gas. | 1. An apparatus, comprising: a tubular channel having a pathway extending between a proximal end and a distal end of the tubular channel, the proximal end being configured to connect directly or indirectly to an end tidal CO2 monitor, and the distal end having an opening;a detection member including: a chamber disposed in fluid communication with the pathway of the tubular channel such that gas entering the tubular channel via the opening on the distal end passes into the chamber, anda detection element disposed within the chamber, the detection element including a component that is sensitive to one or more systemic biomarkers such that, upon exposure to a predetermined concentration level of the one or more systemic biomarkers contained in the gas, a state of the detection element experiences a permanent alteration as an indication that the predetermined concentration level of the one or more systemic biomarkers is present in the gas;a cannula channel that extends from a side of the detection member in a direction transverse to the pathway of the tubular channel, the cannula channel being configured to allow a nasal cannula to be inserted laterally therein such that the nasal cannula extends in a same direction as the pathway of the tubular channel; anda movable obstruction element that, in a first position, permits insertion of the nasal cannula into the cannula channel, and when moved from the first position into a second position, impedes removal of the nasal cannula without pinching the nasal cannula when the nasal cannula is inserted in the cannula channel,wherein the cannula channel includes a slot having a width as wide as a diameter of the nasal cannula and a thickness as thick as the detection member, the slot extending longitudinally from an outer wall of the detection member toward the chamber and curving to extend away from the chamber before terminating within the outer wall of the detection member,wherein the detection member further includes a protrusion on a side thereof, the protrusion extending around the tubular channel and having a threaded external surface,wherein the obstruction element is a cylindrical collar having threads on an internal surface, a first side of the collar having a hole therein through which the first end of the tubular channel extends, and a second side of the collar facing the detection member,wherein, in the first position, the collar is disconnected from the detection member and the nasal cannula is insertable into the slot, andwherein, in the second position, the collar is connected to the detection member, via engagement between the threads on the internal surface of the collar and the threaded external surface of the protrusion on the side of the detection member, and the nasal cannula is fixed in place via pressure between a sidewall of the slot and an exterior side wall of the collar. 2. The apparatus according to claim 1, wherein the chamber is defined by at least a first wall, a second wall fa... | Bohnas LLC | A61B005/097 A61B005/083 A61B005/00 A61M016/08 A61M016/06 A61M016/04 A61M016/16
US09797881 | B2 | 2017-10-24 | Method and system for controlling a passive driver impairment detection system in a vehicle | 2015-11-05 | Methods and systems for controlling a passive driver impairment detection system in a vehicle are provided. The method and system detect states of vehicle component operating systems and/or vehicle occupant monitoring systems and control activation and deactivation of the impairment detection system in response. The methods and systems enable a proactive transition from an inactive state to an active state by detecting the presence of an individual within or approaching the vehicle or an intent of the individual to approach the vehicle in order to overcome a significant functional latency in the impairment detection system. The methods and system further enable a transition from the active state to the inactive state by using multiple systems to corroborate the absence of the individual within or approaching the vehicle and the lack of intent of the individual to approach the vehicle. | 1. A method for controlling a passive driver impairment detection system in a vehicle, comprising the steps of: detecting a state of a first vehicle system, wherein the state of the first vehicle system includes an indication of at least one of the following: i) whether an individual is within the vehicle, ii) if not within the vehicle, whether the individual is approaching the vehicle, or iii) if not approaching the vehicle, whether the individual has an intent to approach the vehicle;transmitting an activation signal to the passive driver impairment detection system when detection of the state of the first vehicle system indicates at least one of i), ii) or iii), wherein the passive driver impairment detection system is provided to detect an impairment of the individual who intends to operate the vehicle and the passive driver impairment detection system has a significant functional latency; andproactively transitioning the passive driver impairment detection system from an inactive state to an active state responsive to the activation signal before the individual operates the vehicle;wherein the transition from the inactive state to the active state results in a reduction, from the perspective of the individual, of the significant functional latency of the passive driver impairment detection system. 2. The method of claim 1, further comprising the steps of: detecting a state of a second vehicle system indicative of the presence of the individual within or approaching the vehicle or the intent of the individual to approach the vehicle; andtransmitting the activation signal to the passive driver impairment detection system upon detection of the state of the second vehicle system indicative of the presence of the individual within or approaching the vehicle or the intent of the individual to approach the vehicle. 3. The method of claim 1, wherein the first vehicle system comprises a vehicle key fob system and the state of the first vehicle system indicative of the presence of the individual within or approaching the vehicle or the intent of the individual to approach the vehicle is one of a proximity of a key fob to the vehicle or actuation of an actuator of the key fob system. 4. The method of claim 1, wherein the first vehicle system comprises a vehicle door system and the state of the first vehicle system indicative of the presence of the individual within or approaching the vehicle or the intent of the individual to approach the vehicle is one of an open position of a door of the vehicle or an actuated state of a door handle of the door of the vehicle. 5. The method of claim 1, wherein the first vehicle system comprises a vehicle infotainment system and the state of the first vehicle system indicative of the presence of the individual within or approaching the vehicle is an active state of the vehicle infotainment system. 6. The method of claim 1, wherein the first vehicle system comprises a vehicle cabin atmospheric monitoring system ha... | GM GLOBAL TECHNOLOGY OPERATIONS LLC | G08B023/00 G01N033/497 B60R016/03 F02N011/08 G01N033/00 B60K028/00 A61B005/00 A61B005/087
US09819138 | B2 | 2017-11-14 | Object information acquiring apparatus, laser apparatus, and medical apparatus | 2014-07-25 | An object information acquiring apparatus is used which includes a laser medium that oscillates laser light, an excitation source that excites the laser medium, a voltage accumulator that applies a voltage to the excitation source, a voltage supplier that supplies a voltage to the voltage accumulator, a voltage controller that limits a maximum supplied voltage from the voltage supplier, a receiver that receives a photoacoustic wave generated by an object irradiated with the laser light, and a constructor that acquires characteristic information relating to the object in use of the photoacoustic wave, wherein the voltage controller compares a measured voltage value obtained by implementing division of a supplied voltage from the voltage supplier with a reference voltage value defining the maximum supplied voltage. | 1. An object information acquiring apparatus comprising: a laser medium that oscillates laser light;an excitation source that excites the laser medium;a voltage accumulator that applies a voltage to the excitation source;a voltage supplier that supplies a voltage to the voltage accumulator;a voltage controller that limits a maximum supplied voltage from the voltage supplier;a receiver that receives a photoacoustic wave generated by an object irradiated with the laser light; anda constructor that acquires characteristic information relating to the object in use of the photoacoustic wave,wherein the voltage controller compares a measured voltage value obtained by implementing division of a supplied voltage from the voltage supplier with a reference voltage value defining the maximum supplied voltage. 2. The object information acquiring apparatus according to claim 1, further comprising a display that displays a result of the comparison performed by the voltage controller. 3. The object information acquiring apparatus according to claim 1, wherein a result of the comparison performed by the voltage controller is a ratio of the measured voltage value to the reference voltage value. 4. The object information acquiring apparatus according to claim 3, wherein the voltage supplier supplies a voltage such that the ratio of the measured voltage value to the reference voltage value is lower than 100% and equal to or more than 90%. 5. The object information acquiring apparatus according to claim 1, further comprising a setting mechanism that sets the reference voltage value. 6. The object information acquiring apparatus according to claim 5, wherein the setting mechanism specifies the reference voltage value at different levels according to a time elapsed since start of using the excitation source. 7. The object information acquiring apparatus according to claim 6, wherein the excitation source is a flash lamp. 8. The object information acquiring apparatus according to claim 1, further comprising a wavelength selector for the laser light to be oscillated, wherein the constructor acquires an oxygen saturation in the object in use of characteristic information on a plurality of wavelengths. 9. A laser apparatus comprising: a laser medium that oscillates laser light;an excitation source that excites the laser medium;a voltage accumulator that applies a voltage to the excitation source;a voltage supplier that supplies a voltage to the voltage accumulator; anda voltage controller that limits a maximum supplied voltage from the voltage supplier;wherein the voltage controller compares a measured voltage value obtained by implementing division of a supplied voltage from the voltage supplier with a reference voltage value defining the maximum supplied voltage prior to an oscillation of laser light, and the voltage controller controls based on a result of the comparison, the voltage supplier to supply a voltage to the voltage accumulator. 10. A medical appara... | CANON KABUSHIKI KAISHA | H01S003/13 H01S003/092 G01N029/24 H01S003/11 H01S003/131 H01S003/102 A61N005/06 G01N029/06 G01N029/44 A61B018/20 A61N005/067 A61B018/00 A61B005/00 H01S003/091 H01S003/10
US09775541 | B2 | 2017-10-03 | Mask | 2014-10-03 | A mask configured to be attached to the face of a subject includes a wall section that defines an internal space and covers at least a portion of a nose and a mouth of the subject, an expired gas introduction section that is disposed in the internal space and introduce the subject's expired gas, and a communication section defining a communication channel through which the subject's expired gas introduced from the expired gas introduction section is introduced into an expired gas concentration detection sensor. The position of the expired gas introduction section in the internal space is variable. | 1. A mask configured to be attached to a face of a subject comprising: a wall section that defines an internal space and is adapted to cover at least a portion of a nose and a mouth of the subject;an expired gas introduction section that is disposed in the internal space and is adapted to introduce expired gas of the subject;a communication section that defines a communication channel through which the expired gas introduced from the expired gas introduction section is introduced into an expired gas concentration detection sensor; anda bridge part that connects the wall section and the expired gas introduction section,wherein the bridge part is stretchable,wherein the bridge part includes a corrugated part, andwherein a position of the expired gas introduction section in the internal space is variable. 2. The mask according to claim 1, wherein the communication section and the expired gas introduction section are coupled so as to be relatively movable with respect to each other. 3. The mask according to claim 1, wherein the expired gas introduction section is supported by the communication section, wherein the mask further comprises an adjusting mechanism that adjusts a supporting position at which the communication section is supported by the wall section. 4. The mask according to claim 3, wherein the adjusting mechanism includes a mechanism that adjusts the supporting position by elastic deformation of the wall section. 5. The mask according to claim 1, wherein the expired gas introduction section includes a partition part which defines a first opening part adapted to open against the face of the subject when the subject wears the mask, and which defines a second opening part opening upwards. 6. A mask configured to be attached to a face of a subject comprising: a wall section that defines an internal space and is adapted to cover at least a portion of a nose and a mouth of the subject;an expired gas introduction section that is disposed in the internal space and is adapted to introduce expired gas of the subject;a communication section that defines a communication channel through which the expired gas introduced from the expired gas introduction section is introduced into an expired gas concentration detection sensor;a sensor support that supports the expired gas concentration detection sensor, which includes a light-emitting device and a light-receiving device, so that light emitted from the light-emitting device passes across a light-passing portion of the communication channel; an oxygen-supply section that communicates with the internal space and holds one end of an oxygen-supply tube; anda holding section that detachably holds a signal line connected to the expired gas concentration detection sensor,wherein a position of the expired gas introduction section in the internal space is variable,wherein the sensor support has a vent hole through which an outside of the wall section and the communication channel communicate with each other... | NIHON KOHDEN CORPORATION | A61B005/08 A61B005/097 A61M016/06 A61M016/08 A61M016/00 A61B005/00 A61M016/10 A61B005/083
US09799187 | B2 | 2017-10-24 | Method and apparatus for limiting range of motion of body | 2013-02-08 | The present invention relates to apparatus and methods for warning a user about undesirable motion of a first portion of the user's body relative to a second portion of the user's body. In one representative embodiment, an apparatus for warning a user about motion of a first portion of the user's body relative to a second portion of the user's body comprises at least one sensor configured to measure the angle of the first portion of the user's body relative to a reference. The sensor is configured to be worn on the user's body, and is in communication with a controller. A warning mechanism is also in communication with the controller, and the controller is operable to activate the warning mechanism upon movement of the first portion of the user's body relative to the second portion of the user's body beyond a predetermined angle relative to the reference. | 1. An apparatus for warning a user about motion of a first portion of the user's body relative to a second portion of the user's body, comprising: a first sensor mountable on the first portion of the user's body and configured to measure an angle of the first portion of the user's body relative to a reference;a second sensor mountable on the second portion of the user's body and configured to measure an angle of the second portion of the user's body relative to the reference;a controller in communication with the first and second sensors;a warning mechanism in communication with the controller;wherein the controller determines a difference between the angle of the first portion of the user's body and the angle of the second portion of the user's body indicative of an amount of movement of the first portion of the user's body relative to the second portion of the user's body, and activates the warning mechanism when the difference is beyond a predetermined angle. 2. The apparatus of claim 1, wherein the predetermined angle is an angle of flexion. 3. The apparatus of claim 1, wherein the predetermined angle is an angle of extension. 4. The apparatus of claim 1, wherein the first portion of the user's body is an upper portion of the torso and the second portion of the user's body is the pelvic region. 5. The apparatus of claim 4, further comprising a load sensor configured to be placed on the bottom of a shoe of the user, the load sensor being configured to detect the weight of a load lifted by the user, wherein the controller is in communication with the load sensor and is operable to activate the warning mechanism if the weight of the load exceeds a predetermined weight. 6. The apparatus of claim 5, wherein the controller is operable to determine a weight limit corresponding to an angle of flexion of the torso and activate the warning mechanism if the weight of the load lifted by the user exceeds the weight limit. 7. The apparatus of claim 1, wherein the warning mechanism comprises at least one of a speaker, a light-emitting diode, or a vibrator motor. 8. The apparatus of claim 1, wherein at least one of the first or second sensors comprises an inclinometer. 9. The apparatus of claim 1, wherein the controller further comprises non-transitory computer-readable storage media for storing data relating to the motion of the first and second portions of the user's body. 10. An apparatus for warning a user about motion of a first portion of the user's body relative to a second portion of the user's body, comprising: a first sensor mountable on the first portion of the user's body and configured to measure an angle of the first portion of the user's body relative to a reference;a second sensor mountable on the second portion of the user's body and configured to measure an angle of the second portion of the user's body relative to the reference;a controller in communication with the first and second sensors;a warning mechanism in communication wit... | - | G08B023/00 G08B021/02 A61B005/00 A61B005/107 A61B005/11 A63B023/02 A63B021/00 A63B021/04 A63B024/00 A63B071/06
US09801551 | B2 | 2017-10-31 | Annular vision system | 2013-07-18 | A vision system that may be used in a catheter or similar guiding instrument includes receptors distributed in an annular area. Each of the receptors has a field of view covering only a portion of an object environment, and the field of view of each of the receptors overlaps with at least one of the fields of view of the other receptors. A processing system can receive image data from the receptors and combine image data from the receptors to construct a visual representation of the entirety of the object environment. | 1. A system comprising: an elongated structure including an instrument lumen extending through the elongated structure, the elongated structure having a distal end including an annular area surrounding an unobstructed opening out of which an end of an instrument inserted through the instrument lumen may extend past the distal end of the elongated structure;a plurality of receptors, each of the receptors comprising an optical system with an optical axis passing through the annular area, the optical systems being distributed around the instrument lumen, wherein each of the receptors captures light from a field of view covering only a portion of an object environment, and wherein the field of view of each of the receptors overlaps with at least one of the fields of view of the other receptors; anda processing system coupled to receive and process image data derived from the light that the receptors capture, wherein the processing system combines the image data to construct a representation of the object environment. 2. The system of claim 1, wherein each of the optical systems comprises a lens at a distal end of an optical fiber. 3. The system of claim 2, further comprising: an image sensor, whereinthe elongated structure comprises a tube having a distal end and a proximal end, wherein the receptors are distributed on the distal end of the tube, and the optical fibers extend from the distal end to the proximal end of the tube where the optical fibers are positioned to project light onto the image sensor. 4. The system of claim 2, further comprising illumination fibers that positioned to project light into the object environment. 5. The system of claim 1, wherein each of the optical system comprises a lens at a distal end of a bundle containing a plurality of optical fibers. 6. The system of claim 1, wherein each of the receptors further comprises an image sensor chip. 7. The system of claim 6, wherein for each of the receptors, the optical axis of the optical system of the receptor is offset from a center of an active region of the image sensor chip of the receptor. 8. The system of claim 1, wherein the processing system processes the image data by stitching together overlapping images respectively from the receptors to construct an image of the object environment. 9. The system of claim 1, wherein the processing system combines the image data from the receptors to construct a stereoscopic view of the object environment. 10. The system of claim 1, wherein the processing system is operable to identify portions of the image data corresponding to components of the system and to remove from the representation the portions of the image data corresponding to the components of the system. 11. The system of claim 1, wherein the plurality of receptors comprises at least three receptors. 12. The system of claim 1, wherein the fields of view of the receptors are centered on the respective optical axes, and the optical axes are at a non-zero angle wi... | Intuitive Sugical Operations, Inc. | A61B006/00 A61B005/00 A61B001/267 A61B001/06 A61B001/00 A61B005/08
US09795423 | B2 | 2017-10-24 | Device and method for normalizing implant strain readings to assess bone healing | 2012-01-23 | A device and method for treating a bone includes a bone plate including first and second portions joined to one another via a connecting portion, a rigidity of the connecting portion being less than rigidities of each of the first and second portions in combination with a first sensor mounted on the first portion measuring strain on the first portion and a second sensor mounted on the second portion measuring strain on the second portion. | 1. A device for treating a bone, comprising: a bone plate including a first portion and a second portion joined to one another via a connecting portion, a rigidity of the connecting portion being less than rigidities of each of the first and second portions;a first sensor mounted on the first portion, the first sensor measuring strain on the first portion; anda second sensor mounted on the second portion, the second sensor measuring strain on the second portion. 2. The device of claim 1, wherein the bone plate has a length extending along a longitudinal axis of the bone plate and a width transverse to the longitudinal axis, the connecting portion having a width smaller than widths of the first and second portions. 3. The device of claim 1, wherein the connecting portion is formed as a frangible link designed to fail when a torsion of at least a predetermined threshold level is exerted on the frangible link. 4. The device of claim 1, wherein the first sensor includes one of a MEM sensor, a strain gauge and a powered chip connected to a printed circuit board. 5. The device of claim 1, wherein the first sensor includes a passively powered chip. 6. The device of claim 1, wherein the first and second sensors wirelessly provide data to an external data gathering unit. 7. The device of claim 1, further comprising a processing unit calculating a ratio of the first strain to the second strain to normalize effects of loads on the bone. | DEPUY SYNTHES PRODUCTS, INC. | A61B005/103 A61B017/80 A61B005/07 A61B005/00 A61B017/00 A61B090/00
US09788762 | B2 | 2017-10-17 | Motion monitor | 2016-04-22 | A system for monitoring the respiratory activity of a subject, which comprises one or more signal generating elements being kinematic or light emitting elements, applied to the thorax of a subject, for generating signals that are indicative of movement of the thorax of the subject; a receiver for receiving the generated signals during breathing motion of the subject; and one or more computing devices in data communication with the receiver, for analyzing the breathing motion. The more computing device is operable to calculate, in response to the received generated signals, the magnitude of a maximum displacement in 3D space of the one or more signal generating elements during a cycle of the breathing motion; and to calculate the magnitude of a current displacement in 3D space of the one or more signal generating elements during the breathing motion with respect to a reference tidal volume associated with the maximum displacement in 3D space. | 1. A system for monitoring the respiratory activity of a subject, comprising: a) one or more signal generating elements being kinematic or light emitting elements, applied to the thorax of a subject, for generating signals that are indicative of movement of the thorax of said subject;b) a camera for receiving said generated signals during breathing motion of said subject; andc) one or more computing devices in data communication with said receiver, for analyzing said breathing motion, wherein said one or more computing devices is operable to: i. calculate, in response to said received generated signals, a magnitude of a maximum displacement in 3D space of said one or more signal generating elements during a cycle of said breathing motion, wherein the maximum displacement is related to first and second reference breathing vectors associated with a same respiration cycle of said subject, said second reference breathing vector being complementary to said first reference breathing vector and originating from a maximal light emitting element location corresponding to exhalation conditions and terminating at a minimal light emitting element location corresponding to inhalation conditions;calculate a reference tidal volume that is related to said first and second reference breathing vectors;iii. store data associated with said reference tidal volume in memory;iv. calculate in a monitoring stage, for each of said one or more light emitting elements, the magnitude of a current displacement in 3D space of said one or more light emitting elements during said breathing motion related to corresponding complementary first and second monitored breathing vectors that are indicative of thorax movement and extend between a minimal light emitting element location and a maximal light emitting element location associated with a same monitored respiration cycle of said subject;v. disregard one or more of said second monitored breathing vectors if a motion deviation of recent sampled light from a single direction exceeds a predetermined threshold; andvi. calculate a value of a current tidal volume for said monitored respiration cycle by determining a function of the magnitude of reference displacements and reference tidal volumes, when the corresponding second monitored breathing vector is non-disregarded, with respect to the magnitude of a corresponding maximum displacement, and multiplying said function by the value of the corresponding reference tidal volume. 2. The system of claim 1, wherein each of the one or more kinematic signal generating elements is an accelerometer, a gyroscope, or a combination thereof. 3. The system of claim 1, wherein each of the one or more signal generating elements is a light source assembly which comprises two or more spaced light sources suitable to generate light, which are mounted on a rigid plate and fed by a power source, and the receiver is a stationary camera for tracking movement of, and imaging, said two or more light source... | BREATHEVISION LTD. | A61B005/08 A61B005/113 A61B005/087 A61B005/11 A61B090/00 A61B003/113 A61B005/024 A61B005/091 A61B005/00
US09795785 | B2 | 2017-10-24 | Cochlear implant fitting via efferent nerve fibers | 2016-01-07 | A system for frequency matching a cochlear implant during fitting includes a tissue stimulation device configured to generate a tonal stimulus to mask efferent nerve fibers in a subject, one or more response measurement contacts configured to measure CAP signals during and outside of a refractory period, a frequency matching module in communication with the response measurement contacts and configured to receive the CAP signals during and outside of the refractory period to determine a stimulation location on the cochlear implant based on a comparison of the received CAP signals both during and outside of the refractory period, and a parameter adjusting module in communication with the frequency matching module and configured to interface with the cochlear implant and to adjust its processing parameters based on the stimulation location. Methods for frequency matching a cochlear implant during fitting are also disclosed. | 1. A method of fitting a cochlear implant in an ipsilateral ear of a subject in order to frequency-match the cochlear implant, the method comprising: applying a tonal stimulus to a contralateral ear in order to mask efferent nerve fibers;measuring evoked compound action potential (CAP) signals in the ipsilateral ear during a refractory period;measuring CAP signals in the ipsilateral ear outside of the refractory period;comparing the measured CAP signals during the refractory period to the measured CAP signals outside of the refractory period in order to determine a stimulation location on the cochlear implant that corresponds to a location of maximum reduced CAP response in the ipsilateral ear for a frequency of the tonal stimulus applied to the contralateral ear; andadjusting processing parameters of the cochlear implant based on the stimulation location. 2. The method of claim 1, wherein the tonal stimulus comprises two or more tonal stimuli applied at the same or different times, each tonal stimulus having the same or different stimulation parameters. 3. The method of claim 1, further comprising applying stimulation to the ipsilateral ear when applying the tonal stimulus to the contralateral ear and before measuring the CAP signals during the refractory period. 4. The method of claim 1, wherein measuring CAP signals in the ipsilateral ear outside of the refractory period includes applying stimulation to the ipsilateral ear before measuring the CAP signals outside of the refractory period. 5. The method of claim 1, wherein comparing includes comparing the measured CAP signals during the refractory period to the measured CAP signals outside of the refractory period in order to determine the stimulation location on the cochlear implant that is between two stimulation contacts on the cochlear implant. 6. The method of claim 1, further comprising: applying the tonal stimulus to the ipsilateral ear in order to mask efferent nerve fibers;measuring the CAP signals in the contralateral ear during a refractory period;measuring CAP signals in the contralateral ear outside of the refractory period; andcomparing the measured CAP signals in the contralateral ear during the refractory period to the measured CAP signals in the contralateral ear outside of the refractory period in order to verify that the processing parameters are correctly adjusted. 7. The method of claim 1, wherein the tonal stimulus is applied by an electrical stimulation. 8. The method of claim 7, wherein the electrical stimulation is applied with a second cochlear implant in the contralateral ear. 9. The method of claim 1, wherein the tonal stimulus is applied by a mechanical stimulation. 10. The method of claim 1, wherein the tonal stimulus is applied by an acoustical stimulation. 11. A system for fitting a cochlear implant in an ipsilateral ear of a subject in order to frequency-match the cochlear implant, the system comprising: a tissue stimulation device configured to genera... | MED-EL Elektromedizinische Geraete GmbH | A61N001/36 A61B005/00
US09824698 | B2 | 2017-11-21 | Wearable emotion detection and feedback system | 2016-11-29 | A see-through, head mounted display and sensing devices cooperating with the display detect audible and visual behaviors of a subject in a field of view of the device. A processing device communicating with display and the sensors monitors audible and visual behaviors of the subject by receiving data from the sensors. Emotional states are computed based on the behaviors and feedback provided to the wearer indicating computed emotional states of the subject. During interactions, the device, recognizes emotional states in subjects by comparing detected sensor input against a database of human/primate gestures/expressions, posture, and speech. Feedback is provided to the wearer after interpretation of the sensor input. | 1. A see through head mounted display apparatus, comprising: a see-through, head mounted display;a plurality of sensors cooperating with the see-through, head mounted display to detect at least one of audible and visual behaviors of a subject in a field of view of the apparatus; andone or more processing devices in communication with the see-through, head mounted display and the sensors, the one or more processing devices:monitor the at least one of audible and visual behaviors of the subject by receiving data from the sensors;receive supplemental data relating to an emotional state of the subject from a second see-through, head mounted display also monitoring at least one of audible and visual behaviors of the subject;compute an emotional state of the subject based on the at least one of audible and visual behaviors monitored by see-through, head mounted display device and the supplemental data received from the second see-through, head mounted display; andprovide at least one of audible and visible feedback to a computing device indicating computed emotional states of the subject. 2. The apparatus of claim 1 wherein to compute an emotional state, the one or more processing devices: recognize emitted characteristics of the subject comprising each one of an emitted posture, expression, gesture, audible expression and word from the subject in the input from the sensors and compare the posture, expression, gesture, audible expression and word to a set of emotional states associated with any one or more similar posture, expression, gesture, audible expression and word. 3. The apparatus of claim 2 wherein the one or more processing devices store computed emotional states based on the characteristics relative to recognized audible and visual sensor data for each subject identified. 4. The apparatus of claim 3 wherein the one or more processing devices compute the emotional state of a subject by identifying the subject, retrieving emotional states stored relative to the subject, and comparing recognized audible and visual sensor data to a set of associated emotional states represented by the behaviors exhibited in the behaviors and the computed emotional states stored relative to the subject. 5. The apparatus of claim 1 wherein the one or more processors identify one or more subjects as an unknown subject, determine that no stored emotional state data exists for the unknown subject, and compute emotional states for each unknown subject, and associate the computed emotional states with each unknown subject thereby creating a known subject. 6. The apparatus of claim 1 wherein the one or more processors prompt a wearer of the display for identification of the subjects and are operable to receiving identification of the subjects. 7. The apparatus of claim 1 wherein the one or more processors are operable to determine the emotional state based on scenario information. 8. The apparatus of claim 7 wherein the scenario information is determined by provi... | MICROSOFT TECHNOLOGIES LICENSING, LLC | G06K009/00 G10L025/63 G06F001/16 G06F003/0482 G02B027/01 A61B005/16 A61B005/00
US09814416 | B2 | 2017-11-14 | Displays for a medical device | 2016-12-13 | Embodiments described herein relate to an analyte monitoring device having a user interface with a display and a plurality of actuators. The display is configured to render a plurality of display screens, including a home screen and an alert screen. The home screen is divided into a plurality of simultaneously displayed panels, with a first panel displays a rate of change of continuously monitored analyte levels in interstitial fluid, a second panel simultaneously displays a current analyte level and an analyte trend indicator, and a third panel displays status information of a plurality of components of the device. When an alarm condition is detected, the display renders the alert screen in place of the home screen, the alert screen displaying information corresponding to the detected alarm condition. Furthermore, the actuators are configured to affect further output of the analyte monitoring device corresponding to the detected condition. | 1. A system, comprising: an analyte sensor for positioning in fluid contact with bodily fluid under a skin surface to measure analyte levels;sensor electronics operatively coupled to the analyte sensor to receive signals corresponding to the measured analyte levels; anda receiving device communicatively coupled to the sensor electronics, the receiving device configured to: receive analyte information corresponding to the measured analyte levels;determine a set of historical analyte levels each measured at a corresponding point in time based on the analyte information;define a trigger window including the corresponding points in time of the historical analyte levels;define a range of acceptable analyte values;compare the historical analyte levels to the trigger window based at least on the range of acceptable analyte values; anddetermine whether an episode of changing analyte level is present based on the comparison of the historical analyte levels to the trigger window. 2. The system of claim 1, wherein the receiving device is further configured to provide an alert including a recommendation for a user to take a corrective action relating to a therapeutic treatment on a display to the user when the episode of changing analyte levels is determined to be present. 3. The system of claim 1, wherein the acceptable analyte values are analyte levels, and wherein the range of acceptable analyte values is a range of acceptable analyte levels. 4. The system of claim 3, wherein the receiving device is configured to associate each of the corresponding points in time of the historical analyte levels with an individual subrange of acceptable analyte levels of the range of acceptable analyte levels. 5. The system of claim 4, wherein the set of historical analyte levels includes a most recent analyte level measured at a most recent point in time, and wherein the receiving device is configured to associate each of the corresponding points in time of the historical analyte levels with the individual subrange of acceptable analyte levels so that corresponding points in time which are further away from the most recent point in time are associated with individual subranges which are larger than the individual subranges associated with corresponding points in time which are closer to the most recent point in time. 6. The system of claim 5, wherein the receiving device is configured to compare the historical analyte levels to the trigger window by determining, for each analyte level of the set of historical analyte levels, in an order starting with the most recent analyte level and finishing with a least recent analyte level, whether the analyte level is within the individual subrange associated with the corresponding point in time of the analyte level, until it is determined that a specific analyte level of the set of historical analyte levels is not within the subrange associated with the corresponding time of the specific analyte level; and wherein determining ... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/145 A61B005/026 A61B005/0285
US09788742 | B2 | 2017-10-17 | Impedance waveform monitoring for heart beat confirmation | 2015-01-29 | Implantable systems and methods directed toward improved accuracy in cardiac signal analysis. An impedance waveform is captured and used to confirm the analysis performed by the system on electrical signals or electrocardiogram. A detected heart beat from the electrocardiogram is either confirmed or identified as a misdetection depending on whether the impedance waveform shows likely correct or incorrect detection. Identified misdetection can then be corrected or otherwise mitigated. | 1. A method of operation in an implantable cardiac stimulus device having operational circuitry for monitoring and analyzing cardiac signals and a plurality of electrodes including at least two electrodes configured for current injection and at least two electrodes configured for electrical monitoring of voltage resultant from the injected current to allow the operational circuitry to generate an impedance waveform for monitoring purposes, the method comprising: the operational circuitry injecting current using the at least two electrodes configured for current injection and sensing resultant voltage using the at least two electrodes configured for electrical monitoring to generate an impedance waveform;the operational circuitry capturing an electrocardiogram (ECG) signal using a combination of the plurality of electrodes selected for sensing an ECG signal;the operational circuitry detecting an event in the ECG;the operational circuitry analyzing a portion of the impedance waveform relative to the detected event; andthe operational circuitry determining whether the detected event is an overdetection using the analysis of the impedance waveform. 2. The method of claim 1 wherein the step of analyzing a portion of the impedance waveform relative to the detected event is performed by the operational circuitry doing the following: calculating an integral of the impedance waveform between the detected event and an adjacent-in-time detected event;comparing the integral to an expected value; andif the integral does not match the expected value, determining that one of the detected event or the adjacent-in-time detected event is an overdetection; orif the integral matches the expected value, determining that a complete cardiac cycle has occurred between the detected event and the adjacent in time detected event. 3. The method of claim 1 wherein the step of analyzing a portion of the impedance waveform relative to the detected event is performed by the operational circuitry doing the following: defining a window within the impedance waveform to correlate, in time, with a peak in the ECG signal associated with the detected event;determine whether a positive peak appears in the impedance waveform in the defined window; andin response to determining there is the positive peak in the impedance waveform in the defined window, determining that the detected event is an overdetection. 4. The method of claim 1 wherein the step of analyzing a portion of the impedance waveform relative to the detected event is performed by the operational circuitry doing the following: defining a window within the impedance waveform to correlate, in time, with a peak in the ECG signal associated with the detected event;determining whether a negative peak appears in the impedance waveform in the defined window; andin response to determining there is no negative peak in the impedance waveform in the defined window, determining that the detected event is an overdetection. 5. The met... | Cameron Health, Inc. | A61B005/04 A61N001/37 A61B005/00 A61N001/365 A61B005/053 A61B005/0452
US09782105 | B2 | 2017-10-10 | Device for measuring the amount of water in a subject's body | 2015-02-13 | A device (100) for measuring the amount of water in a subject's body includes a main body part (110) formed in a linear shape and an insertion part (120) extending in a curved shape from one end of the main body part (100). The insertion part (120) includes a tip side insertion part (120-2) having a tip surface (122) to which a sensor unit (121) is fixed and a base end side insertion part (120-1), wherein the sensor unit (121) is brought into contact with a body surface of a subject through a slide mechanism. A protective member (651) liquid-tightly covers the slide mechanism between the tip side insertion part (120-2) and a base end side insertion part (120-1). The protective member (651) smoothly connects the outer peripheral surface of the tip side insertion part (120-2) and the outer peripheral surface of the base end side insertion part (120-1). | 1. A device for measuring the amount of water in a subject's body comprising: a main body part formed in a linear shape; andan insertion part extending in a curved shape from one end of the main body part, the insertion part including:a tip side insertion part having a tip surface to which a sensor unit is fixed, wherein the sensor unit is brought into contact with a body surface of a subject to measure data regarding water in a subject's body,a base end side insertion part supporting the tip side insertion part through a slide mechanism, wherein the slide mechanism allows the tip side insertion part to slide in a direction substantially perpendicular to the tip surface, anda protective member liquid-tightly covering the slide mechanism between the tip side insertion part and the base end side insertion part, wherein the protective member smoothly connects a first outer peripheral surface of the tip side insertion part, wherein the first outer peripheral surface is a first outer most surface that faces radial outward on an exterior of the tip insertion part, and a second outer peripheral surface of the base end side insertion part to each other, wherein the second outer peripheral surface is a second outer most surface that faces radial outward on an exterior of the base end side insertion part;wherein at least a first groove is formed on the first outer peripheral surface of the tip side insertion part in a first region to which the protective member is connected over the whole circumference, at least a second groove is formed on the second outer peripheral surface of the base end side insertion part in a second region to which the protective member is connected over the whole circumference, and projections formed on ends of the protective member are fitted with the first and second grooves when the protective member is connected. 2. The device for measuring the amount of water in a subject's body according to claim 1, wherein the protective member is made of one of an elastomer and a polyolefin. 3. The device for measuring the amount of water in a subject's body according to claim 1, wherein the first groove has a first depth that corresponds to the thickness of a first end of the protective member, wherein the second groove has a second depth that corresponds to the thickness of a second end of the protective member. 4. The device for measuring the amount of water in a subject's body according to claim 1, wherein the protective member extends and contracts in the slide direction. 5. The device for measuring the amount of water in a subject's body according to claim 1, wherein the sensor unit measures capacitance with a CR circuit. 6. The device for measuring the amount of water in a subject's body according to claim 1, wherein the protective member is bendable at a central position in the slide direction to release a force acting in the slide direction to a direction that differs from the slide direction. 7. The device for measuring the ... | Terumo Kabushiki Kaisha | A61B005/053 A61B005/00 G01N033/18
US09788714 | B2 | 2017-10-17 | Systems and methods using virtual reality or augmented reality environments for the measurement and/or improvement of human vestibulo-ocular performance | 2016-05-23 | A system and method for using a virtual reality or an augmented reality environment for the measurement and/or improvement of human vestibulo-ocular performance can be implemented by combining a video camera based eye orientation sensor, a head orientation sensor, a display, and an electronic circuit that connects the eye sensor, head sensor, and display. The system and method can be operated in the range of frequencies between 0.01 Hertz and 15 Hertz. The system and method can use a Fourier transform to compute a gain and a phase. The system and method can be used for measuring vestibulo-ocular reflex, dynamic visual acuity, dynamic visual stability, kinetic visual acuity, retinal image stability, or foveal fixation stability in a non-clinical setting. The system and method can be completely portable, head-worn, and self-contained. | 1. A vestibulo-ocular performance measuring device wherein: the device is configured for measuring a human physiologic vestibulo-ocular performance characteristic selected from the group of: vestibulo-ocular reflex;dynamic visual acuity;dynamic visual stability;kinetic visual acuity;retinal image stability; andfoveal fixation stability; andthe device comprises: an eye orientation sensor wherein: the eye orientation sensor is attachable to a person's head;the eye orientation sensor comprises a video camera; andthe eye orientation sensor senses vertical movement and horizontal movement of at least one eye;a head orientation sensor wherein: the head orientation sensor is attachable to the person's head;the head orientation sensor senses pitch and yaw of the person's head wherein pitch represents a rotation about a first axis representing up and down movement of the person's face when the rear of the person's head moves in the opposite direction and yaw represents horizontal movement of the face when looked at from the front about a second axis substantially aligned with the spine and perpendicular to the first axis; andthe head orientation sensor senses pitch and yaw in a range of frequencies between 0.01 Hertz and 15 Hertz;the head orientation sensor comprises a micro-electro-mechanical system integrated circuit comprising a module selected from the group consisting of an accelerometer, a magnetometer, and a gyroscope;an electronic circuit wherein: the electronic circuit comprises a battery, a central processing unit, and a memory unit;the electronic circuit is responsive to horizontal and vertical eye movement information received from the eye orientation sensor;the electronic circuit is responsive to pitch and yaw information received from the head orientation sensor;the electronic circuit uses a Fourier transform to generate a vertical gain signal and a vertical phase signal in response to the vertical eye movement information and the pitch information; andthe electronic circuit uses a Fourier transform to generate a horizontal gain signal and a horizontal phase signal in response to the horizontal eye movement information and the yaw information; anda display wherein: the display is attachable to the person's head. 2. The device of claim 1; wherein: the device is a portable head-worn device configured for use outside of a medical facility;the device is self-contained wherein self-contained comprises a configuration that measures the human physiologic vestibulo-ocular performance characteristic without the use of any external devices;the video camera is responsive to motion of an eye feature selected from the group of the pupil, the cornea, the iris, the limbus, and the retina;the eye orientation sensor senses movement of both the left eye and the right eye;the electronic circuit is responsive to both left eye movement information and right eye movement information received from the eye orientation sensor;the electronic circuit compares eye mov... | IARMOURHOLDINGS, INC. | A61B003/14 A61B003/00 A61B005/00 A61B003/032 G06T019/00 G02B027/01 G06F003/01 A61B003/113 A61B003/18 A61B005/11 A61B003/11
US09782069 | B2 | 2017-10-10 | Correcting systematic calibration errors in eye tracking data | 2014-11-06 | Systems and methods are provided for post-hoc correction of calibration errors in eye tracking data, which take into consideration calibration errors that result from changes in user position during a user session in which the user's fixations on a display screen are captured and recorded by an eye tracking system, and which take into consideration errors that occur when the user looks away from a displayed target item before selecting the target item. | 1. A method for correcting systematic error in calibration data of an eye tracking system, comprising: accessing recorded information that is associated with a given session in which a user performs a computer task while selecting known target items that are displayed on a display screen during the session and while using an eye tracking system to record locations of user fixations on the display screen as the user views the display screen during the session, wherein the recorded information comprises (i) fixation data that include start time, display screen location, and duration information for each of a plurality of fixations recorded during the session, (ii) selection event data that include start time and display screen location information for each of a plurality of user selection events recorded during the session; and (iii) known target location data that include display screen location information for each of a plurality of known target items that are displayed during the session;processing the obtained information to identify a known target item selected by the user;processing the obtained information to identify a recorded fixation, if any, which most likely corresponds to the identified known target item selected by the user, wherein a recorded fixation which was not active at a time that the identified known target item was selected by the user, is considered during said processing as a candidate fixation;computing a distance vector between the identified recorded fixation and the corresponding identified known target item; andstoring, in a data store, the computed distance vector as calibration error information for the identified recorded fixation and the corresponding identified known target item,wherein the accessing, processing, computing and storing are automated steps implemented by a computing device executing program code. 2. The method of claim 1, wherein processing the obtained information to identify a recorded fixation, if any, which most likely corresponds to the identified known target item selected by the user comprises determining a recorded fixation (i) which falls within a predefined time window prior to the time that the identified known target item was selected by the user, and (ii) which is nearest to a center region of the identified known target item selected by the user. 3. The method of claim 1, wherein processing the obtained information to identify a recorded fixation, if any, which most likely corresponds to the identified known target item selected by the user comprises determining a recorded fixation (i) which falls within a predefined time window prior to the time of selection of the identified known target item, (ii) which has a start time that is nearest in time to the time of selection of the identified known target item, but which start time is prior to the time of selection of the identified known target item by a predetermined amount of time, and (iii) which has a duration that meets a minimum ... | International Business Machines Corporation | A61B003/113 A61B003/00 A61B003/032 A61B005/16
US09791369 | B2 | 2017-10-17 | Visualization kit comprising a fluorescent agent and a cyanoacrylate, and method of cofumigation of a fluorescent agent and a cyanoacrylate | 2012-03-06 | A visualization kit includes (i) a cyanoacrylate and (ii) a fluorescent agent containing at least one tetrazine of formula 1:; wherein X and Y represent independently a nonionic, electron-withdrawing group selected from halogens, azide, cyano and O-A-R, in which A represents a single bond or a C1-C4 alkyldiyl and R is alkyl, alkenyl, alkynyl, cycloalkyl or epoxyalkyl, or an aziridine group, which is optionally substituted by alkyl or aryl. A method of cofugation and a method of visualizing traces on a substrate utilizing these components are also presented. | 1. A visualization kit comprising: (A) a cyanoacrylate; and(B) a fluorescent agent comprising at least one tetrazine of formula 1:wherein X and Y represent independently a nonionic, electron-withdrawing group selected from halogens, azide, cyano and O-A-R, in which A represents a single bond or a C1-C4 alkyldiyl and R is alkyl, alkenyl, alkynyl, cycloalkyl or epoxyalkyl; oran aziridine group, which is optionally substituted by alkyl or aryl. 2. The kit according to claim 1, wherein the fluorescent agent has a molecular weight of less than 400 g·mol−1, a sublimation temperature of less than 200° C., and an extinction coefficient of less than 2950 L·mol−1·cm−1. 3. The kit of claim 1, wherein the fluorescent agent has a brightness of more than 200 L·mol−1·cm−1. 4. The kit of claim 1, wherein X and Y are selected independently from the group consisting of halogens, azide, cyano and —OR in which R is alkyl, alkenyl, alkynyl, cycloalkyl, and epoxyalkyl. 5. The kit of claim 1, wherein the at least one tetrazine of formula 1 is selected from the group consisting of 3,6-dichloro-1,2,4,5-tetrazine, 3-chloro-6-methoxy-1,2,4,5-tetrazine, 3-chloro-6-ethoxy-1,2,4,5-tetrazine, 3,6-dimethoxy-1,2,4,5-tetrazine, 3-chloro-6-(prop-2-enyloxy)-1,2,4,5-tetrazine, 3-chloro-6-(prop-2-ynyloxy)-1,2,4,5-tetrazine, 3-chloro-6-(adamant-2-yloxy)-1,2,4,5-tetrazine, 3-chloro-6-(adamant-1-ylmethoxy)-1,2,4,5-tetrazine, 3-chloro-6-(3-epoxybutyloxy)-1,2,4,5-tetrazine, the product of formulaand mixtures thereof. 6. The kit of claim 1, wherein the cyanoacrylate is a cyanoacrylate of formula 2:in which R1 is alkyl, alkenyl, alkynyl, cycloalkyl or aryl, and R1 may be optionally substituted by a halogen or a hydroxyl group. 7. The kit of claim 1, wherein the cyanoacrylate is selected from the group consisting of methyl 2-cyanoacrylate, ethyl 2-cyanoacrylate, propyl 2-cyanoacrylate, butyl 2-cyanoacrylate, allyl 2-cyanoacrylate, methoxyethyl 2-cyanoacrylate, ethoxyethyl 2-cyanoacrylate, 2-chloroethyl 2-cyanoacrylate, cyclohexyl 2-cyanoacrylate, ethoxycarbonylmethyl 2-cyanoacrylate, trifluoroethyl 2-cyanoacrylate and mixtures thereof. 8. A composition, comprising (A) a cyanoacrylate, and(B) a fluorescent agent comprising at least one tetrazine of formula 1:wherein X and Y represent independently a nonionic, electron-withdrawing group selected from halogens, azide, cyano and O-A-R in which A represents a single bond or a C1-C4 alkyldiyl and R is alkyl, alkenyl, alkynyl, cycloalkyl or epoxyalkyl; oran aziridine group, which is optionally substituted by alkyl or aryl. 9. A method, comprising cofumigation of (A) a cyanoacrylate, and(B) a fluorescent agent comprising at least one tetrazine of formula 1:wherein X and Y represent independently a nonionic, electron-withdrawing group selected from halogens, azide, cyano and O-A-R in which A represents a single bond or a C1-C4 alkyldiyl and R is alkyl, alkenyl, alkynyl, cycloalkyl or epoxyalkyl; oran aziridine group, which is optionally subst... | CRIME SCIENCE TECHNOLOGY ECOLE NORMALE SUPERIEURE DE CACHAN CENTRE NATIONAL DE LA RECHERCHE SCIENTIFIQUE—CNRS | G01N021/64 C07D257/08 C07D403/04 A61B005/117 G01N033/52 A61B005/1172
US09787274 | B2 | 2017-10-10 | Automatic sound equalization device | 2015-10-20 | A technique for determining one or more equalization parameters includes acquiring, via one or more electrodes, auditory brainstem response (ABR) data associated with a first audio sample and determining, via a processor, one or more equalization parameters based on the ABR data. The technique further includes reproducing a second audio sample based on the one or more equalization parameters, acquiring, via the one or more electrodes, complex auditory brainstem response (cABR) data associated with the second audio sample, and comparing, via the processor, the cABR data to at least one representation of the second audio sample to determine at least one measure of similarity. The technique further includes modifying the one or more equalization parameters based on the at least one measure of similarity. | 1. A non-transitory computer-readable storage medium including instructions that, when executed by a processor, configure the processor to determine one or more equalization parameters, by performing the steps of: acquiring complex auditory brainstem response (cABR) data associated with an audio sample;comparing the cABR data to at least one representation of the audio sample to determine at least one measure of similarity; anddetermining the one or more equalization parameters based on the at least one measure of similarity. 2. The non-transitory computer-readable storage medium of claim 1, wherein comparing the cABR data to the at least one representation of the audio sample comprises: filtering the cABR data based on a plurality of frequency bands; andfor each frequency band, comparing the cABR data associated with the frequency band to the representation of the audio sample associated with the frequency band to determine at least one measure of similarity. 3. The non-transitory computer-readable storage medium of claim 2, wherein, for each frequency band, the representation of the audio sample comprises at least one of a wavelet and spectrographic data associated with the frequency band. 4. The non-transitory computer-readable storage medium of claim 2, wherein, for each frequency band, comparing the cABR data to the representation of the audio sample comprises comparing a first wavelet associated with the cABR data to a second wavelet associated with the representation of the audio sample. 5. The non-transitory computer-readable storage medium of claim 1, wherein the audio sample comprises at least one of a voice sample and a music sample. 6. The non-transitory computer-readable storage medium of claim 1, wherein determining the one or more equalization parameters based on the at least one measure of similarity comprises modifying one or more gain values, and wherein each of the gain values is associated with a different frequency band. 7. The non-transitory computer-readable storage medium of claim 1, further comprising: reproducing a second audio sample based on the one or more equalization parameters;acquiring second cABR data associated with the second audio sample;comparing the second cABR data to at least one representation of the second audio sample to determine at least one additional measure of similarity; andmodifying at least one of the one or more equalization parameters based on the at least one additional measure of similarity. 8. The non-transitory computer-readable storage medium of claim 7, wherein both the first audio sample and the second audio sample correspond to a music sample, and the one or more equalization parameters are dynamically modified as a user listens to the music sample. 9. The non-transitory computer-readable storage medium of claim 1, further comprising, prior to acquiring the cABR data: acquiring auditory brainstem response (ABR) data associated with a second audio sample; anddetermining one or m... | HARMAN INTERNATIONAL INDUSTRIES, INCORPORATED | H03G005/00 H03G005/16 G06F003/01 A61N001/36 H04R005/04 A61B005/12 H04R001/10 H04R003/04 H04R025/00 A61B005/0484
US09801953 | B2 | 2017-10-31 | Nanoparticles carrying nucleic acid cassettes for expressing RNA | 2013-10-14 | This disclosure relates to nanoparticles carrying nucleic acid cassettes for expressing RNA. In certain embodiments, the disclosure relates to improved methods for targeted delivery and expression of siRNAs in vivo using DNA-based siRNA-expressing nanocassettes and receptor-targeted nanoparticles. In certain embodiments, the disclosure relates to methods of targeted delivery of survivin siRNA expressing nanocassettes which enhance sensitivity of human cancer cells to anticancer agents. | 1. A nanoparticle comprising a core comprising iron oxide with polymer coating wherein the polymer coating is conjugated to a double stranded deoxyribonucleic acid having between about 350 and 750 base pairs that encodes a RNA capable of RNA interference in operable combination with a promoter and wherein a cell targeting molecule is amino terminal fragment (ATF) of the urokinase plasminogen activator (uPA) and is conjugated to the polymer coating, wherein a covalent bond is between the nanoparticle and the deoxyribonucleic acid, wherein the nanoparticle has a size of less than 100 nm, and wherein the RNA capable of RNA interference is RNA that forms a hairpin. 2. The nanoparticle of claim 1, wherein the RNA capable of RNA interference comprises a survivin sequence of greater than 15 nucleotides. 3. The nanoparticle of claim 1, wherein the promoter is U6 or H1. 4. The nanoparticle of claim 1 further comprising an anticancer agent. | Emory University | A61B005/055 A61K049/00 A61K047/48 A61K045/06 A61K049/18
US09820683 | B2 | 2017-11-21 | Allowing measurements to be made of a blood sample | 2016-05-31 | An apparatus and a method for measuring a property of a blood sample is shown. The apparatus comprises of a housing having an aperture; a shaft mounted inside the housing; a testing member rotatably mounted on the shaft; an actuating member coupled to the housing and configured to exert a force against the testing member; and a lancet for eliciting a blood sample fixedly coupled to and protruding substantially radially from the testing member and configured to co-rotate with the testing member. The apparatus is further configured such that the lancet is aligned with the aperture in the housing when the testing member is in a first position. The testing member is configured, in the presence of a force exerted against the testing member by the actuator member, to translate to a second position in which the lancet is in a skin penetrating position. | 1. A method of controlling a lancing and blood collection device, the method comprising: providing a shaft mounted inside a housing;providing a testing member rotatable mounted on the shaft;controlling a motor to rotate the testing member to a first position so as to align a lancet protruding substantially radially from the testing member with an aperture in the housing; andcontrolling an actuating member to exert a force against the testing member to translate the testing member to a second position in which the lancet is in a skin penetrating position. 2. A method as claimed in claim 1, further comprising controlling the motor to rotate the testing member to a third position in which a blood collection part of the testing member is aligned with the aperture. 3. A method as claimed in claim 2, further comprising controlling the actuating member to exert a force against the testing member while the testing member is in the third position so as to move the testing member laterally with respect to the shaft towards a fourth position. 4. A method as claimed in claim 2, wherein controlling the motor to rotate the testing member to the first position comprises controlling the motor to rotate the testing member in a clockwise direction. 5. A method as claimed in claim 4, wherein controlling the motor to rotate the testing member to the third position comprises controlling the motor to rotate the testing member in a counter clockwise direction. 6. A method as claimed in claim 1, further comprising: providing a swing arm within the housing, the swing arm supporting a plurality of electrical contact terminals; andcontrolling the swing arm to be rotated towards the shaft such that the plurality of electrical contact terminals contact corresponding contact pads supported on the testing member. 7. A method as claimed in claim 6, further comprising controlling an analyte interface circuit to provide voltages to the electrical contact terminals in order to perform an analyte measurement. 8. A method as claimed in claim 7, further comprising subsequent to performing the analyte measurement, controlling the motor to rotate the testing member in the counter clockwise direction until a notch on the testing member is coincident with a guide in the housing. 9. A method as claimed in claim 8, further comprising causing the testing member to rise upwards in the hosing so that a new testing member can be utilized. 10. A method as claimed in claim 1, further comprising providing an elastic member associated with the testing member or the shaft and configured to allow the testing member to translate into the second position in the presence of the force exerted against the testing member by the actuating member. 11. A method as claimed in claim 10, wherein the elastic member comprises at least part of a first annulus that surrounds a core of the shaft. | SANOFI-AVENTIS DEUTSCHLAND GMBH | A61B005/00 A61B005/151 A61B005/15 A61B005/157 A61B005/145
US09820687 | B2 | 2017-11-21 | Method for determining drowsiness | 2016-12-09 | A method for determining a drowsiness value that represents the drowsiness of a driver of a motor vehicle and is determined cyclically over defined time intervals on the basis of at least one drowsiness characteristic, includes the task of determining a drowsiness value of a time interval to be assessed, in addition to the drowsiness characteristic, at least one drowsiness base value is taken into account. | 1. A method for determining a drowsiness value that represents the drowsiness of a driver of a motor vehicle, the method comprising: determining the drowsiness value cyclically over defined time intervals on the basis of at least one drowsiness characteristic;wherein in determining a drowsiness value of a time interval to be assessed, in addition to the drowsiness characteristic, at least one drowsiness base value is taken into account;wherein in determining the drowsiness value for the time interval to be assessed, ascertained drowsiness values of prior consecutive time intervals are weighted and used for the determination, the ascertained drowsiness values of the prior intervals occurring earlier in time being given less weight than ascertained drowsiness values of the prior intervals occurring later in time;transmitting a signal to trigger a warning to the driver based on the drowsiness value, the warning being at least one of a visual warning, an auditory warning, and a haptic warning. 2. The method of claim 1, wherein an ascertained drowsiness value of a prior time interval is taken into account as drowsiness base value. 3. The method of claim 2, further comprising: determining the drowsiness value for the prior time interval, while taking account of data ascertained with respect to the drowsiness characteristic over the prior time interval;storing the drowsiness value of the prior time interval; anddetermining the drowsiness value for the time interval to be assessed, while taking into account data ascertained with respect to the drowsiness characteristic over the time interval to be assessed, as well as in due consideration of the drowsiness value stored for the prior time interval. 4. The method of claim 1, wherein in determining the drowsiness value for the time interval to be assessed, a characteristic curve of the ascertained drowsiness values of the prior time intervals is taken into account. 5. The method of claim 1, wherein the initial value is defined as a function of the time of day. 6. The method of claim 1, wherein one or more drowsiness characteristics is/are ascertained and/or evaluated in order to determine the drowsiness value for the time interval to be assessed and/or the drowsiness value of a prior time interval. 7. The method of claim 1, wherein data, especially ascertained data, is preprocessed to determine the drowsiness value. 8. The method of claim 1, wherein data, especially ascertained and/or stored and/or defined data, is processed in order to determine the drowsiness value. 9. An apparatus for determining a drowsiness value that represents the drowsiness of a driver of a motor vehicle, comprising: a control device to determine the drowsiness value cyclically over defined time intervals on the basis of at least one drowsiness characteristic;wherein in determining a drowsiness value of a time interval to be assessed, in addition to the drowsiness characteristic, at least one drowsiness base value is taken i... | ROBERT BOSCH GMBH | G08B023/00 A61B005/18 B60Q005/00 B60Q009/00 A61B005/11
US09814381 | B2 | 2017-11-14 | Ophthalmic apparatus and control method therefor, personal digital assistant device and control method therefor, ophthalmic system, and program | 2016-05-03 | An ophthalmic apparatus is capable of communicating with a PDA device provided with a display unit and includes a main body that includes an image-capturing unit, the image-capturing unit obtaining a moving image of an eye to be examined (target eye) based on light returned from the target eye which is illuminated, the moving image of the target eye used by the ophthalmic apparatus for obtaining, based on an examination of the target eye, information regarding the target eye, a driving unit that drives the main body, a transmission unit that transmits, before the information regarding the target eye is obtained, a moving image signal of the obtained moving image to the PDA device, a reception unit that receives a control signal from the PDA device during transmission of the moving image signal, and a control unit that controls the driving unit based on the received control signal. | 1. An ophthalmic apparatus that is capable of communicating with a personal digital assistant (PDA) device, the PDA device including a display unit, the ophthalmic apparatus comprising: a main body that includes an image-capturing unit, the image-capturing unit obtaining a moving image of an eye to be examined using light returned from the eye to be examined which is illuminated, the main body being for obtaining information regarding the eye to be examined;a driving unit that drives the main body;a transmission unit that transmits, before the information regarding the eye to be examined is obtained, a moving image signal of the obtained moving image, to the PDA device;a reception unit that receives a control signal from the PDA device when the moving image signal is transmitted; anda control unit that controls the driving unit using the received control signal. 2. The ophthalmic apparatus according to claim 1, wherein the reception unit receives, from the PDA device, a signal of an instruction issued on the moving image which is displayed on the display unit based on the moving image signal received by the PDA device, as the control signal. 3. The ophthalmic apparatus according to claim 1, further comprising: a sensing unit that detects that a distance between the main body and the PDA device is smaller than a predetermined distance by wireless communication,wherein the transmission unit transmits the moving image signal to the PDA device when it is determined that the distance is smaller than the predetermined distance. 4. The ophthalmic apparatus according to claim 1, further comprising: a sensing unit that detects that the distance between the main body and the PDA device is smaller than a predetermined distance as a result of the main body and the PDA device making contact with each other,wherein the transmission unit transmits the moving image signal to the PDA device when it is determined that the distance is smaller than the predetermined distance. 5. An ophthalmic apparatus that is capable of communicating with a personal digital assistant (PDA) device, the ophthalmic apparatus comprising: an eye-examination unit that performs an examination on an eye to be examined;a reception unit that receives a state-confirmation signal from the PDA device when the ophthalmic apparatus is brought into contact with the PDA device or when the ophthalmic apparatus is not in contact with the PDA device, and a distance between the ophthalmic apparatus and the PDA device is smaller than a predetermined distance; anda transmission unit that transmits, to the PDA device, a moving image signal of the eye to be examined obtained before performing the examination. 6. The ophthalmic apparatus according to claim 5, wherein the reception unit further receives an operation signal from the PDA device after the moving image signal has been transmitted to the PDA device, andwherein the transmission unit further transmits, based on the received operation signal, a... | CANON KABUSHIKI KAISHA | A61B003/14 A61B003/00 A61B003/15 A61B003/18 A61B005/00
US09801546 | B2 | 2017-10-31 | Systems for and methods of diagnosing and treating a sacroiliac joint disorder | 2015-05-27 | Aspects of the present disclosure involve a method of diagnosing and treating a sacroiliac joint of a patient comprising: a) delivering a first member into the ilium via a first posterior approach; b) delivering a second member into the sacrum via a second posterior approach; and c) diagnosing an ailment of the sacroiliac joint by: manipulating the first member relative to the second member; or identifying joint movement via a sensor positioned in or near the sacroiliac joint. | 1. A method of diagnosing and treating a sacroiliac joint of a patient, the sacroiliac joint comprising a sacrum, an ilium, a joint line, an intra-articular region, and an extra-articular region, the method comprising: a) delivering a first member into the ilium via a first posterior approach;b) delivering a second member into the sacrum via a second posterior approach; andc) diagnosing an ailment of the sacroiliac joint by manipulating the first member relative to the second member. 2. The method of claim 1, wherein manipulating the first member relative to the second member comprises rotating the first member relative to the second member. 3. The method of claim 2, wherein rotation of the first member relative to the second member positions the sacroiliac joint in nutation. 4. The method of 2, wherein rotation of the first member relative to the second member positions the sacroiliac joint in counter-nutation. 5. The method of claim 1, wherein manipulating the first member relative to the second member comprises exerting a force on one of the first member or the second member in an anterior direction while exerting a stabilizing force on the other of the first member or the second member. 6. The method of claim 1, wherein manipulating the first member relative to the second member comprises exerting a force on one of the first member or the second member in a posterior direction while exerting a stabilizing force on the other of the first member or the second member. 7. The method of claim 1, further comprises using a diagnostic tool to manipulate the first member relative to the second member. 8. The method of claim 7, wherein the diagnostic tool comprising a first arm and a second arm, the first arm configured to grasp the first member, the second arm configured to grasp the second member. 9. The method of claim 8, wherein the diagnostic tool is configured to limit a degree to which the first member is manipulated relative to the second member. 10. The method of claim 8, wherein the diagnostic tool is configured to limit translational displacement of the first member relative to the second member to substantially parallel trajectories. 11. The method of claim 8, wherein the diagnostic tool is configured to limit rotational displacement of the first member relative to the second member by a predetermined amount. 12. The method of claim 11, wherein the predetermined amount is limited to less than 6 degrees. 13. The method of claim 8, wherein the diagnostic tool is configured to limit rotational displacement of the first member relative to the second member to a single plane of rotation. 14. The method of claim 1, wherein the first member is delivered into the ilium in a first trajectory that is substantially parallel to a second trajectory of the second member being delivered into the sacrum. 15. The method of claim 14, wherein the first member and the second member are delivered cranial of the intra-articular region of the sacroi... | JCBD, LLC | A61B017/70 A61B005/00 A61B017/88 A61B017/17 A61B017/84 A61F002/30 A61B017/02
US09814389 | B2 | 2017-11-14 | Continuous analyte monitoring system | 2014-12-23 | An improved analyte monitoring system having a sensor and a transceiver with improved communication and/or user interface capabilities. The transceiver may communicate with and power the sensor. The transceiver may receive one or more analyte measurements from the sensor and may calculate one or more analyte concentrations based on the received analyte measurements. The transceiver may generate analyte concentration trends, alerts, and/or alarms based on the calculated analyte concentrations. The system may also include a display device, which may be, for example, a smartphone and may be used to display analyte measurements received from the transceiver. The display device may execute a mobile medical application. The system may include a data management system, which may be web-based. | 1. A system for detecting an amount or concentration of an analyte in vivo within a living organism, the system comprising: (A) an analyte sensor comprising: (A1) an analyte indicator configured to exhibit a detectable property based on the amount or concentration of the analyte in proximity to the analyte indicator;(A2) sensor elements configured to generate a data signal based on the detectable property exhibited by the analyte indicator; and(A3) a transceiver interface device configured to receive a power signal and generate power for powering the sensor elements and to convey data signals generated by the sensor elements;(B) a transceiver comprising: (B1) a sensor interface device configured to convey the power signal to the transceiver interface device of the analyte sensor and to receive data signals conveyed by the transceiver interface device of the analyte sensor;(B2) a processor configured to: calculate analyte concentrations based on the received data signals,calculate analyte concentration trends based on the calculated analyte concentrations,generate alerts or alarms based on the calculated analyte concentrations and the calculated analyte concentration trends, andperform an analyte concentration calibration based on one or more user-entered calibration points; and(B3) a display interface device configured to convey the calculated analyte concentrations and the calculated analyte concentration trends to a display device and receive the one or more user-entered calibration points from the display device, wherein the display interface comprises a wireless communication circuit configured to wirelessly convey the calculated analyte concentrations, the calculated analyte concentration trends, and the generated alerts or alarms to the display device; and(C) a display device configured to receive the analyte concentrations and the calculated analyte concentration trends conveyed by the display interface device of the transceiver, to display the received analyte concentrations and the calculated analyte concentration trends, and to facilitate user entry of the calibration points. 2. The system of claim 1, wherein the analyte sensor is an implantable sensor. 3. The system of claim 1, wherein the display device is a smartphone. 4. The system of claim 1, wherein the wireless communication circuit employs a wireless communication standard to wirelessly convey the calculated analyte concentrations. 5. The system of claim 4, wherein the wireless communication standard is a Bluetooth standard. 6. The system of claim 1, wherein the transceiver does not display the calculated analyte concentrations. 7. The system of claim 1, wherein the display device is configured to upload the received analyte concentrations to a web-based data management system. 8. The system of claim 7, further comprising the web-based data management system. 9. The system of claim 1, wherein the analyte sensor further comprises a sensor housing, and the sensor elements ... | Senseonics, Incorporated | A61B005/1455 A61B005/00 A61B005/145 A61B005/1459 G06F019/00
US09775555 | B2 | 2017-10-03 | Inter-ureter signaling device | 2017-01-11 | An inter-ureter signaling device intended to provide a surgeon with a signal to identify when a ureter is grasped by a surgical clamping device, such as a pair of forceps, the device being comprised of a catheter, a signal generator, and wiring, such that when the catheter component of the device, having been inserted into a ureter, is compressed by the surgical clamping device an electrical circuit is closed and the signal generator generates a human perceptible signal, indicating that the clamp has been placed on the ureter. | 1. An inter-ureter signaling device comprising a catheter,a signal generator, andwiring;wherein said catheter being configured to be inserted into a ureter of a human subject,said catheter being a first flexible hollow tube, with said first flexible hollow tube of said catheter having an outer surface and a hollow interior,said catheter having a first conducting member and a second conducting member, with said first and second conducting members each being configured to independently conduct an electrical charge, whereby the first conducting member is located in the interior of the catheter and the second conducting member is located in the interior of the catheter spaced apart from the first conducting member, andsaid catheter having a compressed state and an uncompressed state, with said compressed state of said catheter being achieved when the first flexible hollow tube of the catheter is distorted by external pressure on the outer surface of the first flexible hollow tube such that at least a portion of the first conducting member is brought into direct contact with at least a portion of the second conducting member whereby there is no structure interposed between said portion of the first conducting member and said portion of the second conducting member which are in direct contact with each other and there being no physical separation of said portion of the first conducting member and said portion of the second conducting member, and said uncompressed state of said catheter is achieved when the first conducting member is not in direct contact with the second conducting member whereby there is a physical separation between the first and second conducting members so that no electrical circuit is created between the first and second conducting members;said signal generator being capable of generating a human perceivable signal; andsaid wiring being capable of conducting an electrical current, wherein said wiring connects the first and second conducting members of the catheter to the signal generator, said wiring having a first wire and a second wire, with the first wire being in connection with the first conducting member and the signal generator and the second wire being in connection with the second conducting member and the signal generator;whereby said wiring is configured as an open circuit when the first and second conducting members of the catheter are not in direct contact with each other and said wiring is configured as a closed circuit when at least a portion of the first conducting member of the catheter and at least a portion of the second conducting member of the catheter are in direct contact with each other, such that when the wiring is configured as the closed circuit an electrical current travels through the wiring causing the signal generator to generate a human perceivable signal, and when the wiring is configured as the open circuit the electrical current is interrupted by the physical separation between the first and secon... | - | A61B005/20 A61B005/00 G08B003/00 G08B005/36
US09797920 | B2 | 2017-10-24 | Program setting adjustments based on activity identification | 2015-03-30 | An electronic device monitors accelerations using a motion sensor. The electronic device determines a current motion state based on the accelerations. The electronic device identifies a plurality of applications that subscribe to a motion state identification service and notifies a subset of the applications of the current motion state, the subset meeting notification criteria associated with the current motion state. | 1. A method of monitoring a motion state, comprising: monitoring accelerations by an electronic device using a motion sensor;determining, by the electronic device, a current motion state based on the accelerations;identifying a plurality of applications that subscribe to a motion state identification service; andnotifying a subset of the plurality of applications of the current motion state, the subset meeting notification criteria associated with the current motion state. 2. The method of claim 1, further comprising: determining whether the current motion state is different from a previous motion state; andmodifying one or more settings of one or more of the plurality of applications if the current motion state is different from the previous motion state. 3. The method of claim 1, wherein the current motion state is one of a plurality of potential motion states, the method further comprising: determining a confidence rating for the current motion state that indicates a probability that the current motion state corresponds to an actual motion state of a present user of the electronic device. 4. The method of claim 1, further comprising: identifying a plurality of potential current motion states; andidentifying confidence ratings for each of the identified potential current motion states. 5. The method of claim 1, further comprising: determining additional motion information from the acceleration measurements, the additional motion information including at least one of a user's current cadence, the user's current rolling averages of accelerations, a current dominant axis, and counted periodic human motion counts;identifying specific additional motion information the application is configured to receive; andsending the specific additional motion information to the application. 6. The method of claim 1, further comprising: identifying notification criteria associated with an application; andnotifying the application of the current motion state when the notification criteria are satisfied. 7. A computer readable storage medium including instructions that, when executed by a processor, cause the processor to perform a method comprising: monitoring accelerations by an electronic device using a motion sensor;determining, by the electronic device, a current motion state based on the accelerations;identifying a plurality of applications that subscribe to a motion state identification service; andnotifying a subset of the plurality of applications of the current motion state, the subset meeting notification criteria associated with the current motion state. 8. The computer readable storage medium of claim 7, the method further comprising: determining whether the current motion state is different from a previous motion state; andmodifying one or more settings of one or more of the plurality of applications if the current motion state is different from the previous motion state. 9. The computer readable storage medium of claim 7, wherein the current ... | DPTechnologies, Inc. | G06F019/00 G01P015/00 A61B005/11 A63B024/00 G01C022/00
US09814877 | B2 | 2017-11-14 | Biosignal measurement and electrical stimulation device using mesh structure | 2014-11-13 | A bio-electrical signal measurement and electrical stimulation device using skin resistance is provided. The biosignal measurement and electrical stimulation device includes: a pad in which a plurality of horizontal electrical wires and vertical electrical wires for electrical stimulation are formed to intersect at a constant gap; and a control module that is electrically connected to the pad to acquire body information of a patient by measuring a biosignal of intersections in which the horizontal electrical wires and the vertical electrical wires intersect and that enables generation of electrical stimulation while controlling power that is supplied to the horizontal electrical wires and the vertical electrical wires using the body information. | 1. A biosignal measurement and electrical stimulation device, comprising: a pad including a plurality of horizontal electrical wires and a plurality of vertical electrical wires at a plurality of intersecting locations; anda control module that is electrically connected to the pad adapted to acquire body state information of a patient by measuring a biosignal at the plurality of intersecting locations,wherein the biosignal represents skin resistance, andwherein the control module calculates a first average value of the skin resistance measured at the plurality of intersecting locations, selects reference points having skin resistance of the first average value or less, calculates a second average value of skin resistance of the reference points, determines a treatment point having skin resistance of the second average value or less, and generates electrical stimulation at the treatment point. 2. The biosignal measurement and electrical stimulation device of claim 1, wherein the biosignal measured by the control module is an electrical signal flowing to the plurality of intersecting locations. 3. The biosignal measurement and electrical stimulation device of claim 2, wherein the horizontal electrical wires are separated from the vertical electrical wires at each of the plurality of intersecting locations with a gap, and a predetermined potential difference between the horizontal electrical wires and the vertical electrical wires connected to the plurality of intersecting locations creates a current flow to generate the electrical stimulation. 4. The biosignal measurement and electrical stimulation device of claim 3, further comprising an insulating material to insulate the horizontal and vertical electrical wires at the plurality of intersecting locations, and a conductive material disposed at the gap for reducing resistance when generating the electrical stimulation. 5. The biosignal measurement and electrical stimulation device of claim 2, further comprising a semiconductor element disposed at each of the plurality of intersecting locations between the horizontal electrical wires and the vertical electrical wires, and the control module controls the semiconductor element to create the current flow at a desired intersecting location among the plurality of intersecting locations and to generate electrical stimulation at the desired intersecting location. 6. The biosignal measurement and electrical stimulation device of claim 1, wherein the pad comprises a plurality of intersection electrodes formed of a conductor or a conductive bonding gel disposed at the intersecting locations and a periphery part of the intersection electrodes formed of a non-conductive bonding fabric or a non-conductive bonding gel. 7. The biosignal measurement and electrical stimulation device of claim 1, further comprising perforations for air permeability in the pad. 8. The biosignal measurement and electrical stimulation device of claim 1, wherein the pad has a shap... | SMART MEDICAL DEVICE CO., LTD. | A61N001/18 A61N001/36 A61B005/053 A61N001/04
US09795323 | B2 | 2017-10-24 | Methods and systems for generation and rendering interactive events having combined activity and location information | 2015-05-15 | A computer-implemented method is described. The computer-implemented method is used for populating data of a graphical user interface (GUI). The computer-implemented method includes generating one or more graphical elements of one or more activities captured by a monitoring device. The monitoring device is usable by a user during the capturing. The method further includes generating a timeline including a time period over which the activities are performed. The timeline includes a chronological order of time within the time period. The method further includes generating an activity symbol for one or more of the activities performed during the time period. The activity symbol has an image that is graphically overlaid on the graphical elements of the activities. | 1. A monitoring device comprising: a motion sensor configured to measure an amount of movement of a user;a heart rate measurement sensor configured to measure a heart rate of the user;a time measurement device configured to determine an amount of time during which the heart rate is measured;a display device; anda processor coupled to the display device, the motion sensor, the heart rate measurement sensor, and the time measurement device, the processor configured to execute a computer readable code accessed from a memory device of the monitoring device to:identify the amount of time measured by the time measurement device, wherein the amount of time comprises:a first portion of the amount of time during which the user is sleeping, anda second portion of the amount of time during which the user is not sleeping:determine a number of calories burned by the user during the amount of time:identify the second portion of the amount of time during which the user is not sleeping based on the amount of movement, the number of calories burned by the user, and the heart rate of the user during the amount of time:determine the heart rate of the user, the number of calories burned by the user, and the amount of movement of the user during the second portion of the amount of time;classify an activity performed by the user during the second portion of the amount of time for which the user is not sleeping as a sedentary activity based on the determined heart rate of the user, the number of calories burned by the user, and the amount of movement of the user during the second portion of the amount of time;determine a non-sedentary activity to be performed by the user upon said classifying the activity as the sedentary activity, wherein the non-sedentary activity increases an activity level of the user; andgenerate a notification for presentation on the display device recommending the user to increase the activity level by performing the non-sedentary activity, wherein the display device is further configured to display an indication of a time interval during the second portion of the amount of time for which the sedentary activity has occurred. 2. The monitoring device of claim 1, wherein the processor is further configured to classify the activity as the sedentary activity upon determining that the user is not walking, not running, not moving around, or not in a vehicle. 3. The monitoring device of claim 1, wherein the motion sensor is an accelerometer. 4. The monitoring device of claim 1, wherein the monitoring device is a wrist watch, or a wristband, or a bracelet. 5. The monitoring device of claim 1, wherein the display device includes a display screen to display graphical user interface data that includes activity levels of the sedentary activity performed by the user. 6. The monitoring device of claim 1, wherein the indication of the time interval comprises a time identifier for the sedentary activity. 7. The monitoring device of claim 1, wherein the ti... | Fitbit, Inc. | G08C019/22 H04Q009/00 A61B005/11 A61B005/024 A61B005/00
US09826932 | B2 | 2017-11-28 | Automated abdominojugular reflux testing | 2016-04-19 | This document describes automated abdominojugular reflux (AJR) testing. To automate AJR tests, a pressure cuff wrapped around a person's abdomen applies pressure while video of their neck is captured. By way of example, a medical professional wraps a pressure cuff around the person's abdomen and records video of the person's neck using a smartphone, which communicates with the pressure cuff to synchronize the application of pressure with video capture. The video is processed to detect and track the response of jugular venous pulse (JVP), which is compared to AJR test thresholds to determine test results. While determining JVP, and thereby results of AJR tests, from reconstructed videos may not result in data that is as accurate as invasive intra-heart tests, it requires little if any risk to patients and is easy for medical professionals to perform. Further, these techniques enable AJR tests to be performed automatically and without relying on estimates made by skilled medical professionals. | 1. A computer-implemented method for automating an abdominojugular reflux (AJR) test, the method comprising: communicating an indication to a pressure cuff wrapped around a person's abdomen to apply pressure to the person's abdomen;capturing video of the person's neck before, during, and after the pressure is applied to the person's abdomen;processing the captured video to measure a dynamic response of a jugular venous pulse (JVP) of the person while application of the pressure is initiated, maintained, and released;comparing measurements indicative of the dynamic response to one or more thresholds associated with AJR tests to determine a result of the AJR test that is indicative of cardiac functional performance of the person. 2. The computer-implemented method as described in claim 1, wherein the one or more thresholds include an initial JVP rise threshold, the measurements include a measurement indicative of an initial rise of the person's JVP, and the comparing includes comparing the measurement indicative of the initial rise to the initial JVP rise threshold. 3. The computer-implemented method as described in claim 2, wherein a negative AJR test result is determined if the initial rise of the person's JVP does not exceed the initial JVP threshold within a set time interval after the application of pressure is initiated. 4. The computer-implemented method as described in claim 2, wherein the initial rise threshold is 3 centimeters from an initial location of the person's JVP detected before the application of pressure. 5. The computer-implemented method as described in claim 1, wherein the one or more thresholds include a sustained JVP rise threshold, the measurements include metrics indicative of sustained rises of the person's JVP, and the comparing includes comparing metrics indicative of the sustained rises to the sustained rise threshold. 6. The computer-implemented method as described in claim 5, wherein a negative AJR test result is determined if the metrics indicate that a sustained rise of the person's JVP falls below the sustained JVP rise threshold within a set time interval while the pressure is applied. 7. The computer-implemented method as described in claim 5, wherein a positive AJR test result is determined based, at least in part, on the metrics indicating that a sustained rise of the person's JVP remains above the sustained JVP rise threshold for a set time interval while the pressure is applied. 8. The computer-implemented method as described in claim 1, wherein the one or more thresholds include a JVP descent threshold, the measurements include a measurement indicative of a fall in the person's JVP after the application of pressure is released, and the comparing includes comparing the measurement indicative of the fall to the JVP descent threshold. 9. The computer-implemented method as described in claim 8, wherein a positive AJR test result is determined based, at least in part, on a comparison that indicates the ... | Google LLC | A61B005/00 A61B005/024 A61H009/00
US09820685 | B2 | 2017-11-21 | Measurement device to capture vital parameters | 2012-11-16 | The invention concerns a measurement device to capture at least one vital parameter of a person in a motor vehicle with a steering wheel. The measurement device includes a finger sensor device with an optical sensor device, where the finger sensor device is attached to the steering wheel in the transition zone between a steering wheel spoke and the hub of the steering wheel. | 1. A device for measuring a health condition parameter of a user, the device comprising: a toggle switch with at least a depressed position; anda finger sensor comprising a finger sensor surface,wherein the finger sensor is configured to measure the health condition parameter when a finger of the user is placed on the finger sensor surface,wherein the toggle switch is configured to be installed on a steering wheel of a motor vehicle,wherein the finger sensor surface comprises at least a portion of an actuating surface of the toggle switch configured to ensure a reproducible contact pressure between the finger of the user and the finger sensor surface, andwherein the toggle switch is electrically connected with and sends an activation signal to a system in the motor vehicle, where the system performs a function other than initiating measurement of the health condition parameter when the toggle switch is depressed. 2. The device of claim 1, wherein the health condition parameter is a plethysmography parameter, a pulse oximetry parameter, or an electrocardiogram parameter. 3. The device of claim 1, wherein the finger sensor is installed in a transitional zone between a spoke and a hub of the steering wheel. 4. The device of claim 3, wherein the finger sensor is installed on a rear side of the steering wheel and faces away from the user, and wherein the finger sensor surface is positioned to make contact with a finger of the user when the user grasps the steering wheel at the transitional zone. 5. The device of claim 3, wherein the finger sensor surface is installed on a front side of the steering wheel and faces towards the user, and wherein the finger sensor surface is configured to be positioned to make contact with a thumb of the user when the user grasps the steering wheel. 6. The device of claim 1, wherein the finger sensor comprises an optical sensor located on the finger sensor surface, the optical sensor comprising: a light source configured to detect the health condition parameter when the user places a finger on the optical sensor; anda light-sensitive element configured to measure the health condition parameter. 7. The device of claim 6, wherein the light source is configured to generate visible light and invisible light, and the health condition parameter comprises blood vessel volume or blood constituents. 8. The device of claim 1, wherein the finger sensor comprises a finger depression recessed on the finger sensor surface and is fingertip-shaped so as to house a fingertip of the user, and wherein an optical sensor is located in the finger depression. 9. The device of claim 1, wherein the toggle switch is configured to initiate measurement of the health condition parameter when the toggle switch is depressed. 10. The device of claim 9, wherein the finger sensor surface is additionally configured to initiate measurement of the health condition parameter when the toggle switch is in a position other than the depressed position. ... | ZF Friedrichshafen AG | A61B005/00 A61B005/1455 B60W040/08 A61B005/18 B62D001/04 A61B005/0295 A61B005/0402 A61B005/024 A61B005/0404
US09776644 | B2 | 2017-10-03 | On-vehicle situation detection apparatus and method | 2015-03-12 | An on-vehicle situation detection apparatus may include a detection unit to identify a driver and acquire driver status data and data about vehicle driving information or vehicle surrounding obstacles, a driving pattern learning unit to learn and store a driving pattern of a driver, based on the data acquired by the detection unit, a weighted value determination unit to determine a weighted value assigned to the information data acquired by the detection unit, based on the driving pattern learned by the driving pattern learning unit, a determination unit to determine a safe driving state of the driver, based on the data to which the weighted value determined by the weighted value determination unit is assigned, and a warning unit to warn the driver when the driver is determined to be not in the safe driving state. | 1. An on-vehicle situation detection method comprising: performing driver recognition of calling pre-stored driver information to compare whether the driver information coincides with a current driver of a vehicle;performing driving pattern calling of calling a pre-stored driving pattern of the driver identified by the performing driver recognition from a memory unit;performing detection using at least one sensor, of collecting driver status information and vehicle driving information or vehicle surrounding obstacle information;performing driving pattern learning of learning the pre-stored driving pattern called by the performing driving pattern calling and a driving pattern derived from the performing detection;performing weight value determination of determining a weight value for data acquired from the performing detection, based on the driving pattern learned in the performing driving pattern learning;performing determination of determining a safe driving state of the driver, based on the data to which the weight value determined in the performing weight value determination is assigned; andperforming warning of warning the driver when the driver is determined to be not in the safe driving state in the performing determination,wherein the performing weight value determination assigns a weight value determined according to impact on a driver's integrated risk index to the data acquired by the performing detection,wherein the performing weight value determination compares the acquired data and the integrated risk index with an information data reference value and an integrated risk index reference value for updating the weight value,wherein the performing weight value determination changes a weight value according to the following equation: α(n+1)=α(n), if W&lt;dW and R≧dR α(n+1)=α(n)+δ, if W≧dW α(n+1)=α(n)−δ, if W&lt;dW and R&lt;dR,where R=integrated risk index,α=weight value,α(n+1)=weight value after update,α(n)=weight value before update,W=acquired data,dR=integrated risk index reference value,dW=information data reference value, andδ=α(n)/R. 2. The on-vehicle situation detection method of claim 1, wherein the performing driver recognition comprises performing driver data acquisition of acquiring image data of pupils or faces of the current driver through an imaging device, and performing driver identification of comparing the image data of pupils and faces acquired by the performing driver data acquisition with the pre-stored driver information in the memory unit to identify the current driver. 3. The on-vehicle situation detection method of claim 1, wherein in the performing driving pattern calling a pre-stored default driving pattern is called when a driver coinciding with the pre-stored driver information in the performing driver recognition is not present. 4. The on-vehicle situation detection method of claim 1, wherein the performing detection comprises performing driver status detection of detecting one or more of a driver eyelid, a... | HYUNDAI MOBIS CO., LTD. | B60W050/14 B60W050/16 G07C005/08 B60C005/00 A61B005/18 A61B005/00
US09775587 | B2 | 2017-10-03 | Spiral biopsy device | 2013-11-22 | A biopsy device is proposed comprising a shaft and a tubular member. The shaft is movably, i.e. rotatably as well as shiftably accommodated within the tubular member and comprises a spiral at a distal end portion of the shaft, wherein the spiral is formed by a twisted sheet. Such a spiral may be considered as a laterally open thread having no solid core. | 1. A biopsy device comprising a shaft, anda tubular member, wherein the shaft is movably accommodated within the tubular member,wherein the shaft comprises a continuous spiral at a distal end portion of the shaft, wherein the continuous spiral is formed by a twisted sheet, andwherein the biopsy device further comprises a blade with a cutting edge, the blade being laterally attached to outer edges of two adjacent windings of a thread of the continuous spiral, and the blade extending proximally with a small angle relative to the longitudial axis of the continuous spiral. 2. The biopsy device of claim 1, wherein the continuous spiral comprises a sharpened distal front surface. 3. The biopsy device of claim 1, wherein the tubular member is a cutting tube having a sharpened distal edge. 4. The biopsy device of claim 1, further comprising an outer tubular member, wherein the tubular member is accommodated within the outer tubular member, and wherein the tubular member is a cutting tube having a sharpened distal end. 5. The biopsy device of claim 1, further comprising an optical fiber arranged in the tubular member, with a distal front surface of the optical fiber located at a distal end of the tubular member. 6. The biopsy device of claim 1, wherein the twisted sheet forms a double thread so that the continuous spiral comprises two windings at each cross section along a longitudinal direction of the continuous spiral. 7. The biopsy device of claim 1, wherein the biopsy device is formed as a flexible catheter, such that at least a portion of each of the shaft and the tubular member is flexible, enabling bending of the at least a portion of each of the shaft and the tubular member. 8. The biopsy device of claim 1, further comprising a console including a light source, a light detector and a processing unit for processing the signals provided by the light detector. 9. The biopsy device of claim 8, wherein one of the light source and light detector provides wavelength selectivity. 10. The biopsy device of claim 8, wherein the light source is one of a laser, a light-emitting diode or a filtered light source. 11. The biopsy device of claim 8, wherein the console further comprises one of a fiber switch, a beam splitter or a dichroic beam combiner. 12. The biopsy device of claim 8, being adapted to perform at least one out of the group consisting of diffuse reflectance spectroscopy, diffuse optical tomography, differential path length spectroscopy, optical coherence tomography and Raman spectroscopy. 13. The biopsy device of claim 8, wherein the console further comprises a monitor. 14. A biopsy device comprising: an outer tubular member;an intermediate tubular member located within the outer tubular member, such that a distal end of the intermediate tubular member extends past a distal end of the outer tubular member, wherein the intermediate tubular member comprises a sharped distal edge at the distal end of the intermediate tubular member; and a... | Koninklijke Philips N.V. | A61B010/02 A61B010/04 A61B005/00 A61B005/1459
US09819103 | B2 | 2017-11-14 | Washable intelligent garment and components thereof | 2013-03-12 | The present relates to a washable interconnection patch, a connection assembly, and an intelligent washable garment equipped therewith. The patch receives and interconnects wires to a cable. The patch comprises two matching pieces interlocking together so as to define there between two opposite apertures. One of the apertures is adapted to receive and hold the wires, and the other aperture is adapted to receive and hold the cable. One of the two matching pieces defines on an interior face a channel to interconnect the wires to the cables. The connection assembly comprises a male connector and a female connector. The male connector defines a series of independent connection points along a length thereof. The female connector is adapted to receive the male connector, and defines along a length of an inner surface thereof a series of contact points. When the male connector is inserted within the female connector, the connection points and the contact points are aligned and in contact together. | 1. A washable interconnection patch for receiving and interconnecting a plurality of wires to a cable, the patch comprising: two matching pieces interlocking together so as to define therebetween two opposite apertures, one of the apertures being adapted to receive and hold the plurality of wires and the other aperture being adapted to receive and hold the cable, an interior face of one of the two matching pieces defining a channel, the channel including a plurality of discrete connections;wherein, within the channel, each of the plurality of wires is connected to a corresponding one of the plurality of discrete interconnections; andwherein, within the channel, the plurality of discrete interconnections is connected to the cable. 2. The patch of claim 1, wherein one of the two matching pieces further comprises a fastening member for fastening the interlocked matching pieces to a fabric. 3. The patch of claim 1, wherein the two matching pieces are further adapted to be filled with a sealant. 4. A connection assembly comprising: a male connector, the male connector defining a series of independent connection points along a length thereof and having a resilient, curved extremity; anda female connector adapted to receive the male connector, the female connector defining along a length of an inner surface thereof a series of spring-loaded contact points;wherein the curved extremity of the male connector is configured to deflect the spring loaded contact points of the female connector during insertion of the male connector within the female connector until, when the male connector is fully inserted within the female connector, the connection points and the contact points are aligned and in contact together. 5. The assembly of claim 4, wherein: the male connector's cross section is shaped like a C and the connection points are located in a concave section thereof; andan entry of the female connector is C-shaped. 6. The assembly of claim 4 wherein: the male connector comprises electrical contacts plated with gold or another conductive material. 7. The assembly of claim 5, wherein the series of independent connection points of the male connector are part of a printed circuit board. 8. The assembly of claim 4, wherein the connection points and contact points are aligned in a direction of insertion of the male connector within the female connector. 9. The assembly of claim 4, wherein the connection points and contact points are electrical points. 10. The assembly of claim 4, wherein the male connector is adapted for use in a washable garment. 11. An intelligent washable garment, the garment comprising: a plurality of sensors, each of the sensors having a corresponding cable; andan electronic device for receiving signals generated by the plurality of sensors, the electronic device comprising: a casing,a connection assembly comprising a male connector and a female connector, the male connector defining a series of independent connection points along... | CARRE TECHNOLOGIES INC. | H01R011/01 H01R004/58 H02G015/013 A41D001/00 H01R013/52 H01R013/58 H02G015/00 H02G015/113 H01R012/71 A61B005/00
US09788751 | B2 | 2017-10-17 | Methods and systems for generating integrated substrate maps for cardiac arrhythmias | 2015-10-15 | An electrophysiology map, for example a map of arrhythmic substrate, can be generated by acquiring both geometry information and electrophysiology information pertaining to an anatomical region, and associating the acquired geometry and electrophysiology information as a plurality of electrophysiology data points. A user can select two (or more) electrophysiological characteristics for display, and can further elect to apply various filters to the selected electrophysiological characteristics. The user can also define various relationships (e.g., Boolean ANDs, ORs, and the like) between the selected and/or filtered characteristics. The user-selected filtering criteria can be applied to the electrophysiology data points to output various subsets thereof. These subsets can then be graphically rendered using various combinations of colorscale, monochrome scale, and iconography, for example as a three-dimensional cardiac electrophysiology model. | 1. A method of generating an electrophysiology map, comprising: acquiring geometry information pertaining to an anatomical region, the geometry information comprising position information for a plurality of points in the anatomical region;acquiring electrophysiology information pertaining to the anatomical region, the electrophysiology information comprising a plurality of electrophysiological characteristics of the anatomical region;associating the geometry information with the electrophysiology information as a plurality of electrophysiology (“EP”) data points;accepting user input to select a first electrophysiological characteristic of the plurality of electrophysiological characteristics and a first filtering criterion for the first electrophysiological characteristic;accepting user input to select a second electrophysiological characteristic of the plurality of electrophysiological characteristics and a second filtering criterion for the second electrophysiological characteristic;applying the first filtering criterion and the second filtering criterion to the plurality of EP data points;outputting a subset of the plurality of EP data points satisfying both the first filtering criterion and the second filtering criterion;accepting user input prioritizing the first electrophysiological characteristic and the second electrophysiological characteristic; andoutputting a three-dimensional graphical representation of the subset of the plurality of EP data points,wherein the three-dimensional graphical representation includes both the first electrophysiological characteristic and the second electrophysiological characteristic, andwherein a higher priority electrophysiological characteristic of the first and second electrophysiological characteristics is rendered preferentially to and to the exclusion of a lower priority electrophysiological characteristic of the first and second electrophysiological characteristics in regions of overlap between the first electrophysiological characteristic and the second electrophysiological characteristic. 2. The method according to claim 1, wherein the first electrophysiological characteristic comprises cycle length mean and the second electrophysiological characteristic comprises cycle length standard deviation. 3. The method according to claim 2, wherein the first filtering criterion comprises a band pass filter with a pass band from 110 ms to 290 ms and second filtering criterion comprises a band pass filter with a pass band from 1 ms to 30 ms. 4. The method according to claim 1, wherein the first electrophysiological characteristic and the first filtering criterion respectively comprise one of: fractionation index and a high pass filter; peak-to-peak voltage and a low pass filter; electrogram sharpness and a high pass filter; conduction velocity consistency index and a high pass filter; and conduction velocity and a low pass filter. 5. The method according to claim 4, wherein the second electrophysiologica... | St. Jude Medical, Cardiology Division, Inc. | A61B005/05 A61B005/053 A61B005/042 A61B005/0452 A61B005/0408 A61B005/00
US09775552 | B2 | 2017-10-03 | Lancing device | 2014-09-25 | A lancing device includes a housing, a lancing unit capable of reciprocating relative to the housing, an elastic member for causing the lancing unit to reciprocate, and a resistance generator for applying resistance to the lancing unit in motion. The reciprocating motion of the lancing unit includes a puncture section or interval and a retreat section following the puncture section. The puncture section includes a puncture point at which a lancing element of the lancing unit can prick the skin of a human subject, for example. The resistance generator is configured to apply a greater resistance to the lancing unit in the retreat section than in the puncture section. | 1. A lancing device comprising: a housing;a lancing unit capable of reciprocating relative to the housing, the reciprocating including an advancing motion and a retreating motion, the lancing unit being provided with a lancing element to prick skin of a subject;an elastic member for causing the lancing unit to reciprocate; anda resistance generator for applying resistance to the lancing unit;wherein a reciprocating motion of the lancing unit includes a puncture section and a retreat section following the puncture section, the puncture section including a puncture point at which the lancing element pricks the skin, andwherein the resistance generator is configured to apply a greater resistance in the retreat section than in the puncture section,wherein the resistance generator includes a protrusion and a sliding member held in sliding contact with the protrusion in the retreat section,wherein one of the protrusion and the sliding member is provided on the housing, and the other of the protrusion and the sliding member is provided on the lancing unit,wherein the sliding member has an inclined surface, and the protrusion is held in sliding contact with the inclined surface so that the resistance applied to the lancing unit becomes greater in the retreat section than in the puncture section. 2. The lancing device according to claim 1, wherein the reciprocating motion of the lancing unit includes an advance section followed by the puncture section, and the resistance generator is configured to apply a greater resistance in the advance section than in the puncture section. 3. The lancing device according to claim 2, wherein the resistance generator is configured to apply a greater resistance in the retreat section than in the advance section when a positional relation between the lancing unit and the housing is same for the retreat section and the advance section. 4. The lancing device according to claim 1, wherein the resistance generator is configured to apply zero resistance in the puncture section. 5. The lancing device according to claim 1, wherein the sliding member is more elastically deformable in a direction in which the sliding member faces the protrusion than is the protrusion. 6. The lancing device according to claim 5, wherein the sliding member includes a fixed end and a free end spaced apart from each other in a reciprocating direction of the lancing unit, the fixed end being adjacent to the protrusion when the lancing unit is at the puncture point, the free end being adjacent to the protrusion when the lancing unit is at an end point of the retreat section, and wherein the sliding member is provided with an inclined surface arranged between the fixed end and the free end and coming into sliding contact with the protrusion, the inclined surface being inclined so as to be closer to the protrusion in a direction perpendicular to the reciprocating direction as proceeding from the fixed end toward the free end. 7. The lancing device ac... | ARKRAY, Inc. | A61B005/151 A61B005/15
US09802060 | B2 | 2017-10-31 | Light therapy monitoring | 2015-09-24 | A light-monitoring apparatus includes a power source, a light detector, a computer processor coupled with the power source and in communication with the light detector and configured to receive and record light exposure detected by the light detector, an output device coupled with the computer processor, and a computer-readable medium coupled with the computer processor and storing instruction code for summing the recorded light exposure from the computer processor over time and communicating a signal to the output device to generate and communicate a signal indicating that a cumulative threshold light exposure for achieving a health benefit has been reached. The apparatus can accordingly be used by an individual to monitor cumulative light exposure from both natural and artificial sources, e.g., in the treatment of seasonal or non-seasonal depression. | 1. A method for monitoring visible-light exposure, comprising: positioning a light monitor including a visible-light detector and at least one of (a) an output device and (b) a communication device;receiving and recording visible-light exposure received from at least one light source with the visible-light detector of the light monitor;integrating the recorded visible-light exposure from each light source to produce a cumulative visible-light exposure value;comparing the cumulative visible-light exposure value with an established cumulative visible-light exposure target for a user to achieve a health benefit; andproviding active and direct feedback from at least one of:(a) the output device and(b) a mobile device transported by the user and in communication with the communication deviceto the user, wherein the feedback is indicative of the cumulative visible-light exposure value in comparison with the established visible-light exposure target to empower the user to change behavior to achieve the benefit that is a consequence of reaching the established cumulative visible-light exposure target. 2. The method of claim 1, wherein the light monitor is worn on or worn as part of the user's clothing, eyeglasses, or body. 3. The method of claim 2, wherein the established cumulative visible-light exposure target is in the range of 12-24 million lux-seconds. 4. The method of claim 1, wherein the cumulative visible-light exposure value is visually displayed by at least one of (a) the output device and (b) the mobile device. 5. The method of claim 1, further comprising: applying a minimum light-intensity threshold to the visible-light exposure to exclude visible-light illuminance below the minimum light-intensity threshold; andproviding active and direct feedback from at least one of (a) the output device and (b) the mobile device to the user indicating whether instantaneous light exposure is at least as great as the minimum light-intensity threshold to empower the user to change behavior so as to achieve the established cumulative visible-light exposure target. 6. The method of claim 1, wherein the light monitor is used by a user suffering from a condition selected from a form of depression, sleep disorder, circadian rhythm disorder, attention-deficit hyperactivity disorder, dementia, bulimia nervosa, severe premenstrual syndrome, and bipolar disorder. 7. The method of claim 6, further comprising: receiving input regarding at least one supplemental factor selected from (a) supplemental treatment received by the user, (b) user physical activity, and (c) mood polling of the user; anddetermining and communicating visible-light exposure targets to the user based, in part, on the supplemental factor. 8. The method of claim 1, further comprising recording current time of day as the visible-light exposure from the light source is recorded. 9. The method of claim 8, further comprising modifying the recorded visible-light exposure received from the light so... | SunSprite | A61N005/06 G01J001/44 H05B037/02 A61B005/024 G01J001/42
US09801684 | B2 | 2017-10-31 | System and method for performing renal denervation verification | 2016-07-27 | A renal denervation feedback method is described that performs a baseline measurement of renal nerve plexus electrical activity at a renal vessel; denervates at least some tissue proximate the renal vessel after performing the baseline measurement; performs a post-denervation measurement of renal nerve plexus electrical activity at the renal vessel, after the denervating; and assesses denervation of the renal vessel based on a comparison of the baseline measurement and the post-denervation measurement of renal nerve plexus electrical activity at the renal vessel. | 1. A renal denervation feedback method comprising: performing a baseline measurement of intrinsic renal nerve plexus electrical activity at a renal vessel by monitoring compound action potentials including baseline afferent and efferent signals;denervating at least some tissue proximate the renal vessel after performing the baseline measurement;performing a post-denervation measurement of intrinsic renal nerve plexus electrical activity at the renal vessel by monitoring post-denervation compound action potentials including post-denervation afferent and efferent signals, after the denervating;assessing denervation of the renal vessel based on a comparison of the baseline measurement and the post-denervation measurement of renal nerve plexus electrical activity at the renal vessel; anddetermining the denervation is complete when the post-denervation compound action potentials include afferent signals below a threshold value. 2. The renal denervation feedback method of claim 1, further comprising, if the denervation is not complete: repeating the steps of denervating, performing a post-denervation measurement, and assessing denervation of the renal vessel until the post-denervation compound action potentials are solely efferent signals. 3. The renal denervation feedback method of claim 2, wherein repeating the steps of denervating, performing a post-denervation measurement, and assessing denervation of the renal vessel comprises: adjusting a level of denervation for denervating at least some tissue proximate the renal vessel based on a result of assessing denervation of the renal vessel. 4. The renal denervation feedback method of claim 1, wherein performing the baseline measurement of renal nerve plexus electrical activity comprises: counting a number of baseline afferent spikes each representing an afferent compound action potential that exceeds a preset threshold during a specified period of time; andcounting a number of baseline efferent spikes each representing an efferent compound action potential that exceeds the preset threshold during the specified period of time, andwherein performing a post-denervation measurement of renal nerve plexus electrical activity comprises: counting a number of post-denervation afferent spikes each representing an afferent compound action potential that exceeds the preset threshold during the specified period of time; andcounting a number of post-denervation efferent spikes each representing an efferent compound action potential that exceeds the preset threshold during the specified period of time. 5. The renal denervation feedback method of claim 1, wherein performing a baseline measurement comprises supplying nerve stimulation to the renal vessel from one side of a denervation location for denervating at least some tissue proximate the renal vessel and measuring a baseline response of the renal vessel to the nerve stimulation on an opposite side of the denervation location; andwherein performing a post-dene... | PACESETTER, INC. | A61B018/14 A61N001/36 A61B005/00 A61B005/04 A61N001/05 A61B018/00
US09814858 | B2 | 2017-11-14 | Methods and apparatus for administering local anesthetic | 2016-01-21 | Apparatus for administering certain nerve blocks includes a sheath constructed from a flexible ultrasound echogenic material, a more rigid introducer/dilator for introducing the sheath into the patient, and an ultrasound echogenic catheter for inserting through the sheath once the distal end of the sheath is in place adjacent the nerve(s) to be blocked and the introducer/dilator has been withdrawn. The catheter has provisions at its proximal end for connecting to a source of local anesthetic. Methods for use of this apparatus are also described. | 1. A method for administering a peripheral nerve block to a patient, the method comprising: penetrating the body of the patient with a first needle to form a tunnel, wherein a distal end of the first needle emerges from the body of the patient at a predetermined marked site,placing a first sheath over the distal end of the first needle,withdrawing the first needle from the first sheath, wherein the first sheath is partially pulled through the tunnel made by the first needle,inserting a first catheter having a proximal end and a distal end through the partially withdrawn first sheath until the first catheter reaches an opposite end of the tunnel,removing the first sheath,introducing a second needle sheathed within a second sheath having a proximal end and an open distal end through which a distal end of the second needle projects into the body of the patient at the predetermined marked site,advancing the sheathed second needle into the tunnel until the sheathed second needle reaches a nerve bundle to be blocked,delivering local anesthetic,withdrawing the second needle, wherein the distal end of the second sheath is adjacent the blocked nerve bundle,inserting the distal end of the first catheter into the proximal end of the second sheath,advancing the distal end of the first catheter down the second sheath to the location of the blocked nerve bundle, andremoving the second sheath, wherein the first catheter forms a loop extending from the predetermined marked site, wherein the loop is pulled back through the tunnel by pulling on the proximal end of the first catheter, andconnecting the proximal end of the first catheter to a source of local anesthetic, andcommencing anesthesia via the first catheter. 2. The method of claim 1 further comprising providing a stimulating clip adjacent a proximal end of the second needle for coupling the second needle to a peripheral nerve stimulator to facilitate passing a current through the second needle and the patient's body to stimulate nerves in the area of the distal end of the first second needle. | Avent, Inc. | A61M019/00 A61B046/00 A61M005/42 A61M025/06 A61B001/313 A61M005/142 A61B046/20 A61M005/46 A61M005/158 A61M005/172 A61M025/00 A61B017/34 A61B005/11 A61B005/00 A61B008/08 A61M029/00
US09821814 | B2 | 2017-11-21 | Device and method for determining drowsiness and non-transitory storage medium | 2016-06-16 | A drowsiness determination device includes an input unit that receives vehicle information indicating the amount of variation in an attitude of a vehicle with respect to a road surface and head information indicating the amount of movement of a head of an occupant with respect to a predetermined position and a determination unit that determines whether the occupant is drowsy, on the basis of the vehicle information and the head information. The determination unit determines that the occupant is drowsy, on the conditions that when the vehicle attitude varies, the occupant head moves in the same direction as the direction in which the vehicle attitude has varied and that a delay time between the timing when the vehicle attitude has varied and the timing when the occupant head has moved is longer than a predetermined time. | 1. A drowsiness determination device comprising: a receiver, which in operation, receives vehicle information indicating an amount of variation in an attitude of a vehicle with respect to a road surface and body information indicating an amount of displacement movement of an upper body or a head of an occupant from a predetermined position in response to the variation in the attitude of the vehicle; anda determiner, which in operation, determines whether the occupant is drowsy, based on the vehicle information and the body information,wherein the determiner determines that the occupant is drowsy, when the vehicle attitude varies while moving in a traveling direction, the occupant upper body or the occupant head moves in a same direction as the traveling direction and a delay time between a first timing when the vehicle attitude has varied and a second timing when the occupant upper body or the occupant head has moved is longer than a first predetermined time,wherein at least one of the receiver and the determiner is included in a processor,wherein the displacement movement is a change in distance away from the predetermined position, andwherein the change in distance is measured by a movement sensor. 2. The drowsiness determination device of claim 1, wherein the determiner determines whether the occupant is drowsy, using a plurality of drowsiness levels corresponding to ranges of the delay time. 3. The drowsiness determination device of claim 1, wherein if the determiner determines that the occupant is drowsy, the determiner controls an alarm to be executed for awaking the occupant. 4. The drowsiness determination device of claim 1, wherein the delay time is a time corresponding to a time difference between a timing when the amount of variation in the vehicle attitude has reached a peak and a timing when the amount of displacement movement of the occupant upper body or the occupant head has reached a peak. 5. The drowsiness determination device of claim 1, wherein the determiner determines whether the occupant is drowsy, when an operation input to a vehicle-mounted device is not been detected within a second predetermined time from the first timing when the vehicle attitude has varied. 6. The drowsiness determination device of claim 1, wherein the delay time is calculated based on a correlation between the amount of variation in the attitude of the vehicle and the amount of displacement movement of the upper body or the head of the occupant. 7. A method for determining drowsiness, comprising: receiving vehicle information indicating an amount of variation in an attitude of a vehicle with respect to a road surface and body information indicating an amount of displacement movement of an upper body or a head of an occupant from a predetermined position in response to the variation in the attitude of the vehicle;determining whether the occupant is drowsy, based on the vehicle information and the body information; anddetermining that the occupan... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | B60W040/08 B60C009/00 G08B021/06 A61B005/18 B60K028/06 A61B005/11 G08B029/18
US09808193 | B2 | 2017-11-07 | Device and method for inducing brain injury in animal test subjects | 2014-07-24 | An apparatus and method for inflicting brain injury on a laboratory animal that employs a platform for supporting the laboratory animal. The platform defines an opening for positioning the head of the laboratory animal over the opening. A projectile is launched from a projectile launching device orientated below the opening of the platform. The projectile launching device has a means for propelling the projectile directly at and/or through the opening of said platform, thereby inflicting brain injury on the animal via either a pressure wave or concussive impact of the projectile. Without helmet, direct impact of the projectile results in severe traumatic brain injury. Use of helmet protects animals from skull fracture, subdural hematoma, intracerebral hemorrhage and contusion yet produces mild concussion-like pathology. | 1. A laboratory rodent sensor helmet comprising: a flexible substrate conformed to fit a head shape of a life size adult rodent's head and provide even pressure distribution to said rodent's head that will permit concussion to said rodent without overt pathology, said substrate having a first side and a second side;sensor films mounted on said first and said second sides of said substrate configured for measuring direct projectile impact force, wherein said sensor films measure pressure distribution and magnitude on an outer surface and on an inner surface of the helmet and wherein the helmet protects said rodent's head against skull fracture;wherein said substrate is about 0.91 mm thick and the sensor films are about 0.18 mm thick. 2. The helmet of claim 1, wherein said substrate is selected from the group consisting of woven glass/carbon, woven glass and woven glass/epoxy. 3. The helmet of claim 2, wherein the sensor films are colorimetric sensor films. 4. The helmet of claim 1, wherein the helmet is molded to fit the rodent's head. 5. The helmet of claim 1, further comprising one or more fasteners for fastening the helmet to the rodent's head. 6. The helmet of claim 1 wherein said sensor film or films are applied to or made integral with the substrate. 7. The helmet of claim 1, wherein the sensor film is a colorimetric pressure sensor film. 8. The helmet of claim 1, wherein the sensor film mounted on the first side of the substrate reveals a pressure distribution profile, contact area and/or pressure and force magnitudes that occurred between the helmet and the rodent's head and the sensor film mounted on the second side of the substrate reveals the pressure distribution profile, contact area and/or pressure and force magnitudes that occurred between helmet and the projectile. 9. The helmet of claim 1, wherein the rodent is a laboratory mouse or laboratory rat. 10. A laboratory rodent sensor helmet wherein said helmet protects said rodent against skull fracture, comprising: a flexible substrate configured to fit a head shape of the rodent, said substrate having a first side and a second side;sensor films mounted on said first and said second sides of said substrate configured for measuring projectile impact;wherein the sensor film mounted on the first side of the substrate reveals a pressure distribution profile, contact area and/or pressure and force magnitudes that occurred between the helmet and the rodent's head and the sensor film mounted on the second side of the substrate reveals the pressure distribution profile, contact area and/or pressure and force magnitudes that occurred between helmet and the projectile;wherein said substrate is about 0.91 mm thick and the sensor films are about 0.18 mm thick. 11. The helmet of claim 10, wherein the sensor films are colorimetric pressure sensor films. 12. A method of measuring projectile impact to a laboratory rodent's head comprising; a) providing a sensor helmet wherein said helmet p... | The United States of America, as represented by the Secretary of the Army | F41B011/32 A61B005/00 F41B011/60 A61B005/103 A01K029/00 G01L005/00 A22B003/02
US09775558 | B2 | 2017-10-03 | Systems and methods for treating sleep apnea | 2012-03-13 | Sleep apnea can be treated using positive airway pressure. Methods and systems for determining a level of airway obstruction allow beneficial adjustments to the level of expiratory positive airway pressure used to treat a subject. | 1. A system for treating sleep apnea of a subject having an airway by enhancing expiratory positive airway pressure, the system comprising: (a) a pressure generator configured to generate a pressurized flow of breathable gas for delivery to the airway of the subject,(b) one or more sensors configured to generate one or more output signals conveying measurements related to one or more parameters of the subject; and(c) one or more processors configured to execute computer program modules, the computer program modules comprising: (1) a control module configured to control the pressure generator to adjust one or more parameters of the pressurized flow of the breathable gas such as that the airway of the subject is provided with an inspiratory positive airway pressure during inhalation and a target expiratory positive airway pressure during exhalation of the subject, the inspiratory positive airway pressure being greater than the target expiratory positive airway pressure,(2) a resistance module configured to estimate upper airway resistance of the subject based on the one or more output signals, and(3) and adjustment module configured to periodically adjust the target expiratory positive airway pressure based on the estimated upper airway resistance, and to determine whether to reverse or accept an individual periodical adjustment based on subsequently estimated upper airway resistance such that the target expiratory positive airway pressure forms a moving target expiratory positive airway pressure comprising a level of expiratory positive airway pressure that facilitates a reduced amount of upper airway resistance, wherein the adjustment module is configured to determine whether to reverse or accept the individual periodical adjustment of the target expiratory positive airway pressure by comparing a change in estimated upper airway resistance with a threshold. 2. The system of claim 1, further comprising an infrasonic wave generator configured to generate infrasonic pressure oscillations at or near the airway of the subject, wherein one or more output signals conveying measurements are further related to a physiological response by the subject to the infrasonic pressure oscillations, wherein the physiologic response includes flow oscillations, and wherein estimation by the resistance module is further based on a ratio of an amplitude of the pressure oscillations and an amplitude of the flow oscillations. 3. The system of claim 1, wherein the threshold for change in estimated upper airway resistance is a percentage of the estimated upper airway resistance. 4. A method of treating sleep apnea of a subject having an airway by enhancing expiratory positive airway pressure, the method comprising; delivering, with a pressure generator, a pressurized flow of breathable gas having an inspiratory positive airway pressure during inhalation and a target expiratory positive airway pressure during exhalation to the airway of the subject, the inspiratory posit... | KONINKLIJKE PHILIPS N.V. | F16K031/02 A61B005/00 A61B005/085 A61B005/087 A61F005/56 A61M016/00 A61B005/08 A61M016/16 A61M016/10
US09801570 | B2 | 2017-10-31 | Auditory stimulus for auditory rehabilitation | 2012-06-22 | Generally, a method performed by one or more processing devices includes generating a graphical user interface that when rendered on a display of the one or more processing devices renders a visual representation of an environment and a visual representation of an object in the environment; retrieving an auditory stimulus with one or more auditory attributes indicative of a location of a virtual target in the environment; receiving information specifying movement of the object in the environment; determining, based on the movement of the object, a proximity of the object to the virtual target; adjusting, based on the proximity, one or more values of the one or more auditory attributes of the auditory stimulus; and causing the one or more processing devices to play the auditory stimulus using the adjusted one or more values. | 1. A method performed by one or more processing devices, the method comprising: starting a program;generating data for a graphical user interface that when rendered on a display of a display device renders a visual representation of an object;identifying an auditory stimulus signal indicative of a location of a virtual target in the graphical user interface;receiving data representing a location of the visual representation of the object in the graphical user interface;intermittently determining whether a location of the visual representation of the object in the graphical user interface corresponds to the location of the virtual target until the location of the visual representation of the object in the graphical user interface corresponds to the location of the virtual target;while the location of the visual representation of the object in the graphical user interface does not correspond to the location of the virtual target, determining, based on processing of the data representing the location of the visual representation of the object, a proximity of the visual representation of the object to the virtual target in the graphical user interface;intermittently modifying, based on the proximity, the auditory stimulus signal by modifying a frequency of the auditory stimulus signal, a signal-to-noise ratio of the auditory stimulus signal, a sound level of the auditory stimulus signal, a binaural characteristic of the auditory stimulus signal, or a modulation of a sound envelope of the auditory stimulus signal; andcausing one or more speaker devices to play the modified auditory stimulus signal;when the location of the visual representation of the object in the graphical user interface corresponds to the location of the virtual target, completing the program. 2. The method of claim 1, further comprising: detecting, following play of the auditory stimulus, that a location of the visual representation of the object in the graphical user interface corresponds to the location of the virtual target in the graphical user interface. 3. The method of claim 1, wherein the actions of receiving, determining, and intermittently modifying are repeated until a location of the visual representation of the object in the graphical user interface corresponds to the location of the virtual target in the graphical user interface. 4. The method of claim 1, further comprising: causing a diagnostic auditory signal to be played on the one or more speaker devices;receiving information indicative of a response of a user to the diagnostic auditory signal;determining, at least partly based on the received information, auditory dysfunction in one or more types of auditory perception of the user;determining, based on the determined auditory dysfunction, the frequency; anddetermining, based on one or more amounts of auditory dysfunction in the one or more types of auditory perception, one or more ranges of values for the frequency. 5. The method of claim 4, wherein the one o... | Massachusetts Eye & Ear Infirmary | A61B005/00 A61B005/12 A61M021/00 G06F019/00 G09B005/06 G09B019/00
US09823182 | B2 | 2017-11-21 | Elastic wave receiving apparatus, elastic wave receiving method, photoacoustic apparatus, and program | 2010-12-09 | An elastic wave receiving apparatus includes: a probe that receives an elastic wave generated from a subject; a plate-like compression plate that supports the subject and whose surface is scanned by the probe; a motor for driving the probe; a controller that supplies a drive signal to the motor so that the probe moves to a predetermined target position on the compression plate; and a load estimating unit that preliminarily acquires and stores a physical value corresponding to a load generated at the time of scanning the compression plate by the probe. The controller corrects the drive signal so that the probe moves to the target position regardless of the load by using the physical value stored in the load estimating unit. | 1. An acoustic wave receiving apparatus comprising: a supporting member configured to support a subject;a scanning light source located opposing the subject via the supporting member and configured to irradiate the subject with pulsed light so as to generate a photoacoustic wave at the subject;a scanning probe acoustically and slidably in contact with the supporting member via an acoustic impedance matching agent and configured to receive an acoustic wave propagated from the subject via the supporting member and output an electric signal in response to receipt of the photoacoustic wave;a driving unit configured to move said scanning probe and said scanning light source along the supporting member;a drive controlling unit that supplies a drive signal to said driving unit such that said scanning probe is moved; andan information acquiring unit configured to, preliminary to acquiring data from the subject, acquire and store a physical value corresponding to a load generated due to an acoustic connection between the scanning probe and the supporting member,wherein said drive controlling unit corrects the drive signal by using the physical value stored in said information acquiring unit to correct for a scanning inaccuracy due to the load such that said scanning probe moves in association with a position of the scanning light source, simultaneously. 2. The acoustic wave receiving apparatus according to claim 1, wherein said information acquiring unit has a position detecting unit that detects a position of said scanning probe, said scanning of the scanning probe is performed in advance, and the difference between a position of said scanning probe detected and the predetermined target position is stored in said information acquiring unit as a physical value corresponding to the scanning load. 3. The acoustic wave receiving apparatus according to claim 2, wherein, at the time of acquiring the physical value, said scanning of the scanning probe is performed in a state in which no acoustic wave is generated from the subject. 4. The acoustic wave receiving apparatus according to claim 2, wherein, when the scanning load corresponding to the physical value is larger than a scanning load which is assumed from the predetermined target position, said information acquiring unit increases driving speed or hastens drive timing of said scanning probe. 5. The acoustic wave receiving apparatus according to claim 2, wherein, when the scanning load corresponding to the physical value is smaller than a scanning load which is assumed from the predetermined target position, said information acquiring unit decreases driving speed or delays drive timing of said scanning probe. 6. The acoustic wave receiving apparatus according to claim 1, wherein said information acquiring unit has a temperature detecting unit that detects a temperature of said supporting member, and the detected temperature is stored in said information acquiring unit as a physical value corresponding... | CANON KABUSHIKI KAISHA | A61B005/00 G01N021/17
US09826921 | B2 | 2017-11-28 | Detection of hypokinetic and hyperkinetic states | 2009-06-12 | The present invention relates to an automated method of determining a kinetic state of a person. The method obtains accelerometer data from an accelerometer worn on an extremity of the person and processes the accelerometer data to determine a measure for the kinetic state. The present invention further relates to a device for determining a kinetic state of a person. The device comprises a processor configured to process data obtained from an accelerometer worn on an extremity of the person and to determine from the data a measure for the kinetic state. In the method and system the kinetic state is at least one of bradykinesia, dyskinesia, and hyperkinesia. | 1. An automated method of determining a bradykinetic state of a person, which is a state characterized by slowness of movement relative to a normally kinetic person and in which movements involving peaks of acceleration have strong low frequency components, the method comprising: obtaining a time series of accelerometer data in real-time from an accelerometer worn on an extremity of the person, the time series of accelerometer data being representative of movement of the person over an extended period of time, wherein the time series of accelerometer data is obtained while the accelerometer is oriented to be sensitive to at least one of pronation and supination movements of the person and is attached to at least a limb of the person;filtering, by a band pass filter, the accelerometer data to extract filtered data for a band of interest, wherein the band of interest has a lower end cut off frequency selected to remove DC;extracting, by a computer processor, a plurality of bins of acceleration data from the time series of the filtered accelerometer data;searching, by the computer processor, the data in each of the plurality of bins to identify a time window of peak acceleration, the time window within a given bin comprising at least two data samples that capture accelerometer data for a single movement having a peak acceleration relative to other movements observed within the plurality of bins; andprocessing, by the computer processor, data in the time windows of peak acceleration to determine: spectral components for each of the single movements having peak acceleration; andan objective measure of the extent to which the single movements having peak acceleration identified within the plurality of bins of acceleration data are slower than expected from a normally kinetic person and indicative of bradykinesia based upon a comparative statistical analysis of the determined sets of spectral components of the single movements of peak acceleration that occur over the extended period of time. 2. The method according to claim 1, wherein the plurality of bins of acceleration data are of a time duration selected to be small enough that relatively regular measures of the extent to which the single movements having peak acceleration are slower than expected from a normally kinetic person and indicative of bradykinesia are determinable, while being long enough to provide a reasonable likelihood of a significant movement by the person during that bin. 3. The method according to claim 2, wherein the plurality of bins have durations within the range of 15 seconds to four minutes. 4. The method according to claim 1, wherein the searching comprises using a moving mean having a window length in a range of 0.02 seconds to 30 seconds in order to find a maxima. 5. The method according to claim 1, further comprising obtaining a sub-bin which comprises a plurality of data points both before and after a peak acceleration PKi, and wherein the plurality of data points i... | Global Kinetics Corporation Limited | A61B005/103 A61B005/11 A61B005/00
US09814438 | B2 | 2017-11-14 | Methods and apparatus for performing dynamic respiratory classification and tracking | 2013-06-18 | A method for performing dynamic classification of a breathing session is disclosed. The method comprises capturing breathing sounds of a subject using a microphone. Further, it comprises recognizing a plurality of breath cycles and a plurality of breath phases within each of the plurality of breath cycles from the breathing sounds. It also comprises detecting characteristics regarding the plurality of breath cycles and the plurality of breath phases. Finally, it comprises extracting metrics concerning a breath pattern quality of the subject using the detected characteristics. | 1. A method for performing acoustic dynamic classification of a breathing session using a processor coupled to a memory, said method comprising: capturing breathing sounds of a subject using a microphone;processing said breathing sounds to generate an audio respiratory signal;recognizing a plurality of breath cycles and a plurality of breath phases within each of said plurality of breath cycles from said audio respiratory signal;detecting characteristics regarding said plurality of breath cycles and said plurality of breath phases; andextracting and outputting metrics concerning a breath pattern quality of said subject using said characteristics wherein said metrics are selected from a group consisting of: respiratory rate; depth; tension; nasal wheeze; tracheal wheeze; pre-apnea; apnea; ramp; flow; variability; and inhale/exhale ratio;wherein said recognizing comprises: obtaining a first audio envelope by filtering an audio respiratory signal, wherein said obtaining comprises calculating a spectral centroid of each block of said audio respiratory signal and filtering said audio respiratory signal with a low pass filter tuned to a minimum value of said spectral centroid;classifying lobes of said first audio envelope into a plurality of classes; anddefining said plurality of breath cycles and said plurality of breath phases using timestamps obtained from said classifying. 2. The method of claim 1, wherein each breath phase within said plurality of breath phases is selected from a group consisting of: inhale; transition; exhale and rest. 3. The method of claim 1, wherein each class within said plurality of classes is selected from a group consisting of: inhalations; exhalations; rest periods. 4. The method of claim 1, wherein said classifying comprises: calculating a second audio envelope using a window, wherein said window has a size set according to a respiratory rate;normalizing said second audio envelope;detecting peaks of said second audio envelope;determining a start timestamp and an end timestamp of each breath cycle using said peaks; andcalculating a low threshold and a high threshold by using a moving average filter on said second audio envelope, wherein said low threshold detects signal presence, and further wherein said high threshold discriminates between inhalation and exhalation events. 5. The method of claim 4, wherein calculating said respiratory rate comprises: calculating an auto-correlation function (ACF) of said first audio envelope; anddetermining a peak of said ACF; andusing said peak to determine said respiratory rate. 6. The method of claim 5, wherein said respiratory rate is used to determine a ventilatory threshold wherein said ventilatory threshold occurs when a rate of change of said respiratory rate is greatest. 7. The method of claim 1, wherein said characteristics are selected from a group consisting of: inhale duration, pause duration, exhale duration, rest duration, a wheeze source, a wheeze type, a cough type... | BREATH RESEARCH, INC. | A61B007/00 A61B007/04 A61B005/08 A61B005/00
US09821141 | B2 | 2017-11-21 | Macroporous catheter | 2013-03-08 | The present invention provides a catheter for use in delivering and/or recovering materials to and/or from a tissue site in the body, the catheter comprising one or more macroporous regions adapted to selectively deliver molecules and/or recover cells from the tissue site, based on one or more physical-chemical-biological characteristics. The macroporous region can be provided in the form of a helical hollow wires, and can be adapted to both recover cells from the tissue site (e.g., bacterial or other cells present in interstitial fluid) and optionally to permit the infusion of medicament to the site, under corresponding conditions. | 1. A catheter for use in delivering and/or recovering materials to and/or from a tissue site in a body, the catheter comprising: a macroporous region formed of a helical hollow strand having a proximal end and a distal end and providing a lumen, the macroporous region being coupled to a catheter body at or near its proximal end; anda tube extending within the lumen of the helical hollow strand from the proximal end to the distal end of the helical hollow strand, the tube comprising a wall having a plurality of openings through the wall of the tube. 2. A catheter according to claim 1, wherein the macroporous region comprises a helical hollow wire strand. 3. A catheter according to claim 2, wherein the plurality of openings through the wall of the tube are distributed along a longitudinal extent of the tube and wherein the macroporous region and the tube are adapted to permit the even flow of material through the helical hollow wire strand. 4. A catheter according to claim 3 wherein the catheter further comprises a sharp end piece securely attached to a distal end of the catheter and adapted for penetrating the tissue. 5. A catheter according to claim 4 wherein the tube comprises a support tube having sufficient rigidity to facilitate the penetration of the catheter into the tissue site. 6. A catheter according to claim 5 wherein the end piece is adapted to be sufficiently sharp and wherein the inner tube is adapted to be sufficiently rigid for the catheter to be a self-introducing catheter which may penetrate and be inserted into the tissue without the use of an introducer. 7. A catheter according to claim 1, wherein the catheter comprises at least one macroporous region that is substantially non-inflatable. 8. A catheter according to claim 1 wherein the tube comprises a rigid support tube. 9. A catheter for use in delivering materials to a tissue site in a body, the catheter comprising: a macroporous region formed of a helical hollow wire strand having a proximal end and a distal end and providing a lumen, the macroporous region being coupled to a catheter body at or near its proximal end, anda tube extending within the lumen of the helical hollow wire strand from the proximal end to the distal end of the helical hollow strand, the tube comprising a wall having a plurality of openings through the wall, the openings distributed along a longitudinal extent of the tube to permit the even flow of material through the helical hollow wire strand to the tissue site. 10. A catheter according to claim 9, wherein the tube comprises a support tube having sufficient rigidity to facilitate the penetration of the catheter into the tissue site. 11. A catheter according to claim 10 further comprising a sharp end piece securely attached to a distal end of the catheter adapted for penetrating the tissue. 12. A catheter according to claim 11 wherein the end piece is adapted to be sufficiently sharp and wherein the inner tube is adapted to be sufficiently... | Twin Star Medical, Inc. | A61M025/00 A61B010/00 A61L029/00 A61M027/00 A61B005/00 A61M001/28 A61M031/00
US09792823 | B2 | 2017-10-17 | Multi-view learning in detection of psychological states | 2014-09-15 | Systems and methods for using multi-view learning to leverage highly informative, high-cost, psychophysiological data collected in a laboratory setting, to improve post-traumatic stress disorder (PTSD) screening in the field, where only less-informative, low-cost, speech data are available. Partial least squares (PLS) methods can be used to learn a bilinear factor model, which projects speech and EEG responses onto latent spaces, where the covariance between the projections is maximized. The systems and methods use a speech representation based on a combination of audio and language descriptors extracted from spoken responses to open-ended questions. | 1. A system for detecting psychological states, the system comprising: a device to obtain sensor data corresponding to a subject's response to stimuli;a feature extraction processor to extract one or more features from the sensor data to generate feature representation data, the feature representation data being associated with a first view;a memory device configured to store a projection matrix modeling a relationship between the first view and a second view, and to store a detection model trained using training data associated with the first view transformed to a latent space using the projection matrix, wherein the first view is associated with low-cost data and the second view is associated with high-cost data; anda detection module to transform the feature representation data to the latent space using the projection matrix, and to apply the transformed feature representation data to the detection model to detect one or more psychological states of the subject,wherein the device to obtain sensor data comprises a first device to obtain first sensor data corresponding to a training subject's response to the stimuli and a second device to obtain second sensor data corresponding to the training subject's response to the stimuli, wherein the feature extraction processor is a first feature extraction processor, wherein the feature representation data is first feature representation data, the system further comprising:a second feature extraction processor to extract one or more features from the second sensor data to generate second feature representation data, the second feature representation data being associated with the second view;a relationship learning module to receive the first and second feature representation data and to learn the projection matrix for projecting the first feature representation to the latent space related to the first and second views; anda training module to transform the first feature representation data to the latent space using the projection matrix, and to train the detection model using the transformed feature representation data. 2. The system of claim 1 wherein the one or more psychological states include post-traumatic stress disorder (PTSD) indicators. 3. The system of claim 1 wherein the sensor data comprises audio signal data responsive to the subject's speech. 4. The system of claim 3 wherein the feature representation data comprises an audio representation and a language representation. 5. The system of claim 4 wherein the feature extraction processor comprises: a frame-level feature extraction processor to receive the audio signal data and to generate frame-level feature vectors;an audio feature extraction processor to generate the audio representation using the frame-level feature vectors;an automatic speech recognition (ASR) processor to apply a speech recognizer to the frame-level feature vectors to transcribe the subject's speech; anda distress indicator extraction processor to extract psychologic... | Raytheon BBN Technologies Corp. | G09B005/00 A61B005/16 G10L025/63
US09814391 | B2 | 2017-11-14 | Intra-extra oral shock-sensing and indicating systems and other shock-sensing and indicating systems | 2014-04-17 | A mouth guard comprises a base member configured to fit inside the mouth of a user, and at least one shock-sensing and indicating device coupled to the base member. In one exemplary embodiment, the shock-sensing and indicating device is a passive shock-sensing and indicating device that detects a shock substantially along a selected axis with respect to the base member. In another exemplary embodiment, the at least one shock-sensing and indicating device detects a shock substantially along a plurality of selected axes with respect to the base member, each selected axis being substantially orthogonal from another selected axis. The shock-sensing and indicating devices can be configured to detect different levels of shock. In one exemplary embodiment, the shock-sensing and indicating device comprises a multi-component chemical-reaction system, such as a chemi-luminescent reaction system. | 1. A mouth guard, comprising: a base member configured to fit inside a mouth of a user;at least one shock-sensing device directly attached to the base member, the at least one shock-sensing device to detect a level of shock applied to the base member that exceeds a predetermined level of shock; andan indicating device directly attached to the base member and coupled to the at one shock-sensing device to indicate that the shock-sensing device has detected a level of shock that has been applied to the base member that exceeds the predetermined level of shock, the indicating device comprising a light or a light emitting diode. 2. The mouth guard according to claim 1, wherein the at least one shock-sensing and indicating device detects a shock substantially along a selected axis with respect to the base member. 3. The mouth guard according to claim 1, wherein the at least one shock-sensing and indicating device detects a shock substantially along a plurality of selected axes with respect to the base member, each selected axis being substantially orthogonal from another selected axis. 4. The mouth guard according to claim 1, wherein the at least one shock-sensing device comprises at least one accelerometer to detect a level of shock applied to the base member that exceeds a predetermined shock level. | - | A61C005/14 A61C019/04 A61B005/00 A63B071/08 A63B071/06 A63B071/10
US09775583 | B2 | 2017-10-03 | Method and device for positioning a Doppler ultrasound transducer for blood flow measurement and a system for blood flow measurement | 2012-02-27 | The method of positioning a Doppler ultrasound transducer for performing blood flow measurement according to the invention comprises the steps of: detecting a pressure oscillation signal from an inflated cuff placed on patient's artery; detecting an ultrasound pulse signal from the Doppler ultrasound transducer placed along the artery; deriving a first signal from the pressure oscillation signal and the ultrasound pulse signal, the first signal indicating the degree of synchronization between the pressure oscillation signal and the ultrasound pulse signal; and outputting an indication signal to indicate the Doppler ultrasound transducer is in a desired position when the first signal satisfies a predefined condition. Since the synchronization property of the cuff pressure oscillation signal and the ultrasound signal caused by the blood flow is utilized to determine whether the transducer is well positioned or not, ultrasound signal, which is a pulse signal but not reflecting the blood flow of the artery, could be determined as not in synchronization with the oscillation signal and therefore the accuracy of the positioning could be improved. | 1. A method of positioning a Doppler ultrasound transducer for blood flow measurement, comprising: with a pressure sensor, detecting a pressure oscillation from an inflated cuff placed on an artery of a patient and outputting a pressure oscillation signal;with a Doppler ultrasound transducer placed along the artery, detecting an ultrasound pulse and outputting an ultrasound pulse signal;with at least one electronic processor, deriving, from the pressure oscillation signal and the ultrasound pulse signal, a first signal indicating the degree of synchronization between the pressure oscillation signal and the ultrasound pulse signal, the first signal being derived from the time intervals measured between two consecutive peaks of the pressure oscillation signal Δtosi(i=1 . . . n) and the time intervals measured between two corresponding consecutive pulses of the ultrasound pulse signal Δtusi(i=1 . . . n); andon a display, indicating that the Doppler ultrasound transducer is in a desired position in response to the first signal satisfying a predefined condition including(Δtusi−Δtosi)/Δtosi≦x %, wherein n is at least two and x is a preselected percent. 2. The method of claim 1, wherein the at least one electronic processor is programmed to: derive the first signal by computing (Δtusi−Δtosi)/Δtosi, and the first signal satisfies the predefined condition in response to (Δtusi−Δtosi)/Δtosi≦8%. 3. The method of claim 2, wherein n is five. 4. The method of claim 1, wherein detecting a pressure oscillation signal from the inflated cuff placed on an artery of a patient comprises, with the at least one electronic processor: obtaining the pressure signal from the pressure sensor connected with the inflated cuff;determine whether the pressure oscillation signal is detected by the Doppler ultrasound transducer; andincreasing the pressure in the inflated cuff in response to the pressure oscillation signal not being detected. 5. The method of claim 4, inflating the cuff to about 80 mmHg. 6. The method of claim 4, wherein detecting the ultrasound pulse signal from the Doppler ultrasound transducer placed along the artery comprises: determining, with the Doppler ultrasound transducer, the pulse and outputting the ultrasound pulse signal in response to the pressure oscillation signal being detected. 7. The method of claim 6, further comprising: with the at least one electronic processor, outputting a prompting signal to the display prompting the operator to relocate the Doppler ultrasound transducer in response to the ultrasound signal from the Doppler ultrasound transducer is not a pulse. 8. The method of claim 1, wherein outputting the indication signal comprises: turning on an indication lamp of the Doppler ultrasound transducer or changing the color of the Doppler ultrasound transducer on the display of the Doppler ultrasound device. 9. A device for positioning a Doppler ultrasound transducer for blood flow measurement, comprising: at least one electronic ... | KONINKLIJKE PHILIPS N.V. | A61B008/06 A61B005/022 A61B008/00 A61B008/08
US09813270 | B2 | 2017-11-07 | Heart rate sensor and medical diagnostics wireless devices | 2016-04-16 | A heart generated signal is provided by a heart sensor of a mobile device to an analog to digital (A/D) converter for A/D converting the sensor provided signal. The A/D converted heart signal is processed to provide heart rate. The heart rate is recorded or stored in the mobile device or is transmitted in a wireless communication system. The mobile device receives sensor provided Electro Cardiogram (ECG) signal. The ECG signal is stored or is provided to an interface unit. The mobile device has transceivers for receiving and transmitting Orthogonal Frequency Division Multiplexed (OFDM) signals and for modulating and transmitting spread spectrum baseband signals. The spread spectrum baseband signals have cross-correlated in-phase and quadrature-phase filtered baseband signals. | 1. A method comprising steps of: generating in a camera of a mobile device a photo signal and processing said photo signal into a processed baseband camera generated photo signal;receiving a location finding modulated signal and demodulating said location finding signal, in said mobile device, into a location finding baseband signal;providing said location finding baseband signal with said processed baseband camera generated photo signal to a first modulator for modulating said location finding baseband signal with said processed baseband camera generated photo signal into a first modulated location finding signal with processed photo signal; andproviding said modulated location finding signal with said processed photo signal to a first transmitter for transmission of said modulated location finding signal with said processed photo signal;receiving in a mobile device from a heart rate sensor a signal, said heart rate sensor is in said mobile device;providing said received heart rate sensor signal to a processor for processing said received heart rate sensor signal into heart rate information;providing said heart rate information to an interface unit of said mobile device, or providing said heart rate information to said first modulator or to a second modulator for modulating said heart rate information into a second modulated signal;connecting said first modulated signal or said second modulated signal to said first transmitter or to a second transmitter for transmission of said first modulated signal or said second modulated signal; andreceiving a fingerprint generated signal and providing from said fingerprint generated signal a fingerprint authentication signal, said fingerprint authentication signal is for authentication of authorized use of said mobile device. 2. The method of claim 1, further comprising a method for processing a video signal into a spread spectrum processed signal, said spread spectrum signal comprises cross-correlated in-phase and quadrature phase Time Constrained (TCS) waveform and cascaded Long Response (LR) filtered baseband signals, modulating said cross-correlated signals into a cross-correlated modulated signal and transmitting said cross-correlated modulated signal. 3. The method of claim 1, further comprising a method for processing a touch screen generated signal into a processed touch screen control signal, said touch screen generated control signal is for control of said mobile device and further comprising step of receiving, in said mobile device, a voice signal and processing said voice signal into transmit baseband filter filtered first GMSK cross-correlated in-phase and quadrature phase filtered baseband first GMSK modulated signal and receiving, demodulating and filtering a second GMSK modulated voice signal into demodulated and receive filter filtered second GMSK signal, said receive filter is mis-matched to said transmit filter. 4. The method of claim 1, further comprising step of receiving, in said mo... | - | H04L027/00 H04L027/26 G06F019/00 H04W064/00 H04W088/06 A61B005/00 H04B001/709 H04L027/20 H04W004/02 H04W004/22 H04B007/26 H04L001/00 H04L005/14 H04B001/707 H04L005/00 H04W012/06 H04B001/10 H04B001/69 H04B007/04 H04M007/00 A61B005/01 A61B005/024 A61B005/145 H04J011/00 H04W088/02
US09775563 | B2 | 2017-10-03 | Integrated introducer and transmitter assembly and methods of use | 2016-09-21 | Method and apparatus for inserting at least a portion of a sensor into a patient is provided. | 1. A method of using a sensor introducer assembly, the sensor introducer assembly comprising an insertion unit operatively coupled to an analyte sensor, wherein the insertion unit includes a compressible introducer mechanism coupled to a sensor introducer, the method comprising: placing a base portion of the sensor introducer assembly on a skin surface of a patient;applying a manual force on a surface of the compressible introducer mechanism to cause the compressible introducer mechanism to move from a non-compressed state to a compressed state;inserting, by the compressible introducer mechanism, at least a portion of the analyte sensor and the sensor introducer through a skin layer to an insertion position under the skin layer; andautomatically withdrawing the at least the portion of the sensor introducer from under the skin layer while retaining the at least the portion of the analyte sensor at the insertion position, wherein the insertion unit and compressible introducer mechanism are contained within a housing when the compressible introducer mechanism is in the non-compressed state and the compressed state, and wherein the sensor introducer includes a sharp portion configured to releasably couple with the at least the portion of the analyte sensor. 2. The method of claim 1, further comprising driving the sharp portion, by the manual force, to pierce through the skin layer and to position the at least the portion of the analyte sensor in fluid contact with a biological fluid under the skin layer at the insertion position. 3. The method of claim 2, wherein the biological fluid includes one of interstitial fluid or blood. 4. The method of claim 1, wherein the analyte sensor is a glucose sensor. 5. The method of claim 1, wherein the analyte sensor is operatively coupled to sensor electronics. 6. The method of claim 5, wherein the sensor electronics include an opening portion, the opening portion configured to selectively provide access to the sensor introducer. 7. The method of claim 5, further comprising communicating, by the sensor electronics, one or more signals corresponding to an analyte level monitored by the analyte sensor. 8. The method of claim 7, wherein the sensor electronics further comprise a wireless communication unit configured to wirelessly transmit signals according to at least one of a radio frequency (RF), a Bluetooth protocol, an infrared protocol, an 802.11x communication protocol, and a Zigbee communication protocol. 9. The method of claim 1, further comprising removing the manual force from the compressible introducer mechanism to cause the compressible introducer mechanism to move from the compressed state to the non-compressed state. 10. The method of claim 1, wherein the compressible introducer mechanism includes a substantially dome-shaped outer surface. 11. A method, comprising: applying a manual force on a surface of a compressible introducer mechanism coupled to a sensor introducer to cause the compressi... | ABBOTT DIABETES CARE INC. | A61B005/145 A61B005/00
US09782128 | B2 | 2017-10-10 | Wearable device and method for controlling the same | 2016-01-08 | Provided are a wearable device and a method for controlling the same. The wearable device may include a distance sensor unit configured to output a first signal for determining a distance between a user and the wearable device and to receive the first signal reflected from the user. The wearable device may also include a biometric sensor unit configured to output a second signal for determining biometric information of the user and to receive the second signal reflected from the user, and a controller configured to determine the distance based on the received first signal and to obtain the biometric information of the user based on a change in an attribute of the received second signal. The controller may control the biometric sensor unit to output the second signal if the determined distance is equal to or less than a reference distance. | 1. A wearable device comprising: a distance sensor unit configured to output a first signal for determining a distance between a user and the wearable device and to receive the first signal reflected from the user, when the wearable device is worn by the user;a biometric sensor unit configured to output a second signal for determining biometric information of the user for at least a portion of a position where the wearable device is worn by the user and to receive the second signal reflected from the user; anda controller configured to determine the distance based on the reflected first signal and to obtain the biometric information of the user based on a change in an attribute of the reflected second signal,wherein the controller controls the biometric sensor unit to output the second signal to the at least a portion of the position until a skin temperature of the user is increased to a reference temperature if the determined distance is equal to or less than a reference distance. 2. The wearable device of claim 1, wherein the distance sensor unit comprises a time of flight (ToF) light source configured to output the first signal and a light-receiving unit configured to receive the reflected first signal. 3. The wearable device of claim 2, wherein the controller is configured to control the ToF light source to output a wearing sensing signal for determining whether the wearable device is worn by the user at preset time intervals. 4. The wearable device of claim 1, wherein the biometric sensor unit comprises a pulse measuring light source configured to output the second signal and a light-receiving unit configured to receive the reflected second signal. 5. The wearable device of claim 1, wherein the reference distance comprises the determined distance between the user and the wearable device at an initial point in time when the wearable device is worn by the user. 6. The wearable device of claim 1, wherein the biometric information comprises at least one of pulse rate, oxygen saturation, and amount of consumed calories, and wherein the controller is configured to control storing of the biometric information obtained during a preset time in a storage unit. 7. The wearable device of claim 6, further comprising a communication unit configured to transmit the biometric information stored in the storage unit to an electronic device connected with the wearable device through wired or wireless communication. 8. The wearable device of claim 1, wherein the controller is configured to display additional information regarding changes to specific biometric information, if the specific biometric information is above a high threshold when a normal reading of the specific biometric information is below the high threshold, or if the specific biometric information is below a low threshold when a normal reading of the specific biometric information is below the low threshold. 9. The wearable device of claim 1, further comprising a temperature sensor unit c... | Samsung Electronics Co., Ltd. | A61B005/00 A61B005/01 A61B005/024 A61B005/0205
US09826910 | B2 | 2017-11-28 | Sensor mounting assembly for sensored guidewire and associated devices, systems, and methods | 2014-11-24 | Intravascular devices, systems, and methods are disclosed. In some embodiments, a guidewire system for treating a patient includes an internal sensor mount having a body and a first cutout extending through a wall of the body; a core wire secured to the internal sensor mount; a sensor assembly for detecting a physiological characteristic of a patient secured to the internal sensor mount; an external housing disposed about the internal sensor mount and including an opening in communication with the cutout of the internal sensor mount; a proximal flexible member secured to at least one of the internal sensor mount and the external housing; a first distal flexible member secured to at least one of the internal sensor mount and the external housing; and at least one conductor electrically coupled to the sensor assembly and extending proximally beyond the internal sensor mount through the proximal flexible member. | 1. A guidewire system for treating a patient, comprising: an internal sensor mount housing having a body with a first cutout extending through a wall of the body;a core wire secured to the internal sensor mount housing;a sensor assembly for detecting a physiological characteristic of a patient secured to and disposed within the internal sensor mount housing and mounted on a first rail defined by the first cutout of the internal sensor mount housing;an external housing disposed about the internal sensor mount housing and including a first opening in communication with the first cutout of the internal sensor mount housing;a proximal flexible member secured to at least one of the internal sensor mount housing and the external housing;a first distal flexible member secured to at least one of the internal sensor mount housing and the external housing; andat least one conductor electrically coupled to the sensor assembly and extending proximally beyond the internal sensor mount housing through the proximal flexible member. 2. The guidewire system of claim 1, wherein the internal sensor mount housing includes a second cutout extending through the wall of the body. 3. The guidewire system of claim 2, wherein the internal sensor mount housing includes a support member disposed between the first and second cutouts. 4. The guidewire system of claim 2, wherein the sensor assembly is further mounted on a second rail defined by the second cutout of the internal sensor mount housing. 5. The guidewire system of claim 4, wherein an interior of the internal sensor mount housing includes a first filling layer including solder, an adhesive, or a combination thereof, and a second filling layer including solder, an adhesive, or a combination thereof, the second filling layer being disposed over the first filling layer and defining a mounting surface such that the mounting surface is disposed higher than the first and second rails, and wherein the sensor assembly is mounted on the mounting surface. 6. The guidewire system of claim 2, wherein the external housing further includes a second opening in communication with the second cutout of the internal sensor mount housing. 7. The guidewire system of claim 6, wherein the first and second openings are laterally positioned relative to the sensor assembly. 8. The guidewire system of claim 1, wherein an interior of the internal sensor mount housing includes a mounting surface defined by solder, an adhesive, or a combination thereof, the sensor assembly being mounted on the mounting surface. 9. The guidewire system of claim 1, wherein the external housing includes a plurality of openings configured to provide fluid communication between the sensor assembly and an environment outside the external housing. 10. The guidewire system of claim 1, wherein the internal sensor mount housing is formed of stainless steel. 11. The guidewire system of claim 1, further comprising a second distal flexible member secured to at leas... | VOLCANO CORPORATION | A61B005/00 A61M025/00 A61B005/0215 A61M025/09
US09781724 | B2 | 2017-10-03 | Method to use auxiliary channel to achieve fast and power-efficient association in wireless networks | 2012-02-03 | A medical system includes one or more MBAN devices that acquire and communicate patient data. The medical system further including one or more medical body area network (MBAN) systems, each MBAN system including the one or more MBAN devices that communicate the patient data with a hub device via short-range wireless communication, the communication of the patient data via the short-range wireless communication being within a predefined spectrum. The hub device receives patient data communicated from the one or more MBAN devices and communicates with a central monitoring station via a longer range communication. The one or more MBAN devices transmit an association request to the hub device on one or more auxiliary channels to associate each MBAN device with the MBAN system, the auxiliary channels being outside the predefined spectrum. | 1. A medical system comprising: one or more medical body area network (MBAN) systems, each MBAN system including: one or more MBAN devices configured to acquire and communicate patient data with a hub device via short-range wireless communication, the communication of the patient data via the short-range wireless communication being within a predefined MBAN spectrum, the MBAN system being a secondary user in the predefined MBAN spectrum;the hub device being configured to receive patient data communicated from the one or more MBAN devices via the short-range wireless communication and communicate with a central monitoring station via a longer range communicationwherein the one or more MBAN devices are configured to transmit an association request to the hub device via the short-range wireless communication on one or more auxiliary channels to associate each MBAN device with the hub device, the auxiliary channels being outside the predefined MBAN spectrum. 2. The medical system according to claim 1, further including: a MBAN coordinator configured to generate MBAN spectrum authorizations for the one or more hub devices to authorize the transmission of patient data in the MBAN spectrum, each hub device authorized for the transmission of patient data in the MBAN spectrum being configured to respond to the association request from a new MBAN device by issuing the new MBAN device a key which enables the new MBAN device to communicate patient data in the MBAN spectrum. 3. The medical system according to claim 1, wherein the one or more MBAN devices actively or passively scan the auxiliary channels to receive information about the one or more MBAN systems, the information including the resources available of the one or more MBAN systems. 4. The medical system according to claim 1, wherein in response to receiving an association request, the hub device is configured to check its available resources to determine whether to accept the association request and if resources are available, issues a key to the MBAN device requesting association to enable the MBAN device to communicate in the MBAN spectrum. 5. The medical system according to claim 1, wherein the hub device transmits an association response and MBAN spectrum authorization to the one or more MBAN devices on the auxiliary channel in response to accepting the association request. 6. The medical system according to claim 1, wherein the MBAN spectrum are within the 2360 MHz-2390 MHz range and the auxiliary channels are outside the MBAN spectrum. 7. The medical system according to claim 1, wherein each MBAN device has a single transceiver configured to transmit and receive the short-range wireless communications in the MBAN spectrum and on the auxiliary channels and each hub device has a single transceiver configured to transmit and receive the short-range wireless communications in the MBAN spectrum and on the auxiliary channels. 8. The medical system according to claim 7, wherein an MBAN spectr... | KONINKLIJKE PHILIPS N.V. | H04W072/04 H04W048/06 H04W084/18 A61B005/00
US09788797 | B2 | 2017-10-17 | Exercise-assisting device for forecasting the change in a physiological parameter as a function of an itinerary | 2011-11-17 | A device (1) for assisting with physical exertion management, includes a physiological sensor, a positioning device, at least one memory in which data representative of the itinerary to be travelled during the physical exertion can be recorded, and at least one data processing unit organised so as to produce forecast data representative of the change in an physical exertion parameter over the remaining itinerary to be travelled by the individual, to compare these forecast data with predetermined data, and to produce and transmit a message, the content of which depends on the result of the comparison of the forecast data and the predetermined data, with a view to communicating the message to a user. | 1. Portable device (1) for assisting the management of a physical exertion provided by a subject, comprising: at least one sensor (10, 11, 12) for a at least one physiological parameter, called exertion parameter, evolving as a function of the physical exertion, said sensor being suitable to provide digital data, called exertion data, that are representative of said exertion parameter,at least one device, called positioning device (20), that is suitable to determine its own position in a given topographical frame of reference and to provide digital data, called positioning data, that are representative of this position,at least one interface (98, 99) for communication with a user,at least one memory (4) in which data, called input data, are capable of being recorded, comprising at least: positioning data provided by said positioning device,exertion data,data that are representative of an itinerary to be travelled in said topographical frame of reference by said subject in the course of the physical exertion,at least one data-processing unit (30), designed: to use said input data at any present point of the itinerary, said input data comprising at least topographic data of a part of the itinerary remaining to be travelled by the subject from the present point, to generate data, called forecast data, that are representative of the evolution of at least one of said exertion parameter over all of the itinerary, from equations and from tables of values recorded in at least one memory of the device,to compare these forecast data with predetermined values recorded in at least one memory of said assistance device,to generate data that are representative of a message, the content of which is a function of the result of the comparison of the forecast data and the predetermined values,to transmit said data that are representative of a message to the interface, with a view to the communication of said message to a user. 2. Device as claimed in claim 1, wherein a first exertion parameter is the heart rate and at least one second exertion parameter is chosen from: the oxygen saturation of the blood, the body temperature, the level of glucose in the blood, the blood pressure, the respiratory rate, the conductivity of the skin. 3. Device as claimed in claim 1, wherein said data-processing unit (30) is designed to generate said forecast data that are representative of the evolution of the exertion parameters as a function of the input data. 4. Device as claimed in claim 1, wherein said data-processing unit (30) is designed to generate at least one forecast datum that is representative of an expenditure of energy as a function of the input data and of data that are representative of the evolution of at least one exertion parameter. 5. Device as claimed in claim 1, wherein said data-processing unit (30) is designed to compare each exertion parameter to at least one predetermined value, called threshold value, that is specific to said exertion parameter. 6. Dev... | CENTRE NATIONAL D'ETUDES SPATIALES (CNES) | A61B005/00 A63B071/00 A61B005/024 A61B005/11 A61B005/01 A61B005/0205 A61B005/145 A63B024/00
US09814449 | B2 | 2017-11-14 | Motorized optical imaging of prostate cancer | 2016-07-06 | A system, method, and storage device storing computer executable instructions for use in tissue analysis and therapy. A motorized optical probe for illuminating tissue generates light fluorescence and/or diffuse reflectance signals corresponding to the illuminated tissue. One or more ultrasound, CT and MRI imaging guidance systems identify the position of the optical probes relative to the tissue. An imaging system generates a three-dimensional image of the tissue based on the generated light signals and the identified position of the optical probe. | 1. A system for evaluating a segment of tissue in a body comprising: a needle having at least one fiber optic cable adapted to transmit light from a light source to illuminate a first location along the segment of tissue and to transmit spectra emanating from the first location;a system console comprising a fluorometer configured to evaluate the spectra and generate signals of selected data of the spectra, anda module executing a classification algorithm wherein the classification algorithm classifies tissue at the first location based only on the generated signals to generate a classification, wherein the classification indicates tissue condition and is provided real-time to a display by correlating the first location to a three-dimensional image of the tissue and generating a three-dimensional optical spectroscopic image showing the first location and classification in relation to the three-dimensional image while the needle is positioned to illuminate the first location; andan autoneedle advancement mechanism configured to automatically advance the needle in a preselected increment after the classification of tissue at the first location has been provided to the display, wherein the preselected increment advances the needle to allow illumination at a next location along the segment of tissue. 2. A system as in claim 1, wherein the autoneedle advancement mechanism automatically advances the needle according to a preset number of preselected increments, wherein each preselected increment has a corresponding location along the segment of tissue and classification of tissue of the corresponding location. 3. A system as in claim 2, wherein the autoneedle advancement mechanism automatically stops advancing the needle after the preset number of increments. 4. A system as in claim 2, wherein the needle includes a collection space for biopsy and the preset number of preselected increments equals a length of the collection space. 5. A system as in claim 2, wherein the preset number of preselected increments equals 18. 6. A system as in claim 1, wherein the preselected increment is 1 mm. 7. A system as in claim 1, wherein the autoneedle advancement mechanism comprises a stepper motor. 8. A system as in claim 1, wherein the needle includes a collection space for biopsy and wherein the needle includes an outer cannula which is advanceable over the collection space to shear and collect the tissue creating a biopsy. 9. A system as in claim 8, wherein advancement of the outer cannula over the collection space while the needle is positioned to illuminate the first location biopsies tissue at the first location and not tissue at the next location. 10. A system as in claim 8, wherein the outer cannula is advanceable of over the collection space by at least one of a motor and a spring. 11. A system as in claim 8, wherein advancement of the outer cannula of over the collection space stops automatic advancement of the needle. 12. A system as in claim 1, ... | Precision Biopsy, Inc. | A61B010/04 A61B001/07 A61B008/12 A61B008/00 A61B001/06 A61B005/00 A61B001/00 A61B008/08 A61B005/06 A61B001/04 A61B006/12 A61B010/02 A61N005/06 A61N005/10 A61B018/04 A61B034/20 A61B017/34 A61B001/307 A61B006/03 A61N007/02 A61B018/14 A61B018/22 A61B018/02 A61B090/00 A61B090/30
USRE046615 | E1 | 2017-11-28 | Accommodating intra-ocular lens system | 2015-11-03 | An implantable, compressible, accommodating intra-ocular lens (IOL) coupled to at least one sensor which detects a signal created by the ciliary muscle. A ciliary sulcus ring can house the at least one sensor, and the sensor can include miniaturized electrodes (ciliary muscle probes) for implanting into the ciliary muscle of the subject. A potentiometer/microcomputer can modulate the ciliary muscle signal detected by the sensor(s) into an electrical signal, and a transmitter sends this electrical signal to a micromotor, which causes compression of the IOL via an annular support ring system, causing a change in the IOL shape. The IOL can be part of an IOL complex including a compressible, accommodating IOL, an external lens membrane, and an annular support ring system. The annular support ring system provides a foundation for the micromotor to compress the IOL. | 1. A system for causing a change in the shape of a compressible, accommodating intra-ocular lens (IOL) of a subject, the system comprising: a) an IOL complex comprising a compressible, accommodating IOL;b) at least one sensor for detecting a signal from the subject's ciliary muscle;c) a signal processor for receiving and processing the ciliary muscle signal to generate a corresponding electrical signal;d) a transmitter for transmitting the electrical signal; ande) at least one micromotor for receiving the transmitted electrical signal and causing a change in shape of the IOL, the IOL complex further comprising an external lens membrane and an annular support ring system for providing a means to cause a change in shape of the IOL, wherein the annular support ring system comprises two concentric rings including an outer annular ring and an inner annular ring, the outer annular ring being rigid for resisting compressive forces of the native capsular bag, the inner annular ring encircling the external lens membrane and providing a foundation for the at least one micromotor to cause a change in shape of the IOL; andf) a ciliary sulcus ring adapted for implantation in the ciliary sulcus of the subject and adapted to be positioned around the IOL complex. 2. The system of claim 1, wherein the at least one micromotor causes a change in shape of the IOL by causing the inner annular ring to shorten in circumference and diameter about the IOL, thereby compressing the IOL. 3. The system of claim 1, further including a ciliary sulcus ring adapted for implantation into the ciliary sulcus of the subject, wherein the at least one sensor is an electromyography (EMG) receiver for detecting the muscle action potential of the ciliary muscle, the EMG receiver being positioned on the ciliary sulcus ring and having miniaturized electrodes for implanting into the ciliary muscle. 4. The system of claim 31, wherein the at least one sensor is a series of EMG receivers positioned about the ciliary sulcus ring, and wherein the signal processor and the transmitter are also positioned about the ciliary sulcus ring. 5. The system of claim 1, wherein the ciliary muscle signal comprisesis in a form comprising a signal selected from the group consisting of a neuro-electrical pulsation, a muscle action potential, a muscle contractile force or movement, and a chemical signal, or a combination of these signal forms. 6. The system of claim 1, wherein the signal processor is selected from the group consisting of a micropotentiometer and a microcomputer. 7. AAn implantable system for causing a change in the shape of a compressible, accommodating intra-ocular lens (IOL),correction of presbyopia, the system, implantable at an eye of a subject, comprising: a) an IOL complex comprising a compressible, accommodating IOL, an external lens membrane, and an annular support ring system for providing a means to cause a change in shape of the IOLa synthetic intra-ocular lens (IOL) sized and c... | - | A61F002/16 A61B005/0488 A61B005/00 A61B005/0492
USRE046614 | E1 | 2017-11-28 | Ultrasonic methods for diagnosis and treatment of stroke | 2011-08-08 | A medical diagnostic method, system and related equipment particularly adapted to diagnose disorders of the blood circulation serving the head and neck, and especially the brain. A preferred use of the system is early, rapid, accurate, diagnosis of stroke, especially whether the stroke is due to blockage of a blood vessel or leakage from the blood vessel. | 1. Using non-ultrasonic detection of symptoms of stroke to dynamically couple an ultrasonic transducer/receiver to a skull, in which the non-ultrasonic detection comprises at least one of the group consisting of computed tomography scanning, magnetic resonance scanning, differential spectrophotometric methods, near-infrared detection of tissue characteristics, detection of a biological material, measurement of a biological material, detection of a chemical, measurement of a chemical, detection of S100β, measurement of S100β, use of biological assay techniques, detection of change in blood pressure, detection of change in pressure within the eye, detection of change in blood flow in arteries serving organs other than the brain, or detection of change in blood flow in the arteries serving the eye. 2. For a patient having a skull containing brain tissue and blood vessels, a process for diagnosing stroke in the patient, comprising: a) dynamically coupling at least one ultrasonic transmitter/receiver to at least one man-made opening in the skull of the patient;b) imaging a region within the skull of the patient; andc) at least one of: determining presence of hemorrhagic stroke by identifying relatively diffuse blood flow within the skull of the patient; determining presence of ischemic stroke by identifying at least one location of inadequate blood flow within the skull of the patient; and diagnosing between hemorrhagic and ischemic stroke by classifying a region within the skull of the patient in terms of adequacy of blood flow. 3. The process of claim 2, in which the imaging comprises generating and receiving ultrasonic signals suitable for processing into information about the region within the skull of the patient. 4. The process of claim 2, in which a region is classified in terms of normal blood flow. 5. The process of claim 2, in which a region is classified in terms of inadequate blood flow. 6. The process of claim 2, in which a region is classified in terms of relatively diffuse blood flow. 7. The process of claim 2, in which locations of hemorrhagic stroke are determined by identifying relatively diffuse blood flow outside the blood vessels of the brain. 8. The process of claim 2 in which locations of ischemic stroke are determined by identifying relatively inadequate blood flow within the blood vessels of the brain. 9. The process of claim 2 in which presence of ischemic stroke is determined by identifying at least one location where there is at least partial blockage of blood flow. 10. The process of claim 2, further comprising applying a vacuum to the skull. 11. The process of claim 2, further comprising applying an acoustic coupling material to the skull. 12. A stroke diagnosis/treatment method comprising: mounting a unit having two transmitter/receivers on the head of a subject with each transmitter/receiver acoustically coupled to an acoustic window of the skull, wherein each transmitter/receiver includes a plurality of ultr... | KONINKLIJKE PHILIPS N.V. | A61B008/02 A61B008/06 A61B005/026 A61B008/08 A61B008/00 A61B005/00 A61B008/14 A61N007/00 G01S015/89 A61B005/024 G01S007/52 A61B006/00
US09801587 | B2 | 2017-10-31 | Heart rate monitor with time varying linear filtering | 2016-10-18 | A system, method, and device for monitoring a cardiac signal of a user includes an emitter (LED) for transmitting light toward skin of the user, a receiver (photodiode) for receiving a reflection of the transmitted light and generating a light intensity signal, and a processor configured to generate a photoplethysmogram (PPG) signal based on the light intensity signal. The PPG signal includes a cardiac component, a motion component and a respiratory component. An inertial sensor provides a motion signal to the processor based on sensed movement of the device and one or more time-variant filters are configured by the processor to filter the PPG signal to isolate the cardiac component of the PPG signal based on determined filter coefficients. | 1. A fitness monitor for monitoring a cardiac signal of a user, the device comprising: an emitter configured to transmit light toward skin of the user;a receiver configured to receive a reflection of the transmitted light and provide a photoplethysmogram (PPG) signal based on the received reflection of the transmitted light, the PPG signal including a cardiac component and a motion component;an inertial sensor configured to provide a motion signal based on sensed movement of the fitness monitor;a time-variant filter configured to isolate a component of the PPG signal based on filter coefficients corresponding to a desired component of the PPG signal, anda processor coupled to the emitter, the receiver, the inertial sensor and the time-variant filter, the processor configured to: determine one or more frequencies associated with the motion signal;transform the PPG signal to the frequency domain;identify, in the frequency domain, the cardiac component in the PPG signal based on the determined one or more frequencies of the motion signal and the transformed PPG signal;determine filter coefficients based on the identified cardiac component in the PPG signal;configure the time-variant filter with the determined filter coefficients;filter, in the time domain via the time-variant filter, the motion component from the PPG signal to isolate the identified cardiac component; andcalculate, in the time domain, the time between successive cycles of the isolated cardiac component of the PPG signal. 2. The fitness monitor of claim 1, wherein the time-variant filter is integrated within the processor. 3. The fitness monitor of claim 1, wherein the filter coefficients are determined upon each occurrence of the cardiac component being identified based on the determined one or more frequencies of the motion signal and the transformed PPG signal. 4. The fitness monitor of claim 3, wherein the processor configures the time-variant filter with the determined filter coefficients upon each occurrence of the cardiac component being identified. 5. The fitness monitor of claim 1, wherein the processor is further configured to periodically reconfigure the time-variant filter based on the cardiac component identified in the frequency domain. 6. The fitness monitor of claim 1, further comprising a bandpass filter, wherein the processor is further configured to configure the bandpass filter to pre-condition the PPG signal by attenuating frequency components of the PPG signal below a predetermined frequency band that includes the cardiac component and/or attenuates frequency components above the predetermined frequency band that includes the cardiac component, the predetermined frequency band spanning the full range of human heart rates. 7. The fitness monitor of claim 1, wherein the cardiac component is associated with heart rate, and wherein the determined time between successive cycles of the cardiac component is the time between successive heartbeats. 8. The fitness ... | Garmin Switzerland GmbH | A61B005/00 A61B005/083 A61B005/16 A61B005/027 A61B005/024
US09779912 | B2 | 2017-10-03 | Inspection device and measurement device | 2014-12-08 | A detection circuit for accurately detecting a very small foreign material and an inspection/measurement device using the same are provided. The inspection/measurement device includes: an irradiation section that irradiates a laser beam to a surface of a specimen; and a detection section that detects scattered light from the surface of the specimen and generates a detection signal. The detection section includes: a photon counting sensor that outputs M output signals from photo-detecting elements of N pixels (M and N are natural numbers, and M | 1. An electron beam device comprising: an irradiation system for supplying an electron beam to a specimen; anda detection system for detecting electrons from the specimen,wherein the detection system includes: a photon counting sensor that outputs M output signals from photo-detecting elements of N pixels, wherein M and N are natural numbers, and M&lt;N;M current-voltage conversion units that execute current-voltage conversion on the output signals of the photon counting sensor respectively;a voltage application unit that applies reference voltages to the current-voltage conversion units; anda detection signal generation unit that generates a detection signal on the basis of the outputs of the current-voltage conversion sections;wherein the reference voltages are input to the positive inputs of the power source voltage conversion units, and the output signals of the photon counting sensor are input to the negative inputs of the power source voltage conversion units, andwherein an adding circuit that inverse-point-adds the output voltages from the voltage application unit is included, and the summation of the output from the adding circuit and the outputs from the current-voltage conversion units are input to the detection signal generation unit. 2. A medical device comprising: a photon counting sensor that outputs M output signals from photo-detecting elements of N pixels, wherein M and N are natural numbers, and M&lt;N;M current-voltage conversion units that execute current-voltage conversion on the output signals of the photon counting sensor respectively;a voltage application unit that applies reference voltages to the current-voltage conversion units; anda detection signal generation unit that generates a detection signal on the basis of the outputs of the current-voltage conversion units;wherein the reference voltages are input to the positive inputs of the power source voltage conversion units, and the output signals of the photon counting sensor are input to the negative inputs of the power source voltage conversion units, andwherein an adding circuit that inverse-point-adds the output voltages from the voltage application unit is included, and the summation of the output from the adding circuit and the outputs from the current-voltage conversion units are input to the detection signal generation unit. 3. The device according to claim 1, wherein the voltage application section applies the same reference voltages to the M current-voltage conversion sections. 4. The device according to claim 2, wherein the detection section includes an addition section that adds the outputs of the voltage application section; andthe detection signal generation section generates a detection signal on the basis of the outputs of the current-voltage conversion sections and the output of the addition section. 5. The device according to claim 1, wherein the voltage application section outputs the total value of voltages applied to the current-voltage conversi... | Hitachi High-Technologies Corporation | H01J037/00 H01J037/244 G01N021/956 G01N021/94 G01N021/95 H01J037/28 A61B005/00 H01J049/00
US09782497 | B2 | 2017-10-10 | Synthesis and composition of amino acid linking groups conjugated to compounds used for the targeted imaging of tumors | 2015-11-30 | The present disclosure relates to compounds that are useful as near-infrared fluorescence probes, wherein the compounds include i) a pteroyl ligand that binds to a target receptor protein, ii) a dye molecule, and iii) a linker molecule that comprises an amino acid or derivative thereof. The disclosure further describes methods and compositions for making and using the compounds, methods incorporating the compounds, and kits incorporating the compounds. | 1. A method of optical imaging of a biological tissue that expresses a folate receptor in an infectious disease, the method comprising: (a) contacting the biological tissue with a composition comprising a compound having the formula:or a pharmaceutically acceptable salt thereof,or isotopes thereof, wherein:X is an single amino acid or a single amino acid derivative thereof, wherein the single amino acid or single amino acid derivative contains an —OH, —NH2, or —SH functional group, andY is a dye that has a fluorescence excitation and emission spectra in the near infra red range, wherein Y is represented by the formula:wherein, R1 is independently selected from the group consisting of O, S, N and C, and R2 is independently selected from the group consisting of CH2 and CH2CH2, and the compound maintains or enhances the fluorescence of the dye, Y;(b) allowing time for the compound in the composition to distribute within the biological target;(c) illuminating the tissue with an excitation light of a wavelength absorbable by the compound; and(d) detecting the optical signal emitted by the compound. 2. The method of claim 1 wherein the amino acid is selected from the group consisting of tyrosine, cysteine, lysine, serine or a derivative thereof. 3. The method of claim 1 wherein the amino acid is tyrosine, wherein the aromatic ring of tyrosine comprises a carbon isotope and/or a hydrogen isotope. 4. The method of claim 1 wherein the amino acid derivative is a derivative of tyrosine selected from the group consisting of:or racemic mixtures thereof. 5. The method of claim 1 wherein the amino acid is cysteine or a derivative of cysteine. 6. The method of claim 1, wherein the compound has the formula:or is a potassium, sodium, or ammonium salt thereof. 7. The method of claim 1, wherein the compound has an absorption and emission maxima between about 600 nm and 800 nm. 8. The method of claim 1, wherein the compound is made to fluoresce after distribution thereof in infectious foci. 9. The method of claim 1 wherein the compound is highly selective for targeting microenvironment in infectious foci. 10. The method of claim 1 wherein the compound is highly selective for targeting activated macrophages in a microenvironment in infectious foci. 11. A method of identifying a target cell type in an infectious disease comprising a) contacting the biological sample with a compound of claim 1 for a time and under conditions that allow for binding of the compound to at least one cell of the target cell type; andb) optically detecting the presence or absence of the compound of in the biological sample, wherein presence of the compound in detecting step b) indicates that the target cell type is present in the biological sample. 12. A method of diagnosing an infectious disease in a subject comprising: c) administering to a subject in need of diagnosis an amount of a compound of claim 1 for a time and under conditions that allow for binding of the compound to at ... | Purdue Research Foundation | A61B005/00 A61B008/00 A61B010/00 A61K049/00 G01N033/569 G01N033/566 G01N033/50 G01N033/58 G01N033/68 C07D475/04 G01N033/574 G01N021/64
US09795331 | B2 | 2017-10-24 | Method and apparatus for providing analyte sensor insertion | 2016-04-28 | Devices and methods for positioning a portion of a sensor at a first predetermined location, displacing the portion of the sensor from the first predetermined location to a second predetermined location, and detecting one or signals associated with an analyte level of a patient at the second predetermined location are disclosed. Also provided are systems and kits for use in analyte monitoring. | 1. A method, comprising: positioning a portion of an analyte sensor at a first predetermined location at a first predetermined speed, the analyte sensor positioned in fluid contact with a bodily fluid under a skin surface;displacing the analyte sensor at a second predetermined speed slower than the first predetermined speed from the first predetermined location to a second predetermined location;retrieving, using one or more processors, a factory determined calibration value stored in a memory storage component, the factory determined calibration value being not modified during a lifetime of the analyte sensor;detecting one or more signals associated with an analyte level in the bodily fluid under the skin surface; andgenerating, using the one or more processors, clinically accurate calibrated analyte sensor data by processing the detected one or more signals with the factory determined calibration value, wherein the detected one or more signals are processed substantially immediately after the positioning the analyte sensor in fluid contact with the bodily fluid, and further, wherein the clinically accurate calibrated analyte sensor data correlates to the analyte level in blood. 2. The method of claim 1, wherein the analyte sensor comprises a plurality of electrodes including a working electrode comprising an analyte-responsive enzyme bonded to a polymer disposed on the working electrode. 3. The method of claim 2, wherein the analyte-responsive enzyme is chemically bonded to the polymer disposed on the working electrode. 4. The method of claim 2, wherein the working electrode comprises a mediator bonded to the polymer disposed on the working electrode. 5. The method of claim 1, wherein the detected one or more signals correspond to monitored analyte levels in at least one of (i) dermal fluid or (ii) interstitial fluid. 6. The method of claim 1, wherein the clinically accurate calibrated analyte sensor data is generated by processing the detected one or more signals only with the factory determined calibration value during an entire lifetime of the analyte sensor spanning from positioning the analyte sensor in contact with the bodily fluid to removal of the analyte sensor from contact with the bodily fluid. 7. The method of claim 1, wherein the memory storage component is operatively coupled to the one or more processors, and wherein the one or more processors are processors of sensor electronics physically coupled to the analyte sensor. 8. The method of claim 1, further including wirelessly communicating the generated clinically accurate calibrated analyte sensor data to a remote location for display. 9. The method of claim 1, wherein the detected one or more signals are processed within 1 to 3 hours after positioning the analyte sensor in fluid contact with the bodily fluid. 10. An apparatus, comprising: an analyte sensor having a portion for positioning at a first predetermined location at a first predetermined speed and in fluid cont... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/1473 A61B005/145 A61B005/1486
US09798320 | B2 | 2017-10-24 | Method and apparatus for alarm monitoring | 2013-09-03 | A method for alarm monitoring abnormal conditions associated with a process operation includes collecting N sample data representing one or more parameters of a process, determining a distribution type of the sample data, obtaining an alarm monitoring strategy associated with the distribution type, and monitoring the process using the alarm monitoring strategy. The distribution type may be single-constant, multi-level discrete, normally distributed, continuous non-normally distributed, cyclical trend-up/down, and drifting away after a period maintenance. The alarm monitoring strategy includes a unilateral or bilateral control chart. The one-side control chart has either an upper control limit (UCL) of (100−p) or a lower control limit (LCL) of (p), and the two-sided control chart has a UCL of (100−p/2) and an LCL of (p/2), where p is a predetermined false alarm rate which determines the number N. The control limit may be a difference between two adjacent sample values of the sample data. | 1. A computer-implemented method for alarm monitoring of a process operation by an alarm monitoring apparatus comprising a sampling unit, a determining unit, an acquisition unit and a monitoring unit, the method comprising: collecting sample data representing one or more parameters of a process by the sampling unit;determining a distribution type of the sample data by the determining unit;obtaining an alarm monitoring strategy corresponding to the distribution type by the monitoring unit; andmonitoring the process using the alarm monitoring strategy by the monitoring unit,wherein determining the distribution type of sample data comprises:determining whether a value level of the sample data is greater than a first threshold value;if the value level is not greater than the first threshold value, determining whether the distribution type is a single constant-type or multi-level discrete type; otherwisedetermining whether the sample data is normally distributed:if the sample data is normally distributed, setting the upper control limit to a maximum value of the sample data and the lower control limit to a minimum value of the sample data, andif the sample data is not normally distributed, determining whether the distribution type is a continuous non-normal distribution type, cyclical trend-up/down-type, or a drift after a periodic maintenance (PM) type. 2. The computer-implemented method of claim 1, wherein obtaining the alarm monitoring strategy comprises: determining a control limit of a control chart associated with the distribution type; andmonitoring the production process using the control limit of the control chart. 3. The computer-implemented method of claim 1, wherein determining the distribution type further comprises: collecting at least a number N of the sample data, wherein N=2/p %, and p % is a predetermined alarm rate. 4. The computer-implemented method of claim 1, wherein determining whether the distribution type is a single constant-type or multi-level discrete comprises: if the value level is equal to 1 (unity), setting the control limit to a predetermined constant; otherwise,setting an upper limit of the control limit to the maximum value of the sample data;setting a lower control limit of the control limit to a minimum value of the sample data. 5. The computer-implemented method of claim 1, wherein determining the distribution type of the continuous non-normal distribution type, cyclical trend-up/down type, or drift after the PM type comprises: determining whether the distribution of the sample data satisfies PM changing conditions;if the sample data distribution does not satisfy the PM changing conditions, the distribution type is the continuous non-normal distribution type; otherwise,the distribution type is the cyclical trend-up/down type, or drift after the PM type. 6. The computer-implemented method of claim 5, wherein determining the continuous non-normal distribution type comprises: determining whether the control char... | Semiconductor Manufacturing International (Shanghai) Corporation | G05B023/00 G05B013/02 G05B023/02 A61B005/0205 F25J003/04 G06F017/18
US09816959 | B2 | 2017-11-14 | Sensor for monitoring of ethanol | 2015-12-29 | The subject invention provides devices, and methods of making and using the same, for the non-invasive detection of ethanol in a sample. In specific embodiments, the fuel cell based ethanol detector of the subject invention is capable of measuring the concentration of ethanol vapor in the presence of water vapor, which is known to confound signal readings in conventional detectors. Advantageously, the electrochemical sensors provided herein are highly stable and accurate, especially suitable for low-cost, continuous monitoring of ethanol content in transdermal perspiration samples. | 1. A fuel cell based sensing device for continuously detecting the concentration of ethanol vapor in a sample, comprising: a sensor unit comprising a working electrode as an anode, a counter electrode as a cathode, a reference electrode, and a polymer electrolyte membrane, the anode and the cathode being separated by, and respectively in electrical contact with, the polymer electrolyte membrane, the reference electrode being in electrical contact with the polymer electrolyte membrane and disposed adjacent to, though separated from, the counter electrode;circuitry capable of controlling the potential applied to the sensor unit; andcircuitry capable of measuring the current output of the sensor unit;wherein the sensor unit is capable of separating signals produced by ethanol from those produced by background humidity. 2. The device according to claim 1, wherein the polymer electrolyte membrane is a proton exchange membrane (PEM). 3. The device according to claim 1, wherein the working electrode and the counter electrode each includes an active surface area, the active surface area of the working electrode being smaller than the active surface area of the counter electrode. 4. The device according to claim 1, wherein the sample is a product of a human transdermal diffusion process. 5. The device according to claim 4, the sample being skin perspiration. 6. The device according to claim 1, wherein the electrodes comprise a material selected from iron, gold, platinum, carbon, and one or more of these materials combined. 7. The device according to claim 6, wherein the working electrode comprises stainless steel. 8. The device according to claim 6, wherein each of the counter and the reference electrodes comprises nickel. 9. The device according to claim 1, further comprising a thin coating of ethanol-permeable membrane in electrical contact with the working electrode. 10. A method of continuously detecting the concentration of ethanol vapor in a skin perspiration sample collected from a human subject, the method comprising: contacting a sample, the sample optionally comprising water vapor, with a fuel cell based sensing device comprising a sensor unit comprising a working electrode as an anode, a counter electrode as a cathode, a reference electrode, and a polymer electrolyte membrane, the anode and the cathode being separated by, and respectively in electrical contact with, the electrolyte, the reference electrode being in electrical contact with the polymer electrolyte membrane and disposed adjacent to, though separated from, the counter electrode; circuitry capable of controlling the potential applied to the sensor unit; and circuitry capable of measuring the current output of the sensor unit, wherein the sensor unit is capable of separating signals produced by ethanol from those produced by background humidity, and wherein the electrodes comprise a material selected from iron, gold, platinum, carbon, and one or more of these materials combi... | THE FLORIDA INTERNATIONAL UNIVERSITY BOARD OF TRUSTEES | G01N027/403 G01N027/407 G01N027/30 G01N027/333 G01N027/406 G01N027/416 G01N033/497 G01N033/487 G01N033/00 A61B005/00 A61B010/00
US09817092 | B2 | 2017-11-14 | Method and magnetic resonance apparatus with a cooling system to cool a superconducting basic magnetic field coil | 2014-05-09 | A magnetic resonance apparatus has a magnet unit that includes at least one superconducting basic magnetic field coil, a magnet housing unit surrounding the at least one superconducting basic magnetic field coil, a cooling system that has at least one cooling loop and a heat absorption unit to cool the at least one superconducting basic magnetic coil, and an additional unit. The cooling system has a switching unit with at least one first cooling mode, and the switching unit couples the at least one cooling loop of the cooling system with the additional unit for a heat exchange in the first cooling mode. | 1. A magnetic resonance (MR) apparatus comprising: a magnet unit configured to operate to acquire MR data, said magnet unit comprising a basic field magnet having at least one superconducting coil requiring cooling;said magnet unit comprising a plurality of additional units each configured to respectively perform an individual operation that participates in the acquisition of said MR data, each individual operation being unrelated to the cooling of said at least one superconducting coil, and at least one of said additional units having a structure that is capable of heat exchange;said magnet unit comprising a cooling system that provides said cooling to said at least one superconducting coil, said cooling system comprising a cooling loop in thermal communication with said at least one superconducting coil, and a heat absorption unit that discharges heat from said cooling loop;said cooling system comprising a switching unit configured to selectively place said cooling loop in thermal communication with either of said heat absorption unit or said at least one of said additional units that is capable of heat exchange; andsaid switching unit being configured to monitor an operating state of said heat absorption unit to detect a switching-triggering condition in said heat absorption unit and, as long as said switching-triggering condition is not detected, to place said cooling system in a first cooling mode in which said cooling loop is in thermal communication with said heat absorption unit and out of communication with said at least one additional unit and, when said switching-triggering condition is detected, to switch said cooling system to a second cooling mode in which said cooling loop is in thermal communication with both said heat absorption unit and said at least one additional unit, with the heat exchanging capability of said at least one additional unit in said second cooling mode augmenting the discharge of heat from said cooling loop by said heat absorption unit. 2. An MR apparatus as claimed in claim 1 wherein said switching unit is configured to detect an occurrence of a fault state of said heat absorption unit as said switching-triggering condition. 3. An MR apparatus as claimed in claim 2 wherein said cooling is a first cooling loop and wherein said heat absorption unit comprises a second cooling loop, and wherein said switching unit is configured to detect a fault state of said second cooling loop as said fault state of said heat absorption unit. 4. An MR apparatus as claimed in claim 3 wherein said switching unit is configured to thermally decouple said second cooling loop from said first cooling loop upon detection of said fault state of said second cooling loop. 5. An MR apparatus as claimed in claim 2 wherein said cooling system comprises a temperature sensor that detects a temperature within said heat absorption unit, and wherein said switching unit is configured to detect, as said switching-triggering condition, when said ... | Siemens Aktiengesellschaft | G01R033/38 H01F006/06 G01R033/3815 H01F006/04 F25D019/00 A61B005/055
US09814415 | B2 | 2017-11-14 | Method and system for non-invasive optical blood glucose detection utilizing spectral data analysis | 2012-09-11 | Systems and methods are disclosed for non-invasively measuring blood glucose levels in a biological sample based on spectral data. This includes at least one light source configured to strike a target area of a sample, at least one light detector, which includes a preamplifier having a feedback resistor, positioned to receive light from the at least one light source and to generate an output signal, having a time dependent current, which is indicative of the power of light detected, and a processor configured to receive the output signal from the at least one light detector and based on the received output signal, calculate the attenuance attributable to blood in a sample present in the target area and eliminate effect of uncertainty caused by temperature dependent detector response of the at least one light detector, and based on the calculated attenuance, determine a blood glucose level associated with a sample. | 1. A system for detecting glucose in a biological sample, comprising: at least one light beam generating structure generating one or more light beams having a wavelength in a range between 800 nm and 1600 nm to strike a target area of a biological sample;at least one photocurrent signal generating light detector, which includes a preamplifier having a feedback resistor, positioned to receive light from the at least one light source and to generate an output photocurrent signal, having a time dependent current, which is indicative of the power of light detected;wherein the preamplifier having a feedback resistor is configured to convert the time dependent current into a time dependent voltage signal; anda light absorbance change determining algorithm implemented processor programmed to calculate a change in a light absorption caused by blood in the biological sample and configured to receive the output photocurrent signal from the at least one photocurrent signal generating light detector and based on the received output photocurrent signal, calculate the attenuance attributable to blood in the biological sample present in the target area, eliminate effect of uncertainty caused by temperature dependent detector response of the at least one photocurrent signal generating light detector by calculating the standard deviation of a logarithm of the time dependent voltage signal, and based on the calculated attenuance, determine a blood glucose level associated with the biological sample present in the target area. 2. The system for detecting glucose in a biological sample according to claim 1, wherein the light absorbance change determining algorithm implemented processor is configured to calculate a normalization factor Qi(C,T) based on the output voltage Vi (t) of the ith preamplifier as a function of time and standard deviation σ according to the equation:Qi⁡(C,T)=σ⁡[log⁢⁢Vi⁡(t)]∑i=1N⁢⁢σ⁡[log⁢⁢Vi⁡(t)];wherein T is the temperature of the biological sample and C is the concentration of blood glucose in the biological sample. 3. The system for detecting glucose in a biological sample according to claim 1, wherein the light absorbance change determining algorithm implemented processor is configured to calculate a ratio factor Yij(C,T) based on the output voltage Vi(t) of the ith preamplifier and Vj(t) of the jth preamplifier as a function of time and standard deviation σ according to the equation:Yij⁡(C,T)=σ⁡[log⁢⁢Vi⁡(t)]σ⁡[log⁢⁢Vj⁡(t)];wherein T is the temperature of the biological sample and C is concentration of blood glucose in the biological sample. 4. The system for detecting glucose in a biological sample according to claim 1, further comprising an analog-to-digital convertor having a digitized voltage output. 5. The system for detecting glucose in a biological sample according to claim 4, wherein the light absorbance change determining algorithm implemented processor is configured to calculate a normalization factor Qi(C,T) based on the volt... | St. Louis Medical Devices, Inc. | A61B005/145 A61B005/1455 A61B005/00
US09801564 | B2 | 2017-10-31 | System and method for determining physiological parameters based on electrical impedance measurements | 2012-02-29 | A system and method for determining physiological parameters based on electrical impedance measurements is provided. One method includes obtaining electrical measurement signals acquired from a plurality of transducers coupled to a surface of an object and spatially pre-conditioning the obtained electrical measurement signals. The method also includes performing multiple-input-multiple-output (MIMO) analog to information conversion (AIC) of the spatially pre-conditioned electrical measurement signals to correlate the spatially pre-conditioned electrical measurement signals to separate the electrical measurement signals. | 1. A method for separating measurement signals, the method comprising: coupling a plurality of transducers on a surface of an object, the plurality of transducers comprising at least two excitation transducers and at least two measuring transducers;generating excitations on at least a subset of the plurality of transducers, the excitations driven in different configurations from at least one of the two excitation transducers;obtaining measurement response signals at a response detector, as acquired from the subset of the plurality of transducers;spatially pre-conditioning the obtained measurement signals to convert the measurement response signals into spatial vectors;performing multiple-input-multiple-output (MIMO) analog to information conversion (AIC) of the spatially pre-conditioned measurement signals to correlate the spatially pre-conditioned measurement signals in order to separate the measurement signals of a sub-space of desired physiological activities, as distinguished from a sub-space of undesired signals;wherein inputs of the MIMO AIC are the spatial vectors, and the MIMO AIC identifies and then tracks the sub-space of the desired physiological activities; andclassifying the measurement signals of the sub-space of desired physiological activities as respiratory information and classifying the sub-space of undesired signals as ambulatory motion information. 2. The method of claim 1, wherein the measurement response signals are electrical measurement response signals, and the MIMO AIC of the spatially preconditioned electrical measurement signals comprises a multi-channel sigma-delta (EA) MIMO AIC. 3. The method of claim 2, wherein the output of the MIMO AIC comprises digitized de-correlation coefficients of the electrical measurement signals and an uncorrected output data stream. 4. The method of claim 2, wherein spatially pre-conditioning the obtained electrical measurement signals comprises using a common mode averaging of all of the electrical measurement signals. 5. The method of claim 2, wherein the electrical measurement signals are obtained from electrodes in an electrocardiography (ECG) electrode placement configuration. 6. The method of claim 1, wherein classifying the sub-space of undesired signals as ambulatory motion information includes identifying at least one of a type of motion or a position. 7. The method of claim 6, wherein the ambulatory information includes an identification of at least one of breathing, bending, twisting, or reaching motion. 8. The method of claim 6, wherein the ambulatory motion information includes an identification of at least one of a supine position, seated position, or standing position. 9. The method of claim 1, further comprising using an output of the MIMO AIC and multi-lead transducer signals to determine an ambulatory motion index. 10. The method of claim 1, wherein the excitations comprises one of an electrical, magnetic, or radio-frequency excitation. 11. The method of claim... | General Electric Company | A61B005/0402 A61B005/053 A61B005/00 A61B005/08 G06F019/00
US09823318 | B2 | 2017-11-21 | Method for correlating magnetic resonance images with histological sections | 2015-01-13 | In a method for correlating magnetic resonance images with histological sections, a target tissue of a living animal is embedded in an enclosed matrix of an optical cutting temperature compound to obtain a packaged specimen on a platform oriented in a first guiding plane. The packaged specimen on the platform is subjected to an MRI examination by scanning along imaging planes parallel to the first guiding plane, and then subjected to a frozen sectioning procedure along sectioning planes in parallel with a second guiding plane which is parallel to the first guiding plane. | 1. A method for correlating magnetic resonance images with histological sections, comprising the steps of: (a) embedding a target tissue of a living animal in an enclosed matrix of an optimal cutting temperature compound to obtain a packaged specimen which is disposed on a platform such that the packaged specimen is oriented in a first guiding plane;(b) subjecting the packaged specimen on the platform to an MRI examination by scanning the packaged specimen along imaging planes oriented parallel to the first guiding plane, thereby obtaining magnetic resonance images of the packaged specimen;(c1) transferring the packaged specimen to a cryostat; and(d) subjecting the packaged specimen in the cryostat to a frozen sectioning procedure along sectioning planes in parallel with a second guiding plane which is parallel to the first guiding plane to thereby obtain histological sections each of which is readily correlatable with a corresponding one of the magnetic resonance images. 2. The method according to claim 1, wherein, in step (a), the enclosed matrix is disposed in an enclosed space defined between a container and the platform, which is disposed in the container and which is removable from the container along an in-plane direction, a flat surface of the platform extending in the in-plane direction and defining both the first and second guiding planes. 3. The method according to claim 2, before step (d), further comprising a step (c2) of removing the platform together with the packaged specimen disposed thereon from the container. 4. The method according to claim 3, wherein, in step (a), a lower half of the living animal is fixed to the flat surface of the platform. 5. The method according to claim 1, wherein, in step (b), the packaged specimen is subjected to a diffusion MRI examination to obtain diffusion weighted images. | Chang Gung Memorial Hospital, Linkou | G01V003/00 G01R033/31 G01R033/30 A61B005/055 G01N001/06 G01N001/31 G01N001/36 G01N001/42
US09799188 | B2 | 2017-10-24 | Intelligent earplug system | 2016-11-15 | The present teachings disclose methods and apparatuses for an intelligent, wireless earplug system comprising a speaker, an earplug, an electronics base, and a cable. The earplug fits into a user's ear canal and is operatively connected to a smart device, such as a smart phone, to intelligently transmit sounds from the smart device to the earplugs in a manner selected by the user, such as for example transmitting binaural beats to the earplugs to facilitate sleep. | 1. An earphone comprising: an earplug,a processor,a wireless interface,a battery,a speaker, anda sensor; whereinwhen the earphone is located in a user's ear, the sensor detects data indicative of the state of sleep of the user and provides that data to the processor, andthe processor is configured to cause the speaker to provide sound to the user based on the sensor data. 2. The earphone of claim 1, wherein the processor is configured to cause the speaker to provide a sound when the sensor data indicates that the user is in REM sleep. 3. The earphone of claim 2, wherein the sensor comprises an accelerometer, and the determination that the user is in REM sleep is based on head movement detected by the accelerometer. 4. The earphone of claim 1, wherein the sensor comprises one or more of a movement sensor, a temperature sensor, or a heart rate sensor. | Hush Technology Inc. | H04R025/00 G08B021/02 H04R001/10 A61M021/00 G08B003/10 G08B021/04 A61B005/01 A61B005/024 A61B005/06 A61B005/00
US09788893 | B2 | 2017-10-17 | Catheter with soft distal tip for mapping and ablating tubular region | 2014-11-20 | A catheter includes an elongated body having a longitudinal axis, a distal assembly distal the elongated body, the distal assembly having a tapered helical form comprising a larger, electrode-carrying proximal loop and a smaller, softer distal loop, and a shape-memory support member extending through at least the proximal loop. For example, the helical loop subtends at least about 720 radial degrees, with the proximal loop subtending about 360 radial degrees, and the distal loop subtending about 360 radial degrees. The softer distal loop with a straight distal end atraumatically guides the distal assembly into a tubular region so that the larger proximal loop can sit on the ostium of the tubular region with improved electrode and tissue contact. | 1. A catheter comprising: an elongated body having a longitudinal axis;a distal assembly distal the elongated body, the distal assembly having a helical form comprising a proximal loop and a distal loop, and a shape-memory support member extending through at least the proximal loop, the proximal loop and the distal loop being oriented obliquely at an angle relative to the longitudinal axis of the elongated body;at least one irrigated ablation ring electrode mounted on the proximal loop;a control handle proximal the elongated body; anda contraction wire having a proximal end in the control handle and a distal end anchored in the proximal loop, the control handle including a first control member configured to actuate the contraction wire to contract the proximal loop,wherein the proximal loop has a first flexibility and the distal loop has a second flexibility, and the second flexibility is greater than the first flexibility. 2. A catheter of claim 1, wherein the helical form subtends at least about 720 radial degrees. 3. A catheter of claim 1, wherein the helical form subtends about 765 radial degrees. 4. The catheter of claim 1, wherein the proximal loop subtends about 360 radial degrees. 5. The catheter of claim 1, wherein the distal loop subtends about 360 radial degrees. 6. The catheter of claim 1, wherein the distal assembly includes a generally straight distal end distal of the distal loop. 7. The catheter of claim 6, wherein the generally straight distal end is on-axis relative to the longitudinal axis of the elongated body. 8. The catheter of claim 1, wherein the proximal loop has a first radius and the distal loop has a second radius, and the first radius is greater than the second radius. 9. The catheter of claim 1, wherein the proximal loop and the distal loop are on-axis relative to the longitudinal axis of the elongated body. 10. The catheter of claim 1, wherein the oblique angle ranges between about 45 degrees to 105 degrees. 11. The catheter of claim 1, wherein the oblique angle ranges between about 75 to 105 degrees. 12. The catheter of claim 1, wherein the oblique angle is about 90 degrees. 13. The catheter of claim 1, wherein the proximal loop includes between about eight to twenty electrodes. 14. The catheter of claim 1, wherein the proximal loop includes about ten electrodes. 15. The catheter of claim 1, wherein the proximal loop includes about six electrodes spanning about 180 radial degrees. 16. The catheter of claim 1, further comprising at least one impedance recording ring electrode mounted on the proximal loop adapted to measure impedance. 17. The catheter of claim 1, further comprising at least one bi-polar recording ring electrode mounted on the proximal loop adapted to measure PV potentials. 18. The catheter of claim 1, further comprising a deflection wire extending through the elongated body, wherein the control handle includes a second control member configured to actuate the deflection wire to defle... | Biosense Webster (Israel) Ltd. | A61B005/042 A61B018/14 A61B005/053 A61B018/00 A61B017/00 A61B018/12 A61B034/20
US09782200 | B2 | 2017-10-10 | Neuromonitoring needle-cartridge gun system | 2014-02-09 | A neuromonitoring needle electrode placement gun which includes a housing, at least a first needle electrode and housing assembly, an elongated channel dimensioned and configured for accommodating axial movement of the at least a first needle electrode and housing assembly within the channel and for directing movement of the at least a first needle electrode and housing assembly and a pusher dimensioned and configured for engaging the at least a first needle electrode and housing assembly and for axially displacing the at least a first needle electrode and housing assembly. | 1. A neuromonitoring needle electrode placement gun which comprises: a main housing;an elongated channel having an axis and an axial extremity which is a discharge end intended for placement proximate to the body of a patient, said elongated channel having an interior dimensioned and configured for receiving a plurality of associated needle electrode cartridges each including at least one needle electrode having an exposed point, the plurality of associated needle electrode cartridges each being disposed at axially successive positions along said axis, each needle electrode cartridge matching the orientation of the preceding or of the following needle electrode cartridge within said elongated channel with each of the exposed points extending toward said discharge end, said elongated channel being carried by said main housing; a trigger mounted for pivotal movement on said main housing;a pusher cooperating with said trigger, said pusher being dimensioned and configured for engaging an associated needle electrode cartridge within said plurality of associated needle electrode cartridges that is nearest to said discharge end. 2. A neuromonitoring needle electrode placement gun as described in claim 1 wherein each cartridge includes a cartridge housing;at least one elongated needle electrode having an axial extent and having a pointed axial extremity carried by said cartridge housing; andat least a first generally planar wing having first and second faces thereof, said first generally planar wing being carried by said cartridge housing. 3. A neuromonitoring needle electrode placement gun as described in claim 2 wherein said at least a first generally planar wing is coated with a releasable adhesive on said first face thereof. 4. A neuromonitoring needle electrode placement gun as described in claim 2 wherein said cartridge has engagement surfaces disposed on opposed sides thereof, said engagement surfaces being dimensioned and configured for cooperation with an elongated track in the associated neuromonitoring needle electrode placement gun. 5. A neuromonitoring needle electrode placement gun as described in claim 4 wherein the respective engagement surfaces are dimensioned and configured with associated elongated channels in an associated neuromonitoring needle placement gun. 6. A neuromonitoring needle electrode placement gun as described in claim 2 further including a wire fixed to said at least one needle electrode. 7. A neuromonitoring needle electrode placement gun as described in claim 2 wherein said at least one needle electrode has insulation surrounding an axial portion of said axial extent of said at least one elongated needle electrode. 8. A neuromonitoring needle electrode placement gun as described in claim 7 wherein said insulation surrounding an axial portion of said axial extent of said at least one elongated needle electrode is spaced a predetermined distance from said pointed axial extremity. 9. A neuromonitoring needle elec... | Payam Andalib Joseph Haddad, III | A61B005/04 A61B017/34 A61B005/00 A61B005/0492 A61N001/05 A61B017/29 A61B050/30
US09832510 | B2 | 2017-11-28 | Deceptive indicia profile generation from communications interactions | 2012-03-06 | Systems, methods, computer-readable storage mediums including computer-readable instructions and/or circuitry for generating deceptive indicia profiles may implement operations including, but not limited to: detecting one or more indicia of deception associated with one or more signals associated with communication content provided by a participant in a first communications interaction; detecting one or more indicia of deception associated with one or more signals associated with communications content provided by the participant in a second communications interaction; generating a deceptive indicia profile for the participant according to indicia of deception detected in the communications content provided by the participant in the first communications interaction and indicia of deception detected in the communications content provided by the participant in the second communications interaction; and providing a notification associated with the deceptive indicia profile for the participant to the participant. | 1. A system comprising: means for detecting one or more indicia of deception associated with one or more signals associated with communication content provided by a participant in a first communications interaction;means for detecting one or more indicia of deception associated with one or more signals associated with communications content provided by the participant in a second communications interaction;means for generating a deceptive indicia profile for the participant according to indicia of deception detected in the communications content provided by the participant in the first communications interaction and indicia of deception detected in the communications content provided by the participant in the second communications interaction; andmeans for providing a notification associated with the deceptive indicia profile for the participant to the participant. 2. The system of claim 1, wherein at least one of the means for detecting one or more indicia of deception associated with one or more signals associated with communication content provided by a participant in a first communications interaction or the means for detecting one or more indicia of deception associated with one or more signals associated with communications content provided by the participant in a second communications interaction comprise: means for detecting one or more micro expressions of the first participant providing the communication content. 3. The system of claim 1, wherein at least one of the means for detecting one or more indicia of deception associated with one or more signals associated with communication content provided by a participant in a first communications interaction or the means for detecting one or more indicia of deception associated with one or more signals associated with communications content provided by the participant in a second communications interaction comprise: means for detecting an eye dilation of the first participant providing the communication content. 4. The system of claim 1, wherein at least one of the means for detecting one or more indicia of deception associated with one or more signals associated with communication content provided by a participant in a first communications interaction or the means for detecting one or more indicia of deception associated with one or more signals associated with communications content provided by the participant in a second communications interaction comprise: means for detecting an eye movement of the first participant providing the communication content. 5. The system of claim 1, wherein at least one of the means for detecting one or more indicia of deception associated with one or more signals associated with communication content provided by a participant in a first communications interaction or the means for detecting one or more indicia of deception associated with one or more signals associated with communications content provided by the participant in a second communications inte... | Elwha, LLC | G06F021/00 H04N021/422 H04N021/439 H04N021/44 H04N021/442 H04N021/4788 H04N021/84 A61B005/16 A61B005/021 A61B005/024
US09788790 | B2 | 2017-10-17 | Optical coherence tomography for biological imaging | 2010-05-28 | Described herein are catheters for use with Optical Coherence Tomography (OCT) that include an optical fiber core having a first refractive index and an interface medium having a second refractive index, where the first and second refractive indexes are mismatched such that receiving electronics configured to receive optical radiation reflected from the reference interface and the target operate in a total noise range that is within 5dB of the shot noise limit. These OCT catheters may include a silicon die mirror having a reflective coating that is embedded in the interface medium. The optical fiber can be fixed at just the distal end of the catheter, and may be managed within a handle that is attached to the proximal end of the catheter body, and is configured to allow rotation of both the catheter body and the optical fiber relative to the handle. | 1. A system for optical coherence tomography, comprising: a source of optical radiation;an optical fiber having a core providing a common path for optical radiation reflected from a reference interface and a target, the core having a first refractive index;receiving electronics configured to receive the optical radiation reflected from the reference interface and the target;an interface medium at the reference interface and in physical and optical contact with the optical fiber, the interface medium having a second refractive index, wherein the first refractive index and the second refractive index are mismatched so as to adjust the optical radiation reflected from the reference interface such that the receiving electronics operate in a total noise range that is within 5 dB of a shot noise limit; anda processor configured to generate an image of the target based upon the optical radiation received by the receiving electronics that are operating within 5 dB of the shot noise limit. 2. The system of claim 1, wherein the first refractive index and the second refractive index are mismatched such that the receiving electronics operate in a total noise range that is within 3 dB of the shot noise limit. 3. The system of claim 1, wherein the first refractive index and the second refractive index are mismatched such that the receiving electronics operate in a total noise range that is within 2 dB of the shot noise limit. 4. The system of claim 1, wherein the source of optical radiation is a swept-frequency source. 5. The system of claim 1, further comprising a mirror in the interface medium, the mirror configured to reflect the optical radiation from the optical fiber to the target. 6. The system of claim 5, wherein the mirror comprises a gold-coated silicon die. 7. The system of claim 1, wherein the interface medium is a solid transparent medium. 8. The system of claim 1, wherein the interface medium is in physical and optical contact with a distal end of the core. 9. The system of claim 1, further comprising a directional element configured to relay the optical radiation from the source to a distal end of the core. 10. The system of claim 1, wherein the first refractive index n1 and the second refractive index n2 are mismatched such that:PoutPin=(n1-n2n1+n2)wherein Pin is the power of the optical radiation at the distal end of the optical fiber prior to entering the interface medium, and wherein Pout is the power of the optical radiation reflected from the reference interface such that the receiving electronics operate in a total noise range that is within 5 dB of the shot noise limit. 11. The system of claim 10, wherein the first refractive index n1 and the second refractive index n2 are mismatched such that: Pdet=Pout(1−L)wherein L is the sum of all the optical losses from the distal end of the probe to the receiving electronics and Pdet is the power at the receiving electronics. 12. A catheter for use with optical coherence tomography, com... | AVINGER, INC. | A61B008/00 A61B005/00
US09820671 | B2 | 2017-11-21 | Devices and methods for assessing motor point electromyogram as a biomarker | 2008-05-16 | Described herein are methods of detecting the presence or progression of neuromuscular disorders (including ALS) based at least in part on diaphragm EMG signals taken at different points in time from the same implanted electrode(s). Signals may be used to diagnose or to monitor progression of a disorder or to track a treatment. Thus, diaphragm EMG signal may be used as a marker (e.g., a “biomarker”) for the detection or progression of a neurological disorder such as ALS. The characteristics or parameters of diaphragm EMG signals may be used to create an activity index, which may be output or compared and further analyzed. These signals may also be analyzed to show a difference between the left and right sides of the diaphragm. | 1. A method of diagnosing or tracking the progression or treatment of a neurological disorder that affects the function of a diaphragm, the method comprising: measuring a first EMG signal during a volitional breathing maneuver at an initial recording time before the start of therapy from one or more electrodes chronically implanted in a diaphragm;measuring a second EMG signal from the same one or more electrodes at a later recording time during a volitional breathing maneuver after the start of the therapy; andcomparing one or more characteristics of the first EMG signal and the second EMG signal and characterizing the presence, progression or treatment of a neurological disorder based on the comparison. 2. The method of claim 1, wherein the disorder comprises any of a centrally-mediated breathing disorder or a neurodegenerative disease. 3. The method of claim 1, wherein the comparing step comprises comparing one or more of amplitude, frequency, power spectrum, time-to-peak rise time, relaxation time, compound action potential rise time, compound action potential fall time, fast twitch muscle contribution, and slow-twitch muscle contribution. 4. The method of claim 1, further comprising: mapping the diaphragm to determine an active region of the diaphragm; andimplanting the one or more electrodes at an active region of the diaphragm. 5. The method of claim 1, further comprising: mapping the diaphragm to determine a phrenic nerve motor point of the diaphragm; andimplanting the one or more electrodes at the phrenic nerve motor point. 6. The method of claim 5, wherein the mapping step includes applying electrical stimuli to an area of diaphragm surface and identifying a site within the area that provokes a maximal physiological response to the stimuli. 7. The method of claim 6, wherein the mapping step includes observing a physiological response to the stimuli including any of: abdominal pressure or airway pressure, and diaphragm contraction. 8. The method of claim 1, wherein measuring the first EMG signal comprises: selecting a duration and/or period of time for collection, wherein the duration and/or period of time for collection for the first EMG signal is the same for the second EMG signal; andquantifying at least one characteristic of the EMG signal over the selected time. 9. The method of claim 8, wherein the time duration of collection of the EMG signal is selected from the group consisting of: a single breath and a particular number of breaths. 10. The method of claim 8, wherein the characteristic of the EMG signal quantified is selected from the group consisting of: amplitude, frequency, power spectrum, time-to-peak rise time, relaxation time, compound action potential rise time, compound action potential fall time, fast twitch muscle contribution, and slow-twitch muscle contribution. 11. The method of claim 1, wherein the steps of measuring the first EMG signal and measuring the second EMG signal comprises filtering the EMG signa... | SYNAPSE BIOMEDICAL, INC. | A61B010/00 A61B005/0488 A61N001/36
US09782145 | B2 | 2017-10-10 | Electronic catheter stethoscope | 2016-11-28 | An electronic catheter stethoscope measures and analyzes acoustic fields and dynamic pressure variations in the gaseous or liquid fluid inside a conventional medical catheter that is positioned in a patient's urologic, digestive, reproductive, cardiovascular, neurological or pulmonary system. Measurement transducers are installed in a housing connectable to multiple preselected medical catheters. The transducers detect bodily functions that are transmitted to the preselected catheter from within the body. The transducers, housing, electrical interface and signal processing electronics are positioned outside the body. | 1. A method for measurement and analysis of bodily functions using an electronic catheter stethoscope configured to connect to a plurality of types of catheters, comprising the steps of: connecting a housing of said electronic catheter stethoscope to a valve of a medical catheter selected from said plurality of types of catheters, wherein the valve is disposed on a proximal end of said medical catheter,wherein a purge/fill valve is mounted on said housing;inserting a distal end of said medical catheter into a patient's body while maintaining said proximal end and said housing outside the patient's body, wherein an interior of said medical catheter is in fluid communication with bodily fluids in said patient's body;said housing enclosing at least one transducer therewithin so that said at least one transducer is in communication with said interior of said medical catheter of said medical catheter and thus also in fluid communication with said bodily fluids in said patient's body,said at least one transducer sensing an acoustic field, a pressure, or a combination thereof of gaseous or liquid fluids within said medical catheter;introducing said gaseous or liquid fluids from said medical catheter into said housing through said purge/fill valve;purging said gaseous or liquid fluids from said medical catheter from said housing through said purge/fill valve;monitoring acoustic fields and pressure variations of said gaseous or liquid fluids within said medical catheter due to said bodily functions affecting said bodily fluids within said patient, wherein said gaseous or liquid fluids within said medical catheter are affected by changes in said bodily fluids within said patient's body and wherein said gaseous or liquid fluids within said medical catheter may include said bodily fluids as well depending on the bodily function being measured and analyzed;providing a data acquisition system for conditioning, amplifying and converting analog measurement signals generated by said at least one transducer to digital signals;positioning said data acquisition system externally of said housing; andproviding electrical communication between said at least one transducer and said data acquisition system and said at least one transducer. 2. The method of claim 1, wherein said gaseous or liquid fluids within said catheter includes said bodily fluids. 3. The method of claim 1, further comprising the step of: mounting a Luer-type connector on a distal end of said housing;opening said valve mounted on said proximal end of said catheter by employing said Luer-type connector to connect said housing to said proximal end of said medical catheter. 4. The method of claim 3, further comprising the steps of: filling or purging said housing and medical catheter with said gaseous or liquid fluids when said valve mounted on said proximal end of said catheter is opened by said Luer-type connector. 5. The method of claim 1, further comprising the steps of: selecting said medical ca... | University of South Florida The Charles Stark Draper Laboratory, Inc. | A61B007/04 A61B007/02 A61B005/03 A61B005/00
US09788763 | B1 | 2017-10-17 | Methods for magnetic particle capture and separation | 2014-07-09 | Methods of exerting magnetic forces to separate magnetic particles disposed in a portion of subsurface vasculature using a wearable device are provided. The magnetic forces can act to attract, slow, speed, or otherwise influence the magnetic particles in various applications. In some examples, different magnetic forces are exerted on respective sets of magnetic particles to separate the respective sets of magnetic particles. In some examples, similar magnetic forces are exerted on sets of magnetic particles, and separation of the sets of magnetic particles is related to properties of the sets of magnetic particles and/or of the environment of the sets of magnetic particles. In some embodiments, the magnetic particles are configured to bind to an analyte of interest. The separation of the magnetic particles can enable detection of one or more properties of the analyte, modification of the analyte, and/or extraction of the analyte bound to the magnetic particles. | 1. A method, comprising: positioning a device proximate to an external body surface that is proximate to a portion of subsurface vasculature, wherein the portion of subsurface vasculature contains first magnetic particles and second magnetic particles, wherein the first and second magnetic particles selectively bind with one or more analytes of interest;separating the first and second magnetic particles in the portion of subsurface vasculature, wherein separating the first and second magnetic particles comprises exerting magnetic forces on the first and second magnetic particles, wherein the magnetic forces are exerted by the device; anddetecting, using a detector, one or more properties of the one or more analytes of interest based on one or more properties of the separated first and/or second magnetic particles and/or the one or more analytes bound thereto. 2. The method of claim 1, further comprising introducing the first and second magnetic particles into the portion of subsurface vasculature. 3. The method of claim 1, wherein the device comprises at least one electromagnet, wherein separating the first and second magnetic particles in the portion of subsurface vasculature comprises operating at least one electromagnet of the device to exert the magnetic forces on the first and second magnetic particles, respectively. 4. The method of claim 1, wherein the device comprises at least one permanent magnet, wherein the magnetic forces are caused by at least one permanent magnet of the device. 5. The method of claim 1, wherein the magnetic forces include forces parallel to a direction of flow of blood in the portion of subsurface vasculature. 6. The method of claim 1, wherein the magnetic forces include forces toward a wall of the portion of subsurface vasculature. 7. The method of claim 1, further comprising: extracting the first magnetic particles from the portion of subsurface vasculature. 8. The method of claim 1, wherein separating the first and second magnetic particles in the portion of subsurface vasculature comprises collecting the first magnetic particles in the portion of subsurface vasculature while allowing the second magnetic particles to flow away from the portion of subsurface vasculature. 9. The method of claim 1, wherein separating the first and second magnetic particles in the portion of subsurface vasculature comprises: collecting, in an upstream region of the portion of subsurface vasculature, the first magnetic particles while allowing the second magnetic particles to flow away from the upstream region to a downstream region of the portion of subsurface vasculature; andcollecting, in the downstream region of the portion of subsurface vasculature, the second magnetic particles. 10. The method of claim 1, wherein the first magnetic particles are bound to the one or more analytes of interest and the second magnetic particles are not bound to the one or more analytes of interest. 11. The method of claim 1, wherein the fi... | Verily Life Sciences LLC | A61B005/00 A61B005/145 A61B005/05
US09788795 | B2 | 2017-10-17 | Magnetic resonance imaging method | 2014-06-06 | Provided is a method for compensating for tissue motion during magnetic resonance (MR) imaging, and an apparatus for use thereof. The method includes acquiring a plurality of short-time MR scan images; selecting a reference scan image from the acquired plurality of short-time MR scan images; defining a set of transformation images based on the acquired plurality of short-time MR scan images other than the selected reference scan image; registering the reference scan image and the defined set of transformation images; calculating an average of aligned, registered images of the defined set of transformation images; and generating a motion-corrected image based on the calculated average. | 1. A method for compensating for tissue motion during magnetic resonance (MR) imaging, the method comprising: acquiring a plurality of short-time MR scan images;selecting a reference scan image from the acquired plurality of short-time MR scan images;defining a set of transformation images based on the acquired plurality of short-time MR scan images other than the selected reference scan image;registering the reference scan image and the defined set of transformation images;calculating an average of aligned, registered images of the defined set of transformation images; andgenerating a motion-corrected image based on the calculated average. 2. The method according to claim 1, wherein registering the reference scan image and the set of defined transformation images includes aligning points in each image of the acquired plurality of short-time MR scan images in the defined set of transformation images to points in the reference scan image. 3. The method according to claim 2, wherein registering the reference scan image and the set of defined transformation images includes determining a global transformation representing an overall motion of the defined set of transformation images and a local transformation representing a local deformation of the defined set of transformation images. 4. The method according to claim 3, wherein the global transformation is a twelve parameter affine transformation. 5. The method according to claim 3, wherein the local transformation is a free form transformation based on B-splines. 6. The method according to claim 5, wherein the B-splines are one of linear B-splines and cubic B-splines. 7. The method according to claim 1, wherein registering the reference scan image and the defined set of transformation images includes performing a cost function analysis of the calculated average based on a negative of α-information between the reference scan image and the defined set of transformation images using Equation:C⁡(μ,α)=-1α⁡(α-1)⁢(∑l,k⁢p(α)⁡(l,k¦μ)pM⁡(l¦μ)⁢pR⁡(k)α-1-1). 8. The method according to claim 7, further including optimizing the cost function analysis using Equation:∂C∂μi=-1α⁡(α-1)⁢∑l,k⁢⁢α⁢⁢pα-1⁡(l,k❘μ)⁢∂p⁡(l,k❘μ)∂μi-(α-1)⁢pα⁡(l,k❘μ)⁢pM-1⁡(l❘μ)⁢∑k⁢⁢∂p⁡(l,k❘μ)∂μi[pM⁡(l❘μ)⁢pR⁡(k)]α-1. 9. The method according to claim 1, further including calculating a quality of the registration of the reference scan image and the defined set of transformation images using a Squared Sum of intensity Difference (SSD) of:SSD=1n⁢∑(FR-T⁡(FM))2. 10. The method according to claim 9, wherein calculating the quality of the registration of the reference scan image and the defined set of transformation images includes using a Correlation of Coefficients (CC) of:CC=∑(FR-F_R)⁢(T(FM⁢)-T⁡(F_M))∑(FR-F_R)2⁢∑(T⁡(FM)-T⁡(F_M))2. 11. The method according to claim 1, further including estimating underlying image intensity distributions of the registered reference scan image and the defined set of transformation images using a Parzen window. ... | The Research Foundation for The State University of New York | G06T007/00 A61B005/00 A61B005/055 G01R033/565 G06T005/00 G06T005/50 A61B005/20 G06T007/33 G06T007/246 G01R033/56
US09775520 | B2 | 2017-10-03 | Wearable personal monitoring system | 2015-11-03 | A system includes one or more sensors to detect activities of a mobile object; and a processor coupled to the sensor and the wireless transceiver to classify sequences of motions into groups of similar postures each represented by a model and to apply the models to identify an activity of the object. | 1. A wearable monitoring apparatus, comprising: a housing adapted to be worn by a person on a wrist of the person;an accelerometer disposed in the housing;a radio frequency transceiver disposed in the housing; anda processor disposed in the housing,wherein the processor analyzes signals from the accelerometer to determine whether the signals are indicative of a predefined exercise of the person, andwherein the processor computes calories burned by the person based on the signals from the accelerometer and the predefined exercise of the person, andwherein the processor transmits the calories burned to a remote device via the radio frequency transceiver. 2. The apparatus according to claim 1, further comprising: at least one bioelectric contact disposed on the housing, wherein signals from the at least one bioelectric contact are used by the processor to compute a pulse of the person. 3. The apparatus according to claim 2, wherein the processor generates an audible, tactile or visual alert if the pulse is greater than a first value or lower than a second value. 4. The apparatus according to claim 1, wherein the radio frequency transceiver is a Bluetooth® transceiver. 5. The apparatus according to claim 1, further comprising: a touch-sensitive display disposed on the housing. 6. The apparatus according to claim 1, further comprising: an angular position sensor disposed in the housing. 7. The apparatus according to claim 6, wherein the processor computes the calories burned by the person based on further signals from the angular position sensor. 8. The apparatus according to claim 6, wherein the processor computes sleep data based on movement of the person during sleep using the signals from the accelerometer and further signals from the angular position sensor. 9. The apparatus according to claim 1, further comprising: flash memory disposed in the housing. 10. The apparatus according to claim 1, further comprising: a vibration mechanism disposed on the housing for vibrating the housing. 11. The apparatus according to claim 1, further comprising: a rechargeable battery disposed in the housing. 12. The apparatus according to claim 1, further comprising: a mini-USB port disposed on the housing. 13. The apparatus according to claim 1, further comprising: a speaker and a microphone disposed on the housing. 14. The apparatus according to claim 1, further comprising: a push-button switch disposed on the housing. 15. The apparatus according to claim 1, further comprising: an optical sensor array disposed on the housing, the optical sensor array comprising a photo-detector and at least one of an infra-red light emitting diode and a visible light emitting diode, wherein signals from the optical sensor array are used by the processor to compute a pulse of the person. 16. The apparatus according to claim 15, wherein the processor generates an audible, tactile or visual alert if the pulse is greater than a first value or lower than a second value.... | Empire IP LLC | G08B001/08 A61B005/00 G06F019/00 G08B021/04 A61B005/01 A61B005/0205 A61B005/0402 A61B005/0476 A61B005/0488 A61B005/11 A61B007/04 A61B008/00 G08B021/02 H04M003/51 G08B025/01 A61B005/053 A61B008/06 A61B008/08 A61B005/026 H04W084/18
US09798921 | B2 | 2017-10-24 | Wrinkle care support device and method for supporting wrinkle care | 2015-10-21 | A wrinkle care support device includes: a wrinkle change detector that detects a change made in a wrinkle after occurrence of the wrinkle, the wrinkle having occurred in skin by provision of predetermined stimulation; a wrinkle care information determiner that determines, based on the detected change, user presentation information relating to wrinkle care; and an information outputter that outputs the determined user presentation information. Note that the wrinkle care support device may further include: an image obtainer that obtains an image where the skin is photographed; and a wrinkle detector that detects the wrinkle from the obtained image. The detection of the change may include a detection of a wrinkle disappearing time starting from when the predetermined stimulation is provided until the wrinkle is no longer detected. | 1. A wrinkle care support device comprising: a memory that stores instructions; anda processor that, when executing the instructions stored in the memory, performs operations comprising:obtaining an image where skin is photographed;detecting a wrinkle in the skin from the obtained image, the wrinkle appearing in response to predetermined stimulation;obtaining time-sequential images of the wrinkle after an appearance of the wrinkle in response to the predetermined stimulation;detecting a wrinkle disappearing time, which is an elapsed time from the appearance of the wrinkle in response to the predetermined stimulation to a disappearance of the wrinkle, based on the obtained time-sequential images;determining wrinkle care user presentation information corresponding to the detected wrinkle disappearing time; andoutputting the determined wrinkle care user presentation information. 2. The wrinkle care support device according to claim 1, wherein the memory further stores a table including wrinkle disappearing times and wrinkle care information associated with the respective wrinkle disappearing times, and the processor further determines wrinkle care information based on the detected wrinkle disappearing time, by referring to the table. 3. The wrinkle care support device according to claim 2, wherein the processor further obtains additional information including at least one of an age of a person having the skin, a temperature, or a humidity, wherein the table further includes the additional information associated with the respective wrinkle care information, andwherein the processor determines the wrinkle care information based on the obtained additional information, by referring to the table. 4. The wrinkle care support device according to claim 2, wherein the wrinkle care information includes types of wrinkles, andwherein the types of wrinkles include a potential wrinkle type. 5. The wrinkle care support device according to claim 2, wherein the wrinkle care information includes a wrinkle care suggestion for a potential wrinkle type. 6. The wrinkle care support device according to claim 1, wherein the processor detects a specific area of a face from the obtained image, and detects the wrinkle in the detected specific area. 7. The wrinkle care support device according to claim 1, wherein the processor evaluates effectiveness of wrinkle care based on the wrinkle disappearing time detected before the wrinkle care user presentation information is output and a wrinkle disappearing time detected after the wrinkle care user presentation information is output. 8. The wrinkle care support device according to claim 1, wherein the predetermined stimulation includes stimulation to displace the skin. 9. The wrinkle care support device according to claim 1, wherein the processor further creates an information display image by superimposing an image relating to a wrinkle care on a user representative image based on the determined wrinkle care user presentat... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | G06K009/00 G06T011/60 A61B005/00
US09788767 | B1 | 2017-10-17 | System and method for monitoring nitric oxide levels using a non-invasive, multi-band biosensor | 2017-06-14 | A biosensor includes a PPG circuit that emits light directed at living tissue at a plurality of wavelengths. A first and second spectral response of light reflected from the tissue is obtained around a first wavelength and a second wavelength. Using absorption coefficients for substances at the plurality of wavelengths, concentration levels of a plurality of substances such as Nitric Oxide may then be determined from the spectral responses. | 1. A biosensor for monitoring nitric oxide (NO) of a patient in vivo, comprising: a PPG circuit configured to: generate at least a first spectral response for light reflected around a first wavelength from skin tissue of the patient;generate at least a second spectral response for light reflected around a second wavelength from the skin tissue of the patient;a processing circuit configured to: obtain a value Lλ1 using the first spectral response, wherein the value Lλ1 isolates the first spectral response due to pulsating arterial blood flow;obtain a value Lλ2 using the second spectral response, wherein the value Lλ2 isolates the second spectral response due to pulsating arterial blood flow;obtain a value Rλ1, λ2 from a ratio of the value Lλ1 and the value Lλ2; andobtain a concentration level of nitric oxide using the value Rλ1, λ2 and a calibration database, wherein the calibration database is used to correlate the value Rλ1, λ2 and the concentration level of nitric oxide (NO). 2. The biosensor of claim 1, further comprising at least one of: a wireless transceiver configured to transmit the concentration level of nitric oxide to another device; ora port for a wired connection to communicate the concentration level of nitric oxide to another device. 3. The biosensor of claim 1, wherein the first wavelength has a high absorption coefficient for nitric oxide in arterial blood flow and is in a range of approximately 370 nm to 410 nm. 4. The biosensor of claim 3, wherein the second wavelength has a lower absorption coefficient for nitric oxide in arterial blood flow. 5. The biosensor of claim 1, wherein the calibration database includes a range of Rλ1, λ2 values and correlated concentration levels of nitric oxide and wherein the calibration database is generated using nitric oxide level measurements obtained from a large sample population. 6. The biosensor of claim 5, wherein the calibration database includes different correlations between the R values and NO concentration levels depending on the underlying skin tissue. 7. The biosensor of claim 1, further configured to: obtain a concentration level of one or more hemoglobin species; andcompensate for the concentration level of the one or more hemoglobin species to determine the value Rλ1, λ2. 8. The biosensor of claim 7, wherein the concentration level of the one or more hemoglobin species is determined by: analyzing spectral response of a plurality of wavelengths to determine differential absorption signals for each hemoglobin species; andusing a Beer-Lambert matrix equation for determining the concentration levels of the one or more hemoglobin species. 9. A method for monitoring nitric oxide (NO) of a patient in vivo, comprising: obtaining at least a first spectral response for light reflected around a first wavelength from skin tissue of the patient;obtaining at least a second spectral response for light reflected around a second wavelength from the skin tissue of the patient;obtaining a v... | SANMINA CORPORATION | A61B005/1455 A61B005/145 A61B005/00
US09782141 | B2 | 2017-10-10 | Motion tracking system for real time adaptive motion compensation in biomedical imaging | 2014-01-29 | The disclosure herein provides methods, systems, and devices for tracking motion of a patient or object of interest during biomedical imaging and for compensating for that motion in the biomedical imaging scanner and/or the resulting images to reduce or eliminate motion artifacts. In an embodiment, a motion tracking system is configured to overlay tracking data over biomedical imaging data in order to display the tracking data along with its associated image data. In an embodiment, a motion tracking system is configured to overlay tracking data over biomedical imaging data in order to display the tracking data along with its associated image data. In an embodiment, one or more detectors are configured to detect images of a patient, and a detector processing interface is configured to analyze the images to estimate motion or movement of the patient and to generate tracking data describing the patient's motion. The detector processing interface is configured to send the tracking data to a scanner controller to enable adjustment of scanning parameters in real-time in response to the patient's motion. | 1. A biomedical system for tracking motion of an object during biomedical imaging and for compensating for motion of the object, the biomedical system comprising: a biomedical imaging scanner configured to perform scanning of the object to generate biomedical images of the object;at least one detector for generating data describing at least a first landmark and a second landmark of the object, wherein the at least one detector is configured to be positioned relative to the object to enable the at least one detector to detect movement of the first landmark and the second landmark during the scanning;one or more computer readable storage devices configured to store a plurality of computer executable instructions; andone or more hardware computer processors in communication with the one or more computer readable storage devices and configured to execute the plurality of computer executable instructions in order cause the biomedical system to: utilize a first motion tracking technique to determine motion of the first landmark;utilize a second motion tracking technique to determine motion of the second landmark;apply a first weighting factor to the determined motion of the first landmark, wherein the first weighting factor is based on a historical accuracy of the first motion tracking technique;apply a second weighting factor to the determined motion of the second landmark, wherein the second weighting factor is based on a historical accuracy of the second motion tracking technique;determine motion of the object based on analyzing the determined motion of the first landmark and the second landmark;generate motion tracking data of the object; andcontrol at least one parameter of the biomedical imaging scanner to adjust scanner parameters based on the motion tracking data, the scanner parameters configured for controlling the biomedical imaging scanner to account for motion of the object during the scanning of the object. 2. The biomedical system of claim 1, wherein the at least one detector is positioned within a bore of the biomedical imaging scanner. 3. The biomedical system of claim 2, wherein the at least one detector only comprises components configured to not interfere with the biomedical imaging scanner. 4. The biomedical system of claim 1, wherein at least one of the first landmark and the second landmark comprises a facial feature of the subject. 5. The biomedical system of claim 4, wherein the facial feature comprises at least one tooth of the upper jawbone. 6. The biomedical system of claim 1, wherein at least one of the first landmark and the second landmark comprises an organ of the subject. 7. The biomedical system of claim 1, wherein at least one of the first landmark and the second landmark comprises an image projected onto the subject. 8. The biomedical system of claim 1, wherein the biomedical system is configured to utilize an atlas-segmentation technique for identifying the at least one landmark of the object. 9. A computer ... | KINETICOR, INC. THE UNIVERSITY OF HAWAII THE QUEEN'S MEDICAL CENTER | A61B005/055 A61B006/03 G06T007/20 A61B006/00 A61B006/04 A61B005/00 G06T007/00 G06T007/73 G06T007/207
US09801580 | B2 | 2017-10-31 | Driving assistance device, driving assistance method, information-providing device, information-providing method, navigation device and navigation method | 2013-03-22 | A heartbeat sensor obtains a heartbeat of a driver. A concentration level computing unit estimates a concentration level of the driver at a time later than a timing of obtaining the heartbeat by the heartbeat sensor, on the basis of the heartbeat obtained by the heartbeat sensor. A state determining unit controls driving assistance for a vehicle on the basis of a comparison of the concentration level of the driver estimated by the concentration level computing unit with a target concentration level of the driver that is set in association with a position or an environment in which the vehicle driven by the driver is predicted to travel. | 1. A driving assistance device comprising: a biometric information obtaining unit configured to obtain biometric information of a driver;a driver state estimating unit configured to estimate, on the basis of the biometric information obtained by the biometric information obtaining unit, a state of the driver at a time later than a timing of obtaining the biometric information by the biometric information obtaining unit; anda control unit configured to control driving assistance of a vehicle on the basis of a comparison of the state of the driver estimated by the driver state estimating unit with a condition for the state of the driver that is set in association with any one of a position and an environment in which the vehicle driven by the driver is predicted to travel,wherein the state of the driver is a concentration level of the driver;wherein the driver state estimating unit, on the basis of the biometric information obtained by the biometric information obtaining unit, estimates a tendency of the concentration level of the driver to decrease over time; andwherein the control unit promotes the driving assistance as a degree to which the state of the driver estimated by the driver state estimating unit does not satisfy the condition for the state of the driver that is set in association with any one of the position and the environment in which the vehicle driven by the driver is predicted to travel is greater. 2. The driving assistance device according to claim 1, wherein the control unit controls the driving assistance to be initiated when the state of the driver estimated by the driver state estimating unit does not satisfy the condition for the state of the driver that is set in association with any one of the position and the environment in which the vehicle driven by the driver is predicted to travel. 3. The driving assistance device according to claim 1, wherein the driver state estimating unit, on the basis of the biometric information obtained by the biometric information obtaining unit, estimates a tendency of the concentration level of the driver to decrease over time by estimating a degree of accumulation of the concentration level of the driver over time. 4. The driving assistance device according to claim 3, wherein the driver state estimating unit, on the basis of the biometric information obtained by the biometric information obtaining unit, estimates a period of time until the degree of accumulation of the concentration level of the driver over time is saturated and estimates a tendency of the concentration level of the driver to decrease over time on the basis of the time period until the degree of accumulation is saturated. 5. The driving assistance device according to claim 1, wherein the control unit, on the basis of the biometric information obtained by the biometric information obtaining unit, associates the state of the driver estimated at the timing of obtaining the biometric information by the biometric informatio... | TOYOTA JIDOSHA KABUSHIKI KAISHA | A61B005/18 A61B005/0245 A61B005/0476 G08G001/16 B60K028/00 B60W050/12 A61B003/11 A61B005/0205 A61B005/00 B60W040/08 B60W050/00
US09799433 | B2 | 2017-10-24 | Superconducting magnet | 2013-07-11 | A superconducting magnet includes a superconducting coil, a refrigerant container, a radiation shield, a vacuum container, a refrigerating machine cooling an interior of the refrigerant container, a tubular current lead passing from outside of the vacuum container to inside of the refrigerant container electrically connected to the superconducting coil, a power source electrically connected to the current lead, a manometer measuring a pressure inside of the refrigerant container, a thermometer to measure a temperature of the current lead, and a control unit connected to each of the power source, the manometer, and the thermometer. The control unit raises an output of the power source to vary a value of a current flowing into superconducting coil only when a measurement value of the manometer is higher than or equal to a set value and a measurement value of the thermometer is lower than or equal to a set value. | 1. A superconducting magnet, comprising: a superconducting coil;a refrigerant container accommodating said superconducting coil which is in a state of being immersed in a liquid refrigerant;a radiation shield surrounding said refrigerant container;a vacuum container accommodating said superconducting coil, said refrigerant container, and said radiation shield;a refrigerating machine cooling an interior of said refrigerant container and said radiation shield;a tubular current lead passing from outside of said vacuum container to inside of said refrigerant container to constitute a flow path of said gasified refrigerant and being electrically connected to said superconducting coil;a power source arranged outside of said vacuum container and electrically connected to said current lead;a manometer measuring a pressure inside of said refrigerant container;a thermometer arranged in said vacuum container to measure a temperature of said current lead; anda control unit connected to each of said power source, said manometer, and said thermometer;a disturbance magnetic field compensating coil arranged outside of said superconducting coil and immersed in said refrigerant in said refrigerant container for suppressing an influence of a disturbance magnetic field with respect to said superconducting coil;a persistent current switch immersed in said refrigerant in said refrigerant container and connected to said disturbance magnetic field compensating coil electrically in serial; anda heater arranged adjacent to said persistent current switch in said refrigerant container, immersed in said refrigerant, and electrically connected to said control unit,said control unit allowing a current from said power source to flow into said heater to gasify said refrigerant of an amount required cool said current lead while resetting an output of said disturbance magnetic field compensating coil by means of said persistent current switch,said control unit raising an output of said power source to vary a value of a current flowing into said superconducting coil only when a measurement value of said manometer is higher than or equal to a set value and a measurement value of said thermometer is lower than or equal to a set value. 2. The superconducting magnet according to claim 1, further comprising a heater arranged adjacent to said current lead in said refrigerant container to heat said current lead. 3. The superconducting magnet according to claim 1, wherein material of said current lead contains phosphorous deoxidized copper as a main component. 4. A superconducting magnet, comprising: a superconducting coil;a refrigerant container accommodating said superconducting coil which is in a state of being immersed in a liquid refrigerant;a radiation shield surrounding said refrigerant container;a vacuum container accommodating said superconducting coil, said refrigerant container, and said radiation shield;a refrigerating machine cooling an interior of said refrigerant containe... | MITSUBISHI ELECTRIC CORPORATION | H01F006/04 A61B005/055 H01F006/00 G01R033/3815 G01R033/38 H01F006/02 H01F006/06
US09820663 | B2 | 2017-11-21 | Correlating brain signal to intentional and unintentional changes in brain state | 2016-06-13 | Methods of analysis to extract and assess brain data collected from subject animals, including humans, to detect intentional and unintentional brain activity and other unexpected signals are disclosed. These signals are correlated to higher cognitive brain functions or unintended, potentially adverse events, such as a stroke or seizure, and to translation of those signals into defined trigger events or tasks. | 1. A system for detecting an intentional brain signal from a subject comprising: (a) a sensor for detecting data indicative of brain activity of the subject;(b) a computer memory module containing instructions configured to be executed by a computer processor, wherein the instructions comprise: (i) obtaining the data indicative of brain activity from the sensor:(ii) analyzing the obtained data indicative of brain activity, wherein analyzing comprises computing a spectrogram of the data, normalizing the spectrogram of the data at least once, time over frequency, and normalizing the spectrogram of the data at least once, frequency over time; and(iii) correlating the analyzed data to an intentional higher cognitive function from the subject; and(c) a computer processor module configured to execute the instructions in the memory module. 2. The system of claim 1, wherein the data is obtained non-invasively by applying the sensor to the subject. 3. The system of claim 2, wherein the sensor is a dry sensor or a wet sensor. 4. The system of claim 2, wherein the data is obtained from at least a single channel of EEG, EMG, EOG, MEG, ECoG, iEEG, fMRI, LFP, or a peripheral channel. 5. The system of claim 4, wherein the data is obtained from at least a single channel of EEG. 6. The system of claim 2, wherein the data is obtained from a multi-channel detector comprising at least a single channel EEG, EMG, EOG, MEG, ECoG, iEEG, fMRI, LFP or a peripheral channel. 7. The system of claim 2, wherein the data is received wirelessly from the sensor. 8. The system of claim 1, wherein the analyzing further comprises performing a principal and/or independent component analysis of the normalized spectrogram. 9. The system of claim 1, wherein analyzing comprises performing a temporal fragmentation analysis. 10. The system of claim 1, wherein analyzing comprises performing a preferred frequency analysis. 11. The system of claim 10, wherein the preferred frequency analysis is performed on the spectrogram. 12. The system of claim 1, wherein analyzing comprises performing a spectral fragmentation analysis. 13. The system of claim 1, wherein the instructions further comprise translating the analyzed data to effect a task associated with the higher cognitive function after the correlating. 14. The system of claim 13, wherein the task is selected from the group consisting of simulating, speech on a display, simulating speech with a voice synthesizer and movement of an artificial prosthesis. 15. The system of claim 1, wherein the higher cognitive function is selected from the group consisting of intent, speech, memory recall, planned motion, thought and imagination. 16. A system for detecting a brain signal from a subject correlated with at least one unintended event comprising: (a) a sensor for detecting data indicative of brain activity of the subject;(b) a computer memory module containing instructions configured to be executed by a computer processor, wherein t... | NeuroVigil, Inc. | A61B005/04 A61B005/16 A61B005/00 A61B005/0496 A61B005/048 A61B005/0476 A61B005/0488
US09820513 | B2 | 2017-11-21 | Depth proximity layering for wearable devices | 2014-12-23 | Techniques to determine depth for a number of wearable devices, which may be worn in layers are provided. A wearable device can receive measurements from sensors to include sensors from a number of other wearable devices. Based on the sensor measurements, the wearable device can determine a depth for the wearable devices. The depth can include an indication of which wearable device is closest to a user, which wearable device is closest to an external environment, or the like. | 1. An apparatus for a wearable computing device configured to determine depth with respect to another wearable computing device, the apparatus comprising: a first sensor coupled to a first wearable device, the first sensor to generate an indication of a first measured quantity;a sensor output collector to receive the indication of the first measured quantity from the first sensor and to receive an indication of a second measured quantity from a second sensor, the second sensor associated with a second wearable device; anda depth determinator to determine a depth with respect to the second wearable device based on the first measured quantity and the second measured quantity. 2. The apparatus of claim 1, the depth determinator to: determine whether the first measured quantity is greater than the second measured quantity; anddetermine the depth with respect to the second wearable device based on whether the first measured quantity is greater than the second measured quantity. 3. The apparatus of claim 1, wherein the first measured quantity and the second measured quantity comprise a temperature, a noise level, a light level, a wireless signal strength, an electrostatic measurement, or a biometric measurement. 4. The apparatus of claim 1, the sensor output collector to receive an indication of a third measured quantity from a third sensor, the third sensor associated with a third wearable device, the depth determinator to determine the depth with respect to the second wearable device and the third wearable device based on the first measured quantity, the second measured quantity, and the third measured quantity. 5. The apparatus of claim 4, the depth determinator to: determine whether the first measured quantity is greater than the second measured quantity;determine whether the first measured quantity is greater than the third measured quantity; anddetermine the depth with respect to the second wearable device and the third wearable device based on whether the first measured quantity is greater than the second measured quantity and whether the first measured quantity is greater than the third measured quantity. 6. The apparatus of claim 4, comprising: a fourth sensor coupled to the first wearable device, the fourth sensor to generate an indication of a fourth measured quantity;the sensor output collector to receive the indication of the fourth measured quantity from the fourth sensor, an indication of a fifth measured quantity from a fifth sensor and an indication of a sixth measured quantity from a sixth sensor, the fifth sensor associated with the second wearable device and the sixth sensor associated with the third wearable device; andthe depth determinator to validate the depth based on the fourth measured quantity, the fifth measured quantity, and the sixth measured quantity. 7. The apparatus for claim 1, the depth determinator to designate one of the first device, the second device, or the third device to use for a wearable device functio... | INTEL CORPORATION | G08B021/00 A41D001/00 A61B005/00 G06F001/16 A41B009/00 A41B001/08 A41D001/02 A61B005/01 A61B005/024
US09775537 | B2 | 2017-10-03 | Bio signal measuring apparatus and user monitoring system including the same | 2015-05-13 | A bio signal measuring apparatus is provided which includes a sampling unit, and a processing unit. The sampling unit samples a portion of a pulse included in a pulse signal penetrating a person to be measured, and portions sampled from a plurality of pulses included in the pulse signal are different from each other. The processing unit processes a sampled signal and measures a bio signal of the person to be measured based on at least one of a center frequency, a bandwidth, or amplitude of a reconstruction pulse obtained through the sampling. | 1. A bio signal measuring apparatus comprising: a processor configured to sample a portion of a pulse included in a pulse signal penetrating a person to be measured, wherein the sampling circuitry samples portions, being different from each other, sampled from a plurality of pulses included in the pulse signal being different from each other; andto process a sampled signal and measure a bio signal of the person to be measured based on at least one of a center frequency, a bandwidth, or amplitude of a reconstruction pulse obtained through the sampling,wherein the processor measures a heartbeat of the person to be measured, based on at least one of a center frequency, a bandwidth, or amplitude of the reconstruction pulse, andwherein the processor calculates a variation in at least one of a center frequency, a bandwidth, or amplitude of the reconstruction pulse and measures the heartbeat by detecting contraction and distension of a heart according to the variation. 2. The bio signal measuring apparatus of claim 1, further comprising an antenna configured to receive the pulse signal penetrating the person to be measured and provide the pulse signal to the sampling unit. 3. The bio signal measuring apparatus of claim 1, further comprising an analog-to-digital converter configured to convert a sampled signal into a digital signal and provide the digital signal to the processing unit. 4. The bio signal measuring apparatus of claim 1, wherein the pulse signal is a signal where the pulse is repeated every predetermined pulse repetition period. 5. The bio signal measuring apparatus of claim 4, wherein the processor samples the pulse in at least one time of the duration of the pulse, and wherein sampling about the pulse is performed every time obtained by adding a predetermined sampling interval to the pulse repetition period. 6. The bio signal measuring apparatus of claim 5, wherein the processor samples the pulse at a plurality of times of the duration of the pulse, and wherein an interval between the times in the duration of the pulse is equal to the sampling interval. 7. The bio signal measuring apparatus of claim 6, wherein the processor obtains an average of sampling values, corresponding to different points in the duration of the pulse, from among sampling values obtained by sampling the plurality of pulses and determines the average as a sampling value of a corresponding point. 8. The bio signal measuring apparatus of claim 1, wherein the duration of the reconstruction pulse is longer than that of the pulse. 9. The bio signal measuring apparatus of claim 1, wherein the processing unit calculates at least one of a center frequency, a bandwidth, or amplitude of the reconstruction pulse and measures the heartbeat by detecting contraction and distension of a heart according to at least one of the center frequency, the bandwidth, or the amplitude. | Industry-Adademic Cooperation Foundation, Yonsei University | A61B005/02 A61B005/05 A61B005/00 A61B005/024 A61B005/11
US09801584 | B2 | 2017-10-31 | Method for detection of aberrant tissue spectra | 2012-08-29 | A method is provided for determining contact of a sensor with a patient's tissue. The method comprises comparing the intensity of detected light at a first wavelength to a threshold, wherein the first wavelength is not used to determine a physiological characteristic of the patient, and determining if the sensor is in contact with the patient's tissue based on the comparison. In addition, a method is provided for determining the amount of light shunting during operation of the sensor. The method comprises comparing the intensity of detected light at a first wavelength to a threshold, wherein the first wavelength is not used to determine a physiological characteristic of the patient, and determining the amount of light shunting based on the comparison. | 1. A method for correcting light shunting, comprising: emitting light at three or more wavelengths from at least one emitter of a sensor into a patient's tissue, wherein the three or more wavelengths includes a first wavelength that is not used to determine a physiological characteristic of the patient;detecting the light with at least one detector of the sensor;comparing the intensity of the detected light at the first wavelength to a threshold;determining the amount of light shunting between the at least one emitter and the at least one detector based on the comparison; andprocessing the detected light at one or more of a second wavelength or a third wavelength to compensate for the amount of light shunting. 2. The method of claim 1, wherein processing the detected light comprises comprising correcting the intensity of the detected light at one or more of the second wavelength or the third wavelength based on the amount of light shunting. 3. The method of claim 2, wherein correcting the intensity comprises subtracting the amount of light shunting from the intensity of the detected light at the second wavelength or the third wavelength. 4. The method of claim 2, comprising determining a physiological characteristic based upon the corrected intensity of the detected light at the second or third wavelengths. 5. The method of claim 1, wherein comparing the intensity comprises comparing a ratio of the intensity of the detected light at the first wavelength to a reference intensity of light at the first wavelength to the threshold. 6. The method of claim 1, wherein comparing the intensity comprises comparing a ratio of the intensity of the detected light at the first wavelength to a reference intensity of light at a different wavelength to the threshold. 7. The method of claim 1, wherein the first wavelength is primarily absorbed by water. 8. The method of claim 7, wherein the first wavelength of the detected light is between about 1200 nanometers to about 1600 nanometers. 9. The method of claim 1, wherein the sensor comprises a sensor configured to measure tissue hydration. 10. The method of claim 1, wherein the sensor comprises a pulse oximetry sensor. 11. The method of claim 1, determining if the sensor is in contact with the patient's tissue based on the comparison. 12. A patient monitor, comprising: a port configured to connect to a sensor; anda processor configured to receive signals from the port and to use the signals when executing an algorithm configured to compare an intensity of a first wavelength of detected light to a threshold, wherein the first wavelength is not used to determine a physiological characteristic of a patient, and wherein the processor is also configured to execute a second algorithm configured to determine an amount of light shunting between at least one emitter of the sensor and at least one detector of the sensor based on the comparison, and wherein the processor is configured to process the signals related... | Covidien LP | A61B005/1455 A61B005/00
US09779500 | B2 | 2017-10-03 | Image processing device, pathological diagnosis support system, image processing program, and pathological diagnosis support method | 2014-06-27 | An image processing device includes an input unit and an alignment unit. The input unit inputs a cell shape image and a fluorescence image. The cell shape image shows a shape of a cell in a tissue section. The fluorescence image shows expression of a specific protein as a fluorescent bright point in a region same as a region in the tissue section. The alignment unit aligns the cell shape image and the fluorescence image based on an information source detected in both the cell shape image and the fluorescence image. | 1. An image processing device comprising: an interface which inputs a cell shape image showing a shape of a cell in a tissue section and a fluorescence image showing expression of a specific protein as a fluorescent bright point in a region same as a region in the tissue section; anda central processing unit (CPU) which aligns the cell shape image and the fluorescence image based on an information source detected in both the cell shape image and the fluorescence image,wherein the information source is staining information with a certain staining reagent, and the CPU calculates image feature amounts in the cell shape image and the fluorescence image from the staining information on the cell shape image and the staining information on the fluorescence image, respectively, and aligns the cell shape image and the fluorescence image based on the image feature amounts. 2. The image processing device according to claim 1, wherein the information source includes other staining information with a certain staining reagent different from the certain straining reagent, andthe CPU calculates image feature amounts in the cell shape image and the fluorescence image from the staining information on the cell shape image and the other staining information on the fluorescence image, respectively, and aligns the cell shape image and the fluorescence image based on the image feature amounts. 3. The image processing device according to claim 2, wherein the another staining reagent contains a fluorescent particle different from a fluorescent particle as a source of the fluorescent bright point in the fluorescence image at least in one of emission wavelength, particle size and shape. 4. The image processing device according to claim 1, wherein the CPU calculates from the cell shape image a cell feature amount which includes at least one factor of: area of a cell nucleus; average density of a cell nucleus; brightness variation of a cell nucleus; circularity of a cell nucleus; oblateness of a cell nucleus; and thickness of a nuclear membrane of a cell nucleus to area of the cell nucleus;the CPU specifies an observation target cell to observe based on the cell feature amount; andthe CPU which determines a status of the expression of the specific protein in the observation target cell based on the number of fluorescent bright points in the aligned fluorescence image on the observation target cell. 5. A pathological diagnosis support system comprising: the image processing device according to claim 1; andan image acquisition device which acquires the cell shape image and the fluorescence image used by the image processing device. 6. The image processing device according to claim 1, further comprising a display, wherein the CPU obtains analysis results based on the image feature amounts and the display displays the analysis results. 7. A non-transitory computer readable storage medium having an image processing program for execution on a computer the program causing the... | KONICA MINOLTA, INC. | G06K009/00 A61B005/00 G06T007/00 G06T003/00 G06K009/46 G06K009/52 G06K009/62 G06T007/60 G06T007/33 G01N021/64
US09788743 | B2 | 2017-10-17 | Methods for using a multiple electrode composite system in electrocardiogram detection systems | 2014-06-13 | A biomedical sensor system is disclosed that includes a plurality of electrodes and a contiguous adhesive material that is in contact with each of the plurality of electrodes. In certain embodiments a method is provided that includes the step of applying a first surface of adhesive material to a patient wherein the adhesive material includes at least two electrodes on second surface thereof that is opposite the first surface. The method also includes the step of receiving a time varying signal a first electrode of the at least two electrodes at a first location such that the time varying signal is not received at a second electrode of the at least two electrodes. | 1. A method of detecting a time varying signal from a subject, said method comprising the steps of: applying a first surface of a polymeric material to a subject, wherein said polymeric material is anisotropic and includes a polar material that is responsive to the presence of an alternating electric field, said polymeric material having at least two electrodes on a second surface thereof that is opposite the first surface; and receiving a time varying signal at a first electrode of the at least two electrodes at a first location, wherein the time varying signal is not received at a second electrode of the at least two electrodes. 2. The method as claimed in claim 1, wherein said polymeric material is an adhesive. 3. The method as claimed in claim 1, wherein said at least two electrodes are part of an array of electrodes that are provided on the polymeric material. 4. A method of detecting a time varying signal from a patient, said method comprising the steps of: applying a first surface of an adhesive material to a patient, wherein said adhesive material is anisotropic and is provided at a thickness of less than about 500 microns yet provides total peel strength of at least about 600 gram/inch, said adhesive material including a polar material that is responsive to the presence of an alternating electric field; and receiving a time varying signal at a first electrode that is in contact with a second surface of the adhesive material that is opposite the first surface, and not receiving the time varying signal at a second electrode that is in contact with the second surface of the adhesive. 5. The method as claimed in claim 4, wherein said first and second electrodes are part of an array of electrodes. 6. The method as claimed in claim 5, wherein said method further includes the steps of detecting any input signals from the array of electrodes, and selecting certain signals for further analysis in an ECG system. 7. A method of detecting a time varying signal from a subject, said method comprising the steps of: receiving the time varying signal from the subject at a first side of an anisotropic polymeric material, said anisotropic polymeric material including a polar material that is responsive to the presence of an alternating electric field; and providing a first output signal to a first conductive electrode on an opposing second side of the polymeric material, while not providing the first output signal to a second conductive electrode on the opposing second side of the polymeric material responsive to the first time varying signal even though the polymeric material is contiguous with both the first and second conductive electrodes. 8. The method as claimed in claim 7, wherein said polymeric material is an adhesive. 9. The method as claimed in claim 7, wherein said polar material aligns within the polymeric material in the presence of the time varying signal from the subject. 10. The method as claimed in claim 7, wherein said first and ... | FLEXcon Company, Inc. | A61B005/0408 A61B005/0428
US09782601 | B2 | 2017-10-10 | Pacing device with autonomous anti-tachycardia pacing | 2015-02-12 | In an example, an apparatus is described that includes an implantable housing, a heart signal sensing circuit configured to sense intrinsic electrical heart signals, a ventricular tachyarrhythmia (VT) detector circuit, operatively coupled to the heart signal sensing circuit, the detector circuit operable to detect a VT based on the sensed heart signals, a processor configured to control delivery of an anti-tachyarrhythmia pacing (ATP) therapy based on the detected VT, and an energy delivery circuit configured to deliver the ATP therapy in response to the detected VT, wherein the apparatus does not include a shock circuit capable of delivering a therapeutically-effective cardioverting or defibrillating shock. | 1. An anti-tachyarrhythmia system comprising: an anti-tachyarrhythmia pacing (ATP) apparatus that includes:an implantable housing;a heart signal sensing circuit configured to sense intrinsic electrical heart signals;a ventricular tachyarrhythmia (VT) detector circuit, operatively coupled to the heart signal sensing circuit, the detector circuit operable to detect a VT based on the sensed heart signals;a processor configured to control delivery of an anti-tachyarrhythmia pacing (ATP) therapy based on a VT detected by the detector VT, wherein the processor is configured to disable the delivery of the ATP therapy based on an indication of the absence from the system or unavailability of a defibrillation energy delivery circuit configured to deliver shock therapy;andan energy delivery circuit configured to deliver the ATP therapy in response to the detected VT. 2. The anti-tachyarrhythmia system of claim 1, wherein the processor is configured to control delivery of the ATP therapy without a communicated request from another implanted device. 3. The anti-tachyarrhythmia system of claim 1, wherein the processor is configured to receive at least one ATP therapy parameter, and wherein the processor is configured to enable or disable the delivery of the ATP therapy based on the at least one received ATP therapy parameter, the at least one ATP therapy parameter including an indication of the presence, absence or unavailability of a subcutaneous implantable cardioverter/defibrillator in the system. 4. The anti-tachyarrhythmia system of claim 1, further comprising: a shock detection module to detect delivery of the shock therapy. 5. The anti-tachyarrhythmia system of claim 4, wherein the processor is configured to inhibit or stop delivery of the ATP therapy when the shock detection module detects delivery of the shock therapy. 6. The anti-tachyarrhythmia system of claim 4, wherein the processor is configured to receive at least one post-shock pacing therapy parameter, and wherein the processor is configured to control the delivery of the post-shock pacing therapy based on the at least one post-shock pacing therapy parameter. 7. The anti-tachyarrhythmia system of claim 4, wherein the processor is configured to receive at least one post-shock pacing therapy parameter, and wherein the processor is configured to enable or disable the delivery of the post-shock pacing therapy based on the at least one received post-shock pacing therapy parameter. 8. The anti-tachyarrhythmia system of claim 1, wherein the processor of the anti-tachyarrhythmia pacing apparatus is further configured to control delivery of a bradycardia therapy and the energy delivery circuit is configured to deliver the bradycardia therapy. 9. The anti-tachyarrhythmia system of claim 1, further comprising: an implantable defibrillation device including:a implantable defibrillator housing;a plurality of subcutaneous electrodes configured to sense a plurality of electrical signals;a defibrilla... | Cardiac Pacemakers, Inc. | A61N001/39 A61N001/362 A61B005/042 A61B005/0464 A61B005/00 A61N001/05
US09775556 | B2 | 2017-10-03 | Apparatus and method for uroflowmetry | 2011-05-26 | The present invention generally pertains to an apparatus and a method for uroflowmetry. Uroflowmetry can generally be defined as the collection and measurement of a patient's urination for the purpose of determining urine flow data that are necessary to analyze the patient's urination by comparing the patient's urine flow data to standardized urine flow data of healthy individuals. The present invention allows medical professionals to non-invasively diagnose a patient's urination by collecting, measuring and computing urine flow data to determine if any medical conditions and symptoms that affect urine flow are present. The within invention is designed and constructed in such a manner as to allow for the collection and measurement of urination to be done quickly, easily and at a reduced cost. | 1. A method for uroflowmetry comprising: entering patient data into a data entry area of a uroflowmetry device with user input keys, wherein the patient data includes name, date of birth and sex;positioning a cup onto a cup support of the uroflowmetry device, wherein the cup support includes a force sensor;positioning a funnel into a funnel support in a body of the uroflowmetry device to drain into the cup;adjusting the distance of the funnel from the ground with an adjustment mechanism of the uroflowmetry device in response to a signal from user input keys; urinating into the funnel such that the urine flows through the funnel into the cup thereby activating the force sensor;measuring the increase in gravitational force on the cup by the force sensor at a first time interval due to the increase in volume of urine by converting the gravitational force obtained by the force sensor into a first electronic output;measuring the increase in gravitational force on the cup by the force sensor at a second time interval due to the increase in volume of urine by converting the additional gravitational force obtained by the force sensor into a second electronic output; andcomputing flow versus time urination data using a central processing unit that is electronically connected to the force sensor and digital memory, processing the first electronic output received from the force sensor, the second electronic output received from the force sensor and flow versus time equations previously stored in the digital memory,wherein the funnel includes a narrow end, a funnel body, a wide end, at least two scored seams originating at the narrow end and extending along the funnel body to the wide end and two elongated tabs adjacent to one of the at least two scored seams at the wide end which protrude above a top of the wide end,wherein the funnel body is positioned substantially horizontal to the funnel support in the body of the uroflowmetry device such that the funnel has a downgrade which is less than 45 degrees from the horizontal and slopes downward front the wide end of the funnel to the narrow end, andwherein the funnel support is formed from the body of the uroflowmetry device and has an internal shape that is substantially the same as an external shape of the funnel so than when the funnel is positioned into the funnel support the adjacent elongated tabs on the wide end of the funnel are positioned downward relative to the ground to catch and prevent bodily fluids from contaminating the uroflowmetry device. 2. The method of claim 1, wherein the flow versus time urination data include at least one of maximum flow, average flow, total volume, total test time, and the time to peak flow. 3. The method of claim 1, wherein an interior of the funnel is coated with a friction reducing substance, a pH testing substance and a blood presence indicator. 4. The method of claim 1, further comprising: displaying the flow versus time urination data and patient data in the ... | - | A61B005/20
US09801565 | B2 | 2017-10-31 | Methods and workflow design for improving specificity of breast MR | 2007-01-08 | In a breast imaging method, magnetic resonance data are acquired of one or both breasts of a subject (16) using a breast coil (20, 20′) coupled with said one or both breasts. The acquiring employs at least four independent channels (41, 42, 43, 44, 45, 46; 61, 62, 63, 64, 65, 66) per breast. The acquired magnetic resonance data are processed to generate at least one of (i) an image, (ii) a spectrum, and (iii) elasticity data. The at least four independent channels per breast are suitably embodied as a dual breast coil (20) including left breast conductors (31, 32, 33, 34, 35, 36) defining at least four independent left breast acquisition channels (41, 42, 43, 44, 45, 46) and right breast conductors (51, 52, 53, 54, 55, 56) defining at least four independent right breast acquisition channels (61, 62, 63, 64, 65, 66). | 1. A breast imaging method comprising: acquiring magnetic resonance data of one or both breasts of a subject using a dual breast coil coupled with said one or both breasts, the acquiring employing at least four independent channels per breast, the dual breast coil including at least four independent acquisition channels per breast wherein the at least four independent acquisition channels for each breast include acquisition channels defined by (i) at least one conductor positioned anterior of the breast, (ii) at least one conductor positioned posterior of the breast, (iii) at least one conductor positioned above the breast, (iv) at least one conductor positioned below the breast, (v) at least one conductor positioned lateral of the breast, and (vi) at least one conductor positioned medial of the breast, wherein at least one of the four independent acquistion channels is defined by an electrical or inductive coupling between two of the conductors; andprocessing the acquired magnetic resonance data to generate elasticity data. 2. The breast imaging method as set forth in claim 1, wherein the acquiring includes (i) acquiring at least some magnetic resonance data as undersampled magnetic resonance data using the at least four independent acquisition channels, and the processing includes (ii) performing image reconstruction including combining the undersampled magnetic resonance data from the at least four independent acquisition channels using SENSE to generate an image. 3. The breast imaging method as set forth in claim 2, wherein the at least four independent acquisition channels for each breast include at least six independent acquisition channels per breast, and the acquiring (i) and combining (ii) operate with a SENSE factor of at least four. 4. The breast imaging method as set forth in claim 1, wherein the processing includes: generating at least two inherently aligned images including at least two of (i) a T1 weighted image, (ii) a T2 weighted image, (iii) a diffusion weighted image, and (iv) a contrast dynamic image; andcombining the at least two generated images to produce a fused image. 5. The breast imaging method as set forth in claim 4, wherein the generating includes one of: generating at least the T2 weighted image, the diffusion weighted image, and the contrast dynamic image, andgenerating the T1 weighted image, the T2 weighted image, and the diffusion weighted image, but not the contrast dynamic image. 6. The breast imaging method as set forth in claim 4, wherein the acquiring includes: acquiring magnetic resonance data for the at least two of (i) the T1 weighted image, (ii) the T2 weighted image, (iii) the diffusion weighted image, and (iv) the contrast dynamic image over the same imaged volume, the generating producing images spanning the same volume, the combining operating on a voxel-by-voxel basis. 7. The breast imaging method as set forth in claim 4, wherein the acquiring further includes: acquiring magnetic resonance ela... | Koninklijke Philips N.V. | A61B005/055 G01R033/3415 G01R033/561
US09795320 | B2 | 2017-10-24 | Device for treatment of aneurysm | 2008-10-10 | In a method, system and device a member is provided around an aneurysm enabling treatment and monitoring of the aneurysm. In accordance with one embodiment the device is adapted to be adjusted postoperatively. Hereby the treatment can be efficiently carried out without having to perform surgery when adjusting the member. | 1. A device for treating a vascular aneurysm of a human or mammal patient, comprising: an implantable member adapted to hold fluid, the implantable member being adapted to be placed against an outside of a blood vessel having the aneurysm and to exercise a pressure on the aneurysm to prevent or reduce an expansion of the aneurysm, wherein the implantable member comprises two or more segments that are arranged after each other in an axial direction of the blood vessel and individually adjustable in a radial direction of the blood vessel such that the implantable member can be adjusted to follow an outer contour of the aneurysm, and wherein said two or more segments are adapted to provide a pressure that is equal or less than the diastolic blood pressure of the human or mammal patient. 2. The device according to claim 1, wherein the device is adapted to monitor an expansion of the aneurysm. 3. The device according to claim 1, comprising a control unit and a sensor, wherein the control unit is adapted to control the pressure applied onto the aneurysm based on the signal generated by the sensor. 4. The device according to claim 1, wherein the implantable member is a Y-shaped member, and wherein the implantable Y-shaped member is adapted to be placed at the Aorta Bifurcation. 5. The device according to claim 1, further comprising a pressure regulator adapted to achieve a substantially even pressure affecting the aneurysm from the outside of the blood vessel. 6. The device according to claim 1, wherein the implantable member is an elastic member, and wherein the elastic member is adapted to apply a pressure onto the aneurysm and has an expansion interval wherein the pressure applied is substantially constant or within an interval for treating and reducing expansion of the aneurysm. 7. The device according to claim 1, wherein the implantable member is spring loaded. 8. The device according to claim 1, wherein the implantable member is adapted to exert an essentially constant pressure or a pressure reducing the pressure difference, caused by changes in blood pressure in the blood vessel, on the aneurysm. 9. The device according to claim 1, further comprising a control device adapted to increase the pressure on the blood vessel when the aneurysm expands and being implanted in the patient, the control device adapted to increase the pressure on the blood vessel when at least one of: the aneurysm expands more than a predetermined value, and the aneurysm expands more than a predetermined value during a time period. 10. The device according to claim 1, further comprising logic circuitry for determining when the aneurysm is expanding based on a signal from a sensor or measuring device. 11. The device according to claim 1, further comprising an electrical pulse generator adapted to provide stimulation of the aneurysm wall via electrodes located on the inside of the implantable member, and a control unit adapted to vary the position of the electrical st... | - | A61M029/00 A61B005/07 A61B005/02 A61B005/00 A61B017/12 A61B017/135 A61M001/12 A61M001/10 A61B017/00 A61F002/30 A61B090/00
US09808172 | B2 | 2017-11-07 | Systems and methods for performing surface electromyography and range-of-motion test | 2009-06-01 | A soft-tissue-injury diagnostic system for diagnosing soft tissue injury within a patient includes a set of hand-held inclinometers configured and arranged for measuring angles formed between a first inclinometer disposed in proximity to a patient joint and a second inclinometer disposed distal to the joint during controlled patient movements of the joint. A plurality of measuring electrodes are coupleable in proximity to the patient's spine along the body portion that moves along the joint. The measuring electrodes are configured and arranged for measuring action potentials along patient muscle groups during the controlled patient movements of the joint and transmitting the measured action potentials to a dynamic surface electromyograph (“sEMG”) module. A hub receives and processes data from the inclinometers and the dynamic sEMG module. A visual display is configured and arranged for receiving and displaying the processed data. | 1. A soft-tissue-injury diagnostic system for diagnosing soft tissue injury within a patient, the soft-tissue-injury diagnostic system comprising: a set of hand-held inclinometers comprising a main unit and an auxiliary unit, the inclinometers configured and arranged for placement in proximity to two ends of a patient body portion that moves along a patient joint at one end of the body portion such that the main unit is positioned in proximity to the joint and the auxiliary unit is positioned along the patient body portion that moves along a patient joint, the inclinometers configured and arranged for measuring and recording angles formed between the main unit and the auxiliary unit during controlled patient movements of the joint, wherein the controlled patient movements of the joint comprise moving the joint between a first joint position and a second joint position that is different from the first joint position, wherein at least one controller is disposed on at least one of the main unit or the auxiliary unit, and wherein the at least one controller is configured and arranged for marking the recording of the measured angles at one or more of the first joint position and the second joint position when the at least one controller is actuated by a user, wherein a plurality of spaced-apart light-emitting diodes are disposed along each of the main unit and the auxiliary unit, and wherein the plurality of light-emitting diodes indicate degree markings and are configured and arranged to enable the user to view the measured angles between the main unit and the auxiliary unit during the controlled patient movements of the joint;a static surface electromyograph (“sEMG”) scanner, including: at least one set of measuring electrodes that are electrically coupled to the static-sEMG scanner, the measuring electrodes configured in a housing and arranged for coupling to an exterior portion of the patient in lateral proximity to a plurality of levels of the patient's spine along the body portion that moves along the joint, the at least one set of measuring electrodes configured and for measuring action potentials along patient muscle groups beneath the measuring electrodes and transmitting the measured action potentials to the static sEMG scanner, and wherein the at least one set of measuring electrodes is non-retractable in the housing;at least one ground electrode that is configured in the housing and arranged for placement against patient skin adjacent to the at least one set of measuring electrodes, wherein the at least one ground electrode is retractable in the housing independently of the at least one set of measuring electrodes to facilitate maintaining placement of the at least one ground electrode against patient skin during operation of the static sEMG scanner, wherein the at least one ground electrode is retractable to improve simultaneous contact with the patient skin between the static sEMG scanner and the at least one set of measuring electrodes,... | Precision Biometrics, Inc. | A61B005/0488 A61B005/11 A61B005/00
US09795790 | B2 | 2017-10-24 | Unwanted stimulation detection during cardiac pacing | 2016-11-10 | The disclosure relates to systems and methods for cardiac rhythm management. In some cases, a system may include a pulse generator for generating pacing pulses for stimulating a heart of a patient; a memory; and a sensor configured to sense a response to an unwanted stimulation and to produce a corresponding sensor signal. A processing circuit may receive the sensor signal for a time after one or more pacing pulses, and may derive a time-frequency representation of the sensor signal based on the received sensor signal. The processing circuit may use the time-frequency representation of the sensor signal to help identify unwanted stimulation. Once unwanted stimulation is detected, the processing circuit may change the pacing pulses to help reduce or eliminate the unwanted stimulation. | 1. A method of controlling a cardiac pacemaker to determine if a phrenic nerve of a patient is stimulated by one or more pacing pulses of the cardiac pacemaker, the method comprising: delivering one or more electrical pacing stimulation pulses to the patient;receiving a sensor signal for a time period after the one or more electrical pacing stimulation pulses, wherein the sensor signal is provided by a sensor configured to sense a stimulation of the phrenic nerve of the patient;deriving a power spectrum based on the received sensor signal during the time after the one or more pacing pulses;determining if the phrenic nerve of the patient has been stimulated based, at least in part, on the derived power spectrum; andchanging timing of the pacing stimulation pulses in response to determining that the phrenic nerve of the patient has been stimulated. 2. The method of claim 1, wherein the determining if the phrenic nerve of the patient has been stimulated includes comparing the derived power spectrum to one or more power spectrum templates representative of phrenic nerve stimulation. 3. The method of claim 1, wherein the sensor includes one or more of an accelerometer, a minute ventilation sensor, an electrical signal sensor, a pressure sensor, and an acoustic sensor, that is configured to sense an acceleration indicative of a response to a stimulation of the phrenic nerve of the patient. 4. The method of claim 1, wherein deriving the power spectrum includes deriving sensor power divided by sensor signal frequency versus sensor signal frequency. 5. The method of claim 1, wherein the determining if the phrenic nerve of the patient has been stimulated includes comparing the derived power spectrum to a power spectrum template corresponding to no phrenic nerve stimulation. 6. The method of claim 1, wherein the determining if the phrenic nerve of the patient has been stimulated includes comparing the derived power spectrum to a power spectrum template representative of phrenic nerve stimulation in the frequency domain. 7. The method of claim 1, including switching a pacing electrode configuration in response to determining that the phrenic nerve of the patient has been stimulated. 8. The method of claim 1, including changing a number of sites to which pacing stimulation is delivered in response to determining that the phrenic nerve of the patient has been stimulated. 9. The method of claim 1, including decreasing a pacing pulse energy by a specified amount in response to determining that the phrenic nerve of the patient has been stimulated. 10. The method of claim 1, including decreasing a pacing pulse energy by a specified amount in response to detecting contraction of skeletal muscle. 11. The method of claim 1, including: storing a phrenic nerve stimulation event identifier in memory of the cardiac pacemaker when determining that the phrenic nerve of the patient has been stimulated;generating electrical pacing pulse stimulation to both capture ... | Cardiac Pacemakers, Inc. | A61N001/37 A61N001/368 A61B005/00 A61B005/04 A61B005/11 A61N001/365
US09797972 | B2 | 2017-10-24 | Image processing apparatus, magnetic resonance imaging apparatus, and image processing method | 2015-05-22 | An image processing apparatus according to an embodiment includes processing circuitry. The processing circuitry is configured to generate, from three-dimensional medical image data, a first cross-sectional image and a second cross-sectional image intersecting the first cross-sectional image and is configured to change display locations of the first cross-sectional image and the second cross-sectional image on a display, in conjunction with a change in an intersecting location of the first and the second cross-sectional images. | 1. An image processing apparatus comprising processing circuitry configured to: generate, from three-dimensional medical image data, a first cross-sectional image and a second cross-sectional image intersecting the first cross-sectional image;arrange, on a display, the first cross-sectional image to be located in an extending direction of an intersecting line with the first cross-sectional image drawn in the second cross-sectional image; andchange display locations of the first cross-sectional image and the second cross-sectional image on the display, in conjunction with a change in an intersecting location of the first and the second cross-sectional images. 2. The image processing apparatus according to claim 1, wherein the processing circuitry is further configured to arrange the first cross-sectional image in a display location that does not cause the first cross-sectional image to overlap the second cross-sectional image. 3. The image processing apparatus according to claim 1, wherein the processing circuitry is further configured to display the first and the second cross-sectional images, and to display a guide line indicating the extending direction of the intersecting line. 4. The image processing apparatus according to claim 3, wherein the processing circuitry is further configured to display the guide line so as to connect the first and the second cross-sectional images together. 5. The image processing apparatus according to claim 1, wherein the processing circuitry is further configured to: generate, from the three-dimensional medical image data, the first cross-sectional image that is a basic cross-sectional image used for diagnosis of a heart and the second cross-sectional image that is an auxiliary cross-sectional image used when setting an image taking location of the basic cross-sectional image; anddisplay the basic cross-sectional image so as to be located in an extending direction of an intersecting line with the basic cross-sectional image drawn in the auxiliary cross-sectional image; anddisplay a guide line indicating the extending direction of the intersecting line so as to connect the basic and the auxiliary cross-sectional images together. 6. The image processing apparatus according to claim 1, wherein the processing circuitry is further configured to adjust display sizes of cross-sectional images in such a manner that a display size of at least one selected from between a cross-sectional image being operated on and one or more cross-sectional images related to the cross-sectional image being operated on is relatively larger than display sizes of other cross-sectional images. 7. The image processing apparatus according to claim 1, wherein the processing circuitry is further configured to exercise control so that at least one selected from between a cross-sectional image being operated on and one or more cross-sectional images related to the cross-sectional image being operated on is displayed, while other cross-sectio... | TOSHIBA MEDICAL SYSTEMS CORPORATION | G06K009/62 G01R033/561 A61B005/055 A61B005/00 A61B006/00 G01R033/48 G01R033/54 G01R033/56 G06T007/00 G06T019/00 A61B006/03 A61B008/08
US09818310 | B2 | 2017-11-14 | Assessment of nutrition intake using a handheld tool | 2015-06-22 | Embodiments regard nutrition assessment using a handheld device. An embodiment of an apparatus includes a handle with a controller within the handle, an attachment arm extending from the handle, and a user-assistive device coupled with an end of the attachment arm, wherein the apparatus is to determine a mass held by the user-assistive device, the determination being made during a task by a user of the handheld tool including manipulation of the handheld tool. | 1. A handheld tool, comprising: a handle;a controller within the handle;an attachment arm extending from the handle;a user-assistive device coupled with an end of the attachment arm;wherein the handheld tool is configured to determine a mass held by the user-assistive device, the determination being made during a task by a user of the handheld tool, the task including manipulation of the handheld tool, wherein the handheld tool is configured to determine the mass based at least in part on an operational value for the handheld tool required to maintain the user-assistive device in a certain orientation, wherein the operational value is one of a force value, a current value, or a power value; anda vertical actuator, the vertical actuator coupled to provide for vertical motion of the user-assistive device;wherein the handheld tool is configured to determine the mass based at least in part on an additional operational value for the vertical actuator to maintain the user-assistive device in the certain orientation when holding the mass in comparison with an operation value for the vertical actuator to maintain the user-assistive device in the certain orientation without the mass. 2. The handheld tool of claim 1, wherein a determination of the additional operational value is based at least in part on an additional amount of current required by the vertical actuator. 3. The handheld tool of claim 1, wherein the determination of the mass is further based on a geometry of the user-assistive device. 4. The handheld tool of claim 1, further comprising a memory unit for storage of data regarding determined mass that is consumed. 5. The handheld tool of claim 4, wherein the handheld tool is configured to store data regarding a consumed mass in the memory unit upon identifying an end of the task. 6. The handheld tool of claim 5, wherein the handheld tool is to identify the end of the task based at least in part on a motion of the handheld tool and a loss of mass from the user-assistive device. 7. The handheld tool of claim 5, further comprising a communication interface, the handheld tool to communicate the data stored in the memory unit via the communication interface. 8. The handheld tool of claim 1, wherein the mass is a mass of food, the task is an eating task, and the activity is the eating of a meal by the user. 9. A method comprising: monitoring motions of a handheld tool during an activity;detecting the beginning of a task, the task including manipulation of the handheld tool; anddetermining by the handheld tool a mass held by a user-assistive device coupled with the handheld tool;wherein the determination of the mass is made during the task, wherein determining the mass is based at least in part on an operational value for the handheld tool to maintain the user-assistive device in a certain position, wherein the operational value is one of a force value, a current value, or a power value, and wherein the handheld tool includes a vertical actua... | Verily Life Sciences LLC | G09B019/00 A61B005/00 A61B005/11 G01G019/56 A47G021/02 G01G019/414 G06F019/00
US09775522 | B2 | 2017-10-03 | Integration delayed optical feedback in image guidance | 2013-03-26 | The present invention relates to a system (10) for navigating and positioning an instrument (12, 40, 112) in a body (14). The instrument (12, 40, 112) is detectable by the system (10) and the system (10) displays a view to an operator. The instrument (12, 40, 112) is adapted to identify tissue parameters. The system (10) is adapted to display an image being a combined image of the interior of a body (14) and an indication of the tissue parameter at the appropriate earlier locations. The present invention further relates to a method (44) including determining tissue parameters and displaying the parameters in an image of a body (14). The present invention further relates to a software implemented method (44) for being executed on a digital processor. The present invention further relates to an instrument (12, 40, 112). | 1. A system for navigating and positioning an instrument in a body, wherein the instrument comprises a sensor for optical inspection of tissue in the body, the instrument being configured to generate a tissue signal, the system comprising: a tissue-type determination device configured to receive the tissue signal from the instrument at times that correspond with a timestamp Ts, the tissue-type determination device being further configured to determine a set of parameters indicative of tissue type based on the tissue signal from the instrument;a medical imaging device configured to form an image of an interior of the body, the medical imaging device being configured to record a temporal sequence of images, wherein each image is separated from a neighboring image in the temporal sequence by a temporal distance that corresponds with a timestamp Tf, wherein an interval between neighboring Tf timestamps is smaller than an interval between neighboring Ts timestamps;an image processing unit time-synchronized with the tissue-type determination device, wherein the image processing unit is configured to establish a combined image of the interior of the body at each timestamp Tf that comprises (i) a current instrument position (xc, yc) and (ii) an indication of tissue-type at a previous spatial position of the instrument (xs, ys) that corresponds with a tissue signal retrieval location for a timestamp Ts so that the combined image of the interior of the body comprises both the current instrument position and the indication of tissue-type at the previous spatial position of the instrument that corresponds with the tissue signal retrieval location for the corresponding timestamp Ts; anda display unit for displaying the combined image of the interior of the body. 2. The system according to claim 1, wherein the sensor for optical inspection of tissue is configured for determining blood, water, fat, collagen, bile, beta-carotene content, oxygenation of blood and scattering parameters. 3. The system according to claim 1, wherein a color saturation level of the indication of tissue-type is indicative of a measured quantity of a tissue parameter. 4. The system according to claim 1, wherein the tissue-type determination device is connected to a spectrometer configured for transforming a signal from the tissue-type determination device into a spectrum. 5. The system according to claim 4, wherein the tissue parameter is determined based on the spectrum. 6. The system according to claim 1, wherein the tissue parameter is displayed as a solid line and a width of the solid line is an indicator of a tissue parameter determination inaccuracy. 7. The system according to claim 1, wherein the set of parameters indicative of tissue type based on the tissue signal from the instrument comprises: performing multivariate statistical analysis, wherein the multivariate statistical analysis comprises one or more of (i) principal components analysis or (ii) partial least squares... | Koninklijke Philips N.V. | A61B005/05 A61B005/00 A61B005/06 A61B017/00 A61B006/12 A61B090/00 A61B090/30 A61B034/20
US09814419 | B2 | 2017-11-14 | Pulse oximetry devices and systems and methods of making the same | 2016-04-07 | An array of emitters includes a device substrate having first and second sides, a thermally and electrically conductive layer disposed on the first side of the device substrate, and an interconnect layer disposed on a first plurality of portions of the second side of the device substrate. The array of the emitters further includes a plurality of emitters disposed in a second plurality of portions of the device substrate, where the plurality of emitters is electrically coupled to the thermally and electrically conductive layer. Also, the array of the emitters includes a plurality of wirebond contacts configured to electrically couple a portion of the interconnect layer to a corresponding emitter of the plurality of emitters, and a plurality of encapsulations, where one or more encapsulations of the plurality of encapsulations are disposed on at least a portion of a corresponding wirebond contact of the plurality of wirebond contacts. | 1. A pulse oximetry system, comprising: a pulse oximetry device comprising: a device substrate having a first side and a second side;a thermally and electrically conductive layer disposed on the first side of the device substrate;an interconnect layer disposed on the second side of the device substrate;a plurality of arrays of emitters disposed on the second side of the device substrate, wherein the plurality of arrays of the emitters is electrically coupled to the thermally and electrically conductive layer;a plurality of detectors disposed on the second side of the device substrate and configured to detect emission signals from one or more emitters of the plurality of arrays of the emitters;a plurality of wirebond contacts configured to electrically couple a portion of the interconnect layer to a corresponding emitter of the plurality of arrays of the emitters; anda processor configured to process the detected emission signals. 2. The pulse oximetry system of claim 1, wherein one or more emitters of the plurality of arrays of the emitters are disposed orthogonally opposite to one or more detectors of the plurality of detectors. 3. The pulse oximetry system of claim 1, wherein one or more emitters of the plurality of arrays of the emitters are disposed opposite to one or more detectors of the plurality of detectors, and wherein the one or more emitters of the plurality of arrays of the emitters are disposed in parallel to the one or more detectors of the plurality of detectors. 4. The pulse oximetry system of claim 1, wherein one or more emitters of the plurality of arrays of the emitters are disposed opposite to one or more detectors of the plurality of detectors, and wherein the one or more emitters of the plurality of arrays of the emitters are disposed in parallel and at an offset to the one or more detectors of the plurality of detectors. 5. The pulse oximetry system of claim 1, wherein the pulse oximetry device is configured to operate in a transmission mode, a reflection mode, or both. 6. The pulse oximetry system of claim 1, wherein the pulse oximetry device is configured to be disposed on one side of a body part. 7. The pulse oximetry system of claim 1, wherein the pulse oximetry device is configured to be disposed on two sides of a body part. 8. The pulse oximetry system of claim 1, wherein the pulse oximetry device is configured to conform to a shape of a body part. 9. The pulse oximetry system of claim 1, wherein the pulse oximetry device comprises an external connector. 10. The pulse oximetry system of claim 1, wherein the pulse oximetry device comprises a first plurality of electrical traces configured to provide an electrical connection to each emitter of the plurality of arrays of the emitters, and a second plurality of electrical traces configured to provide an electrical connection to each detector of the plurality of detectors. 11. The pulse oximetry system of claim 10, wherein the first plurality of electrical trace... | General Electric Company | A61B005/1455 A61B005/00 A61B005/029 A61B005/08 H01L025/075 H05K001/18 H01L033/62
US09788733 | B2 | 2017-10-17 | Method and device for detecting and assessing reactive hyperemia using segmental plethysmography | 2014-03-11 | A method for measuring reactive hyperemia in a subject is disclosed. The method includes performing a first segmental cuff plethysmography to generate a baseline arterial compliance curve and/or a baseline pressure-area (P-A) curve, performing a second segmental cuff plethysmography to generate a hyperemic arterial compliance curve and/or a hyperemic P-A curve, and calculating an area between the baseline and the hyperemic curves. The size of the area can be used as an indication of endothelial dysfunction (ED) and ED-related diseases. | 1. A method for measuring reactive hyperemia in a subject, comprising: performing a first segmental cuff plethysmography having an inflation phase and a deflation phase and generating a first cuff compliance curve based only on data collected during the deflation phase of the first segmental cuff plethysmography;generating a baseline arterial compliance curve and/or a baseline pressure-area (P-A) curve on a portion of the body of the subject, wherein the cuff pressure is increased to a first peak cuff pressure during the inflation phase and immediately reduced from the first peak cuff pressure during the deflation phase;performing a second segmental cuff plethysmography having an inflation phase, a holding phase, and a deflation phase and generating a second cuff compliance curve based only on data collected during the deflation phase of the first segmental cuff plethysmography;generating a hyperemic arterial compliance curve and/or a hyperemic P-A curve, wherein the cuff pressure is increased to a second peak level during the inflation phase, maintained at the second peak cuff pressure for a predetermined period of time during the holding phase, and then reduced from the second peak cuff pressure during the deflation phase;calculating the difference between the baseline arterial compliance curve and the hyperemic arterial compliance curve as an area between the arterial compliance curves, and/or the difference between the baseline P-A curve and the hyperemic P-A curve as an area between the P-A curves; andmeasuring a level of reactive hyperemia based on the area between the arterial compliance curves and/or the area between the P-A curves,wherein said first and second cuff compliance curves are generated using a flowmeter that directly measures the volume change in the cuff during the deflation phase. 2. The method of claim 1, wherein said cuff is inflated during the first and second segmental cuff plethysmography using a low-frequency pump. 3. The method of claim 2, wherein said low-frequency pump has an operation frequency of 10 Hertz or lower. 4. The method of claim 2, wherein said low-frequency pump has an operation frequency of 2 Hertz or lower. 5. The method of claim 1, further comprising: determining a level of endothelial dysfunction (ED) in the subject based on the result of the measuring step. 6. The method of claim 1, further comprising: determining a risk of an ED-related disease based on the result of the measuring step. 7. The method of claim 6, wherein the ED related disease is selected from the group consisting of hypertension, atherosclerosis, cardiovascular disease, stroke, atrial fibrillation, congestive heart failure, peripheral vascular disease, septic shock, hypercholesterolemia, type I and II diabetes, erectile dysfunction, rheumatic arthritis, HIV, liver diseases and pre-eclampsia. 8. The method of claim 1, wherein the first peak cuff pressure and the second peak cuff pressure are in the range of 150-180 mmHg. 9. ... | CORDEX SYSTEMS, INC. | A61B005/02 A61B005/022 A61B005/00 A61B005/0295 A61B005/0225
US09830687 | B2 | 2017-11-28 | Image processing device, magnetic resonance imaging apparatus and image processing method | 2014-02-18 | In order to reduce errors in assessment of a structure caused by noise and remove the noise superimposed in an image without causing artificiality in the result image after the removal while retaining significant information, a similarity is calculated by comparing a reference image generated from a plurality of original images with each original image and set as an index of noise determination. The respective original images are smoothed and synthesized using the said index to acquire a final image after the noise removal. | 1. An image processing device comprising: an image processing unit for synthesizing a plurality of original images to generate a reference image, calculating a similarity for each of the plurality of original images by comparing with the generated reference image, smoothing the plurality of original images based on the calculated similarity, and synthesizing the plurality of smoothed original images to acquire a synthesized imagewherein the image processing unit comprises:a band dividing section for dividing the plurality of original images for each predetermined wavelength band to generate a plurality of wavelength-band original images;a reference image generating section for synthesizing the plurality of wavelength-band original images for each of the predetermined wavelength bands to generate the reference image for each of the wavelength bands;a similarity calculating section for comparing each of the plurality of wavelength-band original images with the reference image in the same wavelength band to calculate the similarity for each of the wavelength-band original images;a denoising section for generating denoised images from each of the wavelength-band original images by smoothing each of the plurality of wavelength-band original images using the similarity of the said wavelength-band original image; anda synthesizing section for synthesizing the denoised images to generate the synthesized image. 2. The image processing device according to claim 1, wherein the band dividing section calculates a cut-off wavelength from the from a predetermined maximum noise wavelength and generates the wavelength-band original images from each of the plurality of original images by the said cut-off wavelength. 3. The image processing device according to claim 1, wherein the image processing unit determines whether or not the synthesized image converges each time it is generated and further comprises a convergence determining section for replacing the synthesized images with the original images in case of a negative determination,the band dividing section divides the replaced original images each time the synthesized images are replaced with the original images to newly generate the wavelength-band original images,the reference image generating section generates the reference image each time the wavelength-band original image is generated, and the similarity calculating section calculates the similarity each time the reference image is generated,the denoising section generates the denoised image each time the similarity is calculated,the synthesizing section synthesizes a denoised image generated from the same original image to generate a synthesized denoised image each time the denoised image is generated, and then synthesizes all the synthesized denoised images that were synthesized to generate the synthesized image, andthe convergence determining section replaces the original images with the synthesized denoised images generated just before in case of a ... | HITACHI, LTD. | G06K009/00 G06K009/40 G06T005/00 G06T003/40 A61B005/00 G06T005/50 A61B008/00 A61B005/055
US09778027 | B1 | 2017-10-03 | Apparatus and method for imaging feet | 2015-09-18 | An apparatus and method for determining contours of a patient's foot ankle and lower leg. The apparatus includes an alignment structure that orientates the foot for imaging. The alignment structure includes at least one support member that engages the plantar surface of the foot substantially only in the immediate area of the lateral metatarsal heads of the foot, preferably the fifth metatarsal head. The support generates a dorsally-directed force that locks the midtarsal joint. The alignment structure further includes a saddle that engages the rearfoot and a laser beam for aligning the second metatarsal head with the distal one-third of the lower leg to place the subtalar joint in a neutral condition. The foot is thus suspended in space such that imaging is able to produce an accurate measurement of the subject areas of the foot, ankle and lower leg without distortion of the soft tissues or bone structure. | 1. An apparatus for use in determining contours of a foot and/or an associated ankle and lower leg of a patient, said apparatus comprising: a support and alignment assembly that orientates said foot for imaging, said support and alignment assembly comprising: at least one metatarsal head support member that engages a plantar surface of said foot substantially only beneath a lateral forefoot area of said foot; anda mechanism that provides movement of said lateral forefoot area of said foot relative to said at least one metatarsal head support member so that said lateral forefoot area of said foot is reactively loaded in a dorsal direction by said at least one support member so as to lock a midtarsal joint of said foot. 2. The apparatus of claim 1, further comprising: an imaging device that captures an image of at least one subject area of said foot and/or said associated ankle and lower leg so as to determine contours of said subject area with said midtarsal joint of said foot locked. 3. The apparatus of claim 2, wherein said imaging device comprises: a handheld imaging device that is freely moveable by an operator to capture images of subject areas over a bottom and top of said foot and said associated ankle and lower leg of said patient. 4. The apparatus of claim 1, wherein said at least one metatarsal head support member comprises: at least one support member that engages said plantar surface of said foot substantially only beneath an area of said fourth and fifth metatarsal heads of said foot. 5. The apparatus of claim 4, wherein said at least one metatarsal head support member comprises: a metatarsal head support member that engages said plantar surface of said foot substantially only beneath said fifth metatarsal head of said foot. 6. The apparatus of claim 1, wherein said mechanism that provides movement of said lateral forefoot area of said foot relative to said at least one metatarsal head support member comprises: a foot engagement member that supports a rearfoot portion of said foot and a distal aspect of said associated leg and that is mounted for selective movement relative to said at least one metatarsal head support member. 7. The apparatus of claim 6, wherein said foot engagement member comprises: a cradle that engages a rearfoot portion of said patient's foot and that is translatable towards and away from said at least metatarsal head support member. 8. The apparatus of claim 1, wherein said mechanism that provides movement of said lateral forefoot area of said foot relative to said at least one metatarsal head support member comprises: a base on said support and alignment assembly that is adapted for manual movement of said assembly towards and away from said patient. 9. The apparatus of claim 6, wherein said support and alignment assembly comprises: a collapsible frame having an erected configuration for use and a collapsed configuration for transportation or storage. 10. The apparatus of claim 9, wherein said collapsib... | Northwest Podiatric Laboratory, Inc. | G01B011/24 A43D001/02 A43B017/00 A43B007/14 A61B005/107 A61B005/00 A61D001/02 A61B005/103
US09814637 | B2 | 2017-11-14 | Patient monitoring systems and methods | 2015-06-09 | Various implementations of a patient monitoring assembly may be useful in preventing pressure ulcers, monitoring of air cushion inflation, estimating physical activity and energy expenditure and managing body weight of a patient seated on the assembly. The patient monitoring assembly may be installed on a manual or powered wheelchair, for example, and continuously monitors and provides feedback to the patient and/or caregiver about whether the patient is engaging in pressure relief activities while seated on a cushion of the assembly, whether the patient's activities are in accordance with clinical guidelines for pressure relief, the amount of calories expended over a period of time and the weight and weight distribution of the patient seated on the assembly. The feedback may be provided in substantially real time via an interface on the wheelchair and/or via a remotely located computing device. | 1. A patient activity monitoring system comprising: a rigid support pad having generally planar first and second surfaces, the first and second surfaces being substantially parallel to and opposite each other;a cushion having an upper surface and a lower surface, the lower surface of the cushion being disposed on the first surface of the support pad, wherein the upper surface of the cushion is disposable below at least a portion of the patient;at least three force sensors disposed on the second surface of the rigid support pad;a tilt sensor disposed adjacent a central portion of the first or second surface of the rigid support pad; andan electronic control unit comprising a memory and a processor, the processor being in electrical communication with the one or more force sensors, the tilt sensor, and the memory, the processor executing instructions stored on the memory, the instructions comprising: electronically receiving a force signal from each force sensor, each force signal based on an amount of force exerted by the second surface of the rigid support pad on each force sensor;electronically receiving one or more tilt signals from the tilt sensor;comparing at least a portion of the received force signals and at least a portion of the received tilt signals with an expected range of force signals and tilt signals, respectively, associated with performance of a pressure relieving movement;identifying that the pressure relieving movement is being performed by the patient in response to the portion of received force signals and the portion of received tilt signals being within the expected range of force signals and the expected range of tilt signals, respectively, associated with the performance of the pressure relieving movement; andin response to identifying that the patient performed the pressure relieving movement, storing the identification of the identified pressure relieving movement in the memory. 2. The system of claim 1, wherein the pressure relieving movement comprises one or more of the following: leaning forward, leaning right, leaning left, and tricep push ups. 3. The system of claim 1, wherein the instructions further include identifying whether the patient is disposed on the upper surface of the cushion based on at least a portion of the received force signals. 4. The system of claim 1, wherein the instructions further include measuring a duration of time during which the identified pressure relieving movement is performed and storing the measured duration with the identification of the identified pressure relieving movement in the memory. 5. The system of claim 1, wherein the instructions further include counting a number of times that the identified pressure relieving movement is performed and storing the counted number with the identification of the identified pressure relieving movement in the memory. 6. The system of claim 4, wherein the instructions further include: comparing the measured duration associated with the id... | The Board of Trustees of The University of Alabama | A61G005/10 G01L005/00 A61B005/00 A61B005/103 A61B005/11 A61G007/057
US09823127 | B2 | 2017-11-21 | Systems and methods for deep spectroscopic imaging of biological samples with use of an interferometer and spectrometer | 2015-04-09 | Disclosed herein are systems and methods for deep spectroscopic imaging of a biological sample. In an aspect, a system includes a broad bandwidth light source configured to generate an illumination beam, an interferometer, and a spectrometer. The interferometer includes a first beam splitter configured to split the illumination beam into an incident beam and a reference beam; an optical lens directs the incident beam onto a biological sample at a predefined offset from corresponding optical axis, and receive a beam scattered from the biological sample. The beams are configured to intersect with each other within a focal zone of the optical lens. Photons of the incident beam undergo multiple forward scattering within the biological sample. A second beam splitter configured to receive and superimpose the scattered and reference beams, to generate an interference beam. The spectrometer uses a spectral domain detection technique to assess tissue properties of the biological sample. | 1. A system for depth resolved spectroscopic imaging in deep tissue samples, the system comprising: a broad bandwidth light source configured to generate an illumination beam;an interferometer comprising: a first beam splitter configured to receive and to split the illumination beam into an incident beam and a reference beam;an optical element configured to: receive the incident beam that is offset in distance from an optical axis resulting in a beam converging towards a focal zone situated at a predefined distance beneath the surface of a biological sample, wherein the converging beam is centered at an offset angle relative to the corresponding central optical axis; andreceive light scattered from the biological sample centered about an offset angle substantially equal to that of the incident beam and offset in the opposite direction;a second beam splitter configured to: receive and superimpose the scattered light and reference beam to generate an interference beam such that the spatial extent of the reference beam limits the received scattered light to that which has originated from the focal zone of the optical element; anda spectrometer configured to receive the interference beam; anda processor that uses a spectral domain detection technique to assess one or more depth resolved tissue properties of the biological sample based on a spectral interference pattern of the interference beam. 2. The system as claimed in claim 1, wherein the at least one broad bandwidth light source is one of a supercontinuum light source, a tungsten lamp, a superluminescent diode, and a swept source laser. 3. The system as claimed in claim 1, wherein the spectrometer includes a diffraction grating coupled to a charge coupled device (CCD) camera. 4. The system as claimed in claim 1, wherein the biological sample is burned skin tissue, and the deep spectroscopic imaging of the burned skin tissue generates information about tissue composition and structure. 5. The system as claimed in claim 1, wherein the incident and scattered beams are configured to have distinct illumination and detection apertures respectively. 6. The system as claimed in claim 1, wherein the interferometer is configured to employ coherence gating to restrict path length of one or more detected photons, and to employ spatial gating to localize paths of one or more detected photons, to enable the spectrometer to achieve imaging up to 90 Mean free paths (MFPs) and up to one centimeter depth of the biological sample. 7. The system as claimed in claim 6, wherein the incident and scattered beams are delivered and collected via one or more optical fibers. 8. The system as claimed in claim 1, wherein the data acquired from the spectrometer includes acquisitions of less than 10 milliseconds which are averaged together over a period of one or more seconds to one or more minutes to produce high dynamic range. 9. The system as claimed in claim 1, wherein the spectrometer is further configured to: gen... | Duke University | G01J003/28 G01B009/02 G01J003/453 G01N021/45 G01J003/02 A61B005/00 G01N021/47
US09778395 | B2 | 2017-10-03 | Security scanning device | 2017-01-05 | A method for scanning an individual with a security device includes providing a scanner, engaging the scanner in a first scanning mode to produce an initial scan of the individual with a controller, producing an initial representation based on the initial scan, determining a property of a second scanning mode based on information from the initial representation, and engaging the scanner in the second scanning mode to produce a second scan of the individual. | 1. A method, comprising: engaging a scanner in a first scanning mode that includes a first scanning of a standing individual with a first scanning beam of electromagnetic waves;receiving a return of the first scanning beam of electromagnetic waves from the standing individual;producing a representation of the standing individual based on the returned first scanning beam of electromagnetic waves;determining a property of a second scanning mode based on the representation;engaging the scanner in the second scanning mode that includes a second scanning of the standing individual with a second scanning beam of electromagnetic waves; receiving a return of the second scanning beam of electromagnetic waves from the standing individual;determining a threat that the standing individual carries contraband based upon the returned second scanning beam of electromagnetic waves; andoutputting a human-perceivable notification of the determined threat. 2. The method of claim 1, further comprising determining a maximum allowable dosage for the individual based on a characteristic of the individual and a maximum dose exposure limit for the individual. 3. The method of claim 1, further comprising determining an allowable dosage with the controller based on a characteristic of the individual. 4. The method of claim 1, further comprising determining a maximum allowable dosage for the individual based on a characteristic of the individual. 5. The method of claim 1, wherein the property includes a dose that may be received by the individual. 6. The method of claim 5, wherein the dose that may be received by the individual during the second scanning mode is greater than the dose that may be received by the individual during the first scanning mode. 7. The method of claim 1, wherein the property includes a target area of the individual. 8. The method of claim 7, wherein the target area of the individual corresponds to a portion of the initial representation having an image quality that is at least one of less than other portions of the initial representation and below a threshold level. 9. The method of claim 7, wherein the target area of the individual corresponds to a portion of the initial representation potentially indicating a suspect object. 10. The method of claim 1, further comprising determining a characteristic of the individual with the controller using the initial scan. 11. The method of claim 10, wherein the characteristic includes at least one of age, gender, weight, a medical condition, and body structure. 12. The method of claim 10, wherein the characteristic includes a pregnancy status. 13. The method of claim 1, wherein the first scanning mode comprises a physiological scan. 14. The method of claim 13, further comprising determining at least one of a body temperature of the individual and a stress level of the individual with the physiological scan. 15. The method of claim 1, wherein the first scanning mode comprises at least one of an opt... | Elwha LLC | A61B005/00 A61B006/10 G01V011/00 G01N022/00 G01N029/06 G01V005/00 G01N023/04
US09788822 | B2 | 2017-10-17 | Surgical access system and related methods | 2016-01-13 | A surgical access system including a tissue distraction assembly and a tissue retraction assembly, both of which may be equipped with one or more electrodes for use in detecting the existence of (and optionally the distance and/or direction to) neural structures before, during, and after the establishment of an operative corridor to a surgical target site. | 1. A retractor blade for use in a system for accessing a surgical target site, the retractor blade comprising: an elongated retractor blade extending from a proximal end that is dimensioned to couple to a mounting structure to a distal end that is spaced longitudinally from the proximal end, wherein the retractor blade includes a first edge extending between the proximal end and the distal end, a second edge extending between the proximal end and the distal end, an inner surface defined by the retractor blade between the first and second edges and comprising a portion configured to be positioned within an operative corridor established within a patient for accessing the surgical target site and oriented towards a central longitudinal axis of the operative corridor, and an outer face defined by the retractor blade between the first and second edges and comprising a portion configured to be oriented away from the central longitudinal axis, wherein the inner face of the retractor blade defines a receiving portion dimensioned to couple to an engagement portion of a connectable element, and wherein the inner surface of the retractor blade defines a plurality of grooves spaced longitudinally along the portion of the inner surface of the retractor blade and wherein the plurality of grooves are dimensioned to receive a tooth of the connectable element so as to hold the connectable element in place with respect to the retractor blade, wherein the plurality of grooves is on a concave portion of the inner surface of the retractor blade. 2. The retractor blade of claim 1, wherein the grooves defined by the inner surface of the retractor blade comprise a series of ratchet-like grooves configured to longitudinally position the connectable element with respect to the retractor blade. 3. The retractor blade of claim 1, wherein the grooves are positioned in series along a longitudinal center line of the inner surface of the retractor blade. 4. The retractor blade of claim1, wherein the grooves extend in a direction that is substantially perpendicular to a longitudinal axis of the retractor blade. 5. The retractor blade of claim 1, wherein the inner surface is substantially concave and the outer surface is substantially convex. 6. The retractor blade of claim 1, wherein the grooves comprise at least four grooves configured to receive the tooth engaged with a corresponding one of the at least four grooves to position the connectable element with respect to the retractor blade. 7. The retractor blade of claim 1, wherein the retractor blade tapers along at least part of the retractor blade between the first and second edges. 8. The retractor blade of claim 1, further comprising a receiving portion configured to engage with a corresponding engagement element of a shim element or a retractor extender. | NuVasive, Inc. | A61B001/32 A61B017/02 A61B090/00 A61B005/0488 A61B005/0492 A61B005/05 A61B005/11 A61B005/00 A61N001/05 A61B001/06 A61N001/08 A61B017/00
US09788949 | B2 | 2017-10-17 | Methods of dynamic annuloplasty ring sizing | 2016-05-16 | A dynamic, adjustable annuloplasty ring sizer can include an adjustable ring replica, which can be adjusted through a range of sizes corresponding to available prosthetic annuloplasty repair ring sizes. Actuation of an adjustment trigger on a handle portion of the ring sizer can displace tension wires that extend through a malleable shaft and through a plurality of articulating segments that form the ring replica. Displacement of the tension wires causes flexion of the joints between adjacent articulating segments, thereby reducing the overall size of the ring replica. Releasing the tension wires can allow an elastic extension wire to act on the ring replica, enlarging the ring replica to its maximum, at-rest size. In this manner, the appropriate size of annuloplasty ring prosthesis can be determined with a single device, without requiring a plurality of static ring sizers that require individual insertion and placement for the conventional trial-and-error sizing methods. | 1. A method for sizing a heart valve annulus, comprising: preparing a sizing device including a proximal handle having an actuator, a shaft coupled to and extending in a distal direction from the handle, a generally C-shaped annuloplasty ring replica having a first free end and a second free end and coupled to a distal end of the shaft configured to be adjusted to a plurality of different ring sizes, wherein the annuloplasty ring replica is elongated and defines a shape of the annulus being sized and is biased to a maximum ring size, wherein the shaft has a hollow bore, and wherein the sizing device further includes at least one flexible tension member coupled to the actuator in the handle that passes through the hollow bore that extends along the shaft and then through aligned channels in a plurality of articulating segments defining an arc of the annuloplasty ring replica, wherein tension on the flexible tension member pivots the articulating segments with respect to one another to reduce the ring size of the annuloplasty ring replica to a plurality of different ring sizes smaller than the maximum ring size;reducing the size of the adjustable ring replica to less than the maximum ring size using the actuator;positioning the adjustable ring replica at or near a patient's native mitral valve or tricuspid valve annulus;increasing the ring size of the adjustable ring replica using the actuator until an outer surface contacts the valve annulus; anddetermining an appropriate ring size of prosthetic annuloplasty ring for the valve annulus. 2. The method according to claim 1, wherein the annuloplasty ring replica comprises a backbone segment coupled to the shaft including a first backbone end and a second backbone end opposite the first backbone end, wherein a first plurality of the articulating segments serially coupled to one another extend in an elongated manner from the first backbone end to the first free end of the annuloplasty ring replica, and wherein a second plurality of the articulating segments serially coupled to one another extend in an elongated manner from the second backbone end to the second free end of the annuloplasty ring replica. 3. The method according to claim 1, wherein the annuloplasty ring replica further includes a plurality of aligned extension wire channels in a plurality of articulating segments and a flexible extension wire extending therethrough that biases the annuloplasty ring replica to the maximum ring size, wherein tension on the flexible tension wire reduces the ring size of the annuloplasty ring replica against the bias of the flexible extension wire. 4. The method according to claim 1, wherein the shaft is malleable and couples to the annuloplasty ring replica with a swivel joint, and the step of positioning includes bending the handle and/or re-orienting the annuloplasty ring replica relative to the handle. 5. The method according to claim 1, further comprising a display on the handle configured to provide ... | Edwards Lifesciences Corporation | A61B005/00 A61F002/24 A61B005/107 A61B090/00
US09827449 | B2 | 2017-11-28 | Systems for treating skin laxity | 2017-06-16 | A system and method for ultrasound treatment of skin laxity are provided. Systems and methods can include ultrasound imaging of the region of interest for localization of the treatment area, delivering ultrasound energy at a depth and pattern to achieve the desired therapeutic effects, and/or monitoring the treatment area to assess the results and/or provide feedback. In an embodiment, a treatment system and method can be configured for producing arrays of sub-millimeter and larger zones of thermal ablation to treat the epidermal, superficial dermal, mid-dermal or deep dermal components of tissue. | 1. An ultrasound device for treatment of skin laxity, the device comprising: an ultrasound probe comprising a housing, wherein inside the housing comprises a motion mechanism and an ultrasound transducer,wherein the ultrasound transducer comprises a therapy component,wherein the therapy component comprises: a single active focused piezoelectric ultrasound therapy element,wherein the piezoelectric ultrasound therapy element is configured to provide a single mechanical focus,wherein the single mechanical focus is configured to provide ultrasound therapy energy in a form of a single thermal focus in a tissue at a depth below a skin surface,wherein the depth is a single, fixed depth below the skin surface to treat the tissue,wherein the tissue comprises a combination of any of the group consisting of: an epidermal tissue, a superficial dermal tissue, a mid-dermal tissue, a deep dermal tissue, a muscle tissue, and an adipose tissue, wherein the single thermal focus is formed without electronic focusing and without a lens,wherein a portion of the ultrasound probe is configured for acoustic coupling to the skin surface;wherein the piezoelectric ultrasound therapy element is configured for delivery of energy at a temperature sufficient to tighten at least a portion of the tissue at the depth under the skin surface,wherein the piezoelectric ultrasound therapy element is connected to the motion mechanism,wherein the motion mechanism comprises an encoder,wherein the motion mechanism moves the piezoelectric ultrasound therapy element to form a plurality of thermal lesions at the depth for tighten at least a portion of the tissue for reducing an appearance of skin laxity. 2. The device of claim 1, wherein the piezoelectric ultrasound therapy element is configured to increase the temperature of the tissue in the region of interest to greater than 60° C.,wherein the piezoelectric ultrasound therapy element is a single element that delivers ultrasound energy at a frequency of between 2 MHz to 25 MHz. 3. The device of claim 1, further comprising a control system, an input device, and a power supply, wherein the control system is connected to the motion mechanism and the piezoelectric ultrasound therapy element;wherein the input device is connected to the control system; andwherein the power supply is connected to the control system. 4. The device of claim 3, wherein the control system comprises a spatial control and a temporal control, wherein the spatial control and the temporal control are configured for controlling the delivery of energy at a temperature sufficient to cause denaturation of at least the portion of the tissue at the depth under the skin surface,wherein the spatial control and the temporal control are configured for controlling the delivery of energy at a frequency of between 2 MHz to 25 MHz. 5. The device of claim 1, wherein the housing contains a monitoring sensor and a control system, wherein the control system comprises a processor, softw... | Guided Therapy Systems, L.L.C. | A61N007/02 A61B005/00 A61B008/08 A61B008/12 A61H023/02 A61N007/00 G01S015/89 A61B008/13 A61B017/32 A61B008/00
US09787568 | B2 | 2017-10-10 | Physiological test credit method | 2013-11-04 | A physiological test credit method determines if test credits are available to the monitor and checks if a Wi-Fi connection is available. If test credits are less than a test credit threshold, the monitor connects to a test credit server, processes server commands so as to download test credits and disconnects from the server. In various embodiments, the monitor is challenged to break a server code, the server is challenged to break a monitor code. The server validates monitor serial codes, and saves monitor configuration parameters. | 1. A physiological test credit method programmatically initiates wireless communications between a physiological monitor and a remote server in response to available test credits falling below a predetermined threshold in response to a physiological test so as to download additional test credits from the server to the monitor and therefore enable the monitor to perform additional physiological parameter spot-checks on a per test credit basis, the physiological test credit method comprises: establishing a threshold for test credits stored in an optical sensor attached to a corresponding physiological monitor;securely connecting a server to the monitor when remaining test credits fall below the threshold;downloading test credits from the server to the monitor, andsending from the monitor to the server an acknowledgement for receipt of the test credits, wherein upon the server receiving the acknowledgement, the method further comprises the following steps using a two-way handshake: deducting the test credit from a user account at the server;incrementing a test credit downloading index count at the server, andupdating credits from the monitor to a sensor configured for performing the physiological test. 2. The physiological test credit method according to claim 1 further comprising: defining a quantum of test credits; anddownloading an amount of test credits equal to the quantum each time the remaining test credits fall below the threshold. 3. The physiological test credit method according to claim 2 further comprising: defining a download frequency according to the number of times remaining test credits fall below the threshold in a given period of time; andadjusting the threshold according to the download frequency. 4. The physiological test credit method according to claim 3 further comprising adjusting the quantum according to the download frequency. 5. The physiological test credit method according claim 3 further comprising adjusting the threshold and the quantum according to the download frequency. 6. The physiological test credit method according to claim 1, wherein securely connecting a server to the monitor when remaining test credits fall below the threshold further comprises exchanging challenges between the server and the monitor so as to verify both the server and the monitor. 7. The physiological test credit method according to claim 6 wherein exchanging challenges comprising: sending a server challenge code to the monitor;breaking the server challenge code with the monitor;sending a monitor challenge code to the server; andbreaking the monitor challenge code with the server. 8. The physiological test credit method according to claim 7 further comprising validating the monitor via sales data. 9. The physiological test credit method according to claim 8 wherein validating the monitor comprises: sending a server request for device serial numbers to the monitor; andmatching the serial numbers to sales data. 10. The physiological ... | Cercacor Laboratories, Inc. | G06Q020/08 H04L012/26 G06Q020/14 A61B005/00 G06F019/00 G06Q050/22 A61B005/1455
US09782088 | B2 | 2017-10-10 | Pulse measuring device and method of pulse measurement therewith | 2013-07-02 | A pulse measuring device according to one embodiment of the present invention comprises: a measuring unit which is capable of spherical movement with multiple degrees of freedom; and a movement unit on which the measuring unit is mounted, and which has a spherical movement part for moving the measuring unit spherically with multiple degrees of freedom while also having a linear movement structure for moving the measuring unit linearly with multiple degrees of freedom. The measuring unit comprises: a unit body which is coupled so as to be able to rotate relative to an end part of the movement unit; a linking part which is coupled to a lower end part of the unit body and is able to rotate relative to the unit body; and a sensor part which is equipped with a pulse diagnosis sensor. | 1. A pulse measuring device comprising: a measuring unit configured to move spherically with multiple degrees of freedom; anda movement unit on which the measuring unit is provided, the movement unit comprising a spherical movement portion configured to spherically move the measuring unit with multiple degrees of freedom while having a linear movement structure to linearly move the measuring unit with multiple degrees of freedom,wherein the measuring unit comprises: a unit body rotatably connected to one end portion of the movement unit;a connecting portion connected to a lower end portion of the unit body, and configured to rotate with respect to the unit body; anda sensor portion rotatably connected to the connecting portion, the sensor portion comprising a pulse diagnosis sensor,wherein, when the pulse diagnosis sensor is adapted to be in contact with a pulse position, a position of the measuring unit is adjusted to enable the measuring unit to pressurize the pulse position in a vertical direction by spherically moving the measuring unit with multiple degrees of freedom using the spherical movement portion, andwherein the unit body comprises: a body member;a pair of parallel members disposed to be parallel from the body member toward a lower side; anda pair of extending members configured to extend from the body member toward both sides,wherein the connecting portion is rotatably connected to the pair of parallel members. 2. The pulse measuring device of claim 1, wherein the measuring unit is configured to move spherically with three degrees of freedom, and the movement unit is configured to move linearly with three degrees of freedom. 3. The pulse measuring device of claim 1, wherein a pair of light emitting portions are disposed at the pair of extending members to emit lights toward a skin, respectively, and an initial position of the measuring unit is set by matching the pulse position and the pair of lights emitted from the pair of light emitting portions. 4. The pulse measuring device of claim 1, wherein the connecting portion comprises: a connecting bar to which the sensor portion is rotatably connected; anda pair of joint members connected to both sides of the connecting bar, respectively, and connected to the pair of parallel members, respectively, the pair of joint members comprising two joints configured to enable the pair of joint members to perform bending motions with respect to the pair of parallel members. 5. The pulse measuring device of claim 1, wherein the sensor portion comprises: a sensor housing tiltably connected to the connecting portion, the sensor housing on which the pulse diagnosis sensor is provided at a lower end portion; anda tilting sensor configured to measure a gradient of the pulse diagnosis sensor with respect to the pulse position. 6. The pulse measuring device of claim 5, wherein the sensor housing comprises: a first joint portion tiltably connected to the connecting portion; anda second joint portion ... | KOREA INSTITUTE OF ORIENTAL MEDICINE | A61B005/02 A61B005/024 A61B005/00
US09795329 | B2 | 2017-10-24 | Non-invasive device and method for measuring a substance concentration | 2015-01-12 | A system includes a light source and a light detector. The system further includes a first optical system configured to direct a transmitted portion of light onto the detector. The first optical system is further configured to block a scattered portion of the light. The system also includes a second optical system configured to direct the transmitted and scattered portions of the light onto the detector. | 1. A method comprising: providing a light source and directing light through a body that both transmits and scatters the light;via a first optical system, directing the transmitted light onto a light detector, the first optical system blocking the scattered light from the body with one or more field limiters such that the blocked light does not reach the light detector;measuring the transmitted light from the body in a wavelength band via the first optical system;via a second optical system, directing the transmitted light and a portion of the scattered light onto the light detector;measuring the portion of the scattered light from the body in the wavelength band via the second optical system; anddetermining a substance concentration in the body based on a predetermined correlation between measurement data and the substance concentration, the measurement data being associated with the measured transmitted light and the measured scattered light. 2. The method of claim 1, further comprising: measuring the transmitted light from the body in another wavelength band via the first optical system;measuring the portion of the scattered light from the body in the other wavelength band via the second optical system, the measurement data being further associated with the measured scattered light and the measured transmitted light in the other wavelength band. 3. The method of claim 2, the wavelength band including wavelengths where the substance has an effect on absorption of transmitted light and the other wavelength band including only wavelengths where the substance has no effect or negligible effect on absorption of transmitted light. 4. The method of claim 3, the light source including a near-infrared (NIR) laser diode and a collimating lens generating a collimated NIR light beam in the wavelength band when selected and in the other wavelength band when selected. 5. The method of claim 3, the light source including a first NIR laser diode and a first collimating lens generating a collimated NIR light beam in the wavelength band and including a second NIR laser diode and a second collimating lens generating a collimated NIR light beam in the other wavelength band. 6. The method of claim 3, further comprising an interference filter passing light within the wavelength band and blocking light within the other wavelength band. 7. The method of claim 6, further comprising a motor selectively positioning the interference filter between the light source and the light detector. 8. The method of claim 2, further comprising: calculating a scattered light ratio parameter based on a ratio of the measured scattered light in the wavelength band to the measured scattered light in the other wavelength band;calculating a transmitted light ratio parameter based on a ratio of the measured transmitted light in the wavelength band to the measured transmitted light in the other wavelength band, the measurement data including the scattered light ratio parameter and the... | GlucoVista Inc. | A61B005/1455 A61B005/00 A61B005/01 A61B005/145 G01N021/359 G01N021/49
US09829309 | B2 | 2017-11-28 | Depth sensing method, device and system based on symbols array plane structured light | 2015-07-23 | The present invention discloses a depth sensing method, device and system based on symbols array plane structured light. The coded symbols array pattern is projected by the laser pattern projector to the target object or space, and the image sensor collects and obtains the successive sequence of the encoded image of the input symbols. Firstly, the input image encoded with the symbols is decoded. The decoding process includes preprocessing, symbols location, symbols recognition and symbols correction. Secondly, the disparity of the decoded symbols are calculated by the symbols match calculation between the decoded image of the input symbols with completed symbols recognition and the decoded image of the reference symbols with the known distance. Finally the depth calculation formula is combined to generate depth point cloud information of the target object or projection space that is represented in the form of grid. The present invention can quickly obtain high-resolution and high-precision depth information of the target object or projection space in dynamic scenes, which facilitates the porting or insertion as a module into the intelligent device for 3D modeling, 3D mapping, etc. | 1. A depth sensing method of symbols array plane structured light, comprising: obtaining the encoded image of the input symbols of the symbols array plane structured light;achieving adaptive preprocessing of the encoded image of the input symbols;performing symbols location of the preprocessed encoded image of the input symbols to identify the location of the center of all symbols in the coded image of the input symbols;conducting symbols recognition of the preprocessed encoded image of the input symbols, mapped as the corresponding number;correcting identified symbols in accordance with the corresponding number to complete the decoding of the encoded image of the input symbols;making symbols match between the corrected decoded image of the input symbols and the decoded image of the reference symbols of the known distance to generate the corresponding disparity of the symbols;determining depth of the center of the symbols in the encoded image of the input symbols in accordance with the known distance of the decoded image of the reference symbols and the disparity,wherein making the symbols match between the corrected decoded of the input symbols and the decoded image of the reference symbols of the known distance to generate the corresponding disparity of the symbols, and determining the depth of the center of symbols in the encoded image of the input symbols in accordance with the said known distance and the disparity, comprise: the input symbols and their surrounding symbols forming an input symbols template, searching for an optimal match template in a search window of the decoded image of the reference symbols and obtaining the match symbols of the input symbols, so as to obtain the disparity (Δx, Δy) between the input symbols and the match symbols; andusing the disparity Δm (Δx or Δy) in X or Y direction, in combination with the known distance parameter d of the decoded image of the reference symbols, the baseline distance S between the laser pattern projector and the image sensor, the focal length f and the dot pitch parameter μ of the image sensor, calculating a depth information d′ of a center o of the input symbols according to the following formula:d′=d-Δ⁢⁢m⁢⁢μ⁢⁢d2fS+Δ⁢⁢m⁢⁢μ⁢⁢d=f⁢⁢S⁢⁢dfS+Δ⁢⁢m⁢⁢μ⁢⁢d. 2. The method according to claim 1, wherein performing the symbols location of the said preprocessed encoded image of the input symbols to identify the location of the center of all symbols in the said coded image of the input symbols, comprises: computing the location of the center of symbols one by one in accordance with the design parameters or determining that in accordance with the feature information of the symbols and by means of pattern identification. 3. The method according to claim 1, wherein the symbols recognition contains a template match method that takes the code symbols as the standard match symbols template rotating at a small angle around the center of the symbols or identifying the code symbols of the encoded image of t... | RGBDsense Information Technology Ltd. | G01B011/00 G01B011/22 G01B011/25 G01B011/26 G06T007/50 G06T007/521 G06T007/593 G06K009/20 G01S017/06 H04N013/02 H04N019/597 G06K009/62 A61B005/00 A61B005/107
US09785235 | B2 | 2017-10-10 | Display control apparatus, display control method of display control apparatus, and eye gaze direction detection system | 2014-02-19 | It is determined whether a driver is gazing at a gaze target object in the eye gaze direction, and the eye gaze of the driver is guided by changing a display mode of the gaze target object determined to be gazed at by the driver. Calibration of eye gaze direction detection information is accurately and promptly executed. | 1. A display control apparatus to use a display object displayed by a display unit as a gaze target object, and to control calibration of an eye gaze direction detection apparatus based on the gaze target object, the display control apparatus comprising: an eye gaze state determinator to determine, based on eye gaze direction detection information detected by the eye gaze direction detection apparatus and display information of the gaze target object, whether a subject is gazing at the gaze target object existing in an eye gaze direction of the subject;an eye gaze guider to output display information of the gaze target object and screen control information relating to a display mode of the gaze target object, and to guide an eye gaze of the subject by changing the display mode of the gaze target object when the subject is determined to be gazing at the gaze target object;a display controller to output display control information for controlling the display unit, based on the screen control information output by the eye gaze guider; anda calibration controller to output a calibration control signal toward the eye gaze direction detection apparatus when the subject is determined to be gazing at the gaze target object,wherein the gaze target object is an icon or a button displayed by an intelligent panel functioning as the display unit, or by the display unit of a car navigation apparatus, or a marking displayed by a head-up display as the display unit. 2. The display control apparatus according to claim 1, wherein the eye gaze guider scales down the gaze target object determined to be gazed at by the subject, toward a reference point. 3. The display control apparatus according to claim 1, wherein the eye gaze guider arranges a first gaze target object determined to be gazed at by the subject, and a second gaze target object to be gazed at subsequently to the first gaze target object, at different positions on the display unit. 4. The display control apparatus according to claim 3, wherein the second gaze target object is a determination key to determine a manipulation being selected by using the first gaze target object. 5. The display control apparatus according to claim 3, wherein the eye gaze guider arranges the first and the second gaze target objects at diagonal positions on a display screen of the display unit. 6. The display control apparatus according to claim 3, wherein, when a screen transitions from a first display screen displaying the first gaze target object to a second display screen displaying the second gaze target object, the eye gaze guider respectively arranges the first and the second gaze target objects on the first and the second display screens, at positions diagonal to each other. 7. The display control apparatus according to claim 1, wherein the eye gaze guider moves the gaze target object determined to be gazed at by the subject, in a predefined direction on a display screen of the display unit. 8. The display cont... | Mitsubishi Electric Corporation | G09G005/00 G06F003/01 A61B003/113 A61B003/00 A61B005/18 G06K009/00 G06T007/70 G06F009/44
US09797968 | B2 | 2017-10-24 | System and method for magnetic resonance image acquisition | 2014-09-15 | A method of data acquisition at a magnetic resonance imaging (MRI) system is provided. The system receives at least a portion of raw data for an image, and detects anomalies in the portion of raw data received. When anomalies are detected, the system can correct those anomalies dynamically, without waiting for a new scan to be ordered. The system can attempt to scan the offending portion of the raw data, either upon detection of the anomaly or at some point during the scan. The system can also correct anomalies using digital correction methods based on expected values. The anomalies can be detected based on variations from thresholds, masks and expected values all of which can be obtained using one of the ongoing scan, previously performed scans and apriori information relating to the type of scan being performed. | 1. A method of acquiring an image at a magnetic resonance imaging (MRI) system comprising: receiving an analog signal based on a pulse sequence and a first gain;digitizing, by a single analog to digital converter (ADC), the analog signal into a digitized signal;detecting a potential quantization error in the digitized signal based on a boundary;when the detecting is affirmative: receiving a replacement analog signal based on the pulse sequence, at least one portion of the replacement analog signal being based on an adjusted variable gain, the adjusted gain being a factor of the first gain;digitizing, by the ADC, the replacement analog signal into a replacement digitized signal; andadjusting at least one portion of the replacement digitized signal corresponding to the at least one portion of the replacement analog signal, based on a reversal of the factor. 2. The method of claim 1 wherein the reversal of the factor includes dividing, by the factor, values of the replacement digitized signal corresponding to the at least one portion of the replacement analog signal. 3. The method of claim 1 wherein the factor is based on an operational amplitude range of the ADC. 4. The method of claim 1 wherein the at least one portion of the replacement analog signal includes two portions, each portion being based on a different adjusted gain, each different adjusted gain being a different factor of the first gain. 5. The method of claim 1 wherein the boundary includes upper boundary values and the detection is based on whether one or more values of the digitized signal received exceed corresponding ones of the upper boundary values. 6. The method of claim 1 wherein the boundary includes lower boundary values and the detection is based on whether one or more values of the digitized signal received are below corresponding ones of the lower boundary values. 7. The method of claim 5 wherein the upper boundary is based on an upper boundary of the operational amplitude range of the ADC. 8. The method of claim 6 wherein the lower boundary is based on an upper boundary of the operational amplitude range of the ADC. 9. The method of claim 1 wherein adjusting further comprises: adjusting a phase of the replacement digitized signal based on the adjusted gain. 10. The method of any claim 1 wherein the digitized signal and the replacement digitized signal correspond to a portion of a k-space for the image. 11. The method of claim 1 wherein the digitized signal and the replacement digitized signal are at least one line of data along the frequency axis of the k-space. 12. The method of claim 1 wherein the receiving a replacement analog signal is performed at one of: prior to receiving any additional analog signal, and after all of the analog signal for the image is received. | SYNAPTIVE MEDICAL (BARBADOS) INC. | G01V003/00 G01R033/46 G01R033/54 A61B005/055 H03M001/18
US09787750 | B2 | 2017-10-10 | Universal interface system for MRI applications | 2013-12-27 | An interface system for providing a signal interface between devices in a magnetic resonance imaging (MRI) magnet room and devices in the outside environment to the magnet room. An exemplary embodiment includes a signal transmission link between the MRI magnet room and the outside environment, the link including a wireless local area network. A controller system is located in the outside environment and connected to the devices in the outside environment. A transducer system is located in the MRI magnet room and connected to the devices in the MRI magnet room. | 1. A universal interface system for interfacing between the outside environment and an MRI (magnetic resonance imaging) magnet room, configured to avow communication between the outside environment and different devices in the magnet room, which devices employ different signal interfaces or signal protocols from others of the devices, the interface system comprising: a high speed digital data signal transmission link between the MRI magnet room and a control room outside the MRI magnet room;a signal transducer system module disposed within the magnet room, the signal transducer system module configured for communication with the digital data signal transmission link, and comprising a plurality of transducer connectors configured to connect to corresponding device connectors of a set of the different devices which employ different signal interfaces or signal protocols in the magnet room;said transducer system module includes a transducer housing mounting said transducer connectors on a housing surface;the signal transducer system module configured to translate digital signals received through the digital data signal transmission link into signals delivered through at least some of said plurality of connectors of corresponding different formats or protocols for corresponding ones of said different devices in the magnet room, the signal transducer system module further configured to translate signals received through at least some of the transducer connectors from one or more of the corresponding different devices in the magnet room into digital signals for transmission over the digital data transmission link;a controller system module disposed in the control in signal communication with the high speed digital data transmission link, the controller system module having a plurality of module connectors configured to connect to control room device connectors associated with a plurality of different external devices in the control room, the plurality of control room different devices arranged to deliver video, audio, status or control signals for utilization by one or more of the magnet room devices and which employ different signal interfaces or signal protocols, said plurality of different external devices including a control room computer for controlling patient stimulus and recording patient responses, and said plurality of module connectors including a computer connector for signal connection to the control room computer;said controller system module including a controller housing mounting said controller connectors on a controller housing surface;the controller system module configured to translate signals received through at least some of the controller connectors from one or more of the corresponding control room external devices into digital signals for transmission over the digital data transmission link to said transducer system, the control system module comprising an audio codec for converting digital and analog signals received at one or ... | - | H04L029/08 H04L012/28 A61B005/055 H05K009/00 A61B005/00 H04N013/04
US09826933 | B2 | 2017-11-28 | System for control of a prosthetic device | 2014-10-20 | A system for control of a prosthetic device includes at least one Inertial Measurement Unit detecting orientation of a user's foot. The at least one Inertial Measurement Unit is in communication with a device module configured to command at least one actuator of a prosthetic device. The at least one Inertial Measurement unit sends output signals related to orientation of the user's foot to the device module and the device module controls the at least one actuator of the prosthetic device based on the signals from the at least one Inertial Measurement Unit. | 1. A control system for a prosthetic device comprising: at least one sensor module comprising at least one inertial measurement unit adapted to detect user input and to generate at least one body input signal based on the detected user input, the at least one inertial measurement unit comprising at least one accelerometer;at least one device module including a central processing unit and having at least one prosthetic control mode, the at least one device module being in communication with the at least one sensor module; andat least one actuator of the prosthetic device configured to receive commands from the at least one device module,wherein the at least one device module receives the at least one body input signal from the at least one sensor module and commands the at least one actuator in accordance with the at least one body input signal,wherein the at least one sensor module sends body input signals to the at least one device module according to the accelerations detected by the at least one accelerometer and the at least one device module detects user walking based at least on the body input signals, andwherein the at least one device module stops sending commands to the at least one actuator of the prosthetic device when user walking is detected. 2. The control system according to claim 1 further comprising wherein the at least one device module commands the at least one actuator in accordance with the at least one body input signal and a current control mode of the at least one prosthetic control mode. 3. The control system according to claim 1 further comprising wherein the inertial measurement unit comprises at least one accelerometer that is configured to detect accelerations in three perpendicular axes. 4. The control system according to claim 1 further comprising wherein the inertial measurement unit includes at least three accelerometers, each accelerometer adapted to detect acceleration in a single axis, the at least three accelerometers oriented with respect to one another to detect accelerations in three perpendicular axes. 5. The control system according to claim 1 further comprising wherein the at least one prosthetic control mode includes a bulk mode controlling movement of an end point of the prosthetic device within a movement envelope. 6. The control system according to claim 5 further comprising wherein the at least one device module commands movement of the end point of the prosthetic device within the movement envelope in accordance with the at least one body input signal, the at least one body input signal being indicative of a desired directional movement of the end point of the prosthetic device. 7. The control system according to claim 6 further comprising wherein the at least one device module commands movement of the end point of the prosthetic device along a boundary of the movement envelope when the desired directional movement of the end point of the prosthetic device required by the at least one body in... | DEKA PRODUCTS LIMITED PARTNERSHIP | A61F002/70 A61F002/68 A61B005/00 A61F002/72 A61F002/76 G01C021/16 A61B005/11 G01C009/00 G05B015/02 A61B005/0488 A61B005/103 A61F002/54 A61F002/60
US09820675 | B2 | 2017-11-21 | Mask for analyzed mammals | 2014-09-05 | An encapsulatable life support mechanism (ELSM) for an analyzed animal, including: a cradle or bed adapted by means of size and shape to accommodate the animal; an anesthetization gas mask (AGM) characterized by a cup with conic cross section, comprising a plurality of apertures located at the outer circumference of the cup; a fluid supplying mechanism (FSM) in which the AGM is placed, the FSM is in a continuous fluid communication with (i) an anesthetization gas inlet positioned outside the ELSM and an outlet located within the ELSM; (ii) an air suction scavenging device positioned outside the ELSM and a mask and an air suction outlet located within the ELSM; and a plurality of (iii) air conditioning tubes; and an airtight shell enveloping the same. The airtight ELSM prevent leakage of anesthetization gas. | 1. A mouse handling system (MHS) for anesthetizing and accurately positioning a laboratory animal inside a magnetic resonance device (MRD) while preventing leakage of anesthetization gas and/or accidental spillage of hazardous materials to ambient atmosphere, the MHS comprising: a distal portion that is insertable into said MRD, the distal portion comprises: a) a cradle to accommodate said animal,b) an anesthetization gas mask that can be adjustably placed relative to animal,c) a gas inlet in fluid communication with the anesthetization gas mask and capable of being in fluid communication with a gas supply,d) a gas outlet in fluid communication with the anesthetization gas mask and capable of being in fluid communication with an air suction outlet, ande) an air tight shell to encapsulate the cradle, the anesthetization gas mask, the gas inlet and the gas outlet; anda proximal portion positioned outside said MRD when the distal portion is within said MRD and a shaft, the proximal portion incrementally rotatable about and slidable along said shaft to translate and rotate the laboratory animal when the laboratory animal is positioned within said distal end and said distal end is within said MRD. 2. The system of claim 1, wherein said shaft comprises indicia indicating the relative angular and translational position of said distal portion with respect to said MRD. 3. The system of claim 1, further comprising a plurality of levers and handles for locking said angular and translational position. 4. The system of claim 1, wherein said cradle comprises a heating/cooling mechanism. 5. The system of claim 1, further comprising a mask maneuvering mechanism comprising an adjustment rod in connection with said mask, said adjustment rod adapted to ensure that said mask is located in a predefined location within said housing. | ASPECT IMAGING LTD. | A61B006/04 A61D007/04 A61B005/055 A61D003/00
US09789295 | B2 | 2017-10-17 | Customized skin care and method to provide same | 2016-06-26 | Methods to provide customized skin care by using specimen dispensing device to dispense specimens from removable dispensers for the purpose of treating skin of a user are presented. Methods to utilize the embedded memory and electrical interface of the dispensing device and dispensers to produce customizable skin care products that give better skin treatment results are also presented. The invention may also be applied to health care and personal care needs. | 1. A method to provide customized skin care by using a specimen dispensing device to dispense specimens for treating a skin area of a user comprising the steps of: a first computing device creating user data during a skin analysis of said user, whereas said user data contains a first identification (ID) of a first skin feature with a first grade, and a second ID of a second skin feature with a second grade;said first computing device transferring said user data to a first information storage component included in a control unit enclosed in said specimen dispensing device;said user installing a first dispenser containing a first specimen and a second dispenser containing a second specimen into said specimen dispensing device;wherein said first dispenser comprising a second information storage component storing a first specimen data including said first ID and a first method of dispensing, and said second dispenser comprising said second information storage component storing a second specimen data including said second ID and a second method of dispensing;said control unit acquiring said first and said second specimen data from said second information storage components of said first and second dispensers through a first data connection between said control unit and said first and second dispensers;an information processing component included in said control unit calculating a first dispensing recipe with using said first grade, said first method of dispensing, said second grade, and said second method of dispensing;said control unit commanding the dispensing of said first and second specimens from said first and said second dispensers according to said first dispensing recipe. 2. The method of claim 1, wherein said first computing device transfers said user data to a first database included in a first server through a first data network connecting said first computing device and said first server during said skin analysis. 3. The method of claim 2, wherein said first dispensing recipe is stored in said first information storage component and a second computing device external to said specimen dispensing device acquires said first dispensing recipe from said first information storage component through a second data connection to said control unit; and wherein said second computing device includes one of: a computer, a mobile phone, and a tablet. 4. The method of claim 3, wherein said second computing device modifies said first dispensing recipe into a second dispensing recipe; wherein said second computing device transfers said second dispensing recipe to said first information storage component through said second data connection; and wherein said control unit commands dispensing of said first and second specimens from said first and second dispensers according to said second dispensing recipe. 5. The method of claim 4, wherein said second computing device transfers said first and second dispensing recipes to second database included in a seco... | - | A61M035/00 A61H023/00 A61H023/02 A61B005/00 A61H007/00 A61B005/103 A61M037/00 A61N007/00 G06F019/00 G06Q050/00 A61N001/00 A61N001/04 A61N005/06 A61N001/32 A61F007/00 A61H009/00 G01S019/19
US09814467 | B2 | 2017-11-14 | Personalized tourniquet apparatus | 2015-04-25 | Disclosed is a method and apparatus for regulating tourniquet cuff pressure based on a personalized tourniquet pressure (PTP) to facilitate safe performance of a surgical procedure. The apparatus includes a dual-purpose tourniquet cuff adapted to encircle a region of a patient limb to provide both a sensor during a pre-surgical time period and an effector during a surgical time period. An effector module communicates pneumatically with an inflatable bladder of the cuff for maintaining pressure in the bladder near a PTP during the surgical time period to safely stop penetration of arterial blood past the cuff. A pulsation sensor that communicates pneumatically with the inflatable bladder during the pre-surgical time period senses and characterizes pressure pulsations that are indicative of penetration of arterial blood into the region of the limb encircled by the cuff. A PTP estimator is responsive to the pulsation sensor for producing an estimate of the PTP, such that the estimate of the PTP is a function of the sensed pulsations. | 1. An apparatus for regulating tourniquet cuff pressure based on a personalized tourniquet pressure (PTP), comprising: a sensor module for estimating a limb occlusion pressure (LOP) by determining an estimated LOP equal to a minimum pressure at which penetration of blood past an applied tourniquet cuff is stopped, and measuring a pulsation characteristic associated with the LOP,a controller for establishing a PTP by determining a second pressure that is greater than the LOP and that corresponds to a second pulsation characteristic that differs by a predetermined percentage from the pulsation characteristic associated with the LOP; andan effector module operable during a surgical time period for maintaining pressure in an applied tourniquet cuff near the PTP. 2. The apparatus of claim 1 wherein the pulsation characteristic is indicative of a shape of the pulsation. 3. The apparatus of claim 1 wherein the pulsation characteristic is a maximum amplitude of pulsation. 4. The apparatus of claim 1 wherein the predetermined percentage is a function of a magnitude of the estimated LOP. 5. The apparatus of claim 1 wherein the sensor module is operable also to sense a characteristic of blood that penetrates past the applied tourniquet cuff. 6. The apparatus of claim 1 wherein the sensor module is also operable to sense in an inflatable bladder of the cuff a characteristic of pulsation. 7. An apparatus for regulating tourniquet cuff pressure based on a personalized tourniquet pressure (PTP), comprising a controller operable during a pre-surgical time period for: estimating a limb occlusion pressure (LOP) by determining a minimum pressure at which penetration of blood past an applied tourniquet cuff is stopped;measuring a pulsation characteristic associated with the LOP;establishing a PTP by determining a second pressure that is greater than the LOP and that corresponds to a second pulsation characteristic that differs by a predetermined percentage from the pulsation characteristic associated with the LOP; andregulating the tourniquet cuff pressure during a surgical time period to maintain a surgical pulsation characteristic near the second pulsation characteristic. 8. The apparatus of claim 7 wherein the pulsation characteristic is indicative of a shape of the pulsation. 9. The apparatus of claim 7 wherein the pulsation characteristic is a maximum amplitude of pulsation. 10. The apparatus of claim 7 wherein the predetermined percentage is a function of a magnitude of the estimated LOP. 11. The apparatus of claim 7 wherein the controller for estimating the LOP includes sensing a characteristic of blood that penetrates past the applied tourniquet cuff. 12. The apparatus of claim 7 wherein the controller for estimating the LOP includes sensing in an inflatable bladder of the cuff a pulsation characteristic. 13. The apparatus of claim 7 further comprising an alarm condition detector for producing an alert if the surgical pulsation characteristic is ... | Western Clinical Engineering Ltd. | A61B017/00 A61B017/135 A61B017/132 A61B005/022 A61B090/92 A61B090/98 A61B090/00
US09791528 | B2 | 2017-10-17 | Magnetic resonance imaging system and magnetic resonance imaging method | 2013-11-08 | A magnetic resonance imaging (MRI) method includes defining a plurality of sub-volumes so that each of the sub-volumes includes a plurality of sequential slices of a plurality of slices that make up a volume of a subject, wherein the sub-volumes are divided into a plurality of groups so that any neighboring sub-volumes belong to different groups; applying radio-frequency (RF) pulses including a plurality of frequency components and a selection gradient to the subject to simultaneously excite a plurality of sub-volumes in each of the groups; performing three-dimensional (3D) encoding on each of the excited sub-volumes so that only some slices of the plurality of slices in each of the excited sub-volumes are encoded in a slice direction; acquiring magnetic resonance signals from the encoded sub-volumes; and reconstructing the acquired magnetic resonance signals into image data corresponding to each of the plurality of slices in each of the encoded sub-volumes. | 1. A magnetic resonance imaging (MRI) method comprising: dividing a volume of a subject into sub-volumes so that each of the sub-volumes comprises sequential slices, wherein the sub-volumes are assigned to groups so that adjacent sub-volumes belong to different groups;applying radio-frequency (RF) pulses comprising frequency components and a selection gradient to the subject to excite the sub-volumes in each of the groups, wherein the sub-volumes in each group are excited simultaneously;performing three-dimensional (3D) encoding on each the excited sub-volumes byselecting some slices for encoding in the slice direction from the sequential slices in each the excited sub-volume, andperforming slice direction encoding on the slices selected for encoding;acquiring magnetic resonance signals from the encoded sub-volumes; andreconstructing the acquired magnetic resonance signals into image data corresponding to each of the sequential slices in each of the encoded sub-volumes. 2. The method of claim 1, wherein the some slices are selected such that (1+v×fa)-th slices are selected from the sequential slices in each of the excited sub-volumes, wherein fa is a factor for satisfying the condition and is an integer greater than or equal to 1, and v is an integer that increases by 1 beginning from 0 until a value of (1+v×fa) reaches a number of the sequential slices in each of the excited sub-volumes. 3. The method of claim 1, wherein the reconstructing of the acquired magnetic resonance signals into the image data comprises: reconstructing the acquired magnetic resonance signals into image data corresponding to each of the encoded sub-volumes; andreconstructing the image data corresponding to each of the encoded sub-volumes into the image data corresponding to each of the sequential slices in each of the encoded sub-volumes. 4. The method of claim 3, wherein the reconstructing of the acquired magnetic resonance signals into the image data corresponding to each of the encoded sub-volumes and the reconstructing of the image data corresponding to each of the encoded sub-volumes into the image data corresponding to each of the sequential slices in each of the encoded sub-volumes are performed based on channel information of multi-channel receiver coils configured to acquire the magnetic resonance signals. 5. The method of claim 1, wherein the groups comprise a first group through an N-th group, the sub-volumes are sequentially and cyclically distributed to the first group through the N-th group, and N is an integer greater than or equal to 2. 6. The method of claim 1, further comprising performing the applying of the RF pulses and the selection gradient to the subject, the performing of the 3D encoding on each of the excited sub-volumes, the acquiring of the magnetic resonance signals from the encoded sub-volumes, and the reconstructing of the acquired magnetic resonance signals into image data for all of the groups to obtain reconstructed image data for al... | Samsung Electronics Co., Ltd. | G01V003/00 G01R033/483 A61B005/055 G01R033/561
US09808184 | B2 | 2017-11-07 | Motor function evaluation device and motor function evaluation method | 2013-10-18 | To provide a motor function evaluation device and a motor function evaluation method capable of evaluating a motor function of a subject comprehensively and easily. A motor function evaluation device 1 of the present invention includes a measurement base 11, a load measurement unit 14 that measures load change over time of the subject applied to the measurement base 11, and an arithmetic unit 24 that determines a balance ability indicator of the subject determined by the load change over time measured by the load measurement unit 14. The arithmetic unit 24 determines the balance ability indicator from a time interval from when the subject stands up and the load applied to the load measurement unit 14 is maximized until when the load variation is stabilized. | 1. A motor function evaluation device comprising: a measurement base;a load measurement circuit configured to measure load change over time of a subject applied to the measurement base; andan arithmetic circuit configured to determine a balance ability indicator of the subject determined by the load change over time measured by the load measurement circuit, wherein:the arithmetic circuit is configured to determine the balance ability indicator based on a stabilizing time period, within which variation of the load measured by the load measurement circuit is stabilized, andthe stabilizing time period is measured from a time when the subject stands up from a sitting position on a chair and the load applied to the load measurement circuit becomes a maximum value until a time when the variation of the load is stabilized. 2. The motor function evaluation device according to claim 1, wherein the arithmetic circuit is configured to determine a muscle strength indicator based on the maximum value and a weight of the subject. 3. A motor function evaluation method, comprising steps of: measuring load change over time of a subject applied to a measurement base;determining a balance ability indicator of the subject determined by the measured load change over time; andevaluating a motor function of the subject by using the balance ability indicator, wherein:the balance ability indicator is determined based on a stabilizing time period, within which variation of the load measured is stabilized, andthe stabilizing time period is measured from a time when the subject stands up from a sitting position on a chair and the load applied to the measurement base becomes a maximum value until a time when variation of the load is stabilized. 4. The motor function evaluation device according to claim 1, wherein the standing up motion of the subject is accompanied by an inclination of a back of the subject. 5. The motor function evaluation device according to claim 1, wherein the time when the variation of the load is stabilized is defined as a time after a lapse of two cycles or a time when the load measured by the load measurement circuit has a same values as a body weight value after the lapse of the two cycles, after when the load applied to the load measurement circuit becomes the maximum value. 6. The motor function evaluation method according to claim 3, wherein the time when the variation of the load is stabilized is defined as a time after a lapse of two cycles or a time when the load measured by the load measurement circuit has a same values as a body weight value after the lapse of the two cycles, after when the load applied to the load measurement circuit becomes the maximum value. | TANITA CORPORATION | A61B005/11 A61B005/103 A61B005/00 G01G019/50 A61B005/22
US09782115 | B2 | 2017-10-10 | System and method for monitoring bladder and abdominal pressures, and bladder function recovery system | 2015-08-07 | Here provided are systems and methods for monitoring bladder and/or abdominal pressures, and a bladder function recovery device. In the systems and methods, the bladder pressure can be detected in real time with a bladder pressure detection unit that is communicated with a urinary catheter. In addition, the abdominal pressure can also be obtained in real time according to the corresponding relationship between the bladder pressure and the abdominal pressure. Furthermore, the bladder function recovery device can be used for bladder function recovery for patients in need. | 1. A method for monitoring bladder pressure in a subject, comprising the steps of: a. providing a bladder pressure detection unit that is connected to a first end opening of a capillary tube, said capillary tube having a second end opening connected to a urinary drainage tube which is in turn connected to a urinary catheter that is placed in the subject, wherein the capillary tube is filled with pressure transmission medium so that the bladder pressure detection unit does not directly contact with urine;b. detecting the bladder pressure in a continuous manner by the bladder pressure detection unit. 2. The method according to claim 1, further comprising the step of: determining whether the bladder pressure increases by comparing the detected bladder pressure with the normal pressure of the bladder containing the same volume of pre-stored urine; wherein an alarm unit is turned on if the bladder pressure increases. 3. The method according to claim 2, further comprising the steps of: determining whether the bladder pressure is equal to or higher than a preset pressure if the bladder pressure increases; anddetermining whether the bladder pressure is valid if the bladder pressure is equal to or higher than the preset pressure; wherein the bladder pressure is valid if the duration of the bladder pressure is equal to or more than a preset value. 4. The method according to claim 3, further comprising the steps of: allowing urination through a urination device if the bladder pressure is valid;determining whether the bladder pressure reaches a preset minimum threshold value during the process of urination;stopping urination if the bladder pressure reaches a preset minimum threshold value. 5. The method according to claim 1, wherein the bladder pressure detection unit is a low-pressure sensor. 6. A method for monitoring bladder pressure in a subject by detecting the bladder pressure using a bladder pressure monitoring system, wherein the bladder pressure monitoring system comprises: a bladder pressure detection unit;a capillary tube having a first and a second end openings with the first end opening of the capillary tube connected to the bladder pressure detection unit, wherein the capillary tube is filled with pressure transmission medium so that the bladder pressure detection unit does not directly contact with urine and the bladder pressure detection unit detects the bladder pressure in a continuous manner;a urinary drainage tube that is connected to the second end opening of the capillary tube, wherein the drainage tube has a first end opening, a second end opening and one middle opening, with the first end opening of the urinary drainage tube being connected to a urinary catheter, a second end opening of the urinary drainage tube connected to a drainage bag, and the middle opening of the urinary drainage tube connected to the second end opening of the capillary tube;a first solenoid valve being placed inside the urinary drainage tube at a position ... | SUZHOU DAWEI BIOPHARMA, LTD. | A61B005/03 A61B005/20 A61F005/451 A61B005/00 A61F005/44
US09789309 | B2 | 2017-10-17 | Device and implantation system for electrical stimulation of biological systems | 2015-05-08 | The present specification discloses devices and methodologies for the treatment of transient lower esophageal sphincter relaxations (tLESRs). Individuals with tLESRs may be treated by implanting a stimulation device within the patient's lower esophageal sphincter and applying electrical stimulation to the patient's lower esophageal sphincter, in accordance with certain predefined protocols. The presently disclosed devices have a simplified design because they do not require sensing systems capable of sensing when a person is engaged in a wet swallow and have improved energy storage requirements. | 1. A method of determining a plurality of stimulation parameters for an electrical stimulation device, wherein said electrical stimulation device comprises at least one electrode in electrical communication with a lower esophageal sphincter (LES) of a patient, a waveform generator coupled to the at least one electrode, and a controller configured to control the waveform generator, the method comprising: applying a plurality of stimulation sequences, wherein each of said plurality of stimulation sequences is defined by a pulse frequency, pulse duration, and pulse amplitude and wherein each of said plurality of stimulation sequences is applied via said at least one electrode in electrical communication with the LES of the patient;after each of said plurality of stimulation sequences, evaluating a pressure response of the LES of the patient; andterminating said application of a plurality of stimulation sequences when the LES of the patient exhibits a pressure response of at least 20 mm Hg sustained for at least 5 minutes. 2. The method of claim 1 wherein a first of said plurality of stimulation sequences has a pulse frequency of 20 Hz, a pulse duration of 200 μsec, and a pulse amplitude of 5 mAmp. 3. The method of claim 2 wherein a second of said plurality of stimulation sequences has a pulse frequency of 20 Hz, a pulse duration of 200 μsec, and a pulse amplitude of 10-15 mAmp. 4. The method of claim 3 wherein a third of said plurality of stimulation sequences has a pulse frequency of 20 Hz, a pulse duration of 3 ms, and a pulse amplitude of 5-15 mAmp. 5. The method of claim 4 wherein a fourth of said plurality of stimulation sequences has a pulse frequency of 6 cycles/min, a pulse duration of 375 ms, and a pulse amplitude of 5 mAmp. 6. The method of claim 5 wherein a fifth of said plurality of stimulation sequences has a pulse frequency of 6 cycles/min, a pulse duration of 375 ms, and a pulse amplitude of 5-15 mAmp. 7. The method of claim 1 wherein, after terminating an application of the plurality of stimulation sequences, keeping said plurality of stimulation sequences off until a pressure of the patient's LES drops to less than 10 mm Hg or to a baseline pressure of the patient's LES, whichever is higher. 8. The method of claim 7 wherein, after the pressure of the patient's LES drops to less than 10 mm Hg or to the baseline pressure of the patient's LES, whichever is higher, applying another of said plurality of stimulation sequences via said at least one electrode in electrical communication with the LES of the patient. 9. The method of claim 8 further comprising recording periods during which the plurality of stimulation sequences are being applied and periods during which no stimulation is being applied. 10. A method of determining a plurality of stimulation parameters for an electrical stimulation device, wherein said electrical stimulation device comprises at least one electrode in electrical communication with a lower esophageal sp... | EndoStim, Inc. | A61N001/36 A61B005/145 A61N001/372 A61B005/00 A61B005/03 A61B005/053 A61B005/11
US09820681 | B2 | 2017-11-21 | Reducing nuisance alarms | 2014-06-26 | Embodiments disclosed herein may describe systems and methods for reducing nuisance alarms using probability and/or accuracy of a measured physiological parameter, such as the pulse rate or SpO2 measurement generated by a pulse oximeter. Embodiments may include methods for adjusting a predetermined alarm threshold based on the probability distribution of the estimated pulse rate and/or oxygen saturation of a patient's blood. | 1. A method performed by a pulse oximeter comprising a processor and a memory, the method comprising: calculating an estimated value of an oxygen saturation of a patient's blood from information received from a sensor;calculating a probability distribution of an actual value of the oxygen saturation of the patient's blood based at least in part upon information received from the sensor;calculating, by the pulse oximeter, an adjusted threshold based at least in part upon the probability distribution and a predetermined threshold, wherein when the estimated value of oxygen saturation exceeds the adjusted threshold an alarm is indicated;determining that the estimated value of the oxygen saturation exceeds the adjusted threshold;determining, by the pulse oximeter, whether an accuracy of the estimated value has recently degraded by comparing a current accuracy of the estimated value to a previous accuracy of the estimated value; anddelaying generating, by the pulse oximeter, the alarm based at least in part upon determining that the accuracy of the estimated value has recently degraded. 2. The method of claim 1, further comprising storing one or more of: the predetermined threshold,the adjusted threshold,the estimated value of the oxygen saturation of the patient's blood,the previous accuracy associated with the estimated value,the probability distribution, anda delay time period calculated based at least in part on the comparison of the current accuracy of the estimated value and the previous accuracy of the estimated value. 3. The method of claim 1, further comprising: calculating, based at least in part upon the probability distribution, a confidence interval for the actual value of the oxygen saturation of the patient's blood; andcomparing the confidence interval to the predetermined threshold. 4. The method of claim 1, further comprising: retrieving a confidence level; andcalculating, based at least in part upon the probability distribution, a confidence interval for the oxygen saturation of the patient's blood using the retrieved confidence level. 5. The method of claim 4, wherein the retrieved confidence level is selected from 99%, 98%, 97.5%, 95%, 90% and a user-selected value. 6. The method of claim 1, further comprising: calculating the current accuracy of the estimated value based at least in part upon the information received from the sensor. 7. The method of claim 6, wherein calculating the probability distribution further comprises: calculating the probability distribution of the actual value of the oxygen saturation of the patient's blood based at least in part upon the current accuracy of the estimated value. | Covidien LP | A61B005/1455 A61B005/00 G06F019/00
US09820698 | B2 | 2017-11-21 | Actigraphy methods and apparatuses | 2015-11-06 | An actigraphy method includes receiving a physiological parameter signal as a function of time for a physiological parameter other than body motion (such as electrocardiography or a respiration monitor), computing a body motion artifact (BMA) signal as a function of time from the physiological parameter signal (for example, using a local signal power signal, a local variance signal, a short-time Fourier transform, or a wavelet transform over epochs of duration on order a few minutes or less), and computing an actigraphy signal as a function of time from the BMA signal, for example by applying a linear transform to the BMA signal and optionally applying filtering such as median removal and/or high-pass filtering. | 1. A physiological monitoring device comprising: a sensor configured to generate a non-body motion physiological parameter signal as a function of time for a physiological parameter other than velocity, displacement, and acceleration; andan electronic digital signal processing (DSP) device configured to perform operations including: computing a body motion artifact (BMA) signal as a function of time from the non-body motion physiological parameter signal, andcomputing an actigraphy signal as a function of time from the BMA signal. 2. The physiological monitoring device of claim 1 wherein the sensor includes an electrocardiography (ECG) sensor and the physiological parameter includes one or more of (i) at least one ECG trace and (ii) a heart rate. 3. The physiological monitoring device of claim 1 wherein the sensor includes a respiratory sensor and the physiological parameter includes a respiration rate. 4. The physiological monitoring device of claim 1 wherein the sensor includes a Respiratory Inductive Plethysmography (RIP) sensor. 5. The physiological monitoring device of claim 1 wherein computing a BMA signal as a function of time from the non-body motion physiological parameter signal comprises computing a local signal power signal from the non-body motion physiological parameter signal. 6. The physiological monitoring device of claim 1 wherein computing a BMA signal as a function of time from the non-body motion physiological parameter signal comprises computing a local signal variance signal from the non-body motion physiological parameter signal. 7. The physiological monitoring device of claim 1 wherein computing a BMA signal as a function of time from the non-body motion physiological parameter signal comprises computing a Short-Time Fourier Transform (STFT) signal from the non-body motion physiological parameter signal. 8. The physiological monitoring device of claim 1 wherein computing a BMA signal as a function of time from the non-body motion physiological parameter signal comprises computing a wavelet transform signal from the non-body motion physiological parameter signal. 9. The physiological monitoring device of claim 1 wherein computing a BMA signal as a function of time from the non-body motion physiological parameter signal comprises computing a BMA signal sample for each time window of a succession of time windows of the non-body motion physiological parameter signal. 10. The physiological monitoring device of claim 9 wherein the succession of time windows is a succession of overlapping Hamming time windows. 11. The physiological monitoring device of claim 1 wherein computing an actigraphy signal as a function of time from the BMA signal comprises applying a linear transform to the BMA signal. 12. The physiological monitoring device of claim 1 wherein the DSP device is configured to perform further operations including: filtering the actigraphy signal using a median removal filter. 13. The physiological monitoring ... | Koninklijke Philips N.V. | A61B005/02 A61B005/00 A61B005/0245 A61B005/0402 A61B005/11 A61B005/113 A61B005/08
US09778095 | B2 | 2017-10-03 | Multipurpose weighing device | 2013-12-30 | Weighing device of the digital bathroom scale type, comprising four feet each having a strain gauge, four conducting portions on the top surface, and one electronic control unit, the electronic unit being configured to measure first signals indicative of periodic variations in weight and second signals indicative of periodic variations in impedance caused by the heartbeats of the user, the electronic unit being configured to determine the heart rate of the user from the first and second signals. Method for determining the heart rate of the user from the first and second signals. | 1. A digital scale, comprising at least one strain gauge, at least two conducting portions arranged on a top surface of the scale, and an electronic control unit, the strain gauge being connected to the electronic unit and the electronic unit being configured to determine the weight of a user positioned on the weighing device, and to measure a first signal indicative of the periodic variations in weight caused by the heartbeats of the user,the conducting portions being connected to the electronic unit and the electronic unit being configured to measure an impedance at the user's feet and to measure a second signal indicative of the periodic variations in impedance caused by the heartbeats of the user,wherein the electronic unit is configured to evaluate a quality index for the first signal and a quality index for the second signal, and to determine the heart rate from the signal having the best quality index. 2. The weighing device according to claim 1 comprising four conducting portions, wherein the electronic unit is configured to inject a modulated or pulsed current between two of the conducting portions and to measure the electric potential difference between the two other conducting portions, in order to determine the variations in electric potential difference and thus determine the second signal. 3. The weighing device according to claim 1, wherein the electronic unit is configured to perform a cross-correlation calculation on the first and second signals. 4. The weighing device according to claim 3, wherein the electronic unit is configured to measure a phase difference (dT) between the first and second signals, and to shift the second signal relative to the first signal by said phase difference so that the cross-correlation calculation eliminates noise and maximizes the useful signal portions. 5. A digital scale, comprising at least one strain gauge, at least two conducting portions arranged on a top surface of the scale, and an electronic control unit, the strain gauge being connected to the electronic unit and the electronic unit being configured to determine the weight of a user positioned on the weighing device, and to measure a first signal indicative of the periodic variations in weight caused by the heartbeats of the user,the conducting portions being connected to the electronic unit and the electronic unit being configured to measure an impedance at the user's feet and to measure a second signal indicative of the periodic variations in impedance caused by the heartbeats of the user,wherein the electronic unit is configured to perform a cross-correlation calculation on the first and second signals, so to give a third signal,wherein the electronic unit is configured to evaluate a quality index for the first signal, a quality index for the second signal and a quality index for the third signal,wherein the electronic unit is configured to determine the heart rate from the first, second and third signals by choosing the one of the... | Withings | G01G019/414 G01G019/50 A61B005/0245 A61B005/053 A61B005/024
US09801576 | B2 | 2017-10-31 | Application of electrochemical impedance spectroscopy in sensor systems, devices, and related methods | 2017-03-01 | A diagnostic Electrochemical Impedance Spectroscopy (EIS) procedure is applied to measure values of impedance-related parameters for one or more sensing electrodes. The parameters may include real impedance, imaginary impedance, impedance magnitude, and/or phase angle. The measured values of the impedance-related parameters are then used in performing sensor diagnostics, calculating a highly-reliable fused sensor glucose value based on signals from a plurality of redundant sensing electrodes, calibrating sensors, detecting interferents within close proximity of one or more sensing electrodes, and testing surface area characteristics of electroplated electrodes. Advantageously, impedance-related parameters can be defined that are substantially glucose-independent over specific ranges of frequencies. An Application Specific Integrated Circuit (ASIC) enables implementation of the EIS-based diagnostics, fusion algorithms, and other processes based on measurement of EIS-based parameters. | 1. A method of calculating a single, fused sensor glucose value based on respective glucose measurement signals of a plurality of redundant sensing electrodes, comprising: performing respective electrochemical impedance spectroscopy (EIS) procedures for each of the plurality of redundant sensing electrodes to obtain values of at least one impedance-based parameter for each said sensing electrode;measuring the electrode current (Isig) for each of the plurality of redundant sensing electrodes;calculating a plurality of reliability indices for each said sensing electrode based on said values of the at least one impedance-based parameter, said plurality of reliability indices including a dip reliability index that is calculated by: calculating a first similarity index to determine an amount of divergence between respective Isigs of at least a first one and a second one of said plurality of redundant sensing electrodes;calculating a second similarity index to determine an amount of divergence between respective values of the at least one impedance-based parameter for said at least first one and second one of said plurality of redundant sensing electrodes;mapping said first similarity index to said second similarity index; andmapping said second similarity index to said first similarity index;calculating a single fused Isig based on said plurality of reliability indices and respective Isig of each of the plurality of redundant sensing electrodes; andcalibrating said single fused Isig to obtain said single fused sensor glucose value. 2. The method of claim 1, wherein said impedance-based parameter is 1 kHz real impedance. 3. The method of claim 2, wherein said plurality of reliability indices include a bound check reliability index and a noise check reliability index. 4. The method of claim 3, wherein calculation of said bound check reliability index and noise check reliability index includes determining whether each said measured Isig and said values of the 1 kHz real impedance fall within respective predetermined ranges for a bound check and a noise check. 5. The method of claim 4, wherein said predetermined range for the 1 kHz real impedance bound check is between 0.3e+4 and 2e+4. 6. The method of claim 1, wherein said plurality of reliability indices include a sensitivity-loss index that is calculated based on each said electrode's imaginary impedance. 7. The method of claim 6, wherein said imaginary impedance is measured at about 1 kHz over a period of time. 8. The method of claim 7, wherein said plurality of reliability indices include a bound check reliability index and a noise check reliability index, and wherein calculation of said bound check reliability index and said noise check reliability index include determining whether said values of the 1 kHz imaginary impedance fall within respective predetermined ranges for said bound check and noise check. 9. The method of claim 8, wherein said predetermined range for the 1 kHz imaginary imp... | MEDTRONIC MINIMED, INC. | A61B005/1495 A61B005/145 G01N027/416 G01N027/02 A61B005/053 A61B005/1486 A61B005/00 G01N033/66 G01N033/49
US09795309 | B2 | 2017-10-24 | Wearable mobile terminal | 2015-03-06 | A wearable terminal can include a main body configured to be in contact with a body of a user using one side of the main body, a photo plethysmography including a transmitter combined with one side of the main body and outputs a output signal, and a receiver detecting a reflection signal corresponding to the output signal returned, a cover configured to cover the photo plethysmography and be combined with the main body and a controller configured to measure a heart rate based on a reflection signal measured by the photo plethysmography. The cover is combined with the main body and parts corresponding to the transmitter and the receiver of the photo plethysmography include transparent units through which the output signal and the reflection signal are passing and a barrier unit positioned between the transparent units and shielding the output signal and the reflection signal. | 1. A wearable terminal, comprising: a main body comprising a side configured to contact a body part of a user;a photo plethysmography located at the side of the main body and comprising: a transmitter configured to output a signal; anda receiver configured to detect a reflection signal that is reflection of the output signal;a cover of the main body configured to cover the photo plethysmography, the cover comprising: transparent units aligned with the transmitter and the receiver, the transparent units configured to pass the output signal and the reflection signal; anda barrier unit positioned between the transparent units and configured to shield the output signal and the reflection signal;a rubber located between the cover and the photo plethysmography, the rubber including openings that are aligned with the transparent units, and at least one of the openings also aligned with the transmitter and at least one of the openings also aligned with the receiver such that the rubber is configured to block the signal output from the transmitter directly flowing into the receiver; anda controller configured to measure a heart rate of the user based on the detected reflection signal. 2. The wearable terminal of claim 1, wherein the cover is adhered to the main body via waterproof tape. 3. The wearable terminal of claim 1, wherein the transmitter comprises a green light-emitting diode (LED) configured to output green light as the output signal. 4. A wearable terminal, comprising: a main body comprising a first side configured to contact a body part of a user when the wearable terminal is worn by the user and a second side that is an opposite side of the first side;a photo plethysmography located at the first side of the main body and comprising: a transmitter configured to output a signal, anda receiver configured to detect a reflection signal that is a reflection of the output signal from a blood cell in a blood vessel of the user;an acceleration sensor configured to detect movement of the main body when the wearable terminal moves in a first direction toward a direction that the first side faces or in a second direction toward a direction that the second side faces according to movement of the body part, the second direction being an opposite direction of the first direction; anda controller configured to: reduce a first size of the detected reflection signal in response to detection, by the acceleration sensor, of the movement of the main body in the first direction;amplify a second size of the detected reflection signal in response to detection, by the acceleration sensor, of the movement of the main body in the second direction;calculate reflection signal data based on the first or second direction and a speed of the movement detected by the acceleration sensor by compensating the reflection signal; andcalculate a heart rate of the user based on a number of changes in the reflection signal data per reference time. 5. The wearable terminal of clai... | LG ELECTRONICS INC. | A61B005/024 A61B005/00
US09801543 | B2 | 2017-10-31 | Hand-held medical-data capture-device having detection of body core temperature by a microprocessor from a signal from a digital infrared sensor on a separate circuit board with no A/D converter and having interoperation with electronic medical record static IP address system | 2014-12-29 | In one implementation, an apparatus estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point is described, estimates temperature from a digital infrared sensor and determines vital signs from a solid-state image transducer, or determines vital signs from a solid-state image transducer and estimates body core temperature from an infrared measurement of an external source point using a cubic relationship between the body core temperature and the measurement of an external source point; after which the estimated and/or determined information is transmitted to an external database. | 1. A non-touch thermometer to measure a temperature, the non-touch thermometer comprising: a microprocessor;a wireless communication subsystem that is operably coupled to the microprocessor and that is operable to transmit a representation of the temperature;a battery operably coupled to the microprocessor;a digital infrared sensor operably coupled to the microprocessor with no analog-to-digital converter operably coupled between the digital infrared sensor and the microprocessor, the digital infrared sensor having only digital readout ports, the digital infrared sensor comprising a low noise amplifier, an analog-to-digital converter, a digital signal processor, and a pulse width modulator (PWM) in a single industry standard package; wherein the digital infrared sensor is configured with no need for recalibration after an initial factory calibration with an emissivity range from 0.1 to 1.0,wherein the microprocessor is operable to receive from the digital readout ports a digital signal that is representative of an infrared signal, the infrared signal acquired from a forehead of a user and detected by the digital infrared sensor, and the microprocessor is further operable to determine the temperature from the received digital signal by using the PWM. 2. The non-touch thermometer of claim 1, wherein a display device further comprises: a green traffic light operable to indicate that the temperature is good;an amber traffic light operable to indicate that the temperature is low; anda red traffic light operable to indicate that the temperature is high. 3. The non-touch thermometer of claim 1, the microprocessor further comprising: a pixel-examination-module configured to examine pixel values of the at least two images, a temporal-variation module to determine temporal variation of the pixel values between the at least two images being below a particular threshold, a signal processing module configured to amplify the temporal variation resulting in amplified temporal variation, and a visualizer to visualize a pattern of flow of blood in the amplified temporal variation in the at least two images. 4. The non-touch thermometer of claim 1 further comprising: the digital infrared sensor having no analog sensor readout ports. 5. The non-touch thermometer of claim 1 wherein the wireless communication subsystem transmits through an Internet Protocol tunnel. 6. The non-touch thermometer of claim 1 wherein the non-touch thermometer does not support specific discovery protocols or domain name service. 7. The non-touch thermometer of claim 1 wherein a connection is established and data is pushed from the non-touch thermometer through the wireless communication subsystem, thereafter an external device controls flow of the data between the non-touch thermometer and the external device, wherein the connection further comprises: an authenticated communication channel. 8. The non-touch thermometer of claim 1, the wireless communication subsystem is configured t... | Arc Devices, LTD | G08B023/00 A61B005/00 A61B005/01 G06F019/00 A61B005/021 A61B005/026 A61B005/08 A61B005/024 A61B005/0205 G06T003/40 G06T007/00 H04W048/16 H04W076/02 A61B005/0402 A61B005/03 G01J005/00 H04W080/04 G01J005/02 G01J005/10 G08C017/02 H04L029/08 H04W012/06 H04B007/26 H04L012/46 H04N005/33 H04W048/08 G08B005/36 G08B021/18 G01K013/00 G06K009/62 G06T007/174 G06T007/269 G06T007/90 A61B005/0245 H04W012/08
US09788732 | B2 | 2017-10-17 | Optical measuring device and fiber bundle association method | 2014-09-04 | An optical measuring device includes: a light source unit; a measurement probe including a fiber bundle, an illumination fiber that illuminates a living tissue with a illumination light, and a plurality of light-receiving fibers that receives return light of the illumination light reflected and/or scattered at the living tissue; a detection unit that receives the return light of the illumination light detected by the plurality of respective light-receiving fibers, and performs photoelectric conversion to detect respective signal intensities; and an association unit that associates the respective signal intensities detected by the detection unit with distances from the illumination fiber to the respective light-receiving fibers on an end face of a distal end portion of the measurement probe, when light having an intensity gradient around the illumination fiber is projected to the end face of the distal end portion of the measurement probe. | 1. An optical measuring device comprising: a light source unit that emits light for measuring a living tissue;a measurement probe including a fiber bundle configured by bundling a plurality of optical fibers at random, the fiber bundle including an illumination fiber that propagates light from the light source unit to a distal end as illumination light and illuminates the living tissue with the illumination light, and a plurality of light-receiving fibers that receives return light of the illumination light illumined by the illumination fiber and reflected and/or scattered at the living tissue, wherein each of the plurality of optical fibers is arranged such that a position in an arrangement of the optical fibers on an end face of a proximal end portion and a position arrangement of the optical fibers on an end face of a distal end portion are different from each other;a detection unit that receives the return light of the illumination light detected by the plurality of respective light-receiving fibers, and performs photoelectric conversion to detect respective signal intensities; andan association unit that associates the respective signal intensities detected by the detection unit with distances from the illumination fiber to the respective light-receiving fibers on the end face of the distal end portion of the measurement probe, when light having an intensity gradient around the illumination fiber is projected to two or more light receiving fibers whose distances from the illumination fiber are different from each other on the end face of the distal end portion of the measurement probe in order to estimate characterization of the living tissue. 2. The optical measuring device according to claim 1, wherein the association unit associates the respective signal intensities detected by the detection unit with the distances from the illumination fiber to the respective light-receiving fibers on the end face of the distal end portion of the measurement probe, when light having a spatial profile of symmetric intensity distribution around the illumination fiber is projected to the end face of the distal end portion of the measurement probe. 3. The optical measuring device according to claim 2, wherein the fiber bundle is a light guide. 4. The optical measuring device according to claim 1, wherein the association unit sequentially associates the respective signal intensities detected by the detection unit in descending order of the signal intensity, with the light-receiving fibers of the plurality of the light-receiving fibers in order of closer distance from the illumination fiber on the end face of the distal end portion of the measurement probe. 5. The optical measuring device according to claim 4, wherein the fiber bundle is a light guide. 6. The optical measuring device according to claim 1, wherein the association unit performs grouping by the respective distances from the illumination fiber to the plurality of light-receiving fibers on the... | OLYMPUS CORPORATION | A61B001/00 A61B005/00
US09795315 | B2 | 2017-10-24 | Catheter system for mapping of the left atrium, right atrium and coronary sinus | 2014-01-28 | A pair of new catheters designed to be deployed as a catheter system to allow a simultaneous acquisition of electrograms from widely dispersed electrodes in the left atrium, right atrium, and coronary sinus. The first catheter is the spiral globe catheter which has the primary shape of a spiral globe and has additional modifications to facilitate safe entry into the left atrium, to orient the primary axis of the spiral globe toward the mitral valve, and to maximize contact of electrodes to multiple areas of the left atrium. The second catheter is the right atrial and coronary sinus catheter (RA-CS catheter) which allows for electrogram acquisition from the length of the coronary sinus and dispersed areas of the right atrium. The catheter system is designed to provide adequate electrode sensor information so that panoramic mapping of the both atria and the coronary sinus may be performed. | 1. A medical device, comprising: a handle portion comprising an actuator; anda catheter comprising a proximal end, a distal end and intermediate section therebetween, the catheter being coupled to the handle portion at the proximal end, the distal end of the catheter including a shape comprising a plurality of sequential loops around a central rotational axis, wherein the shape at the distal end of the catheter is a spiral globe shape, each of the plurality of sequential loops including a plurality of sensors, wherein the distal end comprises a shaft, the shaft comprising an actuator wire and one of a metal alloy strip or a metal alloy filament, wherein the actuator is configured to place tension on the actuator wire and the metal alloy strip or metal alloy filament acts as a compression element when tension is applied, such that the diameter of each of the sequential loops of the spiral globe is reduced to maintain the spiral globe shape,wherein the catheter further includes a tip portion that is distal to the shape comprising a plurality of sequential loops, andwherein the tip portion contains an electrode that serves as a unipolar reference. 2. The medical device of claim 1, wherein the plurality of sensors are one of electrodes for recording electrical signals from a heart or magnet sensors for use with a three-dimensional mapping system for recording a position of the catheter. 3. The medical device of claim 1, wherein the tip portion is distal to the spiral globe shape and is deflected at an angle of substantially 90 degrees with respect to a spiral globe surface of the spiral globe shape and oriented toward the center of the globe shape. 4. The medical device of claim 1, wherein the plurality of sequential loops have a constant rate of change of latitude versus longitude such that equidistant loops are formed. 5. The medical device of claim 1, wherein the distal end further includes a mitral valve region that is proximal to the plurality of sequential loops, the mitral valve region including a further plurality of sensors and being shaped to contact a circumference of a mitral valve of a heart. 6. The medical device of claim 1, wherein the distal end further includes preshaped curves to orient the central rotational axis toward a mitral valve relative to a inter-atrial septum and a shaft axis of the catheter. 7. The medical device of claim 1, wherein the proximal end of the catheter is coupled to the actuator such that a diameter of the plurality of sequential loops is controllable by the actuator. 8. The medical device of claim 1, wherein the plurality of sequential loops for a shape that comprises one of a spherical shape, a cylindrical shape, a rounded cylindrical shape, a conical shape, or a rounded conical shape. 9. The medical device of claim 1, wherein the plurality of sensors are positioned on the plurality of sequential loops such that the plurality of sensors form equal sized triangles. 10. The medical device of claim 9,... | - | A61B005/05 A61B005/0408 A61B005/00 A61B005/042 A61B005/06
US09821143 | B2 | 2017-11-21 | Steerable sheath including elastomeric member | 2015-09-28 | An MR compatible steerable sheath with elastomeric member is provided. The elastomeric member is configured to serve as a reservoir and receive contrast media therewithin. The elastomeric member is positioned on the distal end of the steerable sheath and may circumferentially surround the sheath shaft or be offset from a longitudinal axis thereof. In operation, the contrast media allows a user to view the distal tip of the steerable sheath by virtue of the contrast media contained within the elastomeric member. | 1. An MR compatible steerable sheath comprising: a tubular shaft defining a longitudinal axis, said tubular shaft receiving first and second longitudinal movement wires operably coupled to a distal end thereof;a control handle having a main body configured to receive first and second rack screws, said first and second rack screws being mechanically led to each other such that movement of said second rack screw along the longitudinal axis of the tubular shaft causes movement of said first rack screw in an opposite direction along the longitudinal axis of the tubular shaft, said second rack screw including a threaded portion on an outer surface at an end thereof; said first longitudinal movement wire operably coupled to said first rack screw and said second longitudinal movement wire operably coupled to said second rack screw; and a rotatable adjustment knob operably engageable with said control handle said rotatable adjustment knob having an internal threaded portion matingly engageable solely with the threaded portion of said second rack screw, said rotatable adjustment knob solely rotatably moveable between a first position and a second position in which the internal thread is configured to engage solely the thread on the outer surface of said second rack screw, wherein said first position causes said second rack screw to move proximally along the longitudinal axis of the tubular shaft which in turn causes said first rack screw to move distally along the longitudinal axis of the tubular shaft and proximal movement of said second rack screw is configured to tension and cause proximal longitudinal movement of the second longitudinal movement wire, and further wherein said second position is configured to move said second rack screw distally along the longitudinal axis of the tubular shaft which in turn causes said first rack screw to move proximally and distal movement of said second rack screw is configured to release tension on the second longitudinal movement wire;an elastomeric member circumferentially positioned on said tubular shaft and including an elastomeric member lumen therewith;a source of contrast media in fluid communication with said elastomeric member lumen,wherein said elastomeric member is configured to denote the position of the steerable sheath. 2. The MR compatible steerable sheath of claim 1 wherein said elastomeric member lumen is in fluid communication with at least one tube defining a tube lumen configured to extend along a longitudinal axis of the tubular shaft on the outside of said tubular shaft. 3. The MR compatible steerable sheath of claim 1 wherein said elastomeric member comprises a plurality of elastomeric members each having an elastomeric member lumen. 4. The MR compatible steerable sheath of claim 3 wherein said plurality of elastomeric member lumens are each in fluid communication with a single tube lumen positioned on an outside of said tubular shaft. 5. The MR compatible steerable sheath of claim 3 where... | IMRICOR MEDICAL SYSTEMS, INC. | A61M025/01 A61B005/055 G01R033/28 A61M025/00
US09814387 | B2 | 2017-11-14 | Device identification | 2013-06-28 | An eye-mountable device includes a controller embedded in a polymeric material configured for mounting to a surface of an eye. The controller is electrically connected to an antenna included in the eye-mountable device. The controller is configured to: (i) receive an indication of an interrogation signal via the antenna, (ii) responsive to the interrogation signal, output a substantially unique identification sequence; and (iii) use the antenna to communicate the substantially unique identification sequence. The substantially unique identification sequence can then be used by external readers to associate the eye-mountable device with device-specific information without storing such information on the eye-mountable device. | 1. An eye-mountable device comprising: a transparent polymeric material having a concave surface and a convex surface, wherein the concave surface removably mounts over a corneal surface and the convex surface is compatible with eyelid motion when the concave surface is so mounted;a substrate at least partially embedded within the transparent polymeric material;an antenna disposed on the substrate;at least one sensor;a sequence generator that dynamically generates a substantially unique identification sequence, wherein the substantially unique identification sequence is based on process variations in components of the sequence generator; anda controller electrically connected to the antenna, wherein the controller: (i) receives an indication of an interrogation signal via the antenna, (ii) responsive to the interrogation signal, outputs the substantially unique identification sequence, (iii) uses the antenna to communicate the substantially unique identification sequence, (iv) responsive to communicating the substantially unique identification sequence, uses the antenna to receive an indication of device-specific information associated with the eye-mountable device, wherein the device-specific information is a response to said communicating of the substantially unique identification sequence and the device-specific information comprises user-specific information for operating the eye-mountable device, and (v) operates the at least one sensor in accordance with the user-specific information. 2. The eye-mountable device according to claim 1, wherein the sequence generator includes one or more state circuits that output one of multiple possible state outputs based on a difference between one or more circuit components in the one or more state circuits, and wherein the substantially unique identification sequence is based at least in part on the outputs of the one or more state circuits. 3. The eye-mountable device according to claim 1, wherein the sequence generator includes one or more comparator circuits that output a binary state value based on a comparison between an output of circuit components in the one or more comparator circuits, and wherein the substantially unique identification sequence is based at least in part on the binary state value. 4. The eye-mountable device according to claim 1, wherein the controller uses the antenna to indicate a measurement obtained based on the operation of the at least one sensor in accordance with the user-specific information. 5. The eye-mountable device according to claim 4, wherein the at least one sensor includes: an electrochemical analyte sensor, a temperature sensor, a light sensor, a pressure sensor, or an acceleration sensor. 6. The eye-mountable device according to claim 1, wherein the at least one sensor includes an electrochemical sensor disposed on the substrate and including: (i) a working electrode, and (ii) a reference electrode adjacent the working electrode, andwherein the controller... | Verily Life Sciences, LLC | G06K007/10 A61B005/00 A61B005/145 A61B005/1477 G02C007/04
US09788712 | B2 | 2017-10-17 | Device and method for determining the presence of middle ear fluid | 2011-11-04 | A device for determining the presence of abnormal fluid in a middle ear of a subject includes an elongated probe, a first light source housed within the elongated probe, and a second light source housed within the elongated probe. The elongated probe includes a distal end for inspection of an ear. The First light source is configured to convey an optical beam through a tympanic membrane associated with the middle ear of the subject, without puncturing the tympanic membrane. The second light source is configured to convey light through the distal end of the elongated probe and illuminate the tympanic membrane. | 1. A method for determining the presence of abnormal fluid in a middle ear of a subject, said method comprising the steps of: providing a device comprising an elongated probe having a distal end configured for inspection of an ear of the subject, a first light source housed within the elongated probe and configured to operate at a low power;activating the first light source to convey an optical beam through the tympanic membrane without puncturing the tympanic membrane; andcreating a reflectance pattern associated with the tympanic membrane based on the optical beam; andobserving the reflectance pattern through the device;wherein a substantially diffuse reflectance pattern indicates the presence of abnormal fluid in the middle ear of the subject. 2. The method of claim 1, wherein a substantially concentrated reflectance pattern indicates the absence of abnormal fluid in the middle ear of the subject. 3. The method of claim 1, wherein said step of providing a device further includes providing a second light source that is housed within the elongated probe and configured to illuminate the tympanic membrane; wherein the second light source comprises an incandescent light source. 4. The method of claim 1, wherein the low power is less than about 10 mW. 5. The method of claim 1, wherein the low power is less than about 5 mW. 6. The method of claim 1, wherein said step of activating the first light source further comprises activating the first light source to convey a coherent optical beam having a wavelength of about 532 nm. 7. The method of claim 1, wherein said step of activating the first light source further comprises activating the first light source to convey a coherent optical beam having a wavelength of about 635 nm. 8. The method of claim 1, wherein the first light source comprises a low power light source selected from the group consisting of a LED, a laser, and a laser diode. | The Cleveland Clinic Foundation | A61B001/227 A61B001/06 A61B005/00 A61B005/12
US09788766 | B2 | 2017-10-17 | Analyte sensing biointerface | 2014-05-19 | Disclosed herein is an analyte sensing biointerface that comprises a sensing electrode incorporated within a non-conductive matrix comprising a plurality of passageways extending through the matrix to the sensing electrode. Also disclosed herein are methods of manufacturing a sensing biointerface and methods of detecting an analyte within tissue of a host using an analyte sensing biointerface. | 1. A transcutaneous continuous glucose sensor system, comprising: a substantially planar sensor, the sensor comprising: a first non-conductive layer;a first conductive layer comprising a first electrode, wherein the first conductive layer is disposed over the first non-conductive layer;a second non-conductive layer disposed over the first conductive layer;a second conductive layer comprising a second electrode, wherein the second conductive layer is disposed over the second non-conductive layer;a third non-conductive layer disposed over the second conductive layer;a third conductive layer comprising a third electrode, wherein the third conductive layer is disposed over the third non-conductive layer;wherein at least one of the first electrode, second electrode, or third electrode is a working electrode configured to measure a signal indicative of a glucose concentration;a membrane disposed on at least a portion of the sensor; andelectronics electrically coupled to the working electrode. 2. The sensor of claim 1, wherein the membrane comprises an enzyme. 3. The sensor of claim 2, wherein the membrane further comprises a resistance domain disposed over the enzyme. 4. The sensor of claim 1, wherein the membrane further comprises a bioactive agent. 5. The sensor of claim 1, wherein the membrane is substantially planar. | DexCom, Inc. | A61B005/05 A61B005/145 A61B005/1486 C12Q001/00 B33Y080/00 A61B005/00
US09805168 | B2 | 2017-10-31 | Method and system for sensitivity analysis in modeling blood flow characteristics | 2015-12-18 | Embodiments include systems and methods for determining cardiovascular information for a patient. A method includes receiving patient-specific data regarding a geometry of the patient's vasculature; creating an anatomic model representing at least a portion of the patient's vasculature based on the patient-specific data; and creating a computational model of a blood flow characteristic based on the anatomic model. The method also includes identifying one or more of an uncertain parameter, an uncertain clinical variable, and an uncertain geometry; modifying a probability model based on one or more of the identified uncertain parameter, uncertain clinical variable, or uncertain geometry; determining a blood flow characteristic within the patient's vasculature based on the anatomic model and the computational model of the blood flow characteristic of the patient's vasculature; and calculating, based on the probability model and the determined blood flow characteristic, a sensitivity of the determined fractional flow reserve to one or more of the identified uncertain parameter, uncertain clinical variable, or uncertain geometry. | 1. A system for determining cardiovascular information for a patient, the system comprising: at least one computer system configured to: generate an anatomical model of at least a portion of a patient's vasculature using patient-specific images of a geometry of at least a portion of the patient's vasculature;create a computational model of a blood flow characteristic of the patient using the anatomical model;generate or identify at least one patient-specific uncertainty value specifying a range of uncertainty for at least one of an uncertain parameter, uncertain clinical variable, and/or uncertain geometry of the anatomical model;generate or receive a probability model comprising statistics for blood flow characteristic calculations, based on the at least one patient-specific uncertainty value;modify the probability model by iteratively: adjusting one or more of the at least one uncertain parameter, uncertain clinical variable, and/or uncertain geometry of the anatomical model using the at least one patient-specific uncertainty value;using the computational model and the one or more adjusted uncertain parameter, adjusted uncertain clinical variable, and/or adjusted uncertain geometry of the anatomical model to generate a blood flow characteristic value; andusing the generated blood flow characteristic value to determine statistical variation of the blood flow characteristic value as a result of adjustment of the at least one patient-specific uncertainty value;receive a target confidence level;calculate a patient-specific blood flow characteristic value and a range of variation for the patient-specific blood flow characteristic value consistent with the target confidence level, using the computational model of the blood flow characteristic and the modified probability model; andoutput the anatomical model, the calculated patient-specific blood flow characteristic value, and the calculated range of variation for the patient-specific blood flow characteristic value. 2. The system of claim 1, wherein the computational model of the blood flow characteristic of the patient is a reduced order model or a one dimensional model. 3. The system of claim 1 wherein the anatomical model is a three-dimensional model, a four-dimensional model, or a one-dimensional wave propagation model. 4. The system of claim 1, wherein the at least one computer system is further configured to initialize the probability model based on a stochastic collocation method. 5. The system of claim 1, wherein: the uncertain clinical variable is one or more of a myocardial mass, hematocrit, viscosity, pressure, body or surface area, heart-rate, systolic and diastolic brachial blood pressures, patient height, patient weight, patient history, smoking status, and presence/absence of diabetes; andthe uncertain geometry is one or more of maximal stenosis diameter, stenosis length, stenosis location, lesion size, lesion length, and healthy segments with image-based artifacts. 6. The system... | HEARTFLOW, INC. | G06F019/10 G06F019/00 A61B006/00 A61B005/02 A61B005/026 G06T007/00 A61B005/055 A61B006/03
US09814402 | B2 | 2017-11-14 | Electrode systems for use with medical monitoring systems | 2014-02-14 | Electrode systems for use with neuromuscular monitoring systems are provided herein. An example electrode system for use with a monitoring system can include a flexible substrate, a connector interface, one or more stimulating electrodes, one or more recording electrodes and a plurality of conductive traces carried on the flexible substrate. The connector interface can be configured to communicatively connect the electrode system with the monitoring system. In addition, the one or more stimulating electrodes can be configured to deliver an electrical pulse, and the one or more recording electrodes can be configured to receive an electrical signal. The plurality of conductive traces can electrically connect at least one stimulating electrode or at least one recording electrode with the connector interface. Further, each of the one or more stimulating electrodes can have an elongate shape with a length dimension that is substantially greater than a width dimension. | 1. An electrode system comprising: a flexible substrate having a length axis and a width axis, wherein said length axis is substantially greater than said width axis, said flexible substrate comprising no more than three islands and bridges between the islands;a connector interface carried on the flexible substrate and configured to communicatively connect the electrode system with a monitoring system;one or more stimulating electrodes carried on a first island of the flexible substrate, said one or more stimulating electrodes configured to deliver an electrical pulse;a first bridge that connects the first island to a second island and a second bridge that connects the second island to a third island;no more than one recording electrodes carried on the second island of the flexible substrate and configured to detect an electrical signal;one recording electrode carried on the third island of the flexible substrate and configured to detect an electrical signal; anda plurality of conductive traces carried on the flexible substrate, wherein a first one of the plurality of conductive traces routes across the first island, the first bridge, the second island, and the second bridge to electrically connect the one recording electrode carried on the third island with the connector interface, and wherein a second one of the plurality of conductive traces routes across the first island and the first bridge to electrically connect the no more than one recording electrode carried on the second island with the connector interface, and wherein a third one or more of the plurality of conductive traces electrically connect the one or more stimulating electrodes carried on the first island with the connector interface, wherein each of the one or more stimulating electrodes has a rectangular-elongate shape with a length dimension that is at least two times greater than a width dimension and said length dimension is substantially perpendicular to the length axis of the flexible substrate. 2. The electrode system of claim 1, wherein the length dimension is at least five times greater than the width dimension. 3. The electrode system of claim 1, wherein each of the one or more stimulating electrodes has at least one rounded corner. 4. The electrode system of claim 3, wherein each of the one or more stimulating electrodes comprises first and second sides that define the length dimension. 5. The electrode system of claim 4, wherein the first and second sides extend substantially in parallel along the length dimension. 6. The electrode system of claim 4, wherein each of the one or more stimulating electrodes comprises third and fourth sides that define the width dimension. 7. The electrode system of claim 6, wherein one of the third and fourth sides forms the at least one rounded corner, the at least one rounded corner connecting the first and second sides. 8. The electrode system of claim 6, wherein the at least one rounded corner connects one of the first and se... | Acacia Designs BV | A61B005/04 A61B005/0492
US09795326 | B2 | 2017-10-24 | Continuous analyte measurement systems and systems and methods for implanting them | 2010-07-22 | Low profile continuous analyte measurement systems and systems and methods for implantation within the skin of a patient are provided. | 1. A continuous analyte measurement system, comprising: a base unit configured for mounting on a skin surface;an analyte sensor comprising a first side edge and a second side edge defining a width of the analyte sensor, and a proximal portion and a distal portion defining a length of the analyte sensor, the analyte sensor further comprising a first and a second functional sides on opposing surfaces of the analyte sensor, a first conductive layer disposed on the first functional side at the distal portion of the analyte sensor and disposed from the first side edge to the second side edge, and a sensing layer disposed on the first conductive layer in a continuous band orthogonal to the first side edge and the second side edge to define an analyte sensor active area, the active area having a rectilinear polygon shape defined by the intersection of the first conductive layer and the sensing layer disposed on the first conductive layer, the proximal portion configured for positioning within the base unit and the distal portion configured for insertion through the skin surface;wherein the base unit comprises a two-piece conductive member including a first piece and a second piece; andwherein the base unit comprises a side wall including a slit through which the width of the proximal portion of the analyte sensor is positioned so that the first and second pieces of the conductive member are electrically coupled with the corresponding first and the second functional sides of the proximal portion of the analyte sensor, the base unit further including one or more features on at least one orthogonally adjacent wall to the side wall for exerting a compressive force on each of the first piece and the second piece of the two-piece conductive member, the one or more features aligned with the two-piece conductive member and comprising one or more exterior flexures of the at least one orthogonally adjacent wall to the side wall. 2. The system of claim 1, wherein at least the proximal portion of the analyte sensor has a planar configuration and wherein the conductive member is mechanically and electrically coupled to the first and the second functional sides of the analyte sensor. 3. The system of claim 1, further comprising an alignment pin extending through the proximal portion of the analyte sensor. 4. The system of claim 1, wherein the base unit comprises a non-conductive compressible material. 5. The system of claim 4, wherein the non-conductive compressible material comprises silicone. 6. The system of claim 1, wherein the conductive member comprises a conductive compressible material. 7. The system of claim 6, wherein the conductive compressible material comprises carbon-doped silicone. 8. The system of claim 1, further comprising a transmitter coupled to the base unit, wherein the base unit further comprises a first pair of receptacles for receiving a corresponding first pair of conductive pins of the transmitter, wherein the conductive pins contact... | Abbott Diabetes Care Inc. | A61B005/145 A61B005/1486 A61B005/00
US09814407 | B2 | 2017-11-14 | Cardiopulmonary resuscitation compression force indicator | 2007-12-14 | A cardiopulmonary resuscitation (CPR) compression force indicator includes circuitry (20, 26, 28, 30) for monitoring a patient's transthoracic impedance while the patient is being given CPR and generating a corresponding impedance signal, circuitry (32, 34, 36) for processing the impedance signal to provide an ongoing measurement of cardiac hemodynamic output, a microprocessor (24) for determining if the measurement falls outside pre-set limits, and indicator(s) (38) for indicating such determination externally to the person giving the CPR. | 1. A cardiopulmonary resuscitation compression force device, comprising: a processor configured to perform operations comprising: receiving, via a pair of electrodes configured to be in contact with a patient, a transthoracic impedance measurement of the patient;computing a first derivative of the transthoracic impedance measurement, to yield an impedance cardiogram;identifying, within the impedance cardiogram, a peak voltage;comparing the peak voltage to an upper threshold voltage, to yield a first comparison, wherein when the peak voltage exceeds the upper threshold voltage, the first comparison indicates a compression force is excessive;comparing the peak voltage to a lower threshold voltage, to yield a second comparison, wherein when the peak voltage is below the lower threshold voltage, the second comparison indicates the compression force is inadequate;determining based on both the first comparison and the second comparison, how much force to apply in subsequent cardiopulmonary resuscitation to yield a determination; andoutputting, based on the determination, instructions; andan indicator, wherein the indicator receives the instructions from the processor and issues a notification to a user, wherein the notification indicates to the user that a force applied in subsequent cardiopulmonary resuscitation should be maintained or adjusted. 2. The cardiopulmonary resuscitation compression force device as of claim 1, wherein the notification comprises at least one of an audible signal or a visual signal. 3. The cardiopulmonary resuscitation compression force device of claim 2, the processor being further configured to perform operations comprising: calculating, based on the first derivative of the transthoracic impedance measurement, a rate at which compressions are applied during cardiopulmonary resuscitation;comparing the rate to a metronome rate, to yield a speed comparison;second determining based on the speed comparison, whether the rate is too fast or too slow; andissuing, via the indicator, a second notification as a result of the second determining, wherein the second notification instructs the user that a rate applied in subsequent cardiopulmonary resuscitation should be maintained or adjusted. 4. The cardiopulmonary resuscitation compression force device as of claim 1, the processor being further configured to perform operations comprising: calculating, based on the first derivative of the transthoracic impedance measurement, a rate at which compressions are applied during cardiopulmonary resuscitation;comparing the rate to a metronome rate, to yield a speed comparison;second determining based on the speed comparison, whether the rate is too fast or too slow;outputting second instructions as a result of the second determining, wherein the indicator receives the second instructions from the processor and issues a second notification to the user, wherein the second notification instructs the user that a rate applied in subsequent cardio... | HEARTSINE TECHNOLOGIES LIMITED | A61B005/08 A61B005/053 A61H031/00 A61N001/39 G09B023/28 A61B005/00
US09829479 | B2 | 2017-11-28 | Photoacoustic imaging device and oxygen saturation measurement method | 2015-03-06 | Provided is a photoacoustic imaging device including: a light source unit which generates an ultra-broadband pulsed laser beam and outputs the ultra-broadband pulsed laser beam; a filter unit which filters narrowband pulsed laser beams having predetermined different wavelength bands from the ultra-broadband pulsed laser beam to selectively extract the narrowband pulsed laser beams and outputs the narrowband pulsed laser beams as pulsed laser beams for photoacoustic imaging; and a PA (photoacoustic) unit which receives the pulsed laser beams for photoacoustic imaging to irradiate a measurement object with the pulsed laser beams for photoacoustic imaging and receives photoacoustic signals generated from the measurement object. | 1. A photoacoustic imaging device comprising: a light source unit which generates an ultra-broadband pulsed laser beam and outputs the ultra-broadband pulsed laser beam;a filter unit which filters narrowband pulsed laser beams having predetermined different wavelength bands from the ultra-broadband pulsed laser beam to selectively extract the narrowband pulsed laser beams and outputs the narrowband pulsed laser beams as pulsed laser beams for photoacoustic imaging; anda PA (photoacoustic) unit which receives the pulsed laser beams for photoacoustic imaging to irradiate a measurement object with the pulsed laser beams for photoacoustic imaging, receives photoacoustic signals generated from the measurement object, processes the received photoacoustic signals to detect an oxygen saturation of the measurement object, and outputs the oxygen saturation of the measurement object,wherein the filter unit includes:a first prism which receives the ultra-broadband pulsed laser beam as an input, separates the ultra-broadband pulsed laser beam wavelength by wavelength, and emits the ultra-broadband pulsed laser beam;a tunable bandpass filter which includes a slit of which position is changeable, selectively filters the narrowband pulsed laser beams by using the slit, and emits the narrowband pulsed laser beams; anda second prism which receives the narrowband pulsed laser beam emitted from the tunable bandpass filter, generates the pulsed laser beams for photoacoustic imaging, and emits the pulsed laser beams for photoacoustic imaging, andwherein the position of the slit of the tunable bandpass filter is changed so as to correspond to incident positions of the narrowband pulsed laser beams. 2. The photoacoustic imaging device according to claim 1, wherein the light source unit includes: a pulsed laser which generates a pulsed laser beam; andan optical output unit which receives the pulsed laser beam as an input, converts the pulsed laser beam into the ultra-broadband pulsed laser beam, and emits the ultra-broadband pulsed laser beam. 3. The photoacoustic imaging device according to claim 2, wherein the light source unit further includes a collimating lens which collimates the ultra-broadband pulsed laser beam supplied from the optical output unit and supplies the collimated ultra-broadband pulsed laser beam to the filter unit. 4. The photoacoustic imaging device according to claim 1, wherein the filter unit further includes a lens which changes an optical path of the pulsed laser beam emitted from the first prism so that the pulsed laser beam is perpendicularly incident on an incident surface of the tunable bandpass filter. 5. The photoacoustic imaging device according to claim 1, wherein the filter unit further includes a lens which changes optical paths of the narrowband pulsed laser beams emitted from the tunable bandpass filter so that the narrowband pulsed laser beams are perpendicularly incident on an incident surface of the second prism. 6. The phot... | POSTECH ACADEMY-INDUSTRY FOUNDATION | G01N033/48 G01N033/50 G01N033/49 A61B005/00 G02B005/04 A61B005/1455 G01N029/06 G01N029/24 G01N021/17
US09817473 | B2 | 2017-11-14 | Portable device and method of controlling therefor | 2015-02-27 | A method of controlling a portable device according to one embodiment of the present specification may include the steps of displaying a first application including a first object and a second object, receiving a first triggering signal indicating that a front direction of the foot wearable device corresponds to a first direction from the foot wearable device, detecting a front direction of the portable device corresponding to a second direction, displaying the first object based on the first direction and displaying the second object based on the second direction. | 1. A portable device, comprising: a display unit configured to display visual information;a location information sensing unit configured to detect location information of the portable device;a direction information sensing unit configured to detect direction information of the portable device;a communication unit configured to receive information from a foot wearable device; anda processor configured to:display a first application including a first object and a second object,display the first object based on a first direction and the second object based on a second direction when a first triggering signal indicating that a front direction of the foot wearable device corresponds to the first direction is received from the foot wearable device and a front direction of the portable device is detected as the second direction,display a first interface indicating path information when the first application is displayed,display a first indicator indicating a path moving direction in the first interface based on the location information of the portable device, andprovide at least one of audio feedback and tactile feedback when the path moving direction of the first indicator is not identical to the first direction. 2. The portable device of claim 1, wherein the processor is further configured to: display the first object based on changed a third direction and the second object based on the second direction when a second triggering signal indicating that the front direction of the foot wearable device is changed to the third direction from the first direction is received from the foot wearable device. 3. The portable device of claim 1, wherein the processor is further configured to: display the first object based on the first direction and the second object based on changed a third direction when the front direction of the portable device is changed to the third direction from the second direction. 4. The portable device of claim 1, wherein the first object corresponds to an object indicating current location information of the portable device. 5. The portable device of claim 4, wherein the second object corresponds to an object indicating information on a real object positioned adjacent to the portable device based on the current location information of the portable device. 6. The portable device of claim 1, wherein the foot wearable device detects surrounding topographic information of the foot wearable device based on the front direction and transmits the detected surrounding topographic information to the portable device. 7. The portable device of claim 6, wherein the processor is further configured to: receive the surrounding topographic information from the foot wearable device, andset the path information of the first interface based on the received surrounding topographic information. 8. The portable device of claim 1, wherein the processor is further configured to display the first application when a second triggering signal is received fr... | LG ELECTRONICS INC. | G06F003/01 G06F001/16 G01C021/00 A61B005/00 A61B005/11 A61B005/117
US09775547 | B2 | 2017-10-03 | System and method for storing and providing patient-related data | 2015-11-03 | According to various embodiments, a regional oximetry sensor may include a light emitting element configured to emit light, a light detector configured to receive the light, and a memory device configured to store a baseline. The baseline stored by the memory device enables a regional oximetry monitor to display the baseline on a display of the regional oximetry monitor. | 1. A regional oximetry sensor, comprising: a light emitting element disposed on a tissue contacting surface of the regional oximetry sensor and configured to emit light;a first light detector and a second light detector each disposed on the tissue contacting surface and each configured to detect the light, wherein the first light detector is disposed at a first distance from the light emitting element and the second light detector is disposed at a second distance from the light emitting element, and wherein the first distance is greater than the second distance;a memory device; andprocessing circuitry associated with the memory device, wherein the processing circuitry is configured to: determine a baseline based at least in part on the detected light, wherein the baseline comprises a plurality of rSO2 measurements of a patient obtained by the regional oximetry sensor over a period of time, and wherein the plurality of rSO2 measurements are based at least in part on the detected light;store the baseline within the memory device;transmit the baseline to a display of a regional oximetry monitor when the regional oximetry sensor is coupled to the regional oximetry monitor;detect removal or dislodging of the regional oximetry sensor from the patient; andlabel the baseline with a timestamp associated with the detected removal or dislodging of the regional oximetry sensor from the patient and store the labeled baseline within the memory device. 2. The sensor of claim 1, wherein the processing circuitry is configured to determine an average rSO2 value of the plurality of rSO2 measurements obtained by the regional oximetry sensor over the period of time. 3. The sensor of claim 1, wherein the processing circuitry is configured to retain the baseline in the memory device when the regional oximetry sensor is disconnected from the regional oximetry monitor or is removed from the patient. 4. The sensor of claim 1, wherein the processing circuitry is configured to store a plurality of timestamps associated with the plurality of rSO2 measurements in the memory device. 5. The sensor of claim 1, wherein the processing circuitry is configured to store a plurality of baselines over a plurality of time periods. 6. The sensor of claim 5, wherein the regional oximetry monitor prompts the caregiver to select a desired baseline of the plurality of baselines when the regional oximetry monitor is coupled to the regional oximetry sensor. 7. The sensor of claim 5, wherein the processing circuitry is configured to store a plurality of timestamps with each of a plurality of baselines stored in the memory device. 8. The sensor of claim 1, wherein the processing circuitry is configured to calculate a cerebral saturation value or a somatic saturation value based at least in part on the received light. 9. The sensor of claim 1, wherein the processing circuitry is configured to associate a patient identification code with the baseline to correctly identify the patient from ... | Covidien LP | A61B005/1455 A61B005/00
US09782096 | B2 | 2017-10-10 | Textile substrate for measuring physical quantity | 2012-01-31 | A textile substrate with a measuring sensor for measuring a physiological signal. The textile substrate has integrated irremovably therewith a base structure component for electronics, such as for example a plastic base for a transmitter for attaching the transmitter and other electronics to the substrate irremovably by way of said base structure component. In addition, signal transfer elements from the measuring sensor are adapted to extend in a watertight manner to the electronics through said base structure component integrated irremovably with the textile substrate. | 1. A device comprising a textile substrate, wherein the textile substrate comprises: measuring sensors for measuring a physiological signal;signal transfer conductors which are coupled with and extended from the measuring sensors along the textile substrate; anda base structure component of physiological signal processing or transmitting electronics,wherein the base structure component is more rigid than the textile substrate and is irremovably and permanently integrated into the textile substrate,wherein said base structure component comprises signal transfer elements for connecting said electronics to the signal transfer conductors, and,wherein the signal transferring from the measuring sensors to the electronics is implemented by electrically coupling the signal transfer conductors from the textile substrate to the signal transfer elements of the base structure component and thereby enabling signal transferring through said base structure component to said electronics in an irremovably, permanently integrated and watertight manner. 2. A device as set forth in claim 1, wherein the measuring sensors are implemented with electrodes and said electrodes being coupled with said conductors, and wherein said irremovably attached electronics comprises at least a transmitter for the wireless transmission of measured signals. 3. A device as set forth in claim 1, wherein the textile substrate comprises a second component structurally more rigid than the textile substrate, and wherein said second component structurally more rigid than the textile substrate is provided on top of the base structure component. 4. A device as set forth in claim 3, wherein between said substrate and the second component provided on top of the base structure component of the electronics is only a mechanical connection, but no electrical connections. 5. A device as set forth in claim 3, wherein said second component is designed to be removable. 6. A device as set forth in claim 3, wherein said second component is at least partially coated with textile. 7. A device as set forth in claim 1, wherein said textile substrate is a heart rate belt, a band, or a garment. 8. A device as set forth in claim 1, wherein elements integrated irremovably with a region of the textile substrate comprising the electronics or the structure more rigid than the textile substrate, are attached by a process selected from the group chosen from injection molding, gluing, sewing, screws, upsetting, ultrasonic welding, high frequency lamination and heat lamination. 9. The device as set forth in claim 1, wherein the device is a wearable article. 10. The device as set forth in claim 9, wherein the textile substrate comprises a second component structurally more rigid than the textile substrate, and wherein the second component is a protective cover and said protective cover is designed to be removable. 11. The device of claim 9, the wearable article chosen from a heart rate belt, a heart rate band an... | Clothing Plus MBU Oy | A61B005/0408 A61B005/00 A61B005/0245 A61B005/024 A61B005/04 H01R013/52
US09785746 | B2 | 2017-10-10 | Systems and methods for determining blood flow characteristics using flow ratio | 2014-07-03 | Embodiments include a system for determining patient cardiovascular information which includes at least one computer system configured to receive patient-specific data regarding a geometry of an anatomical structure of a patient; create a model representing at least a portion of the anatomical structure of the patient based on the patient-specific data; determine a first blood flow rate at at least one point of interest in the model by using relations of individual-specific anatomic data to functional estimates of blood flow characteristics generated from a plurality of individuals; modify the model; determine a second blood flow rate at a point in the modified model corresponding to the at least one point of interest by using the relations of individual-specific anatomic data to functional estimates of blood flow characteristics; and determine a fractional flow reserve value as a ratio of the second blood flow rate to the first blood flow rate. | 1. A system for determining cardiovascular information for a patient, the system comprising: at least one computer system configured to: receive patient-specific data of the patient's cardiovascular anatomy;obtain a first model representing a portion of the patient's cardiovascular anatomy based on the patient-specific data;determine a first point of interest in the first model;determine, for the first point of interest, a first mean blood flow rate within the patient's cardiovascular anatomy represented by the first model;identify one or more narrowings of vasculature upstream from the first point of interest;determine, for each identified narrowing upstream from the first point of interest, a modified cardiovascular anatomy based on a potential treatment or mitigation of each narrowing;generate a revised model by modifying the first model based on the determined modified cardiovascular anatomy;determine a revised model point of the revised model corresponding to the first point of interest of the first model;determine, for the revised model point, a second mean blood flow rate within the patient's cardiovascular anatomy represented by the revised model; anddetermine and generate an output of a ratio of the second mean blood flow rate of the revised model to the first mean blood flow rate of the first model. 2. The system of claim 1, wherein the system is further configured to: determine functional significance of disease at the one or more narrowings of vasculature based on the determined and generated ratio. 3. The system of claim 1, wherein the modification to the first model based on the determined modified cardiovascular anatomy comprises removing one or more anatomic restrictions proximal to the first point of interest or widening a section of the first model. 4. The system of claim 1, wherein the first model representing the portion of the patient's modified cardiovascular anatomy includes at least a portion of an aorta and at least a portion of a coronary artery emanating from the portion of the aorta, wherein the revised model point is located at a location in the coronary arteries. 5. The system of claim 1, wherein the system is further configured to: determine the second mean blood flow rate for a plurality of revised model points, wherein the second mean blood flow rate for one selected revised model point is different from the second mean blood flow rate for a second selected revised model point; anddetermine the ratio of the second mean blood flow rate of the revised model to the first mean blood flow rate of the first model for each revised model point of the plurality of revised model points. 6. The system of claim 1, wherein the patient-specific data includes imaging data and the at least one computer system is configured to obtain the first model based on the imaging data by locating boundaries of lumens of coronary arteries of the patient's heart using the imaging data. 7. The system of claim 1, wherein the second mean b... | HeartFlow, Inc. | A61B005/02 G06F019/00 G06F017/50 G06F017/10 A61B005/029 A61B005/107 G06T019/00 A61B005/00 A61B005/021 A61B005/026
US09782118 | B2 | 2017-10-10 | Diagnosis of swallowing disorders using high resolution manometry | 2013-05-17 | Parameters extracted from high definition manometric measurements of the pharynx and upper esophagus at preidentified locations and times are applied to a supervised learning machine trained using x-ray fluoroscopy data to provide diagnostic information comparable to that of x-ray fluoroscopy without radiation exposure. | 1. A system for diagnosing swallowing disorders of a patient comprising: a manometer adapted to be inserted along the upper aerodigestive tract in the vicinity of the patient's pharynx and esophagus to provide a series of pressure measurements at different points along the patient's pharynx;a computer receiving the series of pressure measurements at different points and executing a stored program to:(a) link the pressure measurements during a swallowing to different multiple predetermined locations along the pharynx;(b) analyze pressure as a function of time in the predetermined locations to generate a set of pressure based predetermined parameters;(c) process the pressure based predetermined parameters to derive image-describing diagnostic parameters and values that conform to a modified barium swallow impairment profile (MBSImP); and(d) output the image-describing diagnostic parameters and values to a user. 2. The system of claim 1 wherein the linking of pressures to predetermined locations identifies the predetermined locations with respect to pressure pattern having dimensions of distance along the pharynx and time. 3. The system of claim 2 wherein the identification of the predetermined locations receives input from a user viewing time pressure pattern identifying at least one point associated with a predetermined location. 4. The system of claim 3 wherein the stored program combines at least one point received from the user with other data from the pressure pattern to identify the predetermined locations. 5. The system of claim 1 wherein the manometer provides for pressure measurements no greater than 1 cm separation along the pharynx. 6. The system of claim 5 wherein the predetermined locations include the patient's velopharynx, and esophageal sphincter. 7. The system of claim 6 wherein the computer further executes a stored program to determine a time of opening of the upper esophageal sphincter. 8. The system of claim 6 wherein the pressure based predetermined parameters for the predetermined locations are selected from the group consisting of: maximum pressures, rise time, rise rate, fall-time, duration of pressure above a baseline, duration between opening and closure peak pressures, and nadir pressure duration. 9. The system of claim 8 wherein the maximum pressures include peak pressure before opening of the esophageal sphincter and a maximum pressure after opening of the esophageal sphincter. 10. The system of claim 6 wherein the predetermined locations include the patient's mesopharynx. 11. The system of claim I wherein the image-describing diagnostic parameters and values provide an indication of normal or abnormal swallowing function. 12. The system of claim 11 wherein the diagnosis provides an indication of a cause of the swallowing disorder selected from the group consisting of cricopharyngeus abnormality, glottic insufficiency, pharyngeal stiffness/fibrosis, neurological disorder, Zenker's diverticulum, head and neck... | Wisconsin Alumni Research Foundation | A61B005/00 A61B005/03 A61B006/00
US09808342 | B2 | 2017-11-07 | Balloon sizing device and method of positioning a prosthetic heart valve | 2013-03-07 | A sizing device for a collapsible prosthetic heart valve, the sizing device includes a collapsible and expandable balloon having a proximal end, a distal end. At least one microelectromechanical sensor is coupled to the balloon, the at least one sensor being capable of measuring information related to size and stiffness of tissue. | 1. A sizing device for a collapsible prosthetic heart valve, the sizing device comprising: a plurality of microelectromechanical sensors, each of the plurality of microelectromechanical sensors being disposed on a flexible substrate, having a top electrode having a contacting surface, a bottom electrode, and an air gap disposed between the top electrode and the bottom electrode, the top electrode and the bottom electrode being deflectable relative to one another, and configured to measure capacitive changes to determine a relative deflection between the top electrode and the bottom electrode, each of the plurality of microelectromechanical sensors being capable of independently measuring, with respect to others, a stiffness value of a native valve annulus by pressing against portions of the native valve annulus with the contacting surface, each flexible substrate being joined edge-to-edge along seams with adjacent substrates to form a collapsible and expandable balloon. 2. The device of claim 1, further comprising a conduit in fluid communication with the inside of the balloon for delivering a fluid to inflate the balloon. 3. The device of claim 2, wherein the balloon comprises at least one of PET, Nylon, polyurethane or a thermoplastic elastomer. 4. The device of claim 1, wherein the flexible substrate includes a polymer. 5. The device of claim 4, wherein the polymer includes polydimethylsiloxane. 6. The device of claim 4, wherein the polymer includes a polyimide. 7. The device of claim 1, wherein the flexible substrate includes a fabric. 8. The device of claim 1, wherein the plurality of microelectromechanical sensors are in contact with one another. 9. A sizing device for a collapsible prosthetic heart valve, the sizing device comprising a plurality of microelectromechanical sensors, each of the plurality of microelectromechanical sensors being disposed on a flexible substrate, the flexible substrate being directly coupled edge-to-edge along seams to adjacent substrates to form a collapsible and expandable balloon, each of the plurality of microelectromechanical sensors having a top electrode having a contacting surface, a bottom electrode, and an air gap disposed between the top electrode and the bottom electrode, the top electrode and the bottom electrode being deflectable relative to one another, and being capable of independently measuring, with respect to others, a property of tissue. 10. The device of claim 9, wherein each of the plurality of microelectromechanical sensors is capable of measuring data relating to native valve annulus diameter. 11. The device of claim 9, wherein each of the plurality of microelectromechanical sensors is capable of measuring data relating to the extent of calcification of tissue. | St. Jude Medical, Cardiology Division, Inc. | A61F002/24 A61B005/02 A61B005/107
US09788775 | B2 | 2017-10-17 | Methods and apparatus for selectively occluding the lumen of a needle | 2014-09-25 | A fluid transfer device for parenterally transferring fluid to and/or from a patient includes a housing, a needle, and an occlusion mechanism. The housing defines a fluid flow path and is couplable to a fluid reservoir. The needle has a distal end portion that is configured to be inserted into the patient and a proximal end portion that is configured to be fluidically coupled to the fluid flow path of the housing, and defines a lumen therebetween. The occlusion mechanism selectively controls a fluid flow between the needle and the fluid flow path. The occlusion mechanism includes an occlusion member that is movable between a first configuration where the lumen of the needle is obstructed during insertion into the patient and a second configuration where the lumen of the needle is unobstructed after the needle has been inserted into the patient allowing fluid transfer to or from the patient. | 1. A device for parenterally transferring fluid to or from a patient, the device comprising: a housing that defines a fluid flow path, the housing configured to be coupleable to a fluid reservoir;a needle having a distal end portion, a proximal end portion, and defining a lumen therebetween, the distal end portion having a distal most surface and configured for insertion into the patient, the proximal end portion configured to be fluidically coupleable to the fluid flow path; andan occlusion mechanism operable to selectively control fluid flow between the needle and the fluid flow path, the occlusion mechanism configured to be moved between a first configuration such that the lumen of the needle is obstructed during insertion into the patient, and a second configuration such that the lumen of the needle is unobstructed after the needle has been inserted into the patient allowing fluid transfer to or from the patient,the occlusion mechanism including an occlusion member operably coupled to the needle, the occlusion member configured to move from a first position to a second position to place the occlusion mechanism in the second configuration, the occlusion member having a distal most surface, the distal most surface of the occlusion member being substantially coplanar to the distal most surface of the needle when the occlusion member is in the first position, the distal most surface of the occlusion member being nonparallel to the distal most surface of the needle when the occlusion member is in the second position. 2. The device of claim 1, wherein the occlusion member is movably disposed on at least a portion of the needle. 3. The device of claim 2, wherein the needle has a closed distal end and a port in fluid communication with the lumen disposed in the distal end portion of the needle. 4. The device of claim 3, wherein the port is obstructed when the occlusion member is in the first position and substantially unobstructed when the occlusion member is in the second position. 5. The device of claim 3, wherein the port includes a plurality of apertures disposed in the distal end portion of the needle. 6. The device of claim 1, wherein the occlusion member is configured to rotate about the needle from the first position to the second position. 7. The device of claim 1, wherein the distal end portion of the needle is a first portion, the first portion having a first outer diameter, the needle having a second portion having a second outer diameter, the first outer diameter greater than the second outer diameter. 8. The device of claim 7, wherein the occlusion member is disposed on the second portion of the needle and has a third outer diameter, the third outer diameter substantially equal to the first outer diameter. 9. The device of claim 1, wherein the distal end portion of the needle defines a recessed portion fluidically isolated from the needle lumen, the recessed portion configured to receive at least one of tissue, bodily fluid, and... | Magnolia Medical Technologies, Inc. | A61M005/158 A61B005/15 A61M039/22 A61M005/168 A61M005/32
US09795800 | B2 | 2017-10-24 | Transmitting treatment information | 2016-09-21 | A system includes a first computing device comprising a processor coupled to a memory. The processor and the memory are configured to receive at least one of (i) information indicative of treatment of a victim by a first caregiver using the first computing device and (ii) information indicative of a health status of the victim; determine that treatment of the victim by the first caregiver using the first computing device is completed; and transmit the received information to a second computing device. | 1. A defibrillating system comprising: one or more electrodes configured to measure an electrocardiogram (ECG) signal from a patient and deliver shock therapy to the patient; anda processor coupled to a memory, the processor and the memory configured to perform operations comprising:analyzing a portion of the electrocardiogram (ECG) signal measured by the one or more electrodes of a defibrillating device to determine whether to deliver shock therapy to the patient using the defibrillating device;in response determining whether to deliver the shock therapy, generating data representing a treatment event;associating the analyzed portion of the ECG signal with the treatment event data;determining that a treatment of the patient by a first caregiver using the defibrillating device is completed; andstoring, in response to the determination, the treatment event data in a profile associated with the patient, the profile configured to be accessible by a second caregiver. 2. The defibrillating system of claim 1, wherein the operations further comprise: delivering shock therapy when a shockable rhythm identified from the ECG signal, wherein shock therapy is not delivered if a non-shockable rhythm is identified from the ECG signal. 3. The defibrillating system of claim 1, wherein the treatment event data comprises: an identifier of the patient being treated by the shock therapy;an indicator indicating whether the portion of the ECG signal is indicative of a shockable rhythm or whether the portion of the ECG signal is indicative of a non-shockable rhythm; anda timestamp representing a time at which the treatment event is generated. 4. The defibrillating system of claim 3, wherein the treatment event data is associated with additional data, the additional data including one or more of SpO2 data, CO2 data, ventilation data, chest compression data, and narrative data, the additional data being contemporaneous with a time period during which the portion of the ECG signal was measured by the one or more electrodes. 5. The defibrillating system of claim 1, wherein the operations further comprise: storing the treatment event data in a storage module of the defibrillating system; andassociating the treatment event data with a patient profile stored in the storage module,wherein storing the treatment data in a storage module includes cataloging a type of the portion of the ECG signal. 6. The defibrillating system of claim 1, wherein the operations further comprise: associating a second portion the ECG signal measured by the one or more electrodes to the treatment event, the second portion of the ECG signal corresponding to a time period after the first portion of the ECG signal. 7. The defibrillating system of claim 1, wherein a length of the portion of the ECG signal corresponds to a predetermined length of time preceding the determination of whether to deliver the shock therapy. 8. The defibrillating system of claim 7, wherein the predetermined length of tim... | ZOLL Medical Corporation | A61B005/04 A61N001/00 A61N001/39 A61B005/0404 A61N001/02 A61B005/00 G06Q050/22 A61B005/053
US09814427 | B2 | 2017-11-14 | Dental appliance for use in supporting sensing devices within an oral cavity | 2015-08-18 | A dental appliance for use in supporting a sensing device within an oral cavity is described herein. The dental appliance includes a sensor assembly and a tooth cap. The sensor assembly includes a support member and a sensor support assembly extending outwardly from the support member. The sensor support assembly includes a recessed portion that defines a sensor chamber that is configured to receive the sensing device therein. The tooth cap is coupled to the sensor assembly and adapted to be positioned within the oral cavity and orientated along a biting surface. | 1. A dental appliance for use in positioning a sensing device within an oral cavity, comprising: a sensor assembly including:a support member including a support ring configured to extend radially about an outer surface of an abutment device to facilitate coupling the sensor assembly to the abutment device:a sensor support assembly coupled to the support ring and extending radially outwardly from an outer surface of the support ring, the sensor support assembly including a recessed portion defining a sensor chamber that is configured to receive the sensing device therein; anda chamber cover removably coupled to the sensor support assembly to enclose the sensor chamber, the chamber cover including a plurality of orifices configured to couple the sensor chamber in flow communication with the oral cavity; anda tooth cap coupled to the sensor assembly and adapted to be positioned within the oral cavity and orientated along a biting surface, the tooth cap including an inner surface defining a cavity that is configured to receive the sensor assembly therein and a cap opening extending through an outer surface of the tooth cap and configured to receive the sensor support assembly therein such that an outer surface of the chamber cover is substantially flush with the outer surface of the tooth cap. 2. A dental appliance in accordance with claim 1, the support member including a plurality of fasteners configured to couple the chamber cover to the sensor support assembly. 3. A dental appliance in accordance with claim 1, the tooth cap including a contoured outer surface configured to form a portion of the biting surface. 4. A dental appliance in accordance with claim 1, wherein the abutment device includes a supporting ridgeline, the tooth cap including an outer ridge configured to contact the supporting ridgeline to support the tooth cap from the abutment device. 5. A dental appliance in accordance with claim 4, wherein the outer ridge of the tooth cap is adapted to be positioned below a gumline. 6. A dental appliance in accordance with claim 1, wherein the sensing device includes a pH sensor. 7. A dental appliance in accordance with claim 1, wherein each of the plurality of orifices are configured to allow saliva to enter the sensor chamber from the oral cavity. 8. A dental appliance in accordance with claim 1, wherein the abutment device is adapted to be coupled to a jawbone. 9. A dental appliance in accordance with claim 8, including a jawbone anchoring member that is adapted to be coupled to the jawbone, the abutment device removably coupled to the jawbone anchoring member and extending outwardly from the jawbone anchoring member and into the oral cavity. 10. A dental appliance for use in positioning a sensing device within an oral cavity, comprising: an abutment device adapted to be coupled to a jawbone;a sensor assembly including:a sensor support assembly extending outwardly from an outer surface of the abutment device, the sensor support a... | KLAB LLC | A61C005/70 A61B005/00 A61C008/00 A61B005/145
US09795316 | B2 | 2017-10-24 | Radiolucent electrode or sensor assembly | 2015-12-14 | A lightweight, disposable and substantially radiolucent electrode or sensor assembly for that universally connects to separate, non-integrated electrodes or sensors for the monitoring of the physiological parameters of a live subject wherein the electrode assembly is comprised of one or more radiolucent electrical connectors for connecting the electrode assembly to the sensors. The present invention also discloses a method of positioning the electrode assembly on a patient whose physiological signs are being monitored such that access to the patient's chest is substantially unimpeded so as not to obstruct the electromagnetic imaging of the patient's chest, the application of defibrillation paddles or surgical procedures that require access to the chest area. | 1. A method of positioning a substantially radiolucent electrode assembly on a patient that is being monitored for physiological signs such that access to the patient's chest is substantially unimpeded, comprising: applying the electrode assembly around the outer periphery of the patient's torso and proceeding around the back of the patient's neck, wherein the electrode assembly does not cross the patient's chest. 2. The method of positioning a substantially radiolucent electrode assembly on a patient that is being monitored for physiological signs such that access to the patient's chest is substantially unimpeded of claim 1 wherein the physiological sign is an ECG or EEG. 3. The method of positioning a substantially radiolucent electrode assembly on a patient that is being monitored for physiological signs such that access to the patient's chest is substantially unimpeded of claim 1 wherein the electrode assembly comprises a plurality of electrical connectors wherein each electrical connector is in electrical contact with a physiological sensor. 4. The method of positioning a substantially radiolucent electrode assembly on a patient that is being monitored for physiological signs such that access to the patient's chest is substantially unimpeded of claim 1 wherein a portion of the electrode assembly is routed around the back side of a patient. 5. The method of positioning a substantially radiolucent electrode assembly on a patient that is being monitored for physiological signs such that access to the patient's chest is substantially unimpeded of claim 4 wherein a portion of the electrode is routed around a patient's neck. 6. The method of positioning a substantially radiolucent-electrode assembly on a patient that is being monitored for physiological signs such that access to the patient's chest is substantially unimpeded of claim 1 wherein said electrode assembly comprises: a flexible radiolucent non-conductive substrate having a proximal end and a distal end;at least one aperture formed in the substrate and sized to receive a sensor; anda conductive trace comprising a substantially radiolucent ink printed on the substrate for conducting a current between a sensor and the distal end of the substrate. 7. The method of positioning a substantially radiolucent-electrode assembly on a patient that is being monitored for physiological signs such that access to the patient's chest is substantially unimpeded of claim 6 wherein the conductive trace comprises at least one linear conductor. 8. The method of positioning a substantially radiolucent-electrode assembly on a patient that is being monitored for physiological signs such that access to the patient's chest is substantially unimpeded of claim 6 wherein the conductive trace comprises a plurality of substantially linear conductors. 9. The method of positioning a substantially radiolucent-electrode assembly on a patient that is being monitored for physiological signs such that access to the p... | - | A61B005/04 A61B005/0428 A61B005/0408 A61B005/0424 A61B005/0478 A61B005/00 A61N001/00
US09826904 | B2 | 2017-11-28 | System and method for detecting tissue surface properties | 2014-03-20 | A system and method for detecting a tissue property. The system comprises a first unit positioned outside a patient body and a second unit positioned inside the patient's body. The first unit includes a first housing, and a magnetic field source supported by the first housing. The second unit includes a second housing, a pressure sensor supported by the second housing, a localization module supported by the second housing, a controller, and a power source. The pressure sensor is configured to detect an indentation force applied to the tissue, and the second unit is configured to wirelessly transmit the indentation force data and localization data to a computer to generate a volumetric stiffness map for the tissue. | 1. A system for detecting a tissue property, the system including: a first unit positioned outside a patient body, the first unit including a first housing, anda first magnetic field source supported by the first housing; anda second unit positioned inside a patient body, the second unit including a second housing,a second magnetic field source supported by the second housing,a pressure sensor supported by the second housing,a localization module supported by the second housing,a controller, anda power source;wherein the pressure sensor is configured to detect an indentation force applied to the tissue, wherein the indentation force is applied when the first magnetic field source and the second magnetic field source are magnetically coupled;wherein the second unit is configured to wirelessly transmit the indentation force data and localization data to a computer to generate a volumetric stiffness map for the tissue. 2. The system of claim 1, wherein the localization module comprises a plurality of magnetic field sensors and an inertial sensor. 3. The system of claim 2, wherein the magnetic field sensors and the inertial sensor track a depth of the indentation applied to the tissue and a position of the second housing relative to the first magnetic field source. 4. The system of claim 2, wherein the magnetic field sensors comprise Hall effect sensors, and wherein at least three of the Hall effect sensors are mounted on three surfaces orthogonal to each other. 5. The system of claim 2, wherein the magnetic field sensor detects a magnetic field generated by the magnetic field source to determine a location and orientation of the second housing with respect to the magnetic field source. 6. The system of claim 1, wherein the localization data comprises data related to a position of the second unit when the indentation force is applied. 7. The system of claim 1, wherein the second housing comprises a cylindrical-shaped body having a first end and a second end, and wherein the pressure sensor is adjacent the first end. 8. The system of claim 7, wherein the second housing further comprises a grasping site for a tool to grip the second housing. 9. The system of claim 8, wherein the grasping site comprises a reduced diameter area on the second housing closer to the first end than the second end. 10. The system of claim 1 further comprising a computer system in communication with the controller, and wherein the computer system is configured to receive signals transmitted by the controller. 11. The system of claim 10, wherein the computer system further comprises a processor and a software program configured to process the signals from the controller to generate the volumetric stiffness map. 12. The system of claim 1, wherein the first unit further comprises an inertial sensor supported by the first housing. 13. A device insertable through an incision in a human body to determine tissue stiffness, the device comprising: an enclosed housing having... | Vanderbilt University | A61B005/00 A61B005/05 A61B005/11 A61B034/00 A61B017/00 A61B090/00
US09782097 | B2 | 2017-10-10 | Electrocardiograph monitoring device and connector | 2014-04-18 | The present invention relates to electrocardiography and to electrode arrangements used in electrocardiographic monitoring devices, and is more particularly related to a pad or patch containing said electrodes which may be used to passively and non-invasively monitor electrical activity generated by a patient's heart from the surface of that patient's chest, and to a connector which allows for fast and simple connection between the pad containing said electrodes and the devices and equipment typically used to monitor and view electrocardiographic information. | 1. An EKG patch system, comprising: an EKG patch having a base layer and a plurality of EKG patch electrodes protruding from the base layer, said plurality of EKG patch electrodes being capable of receiving an electrical signal generated by the myocardium of a human heart;a connector section having a first sheet comprising a plurality of connector section electrodes to match the number and location of the electrodes in the EKG patch and a second sheet comprising a plurality of electrical leads, a perimeter of the first sheet being substantially equal to a perimeter of the base layer of said EKG patch, the connector section having a shape that corresponds to the EKG patch; anda single connector cable electrically connected to said plurality of electrical leads at a first end through a single port in said connector section, said single connector cable being suitable to be electrically connected to an EKG monitoring device at a second end, and the single connector cable being detachable partway along its length with a male and a female connection,wherein said base layer is configured to contact skin of a human patient and attach thereto,wherein said plurality of EKG patch electrodes are held in a fixed configuration with respect to said electrical leads of said connector section and are also configured to contact the skin of a human patient concurrently with said base layer, andwherein said EKG patch is constructed from nonmagnetic materials such that said EKG patch is suitable for use in conjunction with magnetic fields. 2. The EKG patch system according to claim 1, wherein the plurality of EKG patch electrodes and connector section electrodes are snap electrodes, via which the EKG patch and the connector section are attachable. 3. An EKG-patch connector system, comprising: an EKG patch comprising a base layer and a plurality of EKG patch electrodes protruding from the base layer, said plurality of EKG patch electrodes being configured to contact skin of a human patient concurrently with said base layer;a connector section adapted to be electrically connected to the EKG patch, the connector section having a first sheet comprising a plurality of connector section electrodes to match the number and location of the electrodes in the EKG patch and a second sheet comprising a plurality of electrical leads, a perimeter of the first sheet being substantially equal to a perimeter of the base layer of said EKG patch, the connector section having a shape corresponds to the EKG patch; anda single connector cable electrically connected to said plurality of electrical leads at a first end through a single port in said connector section, said single connector cable also being suitable to be electrically connected to an EKG monitoring device at a second end, and the single connector cable being detachable partway along its length with a male and a female connection,wherein said plurality of EKG patch electrodes are held in a fixed configuration with respect to ... | CARDIAC LEAD TECHNOLOGIES, LLC | A61B005/0408 A61B005/0416 A61B005/0428 A61B005/00 A61N001/04
US09775550 | B2 | 2017-10-03 | Cord blood collection vessel and method of using same | 2015-05-12 | An apparatus and method for collecting blood from the umbilical cord of a newborn into a standard vacuum collection tube is disclosed wherein the apparatus comprises a vessel in the general shape of a cup with an open top, a receiving cavity and a tapered bottom surface wherein the taper of the tapered bottom surface is away from the receiving cavity; a pick up tube with open top end and closed bottom, and side tabs, a needle inserted through an opening in the closed bottom of said tube wherein the diameter of the opening is of sufficient size to only allow the tip of the needle to be inserted, and a means for providing an air tight seal between the inserted needle and said opening. | 1. An apparatus for collecting umbilical cord blood comprising: a vessel with an open top, a receiving cavity positioned along an edge of the vessel at the open top of the vessel, and a tapered bottom surface, wherein the taper of the tapered bottom surface is away from the receiving cavity to allow blood within the vessel to flow and collect within a recess along a side of the vessel directly opposite the receiving cavity;a pick up tube with an open top end and a closed bottom, a connecting member in the form of side tabs for engaging the receiving cavity to allow for selective engagement with the vessel wherein the length of the pick up tube is sufficient to allow the closed bottom of the pick up tube to rest within the recess against the side of the vessel opposite the receiving cavity while the open top of the pick up tube rests in the receiving cavity such that the pick up tube extends diagonally across the vessel extending from the receiving cavity along the edge of the vessel to the recess at the bottom surface of the vessel, a needle inserted through an opening in the closed bottom of said pick up tube, and a means for providing an air tight seal between the inserted needle and said opening, wherein the pick up tube is separate and distinct from the vessel so as to be separable from the vessel. 2. The apparatus of claim 1 wherein the side tabs are a clip that attaches to the receiving cavity. 3. The apparatus of claim 1 wherein the diameter of the pick up tube is sufficient to allow insertion of a vacuum collection tube. 4. A method for collecting a sample of umbilical cord blood comprising the steps of: holding a section of said cord over an apparatus for collecting umbilical cord blood comprising a vessel with an open top, a receiving cavity positioned along an edge of the vessel at the open end of the vessel and a tapered bottom surface wherein the taper of the tapered bottom surface is away from the receiving cavity and the bottom surface includes a recess along a side of the vessel directly opposite the receiving cavity;allowing blood to drip from said cord into the open top of said apparatus wherein the taper of the tapered bottom surface is away from the receiving cavity to allow blood within the vessel to flow and collect in the recess along the side of the vessel directly opposite the receiving cavity;placing a pick up tube with an open top end and a closed bottom, a connecting member in the form of side tabs for engaging the receiving cavity to allow for selective engagement with the vessel, a needle inserted through an opening in the closed bottom of said pick up tube, and a means for providing an air tight seal between the inserted needle and said opening into said vessel such that the pick up tube extends diagonally across the vessel extending from the receiving cavity along the edge of the vessel to the recess at the bottom surface of the vessel the length of the pick up tube being sufficient to allow the closed bottom of... | - | A61B005/00 B65D081/00 A61B005/15 B01L003/00 A61B010/00
US09795442 | B2 | 2017-10-24 | Ablation catheters | 2016-10-31 | Cardiac tissue ablation catheters including an inflatable and flexible toroidal or spherically shaped balloon disposed at a distal region of an elongate member, a flexible circuit carried by an outer surface of the balloon, the flexible circuit including, a plurality of flexible branches conforming to the radially outer surface of the balloon, each of the plurality of flexible branches including a substrate, a conductive trace carried by the substrate, and an ablation electrode carried by the substrate, the ablation electrode in electrical communication with the conductive trace, and an elongate shaft comprising a guidewire lumen extending in the elongate member and extending from a proximal region of the inflatable balloon to distal region of the inflatable balloon and being disposed within the inflatable balloon, wherein a distal region of the elongate shaft is secured directly or indirectly to the distal region of the inflatable balloon. | 1. A cardiac tissue ablation catheter, comprising: an inflatable and flexible toroidal or spherically shaped balloon disposed at a distal region of an elongate member;a plurality of flexible substrate branches attached to an outer surface of the balloon along the flexible substrate branches, wherein each of the plurality of flexible substrate branches has a top surface and a bottom surface, and a plurality of holes each extending through the top and bottom surfaces, the plurality of holes located at a periphery of the respective flexible substrate branch adjacent an outer edge of the respective flexible substrate branch, wherein the plurality of holes do not extend through the balloon and do not extend through the electrode, and wherein the plurality of holes have an adhesive therein;a plurality of ablation electrodes, each one of the plurality of ablation electrodes carried by one of the plurality of flexible substrate branches;a plurality of conductive traces, each of the plurality of conductive traces in electrical communication with one of the plurality of ablation electrodes; andan elongate shaft comprising a guidewire lumen extending in the elongate member and extending from a proximal region of the inflatable balloon to a distal region of the inflatable balloon and being disposed within the inflatable balloon, wherein a distal region of the elongate shaft is secured directly or indirectly to the distal region of the inflatable balloon,the toroidal or spherically shaped inflatable balloon having, in a side view, a location with the greatest orthogonal linear dimension measured orthogonal to the longitudinal axis of the elongate shaft,wherein each of the plurality of flexible substrate branches extends proximal to the location, over the location, and distal to the location. 2. The catheter of claim 1 wherein each of the plurality of ablation electrodes has more surface area distal to the location than proximal to the location. 3. The catheter of claim 1 wherein at least one of the balloon and the plurality of flexible substrate branches include a plurality of irrigation holes therethrough to allow irrigation from inside the balloon to outside the balloon. 4. The catheter of claim 3 wherein at least some of the plurality of irrigation holes surround at least a portion of the border of the plurality of ablation electrodes. 5. The catheter of claim 1, wherein, in a side view, proximal ends of all of the plurality of ablation electrodes are disposed in a first plane orthogonal to the longitudinal axis of the elongate shaft, and distal ends of all of the plurality of ablation electrodes are disposed in a second plane orthogonal to the longitudinal axis of the elongate shaft, the second plane being different than the first plane. 6. The catheter of claim 5, wherein, in the side view, the catheter only has a single row of ablation electrodes, and wherein the plurality of ablation electrodes are the single row of electrodes. 7. The catheter of... | Shifamed Holdings, LLC | A61B018/12 A61B018/14 A61B005/042 A61B005/00 A61M025/10 A61B001/00 A61B001/06 A61B018/00 A61N005/02 A61N001/05 A61B005/01 A61B018/18 A61M025/01 A61M025/04 A61B017/00 A61B017/22 A61B018/02 A61M025/06 A61B090/00
US09775524 | B2 | 2017-10-03 | Apparatus and method of assessing a narrowing in a fluid filled tube | 2015-04-30 | An apparatus and method of assessing a narrowing in a fluid filled tube having a fluid flow pressure wave having a backward-originating pressure component and a forward-originating pressure component without taking a flow velocity measurement, comprising: taking pressure measurements in the tube; separating the pressure components into the backward-originating pressure component and the forward-originating pressure component; identifying a time window when the differential of flow velocity (dU) is minimal or absent; and deriving the backward and forward pressure components for pressure measurements taken in at least the time window. | 1. A method of assessing a narrowing in a blood vessel, the method comprising: positioning at least one pressure-sensing probe within a blood vessel;obtaining, with an intravascular processing system having a processor in communication with the at least one pressure-sensing probe, pressure measurements from within the blood vessel using the at least one pressure-sensing probe, the pressure measurements being obtained not during hyperaemia;identifying, with the intravascular processing system, a wave free period corresponding to a time window when a differential of flow velocity is minimal or absent; andcalculating, with the intravascular processing system, a pressure ratio using the pressure measurements obtained during the wave free period to provide an assessment of a severity of the narrowing in the blood vessel. 2. The method of claim 1, wherein the wave free period is identified heuristically. 3. The method of claim 1, wherein the wave free period is identified based on a characteristic of a pressure waveform of the obtained pressure measurements. 4. The method of claim 1, wherein the wave free period is identified as a certain time window before or after a characteristic of the pressure waveform. 5. The method of claim 1, wherein the wave free period is identified by: identifying a peak pressure time (tPmax);identifying an end of pressure waveform time (tPend); andspecifying the wave free period as a time window between tPmax and tPend. 6. The method of claim 5, wherein the time window extends from tPmax+150 ms to tPend−50 ms. 7. The method of claim 5, wherein the time window is a mid 3/5 window between tPmax and tPend. 8. The method of claim 1, wherein the wave free period is identified by: calculating a standard deviation of a pressure ratio of the obtained pressure measurements; andspecifying the wave free period as a time window after a peak pressure time point where the standard deviation is below a predetermined threshold. 9. The method of claim 8, wherein the predetermined threshold is a bottom 10% or less of the calculated standard deviation. 10. The method of claim 8, wherein the time window has a length of at least 100 ms. 11. A system of assessing a narrowing in a blood vessel, the system comprising: at least one pressure-sensing probe sized and shaped for positioning within the blood vessel; anda processor in communication, with the at least one pressure-sensing probe, the processor configured to: receive pressure measurements obtained by the at least one pressure-sensing probe positioned within the blood vessel, the pressure measurements being obtained not during hyperaemia;identify a wave free period corresponding to a time window when a differential of flow velocity is minimal or absent; andcalculate a pressure ratio using the pressure measurements obtained during the wave free period to provide an assessment of a severity of the narrowing in the blood vessel. 12. The system of claim 11, wherein the processor is co... | Medsolve Limited Imperial College of Science, Technology and Medicine | G06F017/00 A61B005/02 A61B005/0215 A61B005/0285 A61B005/021 A61B005/00 G06F011/30
US09827429 | B2 | 2017-11-28 | Medical implant with contact portions that converts movement from the contact portions into an electrical signal | 2015-07-17 | A medical implant including an implant body for insertion into a human and/or animal body. The implant body includes at least one first and at least one second contact portion, wherein the at least one first and the at least one second contact portions contact two tissue regions performing a relative movement with respect to one another. The at least one first and the at least one second contact portions are movable relative to one another, wherein a relative movement of the contact portions may be converted into an electrical signal. | 1. A medical implant for insertion into the human and/or animal body, comprising: an implant body, wherein the implant body comprises at least one first and at least one second contact portion;wherein the at least one first and the at least one second contact portions are configured to contact two tissue regions performing a relative movement with respect to one another,wherein the at least one first and the at least one second contact portions are movable relative to one another, and a relative movement of the at least one first and the at least one second contact portions is converted into electrical energy,wherein one or more of the at least one first contact portion and the at least one second contact portion is coupled directly or indirectly to an electric generator in or on the implant body,wherein the electric generator comprises a magnet and a coil,wherein the at least one second contact portion is coupled directly to the electric generator and moves the magnet of the electric generator relative to the coil when the two tissue regions move relative to one another,wherein the magnet moves laterally relative to the coil when the two tissue regions move laterally relative to one another as a lateral relative movement,wherein the magnet is electromagnetically connected to the coil and induces an electric voltage in the coil to operate the medical implant,wherein the implant body further comprises an electric store, and,wherein the electrical energy is stored in the electric store and called up as necessary. 2. The medical implant as claimed in claim 1, wherein the at least one first contact portion comprises a fixing element to fasten the implant body to one of the two tissue regions. 3. The medical implant as claimed in claim 1, wherein one or more of the at least one first contact portion and the at least one second contact portion is mechanically connected to a movable element of the electric generator provided in or on the implant body. 4. The medical implant as claimed in claim 1, wherein the implant body comprises a number of parts, wherein at least two parts of the number of parts are movable relative to one another and each of the number of parts of the implant body is associated with a contact portion. 5. The medical implant as claimed in claim 4, wherein the number of parts are coupled using at least one piezoelectric element. 6. The medical implant as claimed in claim 4, further comprising a mechanical oscillation generator arranged between the number of parts of the implant body. 7. The medical implant as claimed in claim 1, wherein the electric generator is an electrostatic generator, and wherein the electric generator is driven via mechanical waves. 8. The medical implant as claimed in claim 4, wherein the at least one second contact portion comprises a fixing element to fasten the implant body to one of the two tissue regions. 9. The medical implant as claimed in claim 1, further comprising a control and/or regulation u... | Biotronik SE & Co. KG | A61N001/378 A61B005/00 A61N001/02 A61N001/362 A61B005/07 A61N001/375 A61N001/05 H02N002/18
US09804092 | B2 | 2017-10-31 | Integrated spectral probe for raman, reflectance and fluorescence spectral measurements | 2014-02-11 | Examples of a spectroscopy probe for performing measurements of Raman spectra, reflectance spectra and fluorescence spectra are disclosed. The integrated spectral probe can comprise one or more light sources to provide a white light illumination to generate reflectance spectra, an excitation light to generate an UV/visible fluorescence spectra and a narrow band NIR excitation to induce Raman spectra. The multiple modalities of spectral measurements can be performed within 2 seconds or less. Examples of methods of operating the integrated spectroscopy probe disclosed. | 1. An integrated spectroscopy system to facilitate integrated measurements of Raman spectra, reflectance spectra and fluorescence spectra, the system comprising: a light source configured to provide a white light illumination beam to generate reflectance spectra, an excitation light beam to generate an UV/visible fluorescence spectra and a narrow band excitation light beam to induce Raman spectra;an integrated spectroscopy probe coupled to the light source, the spectroscopy probe comprising:at least one narrow band excitation fiber configured to transmit the narrow band excitation light beam to a target;at least one white light illumination fiber configured to transmit the white light illumination beam to the target;at least one excitation illumination fiber configured to transmit the excitation light beam to the target;a plurality of Raman collection fibers configured to collect a returning radiation reflected and emitted from the target in response to the narrow band excitation light beam;at least one reflectance collection fiber configured to collect a returning radiation reflected from the target in response to the white light illumination beam; andat least one fluorescence collection fiber configured to collect returned radiation emitted from the target in response to the excitation light beam; a first detector coupled to the plurality of Raman collection fibers to detect a Raman spectrum; andat least one additional detector coupled to the at least one reflectance collection fibers and/or the at least one fluorescence collection fibers to detect a reflectance spectrum and/or a fluorescence spectrum,wherein the at least one white light illumination fiber, the at least one excitation illumination fiber, the at least one reflectance collection fiber and the at least one fluorescence collection fiber are placed between the at least one narrow band excitation fiber and the plurality of Raman collection fibers and are each spaced apart from the at least one narrow band excitation fiber and the plurality of Raman collection fibers. 2. The integrated spectroscopy system of claim 1, wherein pluralities of spectral measurements are being performed within 2 seconds or less. 3. The integrated spectroscopy system of claim 1, further comprise an optical modulator mounted in front of the light source configured to modulate the illumination light from the light source and provide the white light illumination light beam, the excitation light beam and the narrow band excitation beam. 4. The integrated spectroscopy system of claim 3, wherein the optical modulator comprises a first filter configured to pass visible light below 750 nm, a second filter configured to pass excitation light in a blue wavelength range and a third filter configured to pass light in a 785±5 nm wavelength range. 5. The integrated spectroscopy system of claim 1, wherein the light source comprises a monochromatic light source configured to provide the narrow band excitation light beam... | - | G01N021/65 G01N021/64 G01J003/44 A61B005/00 G01J003/02
US09828060 | B2 | 2017-11-28 | Automated e-assist adjustment to prevent user perspiration | 2015-10-13 | A method may include determining whether a sweat sensor is available. E-assist may be provided when a sweat sensor is not available by determining a sweat threshold and determining whether the sweat threshold is surpassed. When a sweat sensor is available, e-assist may be provided when the sweat sensor senses sweat. | 1. A method for assisting propulsion of a cycle adapted to carry a rider, the method comprising operating a controller configured to: determine whether a sweat sensor is available as indicated by whether the sweat sensor is actively communicating with the controller, wherein the sweat sensor is configured to determine whether a first sweat threshold is surpassed by sensing rider sweat;when the sweat sensor is not actively communicating with the controller and is therefore not available, determine a second sweat threshold by categorizing, in the controller, user input setting options and an environmental data set for each user input setting option wherein the categorized user input setting options and environmental data sets provide an indication of an exertion level of the rider where a propulsion assistance to the cycle is needed to avoid sweating by the rider as the second sweat threshold;read a user input device to obtain a selected user input setting selected by the rider;read a sensor to obtain sensed environmental data indicative of current conditions; anddetermine whether the second sweat threshold is surpassed by comparing the selected user input setting and the sensed environmental data to the user input setting options and environmental data sets; andwhen the second sweat threshold is surpassed, operate a motor to provide the propulsion assistance to the cycle when the sweat sensor is not available. 2. The method according to claim 1, wherein the sweat sensor is configured to be worn by the rider and senses the presence of moisture, and further comprising providing the propulsion assistance when the sweat sensor is actively communicating with the controller and therefore available by reading, by the controller, the sweat sensor; and determining, by the controller, whether the sweat sensor senses the presence of moisture. 3. The method according to claim 2 further comprising receiving, by the controller, the selected user input setting from the input device that is in communication with the controller, wherein the input device is in the form of an input wristband that is configured to receive the selected user input setting from the rider; and effecting communication between the input device and the controller, wherein the step of determining the second sweat threshold comprises reading, by the controller, the selected user input setting from the input wristband. 4. The method according to claim 3 comprising incorporating the sweat sensor into the input device. 5. The method according to claim 1 wherein determining the second sweat threshold comprises looking up, by the controller, the second sweat threshold in a table stored in the controller and based on the user input setting options and the environmental data sets. 6. The method according to claim 1 wherein when it is determined that the second sweat threshold is surpassed the method further comprises repeating determining whether the second sweat threshold is surpassed after pr... | GM GLOBAL TECHNOLOGY OPERATIONS LLC | B62M006/50 A61B005/00 B60W040/08 B60W050/10 A61B005/024
US09805467 | B2 | 2017-10-31 | Lung measurements | 2014-06-19 | The present invention relates to lung measurement. In order to provide enhanced information about a patient that facilitates further assessment steps, 2D X-ray image data of a patient's chest is provided, and the image data is segmented to identify lung structures to provide segmented image data separated from un-segmented areas. Further, spatial lung volume information is extracted from the image data using the segmented image data derived from the image data. Still further, lungs symmetry information is determined using the extracted spatial lung volume information. Finally, the lungs symmetry information is provided to a user. For example, a 2D X-ray image data of a patient's chest is provided (84) and a lungs mask image is formed (86) after the step of segmenting the input image data. Then, the lungs mask image is used to define areas, within which a predetermined adaptation is applied (88) to the original 2D X-ray image data producing a thorax mask image. Next, left and right images are provided (90) showing the left and the right spatial lungs volume information of the regions defined originally by the lungs mask image. Finally, based on the spatial lungs volume information, lungs symmetry information or total lung volumes may be calculated and provided (92). | 1. A method for providing lungs symmetry information, comprising the following steps: a) providing 2D X-ray image data of a patient's chest;b) segmenting the image data to provide segmented image data separated from un-segmented areas, the segmented image data comprising a lungs mask image having first and second selected regions corresponding to lung structures;c) extracting spatial lung volume information from the image data using the segmented image data derived from the image data;d) determining lungs symmetry information using the extracted spatial lung volume information; ande) providing the lungs symmetry information to a user,wherein the extracting step c) further comprises:c1) generating a thorax mask image wherein, for the first and second selected regions, the image data is replaced by applying a pixel value representing soft tissue;c2) subtracting the thorax mask image from the 2D X-ray image data to achieve a lungs-only image comprising the spatial lung volume information, andwherein the determining step d) further comprises:d1) summing the spatial lung volume information of the first and second selected regions. 2. Method according to claim 1, wherein in step a) the 2D X-ray image data is provided as image data acquired in: a1) a posterior-anterior or an anterior-posterior direction; ora2) in an oblique direction provided in an angular alignment to the posterior-anterior or anterior-posterior direction, wherein the image data comprises separate image regions of the right and the left lung respectively. 3. Method according to claim 1, wherein the step b) further comprises: bi) identifying left and right lung of the lungs; orbii) identifying a lateral boundary of the lungs enclosing the left and right lungs; and wherein the extraction in step c) is provided based on:ci) the left and the right lung; orcii) the lateral boundary of the lungs. 4. Method according to claim 1, wherein step d) further comprises:d2) applying a conversion factor to the summed spatial lungs volume information to obtain left and right lung volume representations;d3) calculating a total representation of the left and right lung volume representations; andd4) generating the lungs symmetry information using the left and right lung volume representations and the total representation. 5. Method according to claim 1, wherein: step c1) further comprises the step of interpolating over the image data corresponding to the first and second selected regions of the lungs mask image to remove lungs information, producing a first area (Aright) and a second area (Aleft) of the thorax, wherein the first and second areas are located within the thorax mask image; andstep c2) further comprises the step of subtracting the thorax mask image from the image data to produce first spatial lung volume data (Lright) and second spatial lung volume data (Lleft). 6. Method according to claim 1, wherein the conversion factor is a scale using an inverse of a half value layer. 7. A lungs m... | KONINKLIJKE PHILIPS N.V. | G06K009/00 G06T007/00 G06K009/46 G06T007/68 A61B005/08
US09806536 | B2 | 2017-10-31 | Method and apparatus for wireless magnetic power transmission | 2013-10-11 | Systems for wirelessly transmitting power to an implanted medical device. The wireless transmission system including a first and second transmitting coil both the first and second coil having substantially equal diameters and at least one conductor winding. A gap between the first transmitting second transmitting coil extending along a common axis by a distance equal to the radius of the first transmitting coil. A plurality of capacitors connected in series along the at least one conductor of the transmitting coils to divide the transmitting coils into a plurality of coil segments. An input connection is electronically coupled to the transmitting coils to deliver an excitation voltage to the transmitting coils to produce a substantially uniform magnetic field between the first transmitting coil and the second transmitting coil. | 1. A system for wirelessly transferring energy to an implanted medical device, the system comprising: a wireless power transmitting device, the wireless power transmitting device including: a first transmitting coil having at least one conductor winding, a first radius and a first diameter;a second transmitting coil having at least one conductor winding, a second radius, and a second diameter and arranged to extend along a common axis with the first transmitting coil, wherein the first diameter and the second diameter are substantially equal;a gap arranged between the first transmitting coil and the second transmitting coil and extending along the common axis a distance equal to the first radius of the first transmitting coil;a first plurality of capacitors connected in series along the at least one conductor winding of the first transmitting coil to divide the coil conductor winding into a first plurality of coil segments;a second plurality of capacitors connected in series along the at least one conductor winding of the second transmitting coil to divide the coil conductor winding into a second plurality of coil segments; andan input connection electrically coupled to the first transmitting coil and the second transmitting coil to deliver an excitation voltage to the first transmitting coil and the second transmitting coil substantially equal to a theoretical excitation voltage required to produce a substantially uniform magnetic field between the first transmitting coil and the second transmitting coil divided by a sum of the first plurality of coil segments and the second plurality of coil segments. 2. The system of claim 1, wherein the first transmitting coil and the second transmitting coil have a resonant frequency based on an operating frequency of less than 15 MHz. 3. The system of claim 1, wherein a number of the first plurality of capacitors and a number of the second plurality of capacitors are equal. 4. The system of claim 3, wherein the first plurality of coil segments and the second plurality of coil segments each include at least four equal coil segments. 5. The system of claim 1, wherein a voltage across each of the first plurality of capacitors and the second plurality of capacitors is equal and in opposite polarity to a voltage across each of the first plurality of coil segments and the second plurality of coil segments to substantially cancel the voltage across each of the first plurality of coil segments and second plurality of coil segments. 6. The system of claim 1, wherein a value of one of the first and second plurality of capacitors is determined by:f=12⁢π⁢LTX⁢CTX;where LTX is an inductance of the first and second transmitting coil, CTX is a resultant capacitance of the first and second plurality of capacitors arranged in series, and f is a resonance frequency of the wireless power transmitting device. 7. The system of claim 1, further comprising a power source coupled to the input connection and configured to deli... | THE BRIGHAM AND WOMEN'S HOSPITAL, INC. | H02J005/00 A61N001/378 A61N001/372 H01F038/14 H02J007/02 A61B001/00 A61B005/026
US09820657 | B2 | 2017-11-21 | Mobile wireless appliance | 2015-05-20 | A wireless system for a person includes a wearable appliance with an accelerometer; a wireless device in communication with the wearable appliance; and a remote computer coupled to the wireless device to provide information to an authorized remote user. | 1. A wearable appliance comprising: a processor;an accelerometer in communication with the processor to capture movement data; a microphone in communication with the processor to receive audio input;a speaker in communication with the processor to output audio;a heart rate sensor in communication with the processor to capture heart rate data;a transceiver in communication with the processor to transmit and receive data via a wireless network, the transceiver further in communication with a plurality of wireless nodes forming a mesh network, wherein at least one of the plurality of wireless network nodes comprises a home security device; anda display in communication with the processor to display at least: (i) information associated with the heart rate data; and (ii) information associated with data received by the transceiver. 2. The wearable appliance of claim 1 further comprising at least one of a temperature sensor, blood oxygen sensor and light sensor. 3. The wearable appliance of claim 1 further comprising software code to generate information regarding steps taken using the movement data captured by the accelerometer. 4. The wearable appliance of claim 1 further comprising software code to generate information regarding steps taken using the movement data captured by the accelerometer. 5. The wearable appliance of claim 1 further comprising a vibration motor in communication with the processor. 6. The wearable appliance of claim 5 further comprising software code executed by the processor to activate the vibration motor in response to receipt of data by the transceiver. 7. The wearable appliance of claim 1 further comprising software code to record audio input received by the microphone, and software code to output audio associated with the data received by the transceiver via the wireless network. 8. The wearable appliance of claim 7 further comprising software code to transmit, using the transceiver, the recorded audio input via the wireless network. 9. The wearable appliance of claim 1 further comprising at least one of a gyroscope and a compass. 10. A wearable appliance comprising: a processor; an accelerometer in communication with the processor to capture movement data; a heart rate sensor in communication with the processor to capture heart rate data; a vibration motor in communication with the processor;a transceiver in communication with the processor to receive data via a wireless network, the transceiver further in communication with a plurality of wireless nodes forming a mesh network, wherein at least one of the plurality of wireless network nodes comprises a home security device;a display in communication with the processor to display at least: (i) informaion associated with data received by the transceiver; and (ii) information associated with the heart rate data. 11. The wearable appliance of claim 10 wherein the transceiver comprises a transceiver in communication with the processor to receive at least telephonesa... | KONINKLIJKE PHILIPS N.V. | A61B005/04 A61B005/0205 A61B005/00 A61B005/021 A61B005/11 A61B005/145 A61B008/02 A61B008/06 A61B008/00 H04M003/42 H04Q009/14 A61B005/024 A61B008/08
US09804066 | B2 | 2017-10-31 | Simultaneous acquisition of biometric data and nucleic acid | 2015-06-15 | Systems, methods, and kits are disclosed for collection, labeling and analyzing biological samples containing nucleic acid in conjunction with collecting at least one ridge and valley signature of an individual. Such devices and methods are used in forensic, human identification, access control and screening technologies to rapidly process an individual's identity or determine the identity of an individual. | 1. A method for collection of a biological sample comprising at least one nucleic acid and at least one ridge and valley signature of an individual, comprising: a. providing at least a first imaging component configured to provide an energy wave and comprising a scanning surface configured to permit the energy wave to penetrate the scanning surface;b. providing a substrate configured to collect the biological sample from an appendage of the individual wherein the substrate is chemically modified with a chemical functional group selected from the group comprising an amino group, a chloromethyl group, a hydroxyl group, a carboxyl group and a quaternary amino group;c. positioning the appendage of the individual upon the scanning surface and in contact with the substrate;d. collecting the at least one ridge and valley signature from the appendage imaged by the energy wave; ande. withdrawing the appendage from the scanning surface, thereby collecting the biological sample upon the substrate. 2. The method of claim 1, further comprising the step of transmitting the at least one ridge and valley signature of the individual to at least one database comprising ridge and valley signatures selected from the group consisting of a forensic, criminal, identity, child identity, missing persons, immigration, department of motor vehicles, terrorist, paternity, arrestee, convict database, access control or any combination thereof. 3. The method of claim 1, further comprising the step of providing an identifier to associate the biological sample of the individual with the at least one ridge and valley signature of the individual, wherein the identifier is alphanumeric, graphical, magnetic, or electromagnetic. 4. The method of claim 3, wherein the identifier is a barcode. 5. The method of claim 1, further comprising the step of subjecting the biological sample to at least one of a polymerase chain reaction, a DNA sequencing reaction, STR analysis, SNP analysis, or Indel analysis. 6. The method of claim 1, further comprising the step of archiving the substrate containing the biological sample. 7. The method of claim 1, wherein the at least first imaging component is an optical scanner or a capacitance scanner. 8. The method of claim 7, wherein the optical scanner comprises a LED, laser diode, incandescent light source, or a multispectral imager. 9. The method of claim 1, wherein the scanning surface is formed from a material selected from a plastic, a polyolefin, a polystyrene, a metal, a metal alloy, a glass, a silicon or combinations thereof, optionally wherein the material of the scanning surface is transparent or translucent. 10. The method of claim 1, wherein the substrate comprises a gel. 11. The method of claim 1, wherein the substrate further comprises a support. 12. The method of claim 11, wherein the support is formed from a material selected from a plastic, a polyolefin, a polystyrene, a metal, a metal alloy, a glass, a silicon or combinations t... | LIFE TECHNOLOGIES CORPORATION | C12Q001/68 C12M001/00 C12M001/34 G01N001/30 G06T007/00 A61B005/15 A61B005/151 A61B005/00 A61B005/1172 A61B005/145 G06K009/00 A61B010/00 A61B010/02 A61B005/157
US09811730 | B2 | 2017-11-07 | Person identification using ocular biometrics with liveness detection | 2015-07-13 | A method of assessing the identity of a person by one or more of: internal non-visible anatomical structure of an eye represented by the Oculomotor Plant Characteristics (OPC), brain performance represented by the Complex Eye Movement patterns (CEM), iris patterns, and periocular information. In some embodiments, a method of making a biometric assessment includes measuring eye movement of a subject, making an assessment of whether the subject is alive based on the measured eye movement, and assessing a person's identity based at least in part on the assessment of whether the subject is alive. In some embodiments, a method of making a biometric assessment includes measuring eye movement of a subject, assessing characteristics from the measured eye movement, and assessing a state of the subject based on the assessed characteristics. | 1. A method of making a biometric assessment, comprising: measuring eye movement of a subject;making an assessment of whether the subject is alive based at least in part on the measured eye movement;estimating one or more anatomical characteristics of an oculomotor plant of the subject based, at least in part, on at least a portion of the measured eye movement, andassessing a person's identity based at least in part on the assessment of whether the subject is alive and on the estimated anatomical characteristics of the oculomotor plant of the subject. 2. The method of claim 1, further comprising assessing one or more complex oculomotor characteristics of the subject based on the measured eye movement. 3. The method of claim 1, further comprising detecting, based at least in part on the assessment of whether the subject is alive, a spoof to attempt an authentication of a person. 4. The method of claim 1, further comprising estimating one or more one or more brain's control strategies in guiding visual attention via exhibition of complex eye movement patterns represented in part by spatial paths of the eye computed from the measured eye movement, wherein the assessment of the subject's identity is based at least in part on at least one of the brain control strategies. 5. The method of claim 1, wherein measuring eye movement of the subject comprises eye tracking. 6. The method of claim 1, wherein assessing the person's identity comprises matching biometric templates related to complex eye movement patterns and oculomotor plant characteristics with previously acquired templates for the subject. 7. The method of claim 1, further comprising authenticating the person based on the assessment of the subject's identity. 8. The method of claim 1, further comprising measuring one or more external characteristics of the eye of the subject, wherein the external characteristics comprise the characteristics of an iris or a periocular region of the eye of the person. 9. The method of claim 1, wherein at least one of the oculomotor plant characteristic templates is based, at least in part, on previous measurements of eye movement of the subject. 10. The method of claim 1, wherein assessing the subject's identity comprises applying one or more statistical tests to one or more oculomotor plant characteristics. 11. A method of making a biometric assessment, comprising: measuring eye movement of a subject;assessing one or more characteristics of the measured eye movement;estimating one or more anatomical characteristics of an oculomotor plant of the subject based, at least in part, on at least a portion of the measured eye movement;assessing a state of the subject based at least in part on the assessed characteristic; andassessing a person's identity based at least in part on the estimated anatomical characteristics of the oculomotor plant of the subject. 12. The method of claim 11, wherein at least one of the assessed characteristics comprises a complex oc... | Texas State University | G06K009/00 A61B005/1171 A61B005/117 A61B005/16 A61B003/113 A61B005/00
US09801581 | B2 | 2017-10-31 | Devices, systems and methods for detecting a bilateral differential in olfactory threshold for pure odorants | 2014-05-20 | The present system is directed in various embodiments to devices, systems and methods for detection, evaluation and/or monitoring olfactory dysfunction by measuring and determining the patient's olfactory detection threshold for the left and the right nostril. More specifically, the present invention relates to devices, systems and methods for detecting an asymmetric differential in a patient's olfactory detection threshold (left vs right nostril) which, when present, may be used as a device to detect, diagnose and/or monitor olfactory deterioration resulting from certain neurodegenerative disorders such as Alzheimer's disease. | 1. A nasal testing device for detecting a patient's pure odorant bilateral olfactory detection threshold for a nostril, comprising: a housing, the housing comprising:a pure odorant compartment within the housing and comprising a pure odorant within the pure odorant compartment;a left airway passage defined within the housing and in fluid communication with a clear air source and in switchable communication with the pure odorant compartment;a right airway passage defined within the housing and in fluid communication with the clear air source and in switchable communication with a pure odorant located within the pure odorant compartment;a left nasal tip having a lumen therethrough and in fluid communication with the left airway passage;a right nasal tip having a lumen therethrough and in fluid communication with the right airway passage;a left airway one-way valve in fluid communication with the lumen of the left nasal tip and the left airway passage and a right airway one-way valve in fluid communication with the lumen of the right nasal tip and the right airway passage, the left and right airway one-way valves allowing exhalation but not inhalation therethrough;a floating ball within the left airway; and a floating ball within the right airway, the floating balls adapted to indicate relative inhalation force, wherein the inhalation force through the left and right airways is generated exclusively from the patient's inhalation at the right and left nasal tips;a left airway activation switch, adapted to switch the left airway out of fluid communication with the clear air source into fluid communication with the pure odorant compartment, thereby introducing pure-odorant infused air into the left airway passage when the right airway remains in fluid communication with the clear air source;a right airway activation switch, adapted to switch the right airway out of fluid communication with the clear air source and into fluid communication with the pure odorant compartment, thereby introducing pure-odorant infused air into the right airway passage when the left airway remains in fluid communication with the clear air source,wherein the left airway activation switch is adapted to further switch the left airway passage out of fluid communication with the pure odorant compartment and into fluid communication with the clear air source when the right airway is switched into fluid communication with the pure odorant compartment; andwherein the right airway activation switch is adapted to further switch the right airway passage out of fluid communication with the pure odorant compartment and into fluid communication with the clear air source when the left airway is switched into fluid communication with the pure odorant compartment. 2. The device of claim 1, wherein the pure odorant is selected from the group consisting of: peanut butter, coffee, vanilla, cinnamon, and lavender. 3. The device of claim 1, wherein the clear air source is selected from the grou... | Inspired Technologies, Inc. | A61B005/00 A61B005/087
US09795308 | B2 | 2017-10-24 | Device for measuring brachial blood pressure in an individual | 2013-09-24 | A device for measuring brachial arterial blood pressure in an individual has a blood pressure cuff attachable to an upper arm, means, separate to the blood pressure cuff, for measuring the heart rate, and a control unit connected to both the blood pressure cuff and the heart rate monitoring means, such that, in use, the control unit monitors the heart rate for the establishment of a stable resting heart rate, initiates a blood pressure measurement, and calculates the pulse pressure (PP) to establish the status of the brachial artery during the blood pressure measurement. Where the heart rate drops to within +12 bpm of a reference, resting heart rate for the individual the device will initiate a blood pressure measurement in accordance with a reference, resting heart rate protocol. However, where an irregular heart rate (IRHR) is found the device will initiate a blood pressure measurement in accordance with a timed protocol. | 1. A device for measuring brachial arterial blood pressure, comprising: a cuff including a brachial pressure sensor for generating a brachial pressure signal corresponding to a brachial pressure of a first brachial artery;a heart rate sensor for generating a heart rate signal of a heart rate at a vessel not downstream of the first brachial artery so that a heart rate measured by the heart rate sensor is not impacted by a pressure applied by the cuff; anda control unit connected to both the brachial pressure sensor and the heart rate sensor for collecting signals therefrom, the control unit including: a storage medium for storing a reference, resting heart rate of a user; anda processor configured to determine a stable parasympathetic state of the brachial artery by comparing signals from the heart rate sensor and comparing the signals against the reference resting heart rate, which reference, resting heart rate has been previously established for the user during initial assessment by monitoring a heart rate for a period of 1000 to 1400 s and identifying a minimum heart rate for the period, which minimum heart rate is defined as the reference, resting heart rate for the user;the processor further configured to initiate a blood pressure evaluation using a resting heart rate protocol, once the measured heart rate has dropped to within +12 bpm of the reference, resting heart rate for the user;the processor further configured to initiate a blood pressure evaluation using a timed protocol, where the measured heart rate does not remain at or below +12 bpm of the reference, resting heart rate for a period within the range of 100 to 140 s, due to the presence of arrhythmia;the processor further configured to receive signals from the brachial pressure sensor and establish a mean arterial pressure (MAP), a pulse pressure (PP), and any arrhythmia from the signals, and write to a storage the pulse pressure, mean arterial pressure, and blood pressure readings in a stored data set; andthe processor further configured to send an alert if the pulse pressure is greater than 55 mmHg unless an arrhythmia is determined by the processor from the heart rate signals. 2. The device according to claim 1, wherein the heart rate sensor is a finger clip heart rate monitor. 3. The device according to claim 1, wherein four sets of data are written to the storage. 4. The device according to claim 3, wherein the control unit analyses the four sets of data and classifies the blood pressure into one of six WHO (World Health Organization) classifications using the systolic or diastolic pressure value, whichever value results in the higher classification. 5. A system for measuring brachial arterial blood pressure, comprising: a cuff including a brachial pressure sensor for generating a brachial pressure signal corresponding to a brachial pressure of a first brachial artery;a heart rate sensor for generating a heart rate signal of a heart rate at a vessel not downstream of the fi... | - | A61B005/02 A61B005/022 A61B005/00 A61B005/024 A61B005/021
US09814787 | B2 | 2017-11-14 | Luciferin amides | 2016-06-03 | This disclosure relates to luciferin amides of formula (I) shown below:; Each variable in formula (I) is defined in the specification. These compounds can be used to detect fatty acid amide hydrolase activities in vitro, in live cells, or in vivo. | 1. A compound of formula (I):wherein n is 0 or 1; is deleted, a single bond, or a double bond; provided that, when n is 0, is deleted;each of X1and X2, independently, is O, S, Se, N, N(Ra), C(Ra), or C(RaRa′), in which each of Ra and Ra′, independently, is H or C1-C10 alkyl; provided that, when is a double bond, X1 is N or C(Ra);each of Y1 and Y2, independently, is N or C(Rb), in which Rb is H or C1-C10 alkyl;Z is C(Rc) or C(RcRc′), in which each of Rc and Rc′, independently, is H or C1-C10 alkyl; provided that, when is a single bond, Z is C(RcRc′) and that, when is a double bond, Z is C(Rc);R2 is N(RdRd′), in which Rd is H, C1 -C10 alkyl optionally substituted with OH, C2-C10 alkenyl, C2-C10 alkynyl, C3-C20 cycloalkyl, C3-C20 cycloalkenyl, C1-C20 heterocycloalkyl, C1-C20 heterocycloalkenyl, aryl, or heteroaryl; Rd′ is C1-C10 alkyl optionally substituted with OH, C2-C10 alkenyl, C2-C10 alkynyl, C3-C20 cycloalkyl, C3-C20 cycloalkenyl, C1-C20 heterocycloalkyl, C1-C20 heterocycloalkenyl, aryl, or heteroaryl; or R2 and R1, together with the carbon atoms to which they are attached, are a substituted or unsubstituted piperidinyl ring; or when R2 and R1, together with the carbon atoms to which they are attached, are a substituted or unsubstituted piperidinyl ring, R2 and R3, together with the carbon atoms to which they are attached, are a substituted or unsubstituted 5, 6, or 7-membered ring optionally fused with one or more substituted or unsubstituted 5, 6, or 7-membered rings;each of R1, R3, R4, R5, and R6, independently, is H, halo, CN, ORe, SRe, NO2, N(ReRe′), C1-C10 alkyl, C2-C10 alkenyl, C2-C10 alkynyl, C3-C20 cycloalkyl, C3-C20 cycloalkenyl, C1-C20 heterocycloalkyl, C1-C20 heterocycloalkenyl, aryl, heteroaryl, in which each of Re and Re′, independently, is H or C1-C10 alkyl; or R1and R2, together with the carbon atoms to which they are attached, are a substituted or unsubstituted piperidinyl ring; or when R1 and R2, together with the carbon atoms to which they are attached, are a substituted or unsubstituted piperidinyl ring, R2 and R3, together with the carbon atoms to which they are attached, are a substituted or unsubstituted 5, 6, or 7-membered ring optionally fused with one or more substituted or unsubstituted 5, 6, or 7-membered rings; andeach of R7 and R8, independently, is H or C1-C3 alkyl optionally substituted with OH. 2. The compound of claim 1, wherein n is 0 and is deleted. 3. The compound of claim 2, wherein each of X1 and X2 is S and each Y1 and Y2 is N. 4. The compound of claim 3, wherein R2 is N(RdRd′), in which each of Rd is H or C1-C10 alkyl optionally substituted with OH, and Rd′ is C1-C10 alkyl optionally substituted with OH. 5. The compound of claim 4, wherein each of Rd is H, CH3, CH2CH3, CH(CH3)2, (CH2)3-OH, or CH2CH(CH3)2, and Rd′ is CH3, CH2CH3, CH(CH3)2, (CH2)3-OH, or CH2CH(CH3)2. 6. The compound of claim 5, wherein each of R1, R3, R4, R5, and R6 is H. 7. The compound of claim 6, wherein each of R7 and R8, i... | University of Massachusetts | A61B005/00 A61B008/00 A61B010/00 A61K049/00 C07D277/66 C07D513/04 G01N021/76 C07D417/04 C07D513/06
US09782624 | B2 | 2017-10-10 | Interchangeable grip and pinch strength assessor and exerciser | 2016-02-02 | A hand strength appliance is provided that may be used in multiple or selected configurations for grip and pinch strength assessment as well as grip and pinch exercises. The appliance includes a frame and a sliding internal handle fitted within and slidable within the frame, and may be positioned for multiple test positions and exercises. The appliance may include and be used with a force sensing device that may transmit information to a computing device. | 1. A multiple configuration hand assessment and exercising appliance comprising: a frame having a connector end with a connector socket, a handle end, and a channel on an interior surface of the frame;an internal handle having a bar and two sliding portions, the internal handle fitted within the frame and slidable within the channel:a stabilizing member adjustably connected to the two sliding portions, the stabilizing member including a stabilizing socket for positioning the bar with respect to the handle end of the frame to provide multiple test positions; anda force sensor assembly secured within the connector socket on the connector end of the frame and within the stabilizing socket, the force sensor assembly capable of processing and transmitting force measurement data to a computing device. 2. The appliance of claim 1, wherein the force sensor assembly further comprises a force sensing device including a body having a neck, a base section, a socket section, and a force sensor within the body, such that the force sensor is capable of measuring tension between the base section and the socket section when force is exerted on the socket section, resulting in a measurable force signal. 3. The appliance of claim 2, wherein the force sensor assembly further includes an anchoring insert, the anchoring insert comprising: an anchoring connector removably secured within a socket passage of the force sensing device, an anchoring base removably secured within the stabilizing socket of the stabilizing member, and an anchoring shaft; and wherein the neck of the force sensing device is secured within the connector socket and wherein the anchoring base is secured within the stabilizing socket. 4. The appliance of claim 3, wherein the stabilizing member is capable of being positioned in five different grip test positions. 5. The appliance of claim 4 wherein the stabilizing member is capable of being positioned in a pinch test position. 6. The appliance of claim 2, wherein the force sensor assembly further comprises: a positioning cable having a length and having a positioning cable plug that may be removably secured within a socket passage of the force sensing device;multiple balls on the length of the positioning cable;a catch assembly that may be removably secured within the connector end of the frame, the catch assembly having stops that provide positions for securing one or more balls on the positioning cable, such that adjustment of the positioning cable and thus the positions of the balls within the catch assembly provides multiple test positions for the appliance. 7. The appliance of claim 6, wherein the stabilizing member may be positioned in five different grip test positions by positioning the balls of the positioning cable in different positions within the catch assembly. 8. The appliance of claim 7 wherein the stabilizing member may additionally be positioned in a pinch test position. 9. The appliance of claim 1, wherein the appliance is u... | Kiio Inc. | A63B023/16 A63B021/055 A63B023/035 A63B021/04 A61B005/22 A63B021/002 A63B021/00 A63B071/06 A61B005/00 G06Q010/06
US09775725 | B2 | 2017-10-03 | Systems and methods for joint replacement | 2016-02-24 | Systems and methods for joint replacement are provided. The systems and methods include a surgical orientation device, a reference sensor device, and at least one orthopedic fixture. The surgical orientation device, reference sensor device, and orthopedic fixtures can be used to locate the orientation of an axis in the body, to adjust an orientation of a cutting plane or planes along a bony surface, or otherwise to assist in an orthopedic procedure(s). | 1. A method of orthopedic surgery, comprising: exposing a distal aspect of a femur of a leg of a patient;inserting a pin into the distal aspect of the femur;coupling a femoral jig assembly to the pin, wherein the femoral jig assembly is mounted at a distal point; andcoupling a surgical orientation device to the femoral jig assembly, the surgical orientation device comprising an inertial sensor, wherein the surgical orientation device is configured to calculate a proximal point. 2. The method of claim 1, wherein the surgical orientation device is configured to calculate a mechanical axis distal point based, at least in part, on the distal point. 3. The method of claim 1, wherein the surgical orientation device is configured to receive an input related to an offset between the distal point and an approximate center of the distal aspect of the femur. 4. The method of claim 1, wherein the surgical orientation device is configured to calculate a mechanical axis proximal point based, at least in part, on a calculated center pivot point of the femur. 5. The method of claim 1, wherein the surgical orientation device is configured to calculate a center of rotation of a head of the femur. 6. The method of claim 1, further comprising swinging the leg of the patient in various directions and/or planes with the surgical orientation device attached to the distal aspect of the femur. 7. The method of claim 6, wherein surgical orientation device is configured to process inertial data as the leg is swung. 8. The method of claim 1, further comprising performing distal femoral resection. 9. The method of claim 1, further comprising cutting the distal aspect of the femur, wherein the pin is located distally to the cut. 10. The method of claim 1, wherein inserting the pin into the distal aspect of the femur further comprises inserting a midline pin at the approximate center of the intercondylar notch. 11. The method of claim 1, further comprising probing an anatomical landmark with a probe extending from the femoral jig assembly. 12. The method of claim 1, further comprising adjusting a cutting block relative to a location of a mechanical axis, the mechanical axis calculated based, at least in part, on the distal point and the proximal point. 13. The method of claim 1, further comprising securing the femoral jig assembly to the femur with one or more additional pins. 14. A method of orthopedic surgery, comprising: exposing a femur of a leg of a patient;mounting a femoral jig assembly to an articular surface of the femur, wherein the femoral jig assembly is mounted at a distal point;coupling a surgical orientation device to the femoral jig assembly, the surgical orientation device comprising an inertial sensor, wherein the surgical orientation device is configured to calculate a proximal point; andresecting the articular surface of the femur. 15. The method of claim 14, wherein the surgical orientation device is configured to calculate a mechanical axis ... | OrthAlign, Inc. | A61F002/46 A61B005/11 A61B017/17 A61B034/10 A61B017/56
US09801588 | B2 | 2017-10-31 | Method and apparatus for reducing coupling between signals in a measurement system | 2015-07-14 | A method and an apparatus for separating a composite signal into a plurality of signals is described. A signal processor receives a composite signal and separates a composite signal in to separate output signals. Pre-demodulation signal values are used to adjust the demodulation scheme. | 1. A system which measures one or more physiological characteristics of a living subject non-invasively by measuring pulsing light of two or more wavelengths after attenuation by body tissue, the system comprising: at least one emitter configured to emit light of at least two or more different wavelengths using a modulation scheme that cycles the at least two or more different wavelengths, the at least one emitter positioned to emit light in the direction of a living subject;at least one detector configured to detect the light of the at least two or more different wavelengths emitted from the at least one emitter after the light has been attenuated by tissue of the living subject, the detector further configured to generate a detector signal representative of the emitted light;a demodulation system configured to demodulate the detector signal, the demodulation system comprising an optimizing demodulator configured to reduce crosstalk by calculating updated coefficients that minimize a predetermined error using baseline coefficients,wherein the demodulation system is configured to use the updated coefficients to determine a plurality of demodulation functions used to demodulate the detector signal. 2. The system of claim 1, the demodulation system is further configured to optimize the demodulation based on previously demodulated signal values. 3. The system of claim 2, wherein the previously demodulated signal values comprises a plurality of demodulated signal values corresponding respectively to the at least two or more different wavelengths. 4. The system of claim 3, wherein each of the plurality of demodulated signal values is determined using a demodulation function. 5. The system of claim 1, wherein the demodulation system is further configured to optimize during an initial calibration period. 6. The system of claim 5, wherein the demodulation system is further configured to optimize during periodic calibration periods during measurement. 7. The system of claim 1, wherein at least one demodulation function is used for each of the at least two or more different wavelengths. 8. A method of demodulating a composite signal generated by applying at least first and second periodic pulses of electromagnetic energy to a system having a physiological parameter to be measured and by non-invasively receiving signals responsive to said electromagnetic energy after attenuation by body tissue of a living subject, said signals received as a composite signal having at least first and second components responsive to said at least first and second pulses respectively, said method comprising: receiving, at one or more signal processors, a composite signal responsive to electromagnetic energy of at least first and second periodic pulses of electromagnetic energy after attenuation by body tissue;applying, using the one or more signal processors, a first demodulation scheme to said composite signal to generate at least a first demodulated signal; andautomat... | CERCACOR LABORATORIES, INC. | A61B005/1455 A61B005/00 A61B005/1495
US09829471 | B2 | 2017-11-28 | pH indicator device and formulation | 2014-10-08 | Disclosed herein are devices and methods for determining the pH of fluid. Example devices include a device comprising a surface configured to contact the fluid and a pH indicator covalently bound thereto, wherein the pH indicator has a first color prior to contact with the fluid and changes color as a function of the pH of the fluid. | 1. A wound dressing configured to determine the pH of a fluid, the wound dressing comprising: an absorbent layer configured to be positioned over a wound;a pH indication layer over the absorbent layer, the pH indication layer comprising a pH indicator covalently immobilized thereon;wherein the pH indicator is configured to have a first color prior to contact with the fluid and changes color as a function of the pH of the fluid; andwherein the pH indicator comprises 2-[(4-(2-hydroxyethylsulfonyl)-phenyl)diazenyl]-4-methylphenol and 4-[4-(2-hydroxyethylsulphonyl)-phenylazo]-2,6-dimethoxy phenol in about a 2:1 ratio. 2. The wound dressing according to claim 1, wherein the color change in the pH indicator is detectable at a 0.1 unit interval change in pH. 3. The wound dressing according to claim 2, wherein the color change in the pH indicator is detectable between about pH 5 and about pH 10. 4. The wound dressing according to claim 3, wherein the color change in the pH indicator is detectable between about pH 5.5 and about pH 9.5. 5. The wound dressing according to claim 1, wherein the wound dressing comprises a cellulosic material. 6. A wound dressing, comprising: a wound-contacting surface,an opposing non-wound contacting surface,a pH indication zone comprising a pH indicator covalently bound therein, which indicates the pH of a fluid, wherein the color of the pH indicator changes in response to a change in the pH of the fluid;an adhesive layer;at least one conduit for directing fluid towards the pH indication zone, the conduit configured to wick fluid directly to the pH indication zone without altering the pH of the fluid; andwherein the pH indicator comprises 2-[(4-(2-hydroxyethylsulfonyl)-phenyl)diazenyl]-4-methylphenol and 4-[4-(2-hydroxyethylsulphonyl)-phenylazo]-2,6-dimethoxy phenol in about a 2:1 ratio. 7. The wound dressing according to claim 6, wherein the device has an outer surface and wherein a pH indication zone is located at or near the outer surface. 8. The wound dressing according to claim 7, wherein the device has a peripheral edge extending between the fluid-contacting surface and the opposing non-fluid contacting surface and wherein the outer surface is the peripheral edge. 9. The wound dressing according to claim 6, wherein the at least one conduit directs fluid laterally towards the pH indication zone. 10. The wound dressing according to claim 6, wherein the color change in the pH indicator is detectable at a 0.1 unit interval change in pH. 11. The wound dressing according to claim 10, wherein the color change in the pH indicator is detectable between about pH 5 and about pH 10. 12. The wound dressing of claim 1, further comprising a flexible transparent layer positioned over the pH indication layer. 13. The wound dressing of claim 12, wherein the pH indication layer comprises an adhesive. 14. The wound dressing of claim 13, wherein the pH indicator is covalently bound to the adhesive. 15. The wound dressing of cl... | Smith & Nephew PLC | G01N031/22 G01N027/416 G01N021/80 A61B005/145 A61F013/42
US09814883 | B2 | 2017-11-14 | Patient posture determination and stimulation program adjustment in an implantable stimulator device using impedance fingerprinting | 2016-09-15 | Methods and circuitry for determining an implanted-neurostimulator patient's position, and adjusting a situation program delivered by the neurostimulator based on the determined position, is disclosed. Impedance measurements of the patient's tissue are taken at the neurostimulator's electrodes, which measurements can comprise complex impedance measurements (magnitude and phase) taken at different frequencies. Such impedance measurements, which can be taken interleaved with stimulation therapy, are used to determine an “impedance fingerprint.” This fingerprint can be compared to other known fingerprints stored in the IPG, which known fingerprints are associated with particular stimulation programs. When a measured fingerprint matches one stored in the IPG, the stimulation program associated with the stored fingerprint is automatically used for patient therapy. As different measured fingerprints are encountered, the IPG can learn and store a new stimulation program for such fingerprint by remembering stimulation parameters selected by the patient when such fingerprint is encountered. | 1. An implantable stimulator device, comprising: a plurality of electrodes configured to provide electrical stimulation therapy to a tissue of a patient; andimpedance monitoring circuitry configured to apply at least one AC measurement signal different from the electrical stimulation therapy to a plurality of combinations of the electrodes and to determine an impedance between each of the plurality of combinations of the electrodes using the at least one AC measurement signal,wherein the impedance monitoring circuitry is configured to apply the at least one AC measurement signal during quiet periods when electrical stimulation therapy is not occurring at any of the plurality of electrodes. 2. The device of claim 1, wherein the determined impedances are complex impedances. 3. The device of claim 1, wherein the impedance monitoring circuitry is configured to apply a plurality of AC measurement signals to each of the plurality of combinations of the electrodes. 4. The device of claim 3, wherein the plurality of AC measurement signals are each of a different frequency, and wherein the impedance between each of the plurality of combinations of the electrodes is determined using the AC measurement signals at each of the different frequencies. 5. The device of claim 1, wherein the impedance monitoring circuitry is further configured to set a stimulation program to be provided to the electrodes in accordance with the determined impedances. 6. The device of claim 5, wherein the determined impedances are indicative of a posture of the patient. 7. The device of claim 1, further comprising a database configured to store a plurality of stimulation programs, wherein the determined impedances are used to select a stimulation program to be provided to the electrodes from the database. 8. The device of claim 7, wherein the plurality of stimulation programs are stored in the database during a fitting procedure. 9. The device of claim 7, wherein the stimulation programs are learned as electrical stimulation is provided to the patient. 10. The device of claim 7, wherein the selected stimulation program comprises one or more stimulation parameters, including which electrodes are active, a polarity of the active electrodes, a magnitude of the electrical stimulation therapy at the active electrodes, a duration of electrical stimulation therapy, and a frequency of the electrical stimulation therapy. 11. The device of claim 1, wherein the impedance monitoring circuitry is configured to determine an impedance between each of the plurality of combinations of the electrodes by assessing responses to the at least one AC measurement signal at each of the plurality of combinations of the electrodes. 12. The device of claim 11, wherein the at least one AC measurement signal comprises an AC current, and the responses comprise AC voltages. 13. The device of claim 1, wherein the impedance monitoring circuitry is further configured to use the impedance determined betwee... | Boston Scientific Neuromodulation Corporation | A61N001/36 A61N001/37 A61B005/053 A61B005/11 A61B005/04 A61N001/08 A61B005/0492
US09794474 | B2 | 2017-10-17 | Method for control of camera module based on physiological signal | 2016-08-17 | A method is provided. The method includes executing a camera application at an electronic device; activating a heart rate monitor (HRM) sensor operatively coupled with a first surface of the electronic device in response to the execution of the camera application; receiving an input signal via the HRM sensor; and capturing, using a processor operatively coupled with the electronic device, an image via an image sensor operatively coupled with a second surface of the electronic device based at least in part on the input signal. | 1. A method comprising: executing a camera application at an electronic device;activating a heart rate monitor (HRM) sensor disposed in a first surface of the electronic device based on the execution of the camera application;receiving an input signal via the HRM sensor;identifying, using a processor operatively coupled with the electronic device, a movement of a user or an object external to the electronic device based at least in part on the input signal; andcapturing, using the processor, an image via an image sensor disposed in a second surface of the electronic device based at least in part on a determination that the movement satisfies a specified condition. 2. The method of claim 1, wherein activating the HRM sensor comprises: generating a signal via the HRM sensor,wherein at least part of the signal is received via the input signal as reflected by a user or the object external to the electronic device. 3. The method of claim 1, wherein the input signal comprises a photoplethysmography (PPG) measurement signal corresponding to a user of the electronic device. 4. The method of claim 3, wherein the capture of the image is performed based at least in part on a determination that the PPG measurement signal falls into a specified range or a specified pattern. 5. The method of claim 1, wherein capturing the image comprises: setting the image sensor with a first setting based at least in part on a determination that the input signal corresponds to an input;setting the image sensor with a second setting based at least in part on a determination that the input signal corresponds to another input; andwherein the image is captured according to a corresponding setting of the first setting and the second setting. 6. The method of claim 5, wherein the corresponding setting comprises adjustment of a focus or a white balance with respect to the image sensor. 7. The method of claim 1, further comprising: presenting the image as a preview via a display operatively coupled with the electronic device; andreplacing the image presented via the display with another image captured via another image sensor operatively coupled with another image sensor in response to another input signal. 8. An apparatus comprising: an image sensor disposed in a first surface of the apparatus;a heart rate monitor (HRM) sensor disposed in a second surface of the apparatus; anda processor adapted to: execute a camera application;activate the HRM sensor based on the execution of the camera application;receive an input signal via the HRM sensor;identify, based at least in part on the input signal, a movement of a user or an object external to the apparatus; andcapture an image using the image sensor based at least in part on a determination that the movement satisfies a specified condition. 9. The apparatus of claim 8, wherein the processor is adapted to: emit a signal using the HRM sensor towards the user or the object external to the apparatus; andreceive, using the HRM senso... | Samsung Electronics Co., Ltd | H04N005/232 A61B005/00 A61B005/024 H04N009/73 G06F003/01 A61B005/0205 A61B005/1455 A61B005/16
US09801569 | B2 | 2017-10-31 | Method of detecting activity with a motion sensor, and corresponding device and computer program | 2012-10-19 | This method for detecting activity (A) of a physical system carrying a motion sensor comprises the following steps: extracting (100) a sequence (M) of measurements provided by the motion sensor; deducing (102) therefrom an observation sequence (O) calculated on the basis of the measurement sequence (M); determining (104) the activity (A) of the physical system in the form of a sequence of states corresponding to the observation sequence, using a statistical model of components of the observation sequence in view of a plurality of possible predetermined states of the physical system.;At least one component of the observation sequence (O) is calculated on the basis of a reference value ({circumflex over (λ)}) of a vector representing a reference axis. In addition, the method comprises the following calibration steps: calculation (106) of a likelihood value of the observation sequence (O) in view of said statistical model; comparison (106) of this likelihood value with a predetermined threshold; on the basis of this comparison, activating an estimation (108) of a new reference value ({circumflex over (λ)}) of the vector representing the reference axis. | 1. A method for detecting activity of a physical system carrying a motion sensor, the method comprising: obtaining a reference value ({circumflex over (λ)}) of a vector representing a reference axis that indicates a preferred direction;extracting a time sequence of measurements provided by the motion sensor;deducing therefrom a time observation sequence calculated using the time sequence of measurements, at least one component of the time observation sequence being calculated from the time sequence of measurements using the obtained reference value ({circumflex over (λ)});using a statistical model of components of the time observation sequence in view of a plurality of possible predetermined states of the physical system, determining the activity of the physical system in the form of a time sequence of states corresponding to the time observation sequence; {circumflex over (λ)}determining whether a calibration is needed by calculating a likelihood value for the time observation sequence in view of said statistical model, andcomparing said likelihood value with a predetermined threshold (δ); andon the basis of the comparison, performing the calibration, in response to determining that the calibration is needed, by updating the obtained reference value ({circumflex over (λ)}). 2. The detection method as claimed in claim 1, wherein the reference axis is estimated, via its representative vector, with measurements provided by the motion sensor. 3. The detection method as claimed in claim 1, further comprising invalidating at least some of the time sequence of states corresponding to the time observation sequence, on the basis of the likelihood value. 4. The detection method as claimed in claim 1, wherein the time observation sequence is deduced from the time sequence of measurements by calculating a first component, on the basis of a low-pass filtering with a predetermined cutoff frequency of each component of the time sequence of measurements, then projecting the resulting components onto the reference value {circumflex over (λ)} ({circumflex over (λ)}); andcalculating at least one second component, on the basis of a band-pass filtering with predetermined cutoff frequencies of each component of the time sequence of measurements. 5. The detection method as claimed in claim 1, wherein the statistical model is a single Markov model with hidden states, each hidden state of which corresponds to one of the possible predetermined states of the physical system, the law of conditional probability of the observation of each hidden state being modeled by a sum of normal laws. 6. The detection method as claimed in claim 1, wherein the calibration comprises: calculating the likelihood value for the time observation sequence in view of said statistical model in the form:V=1N⁢log⁡(p⁡(O1:N)),where N is the number of observation samples in the time observation sequence (O), O1:N designating the N successive samples and p(O1:N) the probability thereof; comparing ... | Commissariat à l'énergie atomique et aux énergies alternatives MOVEA | G01P015/00 A61B005/11 A61B005/00 G01P015/18
US09827430 | B1 | 2017-11-28 | Injected conductive tattoos for powering implants | 2017-02-02 | The present disclosure describes aspects of injected conductive tattoos for powering implants. In some aspects, a system comprises a conductive tattoo used to efficiently transfer power wirelessly received from a transmitter outside a body to an electronic device in the body. The conductive tattoo is formed from a conductive material injected into an outermost permanent layer of the body. The conductive tattoo is configured to wirelessly receive and relay the power from the transmitter to the electronic device. In particular, the conductive tattoo may transfer the power to the electronic device over a coupling between the conductive tattoo and the electronic device. | 1. A system for transferring power wirelessly received from a transmitter located outside a body to an electronic device in the body, the system comprising: a conductive tattoo formed from a conductive material injected into an outermost permanent layer of the body and configured to wirelessly receive the power from the transmitter and transmit the power via a coupling to the electronic device. 2. The system as recited in claim 1, wherein the conductive material is injected in a pattern to form the conductive tattoo, the pattern comprising a resonant loop with an inductor and capacitor. 3. The system as recited in claim 2, wherein the resonant loop is configured to resonate at a frequency of a field generated by the transmitter. 4. The system as recited in claim 2, wherein the resonant loop is configured to generate a magnetic field in response to a transmitter-generated magnetic field and inductively couple with the electronic device using the magnetic field to transmit the power. 5. The system as recited in claim 1, wherein the conductive material is injected in a pattern to form the conductive tattoo, the pattern comprising a gamma loop resonant structure. 6. The system as recited in claim 5, wherein the gamma loop resonant structure comprises a smaller resonant loop with a capacitor, the gamma loop resonant structure further comprising a larger resonant loop formed with a portion of the smaller resonant loop, and the larger resonant loop configured to tune the gamma loop resonant structure to a different impedance for transferring the power. 7. The system as recited in claim 6, wherein the smaller resonant loop is configured to match a lower impedance than the larger resonant loop. 8. The system as recited in claim 1, wherein the conductive material is injected in a loop to form the conductive tattoo, the loop configured to allow current to flow around the loop. 9. The system as recited in claim 1, wherein the outermost permanent layer of the body is dermis of the body's skin. 10. The system as recited in claim 1, further comprising an insulating loop formed from insulating ink injected into the body around the conductive tattoo to prevent the conductive material from migrating through the body. 11. The system as recited in claim 1, wherein the conductive material comprises particles having a size of 13-14 microns. 12. The system as recited in claim 1, wherein the conductive material is an inorganic conductive material comprising at least one of copper, gold, silver, carbon, or titanium oxide. 13. The system as recited in claim 1, wherein the conductive material is injected into the outermost permanent layer of the body using a tattoo machine capable of applying ink tattoos. 14. The system as recited in claim 1, wherein the coupling is a wired coupling between the conductive tattoo and the electronic device. 15. The system as recited in claim 1, wherein the coupling is a wireless coupling between the conductive tattoo and the el... | QUALCOMM Incorporated | H01F037/00 A61N001/378 A61M031/00 A61B005/00 A61M037/00 H02J050/00
US09782083 | B2 | 2017-10-10 | Enhancing exercise safety | 2015-09-04 | A method includes determining that a portable apparatus is in an exercise mode; obtaining, by the portable apparatus when in the exercise mode, real-time exercise data of a user of the portable apparatus acquired using a sensor set having one or more sensors, and reference data related to the exercise mode; observing, when in the exercise mode, the real-time exercise data, and detecting, during the observing based on comparison of the real-time exercise data with the reference data, that at least one value of the exercise data is uncharacteristic to the exercise mode; and determining, when in the exercise mode, whether the detection of the at least one uncharacteristic value is related to an unpredicted incident, and as a response to the determining that the detection of the at least one uncharacteristic value is related to the unpredicted incident, causing an output of an emergency signal. | 1. An apparatus comprising at least one processor and at least one memory including a computer program code, wherein the at least one memory and the computer program code are configured, with the at least one processor, to cause a portable apparatus to perform operations comprising: determining that the portable apparatus is in a user-selectable exercise mode among a plurality of different user-selectable exercise modes, each of the plurality of user-selectable exercise modes being associated with a different type of exercise activity;as a response to determining that the portable apparatus is in the user-selectable exercise mode, causing an activation of one or more sensors of a sensor set;obtaining, by the portable apparatus when in the exercise mode, real-time exercise data of a user of the portable apparatus acquired using the sensor set having the one or more sensors, and exercise mode-specific reference data associated with the exercise mode;observing, when in the exercise mode, the real-time exercise data, and detecting, during the observing based on comparison of the real-time exercise data with the exercise mode-specific reference data, that at least one value of the exercise data is uncharacteristic to the exercise mode; anddetermining, when in the exercise mode, whether the detection of the at least one uncharacteristic value is related to an unpredicted incident, and as a response to the determining that the detection of the at least one uncharacteristic value is related to the unpredicted incident, causing an output of an emergency signal,wherein the plurality of exercise modes comprises at least a first exercise mode and a second exercise mode, wherein the first exercise mode is associated with first exercise mode-specific reference data and the second exercise mode is associated with second exercise mode-specific reference data. 2. The apparatus of claim 1, wherein the determining that the portable apparatus is in the exercise mode comprises determining that the portable apparatus is in said exercise mode among the plurality of exercise modes of the portable apparatus based on sensor input. 3. The apparatus of claim 1, wherein the sensor set is specific to the exercise mode among the plurality of exercise modes. 4. The apparatus of claim 1, wherein the at least one memory and the computer program code are configured, with the at least one processor, to cause the portable apparatus further to perform operations comprising: causing, when in the exercise mode, activation of an exercise mode-specific user interface among a plurality of exercise mode-specific user interfaces. 5. The apparatus of claim 1, wherein the exercise mode-specific reference data indicates at least one of at least one measurement range characteristic to the exercise mode, at least one measurement range uncharacteristic to the exercise mode. 6. The apparatus of claim 5, wherein the detecting that the at least one value of the exercise data is uncharacteristic... | POLAR ELECTRO OY | A61B005/02 G08B021/04 A61B005/024 A61B005/11 G06F019/00 G08B025/08
US09824287 | B2 | 2017-11-21 | Method, apparatus and system for biometric identification | 2016-02-18 | Method and apparatus for processing a biometric measurement signal using a computing device, including receiving biometric measurement records associated with a first biometric measurement generated by contact with a single individual, extracting, for each of the biometric measurement records, feature data including periodic features extracted from the biometric measurement records, determining, pairing data comprising candidate pairs between the feature data and defined profiles associated with a known individual, wherein a candidate pair is associated with one of the periodic features and one of the defined profiles associated with the known individual, determining, for the candidate pair, a similarity value based on the one of the periodic features and the one of the defined profiles associated with the known individual, and determining whether a match exists between the single individual and the known individual based on a combination of the similarity values determined for the candidate pairs. | 1. A method for processing biometric measurement data using a computing device, comprising: receiving at least one biometric measurement record associated with a first biometric measurement generated by contact with a single individual;extracting, by the computing device and for each of the at least one biometric measurement record, feature data including at least one periodic feature extracted from the at least one biometric measurement record;determining, by the computing device, pairing data comprising a plurality of candidate pairs between the feature data and defined profiles associated with a known individual, wherein a candidate pair is associated with one of the at least one periodic feature and one of the defined profiles associated with the known individual;determining, for the candidate pair, a similarity value based on the one of the at least one periodic feature and the one of the defined profiles associated with the known individual; anddetermining whether a match exists between the single individual and the known individual based on a combination of the similarity values determined for the plurality of candidate pairs. 2. The method of claim 1, wherein the at least one biometric measurement record comprises at least one electrocardiogram (ECG) record. 3. The method of claim 1, wherein the extracting, by the computing device and for each of the at least one biometric measurement record, the feature data including at least one periodic feature extracted from the at least one biometric measurement record further comprises: removing noises and outliers from each of the at least one biometric measurement record; andextracting a feature data set including the at least one periodic feature from each of the at least one biometric measurement record using at least one technique. 4. The method of claim 3, wherein the plurality of candidate pairs comprises a candidate pair between the feature data set and each of the defined profiles associated with the known individual. 5. The method of claim 3, wherein the at least one biometric measurement record comprises two or more biometric measurement records associated with different biometric modalities, and the feature data set comprises the feature data extracted for each of the two or more biometric measurement records associated with different biometric modalities. 6. The method of claim 5, wherein the two or more biometric measurement records associated with different biometric modalities comprise two or more of an electrocardiograph (ECG) record, a photoplethysmography (PPG) record, and a bioimpedance record. 7. The method of claim 3, wherein the extracting the feature data set including the at least one periodic feature from each of the at least one biometric measurement record using the at least one technique comprises: extracting a first group of features, as a first part of the feature data set, from the at least one biometric measurement record using a fiducial point based approach,... | Huami Inc. | G06K009/00 G06N099/00 G06K009/46 G06K009/62 A61B005/0452 A61B005/02 A61B005/04
US09808670 | B2 | 2017-11-07 | Monitoring fitness using a mobile device | 2016-11-23 | Athletic performance monitoring and tracking may provide multiple ways in which to track athletic movement and activity. Workouts may also be tagged with various parameters including mood, weather, terrain, athletic equipment, friends used and the like. Workout information may be shared to social messaging and networking outlets. Workout information shared may include map information including images of maps, interactive maps, links to maps, route information and the like and/or combinations thereof. Additionally or alternatively, an application may be configured to execute within a context of a social networking system to facilitate athletic activity data transfer and generation of workout entries in the social networking site. | 1. An apparatus comprising: a processor; anda non-transitory computer-readable medium storing computer readable instructions that, when executed, cause the apparatus to: receive, from an athletic activity monitoring system remote from the apparatus, athletic activity data recorded during an athletic activity performed by a user;select an athletic activity descriptor object based on the received athletic activity data, wherein the athletic activity descriptor object describes a type of action performed by the user during the athletic activity;determine an activity condition based on the received athletic activity data;determine whether the activity condition meets a first sharing threshold; andresponsive to determining that the activity condition meets the first sharing threshold, generate an athletic activity entry including generating a message with the athletic activity descriptor object on a network page of a social networking site, wherein the athletic activity entry and the network page is viewable by at least one other user. 2. The apparatus of claim 1, wherein the activity condition includes an athletic activity metric range. 3. The apparatus of claim 1, wherein the activity condition includes a frequency of athletic activity entries shared on the network page of the social networking site. 4. The apparatus of claim 1, wherein the activity condition includes whether the athletic activity was performed alone or with one or more other users. 5. The apparatus of claim 1, wherein the activity condition includes at least one of whether a goal was achieved or whether a personal best was exceeded. 6. The apparatus of claim 1, wherein the activity condition is determined based on the type of action performed by the user during the athletic activity. 7. The apparatus of claim 1, further comprising an activity tracking application configured to execute within the social networking site. 8. The apparatus of claim 1, wherein selection of the athletic activity descriptor object includes: determining a value of an athletic activity metric recorded in the athletic activity data;determining whether the value is above a threshold value; andin response to determining that the value is equal to or above the threshold value, selecting a first athletic activity descriptor object; andin response to determining that the value is below the threshold value, selecting a second athletic activity descriptor object. 9. The apparatus of claim 1, wherein the athletic activity descriptor object is selected from a plurality of athletic activity descriptor objects, wherein each of the athletic activity descriptor objects is specific to a single type of athletic activity. 10. The apparatus of claim 1, wherein generating the athletic activity entry includes: determine a second activity condition based on the received athletic activity data; anddetermine whether the second activity condition meets a second sharing threshold; andwherein the message is generated on the... | NIKE, Inc. | A63B024/00 A61B005/11 G06Q050/00 G06F019/00 G01C021/20 G01C022/00 G06F001/16 A63B071/06
US09820655 | B2 | 2017-11-21 | Systems and methods for spectral analysis of a tissue mass using an instrument, an optical probe, and a Monte Carlo or a diffusion algorithm | 2008-09-29 | Systems and methods for spectral analysis of a tissue mass using an instrument, an optical probe, and a Monte Carlo algorithm or a diffusion algorithm are provided. According to one method, an instrument is inserted into a tissue mass. A fiber optic probe is applied via the instrument into the tissue mass. Turbid spectral data of the tissue mass is measured using the fiber probe. The turbid spectral data is converted to absorption, scattering, and/or intrinsic fluorescence spectral data via a Monte Carlo algorithm or diffusion algorithm. Biomarker concentrations in the tissue mass are quantified using the absorption, scattering, and/or intrinsic fluorescence spectral data. | 1. A system for determining biomarker concentrations in a tissue mass, the system comprising: an instrument for inserting into the tissue mass;a fiber optic probe for interfacing the tissue mass upon being inserted into the instrument and measuring turbid spectral data of the tissue mass, wherein measuring the turbid spectral data includes conducting fluorescence spectral measurements and diffuse reflectance measurements of the tissue mass simultaneously, wherein the fiber optic probe is a forward firing fiber optic probe that includes a probe tip comprising a plurality of centrally grouped illumination fibers that are concentrically surrounded by a plurality of collection fiber rings, wherein each of the collection fiber rings consists of a plurality of ring-specific collection fibers or a plurality of ring-specific collection fibers and a plurality of dead fibers; anda processing unit for converting the turbid spectral data to absorption, scattering, and intrinsic fluorescence spectral data via a Monte Carlo algorithm or a diffusion algorithm and quantifying biomarker concentrations in the tissue mass using the absorption, scattering, and intrinsic fluorescence spectral data, wherein the intrinsic fluorescence spectral data is corrected to account for the absorption and scattering data. 2. The system of claim 1 wherein the fiber optic probe fits within a biopsy needle or a biopsy cannula. 3. The system of claim 2 wherein the biopsy needle includes at least one of a Suros 9G biopsy needle, a Mammotome biopsy needle, a Bard Vacora 10G biopsy needle, a Bard Vacora 14G biopsy needle, and a Cardinal Achieve 14G biopsy needle. 4. The system of claim 1 wherein the instrument comprises a cannula, a biopsy needle, an endoscopic instrument, and a laparoscopic instrument. 5. The system of claim 1 wherein the turbid spectral data comprises diffuse reflectance spectral data and fluorescence spectral data of the tissue mass. 6. The system of claim 1 wherein the biomarker concentrations are used to determine if the tissue mass is malignant, benign, or normal, or are used to quantify drug uptake by the tissue mass. 7. The system of claim 6 wherein a biopsy is conducted on the tissue mass through a cannula or a biopsy needle if the tissue is malignant. 8. The system of claim 1 wherein the fiber optic probe is adapted to acquire the turbid spectral data using a spectrometer having a spectral bandpass of between about 2.5 nm and 5 nm. 9. The system of claim 1 wherein the Monte Carlo algorithm includes either an inverse Monte Carlo reflectance algorithm or an inverse Monte Carlo fluorescence algorithm. 10. The system of claim 1 wherein the processing unit is further configured to determine the concentrations at least one of total hemoglobin, oxyhemoglobin concentration, deoxyhemoglobin concentration, beta carotene concentration, hemoglobin saturation, nicotinamide adenine dinucleotide (NADH), flavin adenine dinucleotide (FAD), collagen, porphyrins, retino... | DUKE UNIVERSITY | A61B005/00
US09775726 | B2 | 2017-10-03 | Electronically controlled prosthetic system | 2016-04-06 | A prosthetic joint system for users comprising a housing having an interior cavity, a center axis in said interior cavity, and an attachment means for fixedly connecting said housing to said user; an inner cylinder disposed in said housing interior cavity wherein said inner cylinder rotates around said center axis of said housing; an appendage attached to said inner cylinder; a sensor system attached to said appendage; and a dampening system, having a power source, in communication with said sensor system, said inner cylinder, and said housing for controlling dampening of the rotation of said inner cylinder around said center axis of said housing. | 1. A prosthetic ankle joint system for a user comprising: a bracket having a first end, a length, and a second end;a connection point positioned on said first end of said bracket;a central axis positioned on said second end of said bracket;a joint axis positioned on said length of said bracket;an anterior base;a posterior base;an anterior section having a first end pivotally attached to said joint axis and a second end pivotally attached to said anterior base;a posterior section having a first end pivotally attached to said joint axis and a second end attached to said posterior base;a first linkage having a first end connected to said anterior base, a second rotationally attached to said central axis and a spring element positioned between said first end of said linkage and said second end of said linkage; anda second linkage having a first end connected to said posterior base and a second end rotationally attached to said central axis. | - | A61F002/66 A61F002/72 A61B005/11 A61B005/00 A61F002/50 A61F002/68 A61F002/70 A61F002/74 A61F002/58 A61F002/60 A61F002/64 A61F002/76
US09775540 | B2 | 2017-10-03 | Endoscope insertion shape observation probe | 2016-04-28 | An endoscope insertion shape observation probe that is elongatedly formed along a longitudinal axis direction and is insertable into a channel provided in an endoscope insertion section or is incorporated in the endoscope insertion section includes an electronic component, a housing section having a recessed shape that houses the electronic component, a pair of electric connection sections provided in the housing section, a pair of electric connection terminal sections provided in the electronic component and engaged with and electrically connected to the electric connection sections, the electric connection terminal sections mechanically holding the electronic component in the housing section, a groove section formed in the housing section, and two attaching-direction defining terminal sections provided in the electronic component, the attaching-direction defining terminal sections engaging into the groove section and defining a direction around a longitudinal axis in which the electronic component is attached to the housing section. | 1. An endoscope insertion shape observation probe that is elongatedly formed along a longitudinal axis direction and is insertable into a channel provided in an endoscope insertion section or is incorporated in the endoscope insertion section, the endoscope insertion shape observation probe comprising:an electronic component;a housing section having a recessed shape that detachably and attaachably houses the electronic component;a pair of electric connection sections provided in the housing section;a pair of electric connection terminal sections provided in the electronic component and engaged with and electrically connected to the electric connection sections, the electric connection terminal sections mechanically holding the electronic component when the detachable and attachable electronic component is attached to the housing section;a groove section formed in the housing section; andtwo attaching-direction defining terminal sections provided in the electronic component, the attaching-direction defining terminal sections engaging into the groove section and defining a direction around the longitudinal axis in which the electronic component is attached to the housing section. 2. The endoscope insertion shape observation probe according to claim 1, wherein the groove section and the attaching-direction defining terminal sections are provided on an end face opposed to the housing section in a longitudinal axis direction in the electronic component. 3. The endoscope insertion shape observation probe according to claim 2, wherein one or another of the pair of electric connection terminal sections and one or another of the two attaching-direction defining terminal sections are formed of conductor members having a same shape respectively forming a pair. 4. The endoscope insertion shape observation probe according to claim 1, wherein the pair of electric connection terminal sections and the two attaching-direction defining terminal sections are provided to project from a distal end face in a longitudinal direction of the electronic component,the pair of electric connection terminal sections is provided at one end of the electronic component to have a predetermined angle with respect to the direction in which the electronic component is attached to the housing section and have centers in point symmetrical positions with respect to a center of the distal end face, andthe two attaching-direction defining terminal sections are provided in parallel in the attaching direction of the electronic component to the housing section and provided at one end of the electronic component to have centers in positions where separation distances from the center of the distal end face are different from each other. 5. The endoscope insertion shape observation probe according to claim 1, wherein the pair of electric connection terminal sections and the two attaching-direction defining terminal sections are provided to project from a distal end face in a longitudinal di... | OLYMPUS CORPORATION | A61B001/00 A61B005/06 G02B023/24 A61B034/20
US09782119 | B2 | 2017-10-10 | Wrinkle detection apparatus and wrinkle detection method | 2015-06-01 | A wrinkle detection apparatus is an apparatus for detecting a wrinkle area of skin included in an image. The wrinkle detection apparatus includes a first threshold processing unit that detects a first edge area included in the image by performing first threshold processing for comparing a gradient value that indicates a degree of change in pixel values in the image with a first threshold, a second threshold processing unit that detects a second edge area included in the image by performing second threshold processing for comparing the gradient value with a second threshold, and an edge selection processing unit that determines a wrinkle area based on an overlapping area in which the first edge area and the second edge area are overlapped. | 1. A wrinkle detection apparatus comprising: a first threshold processor adapted to detect one or more first edge areas included in an image of skin, and adapted for comparing a gradient value that indicates a degree of change in pixel values in the image with a first threshold, each of the one or more first edge areas being a continuous area in the image where the gradient value is larger than the first threshold;a second threshold processor adapted to detect one or more second edge areas included in the image, and adapted for comparing the gradient value with a second threshold that is larger than the first threshold, each of the one or more second edges areas being another continuous area in the image where the gradient value is larger than the second threshold; andan edge selection processor adapted for determining at least one first edge area from among the one or more first edge areas as a wrinkle area in the image, the at least one first edge area partially overlapping with at least one of the one or more second edge areas. 2. The wrinkle detection apparatus according to claim 1, further comprising: an image obtaining unit that obtains the image;a facial component detector that obtains a position of a facial component from the image; anda detection area determination unit that determines a detection area that is a target of first threshold processing and second threshold processing, out of the image, based on the position of the facial component,wherein the first threshold processor and the second threshold processor calculate at least the gradient value of the detection area. 3. The wrinkle detection apparatus according to claim 1, further comprising: an image obtaining unit that obtains the image;an image generation unit that generates a wrinkle image indicating the wrinkle area in the image, and generates a wrinkle-enhanced image obtained through superimposition of a generated wrinkle image on the image; anda display that displays the wrinkle-enhanced image. 4. A wrinkle detection method comprising: detecting one or more first edge areas included in an image of skin by comparing a gradient value that indicates a degree of change in pixel values in the image with a first threshold, each of the one or more first edge areas being a continuous area in the image where the gradient value is larger than the first threshold;detecting one or more second edge areas included in the image by comparing the gradient value with a second threshold that is larger than the first threshold, each of the one or more second edge areas being another continuous area in the image where the gradient value is larger than the second threshold; anddetermining at least one first edge area from among the one or more first edge areas as a wrinkle area in the image, the at least one first edge area partially overlapping with at least one of the one or more second edge areas. 5. A wrinkle detection apparatus comprising: a processor; anda memory having instructions s... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | A61B005/02 G06K009/00 A61B005/00 G06T007/00 G06T007/13 G06T007/136
US09775566 | B2 | 2017-10-03 | Filtering patient signal also for ventilation artifacts | 2017-01-30 | In embodiments, an external medical device is intended to care for a patient. If it receives an input that signifies that ventilation artifact is present in a signal of the patient, it transmits a corrective signal responsive to the received input. In further embodiments, a patient signal is received, which is generated from a patient while the patient is or was receiving chest compressions at a frequency Fc, and also receiving ventilations at frequency Fv. At least one filter mechanism may be applied to the patient signal to substantially remove artifacts at a) frequency Fc, b) a higher harmonic of frequency Fc, and c) a third frequency substantially equaling frequency Fc plus or minus frequency Fv, while substantially passing other frequencies between them. As a result, the patient signal can be cleaner, for diagnosing the patient's state more accurately. | 1. An external device for use when a patient is receiving ventilations from a ventilator that is being operated by a rescuer, comprising: a housing; anda processor within the housing configured to: perform a correlation of a sensed patient impedance signal with an airway signal of the patient;receive, responsive to the performed correlation, an input signifying that a ventilation artifact is present in the sensed patient impedance signal, the ventilation artifact possibly due to the ventilations from the ventilator; andcause to be transmitted a corrective signal responsive to the received input; anda user interface configured to output, responsive to the corrective signal, a prompt to the rescuer. 2. The device of claim 1, in which the prompt indicates the presence of a ventilation artifact. 3. The device of claim 1, in which the prompt is to the effect that the rescuer should discontinue operating the ventilator that presumptively has caused the ventilation artifact. 4. The device of claim 1, in which the airway signal is a capnography signal, an airway pressure signal, or an airflow signal of the patient. 5. The device of claim 1, in which the processor is further configured to: cause a record to be created about one of the input and the corrective signal. 6. A non-transitory computer-readable storage medium storing one or more programs which, when executed by at least one external device that is intended for use when a patient is receiving ventilations from a ventilator that is being operated by a rescuer, the external device including a processor and a user interface, the one or more programs result in: performing a correlation of a sensed patient impedance signal with an airway signal of the patient;receiving, responsive to the performed correlation, an input signifying that a ventilation artifact is present in the sensed patient impedance signal, the ventilation artifact possibly due to the ventilations from the ventilator;causing to be transmitted a corrective signal responsive to the received input; andcausing to be output, responsive to the corrective signal, a prompt to the rescuer. 7. The medium of claim 6, in which the prompt indicates the presence of a ventilation artifact. 8. The medium of claim 6, in which the prompt is to the effect that the rescuer should discontinue operating the ventilator that presumptively has caused the ventilation artifact. 9. The medium of claim 6, in which the airway signal is a capnography signal, an airway pressure signal, or an airflow signal of the patient. 10. The medium of claim 6, in which executing the one or more programs further results in: causing a record to be created about one of the input and the corrective signal. 11. A method for an external medical device intended for use when a patient is receiving ventilations from a ventilator that is being operated by a rescuer, the method comprising: performing a correlation of a sensed patient impedance signal with an airway signal of the... | PHYSIO-CONTROL, INC. | A61B005/05 A61B005/00 A61M016/00 A61B005/046 A61B005/0205 A61H031/00 A61N001/39 A61B005/053 A61B005/083 A61B005/087
US09787676 | B2 | 2017-10-10 | Multi-modal biometric identification | 2015-09-29 | At least two biometric measurement signals are generated by contact with a single individual. At least one feature is extracted from each signal of the at least two biometric measurement signals, the extracted features are combined to generate a combined biometric signal. The combined biometric signal is compared with a defined biometric signal associated with a known individual, responsive to the combined biometric signal matching the defined biometric signal, a signal is transmitted indicating that the single individual is the known individual. The biometric measurement signals can be collected by a biometric identification device worn or carried by the single individual. The processing may be done by the biometric identification device or a remote server. | 1. A method, comprising: receiving, at a computing device, at least two biometric measurement signals generated by contact with a single individual, wherein the at least two bimetric measurement signals comprise an electrocardiograph (ECG) signal and a photoplethysmography (PPG) signal;extracting, at the computing device, at least one feature from each signal of the at least two biometric measurement signals;combining the at least one feature extracted from each signal to generate a combined biometric signal;comparing the combined biometric signal with a defined biometric signal associated with a known individual; andtransmitting a signal, responsive to the combined biometric signal matching the defined biometric signal, indicating that the single individual is the known individual. 2. The method of claim 1, further comprising: pre-processing each of the at least two biometric measurement signals before extracting at least one feature from each signal. 3. The method of claim 1 wherein extracting the at least one feature comprises: extracting a first feature from the ECG signal using a first technique;extracting a second feature from the ECG signal using a second technique;extracting a first feature from the PPG signal using a third technique; andextracting a second feature from the PPG signal using a fourth technique. 4. The method of claim 3 wherein the first technique is a fiducial point based approach, the second technique is an appearance based approach, the third technique is a fiducial point based approach, and the fourth technique is an appearance based approach. 5. The method of claim 1 wherein combining the at least one feature extracted from each signal comprises: applying a weight to each feature extracted from the ECG signal; andapplying a weight to each feature extracted from the PPG signal, wherein each weight is based on a training algorithm applied to data from a plurality of individuals. 6. The method of claim 1, further comprising: receiving, at the computing device, a plurality of features associated with the known individual and extracted during a set up process; andgenerating, using the plurality of features, the defined biometric signal using a same technique used to generate the combined biometric signal before comparing the combined biometric signal with the defined biometric signal. 7. The method of claim 1 wherein receiving the at least two biometric signals comprise wirelessly receiving the at least two biometric measurement signals generated by a biometric identification device worn by the single individual. 8. The method of claim 1 wherein the at least two biometric measurement signals comprise at least a first cyclical biometric measurement signal and a second cyclical biometric measurement signal; and wherein extracting the at least one feature from each signal comprises: extracting a first feature from the first cyclical biometric measurement signal using a fiducial point based approach;extracting a second f... | Anhui Huami Information Technology Co., Ltd. Huami Inc. | A61B005/02 H04L029/06 A61B005/021 A61B005/00 A61B005/04 A61B005/0245
US09795330 | B2 | 2017-10-24 | Device, system and method for in-vivo detection of bleeding in the gastrointestinal tract | 2014-06-13 | In-vivo devices, systems and methods for the detection of blood within in-vivo bodily fluids. The methods include irradiating in-vivo fluids passing through a gap in a housing of an in-vivo device introduced to the GI tract of a subject with a plurality of illumination sources positioned on a first side of a gap; detecting with at least one light detector positioned on the opposite side of the gap and facing the illumination sources, light irradiated by the illumination sources; transmitting a plurality of values representing the light detected over time; converting these values to blood concentration values over time, and comparing the blood concentration values to a predetermined threshold value. Based on the comparison, the method includes determining the type of bleeding profile, such that if a plurality of blood concentration values measured consecutively is above the threshold value, the bleeding profile indicates bleeding. | 1. A system for determining the type of a patient's bleeding profile over time, said system comprising: an in-vivo sensing device comprising: a housing comprising a sensing head in the form of a gap, which remains in contact with bodily fluids and through which the fluids pass when the device is within the GI tract;said sensing head comprising a plurality of illumination sources positioned on a first side of the gap, each illumination source for repeatedly irradiating the fluids within the gap at a different narrow band wavelength, and at least one light detector photodiode positioned on the opposite side of the gap and facing the illumination sources, for detecting light irradiated by the illumination sources and which passed through the fluids;a transmitter for transmitting a plurality of values representing the light detected over time by the at least one light detector to an external receiver;a memory unit for storing the values received by the external receiver;a processing unit configured for processing the stored values and for: converting the values to a series of blood concentration values measured consecutively over time, andidentifying a bleeding event based on the blood concentration values as a function of time and based on a comparison of the blood concentration values to a threshold. 2. The system according to claim 1, wherein the processor is configured to identify a bleeding event indicating major-active bleeding over time if the blood concentration values show a peak above the threshold value after which a plurality of blood concentration values measured consecutively are below the peak but are still above the threshold value. 3. The system according to claim 1, wherein the processor is configured to identify a bleeding event indicating major-non-active bleeding over time if the blood concentration values show a steady peak close to the threshold value. 4. The system according to claim 1, wherein the processor is configured to identify a bleeding event indicating minor bleeding over time if the blood concentration values show at least one peak below the threshold value. 5. The system according to claim 1, wherein the processor is configured to identify a bleeding event indicating no bleeding over time if the blood concentration values are below the threshold value with no distinct peak over time. 6. The system according to claim 1, wherein said memory unit and said processing unit are incorporated within said external receiver. 7. The system according to claim 1, wherein said sensing head comprises a cover placed over the gap, said cover comprising at least one hole located on an axis perpendicular to an axis along which the illumination sources and the light detector are located. 8. The system according to claim 7, wherein said at least one hole comprises holes located on opposite sides of the gap. 9. The system according to claim 7, wherein said cover is shaped as a flexible accordion, said cover defining a space withi... | GIVEN IMAGING LTD. | A61B005/1455 A61B005/1459 A61B005/00 A61B005/07 A61B001/04 A61B001/06
US09820656 | B2 | 2017-11-21 | Personalized nutritional and wellness assistant | 2012-07-06 | The invention pertains to the establishment, implementation and management of a personalized information system pertinent to a user's general health, wellness and/or sport performance. Disclosed is a system capable of transcutaneous measurement of a subject including at least one light source, at least one light detector, and at least one component for generating or storing at least one value of VC02 or at least one value of V02 from the detected signal. Further, disclosed is a portable device for analyzing the composition of the respired gasses of a subject including at least one air flow conduit through which the subject can inspire or expire air through the body of the device, at least one sampling portal, an oxygen sensor, and at least one flow sensor. A dual-battery system is also provided by which an uninterrupted power supply can be provided for electronic components. | 1. A system capable of transcutaneous measurement of a subject, wherein the system comprises: (a) at least one light source for shining light onto or through tissues of and under skin of the subject;(b) at least one light detector for receiving light reflected from the tissues of the subject's skin or from the tissues underlying the subject's skin and converting the reflected light into a detected signal; and(c) at least one component adapted to generate and store at least one value for each of VCO2 and VO2 from the detected signal, wherein VO2 corresponds to an instantaneous rate of oxygen consumption of the subject and wherein VCO2 corresponds to an instantaneous rate of carbon dioxide production by the subject, wherein the at least one component is further adapted to continuously calculate a real-time respiratory quotient of the subject from real-time values of VCO2 and VO2 and to use the real-time respiratory quotient to determine a real-time energy uptake of the subject. 2. The system of claim 1, wherein the at least one light source comprises a plurality of light sources that simultaneously or sequentially direct light of different wavelengths onto or through the tissues of and under the skin of the subject. 3. The system of claim 1, wherein the at least one light detector comprises a plurality of light detectors, each of which receives light of at least one wavelength. 4. The system of claim 1, wherein the system incorporates at least one operational amplifier or at least one microprocessor by means of which at least one digital potentiometer may be iteratively adjusted in order to obtain optimal amplification of the at least one detected signal. 5. The system of claim 1, wherein the at least one component stores, executes, transmits or receives a digital value, at least one mathematical function, at least one parameter of the at least one mathematical function, or any combination for generating at least one value of at least one parameter of physiology for the purpose of further mathematical calculation or for informing the subject of its physiological status. 6. The system of claim 5 wherein the at least one parameter of physiology may be selected from the group consisting of heart rate (HR), breathing rate (BR), hemoglobin concentration (Hb) oxyhemoglobin concentration (HbO2), carbaminohemoglobin concentration (HbCO2), oxygen saturation (SpO), oxygen consumption rate (VO2), carbon dioxide production rate (VCO2), Respiratory Exchange Ratio (RER), Food Quotient (FQ), Metabolic Fuel Composition (i.e. the macronutrient that a subject is utilizing as metabolic fuel at a given instance), Total Energy Expenditure (TEE), Resting Energy Expenditure (REE), Physical Activity Energy Expenditure (PAEE), Energy Balance (EB), Excess Postexercise Oxygen Consumption (EPOC), Body Fat Percentage (% BF) and Current Body Composition (CBC). 7. The system of claim 1, wherein the at least one component is further used for storing or transmitting or recei... | LifeQ Global Limited | A61B005/02 A61B005/0205 A61B005/083 A61B005/087 A61B005/00 A61B005/22 G06Q050/22 A61B005/08 A61B005/1455 A61B005/024 A61B005/053 H02J009/06
US09814410 | B2 | 2017-11-14 | Person support apparatus with position monitoring | 2015-04-22 | A person support apparatus includes one or more thermal image sensors whose outputs are analyzed to perform one or more functions. Such functions include automatically turning on a brake, automatically turning on one or more lights, detecting when a patient associated with the person support apparatus has fallen, enabling a propulsion system of the patient support apparatus to be used, automatically controlling one or more environmental controls, and/or automatically arming an exit detection system after entry of a patient onto the person support apparatus. Multiple thermal images may be generated from multiple sensors to generate stereoscopic thermal images of portions of the person support apparatus and its surroundings. | 1. A person support apparatus comprising: a support surface adapted to support a person thereon;a first sensor adapted to detect electromagnetic waves coming from the person when the person is supported on the support surface;a second sensor adapted to detect a force exerted by the person while the person is supported on the support surface;a memory in which a location of the first sensor with respect to the support surface is stored; anda controller adapted to correlate information from the first and second sensors in order to determine a position of the person on the support surface, the controller using the location of the first sensor with respect to the support surface when determining the position of the person on the support surface. 2. The person support apparatus of claim 1 wherein the first sensor is an infrared sensor and the second sensor comprises a plurality of load cells adapted to detect a weight of the person when positioned on the support surface. 3. The person support apparatus of claim 1 further comprising a user input adapted to arm and disarm a person monitor, and wherein the controller is adapted to issue an alert if the person monitor is armed and the position of the person on the support surface changes by more than a predetermined threshold. 4. The person support apparatus of claim 1 further comprising a third sensor spaced a measured distance from the first sensor, the third sensor adapted to detect electromagnetic waves coming from the person when the person is supported on the support surface, and the controller is adapted to use the measured distance in order to determine a distance of the person from the first and third sensors. 5. The person support apparatus of claim 1 wherein the controller is adapted to use the information from the first and second sensors to determine when the person meets at least one of the following conditions: the person is sitting up; the person is lying on his or her back; the person is lying on his or her side; the person's legs have moved beyond an edge of the support surface; and the person has turned from his or her back to his or her side, or vice versa. 6. The person support apparatus of claim 1 wherein the controller determines a three-dimensional position of the person on the support surface. 7. The person support apparatus of claim 1 wherein the controller is further adapted to use the information from the first and second sensors to determine if a non-human object is positioned on the support surface. 8. The person support apparatus of claim 1 further comprising: a wheel; anda brake for the wheel;wherein the controller is adapted to automatically change the brake to a braked state based at least partially upon the information from the first and second sensors. 9. The person support apparatus of claim 2 further comprising: a plurality of side rails;a footboard; anda headboard, wherein the infrared sensor is positioned on one of the side rails, footboard, and headboard. 1... | Stryker Corporation | A61B005/11 G01J005/00 A61G007/05 A61G013/10 G01D005/40 G01G019/44 G01J005/14 A61G005/10
US09782106 | B2 | 2017-10-10 | Method and magnetic resonance apparatus for determination of magnetic resonance angiography images using time-of-flight angiography | 2015-12-14 | In a method and apparatus, magnetic resonance angiography images of an examination volume of a patient are obtained using time-of-flight angiography in a magnetic resonance scanner. By continuous recording, a number of two-dimensional slice images covering the examination volume along an axial direction are acquired in a slice-by-slice layer-wise, such as with overlapping. The slice images are divided into groups of, in each case, a predetermined number of consecutive slice images in the axial direction. A maximum intensity projection image is determined for each group, and the angiography images are determined as the maximum intensity projection images and/or dependent on the maximum intensity projection images. | 1. A method for generating a magnetic resonance (MR) angiography image, comprising: operating an MR scanner, while a patient is situated therein, to obtain a plurality of successive two-dimensional slice images that cover an examination volume of the patient along a direction slice-by-slice, using a time-of-flight angiography MR data acquisition sequence;providing said plurality of slice images to a processor and, in said processor, dividing said plurality of slice images into respective groups with each group comprising a predetermined number of consecutive slice images along said direction;in said processor, determining a maximum intensity projection image for each group using the slice images of that group; andin said processor, determining at least one diagnostic angiography image for each group by determining a correlation measure for each slice image in a respective group with regard to the maximum intensity projection image that was determined for that group, and thereby obtaining a plurality of correlation measures for each group, and determining said at least one diagnostic angiography image for the respective group dependent on an evaluation in said processor of the respective correlation measures for the slice images in that respective group and making the respective diagnostic angiography images for each group available in electronic form from said processor as respective data files. 2. A method as claimed in claim 1 comprising operating said MR scanner to acquire said plurality of two-dimensional slice images so that at least some successive slice images overlap each other. 3. A method as claimed in claim 1 comprising using a slice image that has the highest correlation measure in each group as the only diagnostic angiography image for that group. 4. A method as claimed in claim 1 comprising using all slice images in each group that have a correlation measure that exceeds a predetermined threshold value as diagnostic angiography images for that group. 5. A method as claimed in claim 1 comprising assigning a recording position to a slice image in each group having a highest correlation measure, and giving said at least one diagnostic angiography image for that group said recording position. 6. A method as claimed in claim 1 comprising determining said predetermined number as a function of an image recording rate of said plurality of two-dimensional slice images and as a function of a duration of the cardiac cycle of the patient. 7. A method as claimed in claim 6 comprising determining said predetermined number as a product of said duration and said image recording rate. 8. A method as claimed in claim 6 comprising using a predetermined maximum duration of said cardiac cycle as said duration. 9. A method as claimed in claim 8 comprising setting said predetermined maximum duration to be in a range between 800 and 1100 ms. 10. A method as claimed in claim 8 comprising setting said predetermined maximum duration to be 1000 ms. 1... | Siemens Aktiengesellschaft | A61B005/00 A61B005/055 G06T011/00 G01R033/563
US09775523 | B2 | 2017-10-03 | Pressure sensing guidewire and methods for calculating fractional flow reserve | 2014-10-10 | Medical devices and methods for making and using medical devices are disclosed. An example method may include a method for calculating fractional flow reserve including providing a pressure sensing guidewire, advancing the pressure sensing guidewire through a blood vessel to a first position distal of an intravascular occlusion, determining a distal pressure within the body lumen with the pressure sensing guidewire, proximally shifting the pressure sensing guidewire to a second position proximal of the occlusion, determining a proximal pressure within the body lumen with the pressure sensing guidewire, measuring an aortic pressure, and calculating a pressure drift. The pressure drift may be the difference between the aortic pressure and the proximal pressure. The method may also include calculating a drift-compensated fractional flow reserve. The drift-compensated fraction flow reserve may correspond to (the distal pressure+the pressure drift) divided by the aortic pressure. | 1. A method for calculating fractional flow reserve, the method comprising: advancing a pressure sensing guidewire through a body lumen to a first position distal of an intravascular occlusion;determining a distal pressure within the body lumen with the pressure sensing guidewire;proximally shifting the pressure sensing guidewire to a second position proximal of the intravascular occlusion;determining a proximal pressure within the body lumen with the pressure sensing guidewire;measuring an aortic pressure;calculating a pressure drift, wherein the pressure drift is the difference between the aortic pressure and the proximal pressure; andcalculating a drift-compensated fractional flow reserve, wherein the drift-compensated fraction flow reserve corresponds to (the distal pressure+the pressure drift) divided by the aortic pressure. 2. The method of claim 1, further comprising displaying the drift-compensated fraction flow reserve on a display unit. 3. The method of claim 1, further comprising shifting the pressure sensing guidewire back to the first position and determining a second distal pressure within the body lumen with the pressure sensing guidewire. 4. The method of claim 3, further comprising shifting the pressure sensing guidewire back to the second position and determining a second proximal pressure within the body lumen with the pressure sensing guidewire. 5. The method of claim 4, further comprising calculating a second pressure drift, wherein the second pressure drift is the difference between the aortic pressure and the second proximal pressure. 6. The method of claim 1, wherein the pressure sensing guidewire includes an optical pressure sensor. 7. The method of claim 1, wherein the pressure sensing guidewire includes a Fabry-Perot pressure sensor. 8. The method of claim 1, wherein the pressure sensing guidewire includes a piezoelectric pressure sensor. 9. The method of claim 1, wherein the pressure sensing guidewire includes a proximal shaft portion and a distal shaft portion, the distal shaft portion having a plurality of slots formed therein. 10. The method of claim 9, wherein the distal shaft portion has a variable inner diameter. 11. The method of claim 1, wherein the pressure sensing guidewire includes a pressure sensor and wherein advancing the pressure sensing guidewire through a blood vessel to a first position distal of an intravascular occlusion includes advancing the pressure sensing guidewire through a guide catheter to a location where the pressure sensor is disposed distal of a distal end of the guide catheter. 12. The method of claim 1, wherein measuring an aortic pressure includes measuring the aortic pressure with a second pressure sensing medical device. 13. A method for calculating fractional flow reserve, the method comprising: advancing a pressure sensing guidewire through a body lumen of a patient to a first position proximal of an intravascular occlusion;determining a first proximal pressure within ... | BOSTON SCIENTIFIC SCIMED, INC. | A61B006/00 A61B005/02 A61B005/0215 A61B005/00
US09782090 | B1 | 2017-10-10 | Method and system enabling photoplethysmograph measurement of volume status | 2014-10-10 | A method enables photoplethysmograph measurement of volume status. The method includes the steps of converting photoplethysmograph voltages to volume measurements and characterizing a local microcirculation as a microcosm in a manner allowing a photoplethysmograph to facilitate noninvasive monitoring of systemic status. | 1. A computer implemented method for photoplethysmograph measurement of volume status and changes, comprising: measuring photoplethysmographic values at a monitoring site under a resting condition;measuring photoplethysmographic values at a monitoring site in conjunction with application of a study agent;converting photoplethysmograph voltages, obtained from the steps of measuring photoplethysmographic values at a monitoring site under a resting condition and measuring photoplethysmographic values at a monitoring site in conjunction with application of the study agent, to measurements of volume change based on AC voltage at rest (ACrestVoltage), wherein the steps of converting includes: referencing voltage change generated by portion of a stroke volume delivered to a given nonrecalibrating photoplethysmograph at a given site in a given subject under resting conditions to obtain an ACrestCalibration Voltage;establishing the ACrestCalibration Voltage as equivalent to one ACrestMultiple and resting stroke volume;applying the ACrestCalibration Voltage to convert future AC values for given photoplethysmograph at a given site to ACrestMultiples; andapplying the ACrestCalibration Voltage to convert DC values and changes thereof for given photoplethysmograph at a given site to ACrestMultiples;comparing ACrestMultiples determined based upon the steps of measuring photoplethysmographic values at a monitoring site under a resting condition and measuring photoplethysmographic values at a monitoring site in conjunction with application of the study agent thereby quantifying changes in pulsatile and non-pulsatile volume status of a local microcirculation to assess changes in local volume induced by the study agent when compared to a resting condition; andmonitoring systemic cardiovascular status based upon the assessed changes in local volume. 2. The method according to claim 1, further including the step of applying a drug micro-patch to an individual prior to the step of measuring and comparing results to those obtained without the drug micro-patch. 3. The method according to claim 1, further including increasing pressure to an untreated control site causing progressive displacement of capillary, venous, arteriole and arterial blood to a point that vascular volume no longer contributes to a photoplethysmograph signal and measuring a photoplethysmograph signal wherein a drop in voltage until prior to loss of the pulsatile signal represents a blood contribution to a DC component of a photoplethysmographic signal and remaining non-pulsatile voltage represents a DCbackground component thereby distinguishing the blood contribution to a DC component of a photoplethysmographic signal from a nonblood background component of a DC signal. 4. The method according to claim 1, further including the step of measuring arterial and venous compliance at a monitoring site. 5. The method according to claim 1, further including the step of providing uniformly scaled data wi... | - | A61K049/00 A61B005/0295 A61B005/029 A61B005/02 A61B005/00
US09808165 | B2 | 2017-11-07 | Multi-function health monitor with integrated cellular module | 2016-01-07 | A system and method for a multi-function remote ambulatory cardiac monitoring system. The system includes a housing and a microprocessor disposed within the housing. The microprocessor controls the remote ambulatory cardiac monitoring system. The system also includes an electrode for sensing ECG signals and the electrode being in communication with the microprocessor. An integrated cellular module also is included in the system, and the cellular module is connected to the microprocessor and disposed within the housing. The integrated cellular module transmits ECG signals to a remote center. | 1. A remote ambulatory cardiac monitoring system for acquiring ECG signals, the system comprising: an ambulatory housing configured to be worn by a patient;a microprocessor for controlling the remote ambulatory cardiac monitoring system, and the microprocessor is disposed within the ambulatory housing;an electrode extending from the ambulatory housing for sensing ECG signals and the electrode being in communication with the microprocessor for performing one of a multiple channel EKG/ECG mode, ECG Holter mode, or an MCT/Event mode; andan integrated cellular module connected to the microprocessor and disposed within the ambulatory housing, wherein the integrated cellular module transmits ECG signals to a remote center. 2. The system of claim 1, further comprising a controller connected to the microprocessor and disposed within the ambulatory housing, the controller switches the functionality of the remote ambulatory cardiac monitoring system by causing the microprocessor to initiate one of a plurality of different modes including the EKG/ECG mode, the ECG Holter mode and the MCT/Event mode. 3. The system of claim 2, wherein the plurality of different modes includes a USB data transfer mode. 4. The system of claim 2, further comprising a non-transitory computer readable medium associated with the microprocessor, the non-transitory computer readable medium including instructions to perform the plurality of different modes. 5. The system of claim 2, further comprising a cable socket disposed on the ambulatory housing and connected to the electrode. 6. The system of claim 2, wherein the controller switches the functionality of the remote ambulatory cardiac monitoring system to the MCT/Event mode when three electrodes are in communication with the microprocessor. 7. The system of claim 2, wherein the controller switches the functionality of the remote ambulatory cardiac monitoring system to the ECG Holter mode when five electrodes are in communication with the microprocessor. 8. The system of claim 2, wherein the controller automatically switches the functionality of the remote ambulatory cardiac monitoring system from the ECG Holter mode to the MCT/Event mode if the ECG Holter mode is negative after a period. 9. The system of claim 8, wherein the controller automatically switches the functionality of the remote ambulatory cardiac monitoring system from the ECG Holter mode to the MCT/Event mode if the ECG Holter mode is negative after at least a 24-hour period. 10. The system of claim 2, wherein the controller switches the functionality of the remote ambulatory cardiac monitoring system to the EKG/ECG mode when a multi-wire electrode is in communication with the microprocessor. 11. A method for controlling the functionality of a monitoring device, the method comprising: sensing ECG signals with an electrode extending from an ambulatory housing configured to be worn by a patient;performing one of a multiple channel EKG/ECG mode, ECG Holter mode... | - | A61B005/0402 A61B005/0245 A61B005/0404 A61B005/046 A61B005/0464 A61B005/00 A61B005/04 A61B005/0424 A61B005/0428 A61B005/0432 A61B005/0452
US09826916 | B2 | 2017-11-28 | Device and method for examining a phase distribution used to determine a pathological interaction between different areas of the brain | 2013-10-02 | A device for examining a pathological interaction between different brain areas, including a stimulation unit, which administers identical stimuli to a patient in a sequential manner, wherein the stimuli stimulate neurons of the patient in the brain areas to be examined, a measuring unit for recording measurement signals that represent a neural activity of the stimulated neurons, and a control and analysis unit for controlling the stimulation unit and for analyzing the measurement signals. The control and analysis unit transforms the measurement signals into the complex plane, examines the distribution of the phases of stimuli of the measurement signals absorbed by the measuring unit in response to the stimuli delivered to the patient, and determines the probability, with which the phase distribution differs from a uniform distribution, in order to ascertain whether a pathological interaction between the brain areas exists. | 1. An apparatus for examining a pathological interaction between different brain areas, comprising: a stimulation unit configured to successively administer identical stimuli to a patient, wherein the stimuli stimulate neurons of the patient in the brain areas to be examined;a measuring unit configured to record measured signals which reproduce a neural activity of the stimulated neurons; anda control and analysis unit configured to control the stimulation unit and analyze the measured signals, wherein the control and analysis unit is configured such that the control and analysis unit: controls the stimulation unit to administer the stimuli to the patient; andtransforms the measured signals into a complex plane, examines the distribution of phases of the measured signals recorded by the measuring unit in the complex plane as a response to the stimuli administered to the patient, and determines the probability with which a phase distribution differs from an equal distribution to determine whether a pathological interaction between the different brain areas is present. 2. The apparatus in accordance with claim 1, wherein the measuring unit comprises at least one of EEG electrodes, MEG sensors, EMG sensors, LFP sensors and implantable sensors. 3. The apparatus in accordance with claim 1, wherein the stimulation unit is non-invasive and is configured to generate stimuli from the group consisting of acoustic, visual, tactile, vibratory, proprioceptive, thermal, olfactory and electrical transcutaneous stimuli. 4. The apparatus in accordance with claim 1, wherein the stimulation unit stimulation unit comprises one or more implantable electrodes for administering electrical stimuli. 5. The apparatus in accordance with claim 1, wherein the control and analysis unit controls the stimulation unit to vary intervals between mutually following stimuli. 6. The apparatus in accordance with claim 1, wherein the control and analysis unit is configured to determine a threshold value from the measured signals recorded before the administering of a stimulus and compare the probability determined from the measured signals recorded after the administering of the stimuli with which the phase distribution differs from an equal distribution with the threshold value to determine whether a pathological interaction between the brain areas is present. 7. The apparatus in accordance with claim 1, wherein the control and analysis unit is configured to use a Kuiper test to determine the probability with which the phase distribution differs from an equal distribution. 8. The apparatus in accordance with claim 1, wherein the stimuli administered to the patient by the stimulation unit effect a phase reset of a pathologically synchronous and oscillatory activity of the stimulated neurons. 9. The apparatus in accordance with claim 8, wherein, on the determination of the probability with which the phase distribution differs from an equal distribution, the measured signals in t... | FORSCHUNGSZENTRUM JÜLICH GMBH | A61B005/04 A61B005/00 A61N001/36 A61B005/0484 A61B005/0488 A61B005/0478 A61N001/05
US09826925 | B2 | 2017-11-28 | Analyte meter with basic and advanced meter preset mode selection based on structured queries | 2012-11-29 | Described are methods and systems to allow users to select a basic mode or an advanced mode in which additional features can be customized for the user based on structured queries presented to the user. | 1. An analyte measurement meter that measures and stores results from analyte measurements to allow a user to select a basic mode or a configurable advanced mode of the meter with a user interface, the meter comprising: a housing;a plurality of user interface buttons proximate the housing;a display proximate the housing; anda microprocessor connected to the user interface buttons and the display, the microprocessor connected to a non-volatile memory stored with logical instructions programmed to cause the microprocessor following powering the meter to: upon a first time activation or a reset of the meter, perform mode and feature selection by: (a) allowing a user to select between a basic mode and a configurable advanced mode of the meter with the user interface,(b) upon selection of the configurable advanced mode with the user interface, querying the user with a series of inquiries, the series of inquiries comprising questions relating to past results of the user, meal patterns of the user, and graphical insights requested by the user,(c) receiving from the user, a series of responses corresponding to the series of inquiries, and(d) customizing the configurable advanced mode of the meter by applying an advanced mode logic of the meter to the series of responses to the series of inquiries to select any one of at least three features not available in the basic mode, the advanced mode logic comprising a correspondence between the series of responses to the series of inquiries and the at least three features not available in basic mode, wherein the meter is configured to add selected advanced mode features of the at least three features not available in the basic mode to the configurable advanced mode;upon any turn-on of the meter subsequent to the first time activation or the reset of the meter and the customizing of the configurable advanced mode, return to a main meter routine without selecting between the basic mode and the advanced configurable mode and using the customized configurable advanced mode to automatically maintain the meter to allow selection and activation of any one of the selected advanced-mode features, and present any one of the selected advanced-mode features to the user with the display; andwherein the user interface comprises a first screen and a second screen, the first screen displaying a distribution of analyte concentrations including counts for low, in-range and high analyte concentration measurements, and the second screen comprising a time-of-day distribution of analyte concentrations including counts for low, in-range and high analyte concentration measurements by time-of-day. 2. The meter of claim 1, in which one of the at least three features comprise: glycemic excursion detection functionality that provides for an indication of at least two analyte measurements outside of a predetermined range in a prior number of days (e.g., 3, 4, 5, 6 . . . n days). 3. The meter of claim 1, in which one of the at least three f... | LifeScan, Inc. | A61B005/145 A61B005/00
US09808204 | B2 | 2017-11-07 | Noninvasive physiological analysis using excitation-sensor modules and related devices and methods | 2012-08-03 | Methods and apparatus for qualifying and quantifying excitation-dependent physiological information extracted from wearable sensors in the midst of interference from unwanted sources are provided. An organism is interrogated with at least one excitation energy, energy response signals from two or more distinct physiological regions are sensed, and these signals are processed to generate an extracted signal. The extracted signal is compared with a physiological model to qualify and/or quantify a physiological property. Additionally, important physiological information can be qualified and quantified by comparing the excitation wavelength-dependent response, measured via wearable sensors, with a physiological model. | 1. A method of assessing blood pressure in an organism, the method comprising: directing light at one or more regions of the organism via at least two optical emitters, wherein at least one optical emitter generates light at a first wavelength, wherein at least one other optical emitter generates light at a second wavelength shorter than the first wavelength, and wherein the second wavelength is in the range of blue-UV;detecting, via an optical detector, light at the first and second wavelengths scattered from the organism and generating, via the optical detector, a first electrical signal associated with the first wavelength and a second electrical signal associated with the second wavelength; andprocessing, via at least one processor, the first and second electrical signals, wherein the processing comprises: filtering each of the first and second electrical signals with respect to skin motion noise to generate a first extracted signal associated with the first wavelength and a second extracted signal associated with the second wavelength that is distinct from the first extracted signal, wherein the first extracted signal comprises information about blood flow through a blood vessel of the organism, and wherein the second extracted signal comprises information about size of the blood vessel; andcomparing the first extracted signal and the second extracted signal, responsive to generation of each by the filtering with respect to the skin motion noise, in context of a physiological model to assess the blood pressure of the organism, wherein the comparing comprises:calculating the blood pressure of the organism based on an area of the blood vessel and based on a volumetric flow rate of blood therein, wherein the area is based on the size of the blood vessel and a change in the size thereof indicated by the second extracted signal, and wherein the volumetric flow rate is indicated by the first extracted signal. 2. The method of claim 1, wherein the at least two optical emitters and the optical detector are integrated into a wearable device worn by the organism. 3. The method of claim 2, wherein the wearable device comprises an earpiece module. 4. The method of claim 2, wherein the wearable device comprises a telemetric device. 5. The method of claim 1, wherein the at least two optical emitters are selected from the group consisting of laser diodes (LDs), light-emitting diodes (LEDs), and organic light-emitting diodes (OLEDs). 6. The method of claim 1, wherein the at least two optical emitters, the optical detector, and the at least one processor are integrated into a wearable device worn by the organism. 7. The method of claim 1, wherein filtering each of the first and second electrical signals with respect to skin motion noise comprises: subtracting a skin-related optical scatter signal from the first electrical signal to generate the first extracted signal containing the information about the blood flow; andsubtracting the skin-related optic... | Valencell, Inc. | A61B005/021 A61B005/1455 A61B005/00 A61B005/026 A61B005/0295 A61B005/02 A61B005/024 A61B006/00
US09782111 | B2 | 2017-10-10 | Hermetic feedthrough assembly for ceramic body | 2014-06-25 | A wire extends through a ceramic body. The wire comprises a material selected from the group consisting of platinum, palladium, rhodium, iridium, osmium and alloys of platinum, palladium, rhodium, iridium, and osmium. The wire directly contacts the ceramic body to form a substantially hermetic seal between the ceramic body and the wire. | 1. A feedthrough assembly formed from a method comprising: providing a green ceramic body;forming a through-hole in the green ceramic body having a first diameter;inserting a corrosion-resistant solid wire through the hole, the wire comprising a material selected from the group consisting of platinum, palladium, rhodium, iridium, osmium and alloys of platinum, palladium, rhodium, iridium, and osmium, and the wire having a second diameter;firing the green ceramic body to sinter it thereby forming a sintered ceramic body; andthe first and second diameters being selected such that cracking of the sintered ceramic body adjacent the wire as a result of dimensional changes of the green ceramic body during firing is avoided and a substantially hermetic seal is formed between the sintered ceramic body and the wire,wherein the green ceramic body has a thickness of at least 0.050 inches thick. 2. The feedthrough assembly of claim 1, further comprising an electrode connected to a first end of the wire. 3. The feedthrough assembly of claim 2, wherein the electrode is substantially coplanar with a surface of the sintered ceramic body. 4. The feedthrough assembly of claim 1, wherein the hermetic seal is a diffusion bond between the metal wire and the sintered ceramic body. 5. The feedthrough assembly of claim 1, wherein the wire has a diameter of less than about 0.020 inches. 6. The feedthrough assembly of claim 1, wherein the ceramic body comprises primarily compounds selected from the group consisting of oxides, carbides, borides, nitrides, and silicides of aluminum, zirconium, beryllium, silicon, titanium, hafnium, magnesium, zinc and yttrium. 7. The feedthrough assembly of claim 1, wherein the green ceramic body has a thickness of less than about 0.10 inches. 8. The feedthrough assembly of claim 1, wherein the wire has a diameter of less than about 0.020 inches and the green ceramic body has a thickness of less than about 0.10 inches. 9. The feedthrough assembly of claim 1, wherein the metal wire has a substantially round cross-section or a substantially conical shape. 10. The feedthrough assembly of claim 1, wherein the segment of wire contacting the ceramic body has a substantially spherical shape. 11. The feedthrough assembly of claim 1, further comprising a contact pad connected to a second end of the wire. 12. A feedthrough assembly formed from a method comprising: providing a green ceramic body that surrounds a corrosion-resistant solid wire extending through the green ceramic body, the wire comprising a material selected from the group consisting of platinum, palladium, rhodium, iridium, osmium and alloys of platinum, palladium, rhodium, iridium, and osmium; andfiring the green ceramic body to sinter it, thereby forming a sintered ceramic body;whereby cracking of the sintered ceramic body adjacent the wire as a result of dimensional changes of the green ceramic body during firing is avoided and a substantially hermetic seal is formed betw... | GlySense, Inc., a California Corporation | A61B005/1473 H01L023/10 H01L023/15 A61B005/145 H05K003/00 H05K003/10 H01R043/00 H01L023/498 H01L021/48
US09814390 | B2 | 2017-11-14 | Insert imaging device for surgical procedures | 2014-03-14 | Insertable imaging devices, and methods of use thereof in minimally invasive medical procedures, are described. In some embodiments, insertable imaging devices are described that can be introduced and removed from an access port without disturbing or risking damage to internal tissue. In some embodiments, imaging devices are integrated into an access port, thereby allowing imaging of internal tissues within the vicinity of the access port, while, for example, enabling manipulation of surgical tools in the surgical field of interest. In other embodiments, imaging devices are integrated into an imaging sleeve that is insertable into an access port. Several example embodiments described herein provide imaging devices for performing imaging within an access port, where the imaging may be based one or more imaging modalities that may include, but are not limited to, magnetic resonance imaging, ultrasound, optical imaging such as hyperspectral imaging and optical coherence tomography, and electrical conductive measurements. | 1. A magnetic resonance imaging probe comprising: a longitudinal body;first, second and third magnetic resonance coils supported by said longitudinal body;wherein said first coil is configured to measure fields having a first direction within a region of interest beyond a distal end of said longitudinal body;said first coil comprising: one or more magnetic resonance coils housed within said longitudinal body, wherein at least one coil is a folded stripline coil comprising: two longitudinal stripline conductors having a ground plane conductor provided therebetween;a folded conductor segment connecting said two longitudinal stripline conductors near a distal portion of said longitudinal body;wherein said second coil is configured to measure fields having a second direction within a region of interest beyond the distal end of said longitudinal body, wherein said first direction and said second direction are approximately orthogonal;said second coil comprising: an additional folded stripline coil comprising:additional longitudinal stripline conductors oriented having said ground plane conductor provided therebetween; andan additional folded conductor segment connecting said two additional longitudinal stripline conductors near said distal portion of said longitudinal body;wherein said two additional longitudinal stripline conductors are oriented such that said folded conductor segment and said additional folded conductor segment are approximately orthogonal, thereby forming a folded stripline quadrature pair;wherein said third coil is configured to measure fields having a third direction within the region of interest beyond said distal end of said longitudinal body, wherein said third direction is approximately orthogonal to said first direction and said second direction;said third coil comprising: a coil loop provided near said distal portion of said longitudinal body, said coil loop having a coil axis that is approximately parallel to a longitudinal axis of said longitudinal body;electrical circuits housed within said longitudinal body for tuning and matching said first, second and third coils and preamplifying signals detected by said first, second, and third coils;said electrical circuits comprising: a pair of matching capacitors, each matching capacitor provided between one of said longitudinal stripline conductors and said ground plane conductor;a tuning capacitor serially provided within one of said longitudinal stripline conductors;a preamplifier circuit housed within said longitudinal body, wherein said preamplifier circuit is operatively coupled to said folded stripline coil;a pair of additional matching capacitors, each additional matching capacitor provided between one of said additional longitudinal stripline conductors and said ground plane conductor; andan additional preamplifier circuit housed within said longitudinal body, wherein said additional preamplifier circuit is operatively coupled to said additional folded stripline coil. 2... | SYNAPTIVE MEDICAL (BARBADOS) INC. | A61B005/05 A61B005/00 A61B005/055 G01R033/34 G01R033/345 A61B008/12 A61B008/00 A61B017/34 A61B008/08 A61B090/00 A61B034/20
US09804083 | B2 | 2017-10-31 | Optical standard for calibration of spectral measuring systems | 2014-02-11 | Examples of an optical standard and a calibration apparatus for calibrating or characterizing a spectroscopy system using such optical standard are disclosed. The optical standard can comprises a mixture of acetaminophen and barium sulfate, wherein a mass of the acetaminophen in the mixture is being less than a mass of the BaSO4. Such optical standard can be used in a calibration device for calibrating or characterizing a spectroscopy system. The calibration device can comprise a substrate base with a top surface and a bottom surface. The top surface can include a section for receiving the optical standard sample. The receiving section can be adhesive. The calibration device can further comprise a film that can be attached to the top surface of the substrate base to cover at least the section of the substrate where the optical standard is being placed. The calibration device can be disposed after the calibration measurements are completed. The optical standard and the calibration apparatus using the optical standard can be used as a wavelength calibration standard to calibrate a Raman system, a reflectance reference standard for a reflectance spectral measurement or for a reliability check in a fluorescence spectral system. | 1. A calibrating device for calibrating or characterizing a spectral measuring system, the device comprising: a substrate base having a top surface and a bottom surface, the base being configured to be detachably placed to a distal end of a spectral probe; anda receiving section positioned at the top surface of the substrate base for receiving an optical standard, the receiving section configured to fix the optical standard thereto, the receiving section configured to cover the distal end of the probe when the base is placed to the end of the probe,wherein the optical standard comprises a mixture of an acetaminophen and a barium sulfate, wherein a mass of the acetaminophen is being less than a mass of the barium sulfate. 2. The calibrating device of claim 1, further comprising a cover sized to protect the optical standard and the receiving section. 3. The calibrating device of claim 1, wherein the substrate base comprises a lip extending from an edge of the base configured to be detachably placed to a side wall of the distal tip such that the receiving section covers the distal end of the probe. 4. The calibrating device of claim 1, wherein the substrate base is selected from a group comprising a black aluminum foil, white aluminum foil and black paper tape. 5. The calibrating device of claim 1, wherein the device is reusable. 6. The calibrating device of claim 1, wherein a ratio of the acetaminophen and the barium sulfate is 1:2. 7. A spectroscopy probe with a detachable calibration device, the probe comprising: a housing having a proximate end and a distal end and longitudinal axis between the proximate end and the distal end;at least one light source for generating illumination light beam;means for receiving the illumination beam from said at least one light source and for directing it toward an object under examination;means for receiving a returning radiation from the object and directing at least part of the returning radiation to a detector;a head position at the distal end of the probe having a body and a tapering tip, an opening formed at the tip of the head for the illumination light beam and the returning radiation to pass there through; anda detachable calibration device attached to the tip of the head to cover the opening, the calibration device comprising an optical standard with a mixture of an acetaminophen and a barium sulfate, wherein a mass of the acetaminophen is being less than a mass of the barium sulfate. 8. The spectroscopy probe of claim 7, wherein the head of the probe being removable and disposable. 9. The spectroscopy probe of claim 7, further comprising a grooved channel formed at the tip, the channel extending around the opening. 10. The spectroscopy probe of claim 7, wherein a ratio of the acetaminophen and the barium sulfate is 1:2. 11. A method for the spectral calibration or characterization of a spectral measuring system comprising the steps of: providing a calibration device with an optical standard c... | - | G01J003/28 G01N021/27 G01N021/65 G01N021/64 A61B005/00
US09776009 | B2 | 2017-10-03 | Non-invasive detection of phrenic nerve stimulation | 2014-03-20 | Systems, methods, and graphical user interfaces are described herein for non-invasively detecting phrenic nerve stimulation during cardiac pacing therapy. Phrenic nerve stimulation information may be generated for one or more electrical pacing vectors at one or more power configurations. The phrenic nerve stimulation information may be displayed to a user for use in configuring and/or evaluating cardiac pacing therapy. | 1. A non-invasive system for detecting phrenic nerve stimulation during pacing therapy comprising: external electrode apparatus comprising a plurality of external electrodes configured to be located proximate tissue of a patient; andcomputing apparatus coupled to the external electrode apparatus and configured to: monitor electrical activity using two or more external electrodes of the plurality of external electrodes during delivery of pacing therapy, wherein such monitoring results in two or more electrical signals measured using the two or more external electrodes indicative of whether or not phrenic nerve stimulation occurred as a result of the pacing therapy delivered, anddetermine whether the patient's phrenic nerve is stimulated by the pacing therapy based on the monitored electrical activity, wherein determination of whether the patient's phrenic nerve is stimulated comprises comparison of each of the electrical signals measured using the two or more external electrodes or a value determined from the electrical signals measured using the two or more external electrodes to a phrenic nerve stimulation threshold. 2. The system of claim 1, wherein the two or more external electrodes of the plurality of external electrodes are configured to be located on the anterior torso of the patient. 3. The system of claim 1, wherein the external electrode apparatus comprises one of a band and a vest configured to be worn about the torso of the patient, wherein the plurality of external electrodes are coupled to the band or vest. 4. The system of claim 1, wherein monitoring electrical activity using two or more external electrodes of the plurality of external electrodes during delivery of pacing therapy comprises monitoring electrical activity using two or more external electrodes of the plurality of external electrodes for a selected time period after delivery of a pacing pulse. 5. The system of claim 1, wherein determining whether the patient's phrenic nerve is stimulated by the pacing therapy comprises comparing a maximum peak-to-peak amplitude of the monitored electrical activity to a threshold value. 6. The system of claim 5, wherein the threshold value is based on baseline measurements monitored using one or more external electrodes of the plurality of external electrodes during one of intrinsic rhythm and atrial pacing, wherein the threshold value is greater than or equal to 3 times a baseline value based on the baseline measurements. 7. The system of claim 1, wherein determining whether the patient's phrenic nerve is stimulated by the pacing therapy comprises determining that the patient's phrenic nerve is stimulated if the monitored electrical activity from a selected percentage of the two or more external electrodes used to monitor electrical activity during delivery of pacing therapy is indicative of phrenic nerve stimulation. 8. A non-invasive method for detecting phrenic nerve stimulation during pacing therapy comprising: monitoring el... | Medtronic, Inc. | A61N001/00 A61N001/37 A61B005/04 A61B005/11 A61N001/372 A61B005/00 A61B005/0408
US09791536 | B1 | 2017-10-17 | Mutually calibrated magnetic imaging array | 2017-04-28 | A mutually calibrated magnetic imaging array system is described. The system includes a non-target magnetic source rigidly attached to a magnetometer, and an attached control unit to measure and adjust several parameters of a magnetic imaging array. A non-target magnetic field source is used to generate a well-defined and distinguishable spatial magnetic field distribution. The source is rigidly attached directly to a magnetometer, while the relative positions of the magnetometers are unknown. The magnetic imaging array is used to measure the strength of the non-target source magnetic fields and the information is used to calibrate several parameters of the array, such as, but not limited to, effective magnetometer positions and orientations with respect to each other and cross-talk between the magnetometers. The system, and method described herein eliminates the need for a separate calibration phantom. | 1. A magnetic imaging system, the system comprising: a. at least one non-target magnetic source capable of creating a known magnetic field pattern;b. at least one magnetometer, wherein the magnetometer is rigidly attached to the at least one non-target magnetic source;c. at least one additional magnetometer, wherein the additional magnetometer is capable of detecting the known magnetic field pattern and creating a magnetic field measurement; andd. a computational unit that uses the magnetic field measurement to generate at least one calibration parameter of the magnetic imaging system. 2. The system of claim 1, wherein the magnetometers are optically-pumped magnetometers. 3. The system of claim 1, wherein the at least one non-target magnetic source is a dipolar source. 4. The system of claim 1, wherein the at least one non-target magnetic source is a field coil of the at least one magnetometer. 5. The system of claim 1, wherein the calibration parameter of the magnetic imaging system is related to the relative position of the magnetometers. 6. The system of claim 1, wherein the calibration parameter of the magnetic imaging system is the relative orientation of the magnetometers. 7. The system of claim 1, wherein the calibration parameter of the magnetic imaging system is the cross-talk between the magnetometers. 8. The system of claim 1, wherein the computational unit is attached to the at least one magnetic source and the at least one magnetometer. 9. A magnetic imaging system, the system comprising: a. a flexible housing;b. at least one device, the device comprising a magnetic source and a magnetometer rigidly attached to one another;c. at least one additional magnetometer capable of taking a magnetic field measurement of the magnetic source; andd. a computational unit that uses the magnetic field measurement to generate at least one calibration parameter of the magnetic imaging system; andwherein, the device and the at least one additional magnetometer are fitted to the flexible housing in a secure manner. 10. The system of claim 9, wherein the magnetometers are optically-pumped magnetometers. 11. The system of claim 9, wherein the at least one non-target magnetic source is a dipolar source. 12. The system of claim 9, wherein the at least one non-target magnetic source is a field coil of the device. 13. The system of claim 9, wherein the computational unit is attached to the device. 14. A method of calibrating a magnetic imaging array, the method comprising the steps of: a. rigidly attaching at least one non-target magnetic source to as least one magnetometer;b. using the least one non-target magnetic source to create a known magnetic field pattern;c. using at least one additional magnetometer to measure the magnetic field of the non-target source to create a magnetic field measurement; andd. using the magnetic field measurement to produce a calibration parameter of the imaging array. 15. The method of claim 10, wherein step c us... | QuSpin, Inc. | G01V003/00 G01R035/00 G01R033/032 A61B005/04
US09804150 | B2 | 2017-10-31 | Method and apparatus for providing data processing and control in a medical communication system | 2014-03-24 | Methods and apparatus for providing data processing and control for use in a medical communication system are provided. | 1. A method, comprising: determining, using one or more processors, a sensitivity value for each of a predetermined number of analyte sensors;determining, using the one or more processors, a sensitivity variation for each of the predetermined number of analyte sensors;determining, using the one or more processors, a mean sensitivity based on the sensitivity value determined for each of the predetermined number of analyte sensors when it is determined that the determined sensitivity variations each do not exceed a tolerance threshold level;associating, using the one or more processors, a sensor code with an analyte sensor batch associated with the predetermined number of analyte sensors when the mean sensitivity is within a predetermined sensitivity range; andproviding a plurality of packagings for the analyte sensor batch, each packaging including the sensor code. 2. The method of claim 1, wherein the sensitivity variation for each of the predetermined number of analyte sensors is determined as a coefficient of variation. 3. The method of claim 1, wherein the sensitivity value for each of the predetermined number of analyte sensors is determined in vitro. 4. The method of claim 1, wherein determining the sensitivity value for each of the predetermined number of analyte sensors is performed during an analyte sensor manufacturing process. 5. The method of claim 1, wherein the sensor code is used to calibrate the predetermined number of analyte sensors associated with the analyte sensor batch associated with the predetermined number of analyte sensors. 6. The method of claim 1, wherein the sensor code is stored in an analyte monitoring device. 7. The method of claim 1, further comprising determining a response time for each of the predetermined number of analyte sensors. 8. The method of claim 7, further comprising determining a mean response time of the predetermined number of analyte sensors. 9. The method of claim 8, wherein the mean response time is used for lag correction of data obtained using analyte sensors included in the analyte sensor batch associated with the predetermined number of analyte sensors. 10. The method of claim 1, wherein each of the predetermined number of analyte sensors comprises a plurality of electrodes including a working electrode, wherein the working electrode comprises an analyte-responsive enzyme and a mediator, wherein at least one of the analyte-responsive enzyme and the mediator is chemically bonded to a polymer disposed on the working electrode, and wherein at least one of the analyte-responsive enzyme and the mediator is crosslinked with the polymer. 11. The method of claim 1, further comprising, upon determining that the determined sensitivity variation for an analyte sensor of the predetermined number of analyte sensors exceeds the tolerance threshold level, rejecting the analyte sensor batch associated with the predetermined number of analyte sensors. 12. The method of claim 1, wherein determining... | Abbott Diabetes Care Inc. | G01N033/49 A61B005/145 A61B005/1473 A61B005/1495 G06F019/00
US09788755 | B2 | 2017-10-17 | Illumination systems and devices for tracheal tubes | 2011-05-26 | Various embodiments of an intubation system include a tracheal tube and an illumination assembly that is removably couplable to a tubular body of the tracheal tube. The tracheal tube may be a double lumen tracheal tube having a first cuff that is adapted to be inflated to seal against the walls of a patient's trachea and a second cuff that is adapted to be inflated to seal against the walls of the patient's bronchial stem. The illumination assembly may have one or more illumination devices that are adapted to produce light within the patient's trachea, the patient's bronchial stem, or both when the illumination assembly is coupled to the tubular body. | 1. An intubation system, comprising: a tracheal tube comprising a first ventilation lumen having a first distal end and a first proximal end and a second ventilation lumen adjacent to the first ventilation lumen, the second ventilation lumen having a second distal end and a second proximal end, the second distal end not being adjacent to the first distal end, wherein a contiguous dividing wall forming part of and dividing the first ventilation lumen and the second ventilation lumen extends beyond the first distal end to a portion of an exterior wall of the second ventilation lumen, and wherein opposing surfaces of the contiguous dividing wall form at least a portion of respective interior surfaces of the first ventilation lumen and the second ventilation lumen that are in contact with a respiratory gas flow;a first cuff disposed about the first ventilation lumen and the second ventilation lumen above the first distal end and a second cuff disposed about the second ventilation lumen above the second distal end and below the first distal end;an imaging device disposed on a collar affixed to an external wall surface of the first and second lumens below the first cuff and proximal to the first distal end, wherein the collar is not in contact with the first cuff, and wherein the imaging device is configured to generate a visual representation of a portion of a patient's airway; andan illumination assembly coupled to the second ventilation lumen and positioned proximally relative to the imaging device configured to generate light when the illumination assembly is coupled to the tracheal tube and the tracheal tube is disposed in the airway of the patient. 2. The intubation system of claim 1, wherein the illumination assembly comprises a flexible circuit board having one or more light emitting diodes (LEDs) disposed thereon. 3. The intubation system of claim 1, wherein the illumination assembly comprises one or more illuminating devices coupled to one or more heat dissipating and/or spreading pads configured to transfer heat from the one or more illuminating devices. 4. The intubation system of claim 1, wherein the second cuff is a bronchial cuff, and the illumination assembly is coupled to the second ventilation lumen below the bronchial cuff and is configured to direct the generated light into the bronchial cuff to illuminate the bronchial cuff. 5. The intubation system of claim 1, wherein the illumination assembly comprises one or more illuminating devices encased in a third cuff configured to be inflated to produce a volume between the illuminating devices and a wall of the third cuff. 6. The intubation system of claim 1, comprising an illumination control system coupled to the illumination assembly via a lumen of the tracheal tube and configured to power the illumination assembly and to control generation of light from the illumination assembly. 7. The intubation system of claim 1, wherein the tracheal tube comprises a lumen originating from t... | COVIDIEN LP | A61B001/04 A61B005/06 A61M016/04
US09782776 | B2 | 2017-10-10 | Collection container assembly with nanocomposite barrier coating | 2009-03-19 | An evacuated container assembly suitable for use in connection with blood collection including: (a) a container member formed of a first polymeric material and having a sidewall and one or more openings; (b) a nanocomposite barrier coating disposed on the container member having a thickness of up to about 30 microns and being derived from an aqueous dispersion including (i) a dispersed barrier matrix polymer; and (ii) a substantially exfoliated silicate filler having an aspect ratio of more than 50; and (c) one or more sealing members disposed in the opening(s) operative to hermetically seal the cavity; wherein the cavity is evacuated and maintains a pressure below atmospheric pressure and exhibits a draw volume loss lower than that of a like assembly without a nanocomposite barrier film by a factor of at least 1.5. | 1. An evacuated container assembly suitable for use in connection with biological sample collection comprising: (a) a container member formed of a first polymeric material and having a sidewall defining a cavity with an inner surface as well as one or more open ends;(b) a nanocomposite barrier coating disposed on a container member surface having a thickness of up to about 30 microns and being derived from an aqueous dispersion including (i) a dispersed barrier matrix polymer; and (ii) a substantially exfoliated silicate filler having an aspect ratio of more than 50; and(c) a sealing member disposed in the opening(s) of the container member operative to hermetically seal the cavity;wherein the cavity is evacuated and maintains a pressure below atmospheric pressure and exhibits a draw volume loss lower than that of a like assembly without a nanocomposite barrier film by a factor of at least 1.5. 2. The evacuated container assembly according to claim 1, wherein the assembly is used in connection with blood collection. 3. The evacuated container assembly according to claim 1, wherein the assembly exhibits a draw volume loss lower than that of like assembly without a nanocomposite barrier film by a factor of at least 2. 4. The evacuated container assembly according to claim 1, wherein the assembly exhibits a draw volume loss lower than that of like assembly without a nanocomposite barrier film by a factor of at least 5. 5. The evacuated container assembly according to claim 1, wherein the assembly exhibits a draw volume loss lower than that of like assembly without a nanocomposite barrier film by a factor of at least 10. 6. The evacuated container assembly according to claim 1, wherein the assembly exhibits a draw volume loss lower than that of like assembly without a nanocomposite barrier film by a factor of at least 25. 7. The evacuated container assembly according to claim 1, wherein the assembly exhibits a draw volume loss lower than that of like assembly without a nanocomposite barrier film by a factor of at least 40 up to a factor of 100. 8. The evacuated container assembly according to claim 1, having an overall length, a predetermined volume, as well as an outside diameter and wherein the sealing member is needle pierceable. 9. The evacuated container assembly of claim 8, wherein the outside diameter is between 5 and 15 mm. 10. The evacuated container assembly of claim 8, wherein the total length is between 40 mm and 140 mm. 11. The evacuated container assembly of claim 8, wherein the maximum fill volume is between 400 ml and 30 cc. 12. The evacuated container assembly of claim 8, wherein the maximum fill volume is between 2 and 10 cc. 13. The evacuated container assembly of claim 1, wherein the sidewall has an average thickness between 5 mm and 1.5 mm, wherein a thickness of the nanocomposite barrier coating is less than 20 microns. 14. The evacuated container assembly of claim 1, wherein an average combined thickness of the bar... | INMAT, INC. | B01L003/14 B01L003/00 A61B005/15 A61B005/154
US09826919 | B2 | 2017-11-28 | Construction of a non-imaged view of an object using acquired images | 2008-09-30 | There is described a method and a system for generating a non-imaged plane view of an anatomical part for assisting a surgery. A model of the anatomical part is provided, comprising features and dimension parameters. A first plane view of the anatomical part is obtained. Features are identified in the first plane view. The dimension parameters are measured from the features identified. The model is displayed in the non-imaged plane view such that the model is dimensioned with the measured dimension parameters. A position and orientation of a surgical object is then tracked relative to the anatomical part while a user navigates the surgical object in space and a representation of the surgical object to be overlaid in register on at least one the first plane view and/or the non-imaged view is generated. | 1. A method for generating a display of a tracked surgical object relative to a non-imaged scaled general model in plane view of a vertebra in computer-assisted surgery, the method comprising: receiving a general model of the vertebra comprising features and dimension parameters related to said features, the general model being non patient specific;obtaining at least a first plane view of the anatomical part;identifying said features in said at least first plane view;measuring values of said dimension parameters from said features identified in said at least first plane view;generating said non-imaged scaled general model in a transverse plane view of the vertebra by scaling said general model with said dimension parameters adjusted to said values;tracking a surgical object relative to said vertebra while a user manipulates said surgical object in space; anddisplaying on a display device of a computer-assisted surgery system by generating a representation of said surgical object to be overlaid in register on said transverse plane view of the vertebra, using the tracking of the surgical object. 2. The method as claimed in claim 1, further comprising obtaining a second plane view of the vertebra, said first plane view and second plane view being obtained along different orientations relative to said vertebra; wherein said identifying said features comprises identifying said features in said first plane view and said second plane view; andwherein said measuring values comprises measuring values of said dimension parameters from said features identified in said first plane view and said second plane view. 3. The method as claimed in claim 1, further comprising: determining whether said first plane view is in alignment with a corresponding plane defined by the object; andif said first plane view is not in alignment:correcting the alignment; andobtaining again said first plane view of the object with the corrected alignment. 4. The method as claimed in claim 2, wherein said first plane view is obtained along an antero-posterior plane of the vertebra and said second plane view is obtained along a sagittal plane of the vertebra. 5. The method as claimed in claim 1, wherein: identifying said features comprises identifying a side vertebral body contour, a pedicle contour and a spinous process contour. 6. The method as claimed in claim 5, wherein measuring values comprises measuring a distance between said side vertebral body contour and said pedicle contour, and a distance between said spinous process contour and said pedicle contour. 7. The method as claimed in claim 2, wherein said generating comprises generating said first plane view, said second plane view and said non-imaged scaled general model in a same graphical user interface of the computer-assisted surgery system. 8. The method as claimed in claim 1, wherein said obtaining a first plane view comprises acquiring said first plane view using a fluoroscopic imaging device. 9. A computer-assi... | ORTHOSOFT, INC. | G06K009/00 A61B005/107 A61B006/00 A61B034/20 A61B005/055 A61B005/00 A61B006/12 G06F019/00 A61B034/10
US09782089 | B2 | 2017-10-10 | Worksheet system for determining measured patient values for use in clinical assessment and calculations | 2011-05-16 | A system determines measured patient values for use in clinical calculations using an electronic form including, a first area including data fields for presenting values of the parameters associated with a first part of a cardiac catheterization study of a patient and a second area including data fields for presenting values of the parameters associated with a different second part of a cardiac catheterization study of the patient. A user interface enables a user to copy at least one of the parameters comprising a measured value from the first area to the second area as a substitute value eliminating a need for a re-measurement of the value. A calculation processor automatically calculates a cardiac flow value for incorporation in the second area in response to the measured value being copied into the second area. | 1. A system for determining measured patient values for use in clinical calculations, comprising: a repository operative to record values of parameters;an electronic form comprising a single display image including labels for values of parameters including blood pressure, ventricular assessments and cardiac flows and a plurality of image areas including, a first area of the single display image including data fields for presenting values of said parameters associated with a first part of a cardiac catheterization study of a patient; anda second area of the single display image including data fields for presenting values of said parameters associated with a different second part of a cardiac catheterization study of said patient, wherein the values presented in the first area of the single display image and the values presented in the second area are associated with the same parameters and at least one value presented in the first area and associated with a first parameter is different from a value presented in the second area and associated with the first parameter;a user interface enabling a user to copy a second measured value presented in the first area and associated with a second parameter to said second area of the single display image as associated with the second parameter in the second part of the cardiac catheterization study, said user interface displaying a visual attribute in said second area to indicate that the second measured value in the second area is a substitute value which was measured in the first part of the cardiac catheterization study and was not measured in the second part of the cardiac catheterization study and has been copied to said second area from said first area, the second measured value to be used as an input in calculations instead of newly measured data; anda calculation processor for automatically calculating a cardiac flow value for the second part of the cardiac catheterization study using the second measured value in response to said second measured value being copied to said second area, and automatically simulating hemodynamic effects in a condition based on the copied second measured value, wherein the calculation processor includes a non-transitory memory operative to store instructions regarding the automatic calculation of the cardiac flow value. 2. The system according to claim 1, wherein said parameters include oximetry assessments. 3. The system according to claim 1, wherein said cardiac flow value comprises at least one of (a) a stroke value and (b) a cardiac output value. 4. The system according to claim 1, wherein the first and second parts of the cardiac catheterization study comprise different portions of a cardiac catheterization study. 5. The system according to claim 1, wherein the first and second areas comprise substantially adjacent first and second columns respectively. 6. The system according to claim 1, wherein the first and second areas comprise substantially adjacent first an... | Siemens Healthcare GmbH | A61B005/026 A61B005/029 G06F019/00 A61B005/0205 A61B005/021 A61B005/0215 A61B005/08 A61B005/1455
US09782107 | B1 | 2017-10-10 | Centering cursor | 2012-12-03 | Aspects of the disclosure relate generally to aligning a patient with a target position of a magnetic resonance imaging system. For example, a computer may receive an input identifying a set of coordinates to be aligned with a target position of the magnetic resonance imaging system. The coordinates may correspond to a position of an anatomical feature of a patient within the magnetic resonance imaging system. The computer may also calculate a first travel distance and a second travel distance. Each of the first and second travel distances may be a distance along first and second axes, respectively, along which a patient handling system of the magnetic resonance imaging system is capable of moving. The computer may further reposition the patient handling system according to both the first travel distance and the second travel distance such that the anatomical feature is aligned with the target position. | 1. A method of aligning a patient positioned on an inclined bed with a target position of a magnetic resonance imaging system, the method comprising: displaying one or more cross-sectional images of the patient, at least one of the cross-sectional images featuring an anatomical feature of interest of the patient, said one of the cross-sectional images forming a plane having first and second axes in a first coordinate space, the first coordinate space further comprising a third axis orthogonal to the plane of said cross-sectional image;receiving, at a computer, a positional input corresponding to a location of the anatomical feature of interest in the first coordinate space of said cross-sectional image;receiving, at the computer, an angular input indicating a degree of incline of the bed on which the patient is positioned, wherein the bed is capable of being translated along a plurality of axes including a first and second axes of a second coordinate space, said first and second axes of the second coordinate space comprising a lengthwise direction of the inclined bed and a lengthwise direction of the bed in an uninclined state;determining, at the computer, a distance between the location of the anatomical feature of interest and the target position of the magnetic resonance imaging system along each of the first and second axes of the first coordinate space based on the input corresponding to the location of the anatomical feature of interest in the cross-sectional image;calculating, at the computer, a coordinate transformation of the determined distance from the first coordinate space to the second coordinate space based on the angular input, wherein the calculation converts the determined distance into first and second travel distances along the first and second axes of the second coordinate space, respectively; andrepositioning the inclined bed within the magnetic resonance imaging system according to both the first travel distance and the second travel distance. 2. The method of claim 1, wherein the first axis of the second coordinate space corresponds to a degree of freedom of the magnetic resonance imaging system along which a carriage supporting the inclined bed is capable of moving. 3. The method of claim 2, wherein the second axis of the second coordinate space corresponds to a degree of freedom of the magnetic resonance imaging system along which the inclined bed is capable of moving independent of the carriage. 4. The method of claim 1, further comprising: calculating first and second maximum ranges along each of the first and second axes of the second coordinate space, respectively;comparing the first maximum range to the first travel distance;comparing the second maximum range to the second travel distance;if neither of the travel distances exceeds the corresponding maximum range, repositioning the inclined bed within the magnetic resonance imaging system according to both the first travel distance and the second travel distance. ... | Fonar Corporation | A61B005/00 A61B005/055
US09788729 | B2 | 2017-10-17 | Image processing apparatus and image processing method | 2009-11-18 | A storage unit stores volume data including a blood vessel region. An extracting unit extracts the blood vessel region from the volume data. A specifying unit specifies a position of a region of interest in the blood vessel region and a deflection direction of a blood vessel region included in the region of interest. A determining unit determines a viewing position and a viewing direction based on the position of the region of interest and the deflection direction. A generating unit generates image data concerning the viewing position and the viewing direction based on the volume data. A display unit displays an image represented by the image data. | 1. An image processing apparatus comprising: a storage memory configured to store volume data including a blood vessel region;a display; andprocessing circuitry configured to: extract the blood vessel region from the volume data stored in the storage memory;allocate to each blood vessel branch included in the extracted blood vessel region a label indicating a class name of the blood vessel branch;specify a position of a region of interest in the extracted blood vessel region and a deflection direction of a blood vessel region included in the region of interest based upon a deflection parameter, indicating the degree of deflection of a blood vessel wall included in the extracted blood vessel region;determine a viewing direction so that the viewing direction is perpendicular to the deflection direction;generate image data concerning the viewing direction based on the volume data stored in the storage memory;display an image represented by the image data on the display;determine whether or not the label should be displayed by comparing a length ratio between the blood vessel region included in the image displayed on the display and the blood vessel region in the volume data with a predetermined value; andarrange, in case where the label is determined to be displayed, the label at the image displayed on the display. 2. The image processing apparatus according to claim 1, wherein the region of interest is designated by instructions of a user or by image processing. 3. The image processing apparatus according to claim 1, wherein the processing circuitry specifies the position of the region of interest and the deflection direction based on the deflection parameter, the deflection parameter being concerned with the blood vessel region included in the region of interest. 4. The image processing apparatus according to claim 3, wherein the deflection parameter is a blood vessel diameter of the blood vessel region, a blood vessel constriction rate defined as a ratio between the diameter of a normal part in the blood vessel region and the diameter of an lesion part of the blood vessel, or an arteriosclerosis index defined as a ratio between outer wall diameter and inner wall diameter of the blood vessel region. 5. The image processing apparatus according to claim 3, wherein the processing circuitry sets a plurality of sections almost orthogonal to an axis of the blood vessel region, calculates a plurality of the deflection parameters for the plurality of section, specifies the largest one of the plurality of the deflection parameters, and specifies the position of the section having the largest one of the deflection parameters, as the position of the region of interest. 6. The image processing apparatus according to claim 3, wherein the processing circuitry sets a section almost orthogonal to an axis of the blood vessel region, specifies a direction of a blood vessel diameter in the section, and specifies the direction of the blood vessel diameter as the... | TOSHIBA MEDICAL SYSTEMS CORPORATION | G06K009/00 A61B005/00 A61B005/02 G06T011/00 G06T015/08 G06T019/00
US09782129 | B2 | 2017-10-10 | Pressure sensing guidewires | 2015-07-30 | Medical devices and methods for making and using medical devices are disclosed. An example medical device may include a medical device for measuring blood pressure. The medical device may include an elongated shaft having a proximal region and a distal region. An optical fiber may extend along the proximal region. An optical pressure sensor may be coupled to the optical fiber. The optical pressure sensor may be disposed along the distal region. A centering member may be coupled to the optical fiber and positioned adjacent to the optical pressure sensor. | 1. A medical device for measuring blood pressure, comprising: an elongated shaft having a proximal region and a distal region;wherein the proximal region has a first wall thickness;wherein the distal region has a second wall thickness different from the first wall thickness;an optical fiber extending along the proximal region;an optical pressure sensor coupled to the optical fiber, the optical pressure sensor being disposed along the distal region; anda centering member coupled to the optical fiber and positioned within the distal region of the shaft at a position adjacent to the optical pressure sensor. 2. The medical device of claim 1, wherein the centering member includes a polymer. 3. The medical device of claim 1, wherein the centering member has a substantially circular cross-sectional shape. 4. The medical device of claim 1, wherein the centering member has a non-circular cross-sectional shape. 5. The medical device of claim 1, wherein the centering member includes a coil. 6. The medical device of claim 1, wherein the centering member is disposed at a proximal end of the optical pressure sensor. 7. The medical device of claim 1, wherein the centering member is spaced from a proximal end of the optical pressure sensor. 8. The medical device of claim 1, wherein the shaft has a plurality of slots formed therein. 9. The medical device of claim 8, wherein at least some of the slots are positioned adjacent to the centering member. 10. The medical device of claim 1, wherein the distal region has a distal inner diameter, wherein the proximal region has a proximal inner diameter, and wherein the distal inner diameter is greater than the proximal inner diameter. 11. A pressure sensing guidewire, comprising: a tubular member having a proximal region and a distal region;wherein the distal region has a distal inner diameter;wherein the proximal region has a proximal inner diameter;wherein the distal inner diameter is larger than the proximal inner diameter;an optical fiber extending within the tubular member along the proximal region;an optical pressure sensor coupled to the optical fiber, the optical pressure sensor being disposed within the tubular member along the distal region;a centering member disposed within the tubular member along the distal region;wherein the centering member is coupled to an outer surface of the optical fiber; andwherein the centering member is coupled to an inner surface of the distal region of the tubular member. 12. The guidewire of claim 11, wherein the centering member includes a polymer. 13. The guidewire of claim 11, wherein the centering member has a substantially circular cross-sectional shape. 14. The guidewire of claim 11, wherein the centering member has a non-circular cross-sectional shape. 15. The guidewire of claim 11, wherein the centering member includes a coil. 16. The guidewire of claim 11, wherein the centering member is disposed at a proximal end of the optical pressure sensor. 17. The gu... | BOSTON SCIENTIFIC SCIMED, INC. | A61B005/00 A61B005/0215
US09778280 | B2 | 2017-10-03 | Methods and systems for identification of event data having combined activity and location information of portable monitoring devices | 2014-12-03 | Systems and methods for segmenting a period of time into identification of locations of a user who performs activities are described. One of the methods includes detecting activity of a monitoring device worn by the user. The activity includes an amount of movement of the monitoring device. The activity is performed for a period of time. The method further includes obtaining geo-location data for the monitoring device and storing, during the period of time, the detected activity and corresponding geo-location data. The method also includes analyzing the detected activity and the corresponding geo-location data to identify one or more events. Each event is associated with a group of activity data and one or more of the groups of activity data is associated with an identifier, which is obtained using the geo-location data. | 1. A method comprising: receiving by a server coupled to a computer network activity data of a monitoring device that is configured to be worn by a user, the activity data indicating a plurality of amounts of movements of the monitoring device while being engaged in different activities during a day;receiving by the server via the computer network geo-location data for the monitoring device, wherein the geo-location data is obtained from a global positioning system;receiving by the server via the computer network times at which the activity data is gathered by the monitoring device;storing by the server the activity data, the times, and the geo-location data; andidentifying by the server two or more events during the day by analyzing the activity data, the times, and the geo-location data for the day, wherein one of the two or more events includes activity levels representing a subset of the activity data and includes an identifier of a location at which the activity data of the subset is gathered, wherein another one of the two or more events includes activity levels representing another subset of the activity data and includes another identifier of one of the different activities,wherein identifying the two or more events includes identifying within the one of the two or more events during the day a time of entry at the location by the monitoring device and a time of exit from the location by the monitoring device, wherein the time of entry is one of the times received and the time of exit is another one of the times received,wherein identifying the time of entry comprises: calculating a speed of the monitoring device along a path to the location based on the geo-location data; andcomputing the time of entry at the location from the speed. 2. The method of claim 1, further comprising combining amounts of time of a common activity over a period of time to indicate a combined amount of time for which the common activity is performed. 3. The method of claim 1, further comprising combining activity levels of one or more same activities over a period of time to indicate a combined activity level of the one or more same activities. 4. The method of claim 1, further comprising: combining amounts of time of one or more activities performed at a common location over one or more time periods to indicate a combined amount of time for which the one or more activities are performed at the common location. 5. The method of claim 1, further comprising: combining activity levels of performance of one or more activities at a common location over one or more periods of time. 6. The method of claim 1, further comprising receiving from the user a boundary of the location of the monitoring device, the boundary including at least one longitude and latitude of the monitoring device other than a longitude and latitude of the monitoring device identified by a device locator of the monitoring device. 7. The method of claim 1, further comprising receiving by the se... | Fitbit, Inc. | G06F019/00 G01P013/00 A61B005/00 A61B005/11 G01C022/00 A61B005/0205 A61B005/021 A61B005/024 A61B005/22
US09820662 | B2 | 2017-11-21 | Catheter systems and methods for determining blood flow rates with optical sensing | 2014-10-24 | Catheter systems and methods for determining blood flow rates based on light reflection measurements. The catheter may include a lumen extending between a proximal end of the catheter and a distal end of the catheter. The catheter may include fluid infusion openings at the distal end region of the catheter that are configured to permit the indicator fluid to exit the catheter from the lumen. The catheter system may include an optical fiber having one or more sensors thereon for sensing light reflected by blood particles in a body vessel lumen. A blood flow rate may be determined based on the sensed light reflected by blood particles in the body vessel lumen. | 1. A system for determining fluid flow rate in a body lumen, the system comprising: a catheter including: a first lumen;a second lumen;a fluid infusion opening in communication with the first lumen and located at a distal end region of the catheter, the fluid infusion opening is configured to permit fluid to exit the catheter from the first lumen;an opening in communication with the second lumen and located at the distal end region of the catheter;an optical fiber having a reflective surface at a distal end thereof, wherein the optical fiber is advanceable through the second lumen of the catheter and positionable distal of the opening in communication with the second lumen; andwherein the reflective surface is substantially transparent to one or more wavelengths of light and the optical fiber is configured to receive light reflecting from particles in the body lumen to determine a flow rate of fluid in the body lumen. 2. The system of claim 1, further comprising: a beam splitter in communication with the optical fiber, wherein the beam splitter separates light reflected by the reflective surface from light reflected by particles flowing in the body lumen. 3. The system of claim 1, further comprising: a receiving detector for measuring reflected light received in the optical fiber. 4. The system of claim 1, further comprising: a Fabry-Perrot cell at the distal end of the optical fiber, wherein the reflective surface is a diaphragm of the Fabry-Perrot cell. 5. The system of any of claim 1, further comprising: a metal tube affixed to a distal end of the optical fiber and extending past a distal end of the optical fiber; anda diaphragm covering a distal opening of the metal tube. 6. The system of claim 5, wherein a space between the distal end of the optical fiber and the diaphragm forms a Fabry-Perrot cavity. 7. The system of claim 1, further comprising: one or more light reflecting surfaces positioned at a distal end of the catheter. 8. The system of any one of claim 7, wherein the light reflecting surfaces are positioned on the catheter at a position configured to reflect light toward the optical fiber. 9. A system for determining fluid flow rate in a body lumen, the system comprising: a catheter including: a first lumen;a second lumen;a fluid infusion opening in communication with the first lumen and located at a distal end region of the catheter, the fluid infusion opening is configured to permit fluid to exit the catheter from the first lumen;an opening in communication with the second lumen and located at the distal end region of the catheter;an optical fiber, wherein the optical fiber is positionable in the second lumen of the catheter and positionable distal of the opening in communication with the second lumen; andwherein a distal end of the optical fiber is substantially transparent to one or more wavelengths of light and the optical fiber is configured to receive light reflecting from particles in the body lumen to determine a flo... | BOSTON SCIENTIFIC SCIMED, INC. | A61B006/00 A61B005/026 A61B005/02 A61B005/00 G01F001/708 A61B005/028 G01N015/00 G01N015/14
US09830577 | B2 | 2017-11-28 | Computer mouse system and associated computer medium for monitoring and improving health and productivity of employees | 2013-09-24 | Provided are embodiments of systems, computer medium and computer-implemented methods for sensing health characteristics of a user using a computer mouse including, a computer medium including program instructions executable by a to cause the steps of: activating set of health sensors integrated within a computer mouse, collecting health sensor data output by the set of health sensors, determining an updated health profile for an employee using the collected health sensor data, updating health information for the employee to reflect the updated health profile for the employee, and displaying the updated health profile for the employee. | 1. A system comprising: a set of computer mouse health sensors integrated with a computer mouse for detecting biometric characteristics of health of an employee, the computer mouse being in communication with a communications network via a computer workstation associated with the employee, the set of computer mouse health sensors comprising one or more computer mouse temperature sensors, one or more computer mouse blood pressure sensors, and one or more computer mouse blood condition sensors, the set of computer mouse health sensors configured to output computer mouse health sensor data comprising computer mouse temperature data output by the one or more computer mouse temperature sensors that is indicative of a body temperature of the employee, computer mouse blood pressure data output by the one or more computer mouse blood pressure sensors that is indicative of a blood pressure of the employee, and computer mouse blood condition data output by the one or more computer mouse blood condition sensors that is indicative of a blood oxygenation of the employee,wherein the one or more computer mouse blood pressure sensors comprise a blood pressure cuff comprising a pneumatic tube extending between a body of the computer mouse and a bladder of the blood pressure cuff, wherein an air pump is located in the body of the computer mouse and is configured to supply air via the pneumatic tube to inflate the bladder of the blood pressure cuff, wherein a pressure transducer is located in the body of the computer mouse and is configured to sense air pressure within the pneumatic tube that is indicative of the air pressure within the bladder, and wherein the computer mouse blood pressure data output by the one or more computer mouse blood pressure sensors corresponds to the air pressure sensed within the pneumatic tube;a set of chair health sensors integrated with a chair for detecting biometric characteristics of the health of the employee, the set of chair health sensors comprising a chair temperature sensor configured to sense the body temperature of the employee, the set of chair health sensors configured to output chair health sensor data indicative of the body temperature of the employee; andnon-transitory computer readable storage medium comprising program instructions stored thereon, the program instructions being executable by a computer processor to cause the steps of:activating the set of computer mouse health sensors and the set of chair health sensors to generate the computer mouse health sensor data and the chair health sensor data;collecting, via the communications network, health sensor data comprising the computer mouse health sensor data and the chair health sensor data;determining a health profile for the employee using the health sensor data collected, the health profile comprising: health characteristics for the employee including one or more of the body temperature, the blood pressure, and the blood oxygenation for the employee determined u... | Saudi Arabian Oil Company | G06Q010/00 G06Q010/10 G06F019/00 A61B005/00 A61B005/01 A61B005/0205 A61B005/021 A61B005/0476 A61B005/053 A61B005/08 A61B005/1455 A61B005/16 G06Q050/22 A61B003/113
US09782098 | B2 | 2017-10-10 | Cardiovascular monitoring device | 2015-02-17 | The present invention is related to a cardiovascular monitoring device including an inflatable cuff for surrounding a limb of a user, at least a first and a second electrodes, a controlling circuitry with a processor accommodated in a housing, and a display element. The controlling circuitry is configured to perform a blood pressure measurement through controlling a pressure inside the inflatable cuff, and perform an electrocardiogram measurement by using the electrodes. The processor is also configured to provide a diastolic blood pressure and a systolic blood pressure when the blood pressure measurement is performed, and to provide a heart rhythm information when the electrocardiogram measurement is performed. Further, for achieving a better and more stable contact between the electrodes and the user's skin, the present invention provides an improved structure with electrodes arranged thereon based on the conventional blood pressure monitor. | 1. A cardiovascular monitoring device, comprising: an inflatable cuff, for surrounding a limb of a user;at least a first and a second electrodes, at least one of which is implemented as a dry electrode;a housing;a control circuitry with a processor, accommodated in the housing, wherein the control circuitry is configured to perform a blood pressure measurement through controlling a pressure inside the cuff, and perform an electrocardiogram measurement by using the electrodes; andthe processor is configured to provide a diastolic blood pressure and a systolic blood pressure when the blood pressure measurement is performed, and to provide a heart rhythm information when the electrocardiogram measurement is performed; anda display element, for showing the diastolic and systolic blood pressures and the heart rhythm information,whereinthe device further comprises a force-receiving structure which is extended from the housing or the cuff and has the first and the second electrodes respectively positioned on opposite contactable surfaces thereof for achieving a simultaneous contact of the first and the second electrodes with different portions of the user's skin as a force is applied thereon. 2. The device as claimed in claim 1, wherein the first electrode and/or the second electrode is further connected to a sensor for detecting and/or notifying the contact thereon. | MD Biomedical, Inc. | A61B005/0408 A61B005/022 A61B005/044 A61B005/0245 A61B005/021 A61B005/024 A61B005/0404 A61B005/0452 A61B005/00
US09808636 | B2 | 2017-11-07 | Electrocardiogram identification | 2016-05-03 | A defibrillating system includes a processor coupled to a memory. The processor and the memory are configured to identify a treatment event associated with treatment of a victim with the defibrillating system, and transmit a representation of a portion of an ECG signal associated with the identified treatment event. In some cases, the processor and the memory are configured to identify the portion of the ECG signal associated with the identified treatment event. In some cases, the portion of the ECG signal is of a predetermined length of time having a start time and an end time based on a time associated with the identified treatment event. | 1. A defibrillating system comprising: a non-transitory computer readable medium storing a catalog of representations of one or more portions of physiological signals; anda processor coupled to a memory, the processor and the memory configured to: determine that treatment of a patient using the defibrillating system is completed; andupon determining that treatment of the patient using the defibrillating system is complete, adding, to the catalog, data representing a portion of a physiological signal obtained by the defibrillating system during treatment of the patient. 2. The defibrillating system of claim 1, wherein the representations of the one or more portions of the physiological signals are associated with respective identification data comprising one or more of a timestamp, a characteristic of the portion of the physiological signal, an identification of the patient, a treatment event, or narrative information. 3. The defibrillating system of claim 2, wherein the representations of the one or more portions of the physiological signals are organized in the catalog based on the identification data associated with each representation. 4. The defibrillating system of claim 3, wherein the representations of the one or more portions of the physiological signals in the catalog are organized as a chronological timeline comprising one or more treatment events. 5. The defibrillating system of claim 2, wherein the characteristic of the portion of the physiological signal comprises an indication of a type of cardiac rhythm represented by the portion of the physiological signal. 6. The defibrillating system of claim 5, wherein the type of cardiac rhythm comprises a shockable cardiac rhythm or a non-shockable cardiac rhythm. 7. The defibrillating system of claim 2, wherein the characteristic of the portion of the physiological signal comprises a type of the physiological signal. 8. The defibrillating system of claim 2, wherein the characteristic of the portion of the physiological signal comprises an indication of a treatment period of the patient for the portion of the physiological signal. 9. The defibrillating system of claim 1, wherein providing a representation of the portion of the physiological signal comprises displaying the representations of the one or more portions of the physiological signals in the catalog on a display interface of the defibrillating system. 10. The defibrillating system of claim 9, further comprising a control for interacting with the representations of the one or more portions of the physiological signals in the catalog. 11. The defibrillating system of claim 9, comprising displaying the identification data on the display interface of the defibrillating system. 12. The defibrillating system of claim 1, further comprising transmitting catalog data of the catalog to a remote computing device. 13. The defibrillating system of claim 1, wherein a type of the physiological signal comprises one or more of an ECG sign... | ZOLL Medical Corporation | A61B005/02 A61B005/04 A61N001/00 A61N001/39 A61B005/0432 A61B005/00 A61B005/0408 A61B005/046 A61B005/0464 A61B005/044 A61B005/0404
US09820729 | B2 | 2017-11-21 | Surgical access system and related methods | 2016-03-02 | A system for accessing a surgical target site and related methods, involving an initial distraction system for creating an initial distraction corridor, and an assembly capable of distracting from the initial distraction corridor to a secondary distraction corridor and thereafter sequentially receiving a plurality of retractor blades for retracting from the secondary distraction corridor to thereby create an operative corridor to the surgical target site, both of which may be equipped with one or more electrodes for use in detecting the existence of (and optionally the distance and/or direction to) neural structures before, during, and after the establishment of an operative corridor to a surgical target site. | 1. A retractor blade for use in a system for accessing a surgical target site, the retractor blade comprising: an elongated retractor blade extending from a proximal end that is dimensioned to couple to a mounting structure to a distal end that is spaced longitudinally from the proximal end, wherein the retractor blade includes a first edge extending between the proximal end and the distal end, a second edge extending between the proximal end and the distal end, an inner face defined by the retractor blade between the first and second edges, and an outer face defined by the retractor blade between the first and second edges, wherein the inner face of the retractor blade defines an engagement groove dimensioned to couple to an engagement portion of a connectable element, wherein the engagement groove is defined by the inner face of the retractor blade between a first ridge extending along a first portion of the engagement groove proximate the first edge of the retractor blade and a second ridge extending along a second portion of the engagement groove proximate the second edge of the retractor blade, and wherein the inner face of the retractor blade defines at least one recess dimensioned to receive a displaceable engagement member of the connectable element so as to hold the connectable element in place with respect to the retractor blade. 2. The retractor blade of claim 1, wherein the at least one recess defined by the inner face of the retractor blade comprises a plurality of recesses configured to longitudinally position the connectable element with respect to the retractor blade. 3. The retractor blade of claim 1, wherein the at least one recess is positioned on a portion of the inner face that is outside of the engagement groove. 4. The retractor blade of claim 1, wherein the engagement groove has a female dovetail configuration. 5. The retractor blade of claim 1, wherein the inner face is substantially concave and the outer face is substantially convex. 6. The retractor blade of claim 1, wherein the engagement groove has a proximal portion near the proximal end of the retractor blade and a distal portion near the distal end of the retractor blade and wherein the proximal portion is wider than the distal portion. 7. The retractor blade of claim 1, wherein the engagement groove extends along less than a full longitudinal length of the retractor blade. 8. The retractor blade of claim 1, wherein the engagement groove is positioned along a longitudinal midline of the retractor blade. 9. The retractor blade of claim 1, wherein the retractor blade tapers along at least part of the retractor blade between the first and second edges. 10. A system comprising: the retractor blade of claim 1; anda connectable element having the engagement portion and the displaceable engagement member; wherein the engagement portion is dimensioned to detachably couple with the engagement groove of the retractor blade and wherein the displaceable engagement mem... | NuVasive, Inc. | A61B001/32 A61B017/02 A61B017/34 A61B005/0488 A61B090/00 A61M029/02 A61B017/00
US09829302 | B2 | 2017-11-28 | System and method for measuring a non-bone tendon | 2015-11-13 | There is disclosed a system and method for measuring the diameter or thickness of a tendon. In an embodiment, the device includes an adjustable tube having a variably sized cross-sectional diameter in a direction perpendicular to an axis extending between the first end and the second end, the adjustable tube being positionable from an open configuration in which the variably sized cross-sectional diameter is larger than a diameter of a tendon so as to allow loading of the tendon, to a closed configuration in which the variably sized cross-sectional diameter approximates the diameter of the tendon so as to allow measurement of the tendon. The device includes a measurement gradient in operable configuration with the adjustable tube so as to allow measurement of the tendon in the closed configuration. Other embodiments are also disclosed. | 1. A tendon measurement device, comprising: an adjustable tube having a first end, a second end, and an inner region having a variably sized cross-sectional diameter in a direction perpendicular to an axis extending between said first end and said second end, said adjustable tube positionable between an open configuration in which said inner region of said variably sized cross-sectional diameter is larger than a diameter of a tendon so as to allow loading of the tendon within the inner region and a closed configuration in which said inner region of said variably sized cross-sectional diameter approximates the diameter of the tendon within said inner region so as to allow measurement of the tendon within said inner region; anda measurement gradient operably associated with said adjustable tube so as to allow measurement of the tendon within said inner region when said adjustable tube is positioned in said closed configuration. 2. The device of claim 1, further comprising an outer tube encompassing said adjustable tube, said outer tube affixed to a first edge of said adjustable tube, and said adjustable tube adjustably positionable with respect to said outer tube. 3. The device of claim 2, wherein said measurement gradient provides an indication of the diameter of the tendon. 4. The device of claim 2, further comprising a ratcheting mechanism operably configured to position said adjustable tube between said open configuration and said closed configuration. 5. The device of claim 4, wherein said ratcheting mechanism comprises a rotatable tube, said rotatable tube affixed to a second edge of said adjustable tube and positioned between said adjustable tube and said outer tube when said adjustable tube is in said open configuration. 6. The device of claim 1, wherein said adjustable tube has a length from said first end to said second end, and wherein said length of said adjustable tube exceeds a length of the tendon under measurement. 7. The device of claim 1, wherein said adjustable tube has a length from said first end to said second end, and wherein said length of said adjustable tube approximates a length of the tendon under measurement. 8. The device of claim 1, wherein said adjustable tube has a length from said first end to said second end, and wherein said length of said adjustable tube equals a length of the tendon under measurement. 9. The device of claim 1, wherein said adjustable tube has a length from said first end to said second end, and wherein said length of said adjustable tube is at least half a length of the tendon under measurement. 10. A system for measuring a diameter of a length of tendon, comprising: a cinchable tube defining an inner region, said cinchable tube configured to move between an open position and a closed position, wherein when in said open position, said cinchable tube is configured to receive at least half the length of the tendon within said inner region and when in said closed position, said cinchable ... | Allosource | A61B005/107 G01B005/08 A61B019/00 A61B005/00
US09817098 | B2 | 2017-11-14 | Scaling of local coils of a magnetic resonance imaging scanner | 2014-09-26 | A scaling unit (1) for reception antennae (A1, A2, A3, A4) of a plurality of local coils (LS) of a magnetic resonance imaging scanner (MRT) includes a plurality of signal inputs (in1, in2, in3, in4) configured for the reception of signals from the reception antennae (A1, A2, A3, A4), and a plurality of signal outputs (out1, out2, out3, out4) configured for the output of unaltered and/or altered signals from the reception antennae (A1, A2, A3, A4). | 1. A scaling unit for reception antennae of local coils of a magnetic resonance imaging scanner, the scaling unit comprising: a plurality of signal inputs configured for receiving signals from the reception antennae; anda plurality of signal outputs configured for outputting unaltered signals, altered signals, or unaltered and altered signals from the reception antennae;wherein, at least in a summed signal mode: the plurality of signal inputs is connected, for signaling purposes, to a single common signal output of the plurality of signal outputs, such that a sum of signals from the plurality of signal inputs is applied to the single common signal output; andeach additional signal output of the plurality of signal outputs is connected, for signaling purposes, to a respective single signal input of the plurality of signal inputs, such that an unaltered signal from the respective single signal input is applied to the additional signal output. 2. The scaling unit of claim 1, further comprising: a switching element configured for changing over, for signaling purposes, the single common signal output to which the sum of signals from the plurality of signal inputs is applied in the summed signal mode to a signal input associated with the single common signal output in a single signal mode;wherein each signal input of the plurality of signal inputs is connected, for signaling purposes, to a single associated signal output, such that a respective unaltered, signal from the associated signal input of the plurality of signal inputs is applied to the single associated signal output. 3. The scaling unit of claim 1, wherein: each signal input of the plurality of signal inputs is connected, for signaling purposes, to a single associated signal output of the plurality of signal outputs, such that a respective unaltered signal from an associated signal input is applied to the single associated signal output, wherein an image processing unit of the magnetic resonance imaging scanner is configured to retrieve signals in a single signal mode; andthe plurality of signal inputs, the plurality of signal outputs, or the plurality of signal inputs and the plurality of signal outputs are connected, for signaling purposes, to an additional common signal output, such that a sum of unaltered signals from the plurality of signal inputs is applied to the additional common signal output, wherein the image processing unit of the magnetic resonance imaging scanner is configured to retrieve signals in a summed signal mode. 4. The scaling unit of claim 2, wherein: each signal input of the plurality of signal inputs is connected, for signaling purposes, to a single associated signal output of the plurality of signal outputs, such that a respective unaltered signal from an associated signal input is applied to the single associated signal output, wherein an image processing unit of the magnetic resonance imaging scanner is configured to retrieve signals in a single signal mode; a... | Siemens Aktiengesellschaft | G01V003/00 G01R033/56 A61B005/055 G01R033/36
US09814424 | B2 | 2017-11-14 | Chronic obstructive pulmonary disease drug titration and management | 2014-11-13 | A system may include a port, at least one sensing circuit, and at least one processor. The port is configured to receive an indication of dosing of medication to treat a pulmonary condition of a heart failure (HF) subject and the at least one sensing circuit configured to sense at least one physiological signal, wherein the physiological signal includes physiological information of the HF subject. The at least one processor includes a parameter module configured to extract values of at least one physiological parameter indicative of health status of the HF subject, and a trending module configured to trend extracted values of the at least one physiological parameter and detect an effect of the dosing of the medication on the HF subject using the trending of the extracted values of the at least one physiological parameter. | 1. A system comprising: a port configured to receive an indication of dosing of medication to treat a pulmonary condition of a heart failure (HF) subject;at least one sensing circuit configured to sense at least one physiological signal, wherein the physiological signal includes physiological information of the HF subject;at least one processor including: a parameter module configured to extract values of at least one physiological parameter indicative of health status of the HF subject; anda trending module configured to trend extracted values of the at least one physiological parameter, to detect a side effect of the dosing of the medication on the HF subject when detecting that a value of the physiological parameter exceeds a specified threshold value of the at least one physiological parameter, and to output an indication of the detected side effect of dosing; andan alert unit configured to notify the detected side effect of dosing to a user. 2. The system of claim 1, wherein the sensing circuit includes a respiration sensing circuit configured to sense a respiration signal that includes respiration information of the HF subject, wherein the parameter module is configured to extract a value of a respiration parameter of the HF subject using the respiration signal, andwherein the trending module is configured to trend values of the respiration parameter and generate an indication of effectiveness of the dosing of the medication on the HF subject using the trend of values of the respiration parameter. 3. The system of claim 1, wherein the at least one physiological signal includes at least one of a cardiac activity signal, a thoracic impedance signal, a heart sound signal, and a blood pressure signal, and wherein the parameter module is configured to extract values of at least one of heart rate, heart contractility, arrhythmia burden, thoracic impedance, at least one heart sound parameter, and blood pressure from the at least one physiological signal, andwherein the trending module is configured to trend values of at least one of heart rate, heart contractility, arrhythmia burden, thoracic impedance, at least one heart sound parameter, and blood pressure, and detect a side effect of the dosing of the medication on the HF subject using the trending. 4. The system of claim 1, wherein the at least one physiological signal includes thoracic impedance information of the HF subject, wherein the parameter module configured to extract values of thoracic impedance from the physiological signal, and wherein the trending module is configured to trend values of thoracic impedance and generate an indication of effectiveness of the dosing of the medication on the HF subject using the trend of the measure of thoracic impedance. 5. The system of claim 1, wherein the trending module is configured to detect a side effect of the dosing of the medication on the HF subject when detecting that a value of the integral of values of the physiological parameter exce... | Cardiac Pacemakers, Inc. | A61B005/0452 A61B005/053 A61B005/08 A61B005/091 A61B007/04 A61M005/172 A61N001/365 A61B005/021 A61B005/00 A61B005/1455 G06F019/00 A61B005/145 A61B005/0245
US09814489 | B2 | 2017-11-14 | Probe insertion device for implanting a probe into tissue | 2014-10-03 | A probe insertion device for implanting a probe into tissue includes a rigid base that selectively attaches to the probe due to a bond between the base and the probe, that provides a structural backbone to the probe, is longitudinally aligned with the probe, and can be adapted to receive a fluid between the base and the probe. The probe insertion device can include a surface covering at least a portion of the base that reduces the bond between the base and the probe in the presence of the fluid. | 1. A method of implanting an insertable end of a flexible neural probe into tissue, comprising the steps: attaching a rigid base being a flat rigid blade to the flexible neural probe, thereby providing a structural backbone to the flexible neural probe;inserting at least a portion of the insertable end of the flexible neural probe into the tissue;decoupling at least a portion of the flexible neural probe from the rigid base by flowing a fluid between the base and the flexible neural probe to induce detachment of the base from the flexible neural probe; andwithdrawing the base from the tissue. 2. The method of claim 1, wherein the step of attaching a base being a flat rigid blade to the flexible neural probe includes attaching the flat rigid base to the flexible neural probe with at least one chemical bond selected from the group consisting of electrostatic forces, covalent bonds, ionic bonds, metallic bonds, van der Waals bonds, and hydrogen bonds. 3. The method of claim 2, wherein the step of flowing a fluid between the base and the flexible neural probe includes introducing a fluid that reduces the chemical bond between the base and the flexible neural probe. 4. The method of claim 1, wherein the step of decoupling at least a portion of the flexible neural probe from the base includes peeling a portion of the flexible neural probe from the base. 5. A method of implanting a flexible neural probe into tissue, comprising the steps: providing a flexible neural probe including an insertable end that is implantable into the tissue;attaching a flat rigid base to at least a portion of the probe, thereby providing a structural backbone to the probe;inserting at least a portion of the probe into the tissue;decoupling at least a portion of the probe from the base directly in response to a fluid flowing between the base and the probe to induce detachment of the base from the probe; andwithdrawing the base from the tissue. 6. The method of claim 5 wherein the step of attaching a rigid base to at least a portion of the probe comprises attaching a rigid base to at least a portion of the probe via hydrophobic interaction. 7. The method of claim 6, wherein the step of decoupling at least a portion of the probe from the base directly in response to a fluid flowing between the base and the probe to induce detachment of the base from the probe comprises a fluid that reduces the hydrophobic interaction between the base and the probe. 8. The method of claim 5, wherein the step of attaching a rigid base to at least a portion of the probe comprises attaching the rigid base to at least a portion of the probe with at least one chemical bond selected from the group consisting of electrostatic forces, covalent bonds, ionic bonds, metallic bonds, van der Waals bonds, and hydrogen bonds. 9. The method of claim 8, wherein the step of allowing a fluid to flow between the base and the probe includes introducing a fluid that reduces the chemical bond between the base an... | THE REGENTS OF THE UNIVERSITY OF MICHIGAN | A61B017/34 A61B005/00 A61N001/05 A61B005/0478 A61B017/16 A61B017/00 A61B090/10
US09814400 | B1 | 2017-11-14 | Method for improving accuracy of pulse rate estimation | 2015-05-26 | Systems and methods are provided for determining the frequency of a cardiovascular pulse based on a photoplethysmographic measurement of blood in a portion of subsurface vasculature. A plurality of samples of the photoplethysmographic signal are obtained and a frequency spectrum determined based on a first set of the plurality of samples. A frequency is determined based on a maximum frequency component of the frequency spectrum. Further, a confidence level of the determined frequency of the maximum frequency component is determined based on the magnitude of the maximum frequency component and the overall energy in the frequency spectrum. A pulse rate is then determined by updating a predicted pulse rate of a second, earlier set of the plurality of samples based on the frequency of the maximum frequency component and the determined confidence level. | 1. A method comprising: receiving, from a sensor, a plurality of samples of a signal, wherein the signal is related to a cardiovascular pulse;determining a predicted pulse rate based at least on a first set of the plurality of samples;determining a frequency spectrum of a second set of the plurality of samples, wherein the frequency spectrum comprises a plurality of frequency components, each having a respective magnitude;determining a frequency of a local maximum of the frequency spectrum;determining an overall energy in the frequency spectrum;determining a confidence level of the determined frequency of the local maximum, wherein the determined confidence level is based on at least a magnitude of the local maximum of the frequency spectrum and the determined overall energy in the frequency spectrum; anddetermining a pulse rate based on the predicted pulse rate, the determined frequency of the local maximum, and the determined confidence level. 2. The method of claim 1, wherein determining a pulse rate based on the predicted pulse rate, the determined frequency of the local maximum, and the determined confidence level comprises determining a weighted combination of the predicted pulse rate and the determined frequency based on a weighting factor, wherein the weighting factor is based on the determined confidence level such that a higher determined confidence level causes the pulse rate to be more weighted toward the determined frequency of the local maximum. 3. The method of claim 1, wherein determining a predicted pulse rate based at least on a first set of the plurality of samples comprises determining a frequency spectrum of the first set of the plurality of samples, wherein the first set of the plurality of samples and the second set of the plurality of samples overlap, and wherein determining a frequency spectrum of the second set of the plurality of samples comprises performing an update on the determined frequency spectrum of the first set of the plurality of samples based on samples of the second set of the plurality of samples that are not part of the first set of the plurality of samples. 4. The method of claim 3, further comprising: receiving an artifact signal;determining a frequency spectrum of the artifact signal; andfiltering the determined frequency spectrum of the second set of the plurality of samples based on the determined frequency spectrum of the artifact signal to generate a plurality of filtered values of the magnitude of the second set of the plurality of samples at the plurality of corresponding different frequencies, wherein determining a frequency of a local maximum of the frequency spectrum comprises determining a frequency of a local maximum of the plurality of filtered values of the magnitude of the second set of the plurality of samples at the plurality of corresponding different frequencies. 5. The method of claim 4, wherein filtering the determined frequency spectrum of the second set of the plurality of sam... | Verily Life Sciences LLC | A61B005/00 A61B005/024 G06F019/00 A61B005/022 A61B005/0245 A61B005/0295
US09775567 | B2 | 2017-10-03 | Intravascular sensing method and system | 2015-12-22 | Described herein are methods and systems for calculating a corrected Fractional Flow Reserve (FFR). An illustrative method includes delivering a first pressure sensing device including a pressure sensor to a location in a blood vessel having a stenosis, positioning the pressure sensor distal to the stenosis, measuring the distal pressure, measuring the proximal pressure, and calculating a corrected FFR using the measured proximal and distal pressures and applying a correction factor or correction equation. The corrected FFR may approximate a FFR that would have been obtained if the pressure measured downstream of the stenosis was measured using a second pressure sensing device having a second maximum cross-sectional area in a distal portion of the second pressure sensing device that is different from the first maximum cross-sectional area. | 1. A method of calculating a corrected Fractional Flow Reserve (FFR) using a first pressure sensing device at a location in a blood vessel having a stenosis, the first pressure sensing device having a pressure sensor, the method comprising: measuring a pressure downstream of the stenosis with the first pressure sensing device while the pressure sensor is positioned downstream of the stenosis, the first pressure sensing device having a first maximum cross-sectional area in a distal portion of the first pressure sensing device;measuring a pressure upstream of the stenosis;calculating with a processing device a corrected FFR using the measured upstream and downstream pressures and applying a correction factor or correction equation;wherein the correction factor or correction equation corrects for errors in the pressure measured downstream of the stenosis caused by the presence of the first pressure sensing device; andwherein the corrected FFR approximates a FFR that would have been obtained if the pressure measured downstream of the stenosis was measured using a second pressure sensing device having a second maximum cross-sectional area in a distal portion of the second pressure sensing device that is different from the first maximum cross-sectional area, the correction factor or correction equation accounting for the second maximum cross-sectional area. 2. The method of claim 1, further comprising selecting the correction factor or correction equation from a group of at least two correction factors or correction equations, wherein the selection of the correction factor or correction equation corresponds to the cross-sectional area of the first pressure sensing device, or by a type of the first pressure sensing device. 3. The method of claim 2, wherein the selection of the correction factor or correction equation is further determined by one or more of a size of the stenosis, a size of a lumen of the blood vessel, and a rate of blood flow. 4. The method of claim 3, wherein the first pressure sensing device comprises one or more sensors configured to measure the size of the stenosis, the size of the lumen, and/or the rate of blood flow. 5. The method of claim 1, wherein the second pressure sensing device is a pressure sensing guidewire and/or the distal portion of the second pressure sensing device has a maximum diameter of about 0.014 inches. 6. The method of claim 1, wherein calculating the corrected FFR comprises applying the correction factor to the measured downstream pressure to get a corrected downstream pressure, and calculating the corrected FFR using the corrected downstream pressure. 7. The method of claim 1, wherein calculating the corrected FFR comprises calculating a measured FFR using the measured downstream pressure, and applying the correction equation to the measured FFR to calculate the corrected FFR. 8. The method of claim 1, further comprising displaying the corrected FFR on a visual display. 9. A system for calculating a... | ACIST Medical Systems, Inc. | A61B005/02 A61B005/00 A61B005/0215 A61B005/026 A61M025/00
US09814548 | B2 | 2017-11-14 | Three-dimensional oral cavity scan device using piezoelectric element based pattern module and variable focus lens | 2016-04-05 | The present invention relates to a three-dimensional oral cavity scan device for a digital impression that obtains a three-dimensional image by processing a two-dimensional image acquired by photographing a three-dimensional subject in an oral cavity. The device includes: a pattern mask converting parallel light incident thereto into structured light; a piezoelectric plate where a mechanical vibration is generated on a surface thereof by an AC voltage applied thereto; and a power supply unit applying the AC voltage to the piezoelectric plate, wherein the piezoelectric plate moves the pattern mask in a longitudinal direction of the piezoelectric plate by using the mechanical vibration. The device can reduce noise and vibration by moving the pattern mask with an electrical method. In addition, the device can accurately measure depth of teeth by electrically controlling a focal length of structured light projected on the teeth. | 1. An oral cavity scan device obtaining a three-dimensional shape of teeth by using structured light, the device comprising: a pattern mask converting parallel light incident thereto into the structured light;a piezoelectric plate where a mechanical vibration is generated on a surface thereof by an AC voltage applied thereto; and a power supply unit applying the AC voltage to the piezoelectric plate,wherein the piezoelectric plate moves the pattern mask in a longitudinal direction of the piezoelectric plate by using the mechanical vibration generated on the surface of the piezoelectric plate. 2. The device of claim 1, wherein the piezoelectric plate is a PZT (plumbum zirconate titanate) ceramic element. 3. The device of claim 1, wherein the pattern mask comprises: a light transmission line through which the parallel light passes; anda light shielding line blocking the parallel light,wherein the light transmission line and the light shielding line are alternately arranged in the pattern mask. 4. The device of claim 3, wherein the piezoelectric plate is a PZT (plumbum zirconate titanate) ceramic element. 5. The device of claim 1, further comprising: an offset lens to which the structured light is incident, with a concave curvature formed in the offset lens; anda reflection mirror reflecting the structured light in a direction toward a subject,wherein the offset lens reduces a focal distance of the structured light. 6. The device of claim 5, wherein the piezoelectric plate is a PZT (plumbum zirconate titanate) ceramic element. 7. The device of claim 5, further comprising: a variable focus lens adjusting the reduced focal distance of the structured light penetrating the offset lens. 8. The device of claim 7, wherein the variable focus lens enables reflected light reflected from the subject to be incident to the offset lens so as to obtain an image of the teeth. 9. The device of claim 7, wherein the piezoelectric plate is a PZT (plumbum zirconate titanate) ceramic element. 10. The device of claim 8, wherein the piezoelectric plate is a PZT (plumbum zirconate titanate) ceramic element. | KOREA PHOTONICS TECHNOLOGY INSTITUTE | A61C009/00 G01B011/25 A61B005/00 A61B001/00 A61B001/247
US09801553 | B2 | 2017-10-31 | System, method, and computer program product for the real-time mobile evaluation of physiological stress | 2014-09-26 | A system is disclosed including a measuring device comprising sensors configured to collect data from a user based on the sensor being at least one of in contact with and proximate to a section of skin of the user, a stress feature analyzer configured to determine a cardiovascular activity feature of the user, a skin conductance feature of the user, movement of the user, and skin temperature of the user based on the data collected with the measuring device, and a stress classifier configured to classify and report stress experienced by the user based on at least one of one feature selected from the cardiovascular activity feature and the skin conductance feature, and two features including a combination of the cardiovascular activity feature and the skin conductance feature, wherein when the skin temperature of the user exceeds a defined threshold, the stress classifier ceases to classify the stress experienced by the user. A method and a non-transitory processor readable storage medium are also disclosed. | 1. A system comprising: a measuring device comprising sensors configured to collect data from a user based on the sensors being at least one of in contact with and proximate to a section of skin of the user;a stress feature analyzer configured to determine a cardiovascular activity feature of the user, a skin conductance feature of the user, movement of the user, and skin temperature of the user based on the data collected with the measuring device; anda stress classifier configured to classify and report stress experienced by the user based on at least one of one feature selected from the cardiovascular activity feature and the skin conductance feature, and two features including a combination of the cardiovascular activity feature and the skin conductance feature, the stress classifier includes: a first stress classifier to classify the stress experienced by the user based on the one feature being the cardiovascular activity feature for a minute time interval when skin conductance data is missing during the minute time interval:a second stress classifier to classify the stress experienced by the user based on the one feature being the skin conductance feature for the minute time interval when cardiovascular activity data is missing; anda third stress classifier to classify the stress experienced by the user based on the two features including the combination of the cardiovascular activity feature and the skin conductance feature during the minute time interval when both the cardiovascular activity data and the skin conductance data are available;wherein when at least one of the skin temperature and movement of the user exceeds a defined threshold, the first, the second and third stress classifiers cease classify the stress experienced by the user. 2. The system according to claim 1, further comprising a database configured to maintain the report. 3. The system according to claim 1, further comprising: an input device configured to allow the user, by self-reporting, to designate when psychological stress is experienced; andan adaptation component including an adaptation algorithm wherein the stress classifier makes a determination of the stress experienced by the user relative to baseline references and to reset a bias threshold to allow lesser or greater sensitivity to heart rate of the cardiovascular activity feature or the skin conductance feature of the user. 4. The system according to claim 1, wherein the measuring device is a part of a wearable device and at least one of the stress feature analyzer and the stress classifier are part of a mobile device which is in communication with the wearable device. 5. The system according to claim 1, wherein the stress feature analyzer and the stress classifier determine the stress experienced by the user in real-time on a minute-to-minute basis. 6. The system according to claim 1, further comprising a display configured to show a stress score generated by the stress classifier wherein the stress ... | Design Interactive, Inc. | A61B005/0205 A61B005/11 A61B005/00 A61B005/16 A61B005/053 A61B005/01
US09795335 | B2 | 2017-10-24 | Data collection for vestibulogram construction | 2013-03-15 | The methods, systems, and computer program products described herein enable the amount of information that can be acquired during a defined period of vestibular testing to be significantly increased while also significantly reducing the time required to collect data during such vestibular testing. These objects are achieved by intelligent data collection methods and systems configured to adaptively control a motion platform in a way that increases the amount of information gleaned from each data collection event and to achieve synergies based on the different types of, and the order of, data analysis and calibration methods used. | 1. A method for assessing a subject's vestibular system, the method comprising: administering a psychometric test to a subject and measuring responses of the subject to obtain a set of values, wherein administering the psychometric test comprises:positioning the subject on a motion platform that includes a device for supporting the subject and is configured to execute motion profiles,obtaining a first estimate of a parameter vector of the subject's psychometric function based on the subject's perception of motion in response to a first motion profile set wherein the subject's perception of motion is received using one or more input devices, andadaptively selecting, using one or more processing devices, a subsequent motion profile based at least in part on the first estimate, determining at least one fit parameter that represents a distribution of the set of values and represents a characteristic of the subject's responses to the psychometric test;estimating a bias quantity associated with each said fit parameter; subtracting the bias quantity from each said fit parameter to obtain a bias-reduced estimate of the fit parameter;scaling the bias-reduced estimate of the fit parameter by a scale-factor to obtain a final estimate of the fit parameter; andgenerating a quantitative measure representing the subject's vestibular system based on the final estimate of the fit parameter. 2. The method of claim 1, wherein the scale factor has a value substantially equal to or less than one. 3. The method of claim 1, wherein the responses of the subject are measured for stimuli corresponding to a plurality of values of a test parameter associated with the motion profiles and the responses to different values of the test parameter are scaled using a corresponding scaling function to obtain the set of values. 4. The method of claim 1, wherein the responses of the subject are obtained by providing a stimulus to the subject for a predetermined number of times to measure the responses of the subject, wherein a level of the stimulus is associated with a corresponding motion profile. 5. The method of claim 1, wherein the distribution is represented as a cumulative distribution function. 6. The method of claim 1, wherein the at least one fit parameter is determined using a numerical maximum-likelihood method. 7. The method of claim 1, wherein the bias quantity represents a first order asymptotic bias. 8. The method of claim 1, further comprising selecting the subsequent motion profile to reduce uncertainty in the estimate. 9. The method of claim 1, further comprising selecting the subsequent motion profile based at least in part on Fisher information associated with the estimate. 10. The method of claim 1, further comprising correcting for lapses in the subject's attention by identifying outlier data and eliminating the outlier data prior to obtaining the first estimate. 11. The method of claim 1, wherein the bias reduced estimate of the fit parameter is obtai... | Massachusetts Eye and Ear Infirmary | A61B005/16 A61B005/18 A61B005/00 G09B005/00
US09820694 | B2 | 2017-11-21 | Devices for diagnosing sleep apnea or other conditions and related systems and methods | 2015-08-10 | A system includes a sensor device having a body configured to be inserted into an airway of a patient and one or more sensors mounted in or on the body. The one or more sensors are configured to collect sensor data associated with the airway of the patient. The system also includes a signal analyzer configured to analyze the sensor data. The one or more sensors could include one or more microphones. The signal analyzer could identify volume and/or pitch characteristics of the sensor data, perform pattern recognition to identify one or more patterns using the volume and/or pitch characteristics, and use the one or more patterns to identify a type, a location, and/or a degree of airway obstruction. This could be done, for instance, to determine if the patient suffers from obstructive sleep apnea or other condition that affects his or her airway. | 1. A system comprising: a sensor device configured to be inserted through a mouth or a nose of a patient into an upper airway of the patient, the sensor device comprising multiple sensors configured to capture and transmit sensor data at multiple spaced-apart locations along the upper airway of the patient, the sensor data comprising vibrational or acoustic data associated with the upper airway of the patient;at least one interface configured to receive the sensor data; andat least one processing device configured to analyze the sensor data to identify one or more obstructions within the upper airway of the patient, wherein the at least one processing device is configured to identify the one or more obstructions within the upper airway of the patient by determining whether one or more patterns of airway noise are present in the upper airway of the patient based on the vibrational or acoustic data captured at the multiple spaced-apart locations along the upper airway of the patient. 2. The system of claim 1, wherein the at least one interface is physically coupled to the sensor device. 3. The system of claim 1, wherein the at least one processing device is configured to identify at least one of: volume characteristics and pitch characteristics of the sensor data. 4. The system of claim 3, wherein the at least one processing device is configured to perform pattern recognition in order to identify the one or more patterns using at least one of the volume characteristics and the pitch characteristics of the sensor data. 5. The system of claim 4, wherein the at least one processing device is configured to use the one or more patterns to identify at least one of: a type, a location, and a degree of the one or more obstructions within the upper airway of the patient. 6. The system of claim 1, wherein the at least one processing device is configured to identify one or more locations of one or more sounds within the upper airway of the patient using the sensor data. 7. The system of claim 1, wherein the at least one processing device is configured to identify one or more areas of greatest obstruction within the upper airway of the patient using the sensor data. 8. The system of claim 1, wherein the at least one processing device is configured to analyze the sensor data to determine if the patient suffers from obstructive sleep apnea. 9. The apparatus of claim 1, wherein the sensor data further comprises airway pressure within the upper airway of the patient. 10. A method comprising: capturing sensor data comprising vibrational or acoustic data using a sensor device, the sensor device configured to be inserted through a mouth or a nose of a patient into an upper airway of the patient, the sensor device comprising multiple sensors configured to capture the vibrational or acoustic data associated with the upper airway of the patient at multiple spaced-apart locations along the upper airway of the patient; andanalyzing the sensor data to identify one... | - | A61B005/087 A61B005/08 A61M016/04 A61B008/12 A61B005/00 A61B007/00 A61M016/00 A61B005/091
US09782130 | B2 | 2017-10-10 | Robotic surgical system | 2006-12-29 | A robotic surgical system includes a track, a catheter holding device including a catheter receiving portion translatably associated with the track, a translation servo mechanism to control translation of the catheter holding device relative to the track, a catheter deflection control mechanism, a deflection servo mechanism to control the catheter deflection control mechanism, and a controller to control at least one of the servo mechanisms. The catheter receiving portion is adapted for quick installation and removal of a catheter. The catheter receiving portion may be rotatable, with a rotation servo mechanism to control the rotatable catheter receiving portion. The controller controls at least one of the deflection and rotation servo mechanisms to maintain a substantially constant catheter deflection as the catheter rotates. An introducer, which may be steerable, and an expandable, collapsible sterile tube may also be provided. | 1. A robotic surgical system, comprising: a track;a catheter holding device translatably associated with said track, said catheter holding device including a catheter receiving portion, wherein said catheter receiving portion comprises a rotatable catheter receiving portion;a translation servo mechanism operatively coupled to said catheter holding device and adapted to control translation of said catheter holding device relative to said track;a catheter deflection control mechanism;a deflection servo mechanism operatively coupled to said catheter deflection control mechanism and adapted to control said catheter deflection control mechanism;a rotation servo mechanism operatively coupled to said rotatable catheter receiving portion and adapted to control said rotatable catheter receiving portion; anda controller operatively coupled to at least one of said translation servo mechanism and said deflection servo mechanism, wherein said controller is adapted to control said at least one of said translation servo mechanism and said deflection servo mechanism,wherein the catheter receiving portion is configured to receive a catheter comprising a control handle configured for manual operation and including a manually-actuable catheter deflection actuator, andwherein said controller is operatively coupled to said rotation servo mechanism and said deflection servo mechanism and is adapted to control at least one of said deflection servo mechanism and said rotation servo mechanism to maintain a substantially constant catheter deflection as said catheter receiving portion rotates. 2. The system according to claim 1, wherein said controller controls both of said deflection servo mechanism and said rotation servo mechanism to maintain a substantially constant catheter deflection as said catheter receiving portion rotates. | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B018/18 A61B005/00 A61B034/20 A61B017/00 A61M025/01 A61B034/00 A61B090/00 A61B034/30
US09801583 | B2 | 2017-10-31 | Magnetometer based physiological monitoring garment | 2013-05-07 | The present invention is directed to systems and methods for monitoring characteristics of a subject. A system according to an exemplary embodiment of the invention includes a sensor subsystem including at least one respiratory sensor disposed proximate to the subject and configured to detect a respiratory characteristic of the subject, wherein the sensor subsystem is configured to generate and transmit at least one respiratory signal representing the respiratory characteristic, and at least one physiological sensor disposed proximate to the subject and configured to detect a physiological characteristic of the subject, wherein the sensor subsystem is configured to generate and transmit at least one physiological signal representing the physiological characteristic, and a processor subsystem in communication with the sensor subsystem, the processor subsystem being configured to receive at least one of the at least one respiratory signal and the at least one physiological signal. | 1. A monitoring garment, comprising: a panel formed of an elastic material;a central communications pathway incorporated into the panel;a garment circuit connected to the central communications pathway; anda magnetometer circuit incorporated into the panel,wherein the magnetometer circuit includes a first pair of integral garment conductors electrically connecting a first magnetometer to the central communications pathway, the first magnetometer including a multi-layered printed circuit including a continuous conductive coil;a second pair of integral garment conductors electrically connecting a second magnetometer to the central communications pathway, the second magnetometer including a multi-layered printed circuit including a continuous conductive coil, wherein the magnetometer circuit is configured to sense relative displacement between the first magnetometer and the second magnetometer, andwherein the garment circuit is configured to enable signal transmission to and from the first magnetometer. 2. The monitoring garment of claim 1, wherein the integral garment conductors extend horizontally parallel to each other, and wherein the central communications pathway extends vertically to a data processing subsystem configured to determine information associated with a wearer as a function of signals received from the first magnetometer. 3. The monitoring garment of claim 2, wherein the determined information includes a respiratory characteristic. 4. The monitoring garment of claim 1, the integral garment conductors being woven into the panel. 5. The monitoring garment of claim 1, the integral garment conductors being knitted into the panel. 6. The monitoring garment of claim 1, wherein the central communications pathway comprises a ribbon comprising a plurality of wires. 7. The monitoring garment of claim 1, further comprising a conductive patch configured to sense electrical activity of a wearer's heart; and an integral garment conductor connecting the conductive patch to the central communications pathway. 8. The monitoring garment of claim 1, wherein the magnetometer circuit comprises: a third pair of integral garment conductors connecting a third magnetometer to the central communications pathway; anda fourth pair of integral garment conductors connecting a fourth magnetometer to the central communications pathway. 9. The monitoring garment of claim 8, wherein the first magnetometer and the third magnetometer are disposed at a front side of the garment and vertically offset from each other, wherein the second magnetometer and the fourth magnetometer are disposed at a back side of the garment and vertically offset from each other,wherein the magnetometer circuit is configured to sense relative displacement between the second magnetometer and the third magnetometer,wherein the first magnetometer and the second magnetometer are vertically offset from each other, andwherein the third magnetometer and the fourth magnetometer are vertically... | adidas AG | A61B005/00 G06F019/00 A61B005/08 A61B005/113 A61B005/01 A61B005/0402 A61B005/0476 A61B005/0488 A61B005/145 A61B005/0205
US09814398 | B2 | 2017-11-14 | Method and apparatus for non-invasively measuring hemodynamic parameters using parametrics | 2013-12-02 | An improved method and apparatus for non-invasively assessing one or more hemodynamic parameters associated with the circulatory system of a living organism. In one aspect, the invention comprises a method of measuring a hemodynamic parameter (e.g., arterial blood pressure) by applanating or compressing portions of tissue proximate to the blood vessel of concern until a desired condition is achieved, and then measuring the hemodynamic parameter. Such applanation effectively mitigates transfer and other losses created by the tissue proximate to the blood vessel, thereby facilitating accurate and robust tonometric measurement. An algorithm adapted to maintain optimal levels of applanation is also described. Methods and apparatus for scaling such hemodynamic parameter measurements based on subject physiology, and providing treatment to the subject based on the measured parameters, are also disclosed. | 1. A method for determining an optimum location for applanating a blood vessel, comprising: placing a transducer proximate to one side of a blood vessel at a starting position laterally offset from said substantially optimum location for applanating said blood vessel;automatically moving said transducer laterally over said blood vessel from said starting position to an ending position at a first rate, said ending position being past said substantially optimum location for applanating said blood vessel on a second side of said blood vessel, said automatic moving occurring while simultaneously measuring pulse pressure of said blood vessel via said transducer in order to collect a first pulse pressure data set;automatically backtracking said transducer laterally over said blood vessel from said ending position toward said starting position at a second rate, said automatic backtracking occurring while simultaneously measuring pulse pressure of said blood vessel via said transducer in order to collect a second pulse pressure data set;based at least in part on a correlation between said first and second pulse pressure data sets, determining said optimum location; andbased at least in part on said determination of said optimum location, automatically positioning said transducer over said blood vessel to place said transducer over said substantially optimum location. 2. The method of claim 1, wherein said act of determining said substantially optimum location of said blood vessel based at least in part on said first and second pulse pressure data sets comprises determining a location at which said measured pulse pressure of said blood vessel reaches a substantial maximum during said acts of automatically moving and backtracking said transducer laterally over said blood vessel. 3. The method of claim 1, wherein said acts of automatically moving and backtracking said transducer laterally over said blood vessel each comprise moving said transducer continuously. 4. The method of claim 1, wherein said first rate is a first substantially constant rate, and said second rate is a second substantially constant rate. 5. The method of claim 1, wherein each of said act of simultaneously measuring said pulse pressure of said blood vessel while moving said transducer and said act of simultaneously measuring said pulse pressure of said blood vessel while backtracking said transducer comprises substantially continuous measurement of said pulse pressure of said blood vessel. 6. The method of claim 1, further comprising attempting to maintain a constant level of compression by said transducer to said blood vessel during said acts of automatically moving and backtracking said transducer laterally over said blood vessel. 7. The method of claim 6, wherein said level of compression comprises a level which is sufficient to at least partially compress said blood vessel. 8. The method of claim 1, wherein said first sweep rate is greater than said second sweep rate. 9. Th... | Tensys Medical, Inc. | A61B005/021 A61B005/00
US09814388 | B2 | 2017-11-14 | Wireless patient monitoring system and method | 2016-02-11 | A patient monitoring system includes at least two wireless sensing devices, each configured to measure a different physiological parameter from a patient and wirelessly transmit a parameter dataset. The system further includes a receiver that receives each parameter dataset, a processor, and a monitoring regulation module executable on the processor to assign one of the at least two wireless sensing devices as a dominant wireless sensing device and at least one of the remaining wireless sensing devices as a subordinate wireless sensing device. The physiological parameter measured by the dominant wireless sensing device is a key parameter and the parameter dataset transmitted by the dominant wireless sensing device is a key parameter dataset. The key parameter dataset from the dominant wireless sensing device is processed to determine a stability indicator. The subordinate wireless sensing device is then operated based on the stability indicator for the key parameter. | 1. A patient monitoring system comprising: at least two wireless sensing devices, each wireless sensing device configured to measure a different physiological parameter from a patient and wirelessly transmit a parameter dataset;a receiver that receives each parameter dataset from each of the at least two wireless sensing devices;a processor;a monitoring regulation module executable on the processor to: assign one of the at least two wireless sensing devices as a dominant wireless sensing device and at least one of the remaining wireless sensing devices as a subordinate wireless sensing device, wherein the physiological parameter measured by the dominant wireless sensing device is a key parameter and the parameter dataset transmitted by the dominant wireless sensing device is a key parameter dataset;determine a stability indicator for the key parameter based on the key parameter dataset from the dominant wireless sensing device; andselectively operate the subordinate wireless sensing device based on the stability indicator for the key parameter, wherein at least one of the subordinate wireless sensing devices is turned off when the stability indicator is within a predetermined range, such that the subordinate wireless sensing device no longer measures any physiological parameter from the patient or transmits any parameter dataset. 2. The patient monitoring system of claim 1, wherein the dominant wireless sensing device is operated to continuously measure the key parameter. 3. The patient monitoring system of claim 1, wherein the dominant wireless sensing device is assigned based on one of a diagnosis or a treatment history for the patient. 4. The patient monitoring system of claim 1, wherein the key parameter and dominant wireless sensing device are rotated amongst at least two of the two or more wireless sensing devices. 5. A method of monitoring a patient, the method comprising: providing two or more wireless sensing devices, each wireless sensing device configured to measure a different physiological parameter from a patient and communicatively connected to a computing system having a processor;assigning at the processor one of the at least two wireless sensing devices as a dominant wireless sensing device and at least one of the remaining wireless sensing devices as a subordinate wireless sensing device;operating the dominant wireless sensing device to measure a key parameter from a patient and wirelessly transmit a key parameter dataset;determining a stability indicator for the key parameter based on the key parameter dataset; andselectively operating the at least one subordinate wireless sensing device based on the stability indicator for the key parameter, wherein selectively operating the subordinate wireless sensing device includes turning off at least one subordinate wireless sensing device when the stability indicator for the key parameter is within a predetermined range, such that the respective subordinate wireless sensing device ... | General Electric Company | H04W052/02 A61B005/00 H04Q009/00 H04W052/00
US09788792 | B2 | 2017-10-17 | System for screening skin condition for tissue damage | 2016-03-04 | The present invention relates to a novel approach means of screening various skin surfaces, including hard to reach areas of the human body, as well as more commonly seen areas such as the foot. The present invention can help patients, especially those who are immobile or those with diabetes self-monitor and transmit the skin condition accurately, so that medical providers may assess the risk or presence of skin tissue breakdown. Thus, provision is made for a remote screening device that can be used in telemedicine or home self-monitoring approaches and offers high resolution image, without artifacts or other distortions that might arise from images taken of feet that have been compressed through standing and the like. | 1. A system for screening skin for tissue damage comprising: a support frame structure;at least one image acquisition device;a retention framework structure for reception of, and stable retention of, said at least one image acquisition device, said retention framework attached to said support frame structure, wherein:said at least one image acquisition device is situated with said retention framework so as to be separate from any other structure of said system for screening skin for tissue damage, so as to acquire images without artifacts or other distortions. 2. The system for screening skin for tissue damage of claim 1 further comprising a reflective-based multiple field imaging solution affixed to said support frame structure. 3. A system for screening skin for tissue damage comprising: a support frame structure;a mobility track imaging solution structure having at least one image acquisition device, said mobility track imaging solution being affixed to said support frame structure;a retention framework structure for reception of, and stable retention of, said at least one image acquisition device, said retention framework structure being affixed to said mobility track imaging solution structure wherein:said at least one image acquisition device is situated with said retention framework structure so as to be separate from any other structure of said system for screening skin for tissue damage, so as to acquire images without artifacts or other distortions. 4. The system for screening skin for tissue damage of claim 1 or 3, said system further comprising a change documentation module. 5. The system for screening skin for tissue damage of claim 4, wherein said image acquisition device is chosen from the group comprising: a thermal imager, IR imager, a photographic still image camera, or a dual use thermal/photographic still image camera, either singly or in tandem;said image acquisition devices outputting at least one of the following images chosen from the group comprising:thermal images, IR images, or photographic still images, viewable independently or together as an overlay composite image. 6. The system for screening skin for tissue damage of claim 5, said system further comprising a 3-D rendering module. 7. The system for screening skin for tissue damage of claim 5, said system further comprising a software based image acquisition module, and includes a bimodal overlay. 8. The system for screening skin for tissue damage of claims 1 or 3, wherein said image acquisition device is chosen from the group comprising: a thermal imager, IR imager, a photographic still image camera, or a dual use thermal/photographic still image camera, either singly or in tandem;said image acquisition devices outputting at least one of the following images chosen from the group comprising:thermal images, IR images, or photographic still images, viewable independently or together as an overlay composite image. 9. The system for screening skin for tissue dam... | The United States of America, as represented by Department of Veterans Affairs, Office of General Counsel, Professional Staff Group IV(024) | G03B029/00 G03B041/00 A61B005/00 A61B005/01
US09819782 | B1 | 2017-11-14 | Neurological communication device | 2017-02-07 | An ear worn brainwave brain activity dual phone communication device that works independently for enabling human beings to communicate with their mind, move objects and machines by way of their monitored brain activity. The communication system is sufficiently independent to initiate and receive communication without requiring a mobile phone, since contact information is stored in the communication device. The communication system performs functions with an audible command or brainwave brain activity such as initiating communication, accessing content, moving and controlling other machines and computer devices, accessing tools. | 1. A system comprising: An ear worn brainwave brain activity dual communication phone device that works independently and detects brainwaves signals from a brain with sensors and wirelessly; converts the detected brainwave signals into electrical signals, then into radio waves, then transmits the radio waves to antennas, then transmits the radio waves to satellites; wherein the satellites converts the detected radio waves into another radio waves; to other antennas of other ear worn brainwave brain activity dual communication phone devices; wherein the other ear worn brainwave brain activity dual communication phone devices detect the radio waves; convert the other radio waves into another electrical signals, then into another brainwaves signals; then transmits the other brainwave signals to a communication apparatus comprising: an ear mounting member being configured to be operable to at least partially engage with an ear, wherein an ear mounting member comprises a hook shape configured to at least partially wrap around the ear, an ear worn brainwave brain activity dual communication phone device being configured to be operable to initiate a communication in response to an audible command or brainwave activity, an ear worn brainwave brain activity dual communication phone device being further configured with speak recognition function to be operable to perform at least one function with an audible command or brainwave brain activity, an ear worn brainwave brain activity dual communication phone device being further configured to exchange data with an external processor, an ear worn brainwave brain activity dual communication phone device further being configured as a remote control that is operable to control at least some features or functions listed below, an ear worn brainwave brain activity dual communication phone device being further configured to comprise:a communication processor configured to initiate a communication in response to an audible command or brainwave activity;a microphone operatively connected with a communication processor, a microphone is configured to receive an audible command or brainwave activity;a task module, a task module receives an audible command or brainwave activity from a microphone to perform a voice activated operation;a network module, a network module is configured to receive a task operation from a task module, wherein a network is further configured to send a task operation to at least one of an external computer for execution;Wherein a microphone comprises an array of microphones that is configured to determine directionality of an audible command or brainwave activity, and wherein an array of microphones is further configured to differentiate multiple audible commands or brainwave activity by position;a speaker configured to transmit communication received from an external communication device, wherein a speaker comprising an adjustable volume configured to be operable with an audible command or brain... | Shavar Daniels | A61B005/0476 H04M001/725 G06F003/01 G06F003/16
US09788781 | B2 | 2017-10-17 | System and method for ligament insertion in knee joint surgeries using adaptive migration of ligament insertion geometry | 2017-03-16 | A method for providing feedback during ligament insertion in knee joint orthopedic surgeries includes providing an imaging device providing knee joint articular surface geometry data, providing a position measuring device providing distal ligament insertion sites, providing a strain measuring device providing ligament reference strain, providing a drive system providing a prescribed kinematic path for the knee joint and providing an adaptive migration application. The adaptive migration application receives inputs including the knee joint articular surface geometry data, the distal ligament insertion sites, the ligament reference strain and the prescribed kinematic path for the knee joint and generates outputs including sets of isotonic, hypertonic and hypotonic points representing proximal insertion sites of a ligament fiber associated with prescribed distal ligament insertion sites. | 1. A method for providing feedback during ligament insertion in knee joint orthopedic surgeries comprising: imaging a knee joint using an imaging device to generate knee joint articular surface geometry data;measuring the position of the ligaments in the knee joint using a position measuring device to generate distal ligament insertion sites data;measuring the strain of the ligaments in the knee joint using a strain measuring device to generate ligament reference strain data;moving the knee joint using a drive system to generate data of a prescribed kinematic path for the knee joint;receiving said knee joint articular surface geometry data, said distal ligament insertion sites data, said ligament reference strain data, and said data of a prescribed kinematic path for the knee joint by an adaptive migration application and generating outputs comprising sets of isotonic, hypertonic and hypotonic points representing proximal insertion sites of a ligament fiber associated with prescribed distal ligament insertion sites, such that the ligament fiber stretched between two isotonic points has minimal change in tension throughout a flexion arc. 2. The method of claim 1, wherein the knee joint articular surface geometry data are represented as two tangentially intersecting circular arcs. 3. The method of claim 1, wherein the adaptive migration application generates said set of isotonic, hypertonic and hypotonic points by iteratively adapting the distal ligament insertion sites based on an integral of a ligament tension vector over the flexion arc. 4. The method of claim 3, wherein said integral of the ligament tension vector trends towards a limit of zero with successive iterations of the adaptive migration application and wherein at said limit a distal ligament insertion site is determined that balances tensile and compressive strain over the flexion arc. 5. The method of claim 1, wherein the prescribed kinematic path for the knee joint comprises passive kinematics of the knee joint through the flexion arc. 6. The method of claim 1, wherein the prescribed kinematic path for the knee joint comprises one of flexion, internal rotation, anterior posterior displacement or medial lateral displacement. 7. The method of claim 1, wherein the prescribed kinematic path for the knee joint is applied while the knee joint is constrained in two degrees of freedom. 8. The method of claim 1, wherein the adaptive migration application further determines a cost function based on applied strain magnitude and direction of a strain induced tension in the ligament fiber. 9. The method of claim 8, wherein the adaptive migration application migrates the distal ligament insertion sites in a direction counterpoised to a direction of the induced tension in the ligament fiber throughout the flexion arc. | Advanced Mechanical Technology, Inc. | A61B005/103 A61B005/00 A61B005/11 A61B034/10 A61B017/17
US09782084 | B2 | 2017-10-10 | Exercise and communications system and associated methods | 2014-09-01 | An exercise and communications system includes an interactive device, a remote device, and an external device, wherein the interactive device is configured to gather data relating to a user of the system and transmit the same to the remote device, and the remote device is configured to provide analyze the data and transmit a response to the interactive device, which in turn communicates the response to the user and additionally communication with an external device for retrieval of instructions, programs, and data, inter alia. An exercise and communications system facilitates communication between a plurality of users, each having an interactive device and a remote device. | 1. An exercise and communications system comprising, in combination: a remote device in communication with an interactive device and an external device;wherein the interactive device comprises one or more sensors configured to sense one or more physiological parameters of a user, including at least end tidal carbon dioxide;wherein the interactive device is integrated into a housing configured to be disposed within an ear of the user, said housing further comprising a sensor mounted on a tube extending from the housing and configured to locate the sensor in front of a mouth of the user when said housing is disposed within the ear of the user;wherein the remote device comprises a memory coupled to a processor and is operable to:receive data for at least the physiological parameter from the interactive device;determine whether the data is within a predetermined criteria; andinitiate a response where the data is within the predetermined criteria;wherein the response includes at least one of contacting a coach, a medical professional and an emergency response center along with creating a record for the user, wherein the interactive device further comprises a GPS device, and wherein the GPS device provides audio message alerts for training zone high and low limits. 2. The exercise and communications system of claim 1 wherein the response includes receiving a user location from the GPS device and sending the user location to at least one of a coach, a medical professional and an emergency response center along with creating a record for the user. 3. The exercise and communications system of claim 1, wherein the one or more physiological parameters include heart rate, temperature, or oxygen saturation. 4. The exercise and communications system of claim 1 wherein the interactive device further comprises a user interface. 5. The exercise and communications system of claim 4 wherein the user interface comprises an audio speaker. | - | A61B005/0205 A61B005/00 A61B005/11 A61B005/1455 A63B024/00 A61B005/083 A63B071/06 A61B005/024
US09802016 | B2 | 2017-10-31 | Respiration system | 2013-01-29 | Respiration system for non-invasive positive-pressure respiration, with a pressure source providing respiratory gas, with a control and evaluation unit connected to sensors detecting a leakage volume, spontaneous respiration frequency, tidal volume and the inspiration time. The control and evaluation unit I) checks the leakage volume and reduces the inspiratory pressure assistance proceeding to ii) or triggers an alarm and returns to I), ii) checks the frequency and triggers an alarm and returns to I) or reduces or increases the inspiratory pressure and returns to I) or proceeds to step iii), iii) checks the volume and reduces or increases the inspiratory pressure and returns to I) or leaves the pressure assistance unchanged proceeding to step iv), iv) adjusts the time period of the pressure assistance, depending on the inspiration time, the time period being left unchanged if the inspiration time lies in the predefined inspiration time interval, and returns to I). | 1. A respiration system for non-invasive positive-pressure ventilation, the respiration system comprising: a pressure source for providing respiratory gas with controllable pressure;sensors to determine a leakage volume, a spontaneous respiration rate, a tidal volume and an inspiration time;a control and analyzing unit configured to control the pressure source connected with the sensors, wherein the control and analyzing unit is configured to automatically perform the following measurements and settings in the sequence indicated:i) to check the leakage volume (MV_LEAK) to determine whether the leakage volume is above a preset limit value, and to reduce an inspiratory pressure assistance if the leakage volume (MV_LEAK) is above a preset limit value and if the spontaneous respiration rate (F_SPON) is within a preset spontaneous respiration rate range and the tidal volume (VT) is within a preset tidal volume range, and to trigger an alarm if the spontaneous respiration rate (F_SPON) is outside the preset spontaneous respiration rate range or if the tidal volume (VT) is outside the preset tidal volume range and to return to a starting point, and to proceed to step ii) only if the leakage volume is below the preset limit value;ii) to trigger the alarm and to return to step i) if the spontaneous respiration rate (F_SPON) is above a preset maximum or below a preset minimum, and to compare the spontaneous respiration rate (F_SPON) with the preset spontaneous respiration rate range, which is within the range between the preset maximum and the preset minimum, and to reduce the inspiratory pressure assistance if the spontaneous respiration rate (F_SPON) is above the preset spontaneous respiration rate range, and to increase the inspiratory pressure assistance if the spontaneous respiration rate (F_SPON) is below the preset range and to return to step i) after the increase or the reduction of the inspiratory pressure assistance, and to leave the inspiratory pressure assistance unchanged and to proceed to step iii) if the spontaneous respiration rate (F_SPON) is within the preset spontaneous respiration rate range;iii) to check the tidal volume (VT) to determine whether the tidal volume is within the preset tidal volume range, and to reduce the inspiratory pressure assistance if the tidal volume is above the preset tidal volume range, and to increase the inspiratory pressure assistance if the tidal volume (VT) is below the preset tidal volume range and to return to step i) after the reduction or the increase, and to leave the inspiratory pressure assistance unchanged and to proceed to step iv) if the tidal volume (VT) is in the preset tidal volume range, andiv) to set a time period during which the inspiratory pressure assistance is carried out as a function of the patient's inspiration time by increasing the time period of the inspiratory pressure assistance if the inspiration time (TI_SPON) is below a preset inspiration time range, and reducing the time per... | Drägerwerk AG & Co. KGaA | A61M016/00 A61B005/091
US09795312 | B2 | 2017-10-24 | Method and apparatus for adjusting a blanking period for selecting a sensing vector configuration in a medical device | 2014-07-24 | A method and medical device for determining sensing vectors that includes sensing cardiac signals from a plurality of electrodes, the plurality of electrodes forming a plurality of sensing vectors, setting a blanking period and a blanking period adjustment window for the plurality of sensing vectors in response to the sensed cardiac signals, determining first signal differences during the blanking period adjustment window, and adjusting the blanking period in response to the determined first signal differences. | 1. A method of determining sensing vectors in a medical device, comprising: sensing a cardiac signal obtained from a plurality of electrodes, the plurality of electrodes forming a plurality of sensing vectors;setting a blanking period and a blanking period adjustment window for the plurality of sensing vectors in response to the sensed cardiac signal as sensed by the sensing vector;determining, for each one of the plurality of sensing vectors, first minimum signal differences between the sensed cardiac signal as sensed by the sensing vector and a sensing threshold during the blanking period adjustment window;setting a quality metric window extending subsequent to the blanking period adjustment window;determining, for each one the plurality of sensing vectors, second minimum signal differences between and the sensed cardiac signal as sensed by the sensing vector and the sensing threshold during the quality metric window;ranking the sensing vectors of the plurality of sensing vectors based on the determined second minimum signal differences; andselecting one or more sensing vectors of the plurality of sensing vectors based on the determined rankings. 2. The method of claim 1, further comprising for each of the sensing vectors, sensing an R-wave in response to the cardiac signal exceeding the sensing threshold, wherein setting the blanking period adjustment window and the quality metric window comprises setting the blanking period adjustment window and the quality metric window individually for each sensing vector based on the sensed R-wave for that sensing vector. 3. The method of claim 2, wherein determining the first minimum signal differences and the second minimum signal differences for each of the sensing vectors comprises: determining a first value of a plurality of values comprising the first minimum signal differences by determining a minimum signal difference between the cardiac signal sensed by that sensing vector and the sensing threshold during the blanking period adjustment window as set based on the sensed R-wave; anddetermining a first value of a plurality of values comprising the second minimum signal differences by determining a minimum signal difference between the cardiac signal sensed by the sensing vector and the sensing threshold during the quality metric window as set based on the sensed R-wave. 4. The method of claim 1, wherein ranking the sensing vectors of the plurality of sensing vectors based on the determined second minimum signal differences comprises: determining a vector selection metric for each of the plurality of sensing vectors based on the determined second minimum signal differences;comparing the determined vector selection metric from each of the plurality of sensing vectors to each other; anddetermining a first ranked sensing vector having a first vector selection metric, a second ranked sensing vector having a second vector selection metric less than the first vector selection metric, and a third ranked s... | Medtronic, Inc. | A61B005/04 A61B005/0408 A61B005/042 A61B005/0456
US09775504 | B1 | 2017-10-03 | Controlling an intraoral mirror with an integrated camera, and applications thereof | 2017-05-24 | Disclosed embodiments integrate a camera into an intraoral mirror. Integrating a camera into an intraoral mirror provides an efficient way to record and display what is visible to the healthcare provider in the mirror. | 1. A method for capturing photographs in a patient's mouth, comprising: (a) receiving an image captured from an image sensor affixed to a dental mirror, the dental mirror having a plurality of light sources;(b) determining a portion of the image to be darker than another portion of the image;(c) selecting, based on which portion was determined in (b) to be darker, a light source from the plurality of light sources such that the light source is selected to illuminate the determined portion more brightly than the other portion; and(d) increasing a brightness of the light source selected in (c). 2. The method of claim 1, further comprising: (e) determining whether a user has overridden automated illumination;(f) when the user is determined not to have overridden automated illumination, performing steps (a)-(d); and(g) when the user is determined to have overridden automated illumination, setting a brightness of the plurality of light sources according to a user setting. 3. The method of claim 1, further comprising: (e) determining whether a software module has overridden automated illumination;(f) when the software module is determined not to have overridden automated illumination, performing steps (a)-(d); and(g) when the software module is determined to have overridden automated illumination, setting a brightness of the plurality of light sources according to output of the software module. 4. The method of claim 1, further comprising: (e) monitoring the dental mirror to determining when it is not in use; and(f) when the dental mirror is determined not to be in use, entering into a standby mode that consumes a reduced amount of power. 5. The method of claim 4, wherein the monitoring (e) comprises monitoring an orientation measuring device on the dental mirror to determine when a time period has elapsed without substantial movement. 6. The method of claim 4, wherein the monitoring (e) comprises monitoring the image sensor on the dental mirror to determine when a time period has elapsed without substantial movement. 7. A program storage device tangibly embodying a program of instructions executable by at least one machine to perform a method for capturing photographs in a patient's mouth, comprising: (a) receiving an image captured from an image sensor affixed to a dental mirror, the dental mirror having a plurality of light sources;(b) determining a portion of the image to be darker than another portion of the image;(c) selecting, based on which portion was determined in (b) to be darker, a light source from the plurality of light sources such that the light source is selected to illuminate the determined portion more brightly than the other portion; and(d) increasing a brightness of the light source selected in (c). 8. The program storage device of claim 7, the method further comprising: (e) determining whether a user has overridden automated illumination;(f) when the user is determined not to have overridden automated illumination, performi... | DENTAL SMARTMIRROR, INC. | A61B001/247 A61B001/05 A61B001/06 A61C013/34 A61B001/00 A61B005/06 A61B005/00 A61C009/00 A61C013/00 H05B037/02
US09820721 | B2 | 2017-11-21 | Aspiration and biopsy needle apparatus and devices and applications thereof | 2015-04-23 | This application discloses aspiration/biopsy needle apparatus, devices containing these needle apparatus, and methods using these needle apparatus and devices for aspiration or biopsy of samples, in particular biological samples from mammalian subjects. The novel aspiration/biopsy needle systems can find wide applications in the manufacture of fine needle aspiration (FNA)/biopsy devices for convenient cell harvesting and tissue sampling and analysis. | 1. An aspiration/biopsy device comprising: a body, a cannula needle, and a sample collection container, wherein: the body comprises a vacuum channel and a plurality of ports;the cannula needle traverses inside the body, comprising an inlet end positioned at and extending from a first port of the body, an outlet end positioned at and extending from a second port of the body, and a middle portion enclosed inside the body;the vacuum channel traverses inside the body, comprising a first opening end positioned at the second port of the body and a second opening end positioned at a third port of the body in communication with a vacuum source; andthe sample collection container is coupled with the second port of the body, enclosing both of the outlet end of the cannula needle and the first opening end of the vacuum channel;wherein the inlet end of the cannula needle is capable of extracting a biological sample from a subject when vacuum is applied on the second opening of the vacuum channel. 2. The aspiration/biopsy device of claim 1, wherein said middle portion of the cannula needle is in airtight or substantially airtight contact with internal walls of the body surrounding the middle portion of the cannula needle. 3. The aspiration/biopsy device of claim 1, wherein the inlet and outlet ends of the cannula needle are two distinct pieces connected with each other through a connection means at the first port of the body. 4. The aspiration/biopsy device of claim 3, wherein the inlet end of the cannula needle comprises a detachable needle. 5. The aspiration/biopsy device of claim 1, wherein the second port of the body comprises a connection means to couple the body and the sample collection container. 6. The aspiration/biopsy device of claim 5, wherein the connection means is airtight or substantially airtight so that when a vacuum is applied on the vacuum channel, a sample can be extracted by the inlet end of the cannula needle to the sample collection container. 7. The aspiration/biopsy device of claim 5, wherein the connection means is a clip mechanism, a screw mechanism, or a combination thereof. 8. The aspiration/biopsy device of claim 1, wherein the sample collection container is a sample vial or centrifugation tube. 9. The aspiration/biopsy device of claim 1, wherein said vacuum source is a vacuum pump, a syringe, a vacuum line, or a combination thereof. 10. The aspiration/biopsy device of claim 1, wherein the vacuum source comprises a syringe. 11. The aspiration/biopsy device of claim 10, wherein the syringe is connected with the third port of the body through a connecting means so that the syringe can directly apply vacuum on the second opening end of the vacuum channel. 12. The aspiration/biopsy device of claim 11, wherein the connecting means is airtight or substantially airtight. 13. The aspiration/biopsy device of claim 11, wherein the connecting means is selected from clip mechanisms, screw mechanisms, and combinations thereof. 1... | Beta Pharma, Inc. | A61B005/00 A61B010/02 A61B010/00
US09814854 | B2 | 2017-11-14 | Method and system with divided cannula having low oxygen flow rate and improved end-tidal CO2 measurement | 2014-03-14 | A nasal cannula having a face piece including first and second nares which each communicate with a respective nostril of a patient. A septum divides the face piece into first and second flow paths such that the first flow path communicates with the first nare and the second flow path communicates with the second nare. A leading end of the first nare is configured to form a fluid barrier within a first nostril and facilitate only supplying an insufflation gas, via the first flow path, to the patient and prevent exhausting any exhaust gas, of the patient, through the first nostril. A leading end of the second nare is configured to facilitate collecting some of the exhaust gas, from the patient, via the second flow path as well as exhausting a remainder of the exhaust gas out through a second nostril of the patient. | 1. A nasal cannula comprising: an elongate face piece comprising first and second nares, and the first nare facilitates communication with a first nostril of a patient while the second nare facilitates communication with a second nostril of the patient;a septum dividing the face piece into first and second flow paths such that the first flow path communicating with the first nare and the second flow path communicating with the second nare;a leading end of the first nare comprising a peripherally extending sealing member which is adapted to engage with adjacent tissue of a nasal passage of the first nostril to form a fluid barrier within the first nostril of the patient to facilitate only supplying an insufflation gas, via the first flow path, to the patient and prevent exhausting any of an exhaust gas, from the patient, past the fluid barrier, the first nare and the peripherally extending sealing member being integrally formed as a single piece from a same material; anda leading end of the second nare being configured to facilitate collecting some of the exhaust gas, from the patient, via the second flow path as well as exhausting a remainder of the exhaust gas, from the patient, out through the second nostril of the patient. 2. The nasal cannula according to claim 1, wherein the nasal cannula is tubular and the septum is disposed generally centered within the face piece, between a pair of respective flow passages, which facilitate communication with the respective first and second nares. 3. The nasal cannula according to claim 2, wherein a first opening is formed in a first end of the face piece while a second opening is formed in an opposite second end of the face piece, and an oxygen supply tube is connected to the first opening at a first end of the oxygen supply tube while an end-tidal CO2 detection tube is connected to the second opening at a first end of the end-tidal CO2 detection tube, a second opposed end of the oxygen supply tube is connected to a source of an oxygen gas while a second opposed end of the end-tidal CO2 detection tube is connected to a carbon dioxide monitoring device. 4. The nasal cannula according to claim 1, wherein a first opening is formed in a first end of the face piece while a second opening is formed in an opposite second end of the face piece, and a supply tube is connected to the first opening, while a detection tube is connected to the second opening. 5. The nasal cannula according to claim 4, wherein the supply tube and the detection tube each comprise a first end connected to the face piece and a second end opposite the face piece, wherein the second end of the supply tube is connected to a source of an insufflating gas while the second end of the detection tube is connected to a carbon dioxide monitoring device. 6. The nasal cannula according to claim 1, wherein the first nare extends along an axis, at least one opening is formed in a sidewall of the first nare and the at least one opening is axially ... | Salter Labs | A61M016/06 A61B005/08 A61B005/097 A61B005/083 A61M016/20 A61M016/08
US09808532 | B2 | 2017-11-07 | Peptides whose uptake in cells is controllable | 2010-07-15 | Disclosed herein, in certain embodiments, is a selective transport molecule with increased in vivo circulation. In some embodiments, a selective transport molecule disclosed herein has the formula (A-X-B-C)n-M, wherein C is a cargo moiety; A is a peptide with a sequence comprising 5 to 9 consecutive acidic amino acids, wherein the amino acids are selected from: aspartates and glutatmates; B is a peptide with a sequence comprising 5 to 20 consecutive basic amino acids; X is a linker; and M is a macromolecular carrier. | 1. A molecule having the formula: (A-X-B-C)n-M wherein: C is a fluorescent moiety;A is a peptide with a sequence comprising a series of 5 glutamate residues;B is a peptide with a sequence comprising a series of 8 arginine residues;X is a cleavable linker comprising PRSFL (SEQ ID NO: 22);M is a macromolecular carrier; andn is an integer between 1 and 20;wherein M is bound to A or B. 2. The molecule of claim 1, wherein X is further attached to 6-aminohexanoyl, 5-amino-3-oxapentanoyl, or a combination thereof. 3. The molecule of claim 1, wherein X comprises a sequence selected from: DPRSFL (SEQ ID NO: 1), or PPRSFL (SEQ ID NO: 2). 4. The molecule of claim 1, wherein M is a macromolecular carrier selected from: a dendrimer, dextran, a PEG polymer, or albumin. 5. The molecule of claim 1, wherein M is a PEG polymer. 6. The molecule of claim 1, wherein the fluorescent moiety is an indocarbocyanine dye. 7. The molecule of claim 1, wherein the fluorescent moiety is selected from Cy5, Cy5.5, Cy7, Alexa 647, IRDYE 800CW, or a combination thereof. 8. The molecule of claim 1, wherein the fluorescent moiety is an MRI contrast agent. 9. A method of imaging thrombin activity in a subject, comprising imaging thrombin activity after the subject has been administered a molecule of the structure (A-X-B-C)n-M, wherein C is a fluorescent moiety;A is a peptide with a sequence comprising a series of 5 glutamate residues;B is a peptide with a sequence comprising a series of 8 arginine residues;X is a cleavable linker comprising PRSFL (SEQ ID NO: 22);M is a macromolecular carrier; andn is an integer between 1 and 20;wherein M is bound to A or B. 10. The method of claim 9, wherein the fluorescent moiety is an indocarbocyanine dye. 11. The method of claim 9, wherein the fluorescent moiety is selected from Cy5, Cy5.5, Cy7, Alexa 647, IRDYE 800CW, or a combination thereof. | The Regents of the University of California | A61B010/00 A61B005/00 A61B008/00 A61K047/48 A61K049/00 C07K007/06
US09782091 | B2 | 2017-10-10 | Opto-electrical device for artifact reduction | 2012-07-25 | An optical electrode having a plurality of electrodes, including a recording electrode having a roughened surface and an optical light source configured to emit light, wherein at least a portion of the light impinges on the recording electrode. Also disclosed are methods of producing an optical electrode and an opto-electronic neural interface system. | 1. An optical electrode device, comprising: a) a carrier;b) a plurality of electrodes supported on the carrier, at least one of which is a recording electrode configured to sense electrical signals from tissue, wherein the recording electrode comprises:i) an electrically conductive material selected from the group consisting of gold, platinum, iridium, platinum-iridium, titanium nitride, and iridium oxide; andii) a roughened external material having a back side attached to the electrically conductive material to thereby provide the recording electrode with a roughened external surface forming a front side of the roughened external material, wherein the roughened external material is selected from the group consisting of PEDOT/PSS (poly(3,4-ethylene-dioxythiophene)/poly(styrene sulfonate), iridium oxide, glass nanospheres, glass microspheres, poly-styrene nanospheres, and poly-styrene microspheres, andiii) wherein the roughened external surface of the recording electrode has an average roughness as an arithmetic average of absolute values of a series of maximum and minimum vertical height measurements in nm on the roughened external surface that is greater than 10 nm; andc) a light source supported by the carrier, the light source configured to emit light that is directed at the roughened external surface of the recording electrode. 2. The optical electrode device of claim 1, wherein the roughened external surface of the recording electrode has an average roughness greater than 20 nm. 3. The optical electrode device of claim 1, wherein the roughened external surface of the recording electrode has an average roughness greater than 100 nm. 4. The optical electrode device of claim 1, wherein the roughened external surface of the recording electrode has a root mean square (RMS) surface roughness greater than 100 nm. 5. The optical electrode device of claim 1, wherein the plurality of electrodes are part of a thin-film structure. 6. The optical electrode device claim 1, wherein the light source comprises a waveguide. 7. A method of producing an optical electrode with a reduced photoelectrochemical artifact, comprising the steps of: a) providing a carrier;b) supporting a plurality of electrodes on the carrier, wherein at least one of the electrodes is configured as a recording electrode comprising:i) an electrically conductive material selected from the group consisting of gold, platinum, iridium, platinum-iridium, titanium nitride, and iridium oxide; andii) a roughened external material having a back side attached to the electrically conductive material to thereby provide the recording electrode with a roughened external surface forming a front side of the roughened external material, wherein the roughened external material is selected from the group consisting of PEDQT/PSS (poly(3,4-ethylene-dioxythiophene)/poly(styrene sulfonate), iridium oxide, glass nanospheres, glass microspheres, poly-styrene nanospheres, and poly-styrene micro spheres;c) s... | NEURONEXUS TECHNOLOGIES, INC. | A61N001/00 A61B005/04 A61B005/00 A61N005/06 A61N001/05
US09788730 | B2 | 2017-10-17 | Optical process and apparatus for non-invasive detection of melanoma | 2013-03-07 | A non-invasive, optical method and device for the detection of melanoma in skin lesions. The detection of the presence of melanoma is accomplished optically by looking for specific changes (signatures) in the spectrum of optical light elastically scattered off melanoma molecules. Elastic scattering spectroscopy (ESS) converts subcellular morphological changes into scattering spectrum signatures. A melanoma discrimination analysis is performed by illuminating the lesion with a handheld device that also collects a portion of the scattered light, converts it into digital signals, analyzes the requisite spectral signatures, and provides a logical output showing the user the presence (or absence) of melanoma in the subject lesion. | 1. A process for optical detection of melanoma lesions comprising: illuminating a volume of a segment of skin sequentially with light at a plurality of wavelengths, wherein each of the plurality of wavelengths are within a range of approximately 300 nm and 950 nm;receiving an elastic scattering spectrum from the segment of skin at each of the plurality of wavelengths;determining a measure of intensity of the received elastic scattering spectrum over the plurality of wavelengths;comparing the measure of intensity of the received elastic scattering spectrum over the plurality of wavelengths to a set of spectral signatures indicative of melanoma, wherein at least one of the spectral signatures indicative of melanoma is derived at least in part from: a ratio of an amount of light scattered by melanoma-containing lesion tissue to an amount of light scattered by healthy skin in each of at least two spectral regions; anda difference between an average rate of change in the amount of light scattered by the melanoma-containing lesion tissue at a first wavelength and a second wavelength and an average rate of change in the amount of light scattered by the healthy skin at the first wavelength and the second wavelength; anddisplaying a message indicative of a presence or absence of melanoma based on the comparing. 2. The process of claim 1, wherein the measure of intensity of the received elastic scattering spectrum comprises a measure of a rate of change of an intensity measure between two of the plurality of wavelengths. 3. The process of claim 1, wherein the measure of intensity of the received elastic scattering spectrum comprises an area determined by two or more of the plurality of wavelengths. 4. The process of claim 1, wherein the plurality of wavelengths are taken from a group consisting of discrete wavelengths in a range of 300-1100 nm at discrete intervals of 1-50 nm. 5. The process of claim 1, wherein illuminating the volume of the segment of skin comprises placing a fiber optic illumination source in contact with the segment of skin and wherein receiving the elastic scattering spectrum from the segment of skin comprises receiving an elastic scattering spectrum from a fiber optic collector. 6. The process of claim 1, wherein displaying the message indicative of the presence or absence of melanoma based on the comparing comprises displaying “yes” to indicate the presence of melanoma and “no” to indicate absence of melanoma. 7. The process of claim 1, wherein displaying the message indicative of the presence or absence of melanoma based on the comparing comprises displaying a numeric value indicating the presence or absence of melanoma. 8. The process of claim 1, further comprising sending result of the comparing to a third-party medical provider. 9. The process of claim 1, wherein the light is generated at the plurality of wavelengths by a discrete narrow-band light source. 10. The process of claim 9, wherein the discrete narrow-band ligh... | DERMASENSOR, INC. | A61B005/00
US09782103 | B2 | 2017-10-10 | Portable monitoring device with hearing aid and EEG monitor | 2012-03-23 | A portable monitoring device comprises a hearing aid having a housing, a microphone (24), an acoustic signal processing means (43), and an acoustic output transducer (33). The portable monitoring device further comprises an EEG monitoring system arranged at least partly in said housing, and comprising a measuring unit (3) having electrodes (12) arranged external to the housing. The EEG monitoring system comprises a processing unit (42) for analyzing the EEG signal for identifying or predicting specific biological incidences, such as a seizure, in said person, decision means for deciding, based on said analyzed EEG signal, when an alarm or information must be provided to said person, and means for providing said alarm or information through said output transducer (33). The invention further provides a method of monitoring an EEG signal of a hearing impaired person. | 1. A portable monitoring device comprising: a hearing aid having a hearing aid housing, at least one microphone for receiving ambient sound, an acoustic signal processor configured to process a signal from the microphone in accordance with a hearing impairment of a person using said hearing aid to obtain a processed sound signal, and an acoustic output transducer configured to produce an acoustic output corresponding to the processed sound signal, andan EEG monitoring system configured to monitor EEG signals of said person, said EEG monitoring system being arranged at least partly in said hearing aid housing, said EEG monitoring system comprising a measuring unit having electrodes for measuring one or more EEG signals from said person, said electrodes being arranged external to the housing or at an outer surface of the housing,a processing unit having an EEG signal processor configured to analyze the EEG signal, said processing unit being arranged in said housing, said EEG signal processor being adapted for, based on the EEG signal, identifying or predicting specific biological incidences, in said person, said processing unit comprising a decision component for deciding, based on said analyzed EEG signal, when an alarm or information must be provided to said person, and a notification component for providing said alarm or information through said output transducer; andan adjustment component receiving a signal representing an acoustic background noise level and configured to adjust a sound level of said alarm or information according to the acoustic background noise level in order to make the alarm or information clearly discernible over background noise. 2. The monitoring device according to claim 1, wherein said adjustment component comprises said acoustic signal processor. 3. The monitoring device according to claim 1, wherein said biological incidence is hypoglycemia. 4. The monitoring device according to claim 1, wherein said EEG monitoring system is adapted for a wireless connection between the measuring unit and the processing unit. 5. The monitoring device according to claim 4, wherein said measuring unit comprises an electronic module, said electronic module being connected with the electrodes, and further being connected with a communication component for transmitting the EEG signal to the processing unit. 6. The monitoring device according to claim 5, wherein said measuring unit is implanted subcutaneously on a head of the person to be monitored. 7. The monitoring device according to claim 1, comprising a battery for providing power to the hearing aid and to the EEG monitor, where the power to the measuring unit of the EEG monitor is transferred wirelessly. 8. The monitoring device according to claim 1, wherein said measuring unit is adapted to be arranged in an ear canal with at least two electrodes adapted to be in contact with a wall of the ear canal in at least two different positions. 9. The monitoring device according to ... | Widex A/S | A61B005/04 A61B005/0484 A61B005/0482 A61B005/16 A61B005/00
US09820718 | B2 | 2017-11-21 | Enhanced electronic external fetal monitoring system | 2013-03-01 | The present invention relates to fetal monitoring and, more particularly, to an electronic external fetal monitoring system that includes a self adhering single use dermal patch including embedded sensors that can be attached to the skin of an expectant maternal patient and is configured to record fetal heart rate, uterine activity, and uterine integrity. | 1. An electronic external fetal monitoring system comprising: a. a planar dermal patch comprising a first side configured to adhere to the surface of the skin of a pregnant patient, said first side further comprising: i. a first portion embedded in said first side and configured to monitor and detect heart rate data;ii. a second portion embedded in said first side and configured to monitor and detect uterine activity data; andiii. a third portion embedded in said first side and configured to monitor and detect uterine integrity data comprising an acoustic ultrasonic Doppler flow imaging sensor configured to detect blood or amniotic fluid leaking outside the uterine wall, and wherein said uterine integrity data is selected from the group consisting of previous cesarean section integrity data and uterine scarring integrity; andb. a computer configured to temporally combine fetal heart rate data acquired from the first portion and uterine activity data acquired from the second portion. 2. The electronic external fetal monitoring system of claim 1, wherein said first side further comprises a self-adhering adhesive material. 3. The electronic external fetal monitoring system of claim 1, wherein at least one of said first portion, said second portion, and said third portion further comprises a layer of acoustic coupling gel. 4. The electronic external fetal monitoring system of claim 1, wherein the heart rate data is selected from fetal heart rate and maternal heart rate. 5. The electronic external fetal monitoring system of claim 4, wherein said first portion is further configured to measure the detected fetal heart rate pursuant to a pitch-catch ultrasound technique. 6. The electronic external fetal monitoring system of claim 4, wherein said first portion is further configured to measure the detected fetal heart rate pursuant to a real-time ultrasound technique. 7. The electronic external fetal monitoring system of claim 1, wherein each of said first portion, said second portion, and said third portion is configured to collect or record at least one selected from the group consisting of the heart rate data, the uterine activity data, and the uterine integrity data. 8. The electronic external fetal monitoring system of claim 7, wherein each of detect, collect, and record the heart rate data, the uterine activity data, and the uterine integrity data, respectively, in real time. 9. The electronic external fetal monitoring system of claim 7, further comprising a fourth portion configured to transmit the heart rate data, the uterine activity data, and the uterine integrity data to a computer device. 10. The electronic external fetal monitoring system of claim 9, wherein the transmission of data is accomplished wirelessly. 11. The electronic external fetal monitoring system of claim 1, wherein said first portion further comprises an ultrasonic transducer comprising a piezoelectric polymer strip. 12. The electronic external fetal monitoring syste... | Syracuse University | A61B008/00 A61B008/08 A61B008/02 A61B005/00 A61B005/024 A61B005/0205 A61B008/06
US09788975 | B2 | 2017-10-17 | Method of estimating soft tissue balance for knee arthroplasty | 2011-10-28 | A method is provided for evaluating the tension or laxity of the soft tissue surrounding a patient's knee joint. Based on this evaluation, a surgeon may determine a desired resection depth for a knee arthroplasty procedure that will achieve an appropriate spacing between adjacent, articulating components of the knee joint. | 1. A method of performing an arthroplasty procedure on a patient's knee joint of a leg, the knee joint including a femur, a tibia, and soft tissue, the method comprising the steps of: obtaining preoperatively a first image of the knee joint in an unloaded state and a second image of the knee joint in a loaded state, the unloaded state including no weight being applied to the tibia by the leg of the patient being laid down, and the loaded state including weight being applied to the tibia, the first image and the second image being at least one of MRI, CT and X-ray of the knee joint;measuring preoperatively at least a distance between at least a first tibial anatomic feature and a first femoral anatomic feature of the knee joint in the first image;measuring preoperatively at least a distance between at least the first tibial anatomic feature and the first femoral anatomic feature in the second image;evaluating preoperatively a difference by comparing with image processing of a computer at least the measured distances of the first image and the second image, the difference being representative of laxity of the knee joint; anddesigning preoperatively a digital model of a patient-specific cut guide that is contoured three-dimensionally to fit against one of the femur and the tibia and configured to guide a cutting tool to a desired resection depth based on the laxity of the knee joint evaluated during the evaluating step, the patient-specific cut guide adapted to be used to resect at least one of the femur and the tibia with the cutting tool guided by the patient-specific cut guide and into one of the femur and the tibia. 2. The method of claim 1, further comprising capturing the first image of the knee joint, wherein the first image is captured with the patient lying down in the supine position. 3. The method of claim 1, further comprising capturing the second image of the knee joint, wherein the second image is captured with the patient standing. 4. The method of claim 1, wherein the measuring steps comprise measuring the first image as a three-dimensional model and measuring the second image as a two-dimensional image. 5. The method of claim 1, wherein the comparing step comprises overlapping corresponding regions of the femur in the first and second images. 6. The method of claim 5, wherein the overlapping is performed automatically by a computer by performing a best-fit calculation between the first and second images. 7. The method of claim 1, wherein said evaluating comprises evaluating a rotation of the tibia medially or laterally relative to the femur as the difference being representative of the laxity of the knee joint. 8. The method of claim 1, wherein said evaluating comprises evaluating a translation of the tibia medially or laterally relative to the femur as the difference being representative of the laxity of the knee joint. 9. The method of claim 1, wherein said designing comprises designing the patient-specific cut guide with ... | ZIMMER, INC. | A61B017/15 A61F002/46 A61B034/10 B33Y080/00 A61B005/107
US09801552 | B2 | 2017-10-31 | Systems and methods for variable filter adjustment by heart rate metric feedback | 2012-07-25 | A physiological signal processing system for a physiological waveform that includes a cardiovascular signal component provides a variable high pass filter that is responsive to the physiological waveform, and that is configured to high pass filter the physiological waveform in response to a corner frequency that is applied. A heart rate metric extractor is responsive to the variable high pass filter and is configured to extract a heart rate metric from the physiological waveform that is high pass filtered. A corner frequency adjuster is responsive to the heart rate metric extractor and is configured to determine the corner frequency that is applied to the variable high pass filter, based on the heart rate metric that was extracted. Analogous methods may also be provided. | 1. A physiological signal processing system comprising: a physiological sensor that is configured to generate a physiological waveform that includes cardiovascular and pulmonary signal components therein;a variable high pass filter and a variable low pass filter that are responsive to the physiological waveform and that are configured to high pass and low pass filter, respectively, the physiological waveform in response to at least one corner frequency that is applied thereto;a heart rate metric extractor that is coupled to an output of the variable high pass filter and is responsive to the variable high pass filter and that is configured to extract a heart rate metric from the cardiovascular signal component of the physiological waveform that is output from the variable high pass filter responsive to high pass filtering the physiological waveform;a respiration rate metric extractor that is coupled to an output of and is responsive to the variable low pass filter and that is configured to extract a respiration rate metric from the pulmonary signal component of the physiological waveform that is output from the variable low pass filter responsive to low pass filtering the physiological waveform; anda corner frequency adjustor that is responsive to the heart rate metric extractor and that is configured to determine the at least one corner frequency that is applied to the variable high pass filter and the variable low pass filter from the heart rate metric that was extracted based on the output from the variable high pass filter. 2. A physiological signal processing system for a physiological waveform that includes a cardiovascular signal component therein, the physiological signal processing system comprising: a variable high pass filter that is responsive to the physiological waveform and that is configured to high pass filter the physiological waveform in response to a corner frequency that is applied thereto;a heart rate metric extractor that is coupled to an output of the variable high pass filter and is responsive to the variable high pass filter and that is configured to extract a heart rate metric from the cardiovascular signal component of the physiological waveform that is output from the variable high pass filter responsive to high pass filtering the physiological waveform; anda corner frequency adjustor that is responsive to the heart rate metric extractor and that is configured to determine the corner frequency that is applied to the variable high pass filter from the heart rate metric that was extracted based on the output from the variable high pass filter. 3. A physiological signal processing system according to claim 2 wherein the physiological waveform comprises an electrical physiological waveform including an electroencephalogram (EEG), an electrocardiogram (ECG) and/or a radio frequency (RF) waveform, an electro-optical physiological waveform including a photoplethysmograph (PPG) waveform, an electro-photoacoustic waveform inc... | Valencell, Inc. | A61B005/00 A61B005/024 A61B005/0245 A61B005/0402 A61B005/0476 A61B005/04 A61B005/08 A61B005/0205
US09788770 | B1 | 2017-10-17 | Continuous monitoring of tumor hypoxia using near-infrared spectroscopy and tomography with a photonic mixer device | 2016-04-14 | A device and a method to measure the concentrations of oxygenated and deoxygenated hemoglobin in tissue around a tumor via near-infrared (NIR) spectroscopy with a photonic mixer device (PMD) is described. | 1. A device for continuous monitoring of a tumor in a tissue region, the device comprising: a PMD camera chip; andat least two amplitude modulated near-infrared light sources horizontally separated from the PMD camera chip such that the PMD camera chip and the at least two near-infrared light sources are in a reflection geometry,wherein a first amplitude modulated near-infrared light source emits near-infrared light at a first wavelength, the first wavelength below an isosbestic point of hemoglobin, andwherein a second amplitude modulated near-infrared light source emits near-infrared light at a second wavelength, the second wavelength above the isosbestic point of hemoglobin. 2. The device of claim 1, wherein the PMD camera chip is a multi-pixel PMD camera chip. 3. The device of claim 1, wherein the at least two modulated near-infrared light sources are included in an array of modulated near-infrared light sources arranged around the PMD camera chip. 4. The device of claim 1, wherein each of the at least two modulated near-infrared light sources is modulated at different frequencies. 5. The device of claim 1, further comprising a motor to move the at least two modulated near-infrared light sources around the PMD camera chip. 6. The device of claim 1, further comprising a wireless communication system to transmit recorded data to a remote processing unit or a cloud server. 7. The device of claim 1, further comprising a wireless power system. 8. The device of claim 1, wherein a distance between the modulated near-infrared light sources and the PMD camera chip is approximately twice the depth of the tumor from the surface of the tissue region. | Verily Life Sciences LLC | G01N021/55 A61B005/1455 H04N005/33 H04N005/372 H04N005/374 A61B005/00 A61B005/145
US09784811 | B2 | 2017-10-10 | Two-channel magnetic resonance imaging | 2012-09-12 | A two-channel magnetic resonance tomography system is provided with a regulation circuit for an amplification system in order to be able to take into account different load situations of the MRI system in a flexible and efficient manner. It is thus possible to improve the MRI measurements greatly if the MRI system is set to the respective load situation beforehand by an idle state measurement. The adaptation may optionally also be carried out during the MRI measurement. Therefore, a multiplicity of completely different load situations may be taken into account in an optimized manner by the regulation circuit. | 1. A method for setting an amplification system for a two-channel magnetic resonance imaging system, the method comprising: determining a load situation, by the amplification system, with aid of an open-loop measurement; andadapting the amplification system by a control circuit, wherein the amplification system is adapted by the control circuit depending on the load situation. 2. The method as claimed in claim 1, wherein the control circuit has a feedback of an output signal of the amplification system to an input of the amplification system. 3. The method as claimed in claim 2, further comprising: setting, with the amplification system, at least one controller of the two-channel magnetic resonance imaging system depending on the load situation. 4. The method as claimed in claim 2, further comprising: carrying out a delay of the output signal out during the feedback. 5. The method as claimed in claim 4, further comprising: setting, with the amplification system, at least one controller of the two-channel magnetic resonance imaging system depending on the load situation. 6. The method as claimed in claim 1, further comprising: setting, with the amplification system, at least one controller of the two-channel magnetic resonance imaging system depending on the load situation. 7. The method as claimed in claim 6, wherein the at least one controller comprises a load-dependent feedforward controller for setting an amplification of the amplification system. 8. The method as claimed in claim 7, wherein the at least one controller comprises a further load-dependent controller having four Single Input Single Output (SISO) proportional-integral (PI) controllers. 9. The method as claimed in claim 6, wherein the at least one controller comprises a further load-dependent controller having four Single Input Single Output (SISO) proportional-integral (PI) controllers. 10. The method as claimed in claim 1, further comprising: decoupling signals upstream of the amplification system. 11. The method as claimed in claim 1, wherein the load situation is determined on the basis of at least one of the following parameters: a size of a patient;a weight of the patient;a region or an organ to be examined;a position of an examination table;a position of a body coil for carrying out a magnetic resonance imaging (MRI) measurement; ora position of the patient in relation to the examination table. 12. The method as claimed in claim 1, further comprising: carrying out the setting of the amplification system during the open-loop measurement. 13. The method as claimed in claim 12, further comprising: carrying out a magnetic resonance imaging (MRI) measurement with the patient after the open-loop measurement. 14. The method as claimed in claim 13, wherein the control circuit is adapted during the MRI measurement. 15. An apparatus for driving a two-channel magnetic resonance imaging system, the apparatus comprising: a control circuit;an amplification system configured ... | Siemens Aktiengesellschaft | G01R033/54 G01R033/36 A61B005/055 A61B005/00
US09779167 | B2 | 2017-10-03 | Management system for skin condition measurement analysis information and management method for skin condition measurement analysis information | 2013-02-15 | Service can be offered free of charge and the cost of a skin condition measuring device can be reduced by effectively using data on the occasion of obtaining an analysis result by transmitting measurement data by the skin condition measuring device to a server of a company providing a service of analyzing the measurement data. When a request is made from a contractor client to acquire data registered in a measurement data database, authentication is executed based on a contractor ID input from the contractor client. Additionally, when the measurement data database is searched from a contractor database based on the contractor ID, a search level and an access level are obtained. The contractor client is permitted to search the measurement data database within a range of the search level and the access level. | 1. A management system for skin condition measurement analysis information, comprising: a user client used by a user of a skin condition measuring device, connected to the skin condition measuring device, so as to be able to transmit data to and receive data from the skin condition measuring device, and also connected to a network to transmit and receive data;a data management server configured to transmit data to and receive data from the user client via the network, to store information related to the user, and to provide the information to the user as a primary user of the information and a secondary user of the information who is different from the user;an analysis result outputting unit configured to receive analysis result data of a skin condition obtained by analyzing measurement data measured by the skin condition measuring device, and to output the analysis result data to be displayable on the user client; anda secondary user client used by the secondary user, configured to transmit data to and receive data from the data management server, and provided with the information related to the user,wherein the user client includes: a user data transmitting unit configured to transmit user data to the data management server based on a request from the data management server, the user data being input by the user and also including personal information leading to specification of individual of the user and accompanying information excluding personal information of the users, correlated to a unique and non-duplicative client ID for each user;a measurement data transmitting unit configured to transmit the input measurement data correlated to the client ID to the data management server when the measurement data of the user measured by the skin condition measuring device is received from the skin condition measuring device; andan analysis result display unit configured to display analysis result data when the analysis result data of the measurement data is received from the analysis result outputting unit,wherein the data management server includes: a user data database in which user data is registered, the user data including the personal information and the accompanying information correlated to the client ID received from each of a plurality of user clients;a measurement data database in which the measurement data received from the user client and the client ID are registered to correlate with one another, the measurement data being registered as data in chronological order based on the acquisition time of the measurement data,a secondary user database in which a secondary user ID set for each of a plurality of secondary users is registered,a data providing unit configured to check the secondary user ID input from the secondary user client with the secondary user database when a request is made from the secondary user client to acquire data registered in the user data database or the measurement data database,extract the measurement data as data ... | HITACHI MAXELL, LTD. | G06F017/30 A61B005/00 A61B005/145 G06Q030/06 G01N021/17 H04L029/08 H04L029/06 G06F019/00
US09808208 | B1 | 2017-11-07 | Carpal tunnel infomatic monitor | 2017-04-05 | A device, system and method for monitoring carpal tunnel syndrome (“CTS”). The device comprises a body configured to be worn by a user, sensors, a processor, a vibration mechanism, and a power source. The sensors monitor a position of the user's hand to prevent CTS. The processor of the device is configured to determine if the user's hand is in a CTS position and the processor is configured to generate an alert signal to alert the user to the CTS position. | 1. A device for monitoring carpal tunnel syndrome (“CTS”), the device comprising: a body configured to be worn on a hand of a user;a plurality of piezoelectric sensors positioned on the body;a processor positioned within the body and in communication with the plurality of piezoelectric sensors;a vibration mechanism positioned within the body and in communication with the processor; anda power source positioned within the body and in communication with the processor;wherein the plurality of piezoelectric sensors monitors a position of the user's hand to prevent CTS wherein a resistance of the plurality of piezoelectric sensors changes when the user's hand moves up or down relative to a base position and left to right relative to a base position;wherein a voltage from the plurality of piezoelectric sensors is monitored by the processor;wherein the processor is configured to determine if the user's hand is in a CTS position and the processor is configured to activate an alert signal when a pre-determined change in the value of the voltage from the plurality of piezoelectric sensors is sent to the processor to alert the user to the CTS position. 2. The device according to claim 1 wherein the alert signal is a vibration pulse of the vibration mechanism for indication of an improper hand position. 3. The device according to claim 1 wherein the device further comprises a plurality of LEDs for indication of proper hand position. 4. The device according to claim 1 wherein the body is a glove or a fingerless sleeve. 5. The device according to claim 1 further comprising a wireless transceiver. 6. The device according to claim 5 wherein the wireless transceiver operates on a communication format selected from the group consisting of WiFi, BLUETOOTH, BLUETOOTH low energy, UWB, and near-field communications. 7. The device according to claim 5 utilized as part of a system for monitoring CTS, the system comprising the device and a computing device comprising a display screen, a processor,a wireless transceiver, anda software application;wherein the alert signal is received at the wireless transceiver of the computing device and the software application is configured to display a warning on the display screen to alert the user to the CTS position. 8. The system according to claim 7 wherein the computing device further comprises a keyboard. 9. The system according to claim 7 wherein the software application is configured to determine a user's incorrect wrist position from at least one of a user's words, keystroke pattern mannerisms, and mouse movements. 10. The system according to claim 7 wherein the alert signal is one of an audio alert, a visual alert, or the vibration mechanism being executed on the device. 11. The system according to claim 7 wherein device usage data and sensor data is transmitted to a remote database. 12. The device according to claim 5 utilized as part of a system for monitoring CTS, the system comprising the device and a smart wa... | Focal Wellness, Inc. | G08B001/08 A61B005/00 A61B005/11
US09808161 | B2 | 2017-11-07 | Body-worn sensor for characterizing patients with heart failure | 2016-02-15 | The invention provides a sensor for measuring both impedance and ECG waveforms that is configured to be worn around a patient's neck. The sensor features 1) an ECG system that includes an analog ECG circuit, in electrical contact with at least two ECG electrodes, that generates an analog ECG waveform; and 2) an impedance system that includes an analog impedance circuit, in electrical contact with at least two (and typically four) impedance electrodes, that generates an analog impedance waveform. Also included in the neck-worn system are a digital processing system featuring a microprocessor, and an analog-to-digital converter. During a measurement, the digital processing system receives and processes the analog ECG and impedance waveforms to measure physiological information from the patient. Finally, a cable that drapes around the patient's neck connects the ECG system, impedance system, and digital processing system. | 1. A sensor for measuring both impedance and ECG waveforms and configured to be worn on a patient's torso, comprising: (i) a first electrode patch module comprising at least first and second electrodes and a backing supporting the at least first and second electrodes; and a second electrode patch module comprising at least third and fourth electrodes and a backing supporting the at least third and fourth electrodes,wherein the first and second electrode patch modules are configured to be spatially separate and positioned on opposite sides of the medial axis of the patient's chest at approximately shoulder level below the collarbone;(ii) a flexible electronics module comprising: a digital processing system comprising a microprocessor and an analog-to-digital converter and a wireless transceiver;a first sensor module configured to reversibly mate to the first electrode patch module and to electrically connect the first and second electrodes to the electronics module via a first conductor segment, and to operably connect a first electrical connector to the electronics module via a second conductor segment;a second sensor module configured to reversibly mate to the second electrode patch module and to electrically connect the third and fourth electrodes to the electronics module via a third conductor segment, and to operably connect a second electrical connector to the electronics module via a fourth conductor segment; anda battery module configured to reversibly mate to the first electrical connector and to the second electrical connector to electrically connect the battery module to the electronics module,wherein the electronics module, the first sensor module, the second sensor module, the first conductor segment, the second conductor segment, the third conductor segment, the fourth conductor segment, the first electrical connector, and the second electrical connector are of unitary construction;the electronics module, the battery module, the first sensor module, and the second sensor module form an analog ECG circuit when the first sensor module is in electrical contact with the first electrode and the second sensor module is in electrical contact with the third electrode, the ECG circuit configured to generate an analog ECG waveform;the electronics module, the battery module, the first sensor module, and the second sensor module form an analog impedance circuit when the first sensor module is in electrical contact with the first and second electrodes and the second sensor module is in electrical contact with the third and fourth electrodes, the impedance circuit configured to generate an analog impedance waveform; andthe digital processing system is configured to receive and process the analog ECG and impedance waveforms and to generate therefrom digital ECG and impedance waveforms, to determine therefrom on a continuous basis a heart rate (HR) value and a stroke volume (SV) value for the patient, and to transmit the HR and SV values to an exter... | TOSENSE, INC. | A61B005/02 A61N001/00 A61B005/04 A61B005/0205 A61B005/0408 A61B005/053 A61B005/00 A61B005/11
US09778181 | B2 | 2017-10-03 | System and method for analyzing a sample liquid | 2013-09-11 | The present invention provides a system for analyzing a sample liquid comprising a test tape having a plurality of test elements, a tape transport device which successively transports the test elements to a sample application site while advancing the tape, and a measuring device which scans the test elements loaded with sample liquid at a measuring site. The measuring site is located at a distance from the sample application site in the direction of tape movement. The tape transport is interrupted in order to exactly position a wetted area of the test elements carrying a sample liquid at the measuring site by means of a positioning device that responds to the presence of sample liquid on the test element. | 1. A method of using a test system of the type having a tape cassette including a test tape carrying a plurality of test elements for analyzing body fluids, the method comprising: applying sample fluid to a test element of the test tape that is positioned at a sample application site, thereby creating a wetted area on the test element;irradiating the test tape with a light source while advancing the test element from the sample application site toward a sample measuring site;detecting light that is reflected from the wetted area of the test element; andusing the detected light from the wetted area of the test element to determine when the wetted area on the test element is positioned at the measuring site. 2. The method of claim 1, further comprising interrupting the advancing of the test tape when the wetted area is determined to be positioned at the measuring site. 3. The method of claim 1, wherein the detecting step further comprises detecting a drop in reflectance followed by a rise in reflectance. 4. The method of claim 3, further comprising interrupting the advancing of the test tape upon detecting the rise in reflectance. 5. The method of claim 4, further comprising irradiating the wetted area with a second light source, detecting light from the second light source that is reflected from the wetted area, and correlating the detected light reflected from the wetted area to presence or concentration of an analyte in the sample fluid. 6. The method of claim 1, further comprising irradiating the wetted area with a second light source, detecting light from the second light source that is reflected from the wetted area, and correlating the detected light reflected from the wetted area to presence or concentration of an analyte in the sample fluid. 7. The method of claim 6, further comprising using a first detector for the step of detecting light from the light source that is reflected from the test tape and using a second detector to detect light from the second light source that is reflected from the wetted area. 8. The method of claim 7, further comprising interrupting the advancing of the test tape when the wetted area is determined to be positioned at the measuring site. 9. A method of using a test system of the type having a tape cassette including a test tape carrying a plurality of test elements for analyzing body fluids, the method comprising: applying sample fluid to a test element of the test tape that is positioned at a sample application site, thereby creating a wetted area on the test element;irradiating the test tape with a light source while advancing the test element with the wetted area from the sample application site toward a sample measuring site;detecting light from the light source that is reflected from the test tape as the test element with the wetted area is advanced; andusing the light detection from the wetted area of the test element to position the test element. 10. The method of claim 9, wherein the step of... | Roche Diabetes Care, Inc. | G01N021/00 G01N021/55 A61B005/145 G01N033/487 G01N035/00 G01N021/84
US09823322 | B2 | 2017-11-21 | Method and apparatus for magnetic resonance data acquisition using a multipoint dixon technique | 2014-09-04 | In a method and magnetic resonance (MR) apparatus to acquire MR data from a subject, a predetermined spectral model of a multipoint Dixon technique is used that includes at least two spectral components with respective associated relaxation rates, a first phase due to field inhomogeneities; and a second phase due to eddy current effects. MR data are acquired using a bipolar multi-echo MR measurement sequence for multiple image points wherein, for each image point, the multi-echo MR measurement sequence alternately uses positive and negative readout gradient fields for the readout of MR signals of the MR data at at least three echo times. The at least two spectral components are determined based on the MR data. | 1. A method for acquiring magnetic resonance data from an examination subject, comprising: operating a magnetic resonance data acquisition unit, that generates a basic magnetic field in which an examination subject is situated, according to a bipolar multi-echo data acquisition sequence, to acquire magnetic resonance data from the examination subject for multiple image points by, for each image point, alternatingly activating positive and negative readout gradient fields with a gradient system of said acquisition unit, during which magnetic resonance signals are read out during at least three echo times, said activation of said gradient fields producing eddy currents in said data acquisition unit by interaction with said basic magnetic field;providing said magnetic resonance data to a computerized processor and, in said computerized processor, operating on said magnetic resonance data with a spectral model of a multipoint Dixon technique that relates said magnetic resonance data to at least two spectral components of said magnetic resonance signals and respective, associated relaxation rates of said at least two spectral components, and a first phase of said magnetic resonance signals that occurs due to field inhomogeneities of said basic magnetic field, and a second phase that occurs due to effects of said eddy currents, to implement an at least partially numerical determination of said at least two spectral components, said relaxation rates, said first phase, and said second phase; andmaking said at least two spectral components available in electronic form at an output of said processor in a format allowing reconstruction of an in-phase image and two out-of-phase images according to the multipoint Dixon technique that is modeled by said spectral model. 2. A method as claimed in claim 1 wherein said spectral model comprises a positive prefactor of said second phase for said magnetic resonance signals that are read out when said positive gradient fields are activated, and a negative prefactor of said second phase for said magnetic resonance signals that are read out when said negative readout gradient fields are activated. 3. A method as claimed in claim 1 comprising: in said computerized processor, determining a computational grid of lower resolution relative to a resolution of said magnetic resonance data, wherein each grid point of said computational grid includes a predetermined number of adjacent image points of said magnetic resonance data; andin said partially numerical determination of said first phase and said second phase, using an equation that is based on at least one of said first phase and said second phase being constant within a grid point of said computational grid. 4. A method as claimed in claim 1 comprising, in said at least partially numerical determination: implementing a first numerical optimization that determines said first phase;purging said magnetic resonance data of influences of said first phase depending on the d... | Siemens Healthcare GmbH | G01R033/485 A61B005/055 G01R033/46 G01R033/48 G01R033/565
US09814887 | B2 | 2017-11-14 | Selection of optimal accelerometer sensing axis for rate response in leadless pacemaker | 2014-02-06 | A medical device and associated method record signals from each real axis of a multi-axis sensor. An optimal axis for monitoring a physiological signal of the patient is identified from the real axes of the multi-axis sensor and multiple virtual axes. Coordinates defining the optimal axis are stored as respective weighting factors of the signals from each real axis of the multi-axis sensor. A metric of the physiological signal is determined using the multi-axis sensor signals and the weighting factors. | 1. A method for automatically monitoring a patient by an implantable medical device, comprising: receiving signals from each real axis of a multi-axis sensor by a processor of the implantable medical device;generating from the real axis signals a virtual signal for each of a plurality of virtual axes of the multi-axis sensor;determining a first value of an activity metric during a first activity state of the patient for each of the real axis signals and each of the virtual signals;determining a second value of the activity metric during a second activity state of the patient for each of the real axis signals and each of the virtual signals;identifying by the processor an optimal signal axis from the real axes and the plurality of virtual axes for monitoring a physiological signal of the patient by identifying one of the real axis signals and the virtual signals having at least one of a lowest variation of the first value of the activity metric during the first activity state and a greatest difference between the first value and the second value of the activity metric;storing a set of coordinates defining the optimal axis as respective weighting factors of the signals received from each real axis of the multi-axis sensor;monitoring the physiological signal of the patient by determining a third value of the activity metric using the multi-axis sensor signals and the stored weighting factors; andcontrolling a rate of delivering a therapy to the patient by a signal generator of the implantable medical device based on the third value of the activity metric. 2. The method of claim 1, wherein the multi-axis sensor is an accelerometer and further comprising: computing by the processor an activity metric for each real axis of the accelerometer and for the plurality of virtual axes of the accelerometer; andidentifying the optimal axis by a comparative analysis of the activity metrics. 3. The method of claim 2, wherein the activity metrics are determined for each real axis and for each of the plurality of virtual axes during at least two different activity states of the patient comprising the first activity state and the second activity state, wherein the first activity state is a resting state. 4. The method of claim 3, wherein the comparative analysis comprises determining at least one of a difference, a ratio and a slope between the activity metrics for the at least two different activity states for each of the real axes and the plurality of virtual axes. 5. The method of claim 1, wherein the first activity state is a resting state and identifying the optimal axis comprises identifying an axis producing a lowest variation of the first value of the activity metric during the resting state of the patient. 6. The method of claim 1, further comprising: sensing a posture signal;identifying by the processor an optimal axis from the real axes and the plurality of virtual axes for monitoring the physiological condition of the patient during the sensed postu... | Medtronic, Inc. | A61B005/0402 A61B005/00 A61B005/11 A61N001/365 A61N001/375
US09808632 | B2 | 2017-11-07 | Implantable medical device with dual-use communication module | 2015-01-23 | An implantable medical device comprises a communication module that comprises at least one of a receiver module and a transmitter module. The receiver module is configured to both receive from an antenna and demodulate an RF telemetry signal, and receive from a plurality of electrodes and demodulate a tissue conduction communication (TCC) signal. The transmitter module is configured to modulate and transmit both an RF telemetry signal via the antenna and a TCC signal via the plurality of electrodes. The RF telemetry signal and the TCC signal are both within a predetermined band for RF telemetry communication. In some examples, the IMD comprises a switching module configured to selectively couple one of the plurality of electrodes and the antenna to the receiver module or transmitter module. | 1. An implantable medical device configured for implantation in a patient comprising: an antenna;a sensing module configured to receive a cardiac electrogram of the patient via a plurality of electrodes;a communication receiver module configured to receive and demodulate signals within a predetermined frequency band for radio-frequency (RF) telemetry communication, wherein the communication receiver module is configured to receive from the antenna and demodulate a first RF telemetry signal emitted by an external device outside of the patient, and receive from the plurality of electrodes and demodulate a first tissue conduction communication (TCC) signal emitted by another implantable medical device implanted within the patient, wherein the first RF telemetry signal and the first TCC signal are within the predetermined band for RF telemetry communication;a switching module comprising one or more switches, wherein the switching module is configured to selectively couple the plurality of electrodes to either the sensing module or the communication receiver module; anda processor configured to: determine a heart rate of the patient based on depolarizations sensed by the sensing module in the cardiac electrogram;determine that the heart rate is greater than a threshold value; andchange the switching module from a first configuration in which the antenna is coupled to the communication receiver module and the plurality of electrodes are decoupled from the communication receiver module to a second configuration in which the antenna is decoupled from the communication receiver module and the plurality of electrodes are coupled to the communication receiver module in response to the determination that the heart rate is greater than the threshold value. 2. The implantable medical device of claim 1, further comprising a communication transmitter module, wherein the switching module is configured to selectively couple either the antenna or the plurality of electrodes to at least one of the communication receiver module or the communication transmitter module, wherein the communication transmitter module is configured to modulate and transmit signals within the predetermined frequency band for RF telemetry communication, wherein the communication transmitter module is configured to modulate and transmit a second RF telemetry signal to the external device via the antenna, and modulate and transmit a second TCC signal to the other implantable medical device via the plurality of electrodes, wherein the second RF telemetry signal and the second TCC signal are within the predetermined band for RF telemetry communication. 3. The implantable medical device of claim 2, wherein the implantable medical device comprises a leadless pacing device comprising: a therapeutic signal generator configured to deliver cardiac pacing pulses to a heart of the patient via the plurality of electrodes; anda housing configured for implantation within the heart of the patient, wherein... | Medtronic, Inc. | A61N001/372 A61N001/375 A61N001/39 A61N001/362 A61N001/365 A61B005/0452 A61B005/00 A61N001/05
US09827007 | B2 | 2017-11-28 | Tissue collection needle | 2015-02-18 | A tissue collection needle with a hollow interior provides two opposing tips with beveled edges on facets that drive tissue toward the hollow interior where the tissue can be gathered and severed from surrounding tissue as the needle advances through a biological specimen. The points and bevels are advantageously formed from a minimal number of cuts amenable to rapid fabrication from readily available hollow needle stock. | 1. A method for fabricating a needle comprising: providing a tube having an axis through a hollow interior and an open end;forming a first cut in the open end at a first angle to the axis;forming a second cut in the open end at a second angle to the axis opposing the first angle so that the second cut meets the first cut to remove a wedge from the end of the tube, thereby providing a first tip and a second tip on opposing sides of a circumference of the tube;forming a third cut at a third angle to the axis, the third cut intersecting the first tip and the second tip; andforming a fourth cut at a fourth angle to the axis, the fourth cut rotated about the axis by a rotation angle such that the third cut intersects the fourth cut in a line between the first tip and the second tip, thereby providing a cutting surface on the first tip and the second tip in the open end of the tube, with a bevel on the first tip directed toward the axis, wherein the first cut and the second cut meet perpendicular to and offset from the axis of the tube, thereby forming an asymmetrical wedge with the first tip extending further from the open end than the second tip. 2. The method of claim 1 wherein at least one of the first cut, the second cut, the third cut, and the fourth cut is formed with a rotary cutting tool. 3. The method of claim 2 wherein the rotary cutting tool includes a rotary blade or a rotary grinding tool. 4. The method of claim 1 wherein at least one of the first cut, the second cut, the third cut, and the fourth cut is formed with an electric discharge machine cutting tool. 5. The method of claim 1 wherein the rotation angle between the third cut and the fourth cut is about sixty degrees around the axis. 6. The method of claim 1 wherein the first cut and the second cut meet to form a slot substantially perpendicular to the axis. 7. The method of claim 1 wherein the first cut is in a plane about fifteen degrees off the axis of the tube. 8. The method of claim 1 wherein the second cut is in a plane about twenty degrees off the axis of the tube. 9. The method of claim 1 wherein the tube has an outside diameter of about 0.050 inches and an inner diameter of about 0.038 inches. 10. The method of claim 1 wherein the tube is formed of surgical stainless steel. 11. The method of claim 1 further comprising providing a stylet shaped and sized to fit within the hollow interior of the tube, placing the stylet in the tube, and shaping a tip of the stylet to coincide with the end of the tube by forming at least the third cut and the fourth cut while the stylet protrudes through the open end. | Massachusetts Institute of Technology | A61B005/00 A61B017/3205 A61B010/02 A61B017/34 A61M005/32 A61B017/00 A61B017/32
US09788741 | B2 | 2017-10-17 | Method and system for evaluation of functional cardiac electrophysiology | 2016-07-27 | An organ evaluation device, system, or method is configured to receive electrophysiological data from a patient or model organism and integrates the data in a computational backend environment with anatomical data input from an external source, spanning a plurality of file formats, where the input parameters are combined to visualize and output current density and/or current flow activity having ampere-based units displayed in the spatial context of heart or other organ anatomy. | 1. A method for evaluating an organ in an individual in need thereof, comprising: a) receiving magnetic field data associated with said organ from an organ electromagnetic sensor;b) determining an electrical current value associated with said organ from said magnetic field data, wherein said electrical current value is determined in Ampere units or equivalent units of current measurement;c) determining a vector representing said electrical current value, wherein a magnitude of said vector comprises said electrical current value;d) identifying a position of said vector in reference to said organ; ande) overlaying said vector onto an image of said organ to generate an electrical current density map of said organ for use in evaluating said organ. 2. The method of claim 1, wherein evaluating said organ is non-invasive. 3. The method of claim 1, wherein said organ is a heart or a brain. 4. The method of claim 3, further comprising identifying and defining an arrhythmia in said heart. 5. The method of claim 4, further comprising identifying an arrhythmogenic focus in said heart. 6. The method of claim 3, further comprising identifying an ischemic focus in said brain. 7. The method of claim 1, further comprising calculating a plurality of distributed, spatially accurate electrical current values having dynamic, real time parameters. 8. The method of claim 1, wherein the plurality of electrical current values are in units of amperes or equivalent units of current measurement. 9. The method of claim 1, further comprising receiving said image of said organ from an imaging system. 10. The method of claim 9, wherein said imaging system comprises an ultrasound system or a fluoroscope. 11. The method of claim 9, wherein said imaging system comprises a computer tomography system or magnetic resonance system. 12. The method of claim 1, wherein said image is a three dimensional image. 13. The method of claim 12, wherein said three dimensional image is generated from two or more two dimensional images. 14. The method of claim 1, further comprising receiving demographic data associated with said individual. 15. The method of claim 14, wherein said demographic data comprises one or more of age, race, gender, and medical history of said individual. 16. The method of claim 15, wherein said received demographic data associated with said individual is used to identify a three dimensional image of an organ of a different individual, wherein the demographic data associated with said individual and demographic data associated with said different individual are essentially identical in one or more of said age, said race, said gender, and said medical history. 17. The method of claim 15, wherein said image comprises said image of said organ of said different individual. 18. The method of claim 1, further comprising receiving biometric data associated with said individual. 19. The method of claim 18, wherein said biometric data comprises one or more of heart... | GENETESIS LLC | A61B005/04 A61B005/00 A61B005/021 A61B005/024 A61B005/055 A61B006/03 A61B006/00 A61B008/08 A61B090/00 A61B005/0205 A61B005/01
US09814432 | B2 | 2017-11-14 | Device and method for assisting the aligning of a dockable patient support | 2015-12-04 | A device for assisting the aligning of a patient support dockable onto an operating table is provided A sensing unit is configured to be arranged on the patient support or the operating table and configured to sense a feature of the patient support and/or of the operating table, said feature lying opposite the sensing unit. The sensing unit configured to sense a position-dependent physical interaction and herefrom to generate a signal that depends on the spatial relative position between the patient support and the operating table. An evaluating unit is configured to receive the signal from the sensing unit and to generate an output that characterizes the relative position between the operating table and the patient support. | 1. A device for assisting the aligning of a moveable patient support dockable with an operating table, wherein the device comprises: a sensor arranged on one of the moveable patient support or the operating table to sense a feature of the patient support and/or of the operating table, said feature lying opposite the sensor on another of the patient support or the operating table, the other different than the one, the sensor configured to sense a position-dependent physical interaction and to generate a signal that depends on a spatial relative position between the patient support and the operating table,a light source arranged on the patient support, wherein the feature is a light signal emitted by the light source and reflected by at least one reflector arranged on the operating table; anda controller configured to receive the signal from the sensor and to generate an output that characterizes the relative position between the operating table and the patient support. 2. The device as claimed in claim 1, wherein the output of the controller is a feedback signal to be perceived by an operator. 3. The device as claimed in claim 2 wherein the feedback signal comprises a visually and/or acoustically perceptible feedback signal. 4. The device as claimed in claim 1, further comprising a motorized displacement unit for moving the patient support, the motorized displacement unit arranged on the patient support, wherein the output of the controller is a control signal that controls the displacement unit such that a movement of the patient support to a target position takes place. 5. The device as claimed in claim 4, wherein the device additionally comprises a dead man's circuit connected to the motorized displacement unit such that activation of the dead man's circuit feeds a signal to the controller. 6. The device as claimed in claim 1, wherein the sensor comprises at least one optical sensor. 7. The device as claimed in claim 6, wherein the light source is arranged on the patient support disposed opposite the sensor or on the operating table disposed opposite the sensor, and wherein the sensor comprises at least one detector matrix configured to sense the light signal. 8. The device as claimed in claim 7, wherein the at least one detector matrix is arrangeable in a horizontal plane, a vertical plane, and/or in a diagonal plane. 9. The device as claimed in claim 6, wherein the sensor comprises a detector matrix arranged on the patient support, the detector matrix is configured to detect the light signal. 10. The device as claimed in claim 1, wherein the physical interaction takes place optically, acoustically, electrically, magnetically, and/or mechanically. 11. A system comprising: a patient support;a sensor arranged on the patient support or an operating table and configured to sense a feature of the patient support and/or of the operating table, said feature lying opposite the sensor on another of the patient support or the operating table, ... | Siemens Aktiengesellschaft | G01B011/26 A61B006/04 A61B005/055 A61B005/00
US09827052 | B2 | 2017-11-28 | Two-part tracking reference structure | 2013-11-04 | The present invention relates to a tracking reference structure for localizing and tracking an object by means of a medical tracking system, said structure comprising: —a first part (1) which forms a support structure for at least one tracking marker (3); and—a second part (2) which is configured to be fixed to said object, wherein a positionally fixed connection between the first part (1) and the second part (2) is established by means of an interface comprising at least one resiliently articulated element (4) on the first part (1) and/or second part (2), which engage(s) with the respective other part (2, 1), and wherein the resiliently articulated element (4) is configured such that its restoring spring force alone is already sufficient to positionally fix the connection. The present invention also relates to a tracking reference system comprising such a tracking reference structure which in turn comprises at least one first part (1), wherein any additional first part(s) (1) support(s) a different type of tracking marker (3) and said different first parts (1) can be interchangeably connected to the second part (2), and wherein the tracking markers (3) of each of said different first parts (1) are in particular placed in the same spatial position when being coupled to the second part (2). | 1. A tracking reference structure for localizing and tracking an associated object by a medical tracking system, the tracking reference structure comprising: a first part which forms a support structure for at least one tracking marker; anda second part which is configured to be fixed to said object;wherein a form-fit connection between the first part and the second part is established by a receptacle formed in the first part, wherein the receptacle receives the second part, and wherein the receptacle comprises at least one resiliently articulated element, which engages around a correspondingly formed structure of the second part;wherein a restoring spring force forces the resiliently articulated element into a first position where it engages around the correspondingly formed structure; andwherein the tracking reference structure further comprises at least one lever for the at least one resiliently articulated element, wherein the lever deflects the resiliently articulated element counter to the restoring spring force and in a direction which is substantially perpendicular to a direction in which the first part and the second part are moved relative to each other in order for the receptacle to receive the second part, and into a second position where the resiliently articulated element does not engage around the correspondingly formed structure. 2. The tracking reference structure according to claim 1, wherein the receptacle allows the first part and the second part to be located on opposite sides of a surgical drape without penetrating the drape when it is positioned between the first part and the second part. 3. The tracking reference structure according to claim 2, wherein the receptacle establishes a predetermined spatial position of the first part relative to the second part, irrespective of whether or not a surgical drape is placed between the first part and the second part. 4. The tracking reference structure according to claim 1, further comprising a positioning aid comprising elements which are specifically formed on the first part to retain drape material, and prevent the drape material from wrinkling in an area in which the at least one resiliently articulated element engages with the second part, and wherein the positioning aid also comprises receiving areas which are specifically formed on the second part and that accommodate and retain the drape material. 5. The tracking reference structure according to claim 4, wherein the elements of the positioning aid are configured to automatically tighten the drape in an area in which the resiliently articulated element engages with the second part when the first part is connected to the second part. 6. The tracking reference structure according to claim 1, wherein the restoring spring force is applied to the second part. 7. The tracking reference structure according to claim 1, wherein the receptacle comprises a plurality of discrete contact areas. 8. The tracking reference structure a... | Brainlab AG | A61B005/05 A61B019/08 A61B019/00 F16B002/18 A61B046/00 A61B005/00 A61B005/055 A61B005/06 A61B006/03 A61B008/08 A61B017/00 A61B090/00 A61B034/20
US09801639 | B2 | 2017-10-31 | Lateral spondylolisthesis reduction cage | 2014-09-25 | An intervertebral fusion device comprising inferior and superior fusion cage devices that provide an ability to correct spondylolisthesis via in-situ adjustment. | 1. An intervertebral fusion device for correcting spondylolisthesis in a patient, comprising: a) an upper cage having an anterior wall, a posterior wall, and a proximal wall and a distal wall connecting the anterior and posterior walls, and an upper surface for contacting an upper vertebral body and a lower surface having a first groove therein,b) a lower cage having an anterior wall, a posterior wall, and a proximal wall and a distal wall connecting the anterior and posterior walls, and a lower surface for contacting a lower vertebral body and an upper surface having a second groove therein,wherein the lower surface of the upper cage contacts the upper surface of the lower cage so that the first and second grooves form a first throughhole running from the proximal wall to about the distal wall; andc) a locking plate disposed in the first throughhole, wherein the locking plate is disposed in the first throughhole when the lower surface of the upper cage contacts the upper surface of the lower cage; and(d) a rotary spreader rotatably disposed between the upper cage and the lower cage such that, when rotated, the spreader causes the upper endplate to translate relative to the lower endplate in an anterior-posterior direction. 2. The device of claim 1 wherein the first groove is present upon the lower surface of the anterior wall of the upper cage, and the second groove is present upon the upper surface of the anterior wall of the lower cage. 3. The device of claim 1 wherein the first groove is present upon the lower surface of the posterior wall of the upper cage, and the second groove is present upon the upper surface of the posterior wall of the lower cage. 4. The device of claim 1 further comprising a third groove present upon the lower surface of the upper cage between the anterior and posterior walls, and a fourth groove present upon the upper surface of the lower cage between the anterior and posterior walls, and wherein the lower surface of the upper cage contacts the upper surface of the lower plate so that the third and fourth grooves form a second throughhole running from the proximal wall to about the distal wall, the second throughhole adapted for insertion of the spreader therein. 5. An intervertebral fusion device comprising: a) an upper endplate having an anterior wall, a posterior wall, and a first side wall and a second side wall connecting the anterior and posterior walls, and an upper surface for contacting an upper vertebral body and a lower surface,b) a lower endplate having an anterior wall, a posterior wall, and a first side wall and a second side wall connecting the anterior and posterior walls, and a lower surface for contacting a lower vertebral body and an upper surface,wherein the anterior wall of the upper endplate is connected to the anterior wall of the lower endplate by a first dual linkage;wherein the posterior wall of the upper endplate is connected to the posterior wall of the lower endplate by a second dual l... | DePuy Synthes Products, Inc. | A61F002/44 A61B017/16 A61B017/17 A61M029/02 A61F002/46 A61B005/00 A61B017/02 A61B017/34 A61M025/00 A61M025/01
US09820653 | B2 | 2017-11-21 | Methods, systems, and devices for imaging microscopic tumors | 2016-03-18 | An imager for in vivo viewing of diseased tissue by way of fluorescently conjugated molecules. A generally planar imaging surface with a microlens array. The imager may be modular, such that a plurality of generally planar imaging surfaces can be used to image various aspects of disease tissue simultaneously. Certain implementations include an angle-selective imager, wherein light from substantially perpendicular to the plane of the imager is received, while incident light is selectively eliminated. | 1. A system for imaging biological material in a patient, comprising: a. a fluorescently conjugated molecule capable of binding to the biological material;b. a light source;c. a fiber optic light guide;d. an elongate probe sized for placement inside the cavity of the patient, further comprising at least one substantially planar detection surface, the detection surface further comprising: i. an imager;ii. a waveguide in luminary communication with the light source; andiii. an optical filter; ande. a visualization system in electrical communication with the probe;wherein the waveguide is capable of emitting light toward the biological material, such that the biological material's emitted fluorescence is received by the imager for display by way of the visualization system. 2. The system of claim 1, further comprising a microlens array. 3. The system of claim 1, wherein the imager further comprises at least one photodiode. 4. The system of claim 3, further comprising a complementary metal oxide semiconductor process. 5. The system of claim 1, further comprising a charge-coupled device process. 6. The system of claim 2, further comprising a waveguide stencil configured to eliminate oblique light from reaching the microlens array. 7. The system of claim 1, wherein the fluorescently conjugated molecule binds biological material selected from the group consisting of: breast cancer cells, prostate cancer cells, cancer cells inside a tumor bed, cancer cells surrounding a tumor bed, disease cells in the microenvironment surrounding a tumor bed. 8. The system of claim 1, wherein optical filter is directly patterned on the imager surface. 9. The system of claim 1, wherein the waveguide further comprises a plurality of optical gratings. 10. The system of claim 1, further comprising a surgical tool, wherein the elongate probe is operationally coupled with the surgical tool. 11. An angle selective imager, comprising: a) a light source further comprising a light guide;b) an elongate probe sized to be positioned within a body cavity of a patient undergoing surgery, wherein the device is operationally coupled with the light source by way of the light guide, said device further comprising at least one substantially planar detection surface, the detection surface further comprising: i. an imager further comprising a plurality of pixels; andii. a microlens array; andc) a visualization system in electrical communication with the probe capable of computing and displaying fluorescencewherein the detection surface is adapted to allow light from substantially perpendicular angles from surface to pass through the microlens to the imager, and exclude light incident form other directions. 12. The imager of claim 11, further comprising a micro-grating. 13. The imager of claim 11, further comprising a nano-grating. 14. The imager of claim 13, further comprising at least one cylinder disposed adjacent to the microlens array adjacent to an imager pixel. 15. The ima... | Regents of the University of California | A61B001/07 A61B005/00 A61K049/00