## Sheet1
公告(開)號 | KindCode | 公告(開)日 | 名稱 | 申請日 | 摘要 | 申請範圍 | 專利權人 | IPC
US10716647 | B2 | 2020-07-21 | Methods and systems for generating color images | 2019-07-08 | Systems for generating in-focus color images are provided. Related methods and devices are also provided. | 1. A system for generating focused color image data, the system comprising: a first light source that produces an array of light beams;a second light source that produces polychromatic light;an optics system optically coupled to the first and second light sources and configured to combine the polychromatic light with the array of light beams and focus each through a probe onto an object, said optics system focusing first and second wavelengths of the polychromatic light to different focal lengths;a detector configured to collect image data of the object for the polychromatic light and the array of light beams; anda processor configured to generate focused color image data for an image location on the object corresponding to a light beam of the array of light beams based on the first and second wavelengths. 2. The system of claim 1, wherein generating the focused color image data comprises: determining, for each of the first and second wavelengths, first and second focal lengths at which the respective first and second wavelengths are focused relative to the object at the image location; andcombining the image data for each of the first and second wavelengths at the determined respective focal lengths, thereby generating the focused color image data for the image location. 3. The system of claim 2, wherein: the polychromatic light further comprises a third wavelength focused by the optics system to a different focal length than the first and second wavelengths, andgenerating the focused color image data further comprises: determining a third focal length at which the third wavelength is focused relative to the object at the image location; andcombining the image data for each of the first, second, and third wavelengths at the determined respective focal lengths, thereby generating the focused color image data for the image location. 4. The system of claim 3, wherein the polychromatic light comprises a first wavelength range including the first wavelength, a second wavelength range including the second wavelength, and a third wavelength range including the third wavelength. 5. The system of claim 4, wherein the first wavelength range is between about 65 nm and about 485 nm, the second wavelength range is between about 500 nm and about 520 nm, and the third wavelength range is between about 640 nm and about 660 nm. 6. The system of claim 1, wherein the second light source illuminates the object with the first and second wavelengths simultaneously. 7. The system of claim 1, wherein the first and second wavelengths are different from a wavelength of the first light source. 8. The system of claim 1, wherein the optics system comprises a movable optical component configured to, for each of the first and second wavelengths, scan through a corresponding plurality of respective focal lengths. 9. The system of claim 1, further comprising a scanning system configured to collect surface topology data of the object. 10. The system of claim 9, wherein ... | Align Technology, Inc. | G01B011/24 A61C009/00 A61B005/00 A61B005/107 G02B007/36 H04N005/232 H04N005/225 H04N009/04 A61C001/08 A61B001/06 A61B001/24
US10722144 | B2 | 2020-07-28 | Information processing apparatus, information processing system, and insole | 2016-05-12 | An information processing apparatus including circuitry configured to acquire weight distribution information of a subject by at least one sensor, analyze the acquired weight distribution information, and initiate a providing of feedback information based on the analyzed weight distribution information. | 1. An information processing apparatus comprising: a communication device; andcircuitry configured to: acquire weight distribution information of a subject measured by at least one sensor;acquire another weight distribution information of another subject measured by another at least one sensor;analyze the acquired weight distribution information by comparing the acquired weight distribution information to the acquired another weight distribution information;generate feedback information based on the analyzed weight distribution information; andinitiate an outputting, by the communication device, of the feedback information to an output device to output the feedback information to the subject,wherein the communication device is configured to output the feedback information to the output device so as to enable the output device to output the feedback information to the subject. 2. The information processing apparatus according to claim 1, wherein the subject is a person. 3. The information processing apparatus according to claim 1, wherein the subject is an object. 4. The information processing apparatus according to claim 1, wherein the feedback information provides guidance for moving or shifting weight of the subject. 5. The information processing apparatus according to claim 1, wherein the feedback information comprises instruction data. 6. The information processing apparatus according to claim 1, wherein the feedback information comprises weight distribution data including strength and distribution of weight. 7. The information processing apparatus according to claim 1, wherein the output device includes a display. 8. The information processing apparatus according to claim 1, wherein the output device includes an eyewear terminal. 9. The information processing apparatus according to claim 1, wherein the output device includes a touch panel. 10. The information processing apparatus according to claim 1, wherein the output device includes a speaker. 11. The information processing apparatus according to claim 1, wherein the output device includes a stimulation mechanism, and wherein the stimulation mechanism generates stimulus to a sole. 12. The information processing apparatus according to claim 1, wherein the acquired weight distribution information is analyzed by comparing the acquired weight distribution information to a model weight distribution information. 13. The information processing apparatus according to claim 12, wherein the feedback information is generated based on a result of the comparison of the acquired weight distribution information to the model weight distribution information. 14. The information processing apparatus according to claim 12, wherein the feedback information is outputted when a difference between the acquired weight distribution information and the model weight distribution information exceeds a predetermined value. 15. The information processing apparatus according to claim 1, wherein the feedba... | SONY CORPORATION | A61B005/22 G16H050/30 G16H020/30 G16H050/20 G16H040/63 G16H010/60 G01L001/16 G01L001/20 A61B005/103 A61B005/00 G06F019/00
US10716508 | B2 | 2020-07-21 | System and method for determining vessel size | 2016-10-07 | A surgical system includes at least one light emitter generating light at a first intensity and an array of light sensors including a least one row of light sensors, individual light sensors in the row of light sensors adapted to generate a signal including a non-pulsatile component. The system also includes a controller coupled to the array of light sensors, the controller including an analyzer to determine the magnitudes of the non-pulsatile components at the individual light sensors in the row of light sensors, to determine if the non-pulsatile component transitions from a higher magnitude to a lower magnitude and from a lower magnitude to a higher magnitude, and if so, to determine if the first intensity should be changed to a second intensity. | 1. A surgical system used to determine a size of a vessel within a region proximate to a working end of a surgical instrument, comprising: at least one light emitter disposed at the working end of the surgical instrument, the at least one light emitter generating light at a first intensity;an array of light sensors disposed at the working end of the surgical instrument opposite the at least one light emitter, the array of light sensors comprising at least one row of light sensors, each individual light sensor in the at least one row of light sensors adapted to generate a signal comprising a non-pulsatile component; anda controller coupled to the array of light sensors, the controller comprising an analyzer to determine the magnitudes of the non-pulsatile components at the individual light sensors in the at least one row of light sensors, to determine if the non-pulsatile component transitions from a higher magnitude to a lower magnitude and from a lower magnitude to a higher magnitude, and if the non-pulsatile component transitions from a higher magnitude to a lower magnitude and from a lower magnitude to a higher magnitude, to determine if the first intensity should be changed to a second intensity, and to change the at least one light emitter to a second intensity, if so determined,wherein the controller determines if the first intensity should be changed to a second intensity in accordance with an angle determined using a line drawn between (a) a point along the non-pulsatile profile where the derivative changes to a non-zero value and (b) the point corresponding to the lowest non-pulsatile value, or an angle determined using a line drawn between (a) a point along the non-pulsatile profile where the derivative changes from a non-zero value and (b) the point corresponding to the lowest non-pulsatile value. 2. The surgical system according to claim 1, wherein the controller determines if the angle is in a range from a first value to a second value, the second value being larger than the first value. 3. The surgical system according to claim 2, wherein the first value is 45 degrees and the second value is 65 degrees. 4. The surgical system according to claim 2, wherein if the angle is below the first value, the first intensity is decreased to the second intensity. 5. The surgical system according to claim 2, wherein if the angle is above the second value, the first intensity is increased to the second intensity. 6. The surgical system according to claim 1, wherein the controller determines if the non-pulsatile component transitions from a higher magnitude to a lower magnitude or from a lower magnitude to a higher magnitude according to a derivative of a non-pulsatile profile for the array of light sensors. 7. The surgical system according to claim 1, wherein the analyzer determines the magnitudes of the non-pulsatile components at the individual light sensors in the row of light sensors, determines a pair of positions corresponding to a fir... | Briteseed, LLC | A61B005/00 A61B005/02 A61B005/107 A61B034/00 A61B018/08 A61B017/00 A61B018/00
US10722385 | B2 | 2020-07-28 | Dynamic ligament balancing system | 2017-01-11 | A method and system for measuring tension, pressure, and distance of knee tissue, the system comprising a prosthetic inlay device comprising at least two platform structures, wherein each of the platforms are supported on a scissor arm structure and a coil spring, a force sensing sensor configured beneath each coil spring, and a connector cable coupled to the force sensing sensors. | 1. A system for measuring tension, pressure and distance of joint tissue, the system comprising: a prosthetic inlay device comprising:at least two platform structures, wherein each of the platforms are supported on a scissor arm structure and a spring, and wherein each of the platforms are configured to variably depress the spring through a range of motion of a joint when the prosthetic inlay device is placed in situ;a force sensing sensor configured beneath each spring, the force sensing sensor detecting pressures from depressing the spring at each of the at least two platform structures; anda communication interface coupled to the force sensing sensors, the communication interface transmitting data corresponding to the detected pressures to a computing device that is configured to generate a set of heights representing displacement of the at least two platform structures and corresponding to medial and lateral tension on the joint tissue based on the detected pressures at a plurality of reference angles of the joint. 2. The system of claim 1 further comprising a sliding surface beneath the scissor arm structures. 3. The system of claim 2 wherein the scissor arm structure includes a fixed leg and a sliding leg. 4. The system of claim 3 wherein the sliding leg is operable to move along the sliding surface. 5. The system of claim 1 wherein the at least two platform structures include indentations configured to conform to a patient's femur. 6. The system of claim 1 wherein the prosthetic inlay device comprises biocompatible material. 7. The system of claim 1, wherein the set of heights comprises a first set of heights representing displacement of the at least two platforms pre-prosthesis and a second set of heights representing displacement of the at least two platforms post-prosthesis. 8. The system of claim 7, the computing device further configured to determine differences between the first and second sets of heights and display an indication for one or more of the plurality of reference angles of the joint. 9. A method, in a data processing system comprising a processor and a memory, for measuring tension, pressure, and distance of knee tissue, the method comprising:receiving, by a computing device, a first set of signals from a ligament balancing device including at least two platform structures, wherein each of the platform structures are supported on a scissor arm structure and a spring, and force sensing sensors that measure force applied to each of the at least two platform structures that variably displace each of the springs through a range of motion of a knee when the ligament balancing device is placed in situ, the first set of signals representative of force applied to each of the at least two platform structures by a femur at one or more knee flexion angles prior to a prosthetic operation;identifying, by the computing device, a first set of medial and lateral forces at the one or more knee flexion angles from the first set of... | - | A61B005/00 A61F002/46 A61B017/02 A61B090/00 A61B034/10 A61B005/11 A61F002/38
US10709336 | B2 | 2020-07-14 | Photoacoustic image generation apparatus and method | 2017-03-13 | After light has been output to a subject to be examined, a photoacoustic wave induced in the subject by the output light is detected. It is assumed that at least one virtual detector element is present outside of a real detector, and dummy data corresponding to the at least one virtual detector element are added to photoacoustic data in which pieces of data of the photoacoustic wave detected by the detector are arranged in accordance with the positions of detector elements. A photoacoustic image is generated by reconstructing the photoacoustic data to which the dummy data have been added by using a Fourier transform method. | 1. A photoacoustic image generation apparatus comprising: an acoustic wave detector that includes a plurality of detector elements arranged in a line, and detects, after light has been output to a subject to be examined, photoacoustic wave induced in the subject by the output light to produce photoacoustic signal;an AD convertor that samples the photoacoustic signal to produce photoacoustic data;a dummy data adder that adds dummy data that corresponds to at least one virtual detector element to the photoacoustic data, the at least one virtual detector element being arranged outside the plurality of the detector elements in a direction of arrangement of the detector elements along with the line of the detector elements by a channel pitch that is the same as a channel pitch of the detector elements, such that pieces of the photoacoustic data and pieces of the dummy data are arranged in accordance with the positions of the detector elements and the positions of the at least one of the virtual detector element respectively, and the photoacoustic data to which the dummy data have been added is two-dimensional data, wherein the dummy data is added to each of both sides of a real data area; andan image generator that generates a two-dimensional photoacoustic image by reconstructing the photoacoustic data to which the dummy data have been added by using a two-dimensional Fourier transform method,the image generator generates a two dimensional photoacoustic image in which a dummy data area corresponding to the at least one virtual detector element is formed outside of an actual data area corresponding to the plurality of detector elements, and wherein the reconstruction by using the Fourier transform method includes performing two-dimensional Fourier transform on the photoacoustic data to which the dummy data have been added and performing two-dimensional inverse Fourier transform on the photoacoustic data on which the two-dimensional Fourier transform has been performed. 2. The photoacoustic image generation apparatus, as defined in claim 1, wherein the reconstruction by using the Fourier transform method includes: performing two-dimensional Fourier transform on the photoacoustic data to which the dummy data have been added,cutting components higher than or equal to a predetermined frequency in a direction of arrangement of the detector elements in the photoacoustic data on which the two-dimensional Fourier transform has been performed, andperforming two-dimensional inverse Fourier transform on the photoacoustic data on which the components has been cut. 3. The photoacoustic image generation apparatus, as defined in claim 1, wherein the two-dimensional inverse Fourier transform is performed by performing one-dimensional inverse Fourier transform in the direction of arrangement of the detector elements after performing one-dimensional inverse Fourier transform in a direction orthogonal to the direction of arrangement of the detector elements. 4. The ph... | FUJIFILM Corporation | A61B005/00
US10716620 | B2 | 2020-07-21 | Expandable balloon mapping and ablation device | 2016-08-03 | The present disclosure relates to the field of tissue mapping and ablation. Specifically, the present disclosure relates to expandable medical devices for identifying and treating local anatomical abnormalities within a body lumen. More specifically, the present disclosure relates to systems and methods of focal treatment for overactive bladders. | 1. A medical device, comprising: a balloon configured to move between a collapsed configuration and an expanded configuration;an expandable-collapsible framework comprising a plurality of flexible splines and struts disposed about an outer surface of the balloon, wherein a distal end of the framework is attached to a distal end of the balloon such that the struts and splines are free-floating about the balloon; anda plurality of electrodes carried about an outer surface of the framework, each of the plurality of electrodes positioned at an intersection of a flexible spline and strut of the plurality of flexible splines and struts. 2. The medical device of claim 1, wherein each of the plurality of electrodes extends outward from the surface of the framework. 3. The medical device of claim 1, further including a spacer disposed between the outer surface of the framework and one of the plurality of electrodes. 4. The medical device of claim 1, wherein each of the plurality of electrodes is electrically coupled to an ablation energy source. 5. The medical device of claim 1, wherein each of the plurality of electrodes is electrically coupled to an electrical activity processing system. 6. The medical device of claim 1, wherein each of the plurality of electrodes includes a sensing electrode, a mapping electrode, a stimulating electrode or an ablation electrode. 7. The medical device of claim 1, wherein a lumen of the balloon is in fluid communication with a fluid source, wherein the balloon moves from the collapsed configuration to the expanded configuration by flowing a fluid from the fluid source into the lumen of the balloon. 8. A medical device, comprising: a balloon configured to move between a collapsed configuration and an expanded configuration;an expandable-collapsible framework comprising a plurality of flexible splines and struts disposed about an outer surface of the balloon, wherein the flexible splines are configured to bend and conform to the shape of the balloon in the expanded configuration, and wherein a distal end of the framework is attached to a distal end of the balloon such that the struts and splines are free-floating about the balloon; anda plurality of electrodes carried about an outer surface of the framework, each of the plurality of electrodes positioned at an intersection of a flexible spline and strut of the plurality of flexible splines and struts. 9. The medical device of claim 8, wherein each of the plurality of electrodes extends outward from the surface of the framework. 10. The medical device of claim 8, further including a spacer disposed between the outer surface of the framework and one of the plurality of electrodes. 11. The medical device of claim 8, wherein each of the plurality of electrodes is electrically coupled to an ablation energy source. 12. The medical device of claim 8, wherein each of the plurality of electrodes is electrically coupled to an electrical activity processing system. 13. Th... | Boston Scientific Scimed, Inc. | A61B018/14 A61B005/00 A61B005/0488 A61B005/20 A61N001/05 A61N001/36 A61B018/00 A61B017/00
US10702189 | B2 | 2020-07-07 | Sensor measuring patient spine vertebra angular orientation | 2015-11-06 | Disclosed is a sensor measuring patient spine vertebra angular orientation, including: a fastener, adapted to be fastened on a specific patient spine vertebra in a unique orientation relative to the specific vertebra, a support, solidary with the fastener in a unique orientation relative to the fastener, a detector, removably secured to the support in a unique orientation relative to the support and adapted to measure one or more parameters representative of the patient spine vertebra angular orientation. | 1. Sensor measuring patient spine (1) vertebra (2) angular orientation, comprising: a fastener (4), adapted to be fastened on a specific patient spine (1) vertebra (2) in a unique orientation relative to said specific vertebra (2),a support (5), solidary with said fastener (4) in a unique orientation relative to said fastener (4),a detector (6), removably secured to said support (5) in a unique orientation relative to said support (5) and adapted to measure one or more parameters representative of said patient spine (1) vertebra (2) angular orientation. 2. Sensor according to claim 1, wherein said fastener (4) is specific to said specific patient spine (1) vertebra (2). 3. Sensor according to claim 2, wherein said fastener (4) shape is complementary to said specific patient spine (1) vertebra (2) shape so that said specific patient spine (1) vertebra (2) can be nested in said fastener (4). 4. Sensor according to claim 3, wherein said fastener (4) shape is complementary to a rear rib (3) shape of said specific patient spine (1) vertebra (2). 5. Sensor according to claim 1, wherein said support (5) is removable from said fastener (4). 6. Sensor according to claim 1, wherein said support (5) extends said fastener (4), perpendicularly to spine (1) axis. 7. Sensor according to claim 1, wherein said support (5) is specific to said specific patient spine (1) vertebra (2). 8. Sensor according to claim 7, wherein said support (5) general shape is generic for all spine (1) vertebrae (2) of all patients and said support (5) orientation is specific to said specific patient spine (1) vertebra (2). 9. Sensor according to claim 1, wherein said support (5) general shape is oriented with a known orientation relative to an upper (8) or a lower (9) average plate of said specific patient spine (1) vertebra (2). 10. Sensor according to claim 9, wherein said support (5) general shape is a plate which is oriented parallel to an upper (8) or a lower (9) average plate of said specific patient spine (1) vertebra (2). 11. Sensor according to claim 1, wherein said detector (6) is generic for all spine (1) vertebrae (2) of all patients. 12. Sensor according to claim 1, wherein said detector (6) can be clipped on said support (5). 13. Sensor according to claim 1, wherein said detector (6) includes one or more gyroscopes, and/or one or more accelerometers, and/or one or more magnetometers, adapted to measure one or more parameters representative of said patient spine (1) vertebra (2) angular orientation. 14. Sensor according to claim 13, wherein said detector (6) also includes a microprocessor to process measured parameters and/or a short range emitter to send out said processed parameters. 15. Sensor according to claim 1, wherein said detector (6) performs real time measurement. 16. Sensor according to claim 1, wherein said fastener (4) and said support (5) orientation have been designed from a 3D image of said patient spine (1). 17. Set of sensors (10) accord... | EOS IMAGING | A61B005/107 A61B005/00 A61F002/46 A61B017/70 A61B005/103 A61F002/44 A61B034/20
US10709382 | B2 | 2020-07-14 | Computing user's physiological state related to physical exercises | 2013-03-04 | A method, apparatus, and computer program for estimating user's physiological state from gait measurements carried out during a physical exercise are disclosed. The physiological state is computed from at least one of step interval variability and stride interval variability acquired from the gait measurements. | 1. A method for determining a physiological state of a user by an apparatus, the method comprising: acquiring, in the apparatus, gait measurement data from at least one motion sensor representing measured gait of the user during a physical exercise;computing, in the apparatus, at least one of step interval variability and stride interval variability from the gait measurement data;determining, in the apparatus standard conditions for testing a level of fatigue;determining, by the apparatus, the user's level of fatigue from said at least one of step interval variability and stride interval variability and using, in the determination of the level of fatigue, exclusively gait measurement data measured under the standard conditions by excluding from the determination of the level of fatigue the gait measurement data that has been measured outside the standard conditions; anddetermining when the standard conditions are satisfied by using the following steps (a) and (b):(a) the standard conditions are defined by heart activity measurement data received from a heart activity sensor and at least one threshold or range for the heart activity measurement data, and the apparatus compares the measured heart activity measurement data with the at least one threshold or range; if the heart activity measurement data falls outside the range defining the standard conditions, the apparatus determines the standard conditions not to have been met; andif the heart activity measurement data falls within the range defining the standard conditions, the apparatus determines the standard conditions to have been met;(b) the standard conditions are defined by running speed measured from the gait measurement data during the exercise with prior knowledge of a step length or a stride length; if the measured running speed falls within a predetermined range defining the standard conditions, the apparatus determines the standard conditions to be fulfilled; and if the measured running speed falls outside the predetermined range defining the standard conditions, the apparatus determines the standard conditions not to be fulfilled. 2. The method of claim 1, wherein the physiological state is determined to be inversely proportional to the step interval variability and stride interval variability such that a higher step interval variability and a higher stride interval variability is associated with a poorer physiological state. 3. The method of claim 1, wherein said determining the user's physiological state comprises determining the user's level of fatigue in training on a basis of the at least one of step interval variability and stride interval variability. 4. The method of claim 1, wherein said determining the user's physiological state comprises making a determination on user's overtraining syndrome on a basis of the at least one of step interval variability and stride interval variability. 5. The method of claim 1, wherein said determining the user's physiological state compr... | Polar Electro Oy | A61B005/00 G06F019/00 A61B005/024 A61B005/11 G01P015/00
US10722221 | B2 | 2020-07-28 | Method and device for statistical tissue sampling using microdevices | 2017-07-31 | The present invention utilizes tetherless microtools to biopsy tissue. The invention provides a device and method for deployment and retrieval of tetherless microtools. The size of the microtools ensures that tissue damage at a site targeted for biopsy is negligible. As such, large numbers of microtools may be deployed ensuring that a true statistical sampling of biologic tissue is performed. | 1. A method of tissue sampling comprising: a) introducing a system into a cavity of a subject, the system comprising: i) a device having an elongated tubular shaft having proximal and distal ends; and an inflatable balloon disposed on the distal end of the shaft, the inflatable balloon having one or more magnets disposed thereon and configured for retrieval or deployment of a plurality of tetherless microtools, wherein the shaft has one or more openings disposed proximal to the balloon for deployment of the plurality of tetherless microtools from a lumen of the shaft; andii) a plurality of tetherless microtools magnetically detachable to the one or more magnets; andb) contacting the plurality of tetherless microtools with a tissue of the subject, each microtool having a first configuration;c) allowing the plurality of tetherless microtools to alter from the first configuration to a second configuration while in contact with the tissue, the second configuration adapted such that each microtool grasps the tissue at a discrete location; andd) retrieving the plurality of tetherless microtools from the subject using the one or more magnets, wherein each microtool retains an individual sample of the tissue upon retrieval, thereby performing tissue sampling. 2. The method of claim 1, wherein the device is configured as a balloon catheter having an elongated shaft of circular cross-sectional shape having proximal and distal ends. 3. The method of claim 1, wherein the plurality of tetherless microtools includes between about 2-1000, 2-500, 2-250, 2-100 or 10-100 microtools. 4. The method of claim 1, further comprising analyzing the individual samples retrieved from the plurality of tetherless microtools. 5. The method of claim 1, wherein a randomized statistical sampling is provided. 6. A method for retrieval of a tetherless microtool from a subject comprising: a) introducing a device into a cavity of the subject, the cavity having a tetherless microtool attached thereto, wherein the device has an elongated tubular shaft having proximal and distal ends; and an inflatable balloon disposed on the distal end of the shaft, the inflatable balloon having one or more magnets disposed thereon and configured for retrieval or deployment of a plurality of tetherless microtools, wherein the shaft has one or more openings disposed proximal to the balloon for deployment of the plurality of tetherless microtools from a lumen of the shaft;b) allowing the one or more magnets of the device to contact the tetherless microtool; andc) retrieving the tetherless microtool from the subject via magnetism by removing the device from the cavity, wherein the microtool retains a tissue sample from the cavity upon retrieval. 7. A method of diagnosing a condition or disease in a subject comprising: a) obtaining a tissue sample from a subject via the method of claim 1;b) analyzing the tissue sample; andc) providing a diagnosis of the condition or disease based on the analysis of t... | The Johns Hopkins University | A61B005/00 A61B010/04 A61B010/02 A61B017/00
US10709377 | B2 | 2020-07-14 | System and methods for monitoring an orthopedic implant and rehabilitation | 2016-03-22 | A system for monitoring a joint of a patient includes multiple sensors to be disposed near a joint and to measure or observe actions or physical quantities associated with the joint; and at least one communications module coupled to the sensors to receive data from the sensors and to transmit sensor information to an external device. In some embodiments, the sensors are implantable near the joint. In other embodiments, the sensors are disposed in a sensor module that is positioned adjacent the skin of the patient near the joint. | 1. A system for monitoring a joint of a patient, the system comprising: a mechanically independent sensor module configured to be positioned on skin of the patient near the joint, the sensor module comprising a single housing, wherein the housing has a first end and a second end opposite the first end, wherein the second end is narrower than the first end indicating to the patient that the second end is to be positioned on the skin of the patient pointed toward the joint;a plurality of sensors disposed within the single housing configured and arranged to measure or observe actions or physical quantities associated with the joint, wherein the plurality of sensors comprises a) a temperature sensor and b) at least one of a first accelerometer, a gyroscope, or a magnetometer; andat least one communications module disposed within the single housing and coupled to the plurality of sensors and configured and arranged to receive data from the sensors and configured and arranged to transmit sensor information to an external device; andeither a) adhesive directly disposed on the single housing or b) an adhesive pad coupled, or coupleable, to the single housing, wherein the adhesive or adhesive pad is configured adhere the sensor module to the skin of the patient. 2. The system of claim 1, further comprising a second sensor module that is mechanically independent from the sensor module, wherein the second sensor module comprises a single second housing;a plurality of second sensors disposed within the single second housing and configured and arranged to measure or observe actions or physical quantities associated with the joint; andat least one second communications module disposed within the single second housing and coupled to the plurality of second sensors and configured and arranged to receive data from the second sensors and configured and arranged to transmit sensor information to the external device. 3. The system of claim 2, wherein the plurality of sensors comprises a first proximity sensor and the second plurality of sensors comprises a second proximity sensor, wherein the first and second proximity sensors, in combination, are configured and arranged to provide sensor data that can be processed to determine a number of repetitions of joint movement. 4. The system of claim 1, wherein the system comprises the adhesive pad and the adhesive pad comprises at least one first locating feature and the housing comprises at least one second locating feature to facilitate alignment of the adhesive pad with the housing. 5. The system of claim 4, wherein the at least one second locating feature has a shape complementary to a shape of the at least one first locating feature. 6. The system of claim 1, wherein the at least one communications module comprises a sensor processor configured and arranged to receive the data from the sensors and a wireless communications unit configured and arranged to wirelessly transmit the sensor information to the external ... | CONSENSUS ORTHOPEDICS, INC. | A61B005/11 A61B005/00 G01C021/12 A61B001/00 A61B001/04 A61B005/107 G01C022/00 G01C009/00 A61B005/07 A61B034/20 A61B005/01 A61B005/024 A61B090/30
US10709358 | B2 | 2020-07-14 | Sizing device, and methods of manufacture and use thereof | 2016-06-24 | The present invention relates to a sizing device and includes a sizing frame and frame handle for the evaluation of the fitting of a therapy device to a patient. The sizing frame corresponds to the contact region of a therapy device and the inner area of the sizing frame corresponds to the target therapy area. The sizing device is designed to provide a visual indication of correct or incorrect fit of the sizing device and the subsequent corresponding therapy device. | 1. A sizing frame for determining the fit of a therapy device, the therapy device configured to mate to an individual at a position corresponding approximately to the anterior triangle of the neck of the individual to create an enclosed space for the administration of negative pressure upon the external neck surface of the individual, the sizing frame comprising: a mating surface configured to engage with a surface on the individual corresponding to the contact surface of the therapy device, the mating surface configured to permit visualization of the individual's skin surface from a position that will be internal to the enclosed space of the desired therapy device. 2. A sizing frame according to claim 1, wherein the sizing frame comprises (i) a first curved element configured to approximately conform to an individual from approximately the gonion on one side of the mandibular body, across the mental protuberance, and to the opposite gonion of the mandibular body, (ii) a second curved element configured to approximately conform to an individual from approximately the gonion on one side of the mandibular body to a position on the neck at the level of the thyroid cartilage, and to the opposite gonion of the mandibular body, and (iii) a handle spanning the region between the first curved element and the second curved element. 3. A sizing frame of claim 2, wherein the sizing frame contains one or more structural elements configured to receive a predetermined anatomical feature of the patient to ensure proper positioning of the sizing frame element on the individual. 4. A sizing frame of claim 3, wherein the structural element(s) comprise a recess to receive the chin of the individual. 5. A sizing frame according to claim 4 wherein the sizing frame further contains a tab extending outward from an outer surface of the sizing frame corresponding to the interface between the first and second curved elements. 6. A sizing frame according to claim 5, wherein a surface of the handle is configured mimic the position of an internal surface of a therapy device. 7. A sizing frame according to claim 1, wherein an element of the frame indicates a corresponding therapy device. 8. A sizing frame according to claim 7, wherein the element is a color. 9. A sizing frame according to claim 7, wherein the element is alphanumeric. 10. A sizing frame according to claim 7, wherein the element is an encoded label. 11. A sizing frame according to claim 1, wherein the at least a portion of the frame material is sufficiently translucent to observe the facial features of the individual through the frame. 12. A sizing frame according to claim 11, wherein the sizing frame contains a sizing grid, and wherein, when the device is applied to the individual, the facial features and applied markings can be observed through the grid. 13. A system for selection of an appropriately sized therapy device for positioning on the external anatomy of the individual, comprising a plura... | SOMMETRICS, INC. | A61B005/107 A61F005/055 A61F005/56 A61B005/00 A61H009/00 B29C049/06 A61B090/00
US10716528 | B2 | 2020-07-21 | Automatic display of previously-acquired endoluminal images | 2011-09-08 | Apparatus and methods are provided for use with an endoluminal data-acquisition device that acquires a set of endoluminal data-points of a lumen of a subject's body at respective locations inside the lumen, a second endoluminal device, and a display configured to display images. At least one processor includes location-association functionality that associates a given data point acquired by the endoluminal data-acquisition device with a given location within the lumen. Location-determination functionality determines, by means of image processing, in an extraluminal image of the second endoluminal device, a current location of at least a portion of the second endoluminal device. Display-driving functionality drives the display to display an indication of the endoluminal data point associated with the location, in response to determining that the portion of the second device is currently at the location. Other applications are also described. | 1. A system for guiding an endoluminal therapeutic procedure, comprising: a display; andat least one computer processor in communication with the display and configured to: receive a plurality of endoluminal images obtained at a first time by an endoluminal imaging catheter at a plurality of locations, respectively, within a lumen of a subject;associate the plurality of endoluminal images with the plurality of locations within the lumen;receive a plurality of extraluminal images obtained at a later second time by an extraluminal imaging device, wherein the plurality of extraluminal images includes a representation of an endoluminal therapeutic device within the lumen;determine at the later second time, with image processing of the plurality of extraluminal images obtained at the later second time, a current location of the endoluminal therapeutic device within the lumen;associate, at the later second time, the current location of the endoluminal therapeutic device with a location of the plurality of locations;identify at the later second time, from the plurality of endoluminal images obtained at the first time, an endoluminal image corresponding to the location; andcause the display to simultaneously display: at least one extraluminal image of the plurality of extraluminal images, wherein the at least one extraluminal image includes the representation of the endoluminal therapeutic device; andthe identified endoluminal image corresponding to the location, wherein the endoluminal image is automatically updated based on the current location of the endoluminal therapeutic device. 2. The system of claim 1, wherein the at least one computer processor is configured to determine the current location of the endoluminal therapeutic device by determining a location of a radiopaque portion of the endoluminal therapeutic device in the at least one extraluminal image. 3. The system of claim 1, further comprising the endoluminal imaging catheter. 4. The system of claim 1, further comprising the endoluminal therapeutic device. 5. The system of claim 1, further comprising an endoluminal data-acquisition device in communication with the at least one computer processor and configured to acquire a set of endoluminal data points along the lumen. 6. The system of claim 5, wherein the at least one computer processor is configured to cause the display to display endoluminal data points acquired by the endoluminal data-acquisition device that are associated with the current location of the endoluminal therapeutic device. 7. The system of claim 1, wherein the endoluminal imaging catheter comprises an intravascular ultrasound (IVUS) imaging catheter. 8. The system of claim 1, wherein the at least one computer processor is configured to cause the display to display an indication of the current location of the endoluminal therapeutic device within the lumen with respect to an endoluminal image stack of the lumen generated using the endoluminal images acquired by the ... | SYNC-RX, LTD. | A61B006/00 G06F003/14 G06T001/00 G09G005/36 G06T007/00 G06T019/20 A61B001/00 A61B006/12 A61F002/82 A61B017/22 A61B017/00 A61F002/958 A61B008/08 A61B005/00 A61B090/00
US10716503 | B2 | 2020-07-21 | Pain level detection and characterization using capacitive sensors | 2018-09-04 | Systems, methods, sensors, and software for providing enhanced measurement and detection of patient pain response are provided herein. In a first example, a measurement system is provided that includes a capacitive system configured to measure a capacitance signal of tissue of the patient using a capacitive sensor applied to the tissue of the patient. The measurement system also includes a patient monitor configured to measure an electrical signal representing brain activity of the patient using a brain activity sensor applied to the tissue of the patient. The measurement system also includes a processing system configured to determine pain metrics based at least on the capacitance signal and the electrical signal, and determine a pain response of the patient based at least on the pain metrics and pain calibration information for the patient. | 1. A measurement system to detect pain response in a patient, the measurement system comprising: one or more patient sensors, applied to the tissue of the patient;a capacitive system configured to measure a capacitance signal for the tissue of the patient, the capacitance signal being a measure of capacitance changing over time, using at least one of the one or more patient sensors;a patient monitor configured to measure an electrical signal representing brain activity of the patient, using at least one of the one or more patient sensors; anda processing system configured to correlate the capacitance signal and the electrical signal to determine pain metrics, and determine a pain response of the patient based at least on the pain metrics and pain calibration information for the patient. 2. The measurement system of claim 1, wherein the capacitive system is configured to measure the capacitance signal with one or more capacitor plates in the one or more patient sensors. 3. The measurement system of claim 1, wherein the one or more patient sensors are applied to the tissue of the patient. 4. The measurement system of claim 1, wherein the processing system is configured to determine the pain response by at least correlating the pain metrics to further pain metrics identified for the patient that are derived from at least one of a respiration rate, heart rate, electroencephalogram (EEG) data, and photoplethysmogram (PPG) data. 5. The measurement system of claim 1, wherein the pain calibration information comprises capacitance measurements of the patient during one or more calibration stimuli. 6. A measurement system to detect pain response in a patient, the measurement system comprising: one or more patient sensors, applied to the tissue of the patient;a capacitive system configured to measure a capacitance signal for the tissue of the patient, the capacitance signal being a measure of capacitance changing over time, using at least one of the one or more patient sensors;a patient monitor configured to measure an electrical signal representing brain activity of the patient, using at least one of the one or more patient sensors; anda processing system configured to determine pain metrics at least in part from a time segment of the capacitance signal and the electrical signal, wherein the pain metrics are determined based on direct current (DC) trends in ones of the capacitance signal and the electrical signal, and determine a pain response of the patient based at least on the pain metrics and pain calibration information for the patient. 7. A measurement system to detect pain response in a patient, the measurement system comprising: one or more patient sensors, applied to the tissue of the patient;a capacitive system configured to measure a capacitance signal for the tissue of the patient, the capacitance signal being a measure of capacitance changing over time, using at least one of the one or more patient sensors;a patient monitor configured to... | Covidien LP | A61B005/04 A61B005/00 A61B005/0476 A61B005/053 A61B005/1455
US10702146 | B2 | 2020-07-07 | Angular multiplexed optical coherence tomography systems and methods | 2018-08-27 | Angle multiplexed optical coherence tomography systems and methods can be used to evaluate ocular tissue and other anatomical structures or features of a patient. The angle multiplexed optical coherence tomography system includes a light source, an optical assembly for obtaining a plurality of sample beams corresponding to respective anatomical locations of the eye of the patient, where individual sample beams are associated with a respective non-zero angle relative to a reference beam, and a detection mechanism that detects individual unique interference patterns respectively provided by the plurality of sample beams, for simultaneous evaluation of the anatomical locations. | 1. An optical coherence tomography system, comprising: a detector comprising an array of pixels;a light source configured to supply a sample light beam to an eye and to further supply a reference light beam;an optical element configured to direct the reference light beam to the array of pixels; andan optical splitter configured to divide light returning from the eye into a plurality of individual signal beams each having a corresponding different path length from an output of the optical splitter to the optical element and each corresponding to a respective different anatomical location of the eye,wherein each of the individual signal beams impinges on the optical element at a corresponding different angle with respect to the reference light beam, andwherein the optical element is configured to direct the plurality of individual signal beams simultaneously to the array of pixels. 2. The optical coherence tomography system of claim 1, wherein the detector is configured to detect individual unique interference patterns respectively provided by the plurality of individual signal beams and the reference beam, for simultaneous evaluation of the different anatomical locations of the eye. 3. The optical coherence tomography system of claim 2, wherein unique interference spatial periods of the unique interference patterns are adjustable in response to changes in respective individual signal beam angles relative to the reference beam. 4. The optical coherence tomography system of claim 1, wherein the optical element comprises one or more collimation lenses that direct combined signal-reference beam pairs toward the detector. 5. The optical coherence tomography system of claim 1, wherein the system provide an accuracy for range finding on an order of 10 microns. 6. The optical coherence tomography system of claim 1, further comprising a filter assembly that transmits interference signals at spatial frequencies about a first signal-reference beam pair and suppresses interference signals at spatial frequencies about a second signal-reference beam pair. 7. A method of optical coherence tomography, comprising: supplying a sample light beam from a light source to an eye, and further supplying a reference light beam from the light source;an optical element directing the reference light beam to a detector comprising an array of pixels; andan optical splitter dividing light returning from the eye into a plurality of individual signal beams each having a corresponding different path length from an output of the optical splitter to the optical element, and each corresponding to a respective different anatomical location of the eye,wherein each of the individual signal beams impinges on the optical element at a corresponding different angle with respect to the reference light beam, andwherein the optical element directs the plurality of individual signal beams simultaneously to the array of pixels. 8. The method of claim 7, further comprising detecting, by the... | AMO Development, LLC | A61B003/14 A61B003/00 A61B003/10 A61B003/107 A61B005/00 G01N021/47 G01B009/02
US10709838 | B2 | 2020-07-14 | Pumping fluid delivery systems and methods using force application assembly | 2011-04-19 | A method of dispensing a therapeutic fluid from a line includes providing an inlet line connectable to an upstream fluid source. The inlet line is in downstream fluid communication with a pumping chamber. The pumping chamber has a pump outlet. The method also includes actuating a force application assembly so as to restrict retrograde flow of fluid through the inlet while pressurizing the pumping chamber to urge flow through the pump outlet. A corresponding system employs the method. | 1. A wearable insulin pump assembly system for coupling to a reservoir that contains insulin, the wearable insulin pump assembly system comprising: a wearable insulin pump housing portion that contains a reservoir, a pump chamber, an outlet end on one side of the pump chamber and a receptacle on a second side of the pump chamber, the receptacle having a size and shape for receiving one end of the reservoir, the receptacle being in fluid flow communication with the pump chamber;a rotary pump blade supported for rotation on the pump chamber, the rotary pump blade comprising a grouping of coupled projections;a shape memory actuator configured to apply a rotating force to the grouping of coupled projectionsa flow communicator located within the receptacle for providing a flow path for communication of the insulin from a reservoir to the pump chamber when one end of the reservoir is received within the receptacle; anda pump housing septum on the outlet end of the pump housing portion for allowing a hollow needle to extend through the septum structure to provide a fluid flow path out from the outlet end of the pump housing portion;wherein the rotary pump blade causes insulin to be urged from the pump chamber, andwherein the wearable insulin pump is for delivering insulin from the reservoir to a user through the outlet end of the pump housing portion, further comprising a shape memory actuator configured to apply a rotating force to the grouping of coupled projections. 2. A wearable insulin pump assembly system according to claim 1, wherein the hollow needle further comprises a head with a surface on which a force may be applied to move the hollow needle through the pump housing septum. 3. A wearable insulin pump assembly system according to claim 1, further comprising a volume sensor assembly further comprising an outlet chamber in fluid flow communication with the pump chamber, for receiving fluid from the pump chamber. 4. A wearable insulin pump assembly system according to claim 1, further comprising a volume sensor assembly further comprising an outlet chamber in fluid flow communication with the pump chamber and having an opening for conveying the insulin out of the pump housing portion. 5. A wearable insulin pump assembly system as recited in claim 1, wherein the pump housing portion comprises a single unitary body having the receptacle located on one end of the single unitary body and the outlet end and septum structure located on an opposite end of the single unitary body relative to the one end at which the receptacle is located. 6. A wearable insulin pump assembly system as recited in claim 1, further comprising: a first housing portion adapted to be secured to a user;a second housing portion configured to be selectively engaged with and disengaged from the first housing portion;wherein the reservoir for containing the insulin supported on the first housing portion, wherein the reservoir comprises a container and a reservoir septum at on... | DEKA Products Limited Partnership | F04B043/02 F04B043/09 F04B043/12 A61M005/142 A61M005/168 A61M005/145 G05D007/06 A61M005/172 G08C017/02 H04B007/26 G05B023/02 A61M005/158 A61B005/00 A61M005/50 G01F022/00 A61M005/14 A61M005/162 A61M005/36 A61B005/155 A61J001/20
US10702696 | B2 | 2020-07-07 | Systems and methods for spinal cord stimulation | 2017-11-08 | Delivering stimulation includes delivering temporal patterns of stimulation pulses to respective transducers of an array of transducers, wherein the delivery of the pattern to a particular transducer of the array is different from at least some of the deliveries of the patterns to the other transducers of the array at least according to a time delay. The patterns delivered may include regular temporal patterns each having a respective constant inter-pulse interval. The constant inter-pulse intervals may be about the same. The patterns may be staggered. The transducers may deliver electrical, optical, acoustic, thermal or magnetic stimulation. | 1. A system for delivering electrical pulse stimulation, the system comprising: a control device configured to provide programming of intermittent, temporal pulse patterns; anda stimulation device comprising a control module, multiple electrodes in an array in electrical communication with the control module, and an input device intermittent communication configured to communicate with the control device to receive the programming, the stimulation device configured to deliver respective asynchronous temporal patterns of electrical stimulation pulses to the multiple electrodes according to the programming,wherein each respective asynchronous temporal pattern of electrical stimulation pulses activates a different sub-population of targeted neurological tissue of a subject, andwherein the control module is configured to deliver the asynchronous temporal patterns by offsetting delivery of some pulse patterns relative to other pulse patterns. 2. The system for delivering electrical pulse stimulation according to claim 1, wherein the programming comprises instructions such that at least some of the patterns delivered comprise regular temporal patterns each having a respective constant inter-pulse interval. 3. The system for delivering electrical pulse stimulation according to claim 2, wherein the constant inter-pulse intervals are about the same. 4. The system for delivering electrical pulse stimulation according to claim 3, wherein offsetting delivery of some pulse patterns relative to other pulse patterns includes offsetting delivery by a time delay, wherein the time delay is approximately a percentage of the constant inter-pulse intervals. 5. The system for delivering electrical pulse stimulation according to claim 3, wherein offsetting delivery of some pulse patterns relative to other pulse patterns includes offsetting delivery by a time delay, wherein the temporal pulse pattern programming comprises instructions such that at least some of the patterns delivered differ from others of the patterns delivered only by the time delay such that the patterns are staggered. 6. The system for delivering electrical pulse stimulation according to claim 3, wherein offsetting delivery of some pulse patterns relative to other pulse patterns includes offsetting delivery by a time delay, wherein the programming comprises instructions such that some of the patterns delivered differ from others of the patterns according to time delays that vary. 7. The system for delivering electrical pulse stimulation according to claim 1, wherein the programming comprises instructions such that at least some of the patterns delivered comprise non-regular temporal patterns each having respective varying inter-pulse intervals. 8. The system for delivering electrical pulse stimulation according to claim 7, wherein offsetting delivery of some pulse patterns relative to other pulse patterns includes offsetting delivery by a time delay, wherein the programming comprises instructi... | Duke University | A61N001/36 A61N001/05 A61N001/372 A61B005/00
US10722133 | B2 | 2020-07-28 | QT interval determination methods and related devices | 2018-05-09 | Described herein is a system and method of automatically monitoring QT intervals in a patient based on one or more EKG signals received from attached monitoring devices. Each EKG signal is analyzed to detect attributes of the first and second EKG signals, including QRS onset information, QRS peak information, and T-wave offset information. A QT interval is calculated based on QRS onset information derived from the first EKG signal and T-wave offset information derived from the second EKG signal. The calculated QT interval is compared to thresholds to detect elongation of the QT interval and an alert is generated in response to a detected elongated QT interval. | 1. A monitoring system for measuring QT intervals, the monitoring system comprising: a monitoring device adhered to a patient at a location and configured to monitor an electrocardiogram (EKG) signal associated with the patient; anda QT interval measurement module configured to: receive the EKG signal from the monitoring device;identify QRS complex attributes and heart rate attributes for each beat within the received EKG signal, wherein the QRS complex attributes include a peak of a QRS complex (iRpk) and a QRS onset (iRon);generate a T-wave offset estimate (iToff_estimate) based at least in part on the identified heart rate attributes;identify a T-wave offset detection window based on the identified peak of the QRS complex (iRpk) and the generated T-wave offset estimate (iToff_estimate);locate a T-wave offset (iToff) within the identified T-wave offset detection window based on one or more detected maximums or minimums within the T-wave offset detection window; andmeasure a QT interval based on the identified QRS onset (iRon) and the located T-wave offset (iToff). 2. The monitoring system of claim 1, wherein the QT interval measurement module is further configured to generate an alert in response to detecting a long QT interval. 3. The monitoring system of claim 2, wherein the QT interval measurement module is further configured to provide the generated alert and the received EKG signal to a remote monitoring center. 4. The monitoring system of claim 1, wherein the QT interval measurement module is further configured to: identify beats within the received EKG signal as noisy or non-noisy; andidentify a beat segment as non-noisy in response to identifying a consecutive number of beats as non-noisy,wherein to measure the QT interval, the QT interval measurement module is configured to select only non-noisy beat segments. 5. The monitoring system of claim 1, wherein the QT interval measurement module is further configured to select a beat segment to measure the QT interval based on one or more of a time of day, an activity level of the patient, an orientation of the patient, or a heart rate of the patient. 6. The monitoring system of claim 1, wherein the QT interval measurement module is further configured to discards beats in which the QT interval measurement module cannot discern at least one of the QRS onset (iRon) or the T-wave offset (iToff) within the EKG signal even if the beat is non-noisy. 7. The monitoring system of claim 1, wherein the T-wave offset detection window starts from the generated T-wave offset estimate (iToff_estimate) and extends by a threshold value. 8. The monitoring system of claim 1, wherein to generate the T-wave offset estimate (iToff_estimate) based at least in part on the identified heart rate attributes, the QT interval measurement module is configured to use the identified heart rate attributes and a look-up table to estimate the T-wave offset detection window. 9. The monitoring system of claim 1, wherein ... | MEDTRONIC MONITORING, INC. | A61B005/05 A61B005/0472 A61B005/00 A61B005/04 A61B005/0245 A61B005/0452 G16H050/30
US10709390 | B2 | 2020-07-14 | Deep learning algorithms for heartbeats detection | 2017-11-21 | A system that detects heartbeats includes a sensor or a transducer and algorithms based on deep learning. The algorithms employ techniques of artificial intelligence that enable the system to extract heartbeat features under low signal-to-noise-ratio (SNR) conditions when a user is exercising. The algorithms can be applied to various technologies for heart rate monitoring such as ultrasound Doppler, photoplethysmogram (PPG), electrocardiogram (EKG), acoustic, pressure/force sensing and laser/RF Doppler, among other types of sensing methods. | 1. A method for determining a heart rate of a subject, comprising the steps of: obtaining temporal sensor data generated by a sensor worn by the subject, the temporal sensor data represented by a plurality of frames of data samples;generating, for each frame of the plurality of frames, a feature vector based on a frequency domain representation of the corresponding data samples of the frame;inputting the feature vectors to a scenario classifier to determine a scenario of a plurality of scenarios for the temporal sensor data, wherein each of the plurality of scenarios corresponds to a different signal-to-noise level;identifying a model from a plurality of trained models according to the determined scenario;inputting the feature vectors to the model to generate a series of heartbeat determinations of the subject for each frame of the plurality of frames, wherein the model comprises a first set of parameters and associated weights that represent likelihoods of heartbeats of the subject as a function of time; andproviding the heart rate of the subject determined based on the series of heartbeat determinations. 2. The method of claim 1, wherein each of the signal-to-noise levels of the plurality of scenarios correspond to a different level of athletic activity of the subject. 3. The method of claim 1, wherein the plurality of scenarios includes at least a first scenario and a second scenario, the first scenario associated with a first set of one or more types of athletic activity and corresponding with a first average signal-to-noise ratio of temporal sensor data, the second scenario associated with a second set of one or more types of athletic activity and corresponding with a second average signal-to-noise ratio of temporal sensor data greater than the first average signal-to-noise ratio. 4. The method of claim 3, wherein the first set of one or more types of athletic activity includes at least one of jogging and running, and wherein the second set of one or more types of athletic activity includes walking. 5. The method of claim 1, wherein the scenario classifier comprises a second set of parameters and associated weights that represent likelihoods of the temporal sensor data corresponding to each of the plurality of scenarios. 6. The method of claim 5, wherein each of the plurality of scenarios is associated with one of a plurality of models including the model, wherein each of the plurality of models comprises a corresponding set of parameters and associated weights that represent likelihoods of heartbeats of the subject as a function of time, and wherein the method further comprises: determining the heart rate of the subject by applying the set of weights to outputs of the plurality of models. 7. The method of claim 1, wherein the first set of parameters and associated weights represent likelihoods of heartbeats for each heartbeat stage of a plurality of heartbeat stages, and wherein the model generates the series of heartbeat determinatio... | LOGOS CARE, INC. | A61B005/00 A61B005/0404 A61B005/0245 A61B005/04 A61B005/024 A61B008/02 A61B005/0255 A61B008/08 A61B008/06
US10702340 | B2 | 2020-07-07 | Image processing and patient-specific modeling of blood flow | 2016-06-15 | Embodiments include a system for determining cardiovascular information for a patient. The system may include at least one computer system configured to receive patient-specific data regarding a geometry of the patient's heart, and create a three-dimensional model representing at least a portion of the patient's heart based on the patient-specific data. The at least one computer system may be further configured to create a physics-based model relating to a blood flow characteristic of the patient's heart and determine a fractional flow reserve within the patient's heart based on the three-dimensional model and the physics-based model. | 1. A method for non-invasive assessment of an arterial stenosis, comprising: receiving medical imaging data pertaining to a patient;segmenting a plurality of mesh candidates for an anatomical model of an artery including a stenosis region of the patient from the received medical imaging data;computing a hemodynamic index for the stenosis region in each of the plurality of mesh candidates, wherein computing the hemodynamic index comprises simulating blood flow and pressure in each of the plurality of mesh candidates for the artery of the patient, and computing a fractional flow reserve value for the stenosis region in each of the plurality of mesh candidates based on the blood flow and pressure simulations; anddetermining whether a variation among values of the hemodynamic index for the stenosis region in each of the plurality of mesh candidates is significant with respect to a threshold associated with a clinical decision regarding the stenosis region. 2. The method as recited in claim 1, wherein segmenting a plurality of mesh candidates for an anatomical model of an artery including a stenosis region of a patient from medical imaging data comprises: segmenting the artery from the medical imaging data to generate the anatomical model;determining one or more candidate locations for each vertex among a plurality of vertices of the anatomical model; andgenerating the plurality of mesh candidates based on the one or more candidate locations. 3. The method as recited in claim 2, wherein generating the plurality of mesh candidates based on the one or more candidate locations comprises: generating the plurality of mesh candidates by enforcing a plurality of connection rules for connecting the plurality of candidate locations of neighboring vertices of the anatomical model, wherein each of the plurality of connection rules results in a corresponding mesh candidate. 4. The method as recited in claim 3, wherein generating the plurality of mesh candidates based on the one or more candidate locations further comprises: projecting each of the plurality of mesh candidates onto a learned shape space of the artery using an active shape model. 5. The method as recited in claim 1, wherein determining whether a variation among values of the hemodynamic index for the stenosis region in each of the plurality of mesh candidates is significant with respect to a threshold associated with a clinical decision regarding the stenosis region comprises: determining whether the variation among values of the hemodynamic index for the stenosis region in each of the plurality of mesh candidates is entirely below the threshold for the clinical decision or is entirely above the threshold for the clinical decision. 6. The method as recited in claim 1, further comprising: in response to determining that the variation among values is not significant, displaying results of the hemodynamic index without receiving user input. 7. The method as recited in claim 1, further comprising:... | HeartFlow, Inc. | A61B034/10 G06T007/11 G06T007/12 G06T007/13 G06T007/149 G06T007/62 A61B005/02 G16H050/50 A61B005/026 A61B005/00 G06F017/10 A61B034/00 G16B005/00 G16B045/00 G06T007/70 G06T007/73 A61B006/00 G16H010/60 G16H050/30 G06G007/60 A61B005/029 A61B006/03 A61B008/06 A61B008/08 G01R033/563 A61B005/107 G06T007/00 G06T017/00 G06K009/00 G06K009/62 A61B005/021 G06F017/50 A61B005/024 A61B005/22 G06F019/00 G06K009/46 G06T007/20 G06T011/00 G06T011/20 A61M005/00 G06K009/52 A61B005/055 A61B005/11 G06T017/20 G06T015/10 G06T011/60 G06T007/60 A61B008/02 A61B008/04 G01R033/56 G16H030/20 G16H030/40 G16H010/40 G16H050/70 A61B090/00
US10717751 | B2 | 2020-07-21 | Near-IR glucose sensors | 2017-12-27 | The present disclosure provides polymerizable luminescent dyes useful for incorporation into polymers. The dyes and the polymers can be used in sensing and imaging applications, for example, to provide accurate and optionally long term measurements of glucose in vivo. The present disclosure also provides sensors including the polymers described herein. The sensors can be implanted into a tissue of a subject and used for long-term or short-term continuous and semi-continuous collection of data of various biochemical analytes, optionally without the use of implantable hardware of any type and/or enzymatic and electrochemical detection methods. | 1. A compound of Formula I:or a tautomer or a salt thereof,wherein the dotted lines denote a bond or absence of a bond;when the dotted line connecting R1 and O is a bond, R1 is CX1X2 and R15 is absent; andwhen the dotted line connecting R1 and 0 is absence of a bond, R15, at each occurrence, is independently H or C1-C6 alkyl, and R1, at each occurrence, is H, optionally substituted C1-C6 alkyl, optionally substituted C2-C6 alkenyl, optionally substituted C2-C6 alkynyl, optionally substituted C2-C10 heteroalkyl, polymerizable moiety, an NIR dye moiety, an electron-withdrawing group, or an electron-donating group;X1 and X2 are independently H or C1-C6 alkyl;R2 is H or C1-C6 alkyl;Z is a C6-C14 arylene optionally substituted with R11, R12, R14, and L2R13;R3, R4, R5, R6, R7, R8, R11, R12, R13 and R14 are independently H, optionally substituted C1-C6 alkyl, optionally substituted C2-C6 alkenyl, optionally substituted C2-C6 alkynyl, optionally substituted C2-C10 heteroalkyl, polymerizable moiety, NIR dye moiety, electron-withdrawing group, or electron-donating group;R9 and R10 are independently H, C1-C6 alkyl, polymerizable moiety, or NIR dye moiety;L1, L2, and L3 are independently a bond or a linker group; andwherein at least one R1, R3, R4, R5, R6, R7, R8, R9, R10, R11, R12, R13, or R14 is an NIR dye moiety and at least one R1, R3, R4, R5, R6, R7, R8, R9, R10, R11, R12, R13, or R14 is a polymerizable moiety. 2. The compound of claim 1, wherein the compound has a structure of Formula IA, IB, or IC:(IA), (IB)(IC), or a tautomer or a salt thereof. 3. The compound of claim 1, wherein the electron-withdrawing group is selected from the group consisting of halogen, C(O)R′, COOR′, C(O)NH2, C(O)NR′R″, CF3, CN, SO3H, SO2CF3, SO2R′, SO2NR′R″, ammonium, alkyl ammonium, and NO2, wherein R′ and R″ are independently H or C1-C6 alkyl. 4. The compound of claim 1, wherein electron-donating group is selected from the group consisting of NRN1RN2, OR′, NHC(O)R′, OC(O)R′, phenyl, and vinyl, wherein RN1, RN2, and R′ are independently H or C1-C6 alkyl. 5. The compound of claim 1, wherein L1, L2, or L3 are, independently, a bond or a linker group selected from optionally substituted amino, optionally substituted amido, —O—, optionally substituted —CH2C6H4O—, C2-C20 PEG linker, optionally substituted C6-C10 arylene, optionally substituted five to ten member heteroarylene, optionally substituted —C1-C6 alkylene-Ar—, optionally substituted —C2-C6 alkenylene-Ar—, optionally substituted —C2-C6 alkynylene-Ar—, optionally substituted —C(O)NH—C1-C6 alkylene-Ar—, optionally substituted C1-C6 alkylene-C(O)NH—C1-C6 alkylene-Ar—, —(CH2CH2O)n—, —(CH2CH2O)p CH2CH2—, optionally substituted C1-C10 alkyl, optionally substituted C2-C10 alkenyl, optionally substituted C2-C10 alkynyl, and optionally substituted C2-C20 heteroalkyl, wherein n is an integer from 1 to 10, p is 2, 3, 4, or 5, and Ar is C6-C10 arylene or five to ten member heteroarylene. 6. The compound of claim 1, wherein L1, ... | Profusa, Inc. | C07F005/02 C09B057/00 C09B001/00 C09B023/00 C09B023/08 C09B069/00 C09B011/28 C09B023/04 C09B057/02 C09B011/24 A61B005/1459 C09K011/06 C09B023/01 A61B005/145
US10709867 | B2 | 2020-07-14 | Gaze-contingent music reward therapy | 2017-11-09 | A computerized method for treating a psychopathological condition, the method comprising using at least one hardware processor for presenting to a patient on a computer display at least one visual media and at least audio media; receiving from an eye tracker a gaze direction of at least one eye of the patient; computing a display location of the gaze direction; and when a distance from the gaze location to a target location of the visual media is greater than a threshold value, pausing the audio media for a period of time. | 1. A computerized method for treating a psychopathological condition, the method comprising using at least one hardware processor for: presenting, using an audio playback device, at least audio media to a patient;receiving, from an eye tracker, a gaze direction of at least one eye of the patient;computing a difference angle between a target direction and the gaze direction; andpausing the audio media for a period of time, based, at least in part, on said computing,wherein subsequent to the pausing, the audio media is at least in part rewound. 2. The computerized method of claim 1, wherein the presenting further comprises presenting a visual media to the patient on a display. 3. The computerized method of claim 2, wherein the visual media is at least one from a video, a digital video file, and at least one digital image. 4. The computerized method of claim 2, wherein the visual media comprises images of faces. 5. The computerized method of claim 4, wherein at least some of the images of faces depict neutral expressions, and at least some of the images of faces are threatening expressions, and the psychopathological condition is a social anxiety disorder. 6. The computerized method of claim 1, wherein the psychopathological condition is at least one of a social anxiety disorder, an anxiety disorder, a mood disorder, a post-traumatic stress disorder, an eating disorders, a borderline personality disorder, an autism spectrum disorder, a substance-related disorder, psychotic disorder, and an addictive disorder. 7. The computerized method of claim 1, wherein the audio media is at least one of a music track, a digital music file, a lecture track, a digital lecture file, an audio book, and a digital audio book. 8. The computerized method of claim 1, wherein the period of time is between 0.1 second and 10 seconds. 9. A computerized system for treating a psychopathological condition, the computerized system comprising: an eye tracker;at least one hardware processor; anda non-transitory computer-readable storage medium having program code embodied therewith, the program code executable by said at least one hardware processor to:(i) present to a patient at least audio media;(ii) receive, from the eye tracker, a gaze direction of at least one eye of the patient;(iii) compute a difference angle between a target direction and the gaze direction; and(iv) pause the audio media for a period of time based, at least in part, on said computing,wherein subsequent to the pausing, the audio media is at least in part rewound. 10. The computerized system of claim 9, further comprising a display for presenting images to said patient, and wherein the program code is further configured to present to said patient, on the display, at least one visual media. 11. The computerized system of claim 10, wherein the visual media is at least one from a video, a digital video file, and at least one digital image. 12. The computerized system of claim 10, wherein the visual me... | RAMOT AT TEL-AVIV UNIVERSITY LTD. | A61M021/00 A61B005/16
US10722135 | B2 | 2020-07-28 | Myogram determination from ECG signal | 2017-03-03 | Systems and methods for measuring signals representative of muscle activity are provided. One method includes detecting an ECG signal through a plurality of electrodes. The ECG signal includes a plurality of ECG sample signals, and each ECG sample signal is a bipolar signal associated with two of the plurality of electrodes and includes a cardiac signal component and a myographic signal component. The method further includes filtering each of the ECG sample signals to remove at least a portion of the cardiac signal component and generate a combined myographic power signal for the two of the plurality of electrodes with which the ECG sample signal is associated. Each combined myographic power signal represents a myographic potential between the two electrodes. The method further includes calculating individual myographic power signals for each of the plurality of electrodes by applying the combined myographic power signals within a covariance matrix. | 1. A method of measuring signals representative of respiratory activity, the method comprising: detecting, using a processing circuit, an electrocardiogram (ECG) signal through a plurality of electrodes, wherein the ECG signal comprises a plurality of ECG sample signals, wherein each ECG sample signal of the plurality of ECG sample signals is a bipolar signal associated with two of the plurality of electrodes and comprises a cardiac signal component and a myographic signal component representative of muscle contractions relating to the respiratory activity;filtering, using the processing circuit to generate combined myographic power signals, each ECG sample signal of the plurality of ECG sample signals to remove at least a portion of the cardiac signal component and generate a combined myographic power signal for the two of the plurality of electrodes with which a particular ECG sample signal is associated, wherein each combined myographic power signal of the combined myographic power signals represents a myographic potential between the two electrodes;calculating, using the processing circuit, individual myographic power signals for each of the plurality of electrodes by applying the combined myographic power signals within a covariance matrix;wherein: the method further comprises designating one of the plurality of electrodes as a reference electrode and designating electrodes of the plurality of electrodes other than the reference electrode as recording electrodes;detecting the ECG signal comprises detecting a first set of ECG sample signals comprising an ECG sample signal for each recording electrode with respect to the reference electrode;filtering each of the ECG sample signals comprises filtering each of the first set of ECG sample signals to generate a first set of combined myographic power signals, wherein the first set of combined myographic power signals comprises a combined myographic power signal for each recording electrode with respect to the reference electrode; andcalculating the individual myographic power signals comprises applying the first set of combined myographic power signals within the covariance matrix and calculating the individual myographic power signal for the reference electrode. 2. The method of claim 1, wherein the individual myographic power signal for the reference electrode is calculated based on one or more off-diagonal combined myographic power signals within the covariance matrix. 3. The method of claim 1, further comprising: changing the reference electrode to a different one of the plurality of electrodes to designate a new reference electrode;designating electrodes of the plurality of electrodes other than the new reference electrode to be new recording electrodes;determining a second set of combined myographic power signals comprising a combined myographic power signal for each new recording electrode with respect to the new reference electrode; andcalculating the individual myographic power signal fo... | Welch Allyn, Inc. | A61B005/0488 A61B005/08 A61B005/04 A61B005/0452 A61B005/0205 A61B005/0408 A61B005/0492 A61B005/00 A61N001/37 A61B005/0402 A61N001/05
US10702745 | B2 | 2020-07-07 | Facilitating dynamic monitoring of body dimensions over periods of time based on three-dimensional depth and disparity | 2018-01-30 | A mechanism is described for facilitating smart monitoring of body dimensions according to one embodiment. A method of embodiments, as described herein, includes receiving a first request to take a first picture of a user, wherein the first picture is taken at a first point in time using a depth-sensing camera; automatically computing first body dimensions relating to a body of the user based on at least one of a first image of the body and first depth information relating to one or more parts of the body, wherein the first image and the first depth information are obtained from the first picture; generate a first three-dimensional (3D) model of the body based on the first body dimensions; and communicating at least one of the first 3D model and the first body dimensions to a display device, wherein the display device to display at least one of the first 3D model and the first body dimensions. | 1. An apparatus for analyzing a body part of a user, the apparatus comprising: a depth-sensing device including one or more sensors to generate first image data with respect to the body part of the user and to generate second image data with respect to the body part, the depth-sensing device to generate the first image data at a first time and to generate the second image data at a second time after the first time;at least one processor; anda storage device including instructions that, when executed by the at least one processor, cause the at least one processor to: generate a first image indicative of a first change in the body part between the first time and the second time based on the first image data and the second image data, the first image including a first portion based on the first image data and a second portion based on the second image data, the second portion to at least partially overlap the first portion in the first image to show a difference in body shape between the first time and the second time;generate a prediction with respect to a second change in the body part at a third time based on the first image data and the second image data, the third time to occur after the first time and after the second time; andgenerate a second image of the body part based on a target body dimension for the body part, the target body dimension defined by a user input; andcause a display device to display at least one of the first image, the second image, or the prediction. 2. The apparatus as defined in claim 1, wherein the at least one processor is to generate a three-dimensional model of the body part based on the first image data. 3. The apparatus as defined in claim 2, wherein the target body dimension is a first body dimension and wherein the at least one processor is to calculate a second body dimension for the body part based on the three-dimensional model. 4. The apparatus as defined in claim 3, wherein the second body dimension includes one or more of a waist size or a chest size. 5. The apparatus as defined in claim 1, wherein the depth-sensing device is associated with an electronic user device different from the display device. 6. The apparatus of claim 3, wherein the at least one processor is to determine a difference between the target body dimension and the second body dimension. 7. The apparatus as defined in claim 5, wherein the depth-sensing device and the electronic user device are to communicate via a wireless communication protocol. 8. The apparatus as defined in claim 1, wherein the one or more sensors include a red-green-blue camera and an infrared sensor. 9. The apparatus as defined in claim 1, wherein the at least one processor is to generate the prediction based on the target body dimension for the body part. 10. The apparatus as defined in claim 1, wherein the body part is the full body of the user. 11. The apparatus as defined in claim 1, wherein the first portion overlapped by the second portion is at lea... | INTEL CORPORATION | A63B024/00 G16H010/60 G06T007/62 A61B005/107 G16H020/30 G16H020/60 G06T005/50 G06T017/00 G06T007/50 A61B005/00 G06T007/00 G06K009/00 G16H040/67 G06Q010/06 G09B019/00 A41H003/00 G06Q030/00 G06F019/00 A61B005/11
US10716437 | B2 | 2020-07-21 | Skin cleanser | 2018-06-13 | A skin cleanser includes a surface, such as a silicone surface, with at least one textured portion for transmitting vibrational tapping to the skin. The skin cleanser includes at least one oscillating motor for generating the tapping motion to the skin. The textured portion includes touch-points or a wave that transmit the tapping motion to skin in contact with the textured portions. The touch-points may include thicker and thinner formations of the touch-points to provide firmer or softer vibrations to the skin. The touch-points are within about 0.5 to 2.5 mm in diameter. One configuration includes multiple oscillating motors configured to provide different vibration frequencies at around 50-300 Hertz and operable simultaneously. | 1. A skin cleanser having a skin cleanser body, comprising: a substantially flat base configured to stand unaided on a substantially flat surface, said skin cleanser body having a cross-sectional shape that is longer in a first direction substantially parallel to the base than in a second direction substantially parallel to the base;a silicone exterior covering substantially parallel of an exterior of said skin cleanser body, the silicone exterior having: a first side defining first and second textured surface areas that are integrally formed with the silicone exterior, the first textured surface area comprising a first set of touch-points and the second textured surface area comprising a second set of touch-points, anda second side defining a third textured surface area that is integrally formed with the silicone exterior, the third textured surface area comprising a third set of touch-points;an oscillating motor disposed within said skin cleanser body configured to produce pulsations of the skin cleanser; andat least one control disposed on said skin cleanser body configured to operate the first oscillating motor. 2. The skin cleanser of claim 1, wherein the first textured surface area is smaller than the second textured surface area. 3. The skin cleanser of claim 2, wherein the third textured surface area is substantially the same size as the first and second textured surface areas combined. 4. The skin cleanser of claim 1, wherein the base is substantially transparent. 5. The skin cleanser of claim 1, wherein the oscillating motor is operable to oscillate within the range of 50 Hertz to 300 Hertz (Hz). 6. A skin cleanser having a skin cleanser body, comprising: a substantially flat base configured to stand unaided on a substantially flat surface, said skin cleanser body having a cross-sectional shape that is longer in a first direction substantially parallel to the base than in a second direction substantially parallel to the base;a silicone exterior covering substantially all of an exterior of said skin cleanser body, the silicone exterior defining a first textured surface area that is integrally formed with the silicone exterior, the first textured surface area comprising a first set of touch-points; andan oscillating motor disposed within said skin cleanser body configured to produce pulsations of the skin cleanser. 7. The skin cleanser of claim 6, wherein the silicon exterior further comprises a second textured surface area that is integrally formed with the silicone exterior, the second textured surface area comprising a second set of touch-points. 8. The skin cleanser of claim 7, wherein each touch-point of the first set of touch-points defines a first diameter; wherein each touch-point of the second set of touch-points defines a second diameter; andwherein the first diameter is greater than the second diameter. 9. The skin cleanser of claim 8, wherein at least one touch-point of the second set of touch-points is disposed closer... | FOREO Inc. | A47K007/02 A61H007/00 A61H023/00 A61H023/02 A61B005/00 A47K007/04
US10716502 | B2 | 2020-07-21 | Method for predicting arousal level and arousal level prediction apparatus | 2019-06-12 | A method for predicting an arousal level used by a computer of an arousal level prediction apparatus that predicts an arousal level of a user includes obtaining current biological information regarding the user detected by a first sensor, calculating a current arousal level of the user on the basis of the current biological information, obtaining current environment information indicating a current environment around the user detected by a second sensor, predicting a future arousal level, which is an arousal level a certain period of time later, on the basis of the current arousal level and the current environment information, and (i) issuing a notification to the user or (ii) controlling another device, on the basis of the future arousal level. | 1. A method for predicting an arousal level used by a computer of an arousal level prediction apparatus that predicts an arousal level of a user, the method comprising: obtaining, by a processor of the computer, current biological information regarding the user detected by a first sensor;calculating, by the processor, a current arousal level of the user based on the current biological information;obtaining, by the processor, current environment information indicating a current environment around the user;predicting, by the processor, a future arousal level, which is an arousal level a certain period of time later, based on the current arousal level and the current environment information; and(i) issuing, by the processor, a notification to the user or (ii) controlling, by the processor, another device, based on the future arousal level,wherein the biological information includes information indicating a body surface temperature of a periphery of a person's body and a deep-body temperature, andwherein the future arousal level becomes lower as a current body surface temperature of a periphery of the user's body becomes closer to the deep-body temperature. 2. The method according to claim 1, wherein, in the predicting of the future arousal level, a rate of decrease in an arousal level is estimated based on how likely it is for a person to become drowsy indicated by the current environment information, andwherein the future arousal level is predicted by correcting the current arousal level to the arousal level the certain period of time later using the rate of decrease. 3. The method according to claim 1, wherein the issuing or the controlling is performed only if the future arousal level is lower than a certain threshold. 4. The method according to claim 1, wherein the arousal level prediction apparatus is mounted on an automobile, andwherein the certain period of time is included in a time taken to arrive at a target destination, the target destination being inputted by the user to a navigation system installed in the automobile. 5. The method according to claim 1, wherein the biological information further includes information indicating a heart rate, andwherein the current arousal level becomes lower as a current heart rate of the user becomes lower. 6. The method according to claim 5, wherein the first sensor detects the current heart rate of the user using any of a milliwave sensor, a pulse oximeter, a speckle camera, and a laser Doppler velocimeter. 7. The method according to claim 1, wherein the biological information further includes information indicating a respiratory rate or an amount of air breathed, andwherein the current arousal level becomes lower as a current respiratory rate of the user becomes lower or a current amount of air breathed by the user becomes smaller. 8. The method according to claim 7, wherein the first sensor detects the current respiratory rate of the user or the current amount of air breathed by the user by m... | PANASONIC INTELLECTUAL PROPERTY CORPORATION OF AMERICA | A61B005/18 A61B005/01 B60K028/06 A61B005/00 A61B005/026 B60H001/00 A61B005/08 G16H040/63 A61B005/024 G08B021/06 A61B003/113 A61B005/103 A61B005/091 A61B005/0285
US10722116 | B2 | 2020-07-28 | Ophthalmic examination and disease management with multiple illumination modalities | 2016-10-17 | Imaging various regions of the eye is important for both clinical diagnostic and treatment purposes as well as for scientific research. Diagnosis of a number of clinical conditions relies on imaging of the various tissues of the eye. The subject technology describes a method and apparatus for imaging of the back and/or front of the eye using multiple illumination modalities, which permits the collection of one or more of reflectance, spectroscopic, fluorescence, and laser speckle contrast images. | 1. A method of imaging a region of tissue of an eye, the method comprising: configuring an ophthalmic imaging device to achieve a desired imaging modality by adjusting one or more optical assemblies, two or more illumination sources, or one or more image acquisition parameters, wherein the two or more illumination sources include a coherent light source and an incoherent light source;illuminating a first region of interest of the eye using one of the illumination sources;acquiring a first image by capturing light coming from the first region of interest based on the image acquisition parameters, wherein the first image is one of a laser speckle contrast image, spectroscopic image, reflectance image, or fluorescence image;reconfiguring the ophthalmic imaging device by adjusting the one or more optical assemblies, the two or more illumination sources, or the one or more image acquisition parameters;illuminating a second region of interest, the second region of interest at least partially overlapping the first region of interest;acquiring a second image by capturing light coming from the second region of interest, wherein second image is one of a laser speckle contrast image, spectroscopic image, reflectance image, or fluorescence image different from the first image;andprocessing the first image and the second image. 2. The method of claim 1, further comprising immobilizing an ophthalmic imaging device with respect to a subject's eye. 3. The method of claim 1, wherein the ophthalmic imaging device is handheld and the method comprises immobilizing the ophthalmic imaging device by resting or pressing against a subject's face or eye. 4. The method of claim 1, comprising: immobilizing, using eyeglasses, goggles, a helmet, or other accessory, the ophthalmic imaging device with respect to a subject's head or eye. 5. The method of claim 1, comprising: immobilizing, using a chin rest, a subject's head or eye. 6. The method of claim 1, comprising: generating a virtual object by the ophthalmic imaging device. 7. The method of claim 1, wherein the image acquisition parameters include at least one of exposure time, gain, pixel sensitivity, number of images, frame rate, timing sequence, pixel resolution, pixel area, or image magnification. 8. The method of claim 1, wherein reconfiguring the ophthalmic imaging device comprises switching from the coherent light source to the incoherent light source, or switching from the incoherent light source to the coherent light source. 9. The method of claim 1, wherein processing the first image or the second image comprises calculating laser speckle contrast values for pixels associated with the region of interest, wherein the calculated laser speckle contrast values are based on a pixel's neighborhood of pixels in spatial or temporal domains. 10. The method of claim 1, wherein the processing the first image and the second image comprises generating a compound image based on at least two of laser speckle contrast i... | Vasoptic Medical Inc. | A61B003/14 A61B003/00 A61B003/12 A61B003/15 A61B005/00 A61B005/02 G02B027/48 A61B003/18
US10722139 | B2 | 2020-07-28 | Navigation of an angioplasty guidewire | 2015-02-16 | A method, including advancing a guidewire having a first diameter through a patient lumen, the guidewire being tracked by an electromagnetic system and an impedance system. While advancing the guidewire, signals of both systems, that are generated in response to differing positions of the guidewire in the lumen, are recorded. The method includes recording correspondences between the signals at each of the differing positions. The method also includes withdrawing the guidewire from the lumen, and then advancing through the lumen a second guidewire, having a second diameter smaller than the first diameter, which is tracked by the impedance system. While advancing the second guidewire, a signal of the impedance system generated in response to advancement of the second guidewire in the lumen is received. The method includes applying the correspondences to the signal in order to determine a position of the second guidewire in the lumen. | 1. A method, comprising: advancing a first guidewire having a first diameter through a lumen in a body of a patient, the first guidewire being configured to be tracked by an electromagnetic tracking system comprising a conductive coil in said first guidewire and an impedance tracking system comprising an electrode on said first guidewire wherein said first guidewire is a mapping guidewire that does not include a catheter;while advancing the first guidewire, recording first signals of the electromagnetic tracking system and second signals of the impedance tracking system, the first and the second signals being generated in response to differing positions of the first guidewire in the lumen;recording respective correspondences between the first and second signals at each of the differing positions;withdrawing the first guidewire from the lumen;after withdrawing the first guidewire, advancing through the lumen a separate and unconnected second guidewire that is a delivery guidewire and includes a balloon catheter thereon, and having a second diameter smaller than the first diameter, and being configured to be tracked by the impedance tracking system;while advancing the second guidewire, receiving a third signal of the impedance tracking system generated in response to advancement of the second guidewire in the lumen; andapplying the respective correspondences to the third signal in order to determine a position of the second guidewire in the lumen,wherein the first guidewire is formed from a hollow elastic metal tube having an internal longitudinal lumen, wherein the material of the hollow elastic metal tube is a nitinol alloy, the hollow elastic metal tube being formed into a helix by cutting a spiral channel into the hollow elastic metal tube,wherein a conductive coil made of insulated wire is provided in the internal longitudinal lumen, said conductive coil having a first end and a second end, said conductive coil being located at a distal end of the hollow elastic metal tube, the conductive coil being wound on an elongate flexible core wire, wherein the core lies along an axis of symmetry of the internal longitudinal lumen, wherein the conductive coil has a common axis of symmetry with the axis of symmetry of the internal longitudinal lumen, wherein two coil-connecting wires respectively connect the first end and the second end of the conductive coil to a proximal end of the first guidewire, wherein said two coil-connecting wires are connected to the electromagnetic tracking system. 2. The method according to claim 1, wherein the conductive coil is configured to generate the first signals in response to a magnetic field generated by the electromagnetic tracking system, and the electrode configured to inject a current into the patient so as to generate the second signals as a set of currents received by respective electrodes on skin of the patient. 3. The method according to claim 1, wherein the second guidewire does not comprise a coil. 4. Th... | Biosense Webster (Israel) Ltd. | A61B005/06 A61M025/09 A61B034/20 A61B005/00
US10705043 | B2 | 2020-07-07 | Multi-dimensional approach to imaging, monitoring, or measuring systems and processes utilizing capacitance sensors | 2017-05-31 | A system and method for imaging, monitoring, or measuring systems and processes utilizing only data provided from capacitance sensors. The present invention combines the multi-frequency method of both ECVT/AECVT and DCPT to image or measure processes and systems more efficiently and accurately than the methods alone. The present system analyzes capacitance and current phase acquired at multiple frequencies to determine a plurality of properties of single and multiphase systems all at once. The combined use of ECVT and DCPT in multiphase flow can also be extended to measure volume fraction and phase distribution of flows involving greater than three phases by using multiple frequencies for capacitance, current phase, or both. | 1. A system for imaging, monitoring, measuring, or otherwise quantifying a material, combination of materials, or flow within a volume of a vessel interior or other object, the system comprising: a sensor comprising a plurality of electrodes for placement near or around the vessel or the object, wherein the sensor is adapted to provide electric field distribution and sensor sensitivity in at least two geometric dimensions;a data acquisition circuit in communication with the sensor for receiving input data from the sensor, the input data including current output from the sensor and for outputting electrical phase data for the current;a processing system in communication with the data acquisition circuit, the processing system programmed with instructions for executing on the processing system to convert capacitance data and electrical phase data for the current into an image or quantification of a volume being sensed;wherein the data acquisition circuit is adapted to output capacitance data and electric phase data at multiple frequencies to the processing system; andwherein the processing system is programmed with instructions for executing on the processing system to use the capacitance data and electrical phase data at different frequencies to measure a volume fraction of each phase of the flow and for using the capacitance or electrical phase data at different frequencies to calculate a velocity for each phase of the flow. 2. The system of claim 1, wherein the processing system is programmed with instructions for executing on the processing system to: determine which combination of capacitance data, electrical phase data for the current, and excitation frequencies to implement to obtain a desired quantification of a volume being sensed; and to implement the determined combination [to process the data acquisition circuit output] into the desired quantification of the volume being sensed. 3. The system of claim 2, wherein the desired quantification of the volume being sensed is a quantification of a predetermined property to be measured. 4. The system of claim 3, wherein the predetermined property to be measured is selected from the group consisting of temperature, volume fraction, phase distribution, velocity, acceleration, mass flow rate, volumetric flow rate, viscosity, conductivity, or charge. 5. The system of claim 1, wherein the processing system is programmed with instructions for executing on the processing system for using the product of velocity and volume fraction for each phase of the flow to determine mass flow rate. 6. The system of claim 1, wherein the processing system is programmed with instructions for executing on the processing system to reconstruct a three-dimensional volume-image from the capacitance data and electrical phase data at multiple frequencies. 7. A system for imaging, monitoring, measuring, or otherwise quantifying a material, combination of materials, or flow within a volume of a vessel interior or other o... | Tech4Imaging LLC | G01N027/22 G01F001/64 G06T011/00 A61B005/05
US10722188 | B2 | 2020-07-28 | X-ray CT apparatus | 2018-05-18 | An X-ray CT apparatus comprises a gantry including an X-ray generator in a rotor, and processing circuitry configured to control rotation of the rotor during a first scan such that a value related to first rotational speed serving as rotational speed of the rotor at a time when the first scan is finished is brought close to a setting value related to second rotational speed serving as rotational speed of the rotor in a second scan, with respect to the first scan and the second scan performed after the first scan is performed. | 1. An X-ray CT apparatus comprising: A gantry including an X-ray generator in a rotor; andprocessing circuitry configured tocontrol rotation of the rotor during a first scan such that a value related to first rotational speed serving as rotational speed of the rotor at a time when the first scan is finished is brought close to a setting value related to second rotational speed serving as rotational speed of the rotor in a second scan, with respect to the first scan and the second scan performed after the first scan is performed,control to accelerate or decelerate rotation of the rotor during the first scan to the setting value related to the second rotational speed, andcontrol a value related to relative movement speed of the gantry and a tabletop on which a subject is placed during the first scan, in accordance with rotation of the rotor during the first scan. 2. The X-ray CT apparatus according to claim 1, wherein the processing circuitry is configured to control to accelerate rotation of the rotor during the first scan to a setting value related to third rotational speed serving as rotational speed of the rotor in the first scan, and thereafter decelerate rotation to the setting value related to the second rotational speed. 3. The X-ray CT apparatus according to claim 2, wherein the processing circuitry is further configured to control a tube current value used for generation of X-rays with the X-ray generator during the first scan, in accordance with rotation of the rotor during the first scan. 4. The X-ray CT apparatus according to claim 3, wherein the processing circuitry is configured to control the X-ray generator to increase the tube current value in accordance with acceleration of rotation of the rotor during the first scan, and decrease the tube current value in accordance with deceleration of rotation of the rotor during the first scan. 5. The X-ray CT apparatus according to claim 1, wherein the processing circuitry is further configured to control a tube current value used for generation of X-rays with the X-ray generator during the first scan, in accordance with rotation of the rotor during the first scan. 6. The X-ray CT apparatus according to claim 5, wherein the processing circuitry is configured to control the X-ray generator to increase the tube current value in accordance with acceleration of rotation of the rotor during the first scan, and decrease the tube current value in accordance with deceleration of rotation of the rotor during the first scan. 7. The X-ray CT apparatus according to claim 5, wherein the processing circuitry is configured to control rotation of the rotor during the first scan such that the value related to the first rotational speed becomes the setting value related to the second rotational speed. 8. The X-ray CT apparatus according to claim 5, wherein the processing circuitry is further configured to control to accelerate rotation of the rotor to a setting value related to third rotational speed s... | CANON MEDICAL SYSTEMS CORPORATION | A61B006/03 A61B006/00 A61B006/02 A61B005/05 A61B006/04
US10722145 | B2 | 2020-07-28 | Devices, systems and methods for health monitoring using circumferential changes of a body portion | 2016-11-22 | Devices, systems, and methods for monitoring health parameters of an individual, including circumferential changes to a portion of the individual's body, are provided herein. A monitoring device includes: a stretchable component configured to fit securely around a body portion, and a sensor module coupled thereto. The sensor module is configured to obtain and transmit a circumference measurement of the body portion. The sensor module may be further configured to obtain and transmit measurements of one or more additional parameters. A related monitoring system includes the described monitoring device and a mobile computing device. The mobile computing device is configured, at least in part, to: process the body circumference measurements to identify and analyze any change in circumference of the body portion, and generate a relevant alert output. Methods performed by the various devices and systems are also provided. | 1. A monitoring system for detecting circumferential changes to a portion of a body, the monitoring system comprising: a sensor system wearable around a portion of an individual's body and configured to obtain measurements for a plurality of parameters of the body portion over a period of time, the plurality of parameters including one or more of: a circumference of the body portion, a movement of the body portion, and a temperature of the body portion;a processor communicatively coupled to the sensor system; anda non-transitory computer-readable medium with instructions stored thereon, wherein the instructions, when executed by the processor, cause the processor to perform a method comprising: receiving the measurements for the plurality of parameters,applying relative weights to the plurality of parameters to generate weighted measurements,receiving inputs specifying which plurality of parameters and one or more additional factors to include in an overall score, the one or more additional factors comprising one or more of: a compliance score and a user input entered by the individual,calculating the overall score based on the weighted measurements and the one or more additional factors, andgenerating an alert output when the overall score exceeds a predefined threshold indicating a change in circumference to the body portion. 2. The monitoring system of claim 1, wherein the compliance score is an indication of a level of compliance by the individual with prescribed instructions, and wherein the compliance score is determined from one or more of: a detected leg orientation, a detected leg movement, and a user-entered input related to compliance. 3. The monitoring system of claim 2, wherein the prescribed instructions comprise one or more of instructions to elevate a leg, instructions to perform an exercise, and instructions to take a medication. 4. The monitoring system of claim 2, wherein the prescribed instructions are customizable by the professional using a supervisor computing device. 5. The monitoring system of claim 1, wherein the plurality of parameters further includes one or more of an orientation and a color of the body portion. 6. The monitoring system of claim 1, wherein the plurality of parameters further includes one or more of a pulse, a heart rate, a blood oxygenation, and a blood volume of the individual. 7. The monitoring system of claim 1, further comprising a mobile computing device comprising the processor and the non-transitory computer-readable medium. 8. The monitoring system of claim 7, further comprising a supervisor computing device communicatively coupled to the mobile computing device, wherein the supervisor computing device is configured to receive inputs from a professional, and transmit instructions to the mobile computing device, specifying which plurality of parameters and additional factors to include in the overall score. 9. The monitoring system of claim 1, wherein the instructions stored on the comp... | Plethy, Inc. | A61B005/01 A61B005/107 A61B005/00 A61B005/0205 A61B005/103 A61B005/11 A61B005/024 A61B005/029 A61B005/145
US10722136 | B2 | 2020-07-28 | Wireless and power-source-free extravasation and infiltration detection sensor | 2012-08-29 | A system detects extravasation or infiltration by segregating active components that drive a passive sensor for economical single use. A receiving antenna of the passive sensor receives a transmitted signal comprising RF electromagnetic power. A first circuit transmits a first portion of the received signal through a body portion. A sensor detects a resultant signal from the body portion. A second circuit combines a reference signal comprising a second portion of the received signal with the resultant signal so as to define an output signal. A transmit antenna transmits the output signal to a receiver. | 1. A system for detecting extravasation, comprising: a signal generator for transmitting a transmitted signal comprising Radio Frequency (RF) electromagnetic power;a passive sensor, comprising: a receiving antenna configured to receive the transmitted signal as a received signal,a first circuit configured to transmit only a first portion of the received signal through the body portion,a sensor configured to detect a resultant signal from the body portion,a second circuit configured to combine a reference signal comprising a second portion of the received signal with the resultant signal so as to define an output signal, anda transmit antenna configured to transmit the output signal as a transmitted output signal; anda signal processor configured to: receive the transmitted output signal, andperform a comparison of peak amplitude of resultant signal and the reference signal of the transmitted output signal to detect extravasation of fluid in the body portion by determining a change in fluid level in the body portion using the transmitted output signal,wherein the transmitted signal further comprises a time domain pulse of electromagnetic power, andwherein the passive sensor further comprises:(i) a first delay circuit configured to delay a selected one of the first portion of the received signal comprising the time domain pulse and the second portion of the received signal comprising the time domain pulse as the reference signal, and(ii) a second delay circuit configured to delay the other one of the first portion and the second portion of the time domain pulse, the first and second delay circuits being configured to separate in time the resultant signal, the reference signal, and a noise pulse at the sensor arising due to a reflected portion of the transmitted signal. 2. The system of claim 1, wherein the receiving antenna is further configured to recieve the transmitted signal in a range of greater than 1.5 GHz to approximately 30 GHz over a period of time. 3. The system of claim 1, wherein the transmitted signal comprises a sweep of continuous wave electromagnetic power across a frequency range. 4. The system of claim 1, further comprising a substrate configured to support said receiving antenna, said first circuit, said second circuit, said sensor, and said transmit antenna. 5. A method for detecting extravasation, comprising: transmitting, by a signal generator, a transmitted signal comprising Radio Frequency (RF) electromagnetic power;received, by a receiving antenna of a passive sensor, the transmitted signal as a received signal;transmitting, by a first circuit of the passive sensor, only a first portion of the received signal through a body portion;detecting, by a sensor of the passive sensor, a resultant signal from the body portion;combing, by a second circuit of the passive sensor, a reference signal comprising a second portion of the received signal with resultant signal to define an output signal;transmitting, by a transmit antenn... | Battelle Memorial Institute | A61B005/05 A61M005/168 A61B005/053 A61B008/08 A61B005/00
US10702197 | B2 | 2020-07-07 | Dual amplitude modulation and polarization frequency modulation as well as compensation for noninvasive glucose monitoring | 2017-07-21 | An embodiment relates generally to an improved method and apparatus for modulating the amplitude and rotation of the plane of polarization of linearly polarized light for multiple uses but primarily as part of a noninvasive glucose monitoring system. As compared to previous monitoring systems, an embodiment provides faster monitoring while maintaining or even reducing noise and minimizing system complexity. Embodiments described herein address these concerns with a modulation and compensation approach that both uses a single high speed device and also modulation of the lasers. | 1. A system to measure glucose levels in an aqueous humor of an eye, the system comprising: two light sources to generate two light beams at different wavelengths;two polarizers to respectively linearly polarize each of the light beams;two polarization modulation modules to respectively modulate the rotation of a plane of polarization of each of the linearly polarized light beams, each of the polarization modulation modules comprising: (a)(i) a magnetically permeable core having an air gap; (a)(ii) a magneto-optic material at least partially disposed within the air gap; and (a)(iii) at least one conductor wrapped around a portion of the magnetically permeable core; anda detector module;wherein with each polarization modulation module when a current is passed through at least one conductor (b)(i) a magnetic field is generated and propagates through the magnetically permeable core, and (b)(ii) the magnetic field rotates the plane of polarization of the linearly polarized light for the light source, when the linearly polarized light is directed through the magneto-optic material, generating a modulated light beam;wherein the detector unit is to (c)(i) receive the modulated light beams after the modulated light beams are passed through the aqueous humor of the eye, and (c)(ii) detect a change in the plane of polarization of each of the modulated light beams in response to a concentration of glucose in the aqueous humor. 2. The system of claim 1, wherein the magneto-optic material includes a terbium-gallium-garnet (TGG) crystal. 3. The system of claim 2, wherein the TGG crystal is approximately 1 centimeter long. 4. The system of claim 3, wherein with each polarization modulation module a first one of the one or more conductors is to carry an alternating current and a second one of the one or more conductors is to carry a direct current. 5. The system of claim 4, wherein the alternating current has a frequency between 1 kHZ and 1 MHz. 6. The system of example 4, wherein the alternating current is to generate an alternating magnetic field. 7. The system of example 6, wherein the alternating magnetic field is to modulate the rotation of the plane of polarization of the light between a lower angle of rotation and an upper angle of rotation. 8. The system of claim 4, wherein the direct current is supplied by a feedback loop to produce a magnetic field that compensates for the rotation from the glucose signal. 9. The system of claim 1 wherein the light sources are modulated at frequencies between 10 KHz to 1 MHz. 10. The system of claim 1 wherein the light sources are heterodyned with the polarization modulation modules. 11. A method of measuring glucose levels in an aqueous humor of an eye comprising: generating a linearly polarized light beam;passing the linearly polarized light beam through a polarization modulation device comprising: (a)(i) a magnetically permeable core; (a)(ii) an aperture extending through at least a portion of the magnetic... | The Texas A&M University System | A61B005/1455 A61B005/145 A61B005/00 G02F001/09
US10709384 | B2 | 2020-07-14 | Wearable heat flux devices and methods of use | 2016-08-16 | Wearable heat flux devices are disclosed that can detect heat flux based on evaporative cooling for determining a core body temperature of a user, and that can heat or cool a surface of a user for reaching a steady-state heat flux to determine the core body temperature of the user. Exemplary heat flux devices can include a heat flux sensor and a wicking layer. The heat flux sensor can be configured to detect heat flux at a location on a user. The wicking layer can be configured to absorb moisture at the location and to transport the moisture above the heat flux sensor. The heat flux subsequently detected by the heat flux sensor includes the evaporative cooling from the evaporation of the moisture. | 1. A heat flux device wearable at a location on a user, the heat flux device comprising: a heat flux sensor configured to be in contact with the user and detect heat flux at the location, the heat flux sensor comprising a plurality of thermocouple junctions; anda wicking layer having a bottom surface configured to be in contact with the user and a top surface configured to be exposed to an environment of the user, the wicking layer including a first portion and a second portion, the first portion surrounding at least part of vertical sides of the heat flux sensor and the second portion covering a top surface of the heat flux sensor, the wicking layer configured to absorb moisture at the location through the first portion and to transport the moisture to the second portion for evaporation above the heat flux sensor, wherein the heat flux detected by the heat flux sensor includes evaporative cooling from the evaporation of the moisture. 2. The heat flux device of claim 1, wherein the heat flux sensor includes a thermal resistance layer, and the heat flux sensor is configured to detect the heat flux based on a temperature differential across the thermal resistance layer. 3. The heat flux device of claim 2, wherein the heat flux sensor includes a first thermocouple junction that detects a first temperature on a first side of the thermal resistance layer and a second thermocouple junction that detects a second temperature on a second side of the thermal resistance layer, opposite the first side, and the temperature differential is based on a difference between the first temperature and the second temperature. 4. The heat flux device of claim 2, wherein the temperature differential generates a voltage based on the See beck effect, and the heat flux is determined based on the voltage. 5. The heat flux device of claim 1, wherein a bottom surface of the heat flux sensor and the bottom surface of the first portion of the wicking layer are substantially equal in surface area. 6. The heat flux device of claim 1, further comprising: an adhesive layer covering a bottom surface of the first portion of the wicking layer,wherein the adhesive layer is at least partially perforated. 7. The heat flux device of claim 6, wherein a bottom surface of the adhesive layer and a bottom surface of the heat flux sensor are substantially coplanar. 8. The heat flux device of claim 1, further comprising an insulation layer above the first portion of the wicking layer and surrounding the second portion of the wicking layer. 9. The heat flux device of claim 8, wherein a top surface of the insulation layer and a top surface of the second portion of the wicking layer are substantially coplanar. 10. The heat flux device of claim 1, further comprising: a temperature sensor configured to determine a skin temperature of the user at the location. 11. The heat flux device of claim 10, wherein a bottom surface of the temperature sensor is substantially coplanar with a bottom surf... | MC10, INC. | G01K003/00 A61B005/00 G01K013/00 G01K001/14 A61B005/01 G01K007/42 G01K017/00
US10702759 | B2 | 2020-07-07 | Systems and methods for coordinating musculoskeletal and cardiovascular or cerebrovascular hemodynamics | 2019-08-26 | Described herein are methods for determining a target musculoskeletal activity cycle (MSKC) to cardiac cycle (CC) timing relationship. The method may include detecting a signal responsive to a cyclically-varying arterial blood flow at a location on a head of a user; providing a recurrent prompt at a frequency of the heart pump cycle using the signal, such that the signal correlates with a magnitude of blood flow adjacent to the location, and the recurrent prompt is provided to guide the user to time performance of a component of a rhythmic musculoskeletal activity with the recurrent prompt; and guiding the user to adjust a timing of the component of the rhythmic musculoskeletal activity to substantially maximize a magnitude of the signal. In some embodiments, the method further includes generating the recurrent prompt by amplifying the sound generated by the blood flow in or in proximity to an ear of the user. | 1. A method for synchronizing a rhythmic musculoskeletal activity cycle (“MSKC”) of a user exercising on an exercise apparatus with a rhythmic cardiac cycle (“CC”) of the user, the method comprising: detecting a first signal feature representative of a CC timing of the user using a first detector, the first signal feature repeating at a heart rate of the user;detecting a second signal feature representative of an MSKC timing of the user using a second detector, the second signal feature repeating at a cadence of the user;determining an actual timing relationship between the CC timing or first signal feature and the MSKC timing or second signal feature using a controller integrated into or coupled to the exercise apparatus;determining a difference value between the actual timing relationship and a target timing relationship using the controller; andadjusting operations of the exercise apparatus in response to a signal from the controller, wherein the operations of the exercise apparatus are adjusted to alter the MSKC timing of the user such that the magnitude of the difference value is reduced. 2. The method of claim 1, wherein the exercise apparatus is one of: a treadmill; an aerobic machine; a bicycle; a stationary bicycle; a stair-stepper; and a rowing machine. 3. The method of claim 1, wherein the exercise apparatus is a treadmill and adjusting operations comprises adjusting a plurality of settings. 4. The method of claim 3, wherein adjusting the plurality of settings comprises adjusting both a speed and an incline of the treadmill. 5. The method of claim 3, wherein adjusting the plurality of settings comprises more than one of: (a) a change in speed; (b) a change in resistance; (c) a change in incline; (d) a change in stride length; (e) a change in stroke length; (f) a change in crank length; (g) a change in frequency of a repetitive motion; and (h) a change in gearing. 6. The method of claim 1, further comprising receiving an input from the user via the controller and providing an output to the user via the controller. 7. The method of claim 1, further comprising identifying, with the controller, the target timing relationship. 8. The method of claim 7, wherein identifying the target timing relationship comprises performing an empirical calibration for a specific user, the empirical calibration comprising: exposing the user to at least two different CC to MSKC timing relationships; using signals from a detector and a processor to measure a physiologic response in the user exposed to the at least two different timing relationships; and using the controller to select the target timing relationship based on the physiologic response. 9. The method of claim 1, wherein the plurality of settings is adjusted to maintain a target value or range of values for at least one of a heart rate and a work output of the user. 10. A system that synchronizes a rhythmic musculoskeletal activity cycle (“MSKC”) of a user exercising on an exercise apparatu... | Pulson, Inc. | A63B071/06 A61B005/11 A61B005/0205 G09B019/00 A63B024/00 G09B023/28 A61B005/0456 G16H020/30 A61B005/026 A61B005/00 G16H050/30 A61B005/024 A63B021/002 A63B023/12 A63B022/06 A63B022/02 A63B022/00 A63B021/06 A63B015/00 A61H003/00 A61H001/02 A63B069/00
US10712417 | B2 | 2020-07-14 | Method and apparatus for executing an accelerated magnetic resonance scan | 2018-08-10 | In a method and apparatus for creating magnetic resonance data of at least two simultaneously manipulated, non-overlapping slices of an examination object by a parallel acquisition technique, reference data are acquired such that, between acquisition of slice scan data of a slice scan data set in which the scan data of all simultaneously manipulated slices are incorporated in an overlaid manner, and its associated reference data, no slice scan data of a different slice scan data set are acquired. A high level of robustness with respect to movements of the examination object is thereby achieved. | 1. A method for generating a magnetic resonance (MR) image of an examination object, comprising: operating a radio-frequency (RF) radiator of an MR scanner in order to radiate a multi-band RF excitation pulse that simultaneously and selectively manipulates magnetization of nuclear spins in at least two non-overlapping slices of a target volume in the examination object;operating the MR scanner so as to acquire echo signals produced by said multi-band RF excitation pulse from said at least two slices, as slice scan data in a slice scan data set, said slice scan data set comprising scan data for all simultaneously manipulated slices of the target volume;operating the MR scanner so as to acquire reference data for said at least two non-overlapping slices, with reference data for each respective slice scan data set being acquired by, between acquisition of the slice scan data of the respective slice scan data set, and acquisition of the reference data associated therewith, no scan data of slices being acquired that are not included in said at least two non-overlapping slices of the respective slice scan data set;entering the slice scan data for each slice scan data set together with the reference data therefor into a memory organized as k-space, by filling a central region of k-space with said slice scan data for a respective slice scan data set and the associated reference data forming a completely scanned data set, according to the Nyquist criterion, in said central region of k-space;in a computer, accessing said completely scanned data set from said memory and generating calibration data from said completely scanned data set;in said computer, using said calibration data to separate the slice scan data set into single-slice scan data sets respectively for the at least two non-overlapping slices; andreconstructing respective image data sets from the respective single-slice scan data sets, and making the image data sets available in electronic form from the computer. 2. A method as claimed in claim 1 comprising reconstructing said image data sets by completing said single-slice scan data sets using said calibration data. 3. A method as claimed in claim 1 comprising acquiring all echo signals for an individual slice scan data set after radiating a single multi-band RF excitation pulse. 4. A method as claimed in claim 1 comprising operating said MR scanner according to a CAIPIRINHA method in order to acquire said slice scan data. 5. A method as claimed in claim 1 comprising operating said MR scanner according to a turbo spin echo method in order to acquire said slice scan data. 6. A method as claimed in claim 5 comprising operating said MR scanner according to a HASTE sequence, as said turbo spin echo sequence. 7. A method as claimed in claim 1 comprising acquiring said reference data interleaved with the slice scan data for the respective slice scan data set associated with the reference data. 8. A method as claimed in claim 1 comprising acquir... | Siemens Healthcare GmbH | G01R033/561 G01R033/483 A61B005/055 G01R033/48 G01R033/54 G01R033/565 A61B090/00 G01R033/24 G01R033/58
US10702180 | B2 | 2020-07-07 | Atrial fibrillation discrimination using heart rate clustering | 2018-01-08 | This document discusses, among other things, apparatus, systems, and methods to determine a first atrial fibrillation (AF) indication using received information about a heart over a first period, to cluster depolarization information about the heart over the first period, and to discriminate between an atrial fibrillation (AF) event and a non-AF event in the first period using the determined first AF indication the clustered depolarization information. | 1. A system, comprising: an atrial fibrillation (AF) detection circuit configured to receive information about a heart over a first period including a number of cardiac intervals, and to discriminate between an AF event and a non-AF event, the AF detection circuit including: a first AF detection circuit configured to determine a first AF indication in the first period using the received information about the heart; anda second AF detection circuit configured to receive depolarization information about the heart over the first period, to bin the cardiac intervals into separate rate bins using the received depolarization information, and to cluster the received depolarization information,wherein the AF detection circuit is configured to discriminate between an AF event and a non-AF event in the first period using the determined first AF indication and the clustered depolarization information, andwherein, to cluster the received depolarization information, the second AF detection circuit is configured to group successive rate bins having bin counts greater than a bin threshold and a group bin count greater than a threshold percentage of the number of cardiac intervals over the first period. 2. The system of claim 1, wherein, when the first AF indication is indicative of an AF event in the first period, the second AF detection circuit is configured to cluster the received depolarization information to verify the first AF indication, and wherein the AF detection circuit is configured to provide an indication of an AF event in the first period when the first AF indication is indicative of an AF event in the first period and the clustered depolarization information is indicative of an AF event in the first period. 3. The system of claim 1, wherein, when the first AF indication is indicative of a non-AF event over the first period, the second AF detection circuit does not cluster the received depolarization information, and the AF detection circuit is configured to provide an indication of a non-AF event over the first period using only the first AF indication. 4. The system of claim 1, wherein the AF detection circuit is configured to receive information about a ventricle of the heart, wherein the depolarization information includes ventricular depolarization information, and wherein the first AF detection circuit is configured to determine a ventricular rate stability over the first period using the ventricular depolarization information, and to determine the first AF indication using the determined ventricular rate stability. 5. The system of claim 1, wherein the AF detection circuit is configured to receive information about a ventricle of the heart, including ventricular depolarizations, over the first period, and wherein the first AF detection circuit is configured to determine the first AF indication using V-V intervals between successive ventricular depolarizations over the first period. 6. The system of claim 1, wherein the received depolar... | Cardiac Pacemakers, Inc. | A61B005/04 A61B005/046 A61B005/00 G16H040/67 A61N001/39 A61B005/024 A61N001/362
US10722710 | B2 | 2020-07-28 | Secretion clearance and cough assist | 2018-03-24 | A device operates for secretion clearance and cough assistance by means of transdermal stimulation of muscle groups including the pectoralis majoris, the serratus anterior, and the abdominal muscles. The various muscle groups may be stimulated in different ways: not just different pulse trains but in addition, different stimulations to accomplish different phases of clearance or cough assistance. In a first phase a vibration action provided by the muscles (for example the pectoralis and serratus) is used to encourage the motion of mucus and other secretions within airways. In a second phase, the abdominal muscles for example might be stimulated differently so as to cause not a vibration but instead a coughing action, especially to detect and assist a natural cough. The invention further teaches a zone stimulator large enough to span and stimulate major muscles or groups of large muscles effectively. | 1. A secretion clearance and cough assistance device configured for use by a patient having a body, skin, a mouth, airways, first, second, third, and fourth pairs of abdominal muscles, and chest muscles including pectoral muscles and serratus anterior muscles, such patient suffering from bodily secretions in such airways, the secretion clearance and cough assistance device comprising: a control module having operative electrical connections to a plurality of dermal electrodes configured to be attached to such skin of such patient, whereby the control module is in communication with the dermal electrodes, the control module configured so as to be worn on such patient body;a first one of the dermal electrodes configured to be disposed on such skin of such patient at one such chest muscle;a second one of the dermal electrodes configured to be disposed on such skin of such patient at one such abdominal muscle;each of the dermal electrodes configured to deliver a plurality of pulse trains to one such respective muscle;the control module having a stimulation module operative to send a first pulse train to such chest muscle and a second pulse train to such abdominal muscle;the first pulse train operative to stimulate such chest muscle so as to cause a first vibration of such chest muscle, the vibration of such chest muscle thereby causing a second vibration of such airways;the second pulse train operative to stimulate such abdominal muscle so as to stimulate at least one cough;whereby the second vibration loosens such bodily secretions in such airways and the at least one cough moves such bodily secretions toward such mouth of such patient. 2. The secretion clearance and cough assistance device of claim 1, configured for use with a shirt worn on such body by such patient, wherein: the control module, the dermal electrodes and the operative electrical connections are configured so as to be worn on such body concealed within such shirt, and the secretion clearance and cough assistance device further comprising a garment distinct from such shirt:the garment configured to be worn about such body by such patient, with the garment concealed within such shirt, and the garment concealing within itself the control module, dermal electrodes and operative electrical connections. 3. The secretion clearance and cough assistance device of claim 2, further comprising: a third one of the dermal electrodes configured to be disposed on such skin of such patient at a second such abdominal muscle;a fourth one of the dermal electrodes configured to be disposed on such skin of such patient at a third such abdominal muscle;a fifth one of the dermal electrodes configured to be disposed on such skin of such patient at a fourth such abdominal muscle;the control module further operative to send the second pulse train to such second, third and fourth abdominal muscles. 4. The secretion clearance and cough assistance device of claim 3, further comprising: a sixth one of the derm... | - | A61N001/36 A61N001/04 A41B001/08 A61N001/02 A61B005/00 A61B005/08
US10725613 | B2 | 2020-07-28 | Radiation imaging apparatus, display processing method of the same, and computer-readable storage medium | 2015-01-09 | A radiation imaging apparatus for imaging a radiation image inputs an operation instruction from a user and performs display control of a screen that has an input region in which one or more information input regions for having various types of information input thereto are disposed, and a workflow region in which workflow information related to imaging of the radiation image is disposed. Here, when a confirmation instruction indicating confirmation of an operation is input, the radiation imaging apparatus moves the one or more information input regions of the input region toward a prescribed region of the workflow region, and performs display using animation that sequences the movement of the one or more information input regions. | 1. A radiation imaging apparatus for imaging a radiation image, comprising: a memory storing a program; andone or more processors which, by executing the program, function as:a display control unit configured to concurrently display a selection region and a workflow region on a screen of a display unit, the selection region being a region for allowing selection of protocol buttons displayed therein, and the workflow region being a region for presenting information related to imaging of the radiation image, wherein a prescribed region of the workflow region is a protocol list region for having a list of protocols selected using the one or more protocol buttons displayed therein, anda selection unit configured to select at least one of the protocol buttons in the selection region,in response to selection of at least one of the protocol buttons: the display control unit configured to move the selected at least one of the protocol button to the protocol list region and displayed therein with the degree of transparency thereof gradually being reduced while moving the selected at least one of the protocol button,an attribute change unit included in the display control unit configured to change attribute information of the selected at least one of the protocol buttons during the moving, wherein a format of the protocol buttons moved from the selection to the workflow region is changed and displayed therein,wherein the protocol buttons in the workflow region and the protocol buttons in the selection region add information of body part for imaging,wherein in the workflow region, the selected at least one of the protocol buttons further add attribute information including a sensor type and a mount type,wherein the selected protocol buttons in the workflow region are arranged in accordance with an imaging order and the imaging order of the selected protocol buttons can be changed, andwherein in response to a start of the imaging of the radiation image, the selection region switches into a radiation imaging region for displaying the radiation image, the selected protocol button in the workflow region becomes an active protocol and a sensor corresponding to the sensor type and the mount type of the selected protocol button in the workflow region becomes an imaging preparation state. 2. The radiation imaging apparatus according to claim 1, wherein the selected protocol button is moved and displayed from a left boundary position of the workflow region. 3. The radiation imaging apparatus according to claim 1, wherein a size of the selected protocol button in the workflow region is changed and attribute information is added to the selected button. 4. The radiation imaging apparatus according to claim 1, wherein the display control unit changes a format of the selected button while performing display, when moving the protocol button using the animation. 5. The radiation imaging apparatus according to claim 1, wherein in the workflow region, the selected protoc... | CANON KABUSHIKI KAISHA | G06F003/0484 G06F003/0482 G06F003/048 G16H040/63 G06F019/00 A61B005/055 G01T001/161 G06Q050/24 G06Q050/22 G06F003/0481
US10716476 | B2 | 2020-07-21 | Methods to improve axial resolution in optical coherence tomography | 2015-10-06 | Methods are proposed to improve axial resolution in optical coherence tomography (OCT). In one aspect, the method comprises: obtaining a k-space interferogram of an OCT spectral image; uniformly reshaping the k-space interferogram to a quasi-stationary interferogram by extracting a source envelope; fitting a spectral estimation model to the quasi-stationary interferogram; and calculating an axial depth profile using the fitted spectral estimation model. | 1. A method of obtaining an axial depth profile of a subject from an Optical Coherence Tomography (OCT) spectral image of the subject, the method comprising: a) obtaining a k-space interferogram of the OCT spectral image of the subject;b) uniformly reshaping the k-space interferogram to a quasi-stationary interferogram by extracting a source envelope;c) fitting a spectral estimation model to the quasi-stationary interferogram; andd) calculating the axial depth profile of the subject using the fitted spectral estimation model. 2. The method according to claim 1 wherein the spectral estimation model comprises a parametric model or a non-parametric model. 3. The method according to claim 1 wherein the step of uniformly reshaping the k-space interferogram comprises using a Hilbert transform to extract the envelope A as a function of k and sample depth, d; averaging A(k, d) across the dimension d to obtain the source envelope A(k); and dividing k-space interferogram signal by A(k) to obtain the quasi-stationary interferogram Sspre(k). 4. The method according to claim 1 wherein the subject is a cellular structure. | NANYANG TECHNOLOGICAL UNIVERSITY | A61B005/00 A61B003/00 G01B009/00 A61B008/13 G01B009/02 A61B003/10 G01N021/47
US10709346 | B2 | 2020-07-14 | Flexible neural probes | 2017-03-30 | Embodiments include microelectrodes including a flexible shank and a bioabsorbable material surrounding the flexible shank. The flexible shank can include a flexible substrate, a circuit, and a plurality of sensors. Embodiments also include a methods of forming flexible active electrode arrays including depositing a flexible polymer on a substrate. The methods also include forming a plurality of sensors on the flexible polymer and attaching a silicon-based chip to the flexible shank. The methods also include coating the flexible shank in a bioabsorbable material and cutting the shank and a portion of the bioabsorbable material from the substrate. | 1. A method of measuring neural activity, the method comprising: providing a flexible substrate comprising a polymeric material over a rigid semiconductor substrate;providing a plurality of active electrodes on the flexible substrate to define a flexible active electrode shank, each active electrode comprising a gate directly on a surface of the polymeric material, a gate dielectric over the gate, an aminopropyltriethoxysilane (APTES) layer on a surface of the gate dielectric and the surface of the polymeric material, and a carbon nanotube monolayer on a surface of the APTES layer;embedding the flexible active electrode shank in a bioabsorbable material;releasing the flexible active electrode shank from the rigid semiconductor substrate;implanting the flexible active electrode shank in a biological tissue;dissolving the bioabsorbable material in the biological tissue; andsensing neural activity using the plurality of active electrodes. 2. The method of claim 1 further comprising routing a neural signal. 3. The method of claim 1 further comprising detecting a pH local to the flexible active electrode shank. 4. The method of claim 1 further comprising detecting action potential. 5. The method of claim 1 further comprising detecting local field potential. 6. The method of claim 1 further comprising amplifying a neural signal received from one of the plurality of active electrodes. 7. The method of claim 1 further comprising providing an output to the biological tissue. 8. The method of claim 7, comprising providing the output to a neural tissue. 9. The method of claim 1, wherein the flexible active electrode shank further comprises a circuit embedded within the flexible substrate. 10. The method of claim 9, wherein the circuit comprises a complimentary metal-oxide semiconductor chip. 11. The method of claim 1, wherein the bioabsorbable material comprises gelatin. 12. The method of claim 1, wherein the plurality of active electrodes comprises a routing sensor. 13. The method of claim 1, wherein the plurality of sensors comprises an action potential sensor. 14. The method of claim 1, wherein the plurality of active electrodes comprises greater than or equal to 25 sensors. 15. The method of claim 14, wherein the plurality of active electrodes comprises greater than or equal to 30 sensors. 16. The method of claim 1, wherein the flexible active electrode shank comprises a multiplexer. 17. The method of claim 1, wherein the flexible active electrode shank comprises a demultiplexer. | INTERNATIONAL BUSINESS MACHINES CORPORATION | A61B005/04 A61B005/00 A61B005/0478 A61B005/145
US10709892 | B2 | 2020-07-14 | Temporal configuration of a motion sensor in an implantable medical device | 2018-03-26 | Methods and devices for configuring the use of a motion sensor in an implantable cardiac device. The electrical signals of the patient's heart are observed and may be correlated to the physical motion of the heart as detected by the motion sensor of the implantable cardiac device in order to facilitate temporal configuration of motion sensor data collection that avoids detecting cardiac motion in favor of overall motion of the patient. | 1. A method of operating an implantable device (ID) adapted for disposition within the heart of a patient, the ID comprising a motion sensor for detecting movement of the patient, the method comprising: initializing the ID by: contemporaneously collecting a set of data points from the motion sensor and from a cardiac electrical signal;correlating time periods within the set of data points from the motion sensor to cardiac activity reflected by the cardiac electrical signal; andconfiguring the ID to capture data from the motion sensor by reference to specific repeatable elements of the cardiac electrical signal;using the ID as configured in the initializing step to capture cardiac signals of the patient and motion signals for the patient. 2. The method of claim 1 further comprising; using the motion signals to determine an activity level of the patient; andsetting and implementing a pacing rate for the patient using the determined activity level. 3. The method of claim 1 wherein the step of configuring the ID to capture data from the motion sensor is performed by: determining when intrinsic mechanical motion of the heart is detected by the motion sensor relative to electrical signals of the heart; andsetting a data capture period for the motion sensor relative to a specified electrical signal of the heart to avoid capturing the intrinsic mechanical motion of the heart with the motion sensor. 4. The method of claim 1 wherein the step of configuring the ID to capture data from the motion sensor is performed by: determining when a quiet period in the intrinsic mechanical motion of the heart is detected by the motion sensor relative to electrical signals of the heart; andsetting a data capture period for the motion sensor relative to a specified electrical signal of the heart to avoid capturing the intrinsic mechanical motion of the heart with the motion sensor. 5. The method of claim 1 wherein the step of configuring the ID to capture data from the motion sensor includes determining a delay after a T-wave in the cardiac electrical signal. 6. The method of claim 5 wherein the delay is configured as a function of cardiac rate. 7. The method of claim 6 wherein the cardiac rate is determined as a rate of pacing output of the ID. 8. The method of claim 6 wherein the cardiac rate is determined as an intrinsic rate of the patient's heart. 9. The method of claim 1 wherein the specific repeatable elements of the cardiac electrical signal are one of an R-wave or a QRS complex. 10. The method of claim 1 wherein the specific repeatable elements of the cardiac electrical signal comprise the T-wave. 11. The method of claim 1 wherein the step of configuring the ID to capture data from the motion sensor by reference to specific repeatable elements of the cardiac electrical signal further includes configuring the ID to maintain the motion sensor in a low power state except when the ID is capturing data therefrom. 12. An implantable medical device (IMD) ada... | CARDIAC PACEMAKERS, INC. | A61N001/365 A61N001/37 A61B005/11 A61B005/00 A61N001/368 A61N001/375 A61B005/0452 A61N001/372
US10722161 | B2 | 2020-07-28 | Implantable sensor devices, systems, and methods | 2013-03-15 | Disclosed herein are devices, systems, and methods for a continuous analyte sensor, such as a continuous glucose sensor. In certain embodiments disclosed herein, various in vivo properties of the sensor's surroundings can be measured. In some embodiments, the measured properties can be used to identify a physiological response or condition in the body. This information can then be used by a patient, doctor, or system to respond appropriately to the identified condition. | 1. A method performed using a continuous glucose sensor, the method comprising: receiving sensor data generated by a continuous glucose sensor, wherein the sensor data is indicative of a concentration of glucose in a user;identifying, using a processor module, a post-implantation loss of sensitivity of the continuous glucose sensor, wherein identifying the post-implantation loss of sensitivity comprises: measuring a first impedance associated with the continuous glucose sensor at a first time point;measuring a second impedance associated with the continuous glucose sensor at a second time point, wherein the first time point is different than the second time point; anddetermining an occurrence of the post-implantation loss of sensitivity by comparing the first impedance measurement with the second impedance measurement;deactivating the continuous glucose sensor responsive to the determination of the occurrence of the post-implantation loss of sensitivity event;based on the comparing of the first impedance measurement and the second impedance measurement, determining that the post-implantation loss of sensitivity event is temporary; andafter the continuous glucose sensor is deactivated, reactivating the continuous glucose sensor. 2. The method of claim 1, wherein the sensor data comprises data indicative of a signal response to at least one event selected from the group consisting of a signal response to cessation of blood flow to a site surrounding sensor implantation, a signal response to reduction of blood flow to a site surrounding sensor implantation, and a signal response to a vasospastic event. 3. The method of claim 1, wherein identifying the post-implantation loss of sensitivity further comprises: deactivating the continuous glucose sensor for a time period, whereby a product from a catalyzed reaction of glucose and oxygen accumulates over the time period;activating the continuous glucose sensor and measuring a signal value of the continuous glucose sensor immediately after the time period; anddetermining an occurrence of a post-implantation loss of sensitivity event if the signal value is greater than a predetermined value. 4. The method of claim 1, wherein identifying the post-implantation loss of sensitivity further comprises determining a pH of a biological fluid surrounding the continuous glucose sensor. 5. The method of claim 1, wherein identifying the post-implantation loss of sensitivity further comprises: measuring a concentration value of a non-glucose analyte; anddetermining an occurrence of a post-implantation loss of sensitivity event if the concentration value of the non-glucose analyte changes more than a predetermined amount. 6. The method of claim 1, wherein the continuous glucose sensor comprises a first electrode and a second electrode. 7. The method of claim 6, wherein the first electrode and second electrode have different dimensions. 8. The method of claim 7, wherein the second electrode is configured to have a... | DexCom, Inc. | A61B005/00 A61B005/1495 A61B005/1473 A61B005/1459 A61B005/145 A61B005/1486
US10709180 | B2 | 2020-07-14 | Protection system and method | 2016-07-01 | This disclosure provides a protective gear and system for cushioning an object against falling. The protective gear includes a housing conforming to a portion of the object, an airbag disposed in the housing, a detector configured to detect a motion of the object; and a controller configured to generate an actuating signal to deploy the airbag if the motion is over a threshold. The protection system includes a primary protective gear and at least one secondary protective gear. | 1. A protective gear for cushioning a subject against an impact, comprising: a housing conforming to a body joint of the subject;an airbag disposed in the housing;a detector comprised of an acceleration sensor, configured to detect an acceleration subject for indicating a motion of the subject; anda controller coupled with the acceleration detector, configured to determine whether falling occurs based on data of the acceleration of the subjection detected by the acceleration sensor, and to generate an actuating signal to deploy the airbag if the motion is over a threshold,wherein the acceleration sensor detects the acceleration of the subject by sensing acceleration of the object along an X-axis, a Y-axis, and a Z-axis in a three-dimensional coordinate system, the threshold is determined by a sum of the acceleration vectors along the X, Y, Z axis and the vector sum is less than 1g. 2. The protective gear according to claim 1, wherein the protective gear is configured to cushion a human body or a delicate device. 3. The protective gear according to claim 1, wherein the impact happens after falling of the object to ground. 4. The protective gear according to claim 1, further comprising a gas generator, wherein the gas generator is disposed in the housing, and is configured to produce gas to inflate and deploy the airbag out of the housing in response to the actuating signal generated by the controller. 5. The protective gear according to claim 4, wherein: the gas generator and the airbag are arranged in an airbag groove in the housing; andan airbag outlet is arranged on an outer surface of the housing opposite to a surface of the joint to be cushioned, is spatially connected with the airbag groove, and is configured to allow the airbag to expand out of the airbag groove therethrough upon inflation and deployment by the gas generator. 6. The protective gear according to claim 4, wherein the gas generator comprises a tip discharge device and a gas generating agent, wherein: the tip discharge device is configured to generate an electrical spark in response to the actuating signal generated by the controller; andthe gas generating agent is configured to generate gas upon contacting the electrical spark. 7. The protective gear according to claim 6, wherein the gas generating agent is selected from a non-azide gas generant or liquid CO2. 8. The protective gear according to claim 5, wherein the airbag is detachably mounted in the airbag groove, and the airbag groove is provided with a retractable window, arranged on a side wall that is close to the outer surface of the housing and configured to allow replacing the airbag therethrough. 9. The protective gear according to claim 1, wherein the acceleration sensor and the controller are integrated in a circuit board. 10. The protective gear according to claim 2, wherein the protective gear is configured to cushion a joint selected from one of elbow, knee, neck, or waist, and the impact happens after ... | BOE TECHNOLOGY GROUP CO., LTD. | A41D013/00 A41D013/018 A41D001/00 A61B005/11 A41D013/05 A41D013/06 A41D013/08
US10716834 | B2 | 2020-07-21 | Intervention device | 2013-09-06 | A device for interventions within the body, the device comprising: an end piece 6 for insertion into the body at a distal end thereof, the end piece 6 including a rigid lumen for holding an instrument 10 and guiding the instrument 10 to the distal end of the end piece; and a body section 4 supporting the lumen and being rigidly connected thereto, the body section including a navigation array 14 for guidance of the device using a surgical navigation system and/or including an anchor point 20 for a standard navigation array. | 1. A device adapted for interventions within a patient's body, the device comprising: an end piece adapted to be inserted into the patient's body at a distal end of the end piece, the end piece including a rigid lumen for holding an instrument and guiding slidable movement of the instrument to the distal end of the end piece;a body section supporting the lumen and being rigidly connected thereto;a navigation array rigidly connected to the instrument and held in a track on the body section that permits slidable movement of the navigation array relative to the body section, the navigation array comprising a plurality of optical markers at known locations relative to the end piece or an electromagnetic location sensor at a known location relative to the end piece, and wherein the navigation array is configured to allow for guidance of the device in three dimensions toward a target site within the patient's body using a surgical navigation system;a proximal holder piece configured to hold a proximal end of the instrument, the proximal holder piece being positioned at a proximal end of the end piece and of the body section;one or more clamps configured to attach to the instrument, wherein the one or more clamps are configured to fix the instrument in place relative to the end piece and the distal end of the end piece thereof; andmoveable parts configured to connect to and move with the instrument as the instrument moves slidably relative to the end piece;wherein an entirety of the body section, with the navigation array connected thereto, is placed/positioned between the end piece and the proximal holder piece, wherein the instrument extends from a proximal end of the proximal holder piece, through the body section and into the lumen of the end piece, and wherein the slidable movement of the instrument relative to the end piece moves the instrument along the lumen and through the distal end of the end piece into the patient's body. 2. The device of claim 1, wherein the distal end of the end piece comprises a tip adapted to pierce the patient's body, the tip comprising a tapered profile narrowing toward a point. 3. The device of claim 1, wherein the end piece comprises a marking adapted to show a depth of insertion of the end piece into the patient's body. 4. The device of claim 1, wherein the lumen is configured to be rigid enough to permit placement of a tip of the end piece with millimeter accuracy without deformation as the lumen penetrates body tissues and while being subject to any bending moments that arise as the lumen is manoeuvred along a selected approach toward the target site. 5. The device of claim 4, wherein the device is adapted to be used for cranial use and the rigidity of the lumen is sufficient to enable the tip to be placed with millimeter accuracy at the patient's sphenopalatine ganglion or otic ganglion when targeted via a lateral approach. 6. The device of claim 5, wherein the device is adapted to be used for cranial use an... | Norwegian University of Science and Technology (NTNU) | A61B005/00 A61K038/48 A61B010/02 A61B017/34 A61P027/02 A61N001/372 A61P025/06 A61M005/20 A61M005/24 A61K009/00 A61M005/28 A61M005/32 A61B017/00 A61B090/00 A61B034/20
US10706610 | B2 | 2020-07-07 | Method for displaying an object | 2017-10-16 | A method and a device for displaying an object, in particular biological tissue. The method has the following steps: a) generating a first image of at least one sub-region of the object using a first device; b) generating a second image of at least one sub-region of the object using a second device; c) ascertaining first coordinates of at least some image points of the second image in a first coordinate system; d) ascertaining second coordinates of the image points of the second image by projecting the first coordinates in a second coordinate system which is different from the first coordinate system and which is assigned to the first device; and e) generating a combined image of the object from the first and the second image using the ascertained second coordinates of the image points of the second image. | 1. A method for determining image points while displaying an object, in particular biological tissue, comprising the steps of: a) generating a first image of at least one sub-region of the object with a first imaging device;b) generating a second image of at least one sub-region of the object with a second imaging device;c) ascertaining first coordinates of at least some of the image points of the second image in a first coordinate system;d) ascertaining second coordinates of at least some of the image points of the second image by projecting the first ascertained coordinates in a second coordinate system which is different from the first coordinate system and which is assigned to the first imaging device; ande) generating a combined image of the object from the first and the second images using the ascertained second coordinates of the image points of the second image, wherein selection of the image points to be combined is achieved by selecting a sub-region of the object which is marked successively in a marking process including: using the first imaging device that comprises a navigated light source to generate a multitude of distinct light spots on the object as supporting points, orusing a separate navigated instrument to mark the object with a plurality of supporting points,so that generation of the combined image is achieved following the marking process by replacing or overlaying image points of the first image within the selected sub-region with image points of the second image. 2. The method of claim 1, wherein the first imaging device is an intraoperative imaging device and the first image is an intraoperative image. 3. The method of claim 1, wherein the first imaging device is selected from the group consisting of: a computed tomography (CT) device, a magnetic resonance tomography (MRT) device, a nuclear radiological device, and a single photon emission computed tomography (SPECT) device. 4. The method of claim 1, wherein the second imaging device is an endoscopy device and the second image is an endoscopy image. 5. The method of claim 1, wherein the second imaging device is an ultrasound device and the second image is an ultrasound image. 6. The method of claim 1, wherein the second imaging device is an intraoperative imaging device and the second image is an intraoperative image. 7. The method of claim 1, wherein the second imaging device is selected from the group consisting of: a computed tomography (CT) device, a magnetic resonance tomography (MRT) device, a nuclear radiological device, and a single photon emission computed tomography (SPECT) device. 8. The method of claim 1, wherein the first imaging device is an ultrasound device and the first image is an ultrasound image. 9. The method of claim 8, wherein the second imaging device is an intraoperative imaging device and the second image is an intraoperative image. 10. The method of claim 9, wherein the second imaging device is selected from the group consisting of: a ... | Scopis GmbH | G06T015/08 G06T007/33 A61B001/04 A61B001/00 A61B008/00 A61B008/08 A61B005/06 G06T011/60 G06T015/50 A61B034/20 A61B090/00 A61B006/03 A61B006/00
US10702672 | B2 | 2020-07-07 | Pressure-sensing catheters and related methods | 2016-02-25 | A multi-lumen catheter can be used to measure pressure at multiple locations within the vasculature. The multi-lumen catheter can include multiple segments, such as a proximal portion, an intermediate portion, and a distal portion. A segment of a multi-lumen catheter may differ from another segment of the same multi-lumen catheter in radiodensity, hardness, and/or some other characteristic. Some multi-lumen catheters are designed to permit measurements of pressure in different lumens. | 1. A catheter comprising: a body portion comprising: a main lumen centrally disposed along the body portion;a first lumen disposed radially outward from the main lumen;a second lumen disposed radially outward from the main lumen opposite the first separate lumen;a distal portion having at least one port that is in fluid communication with the first lumen of the catheter, wherein the distal portion has a first radiodensity;a proximal portion having at least one port that is in fluid communication with the second lumen of the catheter, wherein the proximal portion has a second radiodensity; andan intermediate portion disposed between the proximal portion and the distal portion, the intermediate portion having a third radiodensity; andan integrated pressure readout housing disposed at a proximal end of the body portion;wherein the proximal end of the body portion terminates in a main proximal conduit that comprises the main lumen, a first conduit that comprises the first lumen, and a second conduit that comprises the second lumen, wherein a portion of each conduit is disposed within the integrated pressure readout housing and a portion of each conduit is disposed outside of the integrated pressure readout housing;wherein a first inline pressure sensor is disposed within the integrated pressure readout housing and is in communication with the first lumen of the first conduit and a second inline pressure sensor is disposed within the integrated pressure readout housing and is in communication with the second lumen of the second conduit;wherein the third radiodensity differs from both the first radiodensity and the second radiodensity;wherein the intermediate portion comprises a bend when the catheter is in an unconstrained state; andwherein the third radiodensity facilitates positioning the intermediate portion across a heart valve. 2. The catheter of claim 1, wherein the first lumen and the second lumen are not in fluid communication with one another. 3. The catheter of claim 1, wherein the intermediate portion is less radiodense than both the distal portion and the proximal portion. 4. The catheter of claim 1, wherein the intermediate portion is more radiodense than both the distal portion and the proximal portion. 5. The catheter of claim 1, wherein the bend of the intermediate portion forms an angle of between 130° and 170° when the catheter is in the unconstrained state. 6. The catheter of claim 1, further comprising a first elongate tube and a second elongate tube that is at least partially disposed within the first elongate tube, wherein the first elongate tube and the second elongate tube form at least a portion of the proximal portion of the catheter. 7. The catheter of claim 6, further comprising: a plurality of septal walls extending radially inward from the first elongate tube to the second elongate tube; anda plurality of radial lumens defined, at least in part, by the plurality of the septal walls, wherein the plurality of radial l... | Merit Medical Systems, Inc. | A61B005/02 A61M025/00 A61B005/0215 A61M025/01 A61B005/00 A61B090/00
US10712335 | B2 | 2020-07-14 | Sensor clip for stacked sensor dispensing system, and systems, methods and devices for making and using the same | 2015-06-12 | Sensor clip assemblies, sensor clips, analyte testing systems, and methods for making and using the same are disclosed. A sensor clip assembly is disclosed for storing and dispensing analyte testing sensors. The sensor clip assembly includes numerous test sensors arranged in a stack. Each test sensor is configured to assist in testing an analyte in a fluid sample. The sensor clip assembly also includes a skeletal frame with a top, a bottom, and numerous sides. The top, bottom and sides are interconnected to define an internal chamber within which is stored the stack of test sensors. At least one of the sides includes one or more elongated rails with structural gaps on opposing sides thereof. For some configurations, multiple sides of the skeletal frame comprise at least one or multiple elongated rails, each of which has structural gaps on opposing sides thereof and may be columnar in nature. | 1. A sensor clip assembly for storing and dispensing analyte testing sensors, the sensor clip assembly comprising: a plurality of test sensors arranged in a stack, each of the test sensors being configured to assist in testing an analyte in a fluid sample; anda skeletal frame with a top, a bottom, and a plurality of sides, the top, bottom and sides being interconnected to define an internal chamber within which is stored the stack of test sensors, at least one of the sides including a plurality of elongated rails spaced from one another by a gap disposed between adjacent elongated rails, wherein the bottom of the skeletal frame defines an aperture configured to receive therethrough the stack of test sensors; anda cap disposed on the top of the skeletal frame such that the top-most test sensor from the stack of sensors lies flush against a bottom of the cap, the cap including an elongated slot extending through the cap for receiving an ejection mechanism for advancing the test sensors, one at a time, out of the internal chamber of the skeletal frame. 2. The sensor clip assembly of claim 1, wherein at least a second one of the sides of the skeletal frame comprises a respective plurality of elongated rails spaced from one another by a second gap disposed between adjacent elongated rails. 3. The sensor clip assembly of claim 1, wherein each of the plurality of elongated rails is columnar, extending between and connecting the top and the bottom of the skeletal frame. 4. The sensor clip assembly of claim 1, wherein the bottom of the skeletal frame comprises a pair of opposing flexible tabs configured to retain the stack of test sensors inside the internal chamber. 5. The sensor clip assembly of claim 4, the opposing flexible tabs being configured to flex such that the stack of test sensors can pass through the bottom of the skeletal frame into the internal chamber. 6. The sensor clip assembly of claim 1, wherein at least one of the sides of the skeletal frame comprises one or more compliant alignment rails configured to align the stack of test sensors within the internal chamber. 7. The sensor clip assembly of claim 1, further comprising a push plate on which is seated the stack of test sensors. 8. The sensor clip assembly of claim 1, further comprising a pocket attached to the skeletal frame, the pocket being configured to store therein a desiccant material. 9. The sensor clip assembly of claim 1, further comprising an auto-calibration tab attached to the skeletal frame, the auto-calibration tab including detailed calibration information for the sensor clip assembly. 10. The sensor clip assembly of claim 1, wherein the analyte is glucose and the test sample is blood. 11. The sensor clip assembly of claim 1, wherein the plurality of test sensors includes electrochemical sensors. 12. The sensor clip assembly of claim 11, wherein each of the electrochemical sensors includes a base, one or more electrodes supported by the base, and a reagent in... | ASCENSIA DIABETES CARE HOLDINGS AG | G01N033/487 A61B005/15 A61B005/157 B65D083/08 C12Q001/00 G01N027/327 G01N021/78 G01N021/77
US10709354 | B2 | 2020-07-14 | Generating a breathing alert | 2017-09-06 | For generating a breathing alert, a method captures a video stream of a subject. The method further generates a vector time series that includes a vector for each image frame of the video stream. The method estimates of breathing signal from the vector time series. The method further determines one of large-scale motion and a breathing event of the subject based on the breathing signal. The method generates an alert if no breathing event is identified and no large-scale motion of the subject is identified within an event time interval. | 1. A method comprising: receiving, by a processing apparatus, a video stream of a subject;generating, by the processing apparatus, a vector time series that comprises a vector for each image frame of the video stream;estimating, by the processing apparatus, a breathing signal from the vector time series;determining, by the processing apparatus, one of a large-scale motion and a breathing event of the subject based on the breathing signal; andgenerating, by the processing apparatus, an alert if both no breathing event is identified and no large-scale motion of the subject is identified within an event time interval,wherein estimating the breathing signal comprises: decomposing a matrix X of a windowed video subsequence of the video stream into a sparse matrix S representing moving objects and a low rank matrix L representing non-moving objects;reconstructing a video stream of the moving objects from the sparse matrix S, wherein each pixel of the reconstructed video stream of the moving objects comprises one of a scalar time series and a vector time series; andmeasuring breathing events in the reconstructed video stream. 2. The method of claim 1, further comprising capturing the video stream using a bandpass filter in the range of 0.8-2.5 micrometers and with a Charge Coupled Device (CCD) tuned to 1.5 micrometers. 3. The method of claim 1, wherein generating the vector time series comprises dividing each image frame into regions and summing pixel values in each region as a pixel sum. 4. The method of claim 1, wherein generating the vector time series comprises dividing each image frame into regions and removing a signal mean in each region. 5. The method of claim 1, the method further comprising: generating a breathing report based at least in part on breathing events determined based on the breathing signal, wherein the breathing report comprises one or more of: a breath rate, a maximum inter-breath interval, a minimum inter-breath interval, inter-breath interval statistics, an inter-breath interval histogram, and apnea event data; andgenerating a motion report based at least in part on large-scale motions determined based on the breathing signal, wherein the motion report comprises one or more of: a motion frequency, a motion magnitude, a motion duration, a sleep length, a sleep quality, and sleep intervals. 6. The method of claim 1, wherein estimating the breathing signal comprises applying an adaptive subspace algorithm to each vector of the vector time series. 7. The method of claim 1, wherein estimating the breathing signal further comprises: defining a cursor t that indexes the image frames I of the video stream;defining a window length N;extracting the windowed video subsequence of the video stream comprising N image frames I(t−N+1), I(t−N+2), I(t−N+3), . . . , I(t−2), I(t−1), I(t); andorganizing the windowed video subsequence into the matrix, X, having M rows and N columns, where M is the product of a height, a width and a number of ... | - | G06K009/00 A61B005/087 A61B005/113 A61M016/00 A61B005/00 A61B005/08
US10709356 | B2 | 2020-07-14 | Apparatuses and methods for extracting chemicals from bodily cavities and breath | 2016-09-29 | A chemical extraction apparatus has absorbent material for absorbing or adsorbing volatile chemicals from a breath of a user. The chemical extraction apparatus is positioned within such that breaths from the user flow over the absorbent material thereby causing volatile chemicals in the breaths to absorb or adsorb into the absorbent material. The absorbent material is then analyzed to determine the amount of chemicals absorbed or adsorbed into the absorbent material. Based on such analysis, various diseases or conditions can be detected or diagnosed. | 1. A pacifier, comprising: a support element; anda nipple for insertion into an oral cavity of a user, the nipple coupled to the support element and having absorbent material for absorbing or adsorbing trace levels of volatile chemicals in breaths of the user, wherein the nipple includes an inner wall defining a cavity along a length of the nipple, wherein the cavity includes a hole allowing access to the cavity, wherein the hole is configured to allow breaths to pass through the hole and into the cavity when the pacifier is in use, and wherein the absorbent material is fixedly attached to the inner wall of the nipple by a single arm extending from the wall. 2. The pacifier of claim 1, further comprising a handle coupled to the support element. 3. A method, comprising: providing a pacifier having a nipple mounted on a support element, wherein the nipple has absorbent material for absorbing or adsorbing trace levels of volatile chemicals in breaths of a user, wherein the nipple includes an inner wall defining a cavity along a length of the nipple, wherein the cavity includes a hole allowing access to the cavity, wherein the hole is configured to allow breaths to pass through the hole and into the cavity when the pacifier is in use, and wherein the absorbent material is fixedly attached to the inner wall of the nipple by a single arm extending from the wall;inserting the nipple, including the absorbent material, of the pacifier into an oral cavity of the user such that the trace levels of the volatile chemicals are absorbed or adsorbed into the absorbent material;removing the nipple from the oral cavity after a period of time;analyzing the absorbent material with analytical instrumentation;identifying, with the analytical instrumentation, the trace levels of the volatile chemicals based on the analyzing;providing an output from the analytical instrumentation based on the trace levels identified by the analytical instrumentation; anddiagnosing a condition of the user based on the output. 4. The method of claim 3, wherein the method further comprises removing the absorbent material from the inner wall. 5. The method of claim 4, wherein the removing the absorbent material from the inner wall comprises cutting the single arm. 6. The method of claim 3, wherein the nipple further comprises at least one additional absorbent material. 7. The method of claim 3, wherein the analytical instrumentation comprises a gas chromatograph. 8. The method of claim 3, wherein the analytical instrumentation comprises a mass spectrometer. 9. The method of claim 3, wherein the absorbent material comprises polydimethylsiloxane. | Volatile Analysis Corporation | A61B005/097 G01N001/40 A61J017/00 A61C005/90 A61B010/00 A61B005/08 G01N001/22 A61J007/00
US10722719 | B2 | 2020-07-28 | Vibration-based secure side channel for medical devices | 2016-02-12 | According to some embodiments, a system for securing communications between an implantable wearable medical device (IWMD) and an external device (ED) is disclosed. The system includes a wireless radio frequency (RF) channel configured for communication between the IWMD and the ED. The system further includes a vibration-based side channel configured for verifying communication between the IWMD and the ED such that the RF channel is activated only when the IWMD detects a vibration signal generated by an ED. | 1. A system for securing communications between an implantable or wearable medical device (IWMD) and an external device (ED), the system comprising: a wireless radio frequency (RF) channel configured to communicate medical information between the IWMD and the ED; anda vibration-based side channel configured to verify the communication of medical information between the IWMD and the ED;the RF channel being activated when the IWMD detects a vibration signal generated by the ED via a motion sensing device of the IWMD and determines an amplitude gradient and amplitude mean of the vibration signal are above respective thresholds via a demodulator of the IWMD. 2. The system of claim 1, wherein the IWMD comprises a high-pass filter to eliminate vibration channel noise. 3. The system of claim 1, wherein the IWMD and ED each comprise a RF module configured to transmit, receive, or transmit and receive the medical information. 4. The system of claim 1, wherein the ED is a mobile device. 5. The system of claim 1, wherein the ED comprises a modulator and vibration source for transmitting the vibration signal. 6. The system of claim 1, wherein the ED comprises a speaker configured to generate a masking sound to mask the vibration signal. 7. The system of claim 1, wherein the motion sensing device comprises an accelerometer. 8. The system of claim 1, wherein an encryption key is transmitted through the vibration channel to the IWMD. 9. The system of claim 8, wherein the medical information is encrypted when the IWMD receives the encryption key. 10. The system of claim 9, wherein the IWMD performs a key reconciliation process when the encryption key includes one or more ambiguous bits. 11. A method for securing radio frequency (RF) communication of medical information between an implantable or wearable medical device (IWMD) and an external device (ED) via a vibration-based side channel, the method comprising the steps of: activating the IWMD to detect a vibration signal;placing the IWMD in a measurement mode when a vibration signal is detected;high-pass filtering vibration signals below a predetermined frequency when the IWMD is in the measurement mode, the IWMD to return to a standby mode when no vibration signals above the predetermined frequency are detected; andactivating an RF channel for secure communication of the medical information when a vibration signal detected above the predetermined frequency has an amplitude gradient and amplitude mean above respective thresholds. 12. The method of claim 11, further comprising receiving a vibration signal from an ED. 13. The method of claim 11, further comprising masking the vibration signal. 14. The method of claim 11, further comprising transmitting an encryption key through the vibration-based side channel. 15. The method of claim 14, further comprising encrypting the medical information when the IWMD receives the transmitted encryption key. 16. The method of claim 15, further comprising perform... | THE TRUSTEES OF PRINCETON UNIVERSITY | A61N001/372 A61B005/00 H04L029/06 H04W012/04 H04W012/02 H04W012/06
US10709374 | B2 | 2020-07-14 | Systems and methods for assessment of a musculoskeletal profile of a target individual | 2018-03-12 | There is provided a method of indirectly estimating muscle strength ratio, comprising: receiving images and associated body part locations of a target individual, wherein the images depict the target individual performing a first and a second defined movement, identifying first image(s) depicting the first defined movement, and obtaining an associated first set of body part locations, identifying second image(s) depicting the second defined movement, and obtaining an associated second set of body part locations, computing a first image-metric according to the first set of body part locations, computing a second image-metric according to the second set of body part locations, computing an image-parameter according to the first and second image-metrics, and converting the image-parameter to an estimate of a measured-parameter indicative of strength measurement ratio of the target muscle(s) obtained by a dynamometer, according to correlation between image-parameters and measured-parameters obtained based on the dynamometer performing empirical measurements. | 1. A system for indirectly estimating a value indicative of muscle strength ratio between a first target muscle and a second target muscle of a target individual based on an analysis of a plurality of digital images, comprising: a non-transitory memory having stored thereon a code for execution by at least one hardware processor of a computing device, the code comprising: code for receiving, from at least one sensor, a video including a plurality of images and associated plurality of body part locations of a target individual corresponding to the plurality of images, wherein the video depicts the target individual performing a physical movement test comprising a first defined movement and a second defined movement;code for identifying at least one first image of the plurality of imaged depicting the target individual performing the first defined movement, and obtaining a first set of body part locations associated with the at least one first image;code for identifying at least one second image of the plurality of imaged depicting the target individual performing the second defined movement, and obtaining a second set of body part locations associated with the at least one second image;code for computing a first image-metric according to the first set of body part locations associated with the at least one first image, and computing a second image-metric according to the second set of body part locations associated with the at least one second image;code for computing a single image-parameter as a function of the first image-metric and the second image-metric; andcode for converting the single image-parameter to an estimate of a measured-parameter indicative of a muscle strength measurement ratio between a first target muscle obtained by a dynamometer to a second target muscle obtained by the dynamometer, according to correlation code that correlates between values of the single image-parameter and measured-parameters obtained based on the dynamometer performing empirical measurements of application of an exerted force of the first and second target muscles by each of a plurality of other subjects. 2. The system according to claim 1, wherein the first defined movement is based on no significant resistance of a target muscle, and wherein the second defined movement is based on at least partial resistance of the target muscle. 3. The system according to claim 1, wherein the first defined movement is based on a first partial resistance of a target muscle, and wherein the second defined movement is based on a second resistance of the target muscle greater than the first partial resistance. 4. The system according to claim 1, wherein the first defined movement and the second defined movement are indicative of force proportion of muscle groups. 5. The system according to claim 4, wherein the first defined movement and the second defined movement indicative of force proportion of muscle groups are selected from the group consisting of: flexion, exten... | Physimax Technologies Ltd. | G06K009/00 A61B005/22 A61B005/11 A61B005/00 G06T007/20
US10702154 | B2 | 2020-07-07 | Systems and methods for modulating physiological state | 2019-03-01 | Methods and systems are provided to monitor bio-signals associated with one or more physiological structures of a subject, process the bio-signals to determine neural information, determine a current state of the subject based on the neural information, determine feedback to be provided to the subject based on the current neural state and a desired neural state, and/or to provide the feedback to the subject in order to modify the subject's current neural state to the desired neural state. The neural information may include validated neural indices relating to dynamic regulation of the monitored physiological structures via identifiable neural pathways and/or neural metrics quantified from such neural indices. | 1. A method comprising: A. receiving bio-signal information relating to bio-signals associated with a monitored physiological structure of a subject;B. determining neural information from the bio-signal information, the neural information comprising: one or more neural indices determined from the bio-signal information; andone or more neural metrics determined from at least one neural index of the one or more neural indices;C. determining a current state of the subject based on the neural information;D. upon determining that the current state is different than a desired state: determining feedback to be provided to the subject, based on the current state and the desired state; andproviding the determined feedback to the subject to thereby modify neural regulation of the monitored physiological structure of the subject; andE. repeating steps A-D until the current state is substantially similar to the desired state or a predetermined time period expires. 2. A method according to claim 1, wherein the bio-signals are selected from the group consisting of: heart rate, pulse rate, respiratory rate, blood pressure, cardiac rhythm, oxygen saturation, pH, pulse wave slope, pulse wave amplitude, vascular tone, an electromagnetic field, sudomotor activity, an electrocardiogram (“ECG”), an electrogastrogram (“EGG”), an electroencephalogram (“EEGs”), an electromyogram (“EMG”), an electrooculography (“EOG”), electrodermal activity (“EDA”), galvanic skin response (“GSR”), electrical potential of a body surface, electrical resistance of a body surface, a mechanomyogram (“MMG”), an endocrine level, an immune reaction, cerebral spinal fluid pressure, pupil diameter, vocalization, movement, body temperature, and blood glucose level. 3. A method according to claim 2, wherein the bio-signals comprise a vocalization. 4. A method according to claim 3, wherein the one or more neural indices comprises one or more of: frequency band of the vocalization, modulation of intonation within the frequency band, tempo of the vocalization, volume of the vocalization, and modulation of the volume. 5. A method according to claim 1, wherein the desired state relates to a physiological, emotional or behavioral state. 6. A method according to claim 1, wherein the desired state differentially and selectively supports a desired outcome relating to the subject, the desired outcome selected from the group consisting of: increasing attentiveness, improving cognitive ability, improving concentration, improving critical thinking, improving creativity, improving productivity, improving health, promoting growth, increasing restoration, improving recovery following illness or injury, increasing mobility, decreasing mobility, improving performance of a physical activity, improving ability to engage in spontaneous interactions with others, increasing sexual desire, improving sexual function, improving ability to eat or drink, improving digestion, improving gastrointestinal motility, increasin... | Polyvagal Science LLC | A61B005/00 A61B005/04 A61B005/0402 G06N003/02 A61B005/0496 A61B005/024 A61B005/1455 A61B005/01 G16H010/60 A61B005/02 A61B005/0476 A61B005/145 A61B005/053
US10709449 | B2 | 2020-07-14 | Systems and methods for variable stiffness tethers | 2012-02-17 | Described herein are devices and methods for guide elements configured with variable stiffness, with one or more flexible portions and one or more stiff portions. The flexible portion is be used as a tensioning element of a cinchable implant to tighten or compress tissues while the stiff portion is used to facilitate the insertion or withdrawal of portions of the implant or instruments acting on the implant. The guide element is further configured to be separated or severed between the flexible and stiff portions so that a flexible portion is left within the body as part of the implant, while the stiff portion is withdrawn from the body after implantation is completed. | 1. A guidewire element comprising: a rigid proximal portion;a flexible distal tether portion that is less rigid than the proximal portion;an attachment portion located between the flexible distal tether portion and the rigid proximal portion that releasably attaches the flexible distal tether portion to the rigid proximal portion; andwherein the flexible distal tether portion is configured to be used as a tensioning element and to be detached after tensioning. 2. The guidewire element of claim 1, wherein the flexible distal tether portion has a first torsional stiffness and the rigid proximal portion has a second torsional stiffness that is higher than the first torsional stiffness. 3. The guidewire element of claim 1, wherein a yield force of the rigid proximal portion as determined by compression testing is greater than a yield force of the flexible distal tether portion as determined by compression testing. 4. The guidewire element of claim 1, wherein the rigid proximal portion comprises a first portion and a second portion, wherein the first portion is more rigid than the second portion. 5. The guidewire element of claim 1, wherein the attachment portion comprises a first clasp attached to the rigid proximal portion, a second clasp that interfits with the first clasp attached to the flexible distal tether portion, and a pull-wire slidably insertable within the first and second clasps, wherein in a first protracted configuration the pull-wire secures the engagement between the first and second clasps, and in a second retracted configuration the pull-wire releases the engagement between the first and second clasps. 6. The guidewire element of claim 1, wherein the attachment portion comprises a temperature-sensitive polymer segment that is attached between the rigid proximal portion and the flexible distal tether portion, wherein the temperature-sensitive polymer segment is configured to weaken when heated. 7. The guidewire element of claim 1, wherein the attachment portion comprises a metal segment that is attached between the rigid proximal portion and the flexible distal tether portion, wherein the metal segment is configured to weaken when heated. 8. The guidewire element of claim 7, wherein the attachment portion comprises a metal segment susceptible to electrolytic dissolution that is attached between the rigid proximal portion and the flexible distal tether portion, wherein the metal segment is configured to dissolve when subjected to an electrical charge. 9. The guidewire element of claim 1, wherein the attachment portion comprises a plug attached to the flexible distal tether portion, wherein the rigid proximal portion comprises a lumen that is sized and shaped to frictionally retain the plug. 10. The guidewire element of claim 1, wherein the attachment portion comprises a screw joint, wherein the external thread of the screw joint is attached to the rigid proximal portion and the internal thread of the screw joint is attached ... | Ancora Heart, Inc. | A61B005/00 A61B017/08 A61M025/00 A61M025/09
US10702230 | B2 | 2020-07-07 | Method and system for generating a phase contrast image | 2018-06-29 | A system and method for generating a phase contrast image based on an absorption contrast image are provided. The method may include obtaining an absorption contrast image of a first object. The method may further include obtaining a phase contrast image generation model, wherein the phase contrast image generation model may be associated with at least one sample absorption contrast image and at least one sample phase contrast image of a second object. The method may further include executing the phase contrast image generation model to generate a phase contrast image of the first object based on the absorption contrast image of the first object. | 1. A system for generating a phase contrast image, comprising: a radioactive scanning source configured to emit radioactive rays toward a first object;a detector configured to detect at least some of the radioactive rays emitted from the radioactive scanning source impinging on the detector and to generate a plurality of readings based on the detected radioactive rays;a storage device storing a set of instructions, andat least one processor configured to communicate with the storage device, wherein when the at least one processor executes the set of instructions, the system is directed to: reconstruct an absorption contrast image of the first object based on the plurality of readings;obtain a phase contrast image generation model, wherein the phase contrast image generation model is trained based on at least one sample absorption contrast image and at least one sample phase contrast image of at least one second object; andexecute the phase contrast image generation model to generate a phase contrast image of the first object based on the absorption contrast image of the first object. 2. The system of claim 1, wherein the absorption contrast image of the first object and the at least one sample absorption contrast image of the second object are obtained by scanning with a same scanning parameter and image reconstruction based on a same reconstruction parameter. 3. The system of claim 1, wherein the phase contrast image generation model is a neural network model. 4. The system of claim 1, wherein the system is further directed to: obtain a dark field image generation model; andexecute the dark field image generation model based on the absorption contrast image or the phase contrast image of the first object to generate a dark field image of the first object,wherein the dark field image generation model is trained based on the at least one sample absorption contrast image of the second object and at least one sample dark field image of the at least one second object or based on the at least one sample phase contrast image of the at least one second object and the at least one sample dark field image of the at least one second object. 5. The system of claim 1, wherein the phase contrast image generation model is trained by a model training process, the model training process comprising: obtaining a preliminary model;obtaining at least one pair of training images of the at least one second object including the at least one sample absorption contrast image and the at least one sample phase contrast image of the at least one second object, wherein each of the at least one pair of the training image includes a sample absorption contrast image and a corresponding sample phase contrast image; andtraining the preliminary model based on the at least one pair of training images to generate a phase contrast image generation model. 6. The system of claim 5, wherein the obtaining the at least one pair of training images of the at least one second object com... | SHANGHAI UNITED IMAGING HEALTHCARE CO., LTD. | A61B006/00 A61B005/00 A61B006/04 A61B006/03
US10722214 | B2 | 2020-07-28 | Ultrasonic device, ultrasonic probe, and ultrasonic apparatus | 2017-12-21 | An ultrasonic device includes a substrate that is provided with an opening and a partition wall surrounding the opening, a vibration portion that closes one end side of the opening, and a piezoelectric element that is provided in the vibration portion, in which, in a case where an opening width dimension of the opening is indicated by W, and a thickness dimension of the substrate is indicated by H, a ratio W/H is 0.66 to 0.92. | 1. An ultrasonic device comprising: a substrate having a through hole and a partition wall, the partition wall surrounding the through hole, the through hole having first and second openings opposite to each other;a vibration plate having first and second surfaces opposite to each other, the first surface of the vibration plate closing the first opening of the through hole;a piezoelectric element that is provided on the second surface of the vibration plate;an acoustic layer filled into the through hole; andan acoustic lens disposed on the substrate, the acoustic lens contacting an end of the partition wall located next to the second opening,wherein the ultrasonic device is configured to transmit an ultrasonic wave toward a target object, and the ultrasonic device is configured to receive a reflected ultrasonic wave that is formed by reflecting the ultrasonic wave off of the target object and to output a received signal corresponding to a received voltage, andwherein, where when an opening width dimension of the first opening is indicated by W, a thickness dimension of the substrate is indicated by H, and a ratio W/H is 0.66 to 0.92, a normalized value of the received voltage is in a range of 0.8 to 1.0. 2. An ultrasonic device comprising: a substrate having a through hole and a partition wall, the partition wall surrounding the through hole, the through hole having first and second openings opposite to each other;a support film having first and second surfaces opposite to each other, a vibration portion of the first surface closing the first opening of the through hole;a beam having first and second ends, the first end being fixed to a first position of the second surface of the support film, the partition wall and the beam being overlapped each other at the first positon in a plan view;a piezoelectric element that is provided on a second position of the second surface of the support film, the second positon of the second surface corresponding to the vibration portion of the first surface of the support film;an acoustic layer filled into the through hole; andan acoustic lens disposed on the substrate, the acoustic lens contacting an end of the partition wall located next to the second opening,wherein the ultrasonic device is configured to transmit an ultrasonic wave toward a target object, and the ultrasonic device is configured to receive a reflected ultrasonic wave that is formed by reflecting the ultrasonic wave off of the target object and to output a received signal corresponding to a received voltage, andwherein, when a width dimension of the vibration portion is indicated by S, a thickness dimension of the substrate is indicated by H, and a ratio S/H is 0.73 to 1.16, a normalized value of the received voltage is in a range of 0.8 to 1.0. 3. An ultrasonic probe comprising: the ultrasonic device according to claim 1; anda casing in which the ultrasonic device is stored. 4. An ultrasonic probe comprising: the ultrasonic device according t... | Seiko Epson Corporation | B06B001/06 A61B008/00 A61B008/15 A61B008/08 A61B005/00 A61B008/13
US10716536 | B2 | 2020-07-21 | Identifying anatomical structures | 2016-03-07 | Aspects described herein disclose devices, systems, and methods for use in contexts such as minimally invasive surgery (MIS). A device is provided herein having a proximal portion and a distal portion, and an ultrasound transducer may be disposed within the distal portion and configured to scan tissue and identify certain portions of a patent's anatomy during the scanning process. The results of the detection may be presented to an operator of the device aurally and/or visually, such as in a 3-D volumetric image. By scanning the tissue, identifying the anatomy, and presenting the results to an operator, unnecessary damage to elements of the patients anatomy may be avoided or lessened. In some aspects, multiple transducers may be positioned on the device to increase the scanning range and/or scanning accuracy of the device. The device may provide an inner channel for the passage of surgical tools while scanning tissue. | 1. A system for mapping anatomical tissue, comprising: a device having a shaft, a housing attached to a distal portion of the shaft, and at least one ultrasound transducer disposed in the housing, wherein the at least one ultrasound transducer is configured to provide signal data corresponding to detected tissue in proximity to the at least one ultrasound transducer; anda computing device configured to: identify an anatomical structure by analyzing the signal data by execution of a classification algorithm, wherein the classification algorithm identifies the anatomical structure by comparing a b-mode scan line from the signal data to a plurality of known b-mode scan lines, each known b-mode scan line having been associated with a corresponding type of anatomical structure; andgenerate, for display, a 2-D mapping or a 3-D volumetric image depicting the anatomical structure,wherein by being configured to identify the anatomical structure by analyzing the signal data by execution of the classification algorithm, the computing device is configured to: threshold a b-mode image of the signal data to generate a binary map of intensity values;after thresholding the b-mode image, filter the b-mode image with a smoothing filter;identify contiguous regions in the filtered b-mode image having a pixel count above a minimum threshold and below a maximum threshold;compare the identified contiguous regions to an area of an ellipse outlining the contiguous regions; andclassify the identified contiguous regions as a nerve based on the comparison. 2. The system of claim 1, wherein the shaft further comprises a channel, through the shaft from the distal portion to a proximal portion, wherein the channel is configured to allow passage of a surgical tool through the shaft there through. 3. The system of claim 1, wherein the detected tissue comprises nerve tissue and wherein the 2-D mapping comprises a mapping of the nerve tissue. 4. The system of claim 1, wherein the detected tissue comprises nerve tissue and wherein the 3-D volumetric image comprises a 3-D volumetric image depicting a mapping of the nerve tissue. 5. The system of claim 1 further comprising a glove having at least one finger, wherein the at least one ultrasound transducer is disposed on the at least one finger. 6. A method comprising: receiving, at a computing device, data from at least one ultrasound transducer disposed at a distal portion of a body, wherein the at least one ultrasound transducer is configured to scan a region extending away from a main portion of the body;analyzing, by the computing device by execution of a classification algorithm, the data to identify an anatomical structure located within the region, wherein the classification algorithm identifies the anatomical structure by comparing a b-mode scan line from the data to a plurality of known b-mode scan lines, each known b-mode scan line having been associated with a corresponding type of anatomical structure; andoutputting, ... | Tissue Differentiation Intelligence, LLC | A61B008/12 A61B008/00 A61B008/08 G06T007/00 A61B017/02 A61B005/00 G06K009/62 A61B034/10 G06T007/60 G06K009/46 G06K009/00 A61B006/00 A61B017/00 A61B042/10
US10709334 | B2 | 2020-07-14 | Intraoperative guidance system for tumor surgery | 2017-10-09 | The current invention pertains to a system and methods of identifying a site in a tissue of a patient as neoplastic or normal. The system comprises a source of electromagnetic signals; an optical probe which delivers the electromagnetic signals to a working end of the probe; a spectrometer which acquires diffuse reflectance electromagnetic signals returned from the site probed by the optical probe. The spectrometer processes the diffuse reflectance signals to produce a diffuse reflectance spectra which is transmitted to a system controller programmed to analyze the diffuse reflectance spectra to calculate hemoglobin concentration, hemoglobin oxygenation, and/or diffuse reflectance intensity of signals having wavelength of about 700 nm. These parameters are used to identify the site as neoplastic or normal. The system of current invention can be used in identifying neoplastic sites in brain in an intraoperative manner, for example, during a pediatric brain surgery. | 1. A system for identifying a site in a tissue of a patient as neoplastic or normal, the system comprising: a) a source of electromagnetic signals;b) a probe operably connected to the source of electromagnetic signals, for delivering the electromagnetic signals to a plurality of known normal reference tissue sites and a suspected neoplastic tissue site of the patient;c) a spectrometer that acquires diffuse reflectance electromagnetic signals returned from the tissue sites and produces a diffuse reflectance spectra from the returned diffuse reflectance electromagnetic signals; andd) a system controller operably connected to the spectrometer, wherein the system controller: i) analyzes the diffuse reflectance spectra for the suspected neoplastic tissue site of the patient to calculate the following parameters: a) absolute hemoglobin concentration, normalized to mean hemoglobin concentration normal measurements to obtain n[Hb],b) absolute hemoglobin oxygenation, normalized to mean hemoglobin oxygenation normal measurements to obtain nSatO,c) absolute diffuse reflectance intensity of signals having wavelength of about 700 nm, normalized to mean normal Rd(700) measurements to obtain nRd(700),d) relative ratio of absorption coefficient μa and reduced scattering coefficient μs between 450 nm and 650 nm, normalized to mean normal μa(λ)/μs(λ) measurements to obtain nμa(λ)/μs(λ),ii) analyzes the diffuse reflectance spectra from the plurality of known normal reference tissue sites to calculate the following respective standard deviations for each of the parameters: a) absolute hemoglobin concentration, normalized to an average hemoglobin concentration from known normal reference tissue measurements to obtain n[Hb]N_std,b) absolute hemoglobin oxygenation, normalized to an average hemoglobin oxygenation from known normal reference tissue measurements to obtain nSatO2N_std,c) absolute diffuse reflectance intensity of signals having wavelength of about 700 nm, normalized to an average normal Rd(700) from known normal reference tissue measurements to obtain nRd(700)N_std,d) relative ratio of absorption coefficient μa and reduced scattering coefficient μs between 450 nm and 650 nm, normalized to an average normal μa(λ)/μs(λ) from known normal reference tissue measurements to obtain nμa(λ)/μs(λ)/μs(λ)N_std, andiii) analyzes n[Hb], nSatO2, nRd(700), and/or nμa(λ)/μs(λ) calculated from the patient suspected neoplastic tissue site with respect to the standard deviations of those same parameters calculated from the plurality of known normal reference tissue sites of the same patient,wherein the suspected neoplastic tissue site is identified by the system controller if: nμa(λ)/μs(λ)>1+x×nμa(λ)/μs(λ)N_std or nμa(λ)/μs(λ)&lt;1−x×μna(λ)/μs(λ)N_std,wherein x is an adjustable cutoff for nμa(λ)/μs(λ). 2. The system of claim 1, wherein the electromagnetic signals have wavelengths between about 500 nm to about 900 nm, about 600 nm to about 850 nm, or about 650 nm to about 800 ... | The Florida International University Board of Trustees | A61B005/00 A61B005/1455 A61B034/20 A61B090/13
US10722178 | B2 | 2020-07-28 | Method and apparatus for motion correction in CT imaging | 2017-05-30 | A method for the reduction of motion artifacts in CT imaging, the method including the steps of: (a) reconstructing from the raw CT data an initial estimate of the object of interest; (b) estimating the pose of the object at each projection angle; (c) undertaking a motion-corrected reconstruction from the measured projections, accounting for the pose changes estimated in (b) by employing a modified source/detector orbit that accounts for the object motion; (d) iterating steps (b)-(c) until a predetermined convergence criterion is met; and (e) making a final reconstruction of the required image size using the pose estimates obtained in the previous steps. | 1. A method for a reduction of motion artifacts in computed tomography (CT) reconstruction imagery, the method including the steps of: (a) reconstructing, from a CT data set comprising measured projections acquired by a detector at a plurality of angles around an object of interest, an initial estimate of a 3D attenuation distribution of the object of interest;(b) at each angle of the CT data set, a forward projection of the estimated 3D attenuation distribution within the object of interest and the CT projection measured at the same projection at the same angle are used to estimate a pose of the object of interest;(c) undertaking a further reconstruction from the measured projections, accounting for a pose changes between two successive pose estimates from (b) to produce a motion corrected estimate of the 3D attenuation distribution;(d) iterating steps (b)-(c) until the pose estimates obtained at all angles in two successive iterations differ by less than a predetermined convergence criterion; and(e) producing a final image reconstruction of a desired image size using a final pose estimates obtained at an end of step (d). 2. The method as claimed in claim 1, wherein said steps (b) and (c) include the minimization of a cost function including a summation of differences between the measured projections and the corresponding forward projections of the motion corrected estimate of the 3d attenuation distribution at all projection angles. 3. The method as claimed in claim 1, further comprising performing the method across multiple resolutions of the CT data set and image by increasing a sampling of the CT data with further iterations of steps (b) to (c). 4. The method as claimed in claim 3, wherein the multiple resolutions provide an increasing resolution for successive estimates of the 3D attenuation distribution. 5. The method as claimed in claim 1, wherein the pose includes three dimensional rotation and translation of the object of interest. 6. A method according to claim 5, wherein a coordinate system for the pose is with respect to the detector. 7. The method as claimed in claim 6, wherein one of said translations is perpendicular to the detector and the perpendicular translation is neglected in the pose estimate. 8. The method as claimed in claim 1, wherein said step (c) includes applying an ordered subset expectation maximization (OSEM) image reconstruction algorithm to the measured projections of the CT data, making use of the pose changes to estimate an arbitrary orbit used in the OSEM. 9. A method according to claim 1, wherein the computed tomography (CT) includes at least one of helical CT, circular orbit CT, cone beam CT and dental CT. 10. A method according to claim 1, wherein said step (c) includes applying an image reconstruction algorithm that makes use of the pose changes to estimate a modified orbit to obtain a motion corrected image reconstruction from the measured CT projections, or applying an ordered subset expectation... | Western Sydney Local Health District The University of Sydney Katholieke Universiteit Leuven KU Leuven R&D | A61B006/00 A61B005/00 G06F003/0484
US10723306 | B2 | 2020-07-28 | Curtain airbag device mounting structure and curtain airbag deployment method | 2018-06-21 | A curtain airbag device mounting structure includes: a first pillar forming a part of a front pillar and extends substantially along a vehicle height direction; a second pillar forming another part of the front pillar, the second pillar being disposed on a rear side of a vehicle relative to the first pillar at a predetermined distance from the first pillar and extending substantially along the vehicle height direction; a transparent member bridged between the first pillar and the second pillar; and a curtain airbag device including a curtain airbag stored along a roof side rail and the second pillar, the curtain airbag being configured to inflate and deploy in a curtain-like fashion over a side portion of a cabin of the vehicle in case of a collision of the vehicle. | 1. A curtain airbag device mounting structure comprising: a first pillar that is a part of a front pillar and extends substantially along a vehicle height direction;a second pillar that is another part of the front pillar, the second pillar being disposed on a rear side of a vehicle relative to the first pillar at a predetermined distance from the first pillar and extending substantially along the vehicle height direction;a transparent member that is bridged between the first pillar and the second pillar and enables an outside of the vehicle to be viewed from a driver seat;a curtain airbag device including a curtain airbag stored along a roof side rail and the second pillar, the curtain airbag being configured to be supplied with gas to inflate and deploy in a curtain-like fashion over a side portion of a cabin of the vehicle in case of a collision of the vehicle; anda front pillar garnish attached to a vehicle interior side of the front pillar, the front pillar garnish including a front-side garnish and a rear-side garnish, the front-side garnish being attached to the first pillar, the rear-side garnish being attached to the second pillar, the rear-side garnish being configured to cover the curtain airbag from a vehicle interior side of the curtain airbag and to hold the curtain airbag in a state in which the curtain airbag is apart from the second pillar. 2. The curtain airbag device mounting structure according to claim 1, wherein the rear-side garnish is attached to the second pillar via a tether clip. 3. The curtain airbag device mounting structure according to claim 2, wherein: the front-side garnish includes an extension portion at an upper end portion of the front-side garnish in the vehicle height direction, the extension portion extending toward the rear side of the vehicle, and including a rear end portion covered by an upper end portion of the rear-side garnish in the vehicle height direction; andthe tether clip is attached to the extension portion at a position corresponding to the extension portion in the rear-side garnish. 4. The curtain airbag device mounting structure according to claim 1, wherein each of the front-side garnish and the rear-side garnish has a U-shape in a section orthogonal to a longitudinal direction of the front pillar andat least one of a wire harness and a hose is held between the front-side garnish and the first pillar, and a part of the curtain airbag device is stored between the rear-side garnish and the second pillar. 5. The curtain airbag device mounting structure according to claim 1, wherein the second pillar includes a reaction force surface at a part on a vehicle interior side of the second pillar, the part of the second pillar facing the curtain airbag, a thickness direction of the reaction force surface being substantially along a vehicle width direction. 6. The curtain airbag device mounting structure according to claim 5, wherein the second pillar includes: a pillar inner panel that is a par... | TOYOTA JIDOSHA KABUSHIKI KAISHA | B60R021/232 B60R021/213 B60R021/2338 B23K020/00 B22D011/06 E04G013/02 B01J019/00 E05B083/38 A47F005/04 B62D025/04 B60J005/04 B65D051/16 D06F057/02 E02F009/08 B25D017/30 E04B002/52 F16M011/28 G02B006/44 B60R021/21 H01L021/768 H01L029/739 B63B003/52 H01L021/762 A47F011/02 B21D047/01 E04B002/38 A61B005/151 H01H035/30 B60R013/02 H05K007/20 B62J001/08 E05C007/04 G02B023/18 G05D023/12 A46B011/00 B01J029/04 F16K099/00 G09F015/00 B62K019/36 B62D025/20 B29C065/54 A61B005/15 B60H001/24 A61M005/14 A47F005/05 B66F007/02 G02F001/1339 B60J010/777 B62D027/06 G02B007/12 B60H001/26 B60J007/00 E04B002/24 H01L045/00 A47F005/10 E04C003/30 F16C033/74 B23D007/04 B62D025/14 E04B002/56 A61B005/02 B65G041/00 E04B002/72 B21D037/10 E04C001/00 B60R011/00 B60R021/04
US10717447 | B2 | 2020-07-21 | Visual perception assistance system and visual-perception target object detection system | 2016-12-16 | The present invention includes: a visual perception target determination unit for determining a visual perception target by using a saliency model for determining that an object to be visually perceived at a glance is a visual perception target, a surprise model for determining that an object behaving abnormally is a visual perception target, and a normative model for determining that an object to be visually perceived by a viewing action of a driver serving as a norm is a visual perception target; and a visual guidance unit for determining whether or not an overlooked visual perception to which a line-of-sight direction detected by a line-of-sight detection unit is not directed is present, and guiding a line-of-sight of a driver toward the overlooked visual perception target, when the overlooked visual perception target is present. | 1. A visual perception assistance system, comprising: a surrounding condition acquisition unit for acquiring a surrounding condition of a moving body to be driven by a driver;a visual perception target determination unit for determining a visual perception target being an object that the driver should look at in a surrounding condition acquired by the surrounding condition acquisition unit;a line-of-sight direction detection unit for detecting a line-of-sight direction of the driver; anda visual guidance unit for determining whether or not an overlooked visual perception target to which a line-of-sight direction detected by the line-of-sight direction detection unit is not directed is present among the visual perception target determined by the visual perception target determination unit, and guiding a line-of-sight of the driver toward the overlooked visual perception target, when the overlooked visual perception target is present, whereinthe visual perception target determination unit determines the visual perception target, based on three visual perception models prepared in advance, andthe three visual perception models include a saliency model for determining that an object to be visually perceived at a glance is a visual perception target, a surprise model for determining that an object behaving abnormally is a visual perception target, and a normative model for determining that an object to be visually perceived by a viewing action of a driver serving as a norm is a visual perception target. 2. The visual perception assistance system according to claim 1, wherein the moving body is constituted by a vehicle, andthe surrounding condition acquisition unit is constituted by a camera for photographing an area ahead of the vehicle. 3. The visual perception assistance system according to claim 1, wherein the visual guidance unit performs visual guidance by displaying an index at a position of the overlooked visual perception target. 4. The visual perception assistance system according to claim 1, further comprising: a driver state detection unit for detecting the driver state, whereinthe visual guidance unit changes a degree of intensity of visual guidance depending on a driver state to be detected by the driver state detection unit. 5. The visual perception assistance system according to claim 4, wherein the visual guidance unit changes the degree of intensity of visual guidance by changing a degree of conspicuousness of an index to be displayed at a position of the overlooked visual perception target. 6. The visual perception assistance system according to claim 4, wherein the driver state detection unit detects at least a state that driving load is large as the driver state, andwhen the driver state detection unit detects that the driving load is large, the visual guidance unit emphasizes visual guidance, as compared with a case where the driver state detection unit detects that the driving load is small. 7. The visual perception assista... | MAZDA MOTOR CORPORATION | G06K009/00 B60W050/14 G02B027/00 B60R001/00 G02B027/01 G08G001/0962 G08G001/16 A61B005/18 B60K035/00 B60Q009/00
US10702152 | B2 | 2020-07-07 | Impact monitoring system for players engaged in a sporting activity | 2017-04-10 | The invention relates to a protective sports equipment system with real-time impact monitoring of a player engaged in a contact sporting activity. The system includes a protective equipment assembly and an interface liner disposed within a padding assembly of the protective equipment assembly. The interface liner has a sensing and control unit with an array of sensors to detect impact data, a processor to determine an impact result, and a transmitter to transmit the impact result. The system also includes a remote unit that receives and displays the transmitted impact result. | 1. A protective sports equipment system with real-time impact monitoring, the protective sports equipment worn by a player to protect a body part of the player while the player is engaged in a contact sporting activity, the protective sports equipment system comprising: a protective equipment assembly including a protective outer layer and a padding assembly that resides inward of the protective outer layer;an interface liner removably positioned within the padding assembly of the protective equipment assembly, the interface liner configured to reside against the player's body part while the protective equipment assembly is worn, the interface liner including a sensing and control unit having a plurality of sensors arrayed within the interface liner that detect impact data resulting from an impact to the protective equipment assembly, the interface liner further including a processor that compares the impact data against at least one pre-loaded impact profile to determine an impact result, and a transmitter to transmit said impact result; and,a remote unit operably connected to the sensing and control unit, the remote unit having a receiver and configured to receive and display said impact result transmitted from the sensing and control unit. 2. The protective sports equipment system of claim 1, wherein the sensing and control unit includes at least one filter that provides signal conditioning for the impact data. 3. The protective sports equipment system of claim 1, wherein the processor further determines a direction of the impact to the protective equipment assembly. 4. The protective sports equipment system of claim 1, wherein the interface liner has a flexible multi-layer construction with at least one padding element within an outer liner housing. 5. The protective sports equipment system of claim 1, the system further comprises a graphical user interface operably connected to the remote unit for immediately displaying at least a magnitude of the impact to the protective equipment assembly. 6. The protective sports equipment system of claim 1, wherein the plurality of sensors are arrayed within the interface liner about the periphery of the player's body part while the protective equipment is worn. 7. The protective sports equipment system of claim 6, wherein the plurality of sensors are configured to sense linear acceleration and rotational acceleration of the player's body part resulting from the impact to the protective equipment assembly. 8. The protective sports equipment system of claim 6, wherein the plurality of sensors are accelerometers configured to sense linear acceleration and rotational acceleration of the player's body part resulting from the impact to the protective equipment assembly. 9. The protective sports equipment system of claim 6, wherein the plurality of sensors are tri-axial accelerometers. 10. A protective sports helmet system with real-time impact monitoring, the protective sports helmet system comprisin... | Riddell, Inc. | A61B005/00 A42B003/04 A63B071/06 A63B071/10 A61B005/11 A42B003/12
US10702164 | B2 | 2020-07-07 | Identification of areas of interest during intraoral scans | 2018-09-28 | A processing device performs image registration between a plurality of intraoral images of a dental site. The processing device identifies a candidate intraoral area of interest from a first intraoral image of the plurality of intraoral images. The processing device verifies the candidate intraoral area of interest as an intraoral area of interest based on a comparison of a second intraoral image to the first intraoral image. The processing device displays a view of the dental site where the intraoral area of interest is hidden and an indication of the hidden intraoral area of interest is visible. | 1. A computer readable medium comprising instructions that, when executed by a processing device, cause the processing device to perform operations comprising: performing image registration between a plurality of intraoral images of a dental site;identifying a candidate intraoral area of interest from a first intraoral image of the plurality of intraoral images;making a comparison of a second intraoral image to the first intraoral image;determining whether to verify the candidate intraoral area of interest as an intraoral area of interest or to dismiss the candidate intraoral area of interest based on a result of the comparison;verifying the candidate intraoral area of interest as an intraoral area of interest; anddisplaying a view of the dental site where the intraoral area of interest is hidden and an indication of the hidden intraoral area of interest is visible. 2. The computer readable medium of claim 1, the operations further comprising: determining a value associated with the candidate intraoral area of interest;determining a threshold; andverifying the candidate intraoral area of interest as an intraoral area of interest responsive to determining that a) a second intraoral image confirms the candidate intraoral area of interest and b) the value is above the threshold. 3. The computer readable medium of claim 2, wherein determining the threshold comprises selecting the threshold based on one or more patient case details, wherein the one or more patient case details comprise a procedure for the dental site. 4. The computer readable medium of claim 1, wherein the candidate intraoral area of interest has a first shape and a first size, the operations further comprising: making a determination that a portion of the candidate intraoral area of interest corresponds to a portion of a surface of the second intraoral image based on the comparison of the second intraoral image to the first intraoral image; andmodifying the candidate intraoral area of interest so that it has at least one of a second shape or a second size that is smaller than the first size. 5. The computer readable medium of claim 1, wherein the candidate intraoral area of interest comprises a void in the first intraoral image, the operations further comprising: interpolating a shape for the void based on geometric features surrounding the void in at least one of the first intraoral image or the second intraoral image, wherein the indication of the intraoral area of interest comprises the shape of the void. 6. The computer readable medium of claim 1, wherein the first intraoral image and the second intraoral image are received during an intraoral scan session, and wherein the indication of the intraoral area of interest is provided during the intraoral scan session. 7. The computer readable medium of claim 6, the operations further comprising: computing a virtual model of the dental site after the intraoral scan session is complete based on a plurality of intraoral images, the... | Align Technology, Inc. | A61B005/00 A61C007/00 A61C013/34 A61C009/00 A61B006/14 A61B006/00 G06T007/55 G06T007/579 A61B001/24 G06T007/33 A61B006/03
US10702141 | B2 | 2020-07-07 | Automated perimeter | 2014-09-02 | The invention relates to a perimeter or a campimeter with a visible fixation point and a method used in them. The method comprises at least the following steps: producing a fixation point having a first visual appearance to be shown to a patient;producing a stimulus shown to the patient at a stimulus time-point at a pre-defined location;activating a response device by the patient upon noticing the stimulus at a response time-point;changing the fixation point to have a second visual appearance for a fixation point second visual appearance time interval near the stimulus time-point. | 1. A method used in a standard automated perimetry (SAP) perimeter or campimeter with a visible fixation point, the method comprising at least the following steps: producing a fixation point (101) having a first visual appearance (101a) to be shown to a patient (100);producing a stimulus (105) shown to the patient at a stimulus time-point at a pre-defined location;activating a response device (103) by the patient upon noticing the stimulus at a response time-point;further comprisingchanging the fixation point to have a second visual appearance (101b) for a fixation point second visual appearance time interval near the stimulus time-point by starting the fixation point second visual appearance time interval not earlier than 1 second before the stimulus time-point; andchanging the fixation point back to have the first visual appearance (101a) after the fixation point second visual appearance time interval has ended or after patient has responded. 2. A method according to claim 1, further comprising using ambient light sensors (15) to adjust luminance of the stimuli and fixation points during a test based on illumination of perimeter surface so that luminance contrast can be set at a desired level. 3. A method according to claim 1, wherein starting the fixation point second visual appearance time interval less than 200 ms before the stimulus time-point, and/orending the fixation point second visual appearance time interval not later than 2 seconds after the stimulus time-point, or after the fixation point second visual appearance is shown, or after patient has responded. 4. A method according to claim 1, further comprising recording whether the patient activates the response device during a response time interval or not. 5. A method according to claim 1, characterized in producing a second stimulus shown to the patient at a second stimulus time-point after a stimulus time interval from either the previous stimulus time-point or the previous response time-point. 6. A method according to claim 5, characterized in varying one or more of: the stimulus time intervallight intensity of the stimulusstimulus sizestimulus colorstimulus location. 7. A method according to claim 5, characterized in varying the position of the fixation point shown to the patient. 8. A method according to claim 1, characterized in changing the fixation point to have a third visual appearance for a time interval near the stimulus time-point when the patient should not activate the response device even if he/she saw the stimulus. 9. A standard automated perimetry (SAP) perimeter or campimeter with a visible fixation point comprising means for producing a fixation point (101) having a first visual appearance (101a) to be shown to a patient (100);means for producing a stimulus (105) shown to the patient at a stimulus time-point;a response device (103) adapted to be activated by the patient upon noticing the stimulus at a response time-point;characterized in that the means for ... | Ocuspecto Oy | A61B003/02 A61B003/024 A61B005/16 A61B003/00 A61B003/028 A61B003/032 A61B003/06 A61B005/00 A61B003/18 H04N005/58
US10706262 | B2 | 2020-07-07 | Intelligent body measurement | 2018-01-08 | A method for estimating the body measurements of a subject from at least two photographic images of the subject. The method includes capturing the photographic images using a digital imaging device, for example, a mobile device camera, and estimating the body measurements of the subject using heatmap data generated by an intelligent computing system such as a trained neural network. | 1. A method for estimating body measurements of a subject, comprising: receiving at least two images containing the subject;generating landmark heatmap data from the at least two images containing the subject;relating the landmark heatmap data to an abstract three-dimensional body model to determine a shape of the subject's body,wherein relating the landmark heatmap data to the abstract three-dimensional body model to determine the shape of the subject's body comprises extracting a plurality of contour points from the landmark heatmap data, and wherein the plurality of contour points comprises at least one control point,generating the abstract three-dimensional body model using the plurality of contour points, wherein the abstract three-dimensional body model comprises a plurality of body slices, and wherein each slice of the plurality of body slices represents an anthropometric measurement of a body that influences body shape,projecting each body slice of the plurality of body slices as a corresponding two-dimensional landmark,generating at least one main cross-section value by matching the at least one control point to the corresponding two-dimensional landmark, wherein the at least one main cross-section value is used to construct a cross-section curve,extracting a vertex from the cross-section curve, andconstructing a surface of a part of the subject's body using the vertex from the cross-section curve; andestimating the body measurements of the subject based on the shape of the subject's body. 2. The method of claim 1, wherein the at least two images containing the subject are generated from a digital imaging device. 3. The method of claim 2, wherein the digital imaging device comprises a mobile device camera. 4. The method of claim 1, wherein the landmark heatmap data is generated using an intelligent computing system. 5. The method of claim 4, wherein the intelligent computing system comprises a trained neural network. 6. The method of claim 1, wherein generating the landmark heatmap data from the at least two images containing the subject comprises: generating a first portion of the landmark heatmap data using a first intelligent computing system, wherein the first intelligent computing system is trained to analyze front-view photographic images, wherein one of the at least two images containing the subject is a front-view image containing the subject, and wherein the first intelligent computing system analyzes the front-view image containing the subject to generate the first portion of the landmark heatmap data; andgenerating a second portion of the landmark heatmap data using a second intelligent computing system, wherein the second intelligent computing system is trained to analyze side-view photographic images, wherein another of the at least two images containing the subject is a side-view image containing the subject, and wherein the second intelligent computing system analyzes the side-view image containing the subject to gene... | 3DLOOK Inc. | G06K009/00 A61B005/01 G06K009/46 G06K009/66 A61B005/107 G06K009/62
US10722720 | B2 | 2020-07-28 | Methods and systems for improved communication between medical devices | 2018-06-21 | At least one of a first medical device and a second medical device may be implanted within a patient while the second medical device may optionally be proximate but external to the patient. At least one of the medical devices has an antenna having at least two electrodes and at least one of the medical devices has an antenna having at least three electrodes. The medical devices can communicate via conducted communication through the patient's tissue between a first pair of electrodes and a second pair of electrodes. At least one of the pairs of electrodes can be selected in accordance with the signal strength of the communication vector between the first and second pairs of electrodes. | 1. A communication system, comprising: a first medical device having at least two electrodes that are associated with the first medical device;a second medical device including a plurality of electrodes defining at least a first communication vector using a first pair of electrodes of the plurality of electrodes and a second communication vector using a second pair of electrodes of the plurality of electrodes;the second medical device configured to listen for messages from the first medical device via the first communication vector and to listen for the same messages from the first medical device via the second communication vector;the second medical device monitoring a first measure related to a signal-to-noise ratio of the conducted communication via the first communication vector and monitoring a second measure related to the signal-to-noise ratio of the conducted communication via the second communications vector; andthe second medical device comparing results of the conducted communication via the first communication vector to results of the conducted communication via the second communications vector, and when the results do not match, identifying the correct result based on one or more criteria. 2. The system of claim 1, wherein the second medical device is configured to transmit messages to the first medical device using the first communication vector, the second communication vector, or both the first communication vector and the second communication vector. 3. The system of claim 1, wherein the first medical device comprises an implantable medical device. 4. The system of claim 1, wherein the first medical device comprises a leadless cardiac pacemaker (LCP). 5. The system of claim 1, wherein the plurality of electrodes comprise four electrodes secured to a skin patch, the four electrodes arranged on the skin patch in a kite configuration having a first pair of equal-length adjacent sides and a second pair of equal-length adjacent sides that are of a different length than the first pair of equal-length adjacent sides, the skin patch being configured to be secured to a patient's skin such that each of the four electrodes is in electrical contact with the patient's skin and assume a fixed position relative to the patient for communicating with the first medical device. 6. The system of claim 5, wherein having the four electrodes arranged on the skin patch in a kite configuration provides a consistent and predictable set of communication vectors including the first communication vector and the second communication vector. 7. A method of communicating between a first medical device and a second medical device, the first medical device implanted within a patient and having at least two electrodes and the second medical device being associated with a skin patch that includes a plurality of electrodes, wherein the skin patch is secured to a patient's skin such that each of the plurality of electrodes of the skin patch are in electrical co... | Cardiac Pacemakers, Inc. | A61N001/372 A61B005/00 H04B013/00 A61N001/36 A61B005/0464 A61N001/362 A61N001/375
US10716531 | B2 | 2020-07-21 | Equipment comprising an acoustic measurement device comprising means for linking a sensor to a rigid structure | 2015-07-16 | The invention relates essentially to equipment (100) to be positioned against the body of a subject, said equipment including an acoustic measurement device (105) having at least one acoustic sensor (110) for measuring acoustic signals given off by the subject's body, wherein said measurement device (105) further comprises: a rigid structure (120) that is stationary relative to the equipment; andconnection means (450, 550, 650, 750, 850, 950) for connecting the acoustic sensor (110) to the rigid structure (120);;the connection means (450, 550, 650, 750, 850, 950) and the acoustic sensor (110) being movable relative to the rigid structure (120), and the connection means (450, 550, 650, 750, 850, 950) defining predetermined degrees of freedom of movement for the acoustic sensor (110). | 1. Equipment to be positioned against the body of a subject, the equipment including an acoustic measurement device having: at least one acoustic sensor that measures for measuring acoustic signals given off by the subject's body, the at least one acoustic sensor includes a first outside wall that faces the subject and a first inside wall;a rigid structure, the rigid structure includes an orifice that extends through the rigid structure, and a second outside wall that faces the first inside wall of the at least one acoustic sensor;in a cross-sectional side view, an axis of the orifice intersects a center of the at least one acoustic sensor;connection means that connects the at least one acoustic sensor to the rigid structure in such a manner that the connection means and the at least one acoustic sensor are movable relative to the rigid structure and the connection means defining predetermined degrees of freedom of movement for the at least one acoustic sensor relative to the rigid structure; andan element attached to the first inside wall of the at least one acoustic sensor, the element extending toward the rigid structure and through the orifice, and the at least one acoustic sensor is positioned entirely outside the orifice. 2. Equipment according to claim 1, wherein, for at least one predetermined degree of freedom, the movement of the at least one acoustic sensor is limited by the second outside wall of the rigid structure. 3. Equipment according to claim 2, wherein the movement of the at least one acoustic sensor is limited in rotation by a predetermined angle relative to the second outside wall of the rigid structure and/or in translation along a predetermined linear shift towards the rigid structure. 4. Equipment according to claim 1, wherein the connection means resiliently biases the at least one acoustic sensor in a direction away from the rigid structure. 5. Equipment according to claim 1, wherein the connection means comprise: a ball support fastened to the rigid structure; anda ball connected to the at least one acoustic sensor and positioned inside the ball support. 6. Equipment according to claim 1, wherein the connection means comprise: a spring having a first end positioned against the rigid structure and a second end fastened to the at least one acoustic sensor. 7. Equipment according to claim 1, wherein the connection means comprise: a spring having a first end and a second end;a ball support; anda ball positioned inside the ball support. 8. Equipment according to claim 7, wherein: the ball is fastened to the at least one acoustic sensor;the first end of the spring is fastened to the rigid structure; andthe second end of the spring is fastened to the ball support. 9. Equipment according to claim 7, wherein: the ball has an orifice passing the element connected to the at least one acoustic sensor;the ball support is fastened to the rigid structure;the first end of the spring is fastened to the rigid structure; andthe se... | H4D | A61B005/00 A61B007/02 A61B007/00 A61B007/04
US10716520 | B2 | 2020-07-21 | PET imaging device for observing the brain | 2018-10-24 | The invention relates to a PET imaging device for observing the brain, characterised by comprising a structure with a shape that is capable of accommodating a human head, having independent gamma-ray detection modules, said detection modules having continuous scintillation crystals with a polygonal main section, wherein all together the detection modules form a hollow three-dimensional structure that can surround the head, and with said three-dimensional structure being elongated and having a main axis in the direction corresponding to the forehead-nape direction and a shorter axis in the direction corresponding to the straight line joining the ears, and with the adjacent scintillation crystals fitting together laterally in a precise manner along their entire thickness, forming a mosaic-like structure, i.e. without leaving gaps and without overlapping with one another. | 1. A PET imaging device for observation of a brain, which has a hollow three-dimensional structure with a shape capable of housing a head, the PET imaging device comprising a plurality of independent gamma ray detection modules that together form a structure capable of surrounding the head, said detection modules comprise continuous scintillation crystals, wherein each of the continuous scintillation crystals has a polygonal main cross-section, and wherein said structure is an elongated structure having a major axis in a direction corresponding to a front-nape direction and a shorter axis in a direction corresponding to a straight line joining ears on the head, and the continuous scintillation crystals are positioned adjacent to each other are to fit laterally in an exact manner with each other throughout their entire thickness, building a mosaic without leaving gaps between adjacent crystals and without overlapping each other. 2. The PET imaging device according to claim 1, wherein the main cross-section of each of the scintillation crystals has a square or rectangular shape, and wherein the structure formed from the scintillation crystals is a hollow, rectangular base prism, an anterior surface of said prism, being in a front part of the structure, that can be confronted to a face of a subject, and a rear surface of the prism, being in a posterior part of the structure, that can be confronted to an occipital bone of the subject. 3. The PET imaging device according to claim 2, wherein each side of said prism is covered with detectors with a square main cross-section; an insertion side of said prism, the insertion side having an opening through which the head may be inserted, being covered only partially, such that the neck fits loosely; the anterior surface being covered only partially, such that vision of the subject is not obstructed, and the back base and all remaining sides of the prism are completely covered by the detection modules. 4. The PET imaging device according to claim 2, in which the rear base has been extended with additional detection modules towards an area corresponding to a back of the subject when the device is placed on the head of the subject. 5. The PET imaging device according to claim 2, wherein an insertion side of the prism is only partially covered with detection modules, wherein the insertion side forms with a lower portion of the anterior surface of the prism a fragment thereof in the form of “L”, and this L shaped fragment, that corresponds to a chin area of the subject when the device is in use, is adjustable in position for each subject after the PET imaging device has been placed on the head. 6. The PET imaging device according to claim 1, wherein the main cross-section of each of the scintillation crystals is square or rectangular shaped, and wherein the structure formed by the detection modules is a hollow prism with a base of pentagonal, hexagonal, or octagonal section, an anterior base of said prism be... | General Equipment for Medical Imaging, S.A. Consejo Superior de Investigaciones Cientificas Universidad Politècnica de València | A61B006/03 G01T001/164 G01T001/29 G06T007/00 A61B005/00
US10726702 | B2 | 2020-07-28 | Information processing system, information processing method, and program | 2016-07-05 | Provided is an information processing system capable of effectively preventing a user's failure to take his or her basal body temperature.;The base time setter 301 sets, as a base measurement time, a time used as a basis for a user to take the user's body temperature after awakening, based on the times at which the user's body temperatures were measured. The sleep duration estimator 302 estimates the user U's most recent sleep duration. When the user's most recent sleep duration exceeds a predetermined threshold value, the alarm controller 303 causes an alarm to sound at the base measurement time. | 1. An information processing system comprising: at least one memory configured to store computer program code; andat least one processor configured to access said computer program code and operate as instructed by said computer program code, said computer program code including: base time setting code configured to cause at least one of said at least one processor to, based on times at which body temperatures of a user were measured, set as a base measurement time, a time used as a basis for the user to take the user's body temperature after awakening;sleep duration estimation code configured to cause at least one of said at least one processor to estimate a first estimated sleep duration of the user; andalarm control code configured, at the base measurement time, to cause at least one of said at least one processor to: when the first estimated sleep duration of the user exceeds a predetermined threshold value, cause an alarm to sound at a first time, wherein the first time is the base measurement time, andwhen the first estimated sleep duration of the user does not exceed the predetermined threshold value, configure the alarm to not sound at the first time but to sound at a second time, wherein the second time is a time when a second estimated sleep duration of the user will exceed the predetermined threshold value. 2. The information processing system according to claim 1, wherein the sleep duration estimation code is configured to cause at least one of said at least one processor to estimate the first estimated sleep duration to be a time interval between a most recent operation of the user on a mobile terminal of the user and the base measurement time. 3. The information processing system according to claim 1, wherein when no operation is performed on a mobile terminal of the user for a predetermined duration of time, the sleep duration estimation code causes at least one of said at least one processor to further estimate that the user is in a sleeping state, andif an operation is performed on the mobile terminal before the base measurement time in the sleeping state, the alarm control code causes at least one of said at least one processor to cancel a process for causing the alarm to sound at the base measurement time. 4. The information processing system according to claim 1, wherein when no operation is performed on a mobile terminal of the user for a predetermined duration of time, the sleep duration estimation code causes at least one of said at least one processor to further estimate that the user is in a sleeping state, andwhen an operation is performed on the mobile terminal at a third time, wherein the third time is before the first time, the alarm control code causes at least one of said at least one processor to cause the alarm to sound if a third estimated sleep duration exceeds a predetermined threshold value at the first time, and cancels a process for causing the alarm to sound at the base measurement time if the third estim... | Rakuten, Inc. | G08B021/24 A61B005/00 G01K007/00 A61B005/01 G08B021/18
US10722142 | B2 | 2020-07-28 | Medical apparatus for the introduction of catheters into the human body | 2018-05-14 | A medical apparatus for the introduction of catheters into the human body including: an ECG device having at least three electrodes, positionable on a human body for the detection of surface electrical potential data and adapted to provide at output at least one surface ECG strip; at least one catheter element having a distal ending part positionable inside the human body and having an auxiliary electrode element for the detection of intracavitary electrical potential data, associated with the ECG device for sending to the latter said intracavitary data, the ECG device being adapted to provide at output at least one intracavitary ECG strip; —an electronic device associated with the ECG device for the receipt of at least said intracavitary ECG strip and comprising a data processing module (10) adapted to process at least the intracavitary ECG strip to obtain usage information adapted at least to provide instructions on the insertion of the catheter element in the human body. | 1. A medical apparatus for the introduction of catheters into a human body, comprising: an ECG device having at least three electrodes, the at least three electrodes being positionable on the human body for the detection of surface electrical potential data and configured to output at least one surface ECG strip;at least one catheter element having a distal ending part, the distal ending part being positionable inside the human body and having an auxiliary electrode element for the detection of intracavitary electrical potential data associated with the ECG device, the auxiliary electrode element being configured for sending the intracavitary electrical potential data to the ECG device, the ECG device being configured to output at least one intracavitary ECG strip; andan electronic device associated with the ECG device, the electronic device configured for receiving the at least one intracavitary ECG strip and comprising a data processor system configured to process the at least one intracavitary ECG strip to obtain usage information adapted at least to provide instructions on insertion of the at least one catheter element in the human body,wherein the data processor system comprises a first data processor associated with the ECG device for acquiring the at least one intracavitary ECG strip, the first data processor being configured to obtain recognition data providing analytical information on the at least one intracavitary ECG strip,wherein the first data processor comprises a first data processor unit receiving an input of the at least one intracavitary ECG strip, the first data processor unit being configured to obtain first recognition data of maximum and minimum points of a same ECG strip of the at least one intracavitary ECG strip, the maximum and/or minimum points corresponding to waves of the same ECG strip,wherein the first data processor comprises a second data processor unit receiving an input of at least the intracavitary ECG strip and the first recognition data, the second data processor unit being configured to obtain second recognition data of points of an isoelectric line of the same ECG strip,wherein the electronic device comprises at least one output unit associated at least with the data processor system and adapted to output at least the usage information, the output unit being connectable to a screen configured to display the usage information, andwherein the usage information is at least in part of the height of a P wave assessed with respect to the isoelectric line, orthe first data processor unit is configured to determine a point V corresponding to an absolute maximum of an ECG trace, a point W corresponding to an absolute minimum of the ECG trace, and/or an average of wave heights of the ECG trace, the usage information being at least in part of: a height of the point V,a height of the point W,a height measurable between the point V and the point W, and/oran average of wave heights of the ECG trace. 2. The medical appa... | PILOT TECHNOLOGIES S.R.L | A61B005/04 A61B005/06 A61B008/12 A61B005/0452 A61B005/053 A61B005/00 A61B005/0408 A61B005/0215
US10709342 | B2 | 2020-07-14 | System and method for spatial cardiovascular monitoring | 2016-12-21 | There is disclosed a system and method for measuring arterial and venous blood pulse waveforms (BPWs) of a subject utilizing photoplethysmography (PPG). In an embodiment, the system and method comprises: providing a plurality of virtual sensors positioned to cover a desired field-of-view of the subject, each virtual sensor adapted to detect and measure a BPW signal from an area of the subject's body and provide a BPW signal output; processing the BPW signal outputs of the plurality of virtual sensors to compare the BPWs at multiple areas of the subject's body to perform spatial perfusion analysis; and displaying at least one aggregate output based on the spatial perfusion analysis. At least one aggregate output may include a visualization of one or more perfusion patterns overlaid on a photographic image of the subject, and aggregate statistics including subject heart rate and breathing rate. The system and method may use a signal from one of the virtual sensors as a reference waveform for cardiovascular monitoring in the generation of parametric maps for assessing BPW characteristics at various parts of the body simultaneously. The system and method may also include a contact photoplethysmography (PPG) sensor, which is connected to the DSP and provides a BPW as a reference waveform for improved cardiovascular monitoring in the generation of parametric maps for assessing BPW characteristics at various parts of the body simultaneously. | 1. A method of measuring arterial and venous blood pulse waveforms (BPWs) of a subject utilizing spatially selective lighting changes, comprising: providing a plurality of virtual sensors positioned to cover a desired field-of-view of the subject, each virtual sensor adapted to detect and measure intensity of arterial and venous BPW signals from an area of a subject's body and provide arterial and venous BPW signal outputs using spatially coded lighting and sensing;processing the arterial and venous BPW signal outputs of the plurality of virtual sensors to compare the arterial and venous BPWs at multiple areas of the subject's body in the time and frequency domain to perform spatial perfusion analysis; anddisplaying at least one arterial and venous pulsatility perfusion aggregate output utilizing different colors to show the intensity of arterial and venous BPWs. 2. The method of claim 1, wherein the plurality of virtual sensors are formed from an array of one or more optical sensors positioned around the subject. 3. The method of claim 1, wherein the at least one aggregate output comprises a visualization of one or more perfusion patterns overlaid on an image of the subject. 4. The method of claim 3, wherein the at least one aggregate output further comprises aggregate statistics such as subject heart rate data and breathing rate data. 5. The method of claim 1, further comprising using the signal of one of the virtual sensors as a reference waveform for cardiovascular monitoring in the generation of parametric maps for assessing BPW characteristics at various parts of the body simultaneously. 6. The method of claim 1, further comprising the inclusion of a contact photoplethysmography (PPG) sensor which is connected to a digital signal processor and provides reference arterial and venous BPW signals as reference waveforms for pulsatility strength analysis. 7. The method of claim 1, further comprising combining the arterial and venous BPW signal outputs of the plurality of virtual sensors with one or more reference signals to determine pulsality strength of a subject. 8. The method of claim 1, further comprising monitoring the arterial and venous BPW signal outputs of the plurality of virtual sensors to determine the presence and quality of a BPW during mechanical cardiopulmonary resuscitation. 9. The method of claim 8, further comprising using the presence and quality of a BPW to determine the effectiveness of cardiopulmonary resuscitation. 10. The method of claim 9, further comprising displaying an aggregate visualization to describe the effectiveness of cardiopulmonary resuscitation. 11. The method of claim 1, further comprising utilizing the arterial and venous BPW signal outputs of the plurality of virtual sensors to monitor neonates innaturalistic or intensive care environments. 12. A system for measuring arterial and venous blood pulse waveforms (BPWs) of a subject utilizing spatially selective lighting changes, comprising: a plu... | - | A61B005/024 A61B005/00 A61H031/00 A61B005/1455
US10716537 | B2 | 2020-07-21 | Esophageal electrocardiogram for transesophageal echocardiography | 2015-07-30 | A medical imaging system employs a transesophageal probe (20) including an ultrasound transducer (21) for scanning a patient's heart and atrial electrode(s) (22) for generating atrial electrocardiogram signal(s) predominately indicative of electrical activity of atrium chambers of the patient's heart. The medical imaging system further employs ventricular electrode(s) (23) for generating ventricular electrocardiogram signal(s) predominately indicative of electrical activity of ventricle chambers of the patient's heart. The medical imaging system further employs an electrocardiogram machine (30) for generating an electrocardiogram waveform based on the indicated electrical activities of the patient's heart, and for generating a cardiac gating signal representative of a cyclical cardiac phase period of the electrocardiogram waveform. The medical imaging system further employs an ultrasound machine (40) for reconstructing an ultrasound image of the patient's heart encompassing a time segment or an entirety of the cyclical cardiac phase period. | 1. A medical imaging system for ultrasound imaging a heart of a patient, the medical imaging system comprising: a transesophageal probe operable to be positioned in an esophagus of the patient adjacent the heart, wherein the transesophageal probe includes an ultrasound transducer operable to scan the heart, andwherein the transesophageal probe further includes at least one atrial electrode, each atrial electrode being operable to generate an atrial electrocardiogram signal predominately indicative of electrical activity by atrium chambers of the heart;at least one ventricular electrode operable to be attached to a surface thoracic region of the patient, each ventricular electrode is further operable to generate a ventricular electrocardiogram signal predominately indicative of electrical activity by ventricle chambers of the heart;an electrocardiogram machine operable to generate an electrocardiogram waveform of the heart responsive to a generation of the at least one atrial electrocardiogram signal by the at least one atrial electrode and responsive to a generation of the at least one ventricular electrocardiogram signal by the at least one ventricular electrode, wherein the electrocardiogram machine is further operable to generate a cardiac gating signal indicative of a cyclical cardiac phase period of the electrocardiogram waveform; andan ultrasound machine to reconstruct at least one ultrasound image of the heart encompassing at least a portion of the cyclical cardiac phase period of the electrocardiogram responsive to an ultrasound scanning of the heart by the ultrasound transducer and responsive to a generation of the cardiac gating signal by the electrocardiogram machine. 2. The medical imaging system of claim 1, wherein the electrocardiogram machine includes a cardiac phase monitor operable to derive the cyclical cardiac phase period from a normal atrial phase of the electrocardiogram waveform. 3. The medical imaging system of claim 1, wherein the electrocardiogram machine includes a cardiac phase monitor operable to derive the cyclical cardiac phase period from a normal ventricular phase of the electrocardiogram waveform. 4. The medical imaging system of claim 1, wherein the electrocardiogram machine includes an atrial arrhythmia monitor operable to derive the cyclical cardiac phase period from an atrial arrhythmia phase of the electrocardiogram waveform. 5. The medical imaging system of claim 1, wherein the electrocardiogram machine generates the electrocardiogram waveform as a function of a bipolar signal derived from a pair of atrial electrodes. 6. The medical imaging system of claim 1, wherein the electrocardiogram machine generates the electrocardiogram waveform as a function of a bipolar signal derived from an atrial electrode and a ventricular electrode. 7. The medical imaging system of claim 1, wherein the electrocardiogram machine generates the electrocardiogram waveform as a function of a signal derived from a combination... | KONINKLIJKE PHILIPS N.V. | A61B008/12 A61B005/00 A61B008/08 A61B008/00 A61B005/042
US10702203 | B2 | 2020-07-07 | Prick test single use sterile vial and method | 2017-02-08 | A method for producing prick test single use sterile vials for use in a prick test is provided. The method comprises providing at least one allergen source including therewithin allergen extract, providing a plurality of single dose vials, connecting a tube to the at least one allergen source, connecting a metering device to the tube, drawing by the tube an amount of the allergen extract from one of the at least one allergen source, receiving the amount of allergen extract at the metering device, and dispensing by the metering device a volume of allergen extract into one of the plurality of single dose vials. | 1. A method for producing a prick test kit for use in a prick test, the method comprising: providing a prick test kit body including a tray, wherein the tray includes a plurality of wells;providing at least one allergen source including therewithin allergen extract;providing a plurality of single dose vials;connecting a tube to the at least one allergen source;connecting a metering device to the tube;drawing by the tube an amount of the allergen extract from one of the at least one allergen source;receiving the amount of the allergen extract at the metering device;dispensing by the metering device a volume of the allergen extract into one of the plurality of single dose vials; andproviding a plurality of prick testers each associated with a particular group of single dose vials from among the plurality of single dose vials and each associated with a group of wells of the plurality of wells,wherein each of the plurality of prick testers include a plurality of needles extending thereon,wherein each of the group of wells associated with one of the plurality of prick testers is disposed adjacent to an associated at least one slot having a particular shape, wherein the particular shape of the associated at least one slot is different from shapes of other slots associated with other groups of wells of the plurality of wells, andwherein each of the plurality of prick testers includes at least one tab extending beyond a length of the plurality of needles and that matches the particular shape of the associated at least one slot, such that each of the plurality of prick testers is operable for use with only its associated group of wells, and such that the at least one tab of each of the plurality of prick testers is only insertable in the associated at least one slot of the associated group of wells. 2. The method of claim 1, further comprising: sealing one of the plurality of single dose vials with one of a plurality of stoppers; andsecuring the one of the plurality of stoppers on the one of the plurality of single dose vials with one of a plurality of seals. 3. The method of claim 2, further comprising affixing one of a plurality of labels to the one of the plurality of single dose vials. 4. The method of claim 3, wherein the tray includes disposed therein a plurality of recesses, the method further comprising: placing within one of the plurality of recesses of the tray the one of the plurality of single dose vials. 5. The method of claim 4, further comprising repeating the drawing, receiving, dispensing, sealing, securing, affixing, and placing steps until a desired number of the plurality of single dose vials are produced. 6. The method of claim 5, further comprising: extracting a predetermined amount of the allergen extract from one of the plurality of single dose vials; anddisposing the predetermined amount of the allergen extract into a corresponding one of the plurality of wells. 7. The method of claim 6, further comprising repeating the extr... | ROCA Medical Ltd. | A61B005/00 A61B010/00 A61B017/20 B65D025/10 B65D025/20 B65D045/16 B65B003/00 B65B007/16
US10702204 | B2 | 2020-07-07 | Apparatus and method for use in analyzing a patient's bowel | 2018-09-24 | An apparatus and method are provided for use in studying a patient's bowel which combines recording and analysis of physiologic parameters and patient sensory perception. A pain input detector, a pain transducer, and a processor are provided. A gas pressure transducer and flow meter may also be provided. Output data may be generated to reflect perceived patient pain, volume of gas delivered to the patient's bowel, and bowel pressure. An indication of perceived patient pain is processed by the processor to generate data that may used to validate a scan of a patient or may be used as a diagnostic tool. | 1. An apparatus configured to provide a representation indicative of a scanning condition of a patient's colon, the apparatus comprising: a pressure transducer configured to detect a pressure level of gas within a patient's colon and to generate an output indicative of the pressure level of gas detected;a flow meter configured to detect a volume level of gas delivered to the patient's colon and to generate an output indicative of the volume level of gas detected;a visual display;a processor in communication with the pressure transducer, the flow meter, and the visual display;circuitry in communication with the processor, wherein the circuitry is configured to monitor a period of time over which the pressure level is detected and maintained within a predetermined range of pressures; anda sensory transducer configured to detect a patient's pain level and to generate an output indicative of the pain level detected,wherein the processor is configured to receive the outputs from the pressure transducer the flow meter, the circuitry, and the sensory transducer and to cause the visual display to present a representation of at least one of the pressure level or the volume level detected within the patient's colon and the pain level detected. 2. The apparatus of claim 1, wherein the processor is configured to receive an input from a user that indicates a desired range of the pressure level of gas within the patient's colon. 3. The apparatus of claim 1, wherein the processor is configured to receive an input from a user that indicates a threshold volume level of gas delivered to the patients colon. 4. The apparatus of claim 1 further comprising a signaling unit that generates an indication to a user in response to detection of a predetermined pressure level and volume level. 5. The apparatus of claim 4, wherein the indication provided by the signaling unit comprises at least one of an electronic signal, an alarm, a flashing light, or a sound alarm. 6. The apparatus of claim 1, wherein the representation presented by the visual display comprises plotted data of at least one of the pressure level or the volume level over time. 7. The apparatus of claim 1, wherein at least one of the outputs is stored for further use by a user. 8. The apparatus of claim 1, wherein the processor is further configured to determine that the pressure level is within a predetermined range of pressures for a predetermined period of time in an instance in which the pressure level is within the predetermined range of pressures for the predetermined period of time and that the volume level meets a predetermined threshold volume level in an instance in which the volume level meets the predetermined threshold volume level. 9. A method for providing a representation indicative of a scanning condition of a patient's colon comprising: insufflating a patient's colon with gas;detecting, via a pressure transducer, a pressure level of the gas within the patient's colon;generating, via ... | UNIVERSITY OF MARYLAND, BALTIMORE | A61B005/00 A61B005/03 A61M013/00 G06F019/00 A61B005/055 A61B006/03
US10722303 | B2 | 2020-07-28 | System and method for mapping the functional nerves innervating the wall of arteries, 3-D mapping and catheters for same | 2017-11-03 | Disclosed herein are systems and methods for locating and identifying nerves innervating the wall of arteries such as the renal artery. The present invention identifies areas on vessel walls that are innervated with nerves; provides indication on whether energy is delivered accurately to a targeted nerve; and provides immediate post-procedural assessment of the effect of energy delivered to the nerve. The methods includes evaluating a change in physiological parameters after energy is delivered to an arterial wall; and determining the type of nerve that the energy was directed to (sympathetic or parasympathetic or none) based on the evaluated results. The system includes at least a device for delivering energy to the wall of blood vessel; sensors for detecting physiological signals from a subject; and indicators to display results obtained using said method. Also provided are catheters for performing the mapping and ablating functions. | 1. A system for mapping of the functional nerves innervating the wall of a renal vessel, comprising: a. a first set of electrodes comprising one or more electrodes connected to a first power source for delivering electrical stimulation, said first set of electrodes is adapted for insertion to a distal one-third of the renal vessel adjacent the kidney;b. a second set of electrodes comprising one or more electrodes connected to a second power source for delivering ablation energy, said second set of electrodes is adapted for insertion to a proximal one-third of the renal vessel further from the kidney;c. one or more measuring devices for measuring one or more physiological parameters associated with the stimulation of said renal vessel; andd. a computing device coupled to said one or more measuring devices and is configured for computing any increase or decrease in the physiological parameters against a baseline in response to the stimulation of said renal vessel; wherein said first set of electrodes and second set of electrodes are spaced apart so that, when said first set of electrodes is placed at the distal one-third of said renal vessel, said second set of electrodes will be located at the proximal one-third of said renal vessel. 2. The system of claim 1, wherein said first set of electrodes is arranged in a configuration to contact said distal one-third of the renal vessel wall at multiple sites, said configuration has a diameter in the range of 3.5 mm to 10 mm. 3. The system of claim 1, wherein said first or second set of electrodes is arranged in a configuration comprising a spiral having one or more loops. 4. The system of claim 1, wherein said configuration comprises a spiral pyramid with the loops becoming progressively smaller from a proximal end to a distal end. 5. The system of claim 1, wherein said first set of electrodes is located on a catheter comprising a distal end, wherein the distal end of said catheter comprises a tubular structure that houses one or more resilient members with pre-formed curvatures, wherein said first set of electrodes is disposed on said one or more resilient members, said one or more resilient members are attached at their proximal ends to a controlling shaft, wherein movement of said controlling shaft causes said one or more resilient members to be pushed out of said tubular structure to resume the pre-formed curvature or retracted into said tubular structure. 6. The system of claim 5, further comprises a controlling ring that sheaths said one or more resilient members, wherein movement of said controlling ring along said resilient members controls the extend said one or more resilient members resumes its pre-formed curvature. 7. The system of claim 5, further comprises a controlling rod within said controlling shaft, wherein the distal end of said controlling rod is attached to the distal ends of said one or more resilient members, wherein retracting said controlling rod after said one or more resi... | SYMAP MEDICAL (SUZHOU), LIMITED | A61B018/14 A61B005/00 A61B005/04 A61B005/0205 A61B005/0215 A61B005/024 A61B005/20 A61B018/00 A61B018/18 A61N001/36
US10714218 | B2 | 2020-07-14 | Methods and systems for displaying electrophysiological lesions | 2019-03-15 | The present disclosure provides systems and methods for rendering lesions on a geometric surface model of a geometric structure. The system includes a computer-based model construction system configured to create a three-dimensional (3D) texture map including a plurality of voxels each having a tissue necrosis value, increment the tissue necrosis values as a function of at least one parameter to generate a total tissue necrosis value for each voxel, render at least one lesion on the geometric surface model based on the total tissue necrosis values, and display the geometric surface model and the at least one rendered lesion. | 1. A system for rendering lesions on a geometric surface model of a geometric structure, the system comprising: a computer-based model construction system configured to be coupled to a device that includes at least one sensor configured to acquire a set of original location data points to facilitate generating the geometric surface model, the computer-based model construction system further configured to: create a three-dimensional (3D) texture map including a plurality of voxels each having a tissue necrosis value;determine an ablation location;identify, from the plurality of voxels, all voxels that are within a given distance from the determined ablation location;for each of the identified voxels that are within the given distance from the determined ablation location, increment the tissue necrosis values as a function of both i) distance from the determined ablation location and ii) at least one parameter to generate a total tissue necrosis value for each voxel, wherein the at least one parameter includes at least one of an ablation power level, an ablation delivery duration, a tissue contact force, a lesion-size index, and a force-time integral associated with an ablation procedure, a tissue type, and a total amount of energy delivered;render at least one lesion on the geometric surface model based on the total tissue necrosis values such that a center of the at least one lesion is rendered differently than portions of the lesion further away from the center of the lesion; anddisplay the geometric surface model and the at least one rendered lesion. 2. The system of claim 1, wherein to render at least one lesion on the geometric surface model, the system is configured to render the at least one lesion directly on the geometric surface model. 3. The system of claim 1, wherein to render at least one lesion on the geometric surface model, the system is configured to render the at least one lesion on an overlay layer that is displaced relative to the geometric surface model. 4. The system of claim 1, wherein to create a 3D texture map, the system is configured to create a 3D texture map including voxels each having a tissue necrosis value and a tissue temperature value. 5. The system of claim 4, wherein the system is further configured to: increment the tissue temperature values as a function of temperature measured during an ablation procedure to generate a total tissue temperature value for each voxel; andrender the at least one lesion based on the total tissue temperature values. 6. The system of claim 1, wherein to render at least one lesion on the geometric surface model, the system is configured to render the at least one lesion such that a border zone of the at least one lesion is transparent further from the center of the at least one lesion and opaque closer to the center of the at least one lesion. 7. The system of claim 1, wherein to render at least one lesion on the geometric surface model, the system is configured to render the ... | St. Jude Medical, Cardiology Division, Inc. | G16H050/50 G06T019/20 A61B034/10 A61B005/01 A61B005/00 A61B018/00 A61B090/00
US10716921 | B2 | 2020-07-21 | Methods of using a dual-lumen sheath in intraluminal procedures | 2016-04-25 | Disclosed are sheaths that comprise a first sheath member having a first passageway, a first length, and a first proximal end defined by a first valve apparatus configured to seal the first passageway, the first passageway having a first passageway diameter at a location in the first passageway; and a second sheath member coupled to the first sheath member, the second sheath member having a second passageway and a second length, the second passageway having a second passageway diameter at a location in the second passageway; where the first passageway and the second passageway are separate from each other and not co-axial, the first length is different from the second length, and the first and second sheath members are positioned beside each other for a portion of their first and second lengths. | 1. A method of positioning an elongated member with two fluidly-distinct passageways into a subject, the method comprising: introducing an elongated member into a bodily lumen of a subject, the elongated member comprising a first passageway extending along a single longitudinal axis from a first proximal end to a first distal end, the first distal end located at a distal end of the elongated member, the first proximal end configured to couple to a valve apparatus;wherein the valve apparatus comprises a membrane that extends at least partially across the first passageway to allow other devices to be introduced through the first passageway while preventing fluid from flowing out of the first proximal end of the first passageway;introducing a separate device into the first passageway through the valve apparatus, wherein the valve apparatus is configured to seal around the separate device;wherein the elongated member comprises a second passageway, the second passageway comprising a generally circular cross-section;wherein the first passageway and the second passageway are fluidly separate from each other and located along different longitudinal axes along an entire length of the elongated member; andwherein the first and second passageways are adjacent each other for at least a portion of the elongated member; anddetecting at least one physiologic or biochemical property using a detector positioned at least partially within the elongate member, wherein the detector is located along a distal end of the second passageway. 2. The method of claim 1, wherein the elongated member comprises a sheath. 3. The method of claim 1, wherein the elongated member comprises a catheter. 4. The method of claim 1, wherein a diameter of the first passageway is different than a diameter of the second passageway. 5. The method of claim 1, wherein the separate device comprises at least one of a catheter, a guidewire and an introducer. 6. The method of claim 1, wherein the bodily lumen of a subject comprises a blood vessel. 7. A method of positioning an elongated member with two fluidly-distinct passageways into a subject, the method comprising: introducing an elongated member into a bodily lumen of a subject, the elongated member comprising a first passageway extending along a single longitudinal axis from a first proximal end to a first distal end, the first proximal end being configured to couple to a valve apparatus;wherein the valve comprises a membrane that extends at least partially across the first passageway to allow at least one other device to be introduced through the first passageway while preventing fluid from flowing out of the first proximal end of the first passageway; andwherein the elongated member comprises a second passageway, wherein the first passageway and the second passageway are fluidly separate from each other and are located along different longitudinal axes of the elongated member;introducing a separate device into the first passageway alo... | Endophys Holdings, LLC | A61M025/09 A61B017/12 A61M025/02 A61M025/10 A61M025/06 A61B005/0476 A61B005/145 A61B005/00 A61B010/02 A61F007/12 A61M039/22 A61N001/05 A61N005/10 A61B017/122 A61M025/00 A61M025/01
US10709381 | B2 | 2020-07-14 | Device for determining fat expenditure from levels of ketone bodies that have passed through the skin and methods for determining the same | 2018-07-18 | A wearable sensing device having a first and second opening, a first semipermeable membrane having a first surface and a second surface, and a second semipermeable membrane having a third and fourth surface, a ketone body sensor, and a void. The first opening is juxtaposed to the first surface and the second opening is juxtaposed to the third surface. The space between the first and second openings is the void and wherein the ketone body sensor is positioned within the void. Gasses may permeate through the first opening and into the void to contact the sensor and exit the void through the second opening. The sensor may be in direct skin contact with vapor dissipation being enabled through ventilation. Ventilation will be facilitated by microgrooves on the sensor surface or through vent openings in the encasement. | 1. A sensing device comprising: a. a first semipermeable membrane having a first surface and a second surface, the first semipermeable membrane having a first opening, a second opening, and a third opening, wherein the first opening has a diameter larger than the second opening;b. a ketone body sensor positioned in contact with the first semipermeable membrane; andc. wherein the first surface and the second surface define a void to receive a vapor entering the void through the first opening and the second opening in a first direction, the vapor exiting the void through the third opening in a second direction, the first direction being substantially perpendicular to the second direction. 2. The sensing device of claim 1, having a housing shaped for contacting a skin of a mammal, wherein the first surface is in contact with the skin. 3. The sensing device of claim 2, wherein said mammal is a primate. 4. The device of claim 2, wherein housing is suitable for contacting a body part selected from the group consisting of: a hand, wrist, arm, chest, torso, leg, foot, and ankle. 5. The device of claim 1, further comprising a receiver component suitable for electronic connection between the device and said receiver component. 6. The device of claim 1, wherein said semipermeable membrane is made up of at least one layer of polytetrafluoroethylene or graphene. 7. A device for measuring acetone gasses released from the skin of a mammal comprising: a. an encasement comprising a plurality of vent-in openings on a bottom face, a vent-out opening on at least one face which is adjacent to the bottom face, a void, a first semipermeable membrane, and a ketone sensor;b. wherein said first semipermeable membrane is positioned in contact with said bottom face, the ketone sensor is positioned in the void, a vapor enters the void through the plurality of vent-in openings, the vapor exits the void through the vent-out opening. 8. The device of claim 7, further comprising a microprocessor and a display, wherein said microprocessor processes data received from said ketone sensor and provides an output to said display. 9. The device of claim 7, wherein said first semipermeable membranes are made up of at least two layers. 10. The device of claim 7, wherein said first semipermeable membrane is made up of at least one layer of polytetrafluoroethylene. 11. The device of claim 7, wherein said sensor is a platinum coated Indium Nitride sensor, or an optical sensor. 12. The device of claim 11, wherein said sensor further comprises platinum nanoparticles. 13. The device of claim 7, wherein said encasement is a metallic microlattice, wherein said microlattice is connected to said sensor for dissipation of heat from said sensor. 14. A method of measuring ketone bodies crossing a skin of a mammal, comprising: a. attaching a device to the skin of a mammal, wherein said device comprises: a plurality of vent-in openings on a bottom face, a vent-out opening on at least one f... | - | A61B005/1477 A61B005/00 A61B005/0205 A61B005/145 B01D069/06 A61B005/0225 A61B005/024 G01N033/64 A61B005/026 A61B005/04 G01N033/52 A61B005/01 G01N033/497 B01D071/36 B01D071/02 A61B005/08 B01D063/04
US10712277 | B2 | 2020-07-14 | Raman imaging devices and methods of molecular imaging | 2011-03-09 | In accordance with the purpose(s) of the present disclosure, as embodied and broadly described herein, embodiments of the present disclosure, in one aspect, relate to Raman imaging devices (e.g., Raman endoscope probes), methods of using Raman agents and Raman imaging devices to image or detect a signal, and the like. | 1. A method of imaging, comprising: administering at least two types of Raman agents to a subject, wherein at least one of the two types of Raman agents has a specific affinity for a specific target;introducing a Raman imaging device into the subject;positioning the Raman imaging device adjacent the specific target inside the subject;exposing the specific target inside the subject to a light beam from the Raman imaging device, wherein the light beam is scattered by the first type of Raman agent that associated with the specific target, wherein the light beam that is scattered is referred to as a Raman scattered light energy; anddetecting the Raman scattered light inside the subject using the Raman imaging device, using the Raman scattered light energy to form an image; andanalyzing the scattered light to determine the type of Raman agent, wherein the type of Raman agent determines the specific target detected,wherein one of the at least two types of Raman agent is not target-specific, and wherein the method further comprises analyzing the scattered light to determine the ratio of the specific binding Raman agent to non-specific binding Raman agent, wherein the ratio provides an estimate of the bound Raman agents. 2. The method of claim 1, wherein the Raman agent is selected from the group consisting of: a Raman nanoparticle and a Raman compound. 3. The method of claim 1, wherein the specific target is selected from the group consisting of: cancer, tumor, precancerous cells or tissue, atherosclerosis, and fibrosis. 4. The method of claim 1, wherein introducing the Raman imaging device includes introducing into the subject the Raman imaging device to an anatomical region of the subject, wherein the anatomical region of the subject is selected from the group consisting of: cervix, bladder, bronchioles, esophagus, stomach, colon, and rectum. 5. The method of claim 4, further comprising, mapping of at least a portion of the subject using the Raman scattered light, wherein mapping includes collecting Raman scattered light and correlating the Raman scattered light to a position in at least a portion of the inside of the subject, wherein the portion of the subject is selected from the group consisting of: cervix, bladder, bronchioles, esophagus, stomach, colon, and rectum. 6. The method of claim 1, wherein administering includes administering, locally to the colon, a gold nanoparticle Raman agent having a targeting agent specific for a dysplastic flat lesion in the colon, wherein the specific target is a dysplastic flat lesion, and wherein introducing includes introducing into the colon. 7. The method of claim 1, wherein administering includes administering, locally to the colon, a polyethylene glycol (PEG)-ylated Raman active silica-gold nanoparticle Raman agent having a targeting agent specific for colorectal cancer, wherein the specific target is a colorectal cancer, and wherein introducing includes introducing into the colon. 8. A method of im... | The Board of Trustees of the Leland Stanford Junior University | G01N021/65 A61B005/00 A61K049/00 A61K051/12 B82Y005/00 B82Y015/00 B82Y030/00
US10705014 | B2 | 2020-07-07 | Image diagnosis apparatus, method of controlling image diagnosis apparatus, computer program, and computer readable storage medium | 2019-03-27 | An image diagnosis apparatus determines an effective pulse based on a pulse signal corresponding to the rotation of an imaging core from a pulse signal corresponding to an A line indicating a period of a light source and switches the polarization state of light transmitted from an optical transceiver according to the effective pulse. According to the effective pulse, the image diagnosis apparatus acquires a signal of interference light between the light received by the optical transceiver and reference light. | 1. An image diagnosis apparatus comprising: a motor drive unit configured to be connected to a probe, the probe being configured to rotatably accommodate an imaging core including an optical transceiver configured to transmit and receive light, the motor drive unit being configured to rotate the imaging core; anda processor configured to: determine an effective pulse based on a pulse signal corresponding to the rotation of the imaging core from a pulse signal corresponding to an A line indicating a period of a light source and for switching a polarization state of the light transmitted from the optical transceiver according to the effective pulse; andacquire a signal of interference light between the light received by the optical transceiver and reference light according to the effective pulse. 2. The image diagnosis apparatus according to claim 1, wherein the processor is configured to: determine, as the effective pulse, a pulse immediately after a pulse of the pulse signal corresponding to the rotation of the imaging core out of pulses of the pulse signal indicating the period of the light source. 3. The image diagnosis apparatus according to claim 1, wherein the processor is configured to: generate the pulse signal, which switches between a high and a low for every timing of the effective pulse, and switches the polarization state of the light transmitted from the optical transceiver according to the generated pulse signal. 4. The image diagnosis apparatus according to claim 1, wherein the processor is configured to: switch the polarization state of the light transmitted from the optical transceiver to one of polarization states orthogonal on a Poincare sphere display. 5. The image diagnosis apparatus according to claim 1, wherein the motor drive unit is configured to output a preset number of pulses per one rotation of the imaging core as the pulse signal; and an encoder signal conversion unit configured to convert the pulse signal from the motor drive unit into a converted pulse signal of a preset number of pulses, and outputs the converted pulse signal to a trigger signal generation unit. 6. The image diagnosis apparatus according to claim 5, comprising: an optical unit configured to output the pulse signal indicating a timing of a predetermined state to the trigger signal generation unit as a sweep signal. 7. The image diagnosis apparatus according to claim 6, further comprising: an electro-optic modulator configured to switch the polarization state of the measurement light obtained from the light source as the light transmitted from the optical transceiver according to the sweep signal from the trigger signal generation unit. 8. The image diagnosis apparatus according to claim 7, wherein the polarization state of the light is switched before being divided into the measurement light and the reference light. 9. The image diagnosis apparatus according to claim 1, comprising: a trigger signal generation unit configured to determine th... | TERUMO KABUSHIKI KAISHA | G01N021/21 G02F001/01 A61B001/00 A61B005/00 G01N021/17 A61B001/06
US10722168 | B2 | 2020-07-28 | Dynamic ligament balancing system with pin positioning block | 2018-01-26 | A surgical pin positioning block comprising a top portion comprising a pin position dial including one or more pin hole guides, the pin position dial setting angular pin positions of the one or more pin hole guides, wherein the one or more pin hole guides includes an orifice for drilling of pin holes, one or more dial position notches that secures the pin position dial in a given angular pin position, one or more sliding members coupled with one or more receptacles of a bottom portion, the bottom portion comprising the one or more receptacles, an adapter including one or more prongs that attach to a prosthetic inlay device, and a locking mechanism for securing the one or more sliding members in the one or more receptacles. | 1. A surgical pin positioning block comprising: a top portion and a bottom portion each having a proximal surface and a distal surface opposite the proximal surface; a pin position dial including one or more pin hole guides, the pin position dial pivotally coupled to the top portion for setting angular pin positions of the one or more pin hole guides, wherein the one or more pin hole guides includes an orifice for drilling of pin holes;one or more dial position notches located on the top portion that secures the pin position dial in a given angular pin position relative to the top portion;one or more sliding members extending from the top portion;a bolt that adjustably extends distally through the proximal and distal surfaces of the top portion, the bolt stabilizing the top portion of the pin positioning block to a patient's bone;one or more receptacles in the bottom portion for receiving in a sliding arrangement the one or more sliding members extending from the top portion;an adapter including one or more prongs extending outward from the distal surface of the bottom portion that attach the surgical pin positioning block to a prosthetic inlay device; anda lock for securing the one or more sliding members in the one or more receptacles therewith setting a distance between the bolt and the bottom portion. 2. The surgical pin positioning block of claim 1 wherein the top portion further comprises a pin position dial indicator that indicates the given angular pin position of the pin position dial. 3. The surgical pin positioning block of claim 1 wherein the pin position dial is operable in clockwise and counter-clockwise directions for setting the angular pin positions of the one or more pin hole guides. 4. The surgical pin positioning block of claim 1 wherein the one or more pin hole guides are on opposite distal ends of the pin position dial. 5. The surgical pin positioning block of claim 4 wherein the angular pin positions of the one or more pin hole guides are inversely related. 6. The surgical pin positioning block of claim 1 wherein the top portion is extendable from the bottom portion via the one or more sliding members coupled with the one or more receptacles. 7. The surgical pin positioning block of claim 1 wherein the lock sets a vertical pin position of the one or more pin hole guides relative to the one or more prongs. 8. The surgical pin positioning block of claim 1 wherein the bottom portion further comprises a caliper configured as a tibial plate sizer. 9. The surgical pin positioning block of claim 1 wherein the prosthetic inlay device comprises: at least two platform structures, wherein each of the platforms are supported on a scissor arm structure and a spring, and wherein each of the platforms are configured to variably depress the spring as a function of joint flexion angle when placed in situ;a force sensing sensor configured beneath each spring, the force sensing sensor detecting pressure from depressing the spring at e... | - | A61B005/00 A61B017/17 A61F002/38 A61F002/46 A61B090/00 A61B005/107 A61B017/02
US10709335 | B2 | 2020-07-14 | Infant monitoring system with observation-based system control and feedback loops | 2017-12-31 | A method of optimizing sleep of a subject using smart-home devices may include operating a smart-home system that is configured to operate in a normal mode and a sleep mode. The method may also include determining that the smart-home system should transition into the sleep mode. The smart-home devices may use a set of default parameters when operating in the sleep mode. The method may additionally include monitoring, while in the sleep mode, a sleep cycle of the subject using the smart-home devices. The method may further include detecting behavior of the subject that indicates that the sleep cycle of the subject is being interrupted or about to be interrupted, determining an environmental control that corresponds with the behavior of the subject, and adjusting the environmental control using the smart-home devices to prevent or stop the sleep cycle of the subject from being interrupted. | 1. A method of monitoring and optimizing sleep of a subject using a plurality of smart-home devices, the method comprising: operating a smart-home system comprising the plurality of smart-home devices, wherein the smart-home system is configured to operate in a plurality of modes comprising: a normal operating mode; anda sleep mode;determining that the smart-home system should transition into the sleep mode from the normal operating mode, wherein: the plurality of smart-home devices use a set of default parameters when operating in the sleep mode; andnoise-generating smart-home devices in the plurality of smart-home devices have their sound outputs lowered or disabled when the smart-home system is in the sleep mode;monitoring, while in the sleep mode, a sleep cycle of the subject using the plurality of smart-home devices;receiving an indication of a noise outside of a home in which the smart-home system is installed, wherein the indication is crowd-sourced from of the plurality of smart-home systems in different homes;adjusting an environmental control using the plurality of smart-home devices to counteract the noise and prevent or stop the sleep cycle of the subject from being interrupted;determining whether adjusting the environmental control using the plurality of smart-home devices prevented or stopped the sleep cycle of the subject from being interrupted; andadjusting the set of default parameters for use in subsequent instances where the sleep cycle of the subject is being interrupted, wherein the set of default parameters are adjusted based on a determination of whether adjusting the environmental control using the plurality of smart-home devices prevented or stopped the sleep cycle of the subject from being interrupted. 2. The method of claim 1, wherein the plurality of smart-home devices comprises a camera with a thermal imager. 3. The method of claim 2, further comprising: detecting behavior of the subject that indicates that the sleep cycle of the subject is being interrupted determining, using the thermal imager, that a temperature of the subject is above or below a target temperature. 4. The method of claim 3, wherein the thermal imager captures a thermal image of a face of the subject and compares the thermal image of the face of the subject to a baseline thermal image of the face of the subject. 5. The method of claim 3, wherein the plurality of smart-home devices comprises a thermostat, and the method further comprises: adjusting a temperature setpoint of the thermostat to bring the temperature of the subject closer to the target temperature. 6. The method of claim 1, wherein determining that the smart-home system should transition into the sleep mode comprises determining, using a camera in the plurality of smart-home devices, that the subject is in a sleep environment that is monitored by the plurality of smart-home devices. 7. The method of claim 1, wherein determining that the smart-home system should transition into the... | Google LLC | A61B005/00 G06K009/00 A61B005/01 H04N005/33 A61B005/11 G08B021/02
US10722650 | B2 | 2020-07-28 | System and method for adjusting therapy based on risk associated with a glucose state | 2014-03-28 | The present disclosure relates to a system and method for determining a basal rate adjustment based on risk associated with a glucose state of a person with diabetes. A method may include detecting a glucose state of the person based on a received glucose measurement signal and determining a current risk metric associated with the detected glucose state. The method may include identifying a reference glucose state and a reference risk metric associated with the reference glucose state, and calculating an adjustment to a basal rate of a therapy delivery device based on the current risk metric associated with the detected glucose state and the reference risk metric associated with the reference glucose level. | 1. A method of determining a basal rate adjustment based on risk associated with a glucose state of a person with diabetes, the method comprising: receiving, by at least one computing device, a signal representative of at least one glucose measurement;detecting, by the at least one computing device, a glucose state of the person based on the signal, the detected glucose state including a glucose level of the person and a rate of change of the glucose level;determining, by the at least one computing device, a current risk associated with the detected glucose state based on a target glucose state, the target glucose state being stored in memory accessible by the at least one computing device, the current risk indicating a risk of at least one of a hypoglycemic condition and a hyperglycemic condition of the person;identifying, by the at least one computing device, a first reference glucose state and a first reference risk associated with the first reference glucose state;identifying, by the at least one computing device, a second reference glucose state and a second reference risk associated with the second reference glucose state;in response to the current risk being greater than the first reference risk and less than the second reference risk, calculating, by the at least one computing device, an adjustment to a basal rate of a therapy delivery device as a difference between the current risk and the first reference risk divided by a difference between the second reference risk and the first reference risk; andtransmitting a control signal to instruct the therapy delivery device to adjust the basal rate based on the calculated adjustment. 2. The method of claim 1, wherein calculating an adjustment comprises mapping the current risk to a percent reduction of the basal rate. 3. The method of claim 2, wherein the first reference glucose state includes a glucose level corresponding to a hypoglycemic condition. 4. The method of claim 1, wherein the current risk is determined based on at least one penalty value associated with the detected glucose state and a probability of accuracy of the detected glucose state, and the at least one penalty value is indicative of a hazard associated with the detected glucose state. 5. The method of claim 1, wherein the current risk metric is determined based on a target transition from the current glucose state to the target glucose state, the target transition comprising at least one intermediate glucose state associated with a return to the target glucose state. 6. The method of claim 5, wherein the current risk metric is determined based on a cumulative penalty value weighted with a probability of accuracy of the detected glucose state, the cumulative penalty value including a sum of penalty values associated with a plurality of intermediate glucose states of the target transition to the target glucose state, each penalty value being indicative of a hazard associated with the corresponding intermediate glucose s... | Roche Diabetes Care, Inc. | A61M005/172 A61M005/142 G06F019/00 G16H050/30 A61B005/145 A61B005/00
US10722697 | B2 | 2020-07-28 | Skin piercing tool for locally puncturing a human or an animal skin and hand-held device | 2016-03-03 | The disclosure relates to a skin piercing tool for locally puncturing a human or an animal skin, comprising a piercing needle, which is embodied as solid needle, in the case of which a needle body has a proximal needle section and a distal needle section, in which a piercing tip is arranged, wherein the needle body has a needle section, which is resilient against the piercing tip in response to a compressive stress, wherein the resilient needle section is arranged on the needle body in the area of an angled needle section, in which adjacent needle sections assume an angle of less than 180° with one another. Further, a hand-held device for repeated local puncturing of a human or an animal skin is provided. | 1. A skin piercing tool for locally puncturing a human or an animal skin, comprising a piercing needle, which is embodied as solid needle, wherein a needle body of the piercing needle has needle sections including at least a proximal needle section and a at least a distal needle section, in which a piercing tip is arranged, wherein the body of the piercing needle has at least two resilient sections, each resilient section having a bend and being flexible and resilient in response to a compressive stress against the piercing tip, the bend of at least one of the at least two resilient sections being oriented in a different direction from the bend of at least one of the other at least two resilient sections, wherein the proximal needle section and at least one of the at least two resilient sections assume an angle of less than 180 degrees with respect to each other, and additionally comprising at least one further piercing needle, which forms a piercing needle group with the piercing needle. 2. The skin piercing tool according to claim 1, characterized in that the at least two resilient sections are arranged on the body of the piercing needle between the proximal needle section and the distal needle section. 3. The skin piercing tool according to claim 1, characterized in that the distal needle section is placed at an incline to a longitudinal direction of the body of the piercing needle. 4. The skin piercing tool according to claim 1, characterized in that one of the at least two resilient sections adjoining the distal needle section is placed at an incline to a longitudinal direction of the body of the piercing needle. 5. The skin piercing tool according to claim 1, characterized in that needle sections adjacent to the at least two resilient needle sections are adapted to carry out a folding or pivoting movement due to the resilience of the at least two resilient sections. 6. The skin piercing tool according to claim 1, characterized in that the at least two resilient needle sections are formed with at least one of a thinner material and a thicker material than either needle section adjacent thereto. 7. The skin piercing tool according to claim 1, characterized in that the at least two resilient sections of the piercing needle and at least two resilient sections of the at least one further piercing needle are arranged so as to be located next to one another. 8. The skin piercing tool according to claim 1, characterized in that the at least two resilient sections of the piercing needle and at least two resilient sections of the at least one further piercing needle are arranged so as to be offset relative to one another in a longitudinal direction of the piercing needle group. 9. The skin piercing tool according to claim 1, characterized in that the piercing needle and the further piercing needle are arranged so as to be located next to one another in a plane. | MT DERM GmbH | A61M037/00 A61B005/15 A61B005/151 A61D007/00 A61M005/178 A61B017/20
US10721826 | B2 | 2020-07-21 | Medical devices | 2018-05-02 | In some examples, a medical device may include circuitry on a circuit board, a casing, and a fill material at least partially within the casing. The casing may receive the circuit board and surround at least a portion of the circuit board. The casing may define an aperture configured to allow one or more electrical connectors of the circuit board to extend through the aperture for one or more electrical connections with one or more components outside of the casing. An outer surface of the casing may define a first portion of an outer surface of the medical device. The casing may comprise a material having a first melting temperature. The fill material may define a second portion of the outer surface of the medical device and may have a second melting temperature that is lower than the first melting temperature. | 1. A medical device, comprising: a circuit board having a first side and a second side opposite the first side, the first side including one or more electrical connectors;circuitry on the circuit board configured to at least one of monitor a physiological parameter of a subject or control delivery of a therapy to the subject;a casing, the casing configured to receive the circuit board with the first side of the circuit board facing the casing and surround at least a portion of the circuit board, the casing defining an aperture configured to allow the one or more electrical connectors to extend through the aperture for one or more electrical connections between the one or more electrical connectors of the circuit board and one or more components outside of the casing, an outer surface of the casing defining a first portion of an outer surface of the medical device, the casing comprising a material having a first melting temperature; anda fill material, at least a portion of the fill material being within the casing, the fill material covering at least a portion of the first side of the circuit board and at least a portion of the second side of the circuit board, the fill material defining a second portion of the outer surface of the medical device, and the fill material having a second melting temperature that is lower than the first melting temperature. 2. The medical device of claim 1, wherein the casing defines one or more internal structures configured to support the circuit board after insertion of the circuit board into the casing. 3. The medical device of claim 2, wherein the one or more internal structures comprise one or more protrusions protruding inwardly from an inner surface of the casing. 4. The medical device of claim 2, wherein the one or more internal structures are configured to support the circuit board before insertion of the fill material into the casing. 5. The medical device of claim 4, wherein the one or more internal structures are configured to maintain a gap between at least one portion of the first side of the circuit board and at least one portion of the casing during insertion of the fill material to allow for at least some of fill material to be placed between the at least one portion of the first side of the circuit board and the at least one portion of the casing. 6. The medical device of claim 2, wherein the one or more internal structures comprise an aperture wall defining at least a portion of the aperture. 7. The medical device of claim 1, wherein the casing is formed using a molding process. 8. The medical device of claim 1, wherein the casing is formed using a machining process. 9. The medical device of claim 1, wherein the fill material does not extend outside the casing. 10. The medical device of claim 1, wherein the outer surface of the casing includes one or more markings indicating an identifier for the casing. 11. The medical device of claim 1, further comprising a switch on the second side of... | Medtronic, Inc. | H05K005/00 H01R012/57 H05K005/06 A61B005/0492 H05K005/02
US10716499 | B1 | 2020-07-21 | Physiological monitoring by optical spectroscopy | 2016-03-23 | Optical spectroscopic devices, apparatus, systems and methods useable for physiological monitoring from intraosseous, subosseous, epidural, subdural, intraventricular, intramuscular, sub-adipose and other subcutaneous locations. | 1. A system for measuring a physiological variable from an internal target location within the body of a subject comprising: an optical spectroscopic apparatus operative to measure the physiological variable at the internal target location, configured to be inserted through a full-thickness opening in a bone of the subject and to thereafter operate from an operating location directly aligned with said full-thickness opening, said optical spectroscopic apparatus comprising:an emitter configured to emit light at one or more wavelengths between approximately 250 nm and approximately 2500 nm at an intensity capable of penetrating from said operating location to the internal target location;a detector for detecting lights that has been emitted from the emitter after said light has undergone dispersion and/or reflection by tissue or body fluid; anda processor or processing device configured to: a) receive spectral information from the detector,b) compare the spectral information to a set of spectra determined from a model for light attenuation to identify a member of the set of spectra that corresponds to the measured spectral information,c) correct the identified member of the set of spectra using a set of spectral correction factors that account for intervening matter, andd) use the corrected identified member of the set of spectra to determine a value associated with the particular physiological variable being measured. 2. A system according to claim 1 wherein the physiological variable comprises a variable selected from pH, temperature, oxygen tension, oxygen saturation, partial pressure of oxygen, partial pressure of carbon dioxide, hemoglobin concentration, water concentration, hematocrit, glucose concentration, the presence/absence of a biomarker, anaylte or other substance, the concentration of a biomarker, anaylte or other substance, and flow rate of blood or other fluid based on changes in concentration of an injected, optically-determinable substance such as indocyanine green or other optical dye. 3. A system according to claim 1 wherein the emitter emits light selected from near-infrared light and ultraviolet light. 4. A system according to claim 1 wherein the emitter comprises a plurality of emitters which emit light at a plurality of different wavelengths. 5. A system according to claim 1 wherein the emitter comprises at least one light emitting diode. 6. A system according to claim 1 further comprising an apparatus useable to facilitate introduction of the optical spectroscopic apparatus into or through bone to an intraosseous or subosseous operating location. 7. A system according to claim 6 wherein the apparatus useable to facilitate introduction of the optical spectroscopic apparatus into or through bone comprises at least one of: an introducer device that is advanceable into or through bone and apparatus for cutting or boring into or through bone to facilitate introduction of the optical spectroscopic apparatus into or through ... | ZOLL Medical Corporation | A61B005/1455 A61B005/1459 A61B005/145 A61B005/01 A61B005/00 A61B005/026 A61B005/1495 A61B017/34 A61B017/16 A61B005/03
US10709505 | B2 | 2020-07-14 | Bodily substance detection by evaluating photoluminescent response to excitation radiation | 2016-06-09 | A medical device may include an ablation device configured to deliver ablation energy to a treatment site. The medical device may further include a probe device configured to deliver excitation radiation to the treatment site. Further the medical device may include a radiation-receiving device configured to receive photoluminescence radiation emitted from the treatment site in response to the treatment site being illuminated by the excitation radiation and to generate a detection signal in response to the received photoluminescence radiation. Additionally, the excitation radiation may be different from the ablation energy. | 1. A method, comprising: delivering excitation radiation in a pulsed manner having a wavelength in the visible spectrum of light to a treatment site, wherein a duration of each pulse of the excitation radiation is between 100 picoseconds and 100 milliseconds;receiving radiation from the treatment site in response to the excitation radiation;discriminating a target from a remainder of the treatment site based on the received radiation;damaging the target via an ablation energy having a wavelength of about 1600 nm to 2500 nm; andcontinuing delivery of the excitation radiation during the damaging step. 2. The method of claim 1, wherein the wavelength of the excitation radiation is about 532 nm. 3. The method of claim 1, wherein the ablation energy is pulsed. 4. The method of claim 1, wherein the received radiation from the treatment site is photoluminescence radiation. 5. The method of claim 1, wherein delivering ablation energy further includes delivering ablation energy via a mechanical lithotripsy device. 6. The method of claim 1, wherein delivering ablation energy further includes delivering radiation having a wavelength in the invisible spectrum of light. 7. The method of claim 1, wherein discriminating a target from a remainder of the treatment site based on the received radiation includes comparing an intensity of the received radiation to a predetermined threshold intensity, and the method further comprising stopping damaging the target via the ablation energy when the intensity of the received radiation is below the predetermined threshold intensity. 8. The method of claim 1, wherein damaging the target via the ablation energy is done in response to the received radiation from the treatment site in response to the excitation radiation. 9. The method of claim 1, further comprising selecting an energy of the excitation radiation based on a tissue type of the target to be discriminated. 10. The method of claim 9, wherein the excitation radiation is delivered to a volume of tissue greater than a volume of tissue that receives the ablation energy. 11. A method, comprising: selecting an energy of an excitation radiation based on a tissue type of a target to be discriminated;delivering the excitation radiation in a pulsed manner to a treatment site;receiving radiation from the treatment site in response to the excitation radiation;discriminating the target from a remainder of the treatment site based on the received radiation being greater than a threshold;damaging the target via an ablation energy, wherein the ablation energy is generated by a Ho:YAG laser; andcontinuing delivery of the excitation radiation during the damaging step. 12. The method of claim 11, the received radiation is a single intensity value by integrating all wavelengths contained in the received radiation. 13. The method of claim 11, wherein the wavelength of the excitation radiation is about 532 nm. 14. The method of claim 11, wherein the ablation energy is puls... | Boston Scientific Corporation | A61B018/26 A61B018/24 A61B005/00 A61B018/20 A61B005/06 A61B018/00 A61B005/20 A61B017/00
US10722726 | B2 | 2020-07-28 | Skin treatment system | 2015-11-03 | A skin treatment system including a treatment device to treat a treatment region of a body. The skin treatment system includes a camera to capture at least one image of skin, the image having an imaging region, and to provide captured image data representing the captured image. The imaging region has a predefined spatial relationship with the treatment region. The skin treatment system further includes a storage facility for storing reference image data indicative of a reference image of the skin. The skin treatment system further includes an orientation estimation facility to determine which part of the reference image corresponds to the at least one captured image, using the captured image data and the reference image data, and to derive from a predetermined orientation of said corresponding part with respect to the reference image, an orientation indication indicative of an orientation of the treatment device with respect to the body. | 1. A skin treatment system comprising: a treatment device configured to treat a treatment region of a mammal body by physically acting upon the mammal body in said treatment region;a camera configured to capture at least one image of skin of the mammal body having an imaging region and provide captured image data representative of said at least one captured image, said imaging region having a predefined spatial relationship with said treatment region;a storage facility configured to store reference image data indicative of a reference image of the skin of the mammal body, wherein said reference image includes body features represented in said reference image data, and wherein said body features may change over time and are independent of said treatment device acting upon the mammal body;an orientation estimation facility is configured to determine which part of said reference image corresponds to said at least one captured image, using said captured image data, said reference image data and said body features, and derive from a predetermined orientation of said corresponding part with respect to the reference image an orientation indication indicative of an orientation of the treatment device with respect to the mammal body; andan update unit configured to update said reference image data in response to at least one of a model specifying an expected development over time of said body features represented in said reference image data or a change in said body features between said captured image data and said reference image data. 2. The skin treatment system according to claim 1, further comprising an analysis facility configured to use said orientation indication to provide a treatment indication indicative of an extent to which a treatment was applied to the mammal body as a function of a position on the mammal body. 3. The skin treatment system according to claim 2, further comprising a display configured to provide a visual representation of said treatment indication. 4. The skin treatment system according to claim 2, further comprising a controller configured to control the treatment device in accordance with said treatment indication. 5. The skin treatment system according to claim 1, further comprising an arrangement of a body treatment module and a processing module that are mutually functionally coupled and, at least within an operational range, independently movable with respect to each other, wherein the body treatment module at least comprises the treatment device and the camera integrated therewith, and wherein the processing module at least comprises the storage facility and the orientation estimation facility. 6. The skin treatment system according to claim 1, wherein said treatment device is a photon radiation device which is configured to irradiate the skin with photon radiation to physically affect the mammal body in the treatment region. 7. A skin treatment system comprising: a treatment device configured to treat a treatm... | KONINKLIJKE PHILIPS N.V. | A61B018/18 A61N005/06 A61B005/00 A61B018/20 A61B034/20 A61B018/00 A61B034/10 A61B034/00
US10716462 | B2 | 2020-07-21 | Systems and methods for visualizing ablated tissue | 2018-05-23 | Systems and methods for visualizing ablated tissue are disclosed. In some embodiments, a system for imaging tissue comprising: a catheter having a distal end and a proximal end; an inflatable balloon disposed about the distal end of the catheter; and an optical housing extending from the distal end of the catheter into the balloon, the optical housing being configured to position inside the balloon a light source for illuminating a tissue outside the balloon and a camera for imaging the illuminated tissue. | 1. A system for imaging tissue comprising: an elongated body having a distal end and a proximal end;an ablation device along a portion of the elongated body for ablating a tissue;a light source providing light for illuminating the tissue having at least one wavelength sufficient to excite mitochondrial nicotinamide adenine dinucleotide hydrogen (NADH) in the tissue;a sensor for detecting NADH fluorescence from the illuminated tissue, the sensor being configured to receive light having at least one wavelength to detect the NADH fluorescence from the illuminated tissue;one or more optical fibers to deliver the light from the light source to the tissue and to deliver the NADH fluorescence to the sensor; anda processor programmed to perform the steps of: obtaining the detected NADH fluorescence from the sensor in real time during ablation of the tissue using the ablation device,generating a digital representation of the detected NADH fluorescence for monitoring a progression of the ablation of the tissue, andwhile the tissue is being ablated, determining a decrease in the detected NADH fluorescence and updating the digital representation in real time to show the measured decrease in the detected NADH fluorescence that is indicative of the progression of the ablation of the tissue to enable a user to determine the need for further ablation. 2. The system of claim 1, wherein the light for illuminating the tissue has at least one wavelength between about 300 nm and about 400 nm. 3. The system of claim 2, wherein the sensor is configured to receive light having at least one wavelength between about 435 nm and about 485 nm. 4. The system of claim 1, wherein the ablation device has an energy source selected from the group consisting of radiofrequency energy, microwave energy, electrical energy, electromagnetic energy, cryoenergy, laser energy, ultrasound energy, acoustic energy, chemical energy, and thermal energy. 5. The system of claim 1, further comprising an irrigation port configured to deliver a fluid to the distal end of the elongated body. 6. A system for imaging tissue comprising: an elongated body having a distal end and a proximal end;an ablation device along a portion of the elongated body for ablating a tissue;a light source providing light for illuminating the tissue to excite mitochondrial nicotinamide adenine dinucleotide hydrogen (NADH) in the tissue, wherein the light from the light source has at least one wavelength between about 300 nm and about 400 nm sufficient for NADH excitation;a sensor being configured to receive light to detect NADH fluorescence from the illuminated tissue, wherein the light detected by the sensor has at least one wavelength between about 435 nm and about 485 nm; anda processor programmed to perform the steps of: obtaining the detected NADH fluorescence from the sensor in real time during ablation of the tissue using the ablation device,generating a digital representation of the detected NADH fluorescence fo... | The George Washington University 460Medical, Inc. | A61B001/04 A61B005/00 A61B005/1459 A61B018/00 A61B001/00 A61B001/06 A61B005/145 A61B090/35 A61B001/05 A61B090/30 A61B090/00 A61B018/14 A61B034/30 A61B018/02
US10702210 | B2 | 2020-07-07 | Meal intake detection device, meal intake detection method and computer-readable non-transitory medium | 2018-05-02 | A meal intake detection device includes: a memory; and a processor coupled to the memory and configured to: detect a rise of heart rate; extract as a meal intake candidate a time range from a detected rise to time when the heart rate decreases to a predetermined value in temporal changes in heart rate, and when a second meal intake candidate is extracted in a time range of an extracted first meal intake candidate, separate the first meal intake candidate and the second meal intake candidate along a predetermined line; and detect whether each of the first meal intake candidate and the second meal intake candidate is a meal intake based on feature quantities relating to the first meal intake candidate and the second meal intake candidate that are separated. | 1. A meal intake detection device comprising: a memory; anda processor coupled to the memory and configured to: detect a rise of heart rate,extract as a meal intake candidate a time range from a detected rise to time when the heart rate decreases to a predetermined value in temporal changes in heart rate,when a second meal intake candidate is extracted in a time range of a first meal intake candidate extracted, separate the first meal intake candidate and the second meal intake candidate along a predetermined line obtained by using the first meal intake candidate to a start point of the second meal intake candidate, anddetect whether each of the first meal intake candidate and the second meal intake candidate is a meal intake based on feature quantities relating to the first meal intake candidate and the second meal intake candidate that are separated. 2. The meal intake detection device according to claim 1, wherein the processor is configured to, when a heart rate of the first meal intake candidate tends to decrease at a time of a rising of the second meal intake candidate, obtain as the predetermined line a line that starts from the time of the rising of the second meal intake candidate and has a slope equal to a slope of heart rate of the first meal intake candidate at the time of the rising of the second meal intake candidate. 3. The meal intake detection device according to claim 1, wherein the processor is configured to eliminate from meal intake candidates a meal intake candidate that is detected as not being a meal intake among meal intake candidates extracted. 4. The meal intake detection device according to claim 1, wherein the processor is configured to estimate an index relating to a meal intake for a meal intake candidate that is detected as a meal intake. 5. A meal intake detection method implemented by a computer, the meal intake detection method comprising: detecting a rise of heart rate;extracting as a meal intake candidate a time range from a detected rise to time when heart rate decreases to a predetermined value in temporal changes in the heart rate;when a second meal intake candidate is extracted in a time range of a first meal intake candidate extracted, separating the first meal intake candidate and the second meal intake candidate along a predetermined line obtained by using the first meal intake candidate to a start point of the second meal intake candidate; anddetecting whether each of the first meal intake candidate and the second meal intake candidate is a meal intake based on feature quantities relating to the first meal intake candidate and the second meal intake candidate that are separated. 6. A non-transitory computer readable storage medium storing a meal intake detection program causing a computer to execute a process, the process comprising: detecting a rise of heart rate;extracting as a meal intake candidate a time range from a detected rise to time when heart rate decreases to a predetermined value in t... | FUJITSU LIMITED | A61B005/00 A61B005/024 G16H050/20 G16H020/60
US10722109 | B2 | 2020-07-28 | Apparatus and method to illuminate and orient guidewire | 2018-09-24 | An apparatus includes a power source, a pulse generator, a light source, a guidewire, an optical fiber, and a connector. The pulse generator is in electrical communication with the power source. The pulse generator is also capable of generating an electric pulse. The light source is in electrical communication with the pulse generator in such a way that the light source turns on and off in response to the electric pulse received from the pulse generator. The guidewire includes a proximal end and a distal end. The optical fiber extends through the guidewire from the proximal end of the guidewire towards the distal end of the guidewire. The connector couples with the guidewire while allowing the light source to communicate with the optical fiber. | 1. A guide system comprising: (a) a guide catheter comprising: (i) a proximal end,(ii) a distal end, and(iii) a lumen extending from the proximal end to the distal end; and(b) a guidewire extending within the lumen of the guide catheter, wherein the guidewire comprises: (i) a proximal end,(ii) a distal end,(iii) a preformed bend formed near the distal end of the guidewire, and(iv) a visible marker positioned on or near the preformed bend, wherein the visible marker visually indicates a direction of the preformed bend. 2. The guide system of claim 1, wherein the guide catheter further comprises a longitudinal shaft. 3. The guide system of claim 1, wherein the guide catheter further comprises a preformed catheter bend. 4. The guide system of claim 3, wherein the preformed bend of the guide catheter is approximately 20°. 5. The guide system of claim 1, wherein the visible marker extends proximally from the distal end of the guidewire along the preformed bend. 6. The guide system of claim 1, wherein the guidewire is configured to translate relative to the guide catheter to thereby position the distal end of the guidewire distally in relation to the distal end of the guide catheter. 7. The guide system of claim 6, wherein the guidewire is configured to translate relative to the guide catheter to thereby position the preformed bend of the guidewire distally in relation to the distal end of the guide catheter. 8. The guide system of claim 6, wherein the guidewire is configured to translate relative to the guide catheter to thereby position at least part of the visible marker of the guidewire distally in relation to the distal end of the guide catheter. 9. The guide system of claim 1, further comprising a dilation catheter slidably disposed along the guidewire, wherein the dilation catheter includes an expandable dilator. 10. The guide system of claim 9, wherein the dilation catheter is slidably disposed in the lumen of the guide catheter. 11. The guide system of claim 1, wherein the guidewire further comprises an atraumatic tip. 12. The guide system of claim 1, wherein the visible marker runs along only one side of the guidewire. 13. The guide system of claim 1, wherein the visible marker is located on an inner radius of the guidewire bend. 14. The guide system of claim 1, wherein the visible marker is located on an outer radius of the guidewire bend. 15. The guide system of claim 1, wherein the guide catheter is rigid. 16. The guide system of claim 1, wherein the guidewire is further configured to emit light through the distal end of the guidewire. 17. The guide system of claim 1, wherein the guidewire is rotatable relative to the guide catheter. 18. An apparatus comprising: (a) a guide catheter comprising: (i) a proximal end,(ii) a distal end, and(iii) a lumen extending from the proximal end to the distal end; and(b) a guidewire extending within the lumen of the guide catheter from the first proximal end to the first distal end, where... | Acclarent, Inc. | A61B001/233 A61B001/07 A61B001/06 A61B001/00 A61B017/24 A61B005/00 A61M025/09 A61B090/00 A61B017/00 A61B017/22 A61B090/30 A61M025/00 A61M025/06 A61M029/02
US10702695 | B2 | 2020-07-07 | Wireless implantable taste system | 2016-08-25 | The system includes various embodiments of components, including (1) a sensor array, (2) a processor, (3) a transmitter, (4) a receiver-stimulator, and (5) an implantable electrode array. The gustatory implant system generates tastant fingerprints by detecting tastants with an array of chemical sensors and then transmitting variable spatiotemporal stimulation patterns for an electrode array with electrode stimulating points positioned at different locations in the gustatory cortex (e.g., stimulating the chorda tympanic nerve). Different patterns of activity in the gustatory cortex are thereby generated which mimic the sense of taste in a subject. Once trained the system should be usable by a subject to detect or correctly identify or perceive one or more tastes. The system may also include an external electronic device for processing and displaying information to monitor ingestion of various substances in a subject. | 1. A system for mimicking a sense of taste in a subject, comprising: a sensor array comprising a plurality of sensors for detecting one or more tastant, said sensor array producing sensor output signals in response to detecting said one or more tastants;a processor for processing said sensor output signals into digital encoded information;a transmitter for transmitting said digital encoded information transcutaneously into said subject;a receiver-stimulator configured to receive said transcutaneously transmitted digital encoded information and generate electrical impulses as a function thereof; andan implantable electrode array comprising at least one stimulating electrode configured to stimulate the gustatory cortex in said subject with said electrical impulses,wherein each of said at least one stimulating electrode of said implantable electrode array is individually controllable. 2. The system as claimed in claim 1, wherein said sensor array is part of an intraoral device selected from the group consisting of a dental bridge, dental implant, artificial tooth, a crown, dental appliance, full denture plate, partial dental plate, and a maxillary appliance. 3. The system as claimed in claim 2, wherein said intraoral device is a maxillary appliance. 4. The system as claimed in claim 3, wherein said maxillary appliance comprises said sensor array fabricated or embedded within an ultrathin, flexible and stretchable material that can be integrated with or be removable from said maxillary appliance. 5. The system as claimed in claim 1, wherein said electrical impulses are associated with spatiotemporal electrical stimulation patterns in a gustatory cortex. 6. The system as claimed in claim 1, wherein said sensor array comprises a plurality of groupings of one or more sensors, each grouping configured to detect a different tastant. 7. A system for mimicking a sense of taste in a subject, comprising: a sensor array comprising a plurality of sensors for detecting one or more tastant, said sensor array producing sensor output signals in response to detecting said one or more tastants;a processor for processing said sensor output signals into digital encoded information;a transmitter for transmitting said digital encoded information transcutaneously into said subject;a receiver-stimulator configured to receive said transcutaneously transmitted digital encoded information and generate electrical impulses as a function thereof; andan implantable electrode array comprising at least one stimulating electrode configured to stimulate the gustatory cortex in said subject with said electrical impulses,wherein said sensor array comprises a plurality of the same sensor, wherein a different filter is associated with each sensor of said plurality of the same sensor which differentially affects the passage or flow of different molecules across the filter. 8. The system as claimed in claim 1, further comprising at least one recording electrode for recording one or ... | Virginia Commonwealth University | A61N001/36 A61N001/05 G09B021/00 G16H020/60 A61B005/00 G09B019/00
US10709385 | B2 | 2020-07-14 | Side-loading connectors for use with intravascular devices and associated systems and methods | 2016-12-09 | Intravascular devices, systems, and methods are disclosed. In some embodiments, side-loading electrical connectors for use with intravascular devices are provided. The side-loading electrical connector has at least one electrical contact configured to interface with an electrical connector of the intravascular device. A first connection piece of the side-loading electrical connector is movable relative to a second connection piece between an open position and a closed position, wherein in the open position an elongated opening is formed between the first and second connection pieces to facilitate insertion of the electrical connector between the first and second connection pieces in a direction transverse to a longitudinal axis of the intravascular device and wherein in the closed position the at least one electrical contact is electrically coupled to the at least one electrical connector received between the first and second connection pieces. | 1. An intravascular system, comprising: an intravascular device comprising: a flexible elongate member having a proximal portion and a distal portion,at least one electronic component secured to the distal portion of the flexible elongate member, andat least one electrical connector secured to the proximal portion of the flexible elongate member, wherein the at least one electrical connector is electrically coupled to the at least one electronic component secured to the distal portion of the flexible elongate member; anda connector having at least one electrical contact configured to interface with the at least one electrical connector of the intravascular device, the connector including a first connection piece and a second connection piece, wherein the first connection piece is movable relative to the second connection piece between an open position and a closed position, wherein in the open position an elongated opening is formed between the first and second connection pieces to facilitate insertion of the at least one electrical connector between the first and second connection pieces in a direction transverse to a longitudinal axis of the intravascular device and wherein in the closed position the at least one electrical contact is electrically coupled to the at least one electrical connector received between the first and second connection pieces. 2. The system of claim 1, wherein the first connection piece is movable relative to the second piece about a pivot axis. 3. The system of claim 2, wherein a bias element urges the first and second connection pieces towards the closed position. 4. The system of claim 3, wherein the bias element is a spring. 5. The system of claim 2, wherein the second connection piece includes a recess sized and shaped to receive a portion of the intravascular device that includes the at least one electrical connector. 6. The system of claim 5, wherein the at least one electrical contact is secured to the first connection piece. 7. The system of claim 1, wherein the first piece includes at least one visual indicator for aligning the at least one electrical contact of the connector with the at least one electrical connector of the intravascular device. 8. The system of claim 7, wherein the at least one visual indicator comprises an arrow. 9. The system of claim 7, wherein the at least one visual indicator is a light. 10. The system of claim 9, wherein the light is configured to illuminate when a proper electrical coupling is achieved between the at least one electrical contact and the at least one electrical connector. 11. The system of claim 9, wherein the light is configured to illuminate a first color when a proper electrical coupling is achieved between the at least one electrical contact and the at least one electrical connector and configured to illuminate a second color when the proper electrical coupling between the at least one electrical contact and the at least one electrical connector is not ac... | PHILIPS IMAGE GUIDED THERAPY CORPORATION | A61B005/00 A61M025/09 A61B005/0215 A61B008/08
US10722055 | B2 | 2020-07-28 | Adjustable pillow device and method | 2015-08-13 | An adjustable pillow device and a pillow adjusting method are disclosed, the adjustable pillow device comprising: a headrest body (1), the height of a region thereof being self-adjustable according to the posture of a sleeper; an inflation/deflation mechanism (2) connected to the headrest body (1) for adjusting the height of a region of the headrest body (1); a sensor (3) for collecting information about the sleeper and providing feedback; a central information processor (4) connected to the sensor (3) and the control end of the inflation/deflation mechanism (2) respectively for receiving information from the sensor (3) and sending an adjustment direction to the control end of the inflation/deflation mechanism (2) according to the sleeping posture information and the body data figure of the sleeper. The adjustable pillow device and method ensure that a sleeper has the most natural and physiologic sleeping posture, and enable automatic adjustment of the head and neck of the sleeper in different sleeping postures to allow the head to be in a proper position relative to the neck, thus providing the sleeper with deeper and longer sleep, improving sleep quality, and meeting the needs of the public. | 1. An adjustable pillow device, comprising: a main body, a height of the main body adjustable according to sleeping posture information and body figure data of a sleeper;an inflation-deflation mechanism, connected to the main body, and programmed to perform regional height adjustment of the main body;a sensor, programmed to collect and feed back the sleeping posture information about the sleeper; anda central information processor, connected to the sensor and to a control terminal of the inflation-deflation mechanism respectively, and configured to acquire the body figure data of the sleeper, receive the sleeping posture information from the sensor, and send an adjusting direction to the control terminal of the inflation-deflation mechanism based on both the sleeping posture information and the body figure data of the sleeper, the body figure data including collected dimensional data of at least one of a head, a neck, shoulders, and a chest of the sleeper,wherein the main body includes a cervical spine supporting area and a head supporting area, the cervical spine supporting area having a height higher than the head supporting area, and each of the cervical spine supporting area and the head supporting area is provided with at least one air bag,wherein the sensor includes an accelerometer configured to be located below the main body, or on the chest, back, hip, or head of the sleeper. 2. The adjustable pillow device of claim 1, wherein the cervical spine supporting area is provided with three air bags sequentially arranged along a length of the main body, and the head supporting area is provided with one air bag. 3. The adjustable pillow device of claim 1, wherein the cervical spine supporting area is provided with three air bags sequentially arranged along a length of the main body, and the head supporting area is provided with three air bags sequentially arranged along the length of the main body. 4. The adjustable pillow device of claim 1, wherein the inflation-deflation mechanism includes an air pump, and a pressure tubing connecting the air pump to the air bag, the air pump has a control terminal electrically connected to the central information processor, and the pressure tubing is provided with a switching valve electrically connected to the central information processor. 5. The adjustable pillow device of claim 4, wherein the air pump is a silent air pump. 6. The adjustable pillow device of claim 1, wherein the cervical spine supporting area and/or the head supporting area are provided with a temperature adjuster. 7. The adjustable pillow device of claim 1, further comprising plates arranged in the head supporting area and/or the cervical spine supporting area respectively, the plate arranged in the head supporting area is positioned away from the top of the head supporting area, and the plate arranged in the cervical spine supporting area is positioned away from the top of the cervical spine supporting area. 8. The adjustable pill... | - | A47G009/10 A61B005/00 A61F005/56 A61B005/024 A61B005/145
US10716613 | B2 | 2020-07-21 | Photonic probe apparatus with integrated tissue marking facility | 2013-02-08 | The invention relates to a photonic probe apparatus and a method for probing tissue to detect and mark biological tissue with cancerous or precancerous states. The apparatus involves a probe for illuminating tissue and collecting light from an illuminated tissue region through the probe, a unit for analyzing collected light to determine whether a threshold measure of probability of a cancerous or precancerous lesion in the probed tissue region in contact with the probe is exceeded, and an integrated tissue marking facility which can be activated to mark the probed tissue region through the probe when the threshold measure is exceeded. The photonic probe apparatus and the method are especially suitable for probing regions in squamous and columnar epithelia to detect and mark regions with cervical cancer or cervical intraepithelial neoplasia (CIN). | 1. A photonic probe apparatus for marking biological tissue and comprising: a probe for transmitting light to tissue and collecting light from an illuminated tissue region in contact with the probe;a light source operably connected to the probe for illuminating tissue;a light sensor operably connected to the probe to receive and detect light collected from the tissue region in contact with the probe;an analyzing unit for determining a measure of probability of a cancerous or precancerous lesion in the tissue in contact with the probe and whether a threshold for the measure of probability of a cancerous or precancerous lesion in the tissue region in contact with the probe is exceeded based on an output signal from the light sensor; andan integrated tissue marking facility which can be activated to mark the tissue region in contact with the probe when the threshold measure is exceeded, wherein:the tissue marking facility comprises an electric current source connected to supply an output electrode situated in the probe with sufficient current to coagulate tissue in contact with the probe, and the probe comprises a replaceable probe tip wherein the light source, light sensor, and tissue marking facility works through the replaceable tip of the probe. 2. The photonic probe apparatus according to claim 1, wherein the analyzing unit comprises: an electronic processor connected to receive the output signal from the light sensor; anda memory holding a computer program configured to, when executed by the electronic processor:determine one or more measures of probability of a cancerous or precancerous lesion in the tissue region based on the output signal from the light sensor; anddetermine whether one or more of the determined measures of probability exceeds a predetermined threshold. 3. The photonic probe apparatus according to claim 1, wherein determining a measure of probability of cancerous or precancerous lesions involves distinguishing precancerous and/or cancerous tissue from not precancerous tissue based on the output signal from the light sensor. 4. The photonic probe apparatus according to claim 2, wherein the electronic processor is connected to control the tissue marking facility, and wherein the computer program is configured to automatically activate the tissue marking facility to mark tissue when the threshold measure is exceeded. 5. The photonic probe apparatus according to claim 2, further comprising an audio and/or visual output connected to the electronic processor, and wherein the computer program is configured to control the audio and/or visual output to indicate to a user whether the threshold measure is exceeded. 6. The photonic probe apparatus according to claim 1, wherein the probe further comprises an interface allowing a user to activate the integrated tissue marking facility to mark tissue. 7. The photonic probe apparatus according to claim 1, wherein the light sensor is configured to detect an intensity of collected light... | KONINKLIJKE PHILIPS N.V. | A61B018/14 A61B005/00 A61B001/04 A61B090/00 A61B001/06 A61B001/07 A61B001/00 A61B017/00 A61B018/00
US10706730 | B2 | 2020-07-07 | Perceptual-cognitive-motor learning system and method | 2013-02-22 | The present disclosure relates to a perceptual-cognitive-motor learning system. The system includes an apparatus for evaluating or improving perceptual-cognitive abilities of a subject during a training sequence. The system also has a training sequence controller for adding in at least a part of the training sequence at least one of (a) a first motor load add-on to the subject and (b) a second motor load add-on to the subject, the second motor load being heavier than the first motor load. A variant of the system has a user interface for allowing the subject to change at least one parameter of the training sequence. Methods for evaluating or improving perceptual-cognitive abilities of a subject are also disclosed. | 1. A perceptual-cognitive-motor learning system, comprising: an apparatus for improving perceptual-cognitive abilities of a subject during a plurality of successive tests, the apparatus including:a three-dimensional display apparatus configured to display stereoscopic images of virtual objects moving in a given three-dimensional environment during the plurality of successive tests, the display apparatus providing a visual contact between the virtual objects and a subject during the plurality of successive tests;a user interface for providing commands issued by the subject to a training sequence controller,the training sequence controller being operatively connected to the user interface for receiving the commands issued by the subject and to the three-dimensional display apparatus, the training sequence controller being configured to: define a training sequence, the training sequence including a first plurality of successive tests, a second plurality of successive tests having a first motor load causing no efferent signal from the subject's brain and a third plurality of successive tests having a second motor load, the second motor load being heavier than the first motor load, wherein the virtual objects move in the three-dimensional environment during the first, second and third plurality of successive tests;cause the display apparatus to present the virtual objects in accordance with the first plurality of successive tests to the subject;receive, through the user interface, responses from the subject, the responses identifying the virtual objects during the first plurality of successive tests;receive, through the user interface, commands from the subject, the commands causing the training sequence controller to perform at least one of: causing the display apparatus to change a movement speed of the virtual objects;causing the display apparatus to reset the movement of the virtual objects; andsetting a tracking speed of the virtual objects and presenting again the first plurality of successive tests to the subject, the movement speed of the virtual objects corresponding to the tracking speed;cause the display apparatus to present the virtual objects in accordance with the second plurality of successive tests to the subject;receive, through the user interface, responses from the subject, the responses identifying the virtual objects during the second plurality of successive tests;cause the display apparatus to present the virtual objects in accordance with the third plurality of successive tests to the subject;receive, through the user interface, responses from the subject, the responses identifying the virtual objects during the third plurality of successive tests;adjust the movement speed of the virtual objects in view of the responses received from the subject. 2. A system as defined in claim 1, wherein each test of the training sequence includes a core exercise. 3. A system as defined in claim 1, wherein the second motor load is a specific ... | COGNISENS INC. | G09B005/00 G09B023/28 G09B019/00 A61B005/11
US10709392 | B2 | 2020-07-14 | Sensor fault detection using analyte sensor data pattern comparison | 2018-08-23 | Analyte sensor faults are detected. Datasets of glucose values sensor electronics are coupled to a glucose sensor in fluid contact with interstitial fluid under a skin surface. Baseline median glucose value and glucose variability values are computed, based on the first dataset. A baseline data point is stored. Evaluation median glucose value and variability are computed, based on the second dataset of glucose values. An evaluation data point is stored. A magnitude of a vector that extends between the baseline data point and the evaluation data point is computed. A component of the magnitude of the vector that is parallel to a hypoglycemia risk contour line is computed and compared to a predefined threshold value. An indication that a sensor fault has been detected if the component is greater than a threshold is displayed. | 1. A computer-implemented method, comprising: computing a magnitude of a vector extending from a baseline analyte data point to an evaluation analyte data point;computing a component of the magnitude of the vector between the baseline analyte data point and the evaluation analyte data point that is parallel to a hypoglycemic risk contour line;determining whether the component of the magnitude of the vector is greater than a first threshold value; anddisplaying, based on the determination, an analyte sensor fault indication on a display. 2. The computer-implemented method of claim 1, further comprising: determining a baseline median glucose value and a baseline glucose variability value; anddetermining the baseline analyte data point based on the baseline median glucose value and the baseline glucose variability value. 3. The computer-implemented method of claim 2, wherein determining the baseline median glucose value and the baseline glucose variability value comprises determining the baseline median glucose value and the baseline glucose variability value based on a first dataset of glucose values. 4. The computer-implemented method of claim 2, further comprising: determining an evaluation median glucose value and an evaluation glucose variability value; anddetermining the evaluation analyte data point based on the evaluation median glucose value and the evaluation glucose variability value. 5. The computer-implemented method of claim 4, wherein determining the baseline median glucose value and the baseline glucose variability value comprises determining the baseline median glucose value and the baseline glucose variability value based on a first dataset of glucose values, wherein determining the evaluation median glucose value and the evaluation glucose variability value comprises determining the evaluation median glucose value and the evaluation glucose variability value based on a second dataset of glucose values, and wherein the second dataset of glucose values includes glucose values that are not in the first dataset of glucose values. 6. The computer-implemented method of claim 1, wherein the analyte sensor fault indication is indicative that an analyte sensor fault has been detected. 7. The computer-implemented method of claim 6, wherein determining whether the component of the magnitude of the vector is greater than the first threshold value comprises determining that the component of the magnitude of the vector is greater than the first threshold value. 8. The computer-implemented method of claim 6, wherein the analyte sensor fault indication includes information regarding whether to maintain use of a glucose sensor. 9. The computer-implemented method of claim 6, wherein the analyte sensor fault indication includes a request for a reference glucose measurement. 10. The computer-implemented method of claim 1, wherein the analyte sensor fault indication is indicative that an early signal attenuation (ESA) fault has been detected.... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/145 A61B005/1459
US10709397 | B2 | 2020-07-14 | Mobile X-ray apparatus | 2019-06-14 | A mobile X-ray apparatus includes: an X-ray radiation device; a controller configured to control the X-ray radiation device; a power supply configured to supply operating power to the X-ray radiation device and the controller via a lithium ion battery and control overcurrent occurring during X-ray emission by the X-ray radiation device; and a charger configured to charge the power supply. Each of the controller, the power supply, and the charger is embodied in a physically separate module, and each of the power supply and the charger is encased in a metal case. | 1. A mobile X-ray apparatus comprising: an X-ray radiation device;a controller configured to control the X-ray radiation device;a power supply configured to: supply operating power to the X-ray radiation device and the controller via a lithium ion battery; andcontrol an overcurrent that occurs during an X-ray emission by the X-ray radiation device;a charger configured to charge the lithium ion battery; anda main body configured to house the controller, the power supply, and the charger,wherein each of the controller, the power supply, and the charger is embodied as a physically separate module,wherein the power supply is encased in a metal case,wherein the power supply includes a master BMS circuit configured to operate a protection circuit against the overcurrent, based on a state of the lithium ion battery, andwherein the master BMS circuit is arranged parallel to a first surface of the metal case of the power supply. 2. The mobile X-ray apparatus of claim 1, wherein the power supply further includes a shutdown circuit arranged adjacent to the master BMS circuit and parallel to the first surface of the metal case of the power supply. 3. The mobile X-ray apparatus of claim 1, wherein the power supply further includes a slave BMS circuit configured to manage at least one voltage of the lithium ion battery, temperatures of the lithium ion battery, and unbalancing between cells of the lithium ion battery, and wherein the slave BMS circuit is arranged parallel to a second surface of the metal case that is perpendicular to the first surface. 4. The mobile X-ray apparatus of claim 1, wherein the power supply further includes a slave BMS circuit configured to manage at least one voltage of the lithium ion battery, temperatures of the lithium ion battery, and unbalancing between cells of the lithium ion battery, wherein the power supply further includes a partition wall including an insulation material, andwherein the partition wall is installed between the lithium ion battery and the slave BMS circuit. 5. The mobile X-ray apparatus of claim 1, wherein the power supply further includes a partition wall including an insulation material, and wherein the partition wall is installed between the lithium ion battery and the master BMS circuit. 6. A mobile X-ray apparatus comprising: an X-ray radiation device;a controller configured to control the X-ray radiation device;a power supply configured to: supply operating power to the X-ray radiation device and the controller via a lithium ion battery; andcontrol an overcurrent that occurs during an X-ray emission by the X-ray radiation device;a charger configured to charge the lithium ion battery; anda main body configured to house the controller, the power supply, and the charger,wherein each of the controller, the power supply, and the charger is embodied as a physically separate module,wherein the charger is encased in a metal case,wherein the power supply includes a master BMS circuit configured to operate... | SAMSUNG ELECTRONICS CO., LTD. | A61B006/00 A61B006/10 G01T001/175 H02J007/00 H02J007/02 H02H007/18 H05G001/12 A61B005/00 G01R031/50 H01M010/052 H02H003/08 G01R031/36 H01M010/0525 H01M010/48 H01M010/42
US10728519 | B2 | 2020-07-28 | Method and apparatus for colour imaging a three-dimensional structure | 2019-04-02 | A device for determining the surface topology and associated color of a structure, such as a teeth segment, includes a scanner for providing depth data for points along a two-dimensional array substantially orthogonal to the depth direction, and an image acquisition means for providing color data for each of the points of the array, while the spatial disposition of the device with respect to the structure is maintained substantially unchanged. A processor combines the color data and depth data for each point in the array, thereby providing a three-dimensional color virtual model of the surface of the structure. A corresponding method for determining the surface topology and associate color of a structure is also provided. | 1. A system for determining surface topology and associated color of at least a portion of a three-dimensional structure, comprising: a hand-held device comprising: a scanning system configured to provide depth data of said portion, the depth data corresponding to a plurality of data points defined on a plane substantially orthogonal to a depth direction;an imaging system configured to provide two-dimensional color image data of said portion associated with said plurality of data points;an illumination unit configured to transmit a first array of incident light along a path towards the three-dimensional structure; anda detector configured to measure intensity of each of a plurality of returned light that returns along the path and from the three-dimensional structure; anda processor operably coupled to the hand-held device and configured to associate the depth data with the two-dimensional color image data. 2. The system according to claim 1, wherein the imaging system is configured to provide the tow-dimensional color image data independently from the depth data. 3. The system according to claim 1, wherein the plurality of data points is associated with a two-dimensional reference array. 4. The system according to claim 1, wherein the operation of the scanning system is based on confocal imaging techniques, and wherein the scanning system is configured to: cause the illumination system to transmit the first array of incident light along the path towards the three-dimensional structure;cause the detector to measure intensity of each of the plurality of returned light that return along the path and from the three-dimensional structure. 5. The system according to claim 1, wherein the depth data and the two-dimensional color image data are associated by aligning the plurality of data points and the two-dimensional color image data in the same frame of reference. 6. The system according to claim 1, wherein the associated depth data and the two-dimensional color image data represent the color and surface topology of the portion of the three-dimensional structure. 7. The system according to claim 1, wherein the imaging system is configured to: cause the illumination system to illuminate the portion with a set of three or more, differently-colored illumination radiations, the differently-colored illumination radiations being combinable to provide white light;cause the detector to capture a monochromatic image of the portion for each of the three or more, differently-colored illumination radiations; andcombine each of the monochromatic images to create the two-dimensional color image data. 8. The system according to claim 7, comprising a color illumination system configured to provide the set of three or more, differently-colored illumination radiations. 9. The system according to claim 8, wherein the color illumination system comprises one of the following: (a) a first illumination means for providing the set of three or more, differently-colore... | ALIGN TECHNOLOGY, INC. | H04N013/15 G01B011/25 H04N013/207 H04N013/257 H04N013/271 H04N013/296 A61B001/00 A61B001/06 A61B001/24 A61B001/247 A61B005/00 A61B005/107 G01B011/24 G01J003/02 G01J003/10 G01J003/50 G06T007/12 G06T007/90 A61C009/00 G06T007/00 G01N021/25 H01L027/148 A61C019/04 G01J003/46 G01J003/51
US10709340 | B2 | 2020-07-14 | Automated fitted cuff blood pressure and arm circumference measuring device | 2017-12-29 | A blood pressure device having a cylindrical drum cuff housing adapted to measure the circumference of either a right or left arm and administer a blood pressure test. The drum cuff housing is connected to a knuckle support assembly adapted to provide freedom of movement for postural positioning before and during the blood pressure test. | 1. A blood pressure device, comprising: a cylindrical drum cuff;a direct drive reel assembly having a motor entirely inside of a reel that rotates the reel and an encoder mounted to the motor;a flexible lead band material is connected to and wraps around the reel and a sensor is associated with the flexible lead band material; anda motor control board determines a positive arm circumference stop and a second pre-set stop based on signals from the encoder,wherein the blood pressure device is configured to operate with a blood pressure module to acquire biometric data. 2. The device of claim 1 wherein the device is mounted to one of a group consisting of a table top and a kiosk. 3. The device of claim 1 further comprising a motor control board configured to operate with the blood pressure module. 4. The device of claim 1 wherein the blood pressure module acquires biometric data with a method selected from a group consisting of an oscillometric method, an auscultatory method, and through the presence of arterial fibrillation. 5. A blood pressure device, comprising: a cylindrical drum cuff housing;a direct drive reel assembly having a motor entirely inside of a reel that rotates the reel and an encoder mounted to the motor;a knuckle support assembly attached to the cylindrical drum cuff housing and adapted to provide freedom of movement for postural positioning before and during a blood pressure test; anda flexible bellow that houses the knuckle support assembly, wherein the blood pressure device is configured to operate with a blood pressure module to acquire biometric data. 6. The device of claim 5 wherein the knuckle support assembly is adapted to provide ergonomic body adjustment and alignment of an arm in relation to a level of a heart. 7. The device of claim 5 wherein the knuckle support assembly tilts with spring tension. 8. The device of claim 7 wherein the knuckle support assembly rotates to improve a position of the arm. 9. The device of claim 7 wherein the knuckle support assembly slides to position the arm closer to heart level. 10. The device of claim 5 wherein the blood pressure device is configured to support the arm from an elbow to a wrist to permit arm muscles to relax during the blood pressure test. 11. A blood pressure device, comprising: a cylindrical drum cuff housing;a direct drive reel assembly that includes a motor entirely inside of a reel to rotate the reel and an encoder mounted to the motor;a flexible lead band material connected to and wrapped around the reel;a sensor associated with the flexible lead band material, wherein the sensor is selected from group consisting of an analog bar label, an infrared sensor, and an encoder; anda bladder attached to an inner surface of the flexible lead band material;wherein the blood pressure device is configured to operate with a blood pressure module to acquire biometric data. 12. The device of claim 11 further comprising a protective skirt fitted over the bladder and the... | HIGI SH LLC | A61B005/022 A61B005/021 A61B005/00
US10718842 | B2 | 2020-07-21 | Systems and methods for automated detection in magnetic resonance images | 2017-11-21 | Some aspects include a method of determining change in size of an abnormality in a brain of a patient positioned within a low-field magnetic resonance imaging (MRI) device. The method comprises, while the patient remains positioned within the low-field MRI device, acquiring first and second magnetic resonance (MR) image data of the patient's brain; providing the first and second MR image data as input to a trained statistical classifier to obtain corresponding first and second output; identifying, using the first output, at least one initial value of at least one feature indicative of a size of the abnormality; identifying, using the second output, at least one updated value of the at least one feature; determining the change in the size of the abnormality using the at least one initial value of the at least one feature and the at least one updated value of the at least one feature. | 1. A method of determining change in size of an abnormality in a brain of a patient positioned within a low-field magnetic resonance imaging (MRI) device, the method comprising: while the patient remains positioned within the low-field MRI device: acquiring first magnetic resonance (MR) image data of a patient's brain;providing the first MR image data as input to a trained statistical classifier to obtain corresponding first output;acquiring second MR image data of the patient's brain subsequent to acquiring the first MR image data;providing the second MR image data as input to the trained statistical classifier to obtain corresponding second output;determining the change in the size of the abnormality in the patient's brain using the first output and the second output. 2. The method of claim 1, further comprising: while the patient remains positioned within the low-field MRI device: identifying, using the first output, at least one initial value of at least one feature indicative of a size of an abnormality in the patient's brain;identifying, using the second output, at least one updated value of the at least one feature indicative of the size of the abnormality in the patient's brain,wherein determining the change in the size of the abnormality is performed using the at least one initial value of the at least one feature and the at least one updated value of the at least one feature. 3. The method of claim 2, wherein identifying, using the first output, the at least one initial value of the at least one feature indicative of the size of the abnormality in the patient's brain comprises:identifying one or more first values indicative of a first diameter of the abnormality. 4. The method of claim 3, wherein the identifying further comprises identifying one or more second values indicative of a second diameter of the abnormality orthogonal to the first diameter. 5. The method of claim 2, wherein determining the change in the size of the abnormality using the at least one initial value of the at least one feature and the at least one updated value of the at least one feature comprises: determining an initial size of the abnormality using the at least one initial value of the at least one feature;determining an updated size of the abnormality using the at least one updated value of the at least one feature; anddetermining the change in the size of the abnormality using the determined initial and updated sizes of the abnormality. 6. The method of claim 1, wherein the trained statistical classifier comprises a multi-layer neural network. 7. The method of claim 1, wherein the trained statistical classifier comprises a convolutional neural network. 8. The method of claim 1, wherein the trained statistical classifier comprises a fully convolutional neural network. 9. The method of claim 1, wherein the second MR image data is obtained within one hour of the first MR image data. 10. The method of claim 1, further comprising repeating acquiring MR i... | Hyperfine Research, Inc. | G01R033/00 G01R033/56 G01R033/38 G06K009/62 A61B005/00 A61B005/055 G01R033/44 G01R033/483 G06T007/00 G01R033/48 G01R033/383 G06N003/04
US10716621 | B2 | 2020-07-21 | Stenosis prevention and ablation delivery system | 2017-10-25 | Various aspects of the present disclosure are directed toward apparatuses, systems, and methods for applying ablation therapy to a tissue region. The apparatuses, systems, and methods may include a balloon structure and one or more electrodes arranged on or within the balloon structure and configured to deliver energy to the tissue region. | 1. An apparatus for applying ablation therapy to a tissue region, the apparatus comprising: a catheter sized and shaped for vascular access and including an elongate body extending between a proximal end and a distal end;a balloon structure arranged near the distal end of the elongate body and having a first portion and a second portion, the first portion of the balloon structure being configured to permeate a liquid therethrough and the second portion of the balloon structure being configured to anchor the elongate body at the tissue region;one or more mapping electrodes arranged on or within the balloon structure and configured to determine a target location for the ablation therapy;one or more electrodes arranged on or within the balloon structure and configured to deliver energy to the tissue region; anda tip electrode arranged at the distal end of the catheter and configured to form a ground or a closed-loop with the one or more electrodes. 2. The apparatus of claim 1, wherein an external surface of the first portion of the balloon structure is configured to transfer the energy from the one or more electrodes to the tissue region. 3. The apparatus of claim 2, wherein the one or more electrodes comprises an electrode arranged within the first portion of the balloon structure. 4. The apparatus of claim 3, wherein the electrode is configured to deliver the energy via the first portion of the balloon structure in response to a direct current applied thereto. 5. The apparatus of claim 4, wherein the liquid comprises at least one of saline, a pharmacological agent, and an anti-stenotic agent, and the liquid is configured to mitigate against stenosis at the tissue region. 6. The apparatus of claim 1, wherein the second portion of the balloon structure is arranged within the first portion of the balloon structure. 7. The apparatus of claim 6, wherein the first portion forms a first chamber of the balloon structure, and the second portion forms a second chamber of the balloon structure. 8. The apparatus of claim 7, wherein the elongate body includes a first opening arranged within the first chamber and the elongate body includes a second opening arranged within the second chamber, the first portion is configured to elute a liquid therethrough in response to influx of the liquid into the first chamber through the first opening, and the second portion is configured to expand and anchor the elongate body at the tissue region in response to influx of a second liquid into the second chamber through the second opening. 9. The apparatus of claim 1, wherein the balloon structure is configured to telescope from the elongate body prior to inflation thereof. 10. The apparatus of claim 1, further comprising a steering mechanism configured to direct at least one of the balloon structure and the elongate body. 11. The apparatus of claim 10, wherein the steering mechanism comprises at least one wire coupled to a catheter handle. 12. An apparatus for appl... | Boston Scientific Scimed, Inc. Mayo Foundation for Medical Education and Research | A61B018/14 A61B018/08 A61M025/10 A61B018/00 A61B005/046
US10702153 | B2 | 2020-07-07 | Wound dressing with reusable electronics for wireless monitoring | 2016-04-15 | A wound dressing device with reusable electronics for wireless monitoring and a method of making the same are provided. The device can be a smart device. In an embodiment, the device has a disposable portion including one or more sensors and a reusable portion including wireless electronics. The one or more sensors can be secured to a flexible substrate and can be printed by non-contact printing on the substrate. The disposable portion can be removably coupled to the one or more sensors. The device can include one or more sensors for wireless monitoring of a wound, a wound dressing, a body fluid exuded by the wound and/or wearer health. | 1. A wound dressing system comprising: a disposable substrate portion, which includes a first substrate, a first electrode of a sensor and a bandage pad interposed between a first side of the first substrate and the first electrode; anda reusable portion, which includes a flexible substrate having a first side removably coupled to a second side of the first substrate, wherein sensor electronics and a second electrode of the sensor are arranged on a second side of the flexible substrate and the second electrode is electrically coupled to the sensor electronics. 2. The wound dressing system of claim 1, wherein the first electrode of the sensor is configured to allow bodily fluids to pass through it to the bandage pad. 3. The wound dressing system of claim 1, wherein the first and second electrodes form a capacitive sensor. 4. The wound dressing system of claim 3, wherein the capacitive sensor senses bodily fluids. 5. The wound dressing system of claim 4, wherein the capacitive sensor senses pressure applied to the wound dressing system. 6. The wound dressing system of claim 1, wherein the first electrode is coupled to the sensor electronics to form a resistive sensor. 7. The wound dressing system of claim 1, further comprising: an antenna arranged on the flexible substrate and coupled to the sensor electronics. 8. The wound dressing system of claim 7, wherein the antenna is a loop antenna and the sensor electronics are arranged inside of the loop antenna. 9. The wound dressing system of claim 1, wherein the flexible substrate is an adhesive tape. 10. The wound dressing system of claim 1, wherein the sensor electronics comprise: a wireless transmitter; anda capacitance to digital converter coupled to the first and second electrodes and to the wireless transmitter. 11. The wound dressing system of claim 10, wherein the wireless transmitter includes a microcontroller configured to store readings from the sensor. 12. A method of making a wound dressing system, the method comprising: arranging a first electrode of a sensor on a first side of a bandage pad, and arranging a second side of the bandage pad on a first side of a first substrate;arranging sensor electronics and a second electrode on a first side of a flexible substrate; andremovably coupling a second side of the flexible substrate with a second side of the first substrate, wherein the step of removably coupling the flexible substrate and the first substrate electrically couples the first electrode with the sensor electronics. 13. The method of making a wound dressing system of claim 12, further comprising: arranging a loop antenna on the first side of the flexible substrate and surrounding the sensor electronics. 14. The method of making a wound dressing system of claim 12, wherein the arranging of the first electrode of the sensor further comprises arranging a portion of the first electrode on the second side of the first substrate so that the first electrode is electrically coup... | KING ABDULLAH UNIVERSITY OF SCIENCE AND TECHNOLOGY | A61B005/00 A61F013/00 A61B005/145 A61B005/02 A61B005/1477
US10722181 | B2 | 2020-07-28 | Systems, methods, and computer-readable media for using descriptors to identify when a subject is likely to have a dysmorphic feature | 2018-06-26 | Systems, methods, and computer-readable media are disclosed for identifying when a subject is likely to be affected by a medical condition. For example, at least one processor may be configured to receive information reflective of an external soft tissue image of the subject. The processor may also be configured to perform an evaluation of the external soft tissue image information and to generate evaluation result information based, at least in part, on the evaluation. The processor may also be configured to predict a likelihood that the subject is affected by the medical condition based, at least in part, on the evaluation result information. | 1. A computer-implemented method for determining, from a series of pixels in an image of a subject, whether the subject is likely to be affected by each of one or more medical conditions, the computer-implemented method comprising: using a plurality of images of positive and negative individuals, calculating, with processing circuitry, descriptors derived from features in each of a plurality of cells within a grid imposed on each of the plurality of images of a first set of first positive individuals having a first medical condition and each of the plurality of images of at least one negative individual who lacks the first medical condition;producing, with the processing circuitry, a plurality of vectors by aggregating the descriptors of the plurality of cells in the plurality of images of the first positive individuals and by aggregating the descriptors of the plurality of cells in the plurality of images of the at least one negative individual;using an image of the subject, calculating, with the processing circuitry, descriptors derived from features in each of a plurality of cells within a grid imposed on the image of the subject;producing, with the processing circuitry, one or more vectors by aggregating the descriptors of the plurality of cells in the image of the subject;implementing, with the processing circuitry, a comparison of the plurality of vectors of the images of the first positive individuals and the at least one negative individual with the one or more vectors of the image of the subject, by computing a distance metric between each of the plurality of vectors and each of the one or more vectors;determining, with the processing circuitry, based on the comparison, a subset of the plurality of vectors more similar to the one or more vectors than a remainder of the plurality of vectors; anddetermining, with the processing circuitry, a first probability score indicating a likelihood that the subject has the first medical condition by analyzing whether vectors in the determined subset of the plurality of vectors exceed a threshold of similarity with the one or more vectors derived from the image of the subject. 2. The computer-implemented method of claim 1, further comprising calculating descriptors derived from features in each of a plurality of cells within a grid imposed on each of images of a second set of second positive individuals having a second medical condition, and aggregating the descriptors derived from the images of the second set of second positive individuals to produce the plurality of vectors. 3. The computer-implemented method of claim 2, further comprising determining a second probability score indicating a likelihood that the subject has the second medical condition by analyzing whether the vectors in the determined subset of the plurality of vectors correspond to the second set of second positive individuals and exceed the threshold of similarity with the one or more vectors derived from the image of the subject... | FDNA, INC. | A61B005/00 G06K009/00 G06T007/00
US10702157 | B2 | 2020-07-07 | Magnetic resonance imaging apparatus and magnetic resonance imaging method | 2014-12-03 | A magnetic resonance imaging apparatus acquires first imaging data of a three-dimensional image including a heart, having a plurality of two-dimensional first imaging area data superimposed one on top of another in parallel, and having a resolution at least in one direction different from a resolution in two other directions. A first axis is detected expressed in three dimensions relating to the heart from the three-dimensional first imaging data. A first vector is calculated passing through the first axis and having at least a predetermined resolution. First image data is generated on a plane passing through the first axis and the first vector from the first imaging data and a second axis is detected relating to the heart from the first image data, the second axis being a higher precision axis. | 1. A magnetic resonance imaging (MRI) apparatus comprising: (a) an MRI system configured to acquire MR images;(b) at least one data processor configured to: (i) receive from said MRI system first MR imaging data that includes data representing a three-dimensional image of a heart;(ii) detect data representing a characteristic anatomic portion of the heart included in the first MR imaging data, wherein the at least one data processor is further configured to: acquire the first imaging data so as to include a heart and a plurality of cross-sectional two dimensional images that are parallel to each other;detect a first axis relating to the heart from the first imaging data;generate a first generated image from the first imaging data, on a plane containing the first axis, and detect a second axis relating to the heart from the acquired first image data based on the first generated image when resolution of the first generated image is higher than a predetermined resolution;(iii) determine a plane to be scanned by said MRI system, based on a position of the second axis relating to the heart; and(iv) cause said MRI system to perform a further MRI scan and acquire second imaging data in the plane; and(c) a display configured to display the second imaging data superimposed with the second axis detected from the first imaging data. 2. The MRI apparatus according to claim 1, wherein the characteristic anatomic portion of the heart is an axis of the heart and/or a point on the axis of the heart. 3. The MRI apparatus according to claim 2, wherein said axis of the heart is a long axis of the heart; and said point is a center of mitral valve or a cardiac apex. 4. A magnetic resonance imaging (MRI) method comprising: receiving first MR imaging data that includes data representing a three-dimensional image of a heart;detecting data representing a characteristic anatomic portion of the heart included in the first MR imaging data;acquiring the first imaging data so as to include the heart and a plurality of cross-sectional two-dimensional images that are parallel to each other,detecting a first axis relating to the heart from the first imaging data,generating a first generated image from the first imaging data, on a plane containing the first axis, and detecting a second axis relating to the heart from the acquired first imaging data based on the first generated image, when resolution of the first generated image is higher than a predetermined resolution;determining a plane to be scanned by said MRI system to acquire further imaging data, based on a position of the second axis relating to the heart;performing a further MRI scan with the MRI system acquiring second MR imaging data in the plane; anddisplaying the second imaging data superimposed with the second axis detected from the first imaging data. 5. A non-transitory program stored in a computer-readable medium, said program, when executed by at least one computer coupled to a magnetic resonance imaging (MR... | CANON MEDICAL SYSTEMS CORPORATION | A61B005/00 G01R033/483 A61B005/055 G01R033/48
US10709944 | B2 | 2020-07-14 | Devices, systems, and methods | 2018-02-28 | Devices systems and methods are disclosed. Among other things, the devices systems and methods can facilitate ergonomic gripping of a handle; conform to the shape of a wrist; enhance golf techniques or performance; enhance racket sport technique or performance; protect a head of a user engaged in contact sports or other hazardous activities; identify concussions, in some cases in substantially real time; provide sterilization of garments; control surgical robots; allow for medical treatment of multiple subjects without disrobing; provide biometric or other data; and/or be used to train muscles or body parts for, for example, performing specified tasks using a body part. | 1. A method of training a body part, the method comprising: coupling a muscle stimulus generator to the body part, wherein the muscle stimulus generator is configured to generate a muscle activation stimulus and a muscle relaxation stimulus;generating a muscle activation stimulus and a muscle relaxation stimulus using the generator, wherein the step of generating the muscle activation stimulus and the muscle relaxation stimulus comprises generating an alternating sequence of muscle activation stimuli and muscle relaxation stimuli; andtransmitting the generated alternating sequence of muscle activation stimuli and muscle relaxation stimuli to an agonist muscle and a corresponding antagonist muscle, respectively, of the body part to alternatingly i) activate the agonist muscle and relax the antagonist muscle and ii) relax the agonist muscle and activate the antagonist muscle. 2. A method as set forth in claim 1, wherein the step of generating the muscle activation stimulus and the muscle relaxation stimulus comprises sending programmed control instructions to the muscle stimulus generator. 3. A method as set forth in claim 2, wherein the programmed control instructions are configured to control the muscle stimulus generator to generate a predefined muscle activation stimulus and a predefined muscle relation stimulus. 4. A method as set forth in claim 3, wherein the step of transmitting the muscle activation stimulus to the agonist muscle and the muscle relaxation stimulus to the antagonist muscle activates the agonist muscle and relaxes the antagonist muscle to move the body part in a specified movement. 5. A method as set forth in claim 2, further comprising recording a muscle activation sequence of another body part while performing a specified movement. 6. A method as set forth in claim 5, further comprising programming the control instructions based on the recorded muscle activation sequence. 7. A method as set forth in claim 1, wherein said generating the muscle activation stimulus comprises generating a noxious stimulus. 8. A method as set forth in claim 1, wherein the step of generating the muscle activation stimulus and the muscle relaxation stimulus comprises generating at least one of an electrical stimulus, a thermal stimulus, an optical stimulus, and a mechanical stimulus. 9. A method as set forth in claim 1, wherein the step of transmitting the generated muscle activation stimulus and muscle relaxation stimulus comprises at least one of imparting TENS to the body part, imparting Russian stimulation to the body part, heating the body part, vibrating the body part, and imparting laser energy to the body part. 10. A method as set forth in claim 1, further comprising sensing a response of the body part to the transmitted muscle activation stimulus and muscle relation stimulus. 11. A method as set forth in claim 10, wherein the step of sensing the response of the body part comprises sensing the response of the body part using one ... | P Tech, LLC | A61B034/00 A61N001/36 A61B005/0488 A63B060/12 A63B024/00 A63B053/04 A63B053/14 A63B060/46 A63B071/14 A42B003/12 A63B069/36 A63B049/02 A63B051/08 A63B049/08 A63B060/14 A63B051/00 A63B053/10 A61L002/10 A61L002/22 A63B060/04 A63B060/02 A61B090/00 A63B060/00 A63B071/06 A61B034/30
US10709387 | B2 | 2020-07-14 | Radio frequency coil methods and apparatus | 2016-05-12 | Aspects relate to providing radio frequency components responsive to magnetic resonance signals. According to some aspects, a radio frequency component comprises at least one coil having a conductor arranged in a plurality of turns oriented about a region of interest to respond to corresponding magnetic resonant signal components. According to some aspects, the radio frequency component comprises a plurality of coils oriented to respond to corresponding magnetic resonant signal components. According to some aspects, an optimization is used to determine a configuration for at least one radio frequency coil. | 1. A radio frequency component configured to be responsive to magnetic resonance signals, the radio frequency component comprising: a first coil, including a first conductor arranged in a plurality of turns, oriented to be responsive to first magnetic resonance signal components; anda second coil, including a second conductor arranged in a plurality of turns, oriented to be responsive to second magnetic resonance signal components, wherein the first coil and the second coil are tuned to resonate at a target frequency, and wherein a number of turns of the first conductor and the second conductor are limited so that a self-resonance of the first conductor and a self-resonance of the second conductor are each respectively at a frequency at least twice the target frequency. 2. The radio frequency component of claim 1, wherein the first conductor has a length of at least 1 meter and/or is configured with at least 5 turns, and whereinthe second conductor has a length of at least 1 meter and/or is configured with at least 5 turns. 3. The radio frequency component of claim 2, wherein the first magnetic resonance signal components are oriented along a first principal axis, and wherein the second magnetic resonance signal components are oriented along a second principal axis oriented differently than the first principal axis. 4. The radio frequency component of claim 3, wherein the first principal axis is substantially orthogonal to the second principal axis so that the first coil and the second coil are responsive to substantially orthogonal magnetic resonance signal components. 5. The radio frequency component of claim 4, wherein the first coil and the second coil are configured to detect magnetic resonance signals emitted from within a field of view of a B0 field oriented along a third axis. 6. The radio frequency component of claim 5, wherein the first axis and the second axis are substantially orthogonal to the third axis. 7. The radio frequency component of claim 5, wherein the first axis and the second axis are oriented at approximately 45° relative to the third axis. 8. The radio frequency component of claim 2, wherein the radio frequency component comprises a support structure for the first coil and the second coil, the support structure comprising: a first substrate layer to which the first coil is applied; anda second substrate layer to which the second coil is applied. 9. The radio frequency component of claim 8, wherein the support structure defines a three dimensional surface and wherein the second substrate layer is offset from the first substrate layer substantially along surface normal directions to the three dimensional surface. 10. The radio frequency component of claim 8, wherein the support structure defines a three dimensional surface about a region of interest, and wherein the second substrate layer is offset from the first substrate layer relative to the region of interest. 11. The radio frequency component of claim 8, whe... | Hyperfine Research, Inc. | A61B005/00 A61B005/055 G01R033/34 G01R033/36 G01R033/381 G01R033/385
US10702169 | B2 | 2020-07-07 | Method and electronic device for cuff-less blood pressure (BP) measurement | 2017-02-14 | A method and an electronic device for measuring blood pressure are provided. The method includes illuminating, by a PPG sensor included in the electronic device, skin of a user and measuring a PPG signal based on an illumination absorption by the skin. Further, the method also includes extracting, by the electronic device, a plurality of parameters from the PPG signal, wherein the parameters may comprise PPG features, Heart Rate Variability (HRV) features, Acceleration Plethysmograph (APG) features, and non-linear features. The method also includes estimating, by the electronic device, the BP based on the extracted plurality of parameters. | 1. An electronic device for measuring Blood Pressure (BP), the electronic device comprising: a photoplethysmogram (PPG) sensor configured to illuminate skin of a user and measure a PPG signal based on illumination absorption by the skin; anda processor configured to: extract a plurality of parameters from the PPG signal, wherein the extracted plurality of parameters are derived from features of the PPG signal,estimate a diastolic blood pressure (DBP) based on the extracted plurality of parameters, andestimate a systolic blood pressure (SBP) based on the estimated DBP and the extracted plurality of parameters,wherein the processor is further configured to: estimate the DBP by using an artificial neural network (ANN) to output the DBP, andestimate the SBP, based on the DBP output from the ANN that is fed back into the ANN. 2. The electronic device of claim 1, wherein the plurality of parameters that are extracted comprise at least one of magnitude data and temporal data in the PPG signal. 3. The electronic device of claim 2, wherein the magnitude data of the PPG signal comprises at least one of a mean value of a window of the PPG signal, a variance value of the PPG signal, a skewness value of the window, and a kurtosis value of the window, and the temporal data of the PPG signal comprises at least one of a length of the window, a difference of locations of a peak and a first valley with respect to the length of the window, a difference of a peak and a first valley with respect to time, and a difference of a peak and a second valley with respect to time. 4. The electronic device of claim 1, wherein the processor is further configured to extract Acceleration Plethysmograph (APG) features from an APG signal, which is a second derivative of the PPG signal, wherein at least one of the APG features is related to age or arterial information. 5. The electronic device of claim 1, wherein the processor is further configured to estimate at least one from among the DBP and the SBP, based on the extracted plurality of parameters related to the PPG signal and non-linear features, and wherein the non-linear features are extracted based on cardiovascular circulatory system analysis, statistical significance, and correlation with BP values. 6. The electronic device of claim 1, wherein the processor is further configured to estimate the DBP based on a combinatorial artificial neural network (ANN) based feedback model. 7. The electronic device of claim 1, wherein the electronic device further comprises a recommendation engine implemented by the processor, the recommendation engine being configured to: receive the estimated DBP and the estimated SBP of the user based on a periodic interval;determine whether a stressful situation has occurred, based on at least one from among the estimated DBP and the estimated SBP; andprovide at least one recommendation to the user to manage the stressful situation. 8. The electronic device of claim 7, wherein the at least one ... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/021 A61B005/00 A61B005/026 A61B005/0295 A61B005/024
US10722156 | B2 | 2020-07-28 | Sleeve for a handheld oximetry probe | 2017-04-20 | A sleeve includes a body having a top opening. The body covers a handheld oximeter probe or a portion of the probe. The sleeve has a shape that approximately matches the oximeter probe or portion of the probe, which is covered by the sleeve. The sleeve has a top opening that allows a user to slide the oximeter probe into the sleeve. The sleeve is transparent to radiation emitted and collected by the oximeter probe. The sleeve is formed of a material that prevents patient tissue, fluid, viruses, bacteria, and fungus from contacting the covered portions of the oximeter probe. The sleeve leaves the probe relatively sterile after use so that little or no clearing of the probe is required for a subsequent use, such as when the probe is covered with a new, unused sleeve. | 1. A probe cover for an oximeter device comprising: a first portion of the probe cover, wherein the first portion comprises a first open end and a first closed end, opposite to the first open end, and the first closed end comprises a display panel;a second portion of the probe cover, wherein the second portion comprises a second open end and a second closed end, opposite to the second open end, the second closed end comprises an optical sensor panel, and coupling of the first open end to the second open end forms a sealed probe cover enclosure for the oximeter device,the optical sensor panel is transparent and comprises a thickness of less than about 250 microns,when in the sealed probe cover enclosure, a display of the oximeter device is visible through the display panel of the probe cover, lighted emitted by a probe tip of the oximeter device is transmitted through the optical sensor panel of the probe cover, and light received at the oximeter device is transmitted through the optical sensor panel of the probe cover,the sealed probe cover enclosure prevents contaminants from outside of the enclosure from contacting the oximeter device contained within an interior of the enclosure, andthe second portion of the probe cover comprises a barrier at the second closed end, the barrier is coupled to the optical sensor panel, and the barrier prevents contaminants on a tissue being measured from contacting the oximeter device contained within the interior of the enclosure. 2. The probe cover of claim 1 wherein the optical sensor panel passes light in a wavelength range from about 650 nanometers to about 900 nanometers without attenuation that would affect an oximeter measurement. 3. The probe cover of claim 1 wherein the thickness is less than about 150 microns. 4. The probe cover of claim 1 wherein the thickness is less than about 100 microns. 5. The probe cover of claim 1 wherein the thickness is less than about 50 microns. 6. The probe cover of claim 1 wherein the thickness of the optical sensor panel comprises a uniform thickness. 7. The probe cover of claim 1 wherein a first surface of the optical sensor panel conforms to a surface of a sensor of the probe tip of the oximeter device, and the first surface is configured to be flush against the sensor of the probe tip without any gaps. 8. The probe cover of claim 1 wherein the probe cover comprises polycarbonate. 9. The probe cover of claim 1 wherein the probe cover comprises a rigid polymer. 10. The probe cover of claim 1 wherein the optical sensor panel comprises polycarbonate. 11. The probe cover of claim 1 wherein the second portion of the probe cover comprises rigid polymer and the optical sensor panel comprise a flexible polymer, in relation to the rigid polymer. 12. The probe cover of claim 1 wherein the second portion of the probe cover comprises flexible polymer and the optical sensor panel comprise a rigid polymer, in relation to the flexible polymer. 13. The probe cover of claim... | ViOptix, Inc. | A61B005/1455 A61B005/107 A61B005/00
US10709352 | B2 | 2020-07-14 | Method of using lung airway carina locations to improve ENB registration | 2016-10-18 | Disclosed are systems, devices, and methods for registering a luminal network to a 3D model of the luminal network. An example method comprises generating a 3D model of a luminal network, identifying a target within the 3D model, determining locations of a plurality of carinas in the luminal network proximate the target, displaying guidance for navigating a location sensor within the luminal network, tracking the location of the location sensor, comparing the tracked locations of the location sensor and the portions of the 3D model representative of open space, displaying guidance for navigating the location sensor a predetermined distance into each lumen originating at the plurality of carinas proximate the target, tracking the location of the location sensor while the location sensor is navigated into each lumen, and updating the registration of the 3D model with the luminal network based on the tracked locations of the location sensor. | 1. A method of registering a luminal network to a 3D model of the luminal network, the method comprising: generating a 3D model of a luminal network based on images of the luminal network;identifying a target within the 3D model of the luminal network;determining locations of a plurality of carinas in the luminal network proximate the target;displaying guidance for navigating a location sensor within the luminal network;tracking a location of the location sensor while the location sensor is navigated within the luminal network;registering the 3D model of the luminal network to the luminal network based on a tracked location of the location sensor within the luminal network;displaying guidance for navigating the location sensor into each of a plurality of lumens originating at each of the plurality of carinas proximate the target;tracking the location of the location sensor while the location sensor is navigated into each of the plurality of lumens;locally registering the 3D model of the luminal network to the luminal network proximate the target based on the tracked location of the location sensor within each of the plurality of lumens; andaugmenting a registration of the 3D model of the luminal network to the luminal network with a local registration. 2. The method of claim 1, wherein the luminal network is an airway of a patient. 3. The method of claim 2, wherein the 3D model is a model of the airway of the patient. 4. The method of claim 1, wherein the carinas are used as fiducial markers for identifying the location of the target. 5. The method according to claim 1, further comprising: using thin plate spline interpolation to non-rigidly register the 3D model of the luminal network to the luminal network based on the tracked location of the location sensor within the luminal network; andaugmenting the non-rigid registration of the 3D model of the luminal network to the luminal network with the local registration. 6. The method according to claim 1, further comprising using thin plate spline interpolation to locally register the 3D model of the luminal network to the luminal network proximate the target. 7. The method according to claim 1, further comprising: determining if the local registration of the 3D model of the luminal network to the luminal network proximate the target is complete; anddisplaying further guidance for navigating the location sensor into each of the plurality of lumens originating at each of the plurality of carinas proximate the target if it is determined that the local registration proximate the target is not complete. 8. The method according to claim 1, further comprising: identifying at least one additional target within the 3D model of the luminal network;determining if the local registration of the 3D model of the luminal network to the luminal network proximate the at least one additional target is complete; anddisplaying further guidance for navigating the location sensor into each of a plurality of lumens... | COVIDIEN LP | A61B005/05 A61B005/06 A61B010/04 A61B034/20 A61B001/267 A61B005/08 A61B006/03 A61B006/00 A61B017/00 A61B090/00 A61B034/10
US10716494 | B2 | 2020-07-21 | Method of providing information according to gait posture and electronic device for same | 2016-05-04 | A method of providing information according to a gait posture and an electronic device for the same are provided. The method includes collecting sensor values detected using a plurality of sensors located at the surrounding of a user's feet, determining a user's gait posture by using the detected sensor values, and outputting at least one of information on the user's gait posture, information on muscle fatigue of the user according to the gait, information on joint fatigue of the user according to the gait, and information on a recommended exercise for the user based on the determined user's gait posture. | 1. A method of providing information according to a gait posture using an electronic device, the method comprising: collecting sensor values detected using a plurality of sensors located around a user's feet;determining a user's gait posture by using the detected sensor values; andoutputting information on the user's gait posture, information on a recommended exercise for the user, information on joint fatigue of the user according to the gait based on the determined user's gait posture, and information on muscle fatigue of the user according to the gait,wherein the outputting of the information on the recommended exercise comprises outputting at least one of information on a recommended exercise for a correction of the user's gait posture, information on a recommended exercise for a recovery of the muscle fatigue of the user according to the gait, and information on a recommended exercise for a recovery of the joint fatigue of the user according to the gait based on the determined user's gait posture,wherein the outputting of the information on the muscle fatigue of the user according to the gait comprises: determining a user's gait period based on the determination of the user's gait posture,calculating a pressure change of the user's feet at a moment of landing according to the determined user's gait period,determining the joint fatigue of the user according to the gait based on the calculated pressure change at the moment of landing, andoutputting information on the joint fatigue of the user according to the gait based on a result of the determination, andwherein the outputting of the information on the joint fatigue of the user comprises calculating a force applied to each joint by applying the calculated pressure change of the user's feet to predetermined dynamic human body modeling and determining the joint fatigue according to the calculated force applied to each joint. 2. The method of claim 1, wherein the plurality of sensors comprise a plurality of pressure sensors located at an insole of the user and an inertia sensor located around the user's feet, andwherein the inertia sensor comprise at least one of an acceleration sensor, a gyro sensor, and a geomagnetic sensor. 3. The method of claim 1, wherein the determining of the user's gait posture comprises determining the user's gait posture according to a detection order of the sensor values detected by the plurality of sensors. 4. The method of claim 1, wherein the determining of the user's gait posture comprises determining the user's gait posture according to a size of the sensor values detected by the plurality of sensors. 5. The method of claim 1, wherein the determining of the user's gait posture comprises: determining whether the user's gait posture corresponds to a state of at least one of walking, running, and being stopped by using the detected sensor values; anddetermining the user's gait posture according to the state. 6. The method of claim 1, wherein the determining of... | Samsung Electronics Co., Ltd. | A61B005/11 A61B005/103 A61B005/00 A43B003/00 G01C022/00 G01C021/16
US10716714 | B2 | 2020-07-21 | Soiled diaper detection system | 2018-03-01 | A soiled diaper detection system having a soiled diaper detection device, one or more computer devices in communication with the diaper detection device configured to generate and process requests to the diaper detection device, and store data received from the soiled diaper detection device. | 1. A soiled diaper detection system for a diaper worn by a diaper wearer, the soiled diaper detection system comprising: a soiled diaper detection device with a vent and with two humidity sensors separated by a printed circuit board layer; wherein the humidity sensors are configured to measure humidity inside and outside of the diaper; one or more computer devices in communication with the soiled diaper detection device configured to generate requests to the soiled diaper detection device and process requests from the soiled diaper detection device, and store data received from the soiled diaper detection device;wherein the printed circuit board layer is configured to minimize radiation exposure to the diaper wearer; andwherein the soiled diaper detection device further comprises a learning-based system that uses humidity detection thresholds detected by the humidity sensors to adapt the soiled diaper detection device to the diaper wearer. 2. The soiled diaper detection system according to claim 1, wherein a first humidity sensor has a corresponding gas sensor configured to measure the rate of change of the difference in concentration of one or more gases. 3. The soiled diaper detection system according to claim 1, wherein each humidity sensor is configured to measure the rate of change of the difference in humidity between the inside and outside of the diaper. 4. The soiled diaper detection system according to claim 2, further comprising one or more additional sensors including at least one load or weight sensor or force sensor to measure the rate of change of the difference in weight of the diaper. 5. The soiled diaper detection system according to claim 1, further configured to send a signal when the rate of change of the difference in humidity between the inside and outside of the diaper exceeds a configurable value. 6. The soiled diaper detection system according to claim 2, further configured to send a signal when the rate of change of the difference in concentration of gases between the inside and outside of the diaper exceeds a configurable value. 7. The soiled diaper detection system according to claim 4, further configured to send a signal when the rate of change of the difference in weight of the diaper exceeds a configurable value within a minute. 8. The soiled diaper detection system according to claim 1, wherein-the learning-based system uses urination patterns to automatically adapt the soiled diaper detection device to the diaper wearer. 9. The soiled diaper detection system according to claim 2, further configured to send gas sensor output readings, humidity data and other data from one or more other sensors wirelessly to a hub. 10. The soiled diaper detection system according to claim 1, wherein the learning-based system automatically adapts the soiled diaper detection device based on the type of the diaper. 11. The soiled diaper detection system according to claim 1, wherein the soiled diaper detection device has circ... | 5TECH LLC | A61F013/42 G08B021/20 A61F013/84 A61F005/48 A61B005/00
US10711930 | B2 | 2020-07-14 | Releasable connection assembly | 2016-10-31 | A releasable connection assembly for connecting tubing sections is disclosed. The assembly includes a male bayonet connector including a shaft having a distal end and a proximal end, and a female latch connector. The male bayonet connector includes a shaft having a distal end and a proximal end, and defining a lumen extending between the distal end and the proximal end. A rib extends radially from the distal end of the shaft. The female latch connector defines an exterior opening configured to receive at least part of the distal end and the rib of the shaft. A method of using a releasable connection assembly for connecting tubing sections is also disclosed. | 1. A male bayonet connector for connecting sections of tubing, the male bayonet connector comprising: a first shaft including a first distal end portion comprising: a first sealing surface configured to engage a female receiving member to create a fluid-tight seal;a first annular channel proximal to the first sealing surface;a first proximal portion proximal to the first annular channel;at least one rib extending at least partially along a longitudinal axis of the first proximal portion and extending radially from the first proximal portion;a first lumen being defined between the first sealing surface and the first proximal portion; anda grip extending radially from the first shaft to define a flange around the first shaft, the flange having an outer dimension along a direction that is angularly offset from the longitudinal axis and that is greater than outer dimensions of both the first sealing surface and the first proximal portion along the direction. 2. The male bayonet connector of claim 1, wherein the first annular channel is directly adjacent the first sealing surface. 3. The male bayonet connector of claim 1, wherein the first shaft further includes a first proximal end portion configured to engage a first section of tubing. 4. The male bayonet connector of claim 1, wherein the at least one rib comprises a first rib and an opposing, second rib extending the first length along the longitudinal axis of the first proximal portion and extending radially the first height from the first proximal portion. 5. The male bayonet connector of claim 4, wherein the first rib and the second rib radially extend in a direct orthogonal to the longitudinal axis. 6. The male bayonet connector of claim 5, wherein the first rib and the second rib radially extend horizontally. 7. The male bayonet connector of claim 5, wherein the first rib and the second rib radially extend vertically. 8. The male bayonet connector of claim 1, wherein the rib longitudinally extends from the grip. 9. The male bayonet connector of claim 1, wherein the flange defines an outer edge including a plurality of indentations for facilitating gripping of the grip. 10. The male bayonet connector of claim 1, wherein the grip comprises a tie-bar. 11. The male bayonet connector of claim 1, wherein the first annular channel is defined by a first distal sidewall that is perpendicular with respect to the longitudinal axis, a first proximal sidewall having a first beveled edge, and a first bottom region connecting the first distal sidewall and the first beveled edge, with the first bottom region being concentric and recessed radially inward with respect to the first sealing surface. 12. The male bayonet connector of claim 11, wherein the first distal sidewall defines a surface for interfacing with a first latch structure within the female receiving member to prevent removal of the male bayonet connector from the female receiving member when the first latch structure engages the first a... | Nordson Corporation | F16L037/084 A61B005/021 A61M039/10 F16L037/56 A61B005/022
US10716504 | B2 | 2020-07-21 | Medical system having plug and play function | 2017-08-28 | A medical system is disclosed, which can be useable in particular for monitoring and/or controlling at least one bodily function of a user. The medical system comprises a control device and at least one medical user element embodied separately from the control device. The medical user element and the control device are designed to exchange data wirelessly. The medical system is designed to enable an automatic assignment step, wherein an exchange of personal data between the medical user element and the control device is enabled by the automatic assignment step. The medical system is furthermore designed to automatically initiate the automatic assignment step by means of an assignment coupling between the medical user element and the control device. The medical system is furthermore designed to enable a separation of the assignment coupling for medical operation of the medical system after the assignment step. | 1. A medical system for monitoring and/or controlling at least one bodily function of a user, comprising: a blood glucose measuring device;at least one additional device, wherein either the blood glucose measuring device or the additional device or both serve as a control device; andat least one medical user element embodied separately from the control device and configured to perform a medical function,wherein the medical user element and the control device are configured to exchange data wirelessly,wherein the medical user element and the control device are configured to move with respect to each other,wherein the medical system is configured to perform an assignment coupling between the control device and the medical user element using a coupling interface and without physical contact between the control device and the medical user element;wherein the control device and the medical user element exchange personal data automatically via a wireless communication link, wherein the personal data is exchanged only in response to the assignment coupling;wherein a data exchange between the medical user element and the control device that are assigned to one another is effected upon establishment of proximity between the medical user element and the control device so that the data, including the personal data, is transmitted during a time period within which the medical user element moves past the control device. 2. The medical system of claim 1, wherein the control device is configured to be stationary or move at a different speed with respect to the medical user element. 3. The medical system of claim 1, wherein the medical user element is configured to be carried by the user. 4. The medical system of claim 3, wherein the medical user element is configured to be carried directly on or in the body. 5. The medical system of claim 1, wherein the data exchange between the medical user element and the control device is performed at transmission rates of more than 100 kbit/s. 6. The medical system of claim 1, wherein the data exchange between the medical user element and the control device comprises a connection set-up between the medical user element and the control device and a transmission of personal data between the medical user element and the control device. 7. The medical system of claim 1, wherein the personal data are transmitted between the medical user element and the control device by a single or by a repeated exchange of complementary data records. 8. The medical system of claim 1, wherein temporal profiles of measurement values are transmitted between the medical user element and the control device in compartments. 9. The medical system of claim 1, wherein temporal profiles of measurement values are mapped as a function of time over a total observation period. 10. The medical system of claim 1, wherein the medical system is configured in a manner that it firstly communicates a first partial data record wirelessly upon establishment ... | Roche Diabetes Care, Inc. | A61B005/00 A61B005/145 A61N001/372 A61B005/117 G06F021/44 G06F021/32 A61B090/98 G16H040/63 H04W004/80 H04W004/70 G06F019/00 H04L029/08
US10719740 | B2 | 2020-07-21 | Method and a system for identifying reflective surfaces in a scene | 2019-08-13 | Methods and a system for identifying reflective surfaces in a scene are provided herein. The system may include a sensing device configured to capture a scene. The system may further include a storage device configured to store three-dimensional positions of at least some of the objects in the scene. The system may further include a computer processor configured to attempt to obtain a reflective surface representation for one or more candidate surfaces selected from the surfaces in the scene. In a case that the attempted obtaining is successful, computer processor is further configured to determine that the candidate reflective surface is indeed a reflective surface defined by the obtained surface representation. According to some embodiments of the present invention, in a case the attempted calculation is unsuccessful, determining that the recognized portion of the object is an object that is independent of the stored objects. | 1. A method comprising: recognizing that at least a portion of an object in a scene in at least one image is not located at an approximate position associated with the object stored in a database;attempting to obtain a reflective surface representation for one or more candidate reflective surfaces selected from the surfaces in the scene by back ray tracking at least one optical path from the recognized portion of the object in the scene in the at least one image to the stored approximate position associated with the object stored in the database; andin a case obtaining the at least one optical path from the recognized portion of the object in the scene to the stored approximate position associated with the object in the database is successful, determining that at least one candidate reflective surface is a reflective surface defined by the obtained reflective surface representation. 2. The method according to claim 1, wherein the reflective surface representation is achieved by a numerical approximation of a surface equation based on the back ray tracking. 3. The method according to claim 1, wherein the attempting is preceded by identifying, based on the images, candidate reflective surfaces within the scene, and wherein the attempting is carried out on the identified candidates. 4. The method according to claim 3, wherein the identifying of the candidate reflective surfaces is carried out by the recognizing that the at least a portion of the object in the scene in the at least one image is not located at the approximate position associated with the stored object in the database. 5. The method according to claim 4, wherein the identifying of the candidate reflective surfaces is carried out by identifying a 3D depth pattern that is indicative of a reflective surface. 6. The method according to claim 5, wherein the 3D depth pattern that is indicative of a reflective surface includes a well-bordered well. 7. The method according to claim 1, wherein in a case the attempted to obtain is unsuccessful, determining that the recognized portion of the object is an object that is independent of the stored objects. 8. The method according to claim 7, further adding the object as a newly recognized object to the database. 9. The method according to claim 1, wherein in a case that the candidate reflective surface is determined as a reflective surface, deriving reflectance parameters of said reflective surface by applying image processing techniques to the recognized portion of the object and the respective object stored in the database. 10. The method according to claim 9, wherein the reflectance parameters further include identifying portions of the reflective surface which are not reflective. 11. The method according to claim 9, wherein the reflectance parameters comprise level and type of reflectance. 12. The method according to claim 1, wherein in a case that the candidate reflective surface is determined as a reflective surface, rendering a vir... | Alibaba Technology (Israel) Ltd. | A61B003/16 A61B005/055 G06K009/62 G06K009/00 G06K009/46 G06T019/00 G06T007/73 G06K009/52 G06T007/00 G06T007/593
US10704966 | B2 | 2020-07-07 | Apparel having sensor system | 2018-11-19 | A sensor system configured for use with an article of apparel includes one or a plurality of sensors formed of a polymeric material having a conductive particulate material dispersed therein and conductive leads connecting the sensors to a port. The leads may also be formed of a polymeric material having a conductive particulate material dispersed therein. The conductive material is dispersed in the sensor(s) at a first dispersion density and the conductive material is dispersed in the leads at a second dispersion density that is higher than the first dispersion density. Each of the sensors is configured to increase in resistance when deformed under pressure, which is detected by a module connected to the port. The second dispersion density is such that each of the leads has sufficient conductivity that the leads are configured to conduct an electronic signal between each sensor and the port in any state of deformation. | 1. A sensor system comprising: a plurality of sensors configured to be connected to an article of apparel, wherein the sensors are formed of a polymeric material having a conductive particulate material dispersed therein at a first dispersion density, and wherein each of the sensors is configured to increase in resistance when deformed under pressure; anda plurality of conductive leads connected to the sensors and extending away from the sensors, wherein the leads have sufficient conductivity such that each lead is configured to conduct an electronic signal to or from the sensor in any state of deformation,wherein at least one of the sensors comprises a plurality of generally parallel branches having one or more bridges extending between the branches to connect the branches together. 2. The sensor system of claim 1, wherein the at least one of the sensors comprises three or more branches, wherein the branches and the bridges are arranged in an alternating pattern. 3. The sensor system of claim 1, wherein the leads are formed of the polymeric material having the conductive particulate material dispersed therein at a second dispersion density, and wherein the second dispersion density is higher than the first dispersion density, such that each of the leads has sufficient conductivity that the leads are configured to conduct an electronic signal to or from each sensor in any state of deformation. 4. The sensor system of claim 3, wherein each sensor has two leads extending away from the sensor, and wherein each sensor and the two leads connected thereto are integrally formed as a single extruded member having a sensor segment forming the sensor and conductor segments forming the leads, the sensor segment formed of the polymeric material having the conductive particulate material dispersed therein at the first dispersion density, and the conductor segments formed of the polymeric material having the conductive particulate material dispersed therein at the second dispersion density. 5. The sensor system of claim 1, wherein the leads are formed of the polymeric material having a second conductive particulate material dispersed therein, and wherein each of the leads has sufficient conductivity that the leads are configured to conduct an electronic signal to or from each sensor in any state of deformation. 6. The sensor system of claim 1, further comprising an interface connected to the leads and configured for connection to an electronic module comprising at least one of a memory and a processor, such that the electronic module is in communication with the leads and the sensors through the interface. 7. An article of apparel having the sensor system of claim 6 connected thereto, such that the sensors, the leads, and the interface are connected to the article of apparel. 8. A sensor system comprising: a plurality of sensors configured to be connected to an article of apparel, wherein the sensors are formed of a polymeric material having a conductiv... | NIKE, Inc. | A61B005/11 G01L001/20 A61B005/00 A63B024/00 A41D001/00
US10722220 | B2 | 2020-07-28 | Sampling system capsule | 2018-12-05 | A non-invasive system for sampling gastrointestinal microbiota with their associated environment for discovery, characterization, medical research, diagnostics, and treatment using an ingestible, non-digestible sampling capsule. The capsule device, with ports open for acquiring a sample of gastrointestinal microbiota and content, is placed inside an immediate or delayed release capsule. When the outer capsule dissolves according to its specifications, enteric fluid and content enters the sampling capsule through the ports triggering a hydrophilic stop to release a pulling force to close the ports by enveloping the capsule outer casing over the inner casing, capturing the sample, and holding it sealed until it completes passage through the digestive track, and is recovered from feces for analysis. | 1. A sampling system comprising: an inner capsule further comprising, a ported first casing with an open end, a closed end, an inner surface, and an outer surface, a ported second casing with an open end, a closed end, an inner surface, and an outer surface,a hydrophobic, acid resistant elastomeric filament fixedly attached to the closed end of the ported first casing and to the closed end of the ported second casing, anda hydrophilic strut, wherein the hydrophilic strut spaces the closed end of the ported first casing away from the closed end of the ported second casing and is of such a length as to impart a tension in the hydrophobic elastomeric filament; anda first digestible outer capsule enveloping the inner capsule, the first digestible outer capsule being configured to temporarily isolate the inner capsule from bodily fluids. 2. The sampling system of claim 1, wherein the outer surface of the open end of one of the ported first casing and the ported second casing slideably engages the inner surface of the open end of the other ported casing permitting a sliding transition of the inner capsule from an open configuration to a closed configuration. 3. The sampling system of claim 2, wherein in the closed configuration the inner capsule is fluid tight. 4. The sampling system of claim 2 further comprising a slide ring stop circumscribing the inner surface of one of the ported first casing and the ported second casing, wherein the slide ring stop is configured to engage the outer surface of the other ported casing when the inner capsule is in the closed configuration thereby arresting any further sliding transition thereof. 5. The sampling system of claim 4, wherein the hydrophobic elastomeric filament applies a contracting force that holds the outer surface of the other ported casing against the slide ring when the inner casing is in its closed configuration. 6. The sampling system of claim 2, wherein the hydrophobic elastomeric filament provides a contracting force between the ported first casing and the ported second casing when the inner capsule is in either of the closed configuration or in the open configuration. 7. The sampling system of claim 1, wherein the hydrophilic strut is one of an immediate action strut and a delayed action strut. 8. The Sampling system of claim 1 further comprising a second digestible outer capsule that envelops the inner capsule and is enveloped by the first digestible capsule, wherein the second digestible capsule is formulated to fail based on different environmental factors than the first digestible capsule. 9. A gastrointestinal sampling device comprising: a hollow capsule having a ported inner casing with a closed end and an open end and having a ported outer casing with a closed end and an open end;an elastomeric filament fixedly securing the closed end of the inner casing to the closed end of the outer casing; anda hydrophilic strut spacing the closed end of the inner casing away from the closed e... | - | A61B005/00 A61B010/00
US10716491 | B2 | 2020-07-21 | Apparatus and method for the analysis of the change of body composition and hydration status and for dynamic indirect individualized measurement of components of the human energy metabolism | 2018-02-27 | One embodiment of an apparatus for analysis of body composition and hydration status by detecting resistance of the human subject at zero and infinite frequency including a method for measuring indirectly extracellular water mass, intracellular water mass, lean body mass, and body fat mass; daily changes of extracellular water mass, intracellular water mass, lean body mass, and body fat mass; and acute changes of extracellular water mass and intracellular water mass; and for individualized calibration of these indirect measurements.;In addition, a method for fitting mathematical models to serial measurements of indirectly measured lean body mass and fat mass and for dynamic indirect individualized measurement using minimum variance estimation and prediction of daily changes of the body composition defined as change of glycogen store, change of fat store and change of protein store; daily utilized macronutrient energy intake defined as utilized carbohydrate, fat, and protein caloric intake; daily macronutrient oxidation rate defined as rate of carbohydrate oxidation, fat oxidation, and protein oxidation; daily resting metabolic rate; daily unknown forms of energy losses or gains; daily rate of endogenous lipolysis; daily nitrogen excretion; daily gluconeogenesis from protein; daily determination of extracellular water mass; daily determination of intracellular water mass; and acute change of extracellular water mass and intracellular water mass. | 1. A machine for detecting a resistance of a user at an extrapolated zero and infinite frequency comprising: a measuring circuit comprising a first current source, a first measuring point, a first reference resistance, a second measuring point, a first current excitation electrode, an impedance of the user, a second current excitation electrode, a third measuring point, a second reference resistance, a fourth measuring point, a second current source, and an earth potential;a first voltage sensing electrode operably connected to the impedance of the user and a fifth measuring point, and a second voltage sensing electrode operably connected to the impedance of the user and a sixth measuring point;a sensing circuit comprising the earth potential connected to a computing device comprising a microcontroller unit and the first, the second, the third, the fourth, the fifth, and the sixth measuring points each operably connected to the microcontroller unit;an input logic circuit comprising a power switch and the microcontroller unit for operating the power switch, wherein the microcontroller unit comprises hardware configured to at least one of isolates or short circuits a predetermined portion of the measuring circuit or leaves the measuring circuit intact;a digital signal processor unit operably connected to the microcontroller unit, wherein the digital signal processor unit comprises hardware configured to: run a sine wave fitting algorithm to determine an offset, an amplitude, and a phase of signals measured at the first, the second, the third, the fourth, the fifth, and the sixth measuring points; andperform a curve fitting of a Cole model via extrapolation to determine the resistance of the user at the extrapolated zero and infinite frequency;a first in first out memory for storing a train of digital sine waves comprising a plurality of full periods coming from the microcontroller unit and releasing the train of digital sine waves upon a software trigger to a two-channel-digital-analog converter operably connected to a first constant current source and a second constant current source;a plurality of analog-digital converters configured to measure voltage signals during a first instance between the earth potential and at least two of the first, the second, the third, the fourth, the fifth, and the sixth measuring points; anda programmable gain instrumentation amplifier to amplify and filter voltage signals during a second instance at the first, the second, the third, the fourth, the fifth, and the sixth measuring points. 2. A method for detecting a resistance of a user at an extrapolated zero and infinite frequency, comprising: providing a measuring circuit wherein current flows from a first current source through a first measuring point, a first reference resistance, a second measuring point, a first current excitation electrode, an impedance of the user, a second current excitation electrode, a third measuring point, a second reference resistance... | - | A61B005/053 A61B005/00 G16B040/00
US10722177 | B2 | 2020-07-28 | Cardiovascular monitoring using combined measurements | 2015-12-18 | A system for collecting data for assessment of cardiovascular function includes a plurality of monitoring devices coupled to different respective body parts. Each monitoring device is configured to measure a respective signal at the respective body part in response to cardiovascular activity. The respective signal includes a cardiovascular component attributable to the cardiovascular activity and an artifact component not attributable to the cardiovascular activity. When the monitoring devices measure the respective signals simultaneously over a same time period, the cardiovascular components are correlated, and the artifact components are not correlated. The system also includes a controller configured to: identify the cardiovascular components included in the signal measurements, according to the correlation of the cardiovascular components; reject the artifact components included in the signal measurements, according to the non-correlation of the artifact components; and determine cardiovascular information from the identified cardiovascular components for an assessment of cardiovascular function. | 1. A system for collecting data for assessment of cardiovascular function, comprising: a plurality of cardiovascular monitoring devices configured to be coupled to body parts of a subject and to determine cardiovascular signal measurements at the body parts in response to cardiovascular activity, each monitoring device including a respective first communication interface and configured to determine a respective one of the cardiovascular signal measurements at a respective one of the body parts in response to the cardiovascular activity, the cardiovascular signal measurements including cardiovascular components attributable to the cardiovascular activity and artifact components not attributable to the cardiovascular activity, the artifact components being attributable at least to movement of at least one of the body parts,wherein the plurality of cardiovascular monitoring devices includes at least two of: (i) a first sensor configured to measure electrical signals in response to the cardiovascular activity, (ii) a second sensor configured to measure blood volume in response to the cardiovascular activity, or (iii) a third sensor configured to measure blood flow in response to the cardiovascular activity, andwherein, in response to the plurality of cardiovascular monitoring devices determining the cardiovascular signal measurements at the body parts simultaneously over a same time period, the cardiovascular components included in the cardiovascular signal measurements by the cardiovascular monitoring devices are correlated and the artifact components included in the cardiovascular signal measurements by the cardiovascular monitoring devices are not correlated; anda controller including: a second communication interface configured to be communicatively coupled to the first communication interface of each monitoring device;at least one data storage device; andat least one processor configured to: (i) receive the cardiovascular signal measurements from the cardiovascular monitoring devices via the second communication interface; (ii) store the cardiovascular signal measurements on the at least one data storage device; (iii) determine a correlation of the cardiovascular components and a non-correlation of the artifact components by statistically comparing the cardiovascular signal measurements from the plurality of cardiovascular monitoring devices; (iv) identify the cardiovascular components included in the cardiovascular signal measurements according to the correlation of the cardiovascular components; (v) reject the artifact components included in the cardiovascular signal measurements according to the non-correlation of the artifact components; and (vi) determine cardiovascular information from the cardiovascular components separated from the artifact components for an assessment of cardiovascular function,wherein the at least one processor rejects the artifact components attributable to the movement of at least one of the body parts based on the c... | Verily Life Sciences LLC | A61B008/00 A61B005/00 A61B005/0402 A61B005/04 A61B005/0295 A61B008/08 A61B005/026 A61B005/11 A61B005/024 A61B008/06
US10719796 | B2 | 2020-07-21 | Methods and systems for detecting an abnormal stress condition in a subject or group of subjects | 2016-07-19 | Disclosed are methods and systems for detecting an abnormal stress condition in a subject or group of subjects. Included is a risk evaluation system monitoring people, wherein a control unit in a stress management system obtains mental health information and records same in a measurement information database. The control unit then specifies a pattern in a reference pattern database and obtains the measurement information relating to this pattern from the measurement information database. The control unit in the stress management system (20) then revises a standard pattern based on individual attributes or work attributes. Next, the control unit performs pattern matching and determines whether or not attention is required. If the determination is that attention is required, the control unit presents an alert and an advice. This enables work, organization, and member risks to be appropriately evaluated based on changes in mental health information. | 1. A method for detecting an abnormal condition of mental stress in a group of subjects engaged in a business operation carried out collectively by the subjects of the group, the method comprising: obtaining and storing constituent information of the business operation carried out by the group and each of the subjects of the group engaged in the business operation, said constituent information including progress status of the business operation from a first time period to a second time period;measuring with a measurement device that measures at least one parameter selected from the group consisting of a physiological parameter, an emotional parameter and a behavioral reaction parameter to determine stress information of each of the subjects;determining transition of mental information for each of the subjects in the group from the first period to the second period from the determined stress information of each of the subjects;calculating a transition of a mental statistic of the group using the transition of mental information measured for each of the subjects of the group;comparing the transition of mental information of each of the subjects to a reference pattern for personnel stored in a reference pattern database, wherein the reference pattern for personnel includes a standard pattern for personnel and an attention pattern for personnel, the standard pattern for personnel including a standard pattern waveform regarding a standard stress waveform in the business operation, the attention pattern for personnel including an attention pattern waveform regarding a waveform of a measured value of mental information to which attention should be paid;obtaining, when the variance of the transition of mental information of at least one subject relative to the standard pattern for personnel is greater than or equal to a reference value or when the transition of mental information of at least one subject includes an evaluation index that matches with the attention pattern for personnel, a contact of the at least one object from a contact management database;outputting automatically an alert including an advice to the contact of the at least one object;comparing the transition of the mental statistic of the group to a reference pattern for group stored in the reference pattern database, wherein the reference pattern for group includes a standard pattern for group and an attention pattern for group, the standard pattern for group including a standard pattern waveform regarding a standard stress waveform in the business operation, the attention pattern for group including an attention pattern waveform regarding a waveform of a measured value of mental information to which attention should be paid;obtaining, when the variance of the transition of the mental statistic of the group relative to the standard pattern for group is greater than or equal to a reference value or when the transition of the mental statistic of the group includes an evaluation index that... | IMATEC INC. | G06Q010/06 A61B005/16 G06F019/00
US10719931 | B2 | 2020-07-21 | Methods and systems for assessing image quality in modeling of patient anatomic or blood flow characteristics | 2017-11-03 | Systems and methods are disclosed for assessing the quality of medical images of at least a portion of a patient's anatomy, using a computer system. One method includes receiving one or more images of at least a portion of the patient's anatomy; determining, using a processor of the computer system, one or more image properties of the received images; performing, using a processor of the computer system, anatomic localization or modeling of at least a portion of the patient's anatomy based on the received images; obtaining an identification of one or more image characteristics associated with an anatomic feature of the patient's anatomy based on the anatomic localization or modeling; and calculating, using a processor of the computer system, an image quality score based on the one or more image properties and the one or more image characteristics. | 1. A method for automatically selecting image(s) to produce model(s) for a patient-specific simulation of blood flow, using a computer system, the method comprising: defining a performance parameter for a simulation of a patient's blood flow;receiving one or more images of at least a portion of a patient's anatomy;determining an image quality score associated with the one or more images;determining a value of the performance parameter of the simulation, based on the determined image quality score; andgenerating a prompt to correct, retake, or update an image of the one or more images, based on the determined value of the performance parameters. 2. The method of claim 1, wherein the performance parameter includes a metric measuring time to complete at least a portion of the simulation, cost of at least a portion of the simulation, error of the simulation, percent confidence or confidence interval of the simulation, or a combination thereof. 3. The method of claim 1, further comprising: determining a threshold performance parameter value;if the determined value exceeds the threshold performance parameter value, prompting a simulation, andif the determined value falls below the threshold performance parameter value, prompting a display to refrain from the simulation. 4. The method of claim 1, further comprising: generating a display for a user to review or inspect the one or more images. 5. The method of claim 1, further comprising: determining a simulation technique associated with the determined value of the performance parameters; andguiding performance of a simulation based on the determined simulation technique. 6. The method of claim 1, further comprising: performing a global quality assessment of the one or more images collectively and performing a local quality assessment of a portion of one image of the one or more images, wherein the image quality score is based on the global quality assessment and the local quality assessment. 7. The method of claim 1, further comprising: generating an anatomic model of at least a portion of the patient's anatomy based on the one or more images; andperforming the simulation based on the generated anatomic model. 8. A system for automatically selecting image(s) to produce model(s) for a patient-specific simulation of blood flow, the system comprising: a digital storage device storing instructions for automatically selecting image(s) to produce model(s) for a patient-specific simulation of blood flow; anda processor configured to execute the instructions to perform a method including:defining a performance parameter for a simulation of a patient's blood flow;receiving one or more images of at least a portion of a patient's anatomy;determining an image quality score associated with the one or more images;determining a value of the performance parameter of the simulation, based on the determined image quality score; andgenerating a prompt to correct, retake, or update an image of the one or more images... | HeartFlow, Inc. | G16H050/50 G06T007/00 A61B006/00 A61B005/026 A61B006/03 A61B008/08 G01R033/563
US10720076 | B1 | 2020-07-21 | Closed-loop model-based controller for accelerating memory and skill acquisition | 2017-08-21 | Described is a closed-loop control system for memory consolidation in a subject. During operation, the system encodes information regarding environmental items as memories in both a long-term memory store and a short-term memory store. The system generates an activation level representation of a memory of interest related to at least one of the environmental items. An association strength representation for the memories is also generated. Memory consolidation is simulated when the subject is in NREM sleep or quiet waking by strengthening the association strength representation related to the memory of interest. The system predicts behavioral performance for the memory of interest as a probability that the memory of interest can be recalled on cue. When the behavioral performance is below a threshold, an intervention system can be activated. | 1. A closed-loop control system for memory consolidation in a subject, the system comprising: one or more processors and a memory, the memory being a non-transitory computer-readable medium having executable instructions encoded thereon, such that upon execution of the instructions, the one or more processors perform operations of: encoding information regarding environmental items as memories in both a long-term memory store and a short-term memory store:generating an activation level representation of the first memory, the first memory being a memory of interest related to at least one of the environmental items:generating an association strength representation for the memories:encoding the first memory in a subject while the subject is coupled to an intervention system and exposed to an environmental item by associating a semantic item description of the environmental item with physiological measurements of the subject;predicting behavioral performance for the first memory, the behavioral performance being a probability that the first memory can be recalled on cue; andcontrolling operation of the intervention system with respect to the first memory based on the behavioral performance of the first memory determined by the simulation, wherein controlling operation of the intervention system includes activating at least one intervention selected from a group consisting of transcranial current stimulation, an audio cue, or an odor cue. 2. The closed-loop control system as set forth in claim 1, wherein the encoding information regarding environmental items includes encoding descriptive attribute-value pairs and encoding parameters. 3. The closed-loop control system as set forth in claim 2, wherein simulating memory changes of a first memory in a subject while the subject is coupled to the intervention system is performed when the subject is in non-rapid eye movement (NREM) sleep or waking by strengthening the association strength representation related to the first memory. 4. The closed-loop control system as set forth in claim 3, where information is encoded in both the long-term memory and short-term memory at the same time. 5. The closed-loop control system as set forth in claim 4, wherein an output of the short-term memory is the input of long-term memory, such that when an item is recalled from the short-term memory, the association strength representation for the item are strengthened in the long-term memory. 6. The closed-loop control system as set forth in claim 5, wherein the encoding parameters include both timestamp data and physiological measurements of a subject. 7. The closed-loop control system as set forth in claim 6, wherein the memory of interest is a sequence of events or a single event. 8. The closed-loop control system as set forth in claim 7, wherein the activation level representation is a function of physiological measurements taken at time of encoding the memory of interest. 9. The closed-loop control system as set ... | HRL Laboratories, LLC | G09B019/00 A61B005/00 A61B005/0476
US10722695 | B2 | 2020-07-28 | Dilation system and method | 2018-08-01 | A dilation system and method for accessing a surgical target site to perform surgical procedures. In one version, the dilation system includes a dilation assembly and an actuating mechanism. The dilation assembly comprises plurality of dilator segments. The actuating mechanism is operably associated with the dilator segments so as to cause the dilator segments to move from a collapsed state to an expanded state. | 1. A dilation system, comprising: a plurality of dilator segments, each dilator segment having a length, a proximal end, a distal end, an inner side, and an outer side, the dilator segments forming a tubular dilator having a longitudinal axis, at least one of the dilator segments including an electrode assembly; andan actuating mechanism operably associated with the dilator segments to move the dilator segments radially from a collapsed state to an expanded state,wherein the actuating mechanism has a rod positioned in the tubular dilator, the rod having a longitudinal bore extending therethrough, andwherein the dilator segments are rotatable relative to at least a portion of the actuating mechanism about the longitudinal axis. 2. The dilation system of claim 1, wherein dilator segments are rotatable relative to the rod. 3. The dilation system of claim 1, wherein the actuating mechanism includes an actuating handle, and wherein the dilator segments are rotatable relative to the actuating handle. 4. The dilation system of claim 3, wherein the actuating handle comprises: a first handle; anda second handle pivotally connected to the first handle and configured in such a way that movement of the first handle and the second handle relative to one another causes radial movement of the dilator segments. 5. The dilation system of claim 4, wherein the actuating handle includes an incremental indicator for identifying the diameter of the dilator assembly. 6. The dilation system of claim 4, wherein the actuating handle includes a ratchet having at least two teeth forming a groove, and a finger capable of being slidably received within the groove. 7. The dilation system of claim 1, wherein the dilator segments are biased in the collapsed state. 8. A dilation system, comprising: a plurality of dilator segments, each dilator segment having a length, a proximal end, a distal end, an inner side, and an outer side, the dilator segments forming a tubular dilator having a longitudinal axis, at least one of the dilator segments having a channel formed in the outer surface thereof and extending the length of the dilator segment for receiving a probe; andan actuating mechanism operably associated with the dilator segments to move the dilator segments radially from a collapsed state to an expanded state,wherein the actuating mechanism has a rod positioned in the tubular dilator, the rod having a longitudinal bore extending therethrough, andwherein the dilator segments are rotatable relative to at least a portion of the actuating mechanism about the longitudinal axis. 9. The dilation system of claim 8, wherein dilator segments are rotatable relative to the rod. 10. The dilation system of claim 8, wherein the actuating mechanism includes an actuating handle, and wherein the dilator segments are rotatable relative to the actuating handle. 11. The dilation system of claim 10, wherein the actuating handle comprises: a first handle; anda second handle pivotally conn... | DEPUY SYNTHES PRODUCTS, INC. | A61M029/00 A61B005/04 A61B017/02
US10713920 | B2 | 2020-07-14 | Wearable computing device for monitoring hazards | 2018-11-27 | Embodiments of the present invention relate to wearable computing devices. In some embodiments, a wearable computing devices (“WCD”) for monitoring occupational hazards are disclosed. The WCD may include an apparel item having one or more control circuits affixed thereto. One or more sensors may be communicatively coupled to the computing device and configured to detect or monitor at least one of an aspect of a user of the WCD and an ambient environment of the user. The control circuits can be configured to generate one or more notifications when sensor data comprises a value above a threshold amount. | 1. A wearable computing device for monitoring occupational hazards comprising: an apparel item;a control circuit affixed to the apparel item;a communications device communicatively coupled to the control circuit;an antenna communicatively coupled to the communications device and comprising a component, the component comprising a first conductive composition, the first conductive composition comprising individual sheets of graphene forming a three-dimensional percolated network within the conductive composition;a sensor communicatively coupled to the control circuit and configured to detect or monitor at least one of an aspect of a user of the wearable computing device and an ambient environment of the user; andwherein the antenna is one or more of removable from and/or dynamically positioned on the apparel item;the communications device is configured for two-way communication;the data comprises at least one of voice data, text data, and graphical data;the sensor comprises a second conductive composition;the second conductive composition comprises individual sheets of graphene;the individual sheets of graphene forming a three-dimensional percolated network within the second conductive composition; andthe control circuit is configured to generate a notification when sensor data comprises a value above a threshold amount. 2. The wearable computing device of claim 1, further comprising an input/output device communicatively coupled to the control circuit; andwherein in generating the notification the control circuit is further configured to emit a signal via the I/O device when the notification is generated. 3. The wearable computing device of claim 1, wherein the sensor is configured to detect a chemical. 4. The wearable computing device of claim 3, wherein the chemical comprises one or more of propane, methane, carbon monoxide, and a petrochemical. 5. The wearable computing device of claim 1, wherein the control circuit is further configured to modulate a data transmission of the communications device. 6. The wearable computing device of claim 1, wherein the sensor comprises an accelerometer. 7. The wearable computing device of claim 1, wherein the communications device is configured to one or more of transmits and receives data, and wherein the data comprises one or more of voice data, text data, and graphical data. 8. The wearable computing device of claim 1, wherein the sensor is configured to capture a heartrate value of the user. 9. The wearable computing device of claim 1, wherein the sensor is configured to capture geolocation data of the wearable computing device. 10. The wearable computing device of claim 9, wherein the geolocation data comprises one or more of GPS data, time difference of arrival data, Wi-Fi positioning data, and latitude/longitude data. 11. The wearable computing device of claim 1, wherein the apparel item is a vest. 12. The wearable computing device of claim 1, wherein the apparel item is a form fitting appar... | VORBECK MATERIALS, CORP. | G08B021/00 G08B021/18 A41D001/00 A61B005/00 G06F011/30 A61B005/0205 G06F001/16 G08B021/02 A61B005/024
US10702200 | B2 | 2020-07-07 | Method for monitoring incontinence | 2018-12-12 | According to one aspect of the inventive concept there is provided a method of monitoring incontinence for a user, comprising: determining, using a urine sensitive circuit provided at an absorbent article and arranged to present a changed electrical characteristic when exposed to urine, whether the user has urinated on the absorbent article, performing, using a measurement circuit, a measurement on the urine bladder to determine a parameter which varies with a fill level of the urine bladder, and in response to determining that the user has urinated on the absorbent article, recording, by a processing circuit, data representing said parameter. | 1. A method of monitoring incontinence for a user, comprising: determining, using a urine sensitive circuit provided at an absorbent article and arranged to present a changed electrical characteristic when exposed to urine, whether the user has urinated on the absorbent article, wherein the changed electrical characteristic of the urine sensitive circuit is indicative of an amount of urine released on to the urine sensitive circuit,performing, using a measurement circuit, a measurement on the urine bladder to determine a parameter which varies with a fill level of the urine bladder,determining an orientation and/or a movement of the user using a sensor arranged on the user, andin response to determining that the user has urinated on the absorbent article, recording, by a processing circuit, data representing said parameter, an estimated movement and/or posture of the user based on the orientation and/or the movement determined by the sensor, and an estimate of the amount of urine released on to the urine sensitive circuit. 2. The method according to claim 1, further comprising, in response to determining that the user has urinated on the absorbent article, recording, by the processing circuit, said data and an indication of that the user has urinated on the absorbent article. 3. The method according to claim 1, further comprising: performing, a plurality of times, using the measurement circuit, a measurement on the urine bladder to determine a parameter which varies with a fill level of the urine bladder, andrecording, by the processing circuit, an association between each determined parameter and an indication of whether the user has urinated on the absorbent article. 4. The method according to claim 1, wherein said amount of urine is estimated based on a size of the changed electrical characteristics of the urine sensitive circuit. 5. The method according to claim 1, further comprising: performing, a plurality of times, using the measurement circuit, a measurement on the urine bladder to determine a parameter which varies with a fill level of the urine bladder, andcomparing, by the processing circuit, each determined parameter to one or more parameter thresholds, each threshold being associated with a movement and/or posture. 6. The method according to claim 1, further comprising: estimating, by the processing circuit, a urine bladder fill level based on the determined parameter, andproviding a signal which is indicative of the estimated urine bladder fill level. 7. The method according to claim 1, further comprising, subsequent to detecting that the urine sensitive circuit has been exposed to urine, performing by the processing circuit a comparison between: the at least one parameter represented by the data recorded in response to the detection of the urination event, and a parameter determined by the measurement circuit at an instant subsequent to the detection of the urination event. 8. The method according to claim 1, further compris... | STICHTING IMEC NEDERLAND | A61B005/20 A61B005/1477 A61B005/145 A61F013/42 A61B005/00 A61B005/11 A61B005/01 A61F013/15 A61F013/84
USRE048076 | E1 | 2020-07-07 | Venous pressure measurement apparatus | 2019-04-29 | A non-invasive venous pressure measurement apparatus is provided, including: a first cuff attached to a portion including a vein and an artery in a living body; a pressure control unit that changes a first applied pressure applied by the cuff to the portion; a pulse wave detection unit that detects a pulse wave from a pressure received by the cuff from the portion; another pulse wave detection unit that detects another pulse wave including at least an arterial pulse wave in another portion of the living body; an analyzing unit that analyzes a correlation between the two pulse waves, which are changed as the applied pressure is changed by the pressure control unit changes; and a venous pressure calculation unit that calculates a venous pressure based on the applied pressure and a result of analysis by the analyzing unit. | 1. A venous pressure measurement apparatus comprising: a first cuff configured to be attached to a first portion including a vein and an artery in a living body;a pressure control unit that changes a first applied pressure applied by said first cuff to said first portion by sending a flow of air to the first cuff;a first pulse wave sensor that detects a first pulse wave from the pressure received by said first cuff from said first portion;a second pulse wave sensor configured to detect a second pulse wave including at least an arterial pulse wave in a second portion different from said first portion offrom said living body;a pressure waveform analyzing processor, comprising an analyzing unit that analyzes differences between said second pulse wave detected by said second pulse wave sensor and said first pulse wave detected by said first pulse wave sensor as a correlationa difference between the first pulse wave and the second pulse wave, which are changed as said first applied pressure is changed by said pressure control unit; anda venous pressure calculation unit that calculates a venous pressure based on said first applied pressure and a result of analysis by said analyzing unit. 2. The venous pressure measurement apparatus as claimed in claim 1, wherein said venous pressure calculation unit calculates said venous pressure based on the first applied pressure when said correlation is minimized as a result of the analysis by said analyzing unit. 3. The venous pressure measurement apparatus as claimed in claim 1, wherein said analyzing unit obtains a venous pressure wave component by subtracting a component correlated with said second pulse wave from said first pulse wave. 4. The venous pressure measurement apparatus as claimed in claim 3, wherein the component correlated with said second pulse wave is an arterial pulse wave component of said first pulse wave calculated from an arterial pulse wave component of said second pulse wave and said first pulse wave. 5. The venous pressure measurement apparatus as claimed in claim 3, wherein said venous pressure calculation unit calculates said venous pressure based on said first applied pressure when an amplitude of said venous pulse wave component is maximized. 6. The venous pressure measurement apparatus as claimed in claim 1, wherein said first cuff is configured to be attached to an upper arm in the vicinity of a neck or an axilla. 7. The venous pressure measurement apparatus as claimed in claim 1, wherein said second pulse wave sensor is a second cuff configured to be attached to said second portion distal to said first cuff with respect to the heart of said living body, andsaid pressure control unit further controls a second applied pressure applied by said second cuff to said second portion. 8. The venous pressure measurement apparatus as claimed in claim 7, wherein said first cuff is configured to be attached to an upper arm in the vicinity of an axilla, and said second cuff is configured ... | NIHON KOHDEN CORPORATION | A61B005/02 A61B005/022 A61B005/021 A61B005/024 A61B005/00
US10705071 | B2 | 2020-07-07 | Medical instrument with shutter for sealing a test strip port | 2017-12-18 | A medical instrument/analyzer for measuring a biological sample using a test strip. The instrument comprises: a housing with an exterior surface, an internal surface, and an interior volume; an analytical unit within the interior volume with a test strip mount; a test strip port between the exterior and the interior surfaces; a shutter configured for sealing the test strip port when closed, wherein the shutter comprises a test strip support aligned with the test strip port and mount when the shutter is in the open position, wherein the shutter has a first sealing surface, the test strip port has a second sealing surface, and the first sealing and second sealing surfaces are configured to mate in the closed position, wherein the shutter comprises a mechanism for moving between the open and the closed position; and an actuator configured for moving the shutter between the open and closed position. | 1. A method of operating a medical instrument, wherein the medical instrument is an analyzer for performing a measurement on a biological sample using a test strip, wherein the medical instrument comprises: a housing with an exterior surface, wherein the housing comprises an internal surface surrounding an interior volume;an analytical unit for analyzing the test strip, wherein the analytical unit is within the interior volume, wherein the analytical unit comprises a test strip mount configured for receiving the test strip to perform the measurement;a test strip port between the exterior surface and the interior surface, wherein the test strip port is configured for receiving the test strip, wherein the test strip port is aligned with the test strip mount along an insertion direction;a shutter for sealing the test strip port, wherein the shutter is configured for being in a open position and a closed position, wherein the shutter is within the interior volume, wherein the shutter is configured for sealing the test strip port when in the closed position, wherein the shutter comprises a test strip support, wherein the test strip support is aligned with the test strip port and the test strip mount when the shutter is in the open position, wherein the shutter has a first sealing surface, wherein the test strip port has a second sealing surface, wherein the first sealing surface and the second sealing surface are configured to mate in the closed position, wherein the shutter comprises a mechanism for moving the shutter between the open position and the closed position, wherein moving the shutter between the closed position and the open position comprises a rotation of the shutter perpendicular to the insertion direction; andan actuator for actuating the mechanism to move the shutter between the open position and the closed position;wherein the method comprises the steps of:controlling the actuator to actuate the mechanism to move the shutter in the open position;placing the biological sample on the test strip;inserting the test strip into the test strip port such that the test strip passes through the test strip support and into the test strip mount;analyzing the test strip with the analytical unit to perform the measurement;removing the test strip from the medical instrument;controlling the actuator to actuate the mechanism to move the shutter in the closed position; andcleaning the exterior surface of the medical instrument. 2. A medical instrument, wherein the medical instrument is an analyzer for performing a measurement on a biological sample using a test strip, wherein the medical instrument comprises: a housing with an exterior surface, wherein the housing comprises an internal surface surrounding a interior volume;an analytical unit for analyzing the test strip, wherein the analytical unit is within the interior volume, wherein the analytical unit comprises a test strip mount configured for receiving the test strip to perform the measurement;... | Roche Diagnostics Operations, Inc. | G01N033/487 A61B005/145 G01N035/00
US10722282 | B2 | 2020-07-28 | Methods and devices to treat nasal airways | 2018-07-20 | A method for reshaping a nasal airway in a patient involves advancing an inflatable balloon of a reshaping device in an uninflated configuration into a nostril of the patient and between a nasal septum and a lateral wall of the nasal airway. The method then involves inflating the inflatable balloon to an inflated configuration to cause the inflatable balloon to contact nasal mucosa covering the nasal septum and the lateral wall, delivering energy from an energy delivery member attached to or inside of the inflatable balloon, and removing the reshaping device from the nasal airway. | 1. A system for treating a nasal airway in a patient, without forming an incision, penetrating tissue with a needle, or leaving behind an implant, the system comprising: a nasal airway treatment device, comprising; a treatment element, comprising; an inflatable balloon configured to be advanced into the nasal airway in an uninflated configuration and inflated to assume an inflated configuration; andmultiple longitudinal, flat, expandable, bipolar radiofrequency electrodes disposed on an outer surface of the inflatable balloon around a full circumference of the inflatable balloon at regular intervals, wherein each of the radiofrequency electrodes comprises a conductive metal configured to bend outward with the inflatable balloon when the inflatable balloon expands;a shaft attached to a proximal end of the treatment element; anda handle attached to a proximal end of the shaft;a controller coupled with the nasal airway treatment device to control delivery of radiofrequency energy to the nasal airway via the multiple radiofrequency electrodes; anda radiofrequency energy generator coupled with the nasal airway treatment device to provide the radiofrequency energy to the multiple radiofrequency electrodes. 2. The system of claim 1, further comprising wiring configured to carry radiofrequency energy from the radiofrequency energy generator to the multiple radiofrequency electrodes. 3. The system of claim 1, further comprising a pharmaceutical substance coated on or contained within the inflatable balloon, wherein the inflatable balloon is configured to release the pharmaceutical substance within the nasal airway. 4. The system of claim 1, wherein the shaft comprises: an outer shaft; andan inner shaft that extends through an interior of the inflatable balloon, wherein a proximal end and a distal end of the inflatable balloon are attached to the inner shaft, and wherein the inner shaft comprises an inflation lumen and at least one inflation aperture for inflating the inflatable balloon. 5. The system of claim 1, wherein the controller is housed in the handle. 6. The system of claim 1, wherein the handle comprises a power button to activate delivery of the radiofrequency energy from the radiofrequency energy generator to the at least one radiofrequency electrode. 7. The system of claim 1, wherein the controller and the radiofrequency generator are housed in one console that is coupled with the nasal airway treatment device, and wherein the nasal airway treatment device further comprises a cable attached to the handle to connect the nasal airway treatment device with the console. | AERIN MEDICAL, INC. | A61B018/14 A61F005/08 A61B017/24 A61B018/00 A61B018/02 A61B018/08 A61B018/06 A61N005/06 A61N005/02 A61F005/56 A61B018/12 A61B018/16 A61B017/28 A61N001/40 A61F007/10 A61B005/01 A61B018/18 A61B018/20 A61B017/32 A61N007/02 A61B018/04
US10702186 | B2 | 2020-07-07 | Method and apparatus for identifying an organ structure of an examined object in magnetic resonance image data | 2016-12-06 | In a method and apparatus for identifying an organ structure of an examined object in magnetic resonance image data, magnetic resonance measurement data for the organ structure of the examined object are acquired by operation of a magnetic resonance scanner using a magnetic resonance sequence that specifies a sampling scheme of k-space. Magnetic resonance image data are reconstructed from the magnetic resonance measurement data. The organ structure is identified in the magnetic resonance image data. The sampling scheme of k-space is selected so as to support the subsequent identification of the organ structure in the magnetic resonance image data reconstructed from the magnetic resonance measurement data. | 1. A method for identifying an organ structure of an examination subject in magnetic resonance image data, comprising: operating a magnetic resonance data acquisition scanner to acquire magnetic resonance measurement data from the organ structure of an examination subject, using a magnetic resonance data acquisition sequence that specifies a sampling scheme for k-space to enter said magnetic resonance measurement data into both a central region and an outer region of k-space;in a reconstruction processor, reconstructing the magnetic resonance image data from the magnetic resonance measurement data entered into k-space;in an identification processor, identifying the organ structure in the magnetic resonance image data; andspecifying said sampling scheme of k-space in said magnetic resonance data acquisition sequence to support said identification of the organ structure in the magnetic resonance image data reconstructed from the magnetic resonance measurement data entered into k-space, said sampling scheme of k-space being configured to produce a higher sampling density in the outer region of k-space than in the central region of k-space, wherein said higher sampling density in said outer region emphasizes an outer contour of the organ structure in the reconstructed magnetic resonance image data, said identifying of said organ structure using the emphasized outer contour of the organ structure in the magnetic resonance image data. 2. A method as claimed in claim 1, wherein said sampling scheme of k-space comprises a spiral sampling or radial sampling of k-space, with a higher sampling density in said outer region than in said central region. 3. A method as claimed in claim 1 wherein said specifying of said sampling scheme of k-space is performed in a processor provided with training magnetic resonance image data acquired from at least one subject other than said examination subject, and in which said organ structure of said at least one subject is already present and identified in said training magnetic resonance image data. 4. A method as claimed in claim 3 comprising determining said sampling scheme of k-space in said processor based on said training magnetic resonance data by calculating a sampling mask based on an outer contour of the organ structure of said at least one subject that is identified in the training magnetic resonance image data, and entering said magnetic resonance measurement data into k-space according to said sampling mask to cause an outer contour of the organ structure of the examination subject to be emphasized in the reconstructed magnetic resonance image data, and identifying said organ structure of the examination subject using the emphasized outer contour of the organ structure in the magnetic resonance image data. 5. A method as claimed in claim 1 wherein said specifying of said sampling scheme of k-space is configured to produce a specific emphasis or a specific saturation of the organ structure, compared with sur... | Siemens Healthcare GmbH | A61B005/055 G01R033/56 G01R033/48 G01R033/54 A61B005/00 A61N005/10 G01R033/58
US10722176 | B2 | 2020-07-28 | Methods and systems for automatically identifying detection parameters for an implantable medical device | 2019-04-03 | An initial set of parameters for operating one or more detection tools is automatically derived and subsequently adjusted so that each detection tool is more or less sensitive to signal characteristics in a region of interest. Detection tool(s) may be applied to physiological signals sensed from a patient (such as EEG signals) and may be configured to run in an implanted medical device that is programmable with the parameters to look for rhythmic activity, spike activity, and power changes in the sensed signals, etc. A detection tool may be selected, and parameter values derived in a logical sequence and/or in pairs based on a graphical representation of an activity type which may be selected by a user, for example, by clicking and dragging on the graphic via a GUI. Displayed simulations allow a user to assess what will be detected with a derived parameter set and then to adjust the sensitivity of the set or start over as desired. | 1. A method comprising: defining a first region of interest within an electrographic signal;processing the first region of interest to determine a plurality of detector parameter values that define an instance of a first selected detection tool;running a simulation of the instance of the first selected detection tool on one or more regions of the electrographic signal, to determine if a detection criterion is met; andif the detection criterion is met, implementing the instance of the first selected detection tool in an active implantable medical device. 2. The method of claim 1, wherein defining the first region of interest within the electrographic signal comprises: presenting a representation of the electrographic signal on a graphical user interface;receiving through a user input, a selection of a portion of the electrographic signal; andautomatically selecting a sub portion of the selected portion of the electrographic signal as the first region of interest. 3. The method of claim 2, wherein the sub portion comprises a prefixed duration of the selected portion of the electrographic signal. 4. The method of claim 1, wherein the processing further comprises selecting the first selected detection tool from among a plurality of different types of detection tools based on one or more characteristics of the first region of interest. 5. The method of claim 4, wherein the plurality of different types of detection tools comprises a rhythmic detection tool, a spike detection tool, and a nonlinear energy increase detection tool. 6. The method of claim 1, wherein the plurality of detector parameter values defines how the instance of the detection tools operates. 7. The method of claim 1, wherein the one or more regions of the electrographic signal on which a simulation of the instance of the selected detection tool is run excludes the first region of interest. 8. The method of claim 1, further comprising, if the detection criterion is not met: identifying a second region of interest based on detection outputs resulting from the simulation of the instance of the first selected detection tool on a portion of the electrographic signal on either side of the first region of interest;processing the second region of interest to determine a plurality of detector parameter values that define an instance of a second selected detection tool; andrunning a simulation of the instance of the second selected detection tool on one or more regions of the electrographic signal, to determine if the detection criterion is met. 9. The method of claim 8, wherein identifying the second region of interest comprises: detecting one or more non-detection segments in the portion of the electrographic signal on either side of the first region of interest, wherein the one or more non-detection segments correspond to a time interval where consecutive detection outputs indicate no detection has occurred; anddefining the start of the second region of interest at a predetermined t... | NeuroPace, Inc. | G08B023/00 A61B005/00 A61B005/0476 A61B005/048 A61N001/36 A61B005/04
US10713929 | B2 | 2020-07-14 | Alarm and monitoring system and method of operation thereof | 2019-01-04 | An alarm and monitoring system including a primary device and at least one secondary device, the alarm and monitoring system including at least one controller configured to: determine whether at least one alarm event is set; establish a wireless communication between a primary device and the secondary device, when it is determined that the alarm event has been set; transmit an alarm event signal including alarm information from the primary device to the secondary device in accordance with the alarm event that is determined to have been set; generate an alarm signal by the secondary device in accordance with at least the alarm information; and render the generated alarm signal on a rendering device. | 1. An alarm system comprising: a primary device having a first controller; andat least one secondary device having at least one secondary controller and at least one rendering device;wherein the first controller is configured to (i) determine whether at least one alarm event is set on the primary device, (ii) establish a wireless communication between the primary device and the secondary device when determination is made that the alarm event has been set on the primary device, (iii) generate a first alarm signal including first alarm rendering instructions based on the alarm event set, and (iv) transmit the first alarm signal including the first alarm rendering instructions from the primary device to the secondary device, andwherein the secondary controller is configured to (i) generate second alarm rendering instructions for the rendering device based on the first alarm rendering instructions and (ii) render the second alarm rendering instructions on the rendering device. 2. The alarm and monitoring system of claim 1, wherein the first controller is configured to determine capabilities of the secondary device prior to generating the first alarm signal including the first alarm rendering instructions. 3. The alarm and monitoring system of claim 2, wherein the first controller is configured to generate the first alarm signal in accordance with the determined capabilities of the secondary device. 4. The alarm and monitoring system of claim 3, wherein the first controller is configured to determine whether communication is established between the primary and secondary devices prior to generating the first alarm signal. 5. The alarm and monitoring system of claim 4, wherein the first controller is configured to provide a user interface in accordance with the determined capabilities of the secondary device after the first controller determines that communication is established. 6. The alarm and monitoring system of claim 1, wherein the secondary device includes a memory, and wherein the first controller is configured to generate the first alarm signal in accordance with rendering data that is pre-stored in the memory of the secondary device. 7. The alarm and monitoring system of claim 1, wherein the secondary device is configured to control one or more tertiary devices external to the secondary device to additionally render the second alarm rendering instructions on the tertiary device. 8. The alarm and monitoring system of claim 1, wherein the at least one rendering device comprises at least one of an auditory, visual, and haptic rendering device to render the second alarm rendering instructions. 9. A method of operating an alarm and monitoring system comprising: determining by a first controller on a primary device whether at least one alarm event is set on the primary device;after determining the alarm event is set, establishing by the first controller communication between the primary device and a secondary device;generating by the first c... | JWIN ELECTRONICS CORPORATION | G08B025/10 A61B005/00 H04M001/725 H04M019/04 A61B005/11
US10702202 | B2 | 2020-07-07 | Feedback device with stimulation of sensory nerves for improved balance control | 2019-07-25 | A feedback device for measuring pressure related information, and for providing that information in a way that is useful to the wearer of the device. At least one sensor detects pressure information and transmits at least one pressure information signal to a signal processing subsystem. The signal processing subsystem converts the received pressure information signal into at least one stimulation control signal. The signal processing subsystem then transmits the stimulation control signal to at least one stimulator, which provides stimulation to a wearer of the device reflecting the stimulation control signal received from the signal processing subsystem. | 1. A system for improving balance control, the system comprising: a) at least one sensor positionable against a limb of the user, wherein the at least one sensor is configured to generate user balance information signals and transmit the user balance information signals, wherein the user balance information signals comprise an amount of force applied to the at least one sensor;b) a signal processing subsystem, the subsystem configured to receive the user balance information signals and generate balance control signals comprising temporal and spatial information reflecting user balance information; andc) at least one stimulator configured to be positionable adjacent to at least one sensory neuron or sensory nerve of the user, the at least one stimulator configured to be responsive to the balance control signals to provide stimulation to the at least one sensory neuron or sensory nerve, thereby providing feedback to the user. 2. The system of claim 1, wherein the stimulation comprises electrical stimulation. 3. The system of claim 1, wherein the force is normal force or shear force. 4. A system for improving balance control, the system comprising: a) at least one sensor positionable against a limb of the user, wherein the at least one sensor is configured to generate user balance information signals and transmit the user balance information signals, wherein the user balance information signals comprise time histories of force applied to the at least one sensor;b) a signal processing subsystem, the subsystem configured to receive the user balance information signals and generate balance control signals comprising temporal and spatial information reflecting user balance information; andc) at least one stimulator configured to be positionable adjacent to at least one sensory neuron or sensory nerve of the user, the at least one stimulator configured to be responsive to the balance control signals to provide stimulation to the at least one sensory neuron or sensory nerve, thereby providing feedback to the user. 5. The system of claim 1, wherein the at least one stimulator is configured to be responsive to the balance control signals to provide stimulation to the at least one sensory neuron or sensory nerve when at least one of the user balance information signals exceeds a predetermined threshold. 6. The system of claim 5, wherein the predetermined threshold is a predetermined force threshold or a predetermined temporal or spatial threshold. 7. The system of claim 1, wherein the amount of force comprises the amount of force applied to a sole of the user's foot. 8. The system of claim 1, wherein the at least one sensor is a foot pressure sensor. 9. The system of claim 1, wherein the at least one stimulator is configured to be implantable within the body of the user. 10. A system for improving balance control, the system comprising: a) at least one sensor positionable under a user's foot, wherein the at least one sensor is configured to generate... | RxFunction, Inc. | A61B005/00 A61B005/103 A61B005/11 A61H003/00 A61N001/36 G09B019/00 A61H023/02
US10725435 | B2 | 2020-07-28 | Motion capture apparatus | 2018-12-07 | To allow combining devices having different functions according to an application, and improve convenience for a user. There are provided a wearing type terminal 1 provided with an external case 10 having a display unit 131 on the upper surface side and belt members 4a and 4b which can be connected to the both sides of the external case 10 respectively; and a docking type device 2 which is detachably engaged with the bottom portion of the external case 10 of the wearing type terminal 1, and is electrically connected to the wearing type terminal 1 via a dock side connection terminal 21. The wearing type terminal 1 and the docking type device 2 are provided with a function which carries out transmission and receiving of data to be displayed on the display unit 131 via the dock side connection terminal 21. | 1. A motion capture apparatus, comprising: an information terminal device which is wearable by a wearer or can be mounted on a bicycle; andsensors which are connected to the information terminal device through wireless communication, whereinthe sensors include body motion sensors which are worn on parts of a body of the wearer to detect three-dimensional displacement and acceleration of the parts of the body,the information terminal device includes a body motion data acquisition unit, a body motion calculation unit, a cycle extracting unit, a correction unit and an analysis unit,the body motion data acquisition unit acquires body motion data from the sensors through the wireless communication,the body motion data includes: sensor identification information for identifying the sensors and identifying the parts on which the sensors are worn; andtime information when a detection result is acquired from the sensors,the body motion calculation unit identifies the parts on which the sensors are worn by determining the sensors from which the detection result is outputted,the body motion calculation unit calculates body motion calculation data of a motion and posture of the wearer on the basis of the detection result of the body motion sensors and relative positional relationship among the body motion sensors determined by a skeleton of the wearer and positions of the body motion sensors, the relative positional relationship including mutual distances,the cycle extracting unit extracts a three-dimensional circular motion from the body motion calculation data,the correction unit corrects the body motion calculation data calculated by the body motion calculation unit on the assumption that the three-dimensional circular motion reaches an end point of the current cycle and a start point of the next cycle at a certain reference point of the three-dimensional circular motion on the basis of the detection result of the cycle extracting unit,the analysis unit analyzes body motions on the basis of the body motion calculation data corrected by the correction unit and displays or outputs an analysis result,at least one of the body motion sensors is attached to the parts of the body corresponding to a swinging point defined on the center line of the body of the wearer to detect a three-dimensional displacement or acceleration of the swinging point,the body motion calculation unit calculates reference orbital locuses projected on a vertical plane in parallel with a gravity direction and the direction in which the wearer runs on the basis of the detection result of the cycle extracting unit,the body motion calculation unit calculates three-dimensional free orbital locuses of the body motion sensors on the basis of the detection results of body motion sensors,the body motion calculation unit estimates relative displacements among the body motion sensors on both the left and right sides of the wearer on the basis of the differences between the calculated reference orbi... | LEOMO, Inc. | G04G021/02 A61B005/11 G06F001/16 G04G021/04 G01S019/01 G04G017/08 G06F003/0354 G06F003/041 G06F003/0488
US10702159 | B2 | 2020-07-07 | System and apparatus for capturing and navigating whole body images that includes high resolution body part images | 2016-03-21 | A system and method for capturing and navigating whole body images that includes capturing a total body image with a wide angle lens camera and simultaneously capturing body part images of the same body using a plurality of high resolution cameras. The total body image is used as a navigational grid that is associated with the high resolution body part images. | 1. A total body imaging apparatus comprising: an enclosed interior area;a camera with a wide angle lens positioned to capture a total body image of a user located in the enclosed interior area; andeight high resolution cameras positioned to capture different body portions of the user that together form a total body image of the user, said eight high resolution cameras being located within the enclosed interior area; anda computer processing unit in communication with at least one program application that triggers the wide angle lens camera and the eight high resolution cameras to simultaneously capture images and to arrange eight high resolution body part images taken with the eight high resolution cameras for a particular pose and associate them with the total body image captured with the wide angle lens camera for the same pose. 2. The total body imaging station of claim 1 wherein the enclosed interior area comprises a front panel which includes a plurality of openings therein positioned in front of the wide angle lens camera and the eight high resolution cameras such that a lens of each of said wide angle lens camera and said eight high resolution cameras can be seen through each of said respective plurality of openings. 3. The total body imaging apparatus of claim 1 wherein the wide angle lens camera is positioned between two rows of cameras that comprise the eight high resolution cameras. 4. The total body imaging apparatus of claim 1 wherein the eight high resolution cameras are configured to form a first row of three cameras, a second row of three cameras located below the first row of three cameras, and a third row of two cameras located below the second row of three cameras. 5. The total body imaging apparatus of claim 4 wherein the wide angle lens camera is positioned between the first row of three high resolution cameras and the second row of three high resolution cameras. 6. The total body imaging, apparatus of claim 5 wherein the enclosed interior area comprises a front panel which includes a plurality of openings therein positioned in front of the wide angle lens camera and the eight high resolution cameras such that a lens of each of said wide angle lens camera and said eight high resolution cameras can be seen through each of said respective plurality of openings. 7. The total body imaging apparatus of claim 5 wherein the program application uses the total body image as a navigational grid by sectioning the total body image into tiles that correspond to the location of the eight high resolution cameras and their related high resolution body part images where any of the eight tiles can be selected to magnify and further view in more detail an area of skin within a high resolution body part image that is associated with the selected tile. 8. The total body imaging apparatus of claim 7 wherein the program application further includes a reduced size image of the eight tile navigational grid displayed on a high resolution body p... | Techara LLC | H04N007/18 A61B005/00 H04N005/232 H04N005/247
US10702226 | B2 | 2020-07-07 | System and method for local three dimensional volume reconstruction using a standard fluoroscope | 2018-02-08 | A system for constructing fluoroscopic-based three-dimensional volumetric data of a target area within a patient from two-dimensional fluoroscopic images including a structure of markers, a fluoroscopic imaging device configured to acquire a sequence of images of the target area and of the structure of markers, and a computing device. The computing device is configured to estimate a pose of the fluoroscopic imaging device for at least a plurality of images of the sequence of images based on detection of a possible and most probable projection of the structure of markers as a whole on each image of the plurality of images. The computing device is further configured to construct fluoroscopic-based three-dimensional volumetric data of the target area based on the estimated poses of the fluoroscopic imaging device. | 1. A system for constructing fluoroscopic-based three-dimensional volumetric data of a target area within a patient from two-dimensional fluoroscopic images, comprising: a structure of markers; anda computing device configured to: receive a sequence of fluoroscopic images of the target area captured by a fluoroscopic imaging device, each image comprising at least a portion of the structure of markers;estimate a pose of the fluoroscopic imaging device for a plurality of images of the sequence of fluoroscopic images based on detection of a projection of the structure of markers on each image of the sequence of images;construct fluoroscopic-based three-dimensional volumetric data of the target area based on each estimated pose of the fluoroscopic imaging device, the fluoroscopic-based three-dimensional volumetric data of the target area including a medical device and a target; andanalyze the constructed fluoroscopic-based three-dimensional volumetric data to determine an offset between a location of the medical device within the constructed fluoroscopic-based three-dimensional volumetric data of the target area and a location of the target within the constructed fluoroscopic-based three-dimensional volumetric data of the target area. 2. The system of claim 1, wherein the computing device is further configured to facilitate an approach of the medical device to the target. 3. The system of claim 1, further comprising a locating system configured to determine the location of the medical device within the patient, wherein the computing device comprises a display and is configured to: display the target area and the location of the medical device with respect to the target based on the location of the medical device as determined by the locating system;facilitate navigation of the medical device to the target area via the locating system and the display; andcorrect the display of the location of the medical device with respect to the target based on the determined offset between the location of the medical device within the constructed fluoroscopic-based three-dimensional volumetric data and the location of the target within the constructed fluoroscopic-based three-dimensional volumetric data as determined by the computing device based on the constructed fluoroscopic-based three-dimensional volumetric data of the target area. 4. The system of claim 3, wherein the computing device is further configured to: display a 3D rendering of the target area on the display; andregister the locating system to the 3D rendering,wherein the display of the location of the medical device with respect to the target is corrected by updating a registration between the locating system and the 3D rendering. 5. The system of claim 3, wherein the locating system is an electromagnetic locating system. 6. The system of claim 3, wherein the target area comprises at least a portion of lungs and the medical device is navigated to the target area through a luminal network. 7. Th... | COVIDIEN LP | A61B005/05 A61B006/12 A61B006/00 A61B034/20 A61B090/00
US10712197 | B2 | 2020-07-14 | Optical sensor package | 2018-01-11 | An optical device package is disclosed. The optical device package includes a substrate that passes light at an optical wavelength. The optical device package also includes an optical device assembly that is mounted to the substrate. The optical device assembly comprises an optical device die. The optical device die has a first surface that is mounted to and facing the substrate and a second surface that is opposite the first surface. The optical device package further includes a molding compound that is disposed at least partially over the second surface of the integrated device die. | 1. An optical device package comprising: a substrate that passes light at an optical wavelength, the substrate having a first side, a second side opposite the first side and a side wall extending between the first and second sides;an optical device assembly mounted to the second side of the substrate, the optical device assembly comprising an integrated device die, the integrated device die having a first surface mounted to and facing the substrate and a second surface opposite the first surface; anda molding compound disposed at least partially over the second surface of the integrated device die and at least partially along a side wall of the substrate. 2. The optical device package of claim 1, wherein the substrate comprises glass. 3. The optical device package of claim 1, wherein the integrated device die comprises a sensor die. 4. The optical device package of claim 3, wherein the sensor die comprises an active surface facing the substrate. 5. The optical device package of claim 1, wherein the second surface has one or more electrical contacts configured to electrically connect the integrated device die to an external device through the molding compound. 6. The optical device package of claim 5, further comprising one or more through mold vias (TMVs) that extend from the one or more electrical contacts through the molding compound, the TMVs configured to receive electrical interconnects, the electrical interconnects electrically connect the optical device assembly to the external device. 7. The optical device package of claim 1, wherein the substrate comprises one or more of a filter, a diffuser, a polarizer, and a waveguide or beam steering element. 8. The optical device package of claim 1, wherein the substrate comprises a first portion for mounting the integrated device die and a second portion for mounting a second integrated device die of the optical device assembly. 9. The optical device package of claim 8, wherein the first portion and second portion of the substrate are spaced apart by the molding compound. 10. The optical device package of claim 1, wherein the molding compound disposed around side edges of the integrated device die. 11. An optical device package comprising: a substrate that passes light at an optical wavelength, the substrate having a first side, a second side opposite the first side and a side wall extending between the first and second sides;an optical device die having a first surface mounted to the second side of the substrate, the optical device die configured to transmit or detect light through the substrate at the optical wavelength;a processor die mounted to the second side of the substrate, the processor die electrically connected to the optical device die; anda frame disposed at least partially along the side wall of the substrate. 12. The optical device package of claim 11, wherein the optical device die comprises an optical emitter die configured to emit light through the substrate. 13. The o... | Analog Devices Global Unlimited Company | G01J001/02 G01J001/42 A61B005/024 G01J001/04
US10709357 | B2 | 2020-07-14 | Device, system, and method for assessing intravascular pressure | 2015-05-01 | An apparatus for assessing the severity of stenosis in a blood vessel includes an elongate body including a distal portion and a centering assembly. The centering assembly is actuatable to selectively center the elongate body in the vessel. A pressure sensor is disposed adjacent the centering assembly and is configured to detect fluid pressure in the vessel. A processing system receives the measured pressure from the pressure sensor, receives data representing the cross-sectional area of the vessel, receives data representing the size of the distal portion, calculates a offset correlation based on the size of the distal portion and based on the size of the vessel, and calculates a fractional flow reserve (FFR) for the vessel as an index of stenosis severity taking into account the offset correlation and the measured fluid pressure from the pressure sensor. | 1. An apparatus for assessing the severity of stenosis in a blood vessel, comprising: an elongate body including a proximal portion and a distal portion, the body defining a lumen extending through at least the distal portion, the lumen sized and shaped to allow the passage of a guidewire therethrough;a centering assembly disposed at the distal portion of the elongate body, the centering assembly being actuatable to selectively center the elongate body in the vessel;a pressure sensor disposed adjacent the centering assembly, the pressure sensor being configured to detect fluid pressure in the vessel; anda processing system configured to: receive the measured pressure at a location within the vessel, from the pressure sensor;receive, from at least one of a sensor coupled to the elongate body or an x-ray imaging device configured to image the centering assembly, data representing a degree of expansion of the centering assembly;receive data representing the size of the distal portion;in response to receiving the data representing the degree of expansion of the centering assembly, calculate an offset correlation based on the data representing the size of the distal portion and based on the data representing the degree of expansion of the centering assembly; andcalculate a fractional flow reserve (FFR) for the vessel as an index of stenosis severity taking into account the offset correlation and the measured fluid pressure from the pressure sensor. 2. The apparatus of claim 1, further comprising a user interface, the FFR being displayed to a clinician on the user interface. 3. The apparatus of claim 1, wherein the processing system is configured to receive data representing the degree of expansion of the centering assembly based on measurements taken by the centering assembly. 4. The apparatus of claim 1, wherein the processing system is configured to determine a size of the vessel based on the degree of expansion of the centering assembly, and wherein the offset correlation is based on a ratio of the distal portion size to the size of the vessel. 5. The apparatus of claim 4, wherein the size of the distal portion and the size of the vessel are diametrical measurements. 6. The apparatus of claim 1, wherein the centering assembly comprises an expandable portion configured to contact an inner wall of the blood vessel. 7. The apparatus of claim 6, wherein the expandable portion comprises a proximal expandable portion fixed in place relative to the elongate body and a distal expandable portion moveable with respect to the elongate body. 8. The apparatus of claim 7, wherein the expandable portion comprises legs extending between the proximal expandable portion and the distal expandable portion, the legs being flexible to expand radially outwardly when the distal expandable portion is moved relative to the proximal expandable portion. 9. The apparatus of claim 1, further comprising an IVUS imaging array disposed on the distal portion. 10. The appar... | KONINKLIJKE PHILIPS N.V. | A61B005/0215 A61B005/107 A61B008/08 A61B005/00 A61B008/12 A61B005/02 A61B008/00 A61B008/06
US10720037 | B2 | 2020-07-21 | Wearable alert system | 2019-08-06 | A system that stores information indicative of real-time locations of a wearable alert devices, receives information indicative of a potentially dangerous condition in a geographic area, determines which of the wearable alert devices are in the geographic area, and outputs alerts for transmittal to the wearable alert devices in the geographic area. | 1. A computer-implemented method, comprising: receiving, by a server that is remote from a user wearable alert device, information indicative of a real-time location of a user portable electronic device that is paired with the user wearable alert device;receiving, by the server, information indicative of a potentially dangerous condition in a geographic area;determining, by the server, whether the real-time location of the user portable electronic device is in the geographic area; andoutputting an alert, by the server, to a communications network for transmittal to the user wearable alert device via the user portable electronic device in response to a determination that the user portable electronic device is in the geographic area. 2. The method of claim 1, wherein the information indicative of the potentially dangerous condition is received from an environmental sensor. 3. The method of claim 2, wherein the environmental sensor is included in the user wearable alert device. 4. The method of claim 1, wherein the server receiving the information indicative of the real-time location of the user portable electronic device comprises: repeatedly receiving, by the server, information indicative of a real-time location of the user portable electronic device; andautomatically and repeatedly updating a device location database, by the server, to reflect the real-time location of the user portable electronic device. 5. The method of claim 1, wherein the real-time location of the user portable electronic device is determined by the user portable electronic device. 6. The method of claim 1, wherein the real-time location of the user portable electronic device is determined by a network device. 7. The method of claim 1, further comprising: storing a plurality of user profiles each associating a user wearable alert device with a user. 8. The method of claim 7, wherein the user profiles also include user preferences, the method further comprising: determining, by the server, whether the user preferences of a user indicate a preference to receive an alert based on the potentially dangerous condition; andoutputting the alert for transmittal to the user wearable alert device associated with the user via the user portable electronic device associated with the user in response to determinations that the user portable electronic device associated with the user is in the geographic area and the user preferences associated with the user indicate a preference to receive an alert based on the potentially dangerous condition. 9. The method of claim 7, further comprising: receiving a personal alert for a user, by the server, from a third party service;determining, by the server, the wearable alert device associated with the user; andoutputting the personal alert to a communications network for transmittal to the wearable alert device associated with the user via the user portable electronic device associated with the user. 10. A server, remote from a user wearable... | Locator IP, L.P. | G08B027/00 G08B021/04 A61B005/00 G08B025/01
US10709934 | B2 | 2020-07-14 | Route planning method and wearable device | 2019-08-14 | A route planning method includes obtaining exercise capability information of a wearer and one or more candidate routes, where the candidate routes include attribute features that comprise historical exercise capability information, where the historical exercise capability information is information calculated according to a first preset rule and based on obtained exercise capability information of a plurality of users having exercised along the candidate routes; determining a target route based on the attribute features of the candidate routes and the exercise capability information of the wearer; and outputting the target route information. | 1. A route planning method applied to a wearable device, wherein the method comprises: obtaining historical exercise capability information according to a first preset rule;obtaining exercise capability information of a wearer of the wearable device and information for a plurality of candidate routes, wherein the information comprises attribute features of the candidate routes comprising the historical exercise capability information, wherein the historical exercise capability information is based on exercise capability information of a plurality of users that have exercised along the candidate routes;determining a target route based on the attribute features and the exercise capability information of the wearer;outputting the target route to permit selection of the target route on the wearable device;monitoring a body index of the wearer in real time; andoutputting warning information when the body index is abnormal. 2. The method of claim 1, further comprising: dividing the candidate routes into k subroutes according to a second preset rule;calculating an attribute feature of each subroute of the k subroutes; andoutputting the target route with the attribute feature of the each subroute. 3. The method of claim 1, further comprising: obtaining exercise requirement information of the wearer; anddetermining the target route based on the exercise requirement information of the wearer. 4. The method of claim 1, further comprising: obtaining environment information at a first moment, wherein the first moment is a moment before the wearer exercises with the wearable device; anddetermining the target route based on the environment information at the first moment. 5. The method of claim 4, further comprising outputting device recommendation information and material supplement recommendation information based on the exercise capability information of the wearer, an attribute feature of the target route, and the environment information at the first moment. 6. The method of claim 1, wherein after outputting the target route, the method further comprises: receiving a first operation instruction comprising a selection instruction operation that is performed on a plurality of target routes;displaying, according to the first operation instruction, the target route that is selected on the wearable device; andupdating the attribute feature of the target route based on the exercise capability information of the wearer. 7. The method of claim 6, further comprising: obtaining environment information at a second moment, wherein the second moment is a moment at which the wearer exercises with the wearable device;determining, based on the environment information at the second moment, whether to adjust the target route that is selected on the wearable device; andoutputting adjustment indication information based on determining to adjust the target route that is selected on the wearable device. 8. The method of claim 1, further comprising: completing an exercise ... | HUAWEI TECHNOLOGIES CO., LTD. | H04W024/00 A63B024/00 A61B005/11 G01C022/00 G06F001/16
US10702168 | B2 | 2020-07-07 | Blood pressure estimation apparatus, blood pressure estimation method, and non-transitory computer-readable recording medium | 2017-10-27 | A blood pressure estimation apparatus includes: a pulse wave measurement unit measuring a user's pulse wave signal; an electrocardiogram signal measurement unit measuring a user's electrocardiogram signal; and a processing circuit estimating a user's blood pressure value by using the pulse wave and electrocardiogram signals, wherein the processing circuit (a) acquires a pulse wave signal feature by using the pulse wave signal, (b) acquires an electrocardiogram signal feature by using the electrocardiogram signal, (c) acquires a pulse wave propagation time by using the pulse wave and electrocardiogram signals, (d) selects a blood pressure group indicating a relationship between a pulse wave propagation time and blood pressure of the user by using the pulse wave signal feature, electrocardiogram signal feature, and pulse wave propagation time, and (e) estimates the blood pressure value by using the blood pressure group, pulse wave signal feature, electrocardiogram signal feature, and pulse wave propagation time. | 1. A blood pressure estimation apparatus comprising: a processing circuit that estimates a blood pressure value of the user by using a pulse wave signal of the user acquired from a pulse wave sensor and an electrocardiogram signal of the user acquired from an electrocardiogram sensor,wherein the processing circuit (a) acquires a pulse wave signal feature by using the pulse wave signal of the user, the pulse wave signal feature being a time-frequency feature of the pulse wave signal, (b) acquires an electrocardiogram signal feature by using the electrocardiogram signal of the user, the electrocardiogram signal feature being a time-frequency feature of the electrocardiogram signal, (c) acquires a pulse wave propagation time by using the pulse wave signal of the user and the electrocardiogram signal of the user, (d) selects a blood pressure group by using the pulse wave signal feature, the electrocardiogram signal feature, and the pulse wave propagation time, the blood pressure group being indicative of a relationship between a pulse wave propagation time and blood pressure of the user, and (e) estimates the blood pressure value of the user by using the blood pressure group, the pulse wave signal feature, the electrocardiogram signal feature, and the pulse wave propagation time. 2. The blood pressure estimation apparatus according to claim 1, wherein the blood pressure group includes a first group having a first blood pressure value with respect to the pulse wave propagation time and a second group having a blood pressure value lower than the first blood pressure value with respect to the pulse wave propagation time; andin (d), the processing circuit selects the first group or the second group by using the pulse wave signal feature, the electrocardiogram signal feature, and the pulse wave propagation time. 3. The blood pressure estimation apparatus according to claim 1, wherein the processing circuit acquires a pulse wave wavelet feature by using the pulse wave signal of the user in (a) and acquires an electrocardiogram wavelet feature by using the electrocardiogram signal of the user in (b). 4. The blood pressure estimation apparatus according to claim 1, wherein the processing circuit acquires a pulse wave wavelet feature by using a pulse wave waveform corresponding to a single heartbeat of the user in (a) and acquires an electrocardiogram wavelet feature by using an electrocardiogram waveform corresponding to a single heartbeat of the user in (b). 5. The blood pressure estimation apparatus according to claim 1, wherein in (e), the processing circuit acquires an estimation rule corresponding to the blood pressure group and estimates the blood pressure value of the user by using the estimation rule, the pulse wave signal feature, the electrocardiogram signal feature, and the pulse wave propagation time. 6. The blood pressure estimation apparatus according to claim 5, wherein the estimation rule is an estimation formula using the pulse wave sig... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | A61B005/021 A61B005/0452 A61B005/00 A61B005/024 A61B005/0245 A61B005/0404 A61B005/0408
US10722745 | B2 | 2020-07-28 | Interactive cycling system and method of using muscle signals to control cycling pattern stimulation intensity | 2018-06-05 | Systems and methods for a muscle signal-driven cycling system for persons with disability for rehabilitation are provided. A system comprises integrating both motor power and muscle power to facilitate rehabilitation cycling-based exercises. By using the intensity of real-time muscle activity signals as inputs, a motor applies either assistive or resistive force to rotate a gear at different speeds to facilitate or impede the cycling motion, and the electrical pulses from an electrical stimulation device can be provided to stimulate target muscles to generate muscle contraction to support the continuous cycling movement. | 1. A system for using muscle signals to control a cycling pattern, the system comprising: a pedal;a crank arm connected to the pedal;a motor connected to the crank arm;a stimulator connected to a plurality of stimulation electrodes;a controller connected to the motor and the stimulator; anda data acquisition system connected to the controller and a muscle activity sensor;wherein the motor is configured to apply either an assistive force or a resistive force to the crank arm,wherein the stimulator is configured to connect to the plurality of stimulation electrodes and transmit a stimulation signal,wherein the data acquisition system is configured to receive a real-time muscle signal,wherein the controller is configured to direct the motor to apply either the assistive force or the resistive force based upon the real-time muscle signal, andwherein the controller is configured to direct the stimulator to increase or decrease an intensity of the stimulation signal based upon the real-time muscle signal. 2. The system of claim 1, wherein the pedal is affixed to an ankle orthosis with an L-shaped profile. 3. The system of claim 1, further comprising: a force sensor connected to the pedal;a force sensor connected to the crank arm;a front gear connected to the crank arm;a rear gear connected to the front gear by a chain; anda torque sensor connected to the rear gear, the motor, and the controller. 4. The system of claim 1, wherein for unilateral leg training the system is configured to attach a muscle sensor and stimulation electrodes to a single leg to train only that leg. 5. The system of claim 1, wherein for bilateral leg training the system is configured to attach muscle sensors and stimulation electrodes to both left and right legs. 6. The system of claim 1, wherein the muscle activity sensor and a pair of stimulation electrodes are configured to be attached to a hamstring muscle of a user, a quadriceps muscle of a user, a gastrocnemius muscle of a user, or a tibialis anterior muscle of a user. 7. The system of claim 1, wherein the stimulator is configured to stimulate a hamstring muscle of a user by emitting the stimulation signal when a rotational angle of the crank arm changes as the hamstring muscle contracts. 8. The system of claim 1, wherein the stimulator is configured to stimulate a quadriceps muscle of a user by emitting the stimulation signal when a rotational angle of the crank arm changes as the quadriceps muscle contracts. 9. The system of claim 1, wherein the stimulator is configured to stimulate a gastrocnemius muscle of a user by emitting the stimulation signal when a rotational angle of the crank arm changes as the gastrocnemius muscle contracts. 10. The system of claim 1, wherein the stimulator is configured to stimulate a tibialis anterior muscle of a user by emitting the stimulation signal when a rotational angle of the crank arm changes as the tibialis anterior muscle contracts. 11. A method for using muscle signals to... | The Chinese University of Hong Kong | A63B021/00 A63B022/06 A63B024/00 A61B005/0488 A61N001/04 A61F005/01 A61N001/36 A63B021/005 A63B022/00
US10720078 | B2 | 2020-07-21 | Systems and methods for extracting keywords in language learning | 2019-06-14 | Systems, methods, and products for language learning that may extract text from various resources having text, using various natural-language processing features, which can be combined with custom-designed learning activities to offer a needs-based, adaptive learning methodology. The system may receive a resource, extract keywords pedagogically valuable to non-native language learning and academic exercises. Metadata describing various aspects of resources from which keywords are extracted may be associated with keywords. Metadata describing various aspects of keywords may also be associated with keywords. Extracted keywords may be stored into a keyword store along with any metadata associated with keywords. | 1. A server-implemented method comprising: parsing, by a server, a set of one or more potential keywords from a text;determining, by the server, for a subset of potential keywords in the set of one or more potential keywords, a score of each of the subset of potential keywords associated with a key of n-grams and a set of one or more collocations based on one or more scoring rules determining the score;determining, by the server, a pedagogical value threshold based upon a level score associated with a client device, wherein the level score is based on a plurality of responsive inputs received from the client device during a first learning session associated with the server and the client device;storing, by the server, into a keyword store, one or more potential keywords of the subset of potential keywords having the score that satisfies the pedagogical value threshold; andgenerating, by the server, a learner activity for the client device configured to test the one or more potential keywords. 2. The server-implemented method according to claim 1, further comprising: generating, by the server, a graphical user interface on the client device associated with the first learning session. 3. The server-implemented method according to claim 2, wherein the graphical user interface is configured to display one or more input elements configured to receive the plurality of responsive inputs from the client device. 4. The server-implemented method according to claim 1, further comprising: updating, by the server, the level score associated with the client device upon receiving new responsive inputs in a second learning session. 5. The server-implemented method according to claim 1, further comprising: determining, by the server, the pedagogical value threshold based upon a resource difficulty score of the text. 6. The server-implemented method according to claim 5, further comprising: determining, by the server, the resource difficulty score of the text based upon one or more characteristics of the text. 7. The server-implemented method according to claim 1, further comprising: receiving, by the server, the text from the client device via one or more networks. 8. The server-implemented method according to claim 1, further comprising: discarding, by the server, from the set of one or more potential keywords each potential keyword matching a filtered word in a file containing a stop word list. 9. The server-implemented method according to claim 8, wherein the filtered word in the stop word list is selected from a group consisting of a proper noun, an ordinal number, a number, a preposition, and a conjunction. 10. The server-implemented method according to claim 1, further comprising: discarding, by the server, from the set of one or more potential keywords each potential keyword not satisfying a keyword frequency score threshold. 11. A system comprising: a client device; anda server, wherein the server is configured to: parse a set of one or more pot... | VOXY, INC | G06F040/205 G09B019/06 A61B005/16 A61B005/00 A61B005/048 G06F040/253 G06F040/263 G09B005/02 G09B005/00 G09B007/08 A61B005/11 A61B005/0484 A61B005/0476
US10715216 | B2 | 2020-07-14 | Network of electronic devices assembled on a flexible support and communication method | 2019-02-11 | A smart button for use in a network formed on a garment includes a housing and an antenna carried within the housing to communicate with elements of the network. A functional element is carried within the housing. An electronic circuit is carried within the housing and coupled to the antenna and the at least one functional element. The housing is formed by a stem carrying a head, and the antenna is housed within the head. | 1. An apparatus, comprising: a smart button for use in a network formed on a garment, the smart button comprising: a button housing being substantially smaller in surface area than the garment, the button housing comprising: at least one stem fastened to the garment; anda head connected to and carried by the at least one stem, the head being sized and shaped to pass through a hole in the garment to thereby fasten a first portion of the garment to a second portion of the garment;an antenna carried within the button housing and configured to communicate with elements of the network formed on the garment;at least one non-electronic element carried within the button housing; andat least one integrated circuit carried within the button housing and coupled to the antenna and the at least one non-electronic element;wherein the at least one non-electronic element is a microelectromechanical device, a nanoelectromechanical device, a mechanical sensor, a well containing liquids, a cell for containing liquids, a vial for containing liquids, or a microchannel for containing liquids. 2. The apparatus of claim 1, wherein the at least one non-electronic element is housed within the at least one stem; further comprising at least one additional non-electronic element housed within the head, the at least one additional non-electronic element being a user interface component. 3. The apparatus of claim 2, wherein the antenna is housed within the head of the button housing. 4. The apparatus of claim 1, wherein the at least one integrated circuit comprises at least one of a memory element, a control unit, a converter, an adapter, a digital circuit, an analog circuit, a RF circuit, a transceiver circuit, an integrated circuit, a system on a chip, and a system in a package. 5. The apparatus of claim 1, wherein the antenna is embodied within the at least one integrated circuit. 6. The apparatus of claim 1, wherein, when the at least one element comprises a well for containing liquids, a cell for containing liquids, a vial for containing liquids, or a microchannel for containing liquids, the antenna receives a control signal from at least one other smart button indicating when to dispense the liquids. 7. The apparatus of claim 6, wherein the button housing includes therein an input channel enabling filling of the well for containing liquids, the cell for containing liquids, the vial for containing liquids, or the microchannel for containing liquids. 8. The apparatus of claim 7, wherein the input channel includes closing elements. 9. The apparatus of claim 1, wherein the antenna is shaped so as to form two turns arranged alongside each other and contiguous, each turn surrounding a respective hole defined in the button housing. 10. The apparatus of claim 1, wherein the antenna is formed from first and second concentric circumferences, with the second circumference being interior to the first circumference and passing between first and second holes defined in the bu... | STMicroelectronics S.r.l. | H04B005/00 H02J050/12 H02J050/40 H02J007/02 H02J050/70 H01F038/14 H01L023/538 H01L023/64 H01L023/66 H01L025/065 H01Q001/22 H01Q007/00 H01Q001/27 H01Q001/52 H02J050/10 A61B005/00
US10709338 | B2 | 2020-07-14 | Useful information presentation device | 2016-10-17 | A useful information presentation device includes a facial skin temperature acquisition unit and/or a blood circulation volume acquisition unit, a brain activity estimation unit, and a useful information presentation unit. The facial skin temperature acquisition unit acquires facial skin temperature data in time-series. The blood circulation volume acquisition unit acquires photographic image data by imaging the facial surface in time-series. The brain activity estimation unit decomposes the facial skin temperature data and/or the facial blood circulation volume data into a plurality of components by singular value decomposition, principal component analysis, or independent component analysis, and the brain activity estimation unit estimating brain activity of the subject based on the plurality of components. The useful information presentation unit presents, to the subject, useful information related to an emotional or physical state based on the brain activity of the subject estimated by the brain activity estimation unit. | 1. A useful information presentation device comprising: at least one of an imaging device and a plurality of temperature sensors;a processor communicatively connected to the at least one of the imaging device and the plurality of temperature sensors; anda display device communicatively connected to the processor,the processor being configured to execute at least one of acquiring facial skin temperature data including skin temperature data detected from a facial surface of a subject based on image data captured by the imaging device or temperature data detected by the plurality of temperature sensors in a time series, and acquiring position data of a detection site of the facial skin temperature data in time-series, andobtaining photographic image data of the facial surface of the subject from the imaging device in time-series, obtaining RGB data of the photographic image data by subjecting the photographic image data to RGB processing, and acquiring facial blood circulation volume data based on the RGB data in time-series,the processor being further configured to decompose at least one of the facial skin temperature data and the facial blood circulation volume data into a plurality of components by singular value decomposition, principal component analysis, or independent component analysis, and estimate brain activity of the subject based on the plurality of components; andcontrol the display device to present to the subject useful information related to an emotional or physical state based on the estimated brain activity of the subjectat least a portion the facial skin temperature data and/or the facial blood circulation volume data are obtained from a region around paranasal sinuses and/or a forehead of the subject. 2. The useful information presentation device according to claim 1, wherein the processor is further configured to extract, from the plurality of components, a component having a waveform with an amplitude that has correlation with changes of a brain resting time and a brain activated time as a determination component, andestimate the brain activity of the subject based on the determination component. 3. The useful information presentation device according to claim 1, wherein the at least one of the plurality of temperature sensors and the imaging device includes the imaging device, andthe imaging device and the display device are housed in a smart device. 4. The useful information presentation device according to claim 3, wherein the smart device includes a mirror member that mirrorizes the display device,the at least one of the plurality of temperature sensors and the imaging device includes the imaging device, andthe imaging device is disposed behind the display device. 5. The useful information presentation device according to claim 1, wherein the at least one of the plurality of temperature sensors and the imaging device includes the imaging device, andthe imaging device, the processor, and the display device are housed in... | Daikin Industries, Ltd. Tokyo Institute of Technology | A61B005/01 A61B005/00 A61B005/026 A61B005/11 A61B005/16
US10702182 | B2 | 2020-07-07 | Myoelectric sensor | 2016-01-12 | A myoelectric sensor for detecting myoelectric signals which accompany body movement includes: a wearing band that is elastic, expandable, and circular, and that is worn around a limb to surround the limb tightly; myoelectric detection units a plurality of which are disposed in the circumferential direction on the wearing band with intervals therebetween so as to cause each of a plurality of myoelectric detection electrodes to be in close contact with the surface of the limb, and which detect myoelectric signals from corresponding positions on the limb using the myoelectric detection electrodes; and connection cables that electrically connect mutually adjacent myoelectric detection units and thereby transmit the myoelectric signals. The connection cables each include a bent portion, the bent shape of which changes in response to changes in the distance between the mutually adjacent myoelectric detection units. | 1. A myoelectric sensor for detecting myoelectric signals caused by a body movement comprising: an elastically expandable and contractible annular wearing band adapted to be worn on a limb and tightly surrounding therearound;a plurality of myoelectric detection units arranged in a circumferential direction of said wearing band apart from each other and each having a plurality of myoelectric detection electrodes which adapts to contact with a surface of the limb and to detect myoelectric signals from corresponding parts of the limb; andconnecting cables which electrically connect said myoelectric detection units adjacent to each other and transmit the myoelectric signals therebetween,wherein each of said connecting cables has at least one bent portion whose bent shape changes in accordance with a change in distance between the myoelectric detection units adjacent to each other,wherein each of said plurality of myoelectric detection units comprises a central part and a terminal part each having at least one of the myoelectric detection electrodes, the central part of at least one of the plurality of myoelectric detection units mounting a circuit board with at least an amplifying circuit, and the terminal part being engaged with said central part movably forward and backward while being prevented from twisting with respect to said central part. 2. The myoelectric sensor according to claim 1, wherein at least one of the connecting cables has an S-shape as a whole. 3. The myoelectric sensor according to claim 1, wherein the circuit board has at least an analog-to-digital converter circuit for converting analog myoelectric signals detected by said myoelectric detection electrodes of said plurality of myoelectric detection units to digital signals and outputting them. 4. The myoelectric sensor according to claim 1, wherein the circuit board has at least a USB controller circuit for outputting digital myoelectric signals converted with an analog-to-digital converter circuit, via a USB line. | NIPRO CORPORATION HARADA ELECTRONICS INDUSTRY CO., LTD. | A61B005/0492 G06F003/01 A61B005/00
US10727956 | B2 | 2020-07-28 | Wireless sensors in medical environments | 2018-09-04 | Monitoring a patient can include a vital sign device including a skin contact and a demodulator circuit in communication with the electrically conductive skin contact, the demodulator circuit including: a physiological waveform data processing module configured to process the waveform data received from the electrically conductive skin contact; and a digitally encoded data processing module configured to detect and decode digitally encoded data modulated at the carrier frequency. Also included can be a signal conductive blanket including an extended touch point. A clinician contacts the extended touch point of the signal conductive blanket and the patient monitoring device to connect the vital sign device and the patient monitoring device. | 1. A method for establishing a connection between a wireless vital sign device and a patient monitoring device, the method comprising: attaching the wireless vital sign device to a patient;providing a signal conductive blanket, at least some portion of the signal conductive blanket touching the patient;contacting a portion of the signal conductive blanket with a first clinician body part;contacting a portion of the patient monitoring device with a second clinician body part, the patient monitoring device being detached from the patient; andconfirming, on the patient monitoring device, connection between the vital sign device and the patient monitoring device. 2. The method of claim 1, further comprising contacting an extended touch point on the signal conductive blanket with the first clinician body part. 3. The method of claim 1, wherein the first clinician body part is a first hand of a clinician and the second clinician body part is a second hand of the clinician, the method further comprising: allowing the first hand of the clinician to touch an extended touch point on the signal conductive blanket; andallowing the second hand of the clinician to touch the portion of the patient monitoring device. 4. The method of claim 1, further comprising forming multiple extended touch points on the signal conductive blanket. 5. The method of claim 4, further comprising forming multiple touch pads on the signal conductive blanket. 6. The method of claim 5, further comprising forming the vital sign device to include an electrically conductive skin contact configured to obtain waveform data including waveform data modulated at a carrier frequency, and a demodulator circuit in communication with the electrically conductive skin contact. 7. The method of claim 1, further comprising contacting an extended touch point on the signal conductive blanket with the first clinician body part, the extended touch point located away from the patient, and contacting an electrically conductive skin contact on the patient monitoring device with the second clinician body part. | HILL-ROM SERVICES, INC. | H04B013/00 A61B005/00 A61B005/0205 G08B021/24 G08B025/10 A61B005/08 A61B090/53 A61B005/024
US10709386 | B2 | 2020-07-14 | Electrocardiogram waveform identification and diagnostics via electrophysiological sensor system fusion | 2017-12-12 | An electroencephalography (EEG) sensor is installed on a seat and configured to monitor brain activity of a user to generate a EEG signal. An electrocardiogram (ECG) sensor is installed on the seat and configured to monitor heart activity of the user to generate a ECG signal. A processor of the seat is in communication with the EEG sensor and the ECG sensor, and is configured to process the EEG signal and the ECG signal to determine a stress level of the user. | 1. A vehicle system, comprising: an electroencephalography (EEG) sensor, installed on a seat and configured to monitor brain activity of a user located in the seat to generate a EEG signal;an electrocardiogram (ECG) sensor, installed on the seat and configured to monitor heart activity of the user to generate an ECG signal;a processor, in communication with the EEG sensor and the ECG sensor, configured to process vital data, the vital data including the EEG signal and the ECG signal, to determine a stress level of the user;responsive to the stress level being above a predefined threshold, output a countermeasure to the user;responsive to the stress level persisting for a predefined period of time, upload a first portion of the vital data indicative of the stress to a server for analysis in an anonymous manner;responsive to receiving a request for additional vital data from the server, output a message via a human-machine interface (HMI) to request user permission to grant access to additional vital data; andresponsive to receiving the user permission via the HMI, upload a second portion of the vital data corresponding to the stress to the server in a non-anonymous manner. 2. The vehicle system of claim 1, wherein the countermeasure includes at least one of to: play a previously-recorded audio file intended to elicit a calming effect, invite the user to slow down or stop the vehicle, or send a message to a third party. 3. The vehicle system of claim 1, further comprising a breathing sensor, installed on the seat and configured to monitor breathing activity of the user to generate a breathing signal. 4. The vehicle system of claim 3, wherein the breathing sensor is combined with the ECG sensor as a single unit. 5. The vehicle system of claim 3, wherein the processor is further configured to analyze a cardiorespiratory synchronization using the ECG signal and the breathing signal. 6. The vehicle system of claim 1, wherein the processor is further configured to communicate with a telematics controller to upload one or more of the EEG signal and the ECG signal to a server. 7. The vehicle system of claim 6, wherein the telematics controller is further configured to receive a diagnostic result from a server. 8. The vehicle system of claim 6, wherein the telematics controller is further configured to receive instruction to perform the countermeasure from the server, and the processor is configured to execute the countermeasure responsive to receipt of the instruction. 9. The vehicle system of claim 6, wherein the processor is further configured to communicate with an infotainment system, the telematics controller is further configured to receive a request from the server for additional EEG and ECG signals, and the infotainment system is configured to prompt the user to accept or deny the request. 10. The vehicle system of claim 1, further comprising a storage, installed to the seat and configured to record the EEG signal and the ECG signal. 11.... | Lear Corporation | A61B005/00 A61B005/0452 A61B005/0205 A61B005/18 A61B005/04 A61B005/0476 A61B005/08
US10709339 | B1 | 2020-07-14 | Biometric wearable for continuous heart rate and blood pressure monitoring | 2018-07-03 | A wearable biometric sensing ring apparatus for continuous heart rate and blood pressure monitoring having a ring housing for retention on a finger of a user, a photoplethysmograph (PPG) sensor disposed within the housing on an interior surface of the ring to contact or be placed adjacent the finger to register a pulse rate of the user, a first electrode and second electrode disposed within the housing on said interior surface of the ring to contact the finger on a first side of the user's body, and a third electrode disposed on an exterior surface of the ring housing to contact a contralateral portion of the user's body. A combination of first, second and third electrodes are configured for obtaining electrocardiogram (ECG) measurements. A controller and programming are provided for receiving analog pulse rate data from the PPG sensor and analog ECG data from the first, second and third electrodes, converting the analog pulse rate data and ECG data into digital data, calculating one or more of heart rate (HR), heart rate variability (HRV) and blood pressure (BP) from a combination of the digital ECG and PPG sensor data, and graphically outputting one or more of the calculated HR, HRV and BP. | 1. A wearable biometric sensing ring apparatus for continuous heart rate and blood pressure monitoring, comprising: (a) a ring-shaped housing having a finger aperture with a flexible circuit board having a first surface configured for retention on a skin surface on any finger of a user's hand on a first side of a user's body;(b) a photoplethysmograph (PPG) sensor disposed on said first surface of said flexible circuit board and configured to contact or be placed adjacent the skin surface to register a pulse rate of the user;(c) a rigid sensor surface attached along said first surface of said flexible circuit board upon which are mounted a first electrode and second electrode which are configured to contact the skin of a finger of the user while minimizing motion artifacts;(d) a third electrode disposed on an exterior surface of said housing and having a free sensing surface configured for establishing direct contact with a contralateral portion of the user's body, as a second side of the user's body, wherein said combination of first, second and third electrodes are configured for obtaining electrocardiogram (ECG) measurements;(e) a processor disposed on said flexible circuit board, said processor configured to receive data from the PPG sensor and first, second and third electrodes; and(f) a non-transitory memory storing instructions executable by the processor;(g) wherein said instructions, when executed by the processor, perform steps comprising: (i) receiving analog pulse rate data from the PPG sensor and analog ECG data from the first, second and third electrodes;(ii) converting the analog pulse rate data and ECG data into digital data;(iii) calculating blood pressure (BP) from a combination of the digital ECG and PPG sensor data, wherein BP is calculated by determining a time difference between an ECG wave peak and a PPG wave peak extracted from the digital ECG and PPG sensor data to generate a pulse transit time (PTT), and calibrating a set of PTT measurements against standard BP measurements to determine a user's variance from a standard measure as a function of one or more of the user's due height and weight; and(iv) outputting BP;(h) wherein said first and second electrodes are disposed on said first surface which is a bottom interior surface of the aperture of the ring housing to contact skin on a palmar side of the user's finger upon which the ring-shaped housing is retained; and(i) wherein said third electrode is located on an outer surface of the ring housing at a location to allow for contact with a portion of the user separate from the hand upon which the housing is retained. 2. The apparatus of claim 1, further comprising calculating one or more of heart rate (HR) and heart rate variability (HRV) in addition to BP, and outputting one or more of the HR and HRV along with the BP. 3. The apparatus of claim 2, further comprising: (k) a communication circuit configured for wirelessly communicating between said wearable biometric sens... | Senstream, Inc. | A61B005/00 A61B005/0205 A61B005/0408 A61B005/0245 A61B005/024 A61B005/021 A61B005/053
US10702163 | B2 | 2020-07-07 | Systems and methods for intracavitary temperature measurement and monitoring | 2017-05-01 | Systems and methods for measuring and monitoring intracavitary tissue temperature. The system may include a catheter shaft with a circuit board disposed therein, the circuit board having an array of sensors disposed thereon. The catheter shaft may have an opening and an expandable structure surrounding the opening to provide a field of view of the intracavitary tissue for the array of sensors through the opening. The system may include a software-based programming system run on a computer such that a clinician may review information indicative of temperature of the intracavitary tissue, and be alerted if the temperature exceeds a predetermined threshold. | 1. A system for intracavitary tissue measurement and monitoring, the system comprising: an introducer device configured to be positioned adjacent to an intracavitary tissue, the introducer device comprising: a catheter shaft having a distal end, a longitudinal axis, and a lumen extending therethrough, the catheter shaft having an opening at the distal end along the longitudinal axis such that at least a portion of the lumen is exposed;a circuit board having an array of infrared sensors disposed thereon, the circuit board disposed within the lumen of the catheter shaft at least partially in the opening at the distal end, the array of infrared sensors each having circuitry configured to generate a signal indicative of temperature of the intracavitary tissue; andan expandable structure formed from an infrared transmissive material and disposed along the longitudinal axis on the distal end of the catheter shaft to surround the array of infrared sensors and to provide a field of view through the opening such that the opening is in fluid communication with an interior of the expandable structure; anda non-transitory computer readable media having instructions stored thereon, wherein the instructions, when executed by a processor operatively coupled to the circuit board, cause a graphical user interface to display information indicative of temperature of the intracavitary tissue based on the signal from the array of infrared sensors. 2. The system of claim 1, wherein the introducer device is configured to be positioned adjacent esophageal tissue. 3. The system of claim 1, wherein the catheter shaft comprises one or more support members, each of the one or more support members having end portions coupled to the catheter shaft and a middle portion configured to be positioned parallel to the longitudinal axis of the catheter shaft in a delivery state, and curved outwardly away from the catheter shaft to engage the intracavitary tissue in a deployed state. 4. The system of claim 1, wherein the circuit board is flexible. 5. The system of claim 1, wherein the circuit board is further configured to be fixed within the lumen of the catheter shaft. 6. The system of claim 1, wherein the array of infrared sensors comprises at least one of infrared sensitive photodiodes, infrared sensitive transistors, infrared sensitive photocells, or infrared sensitive thermopiles. 7. The system of claim 1, wherein the circuitry is configured to generate the signal indicative of temperature of the intracavitary tissue by amplifying the signal, filtering the signal, performing compensation for temperature of the array of infrared sensors, and converting the signal to a digital serial stream. 8. The system of claim 1, wherein the circuit board is configured to be slidably inserted within the lumen of the catheter shaft, the circuit board further configured to be rotated within the lumen of the catheter shaft to enhance the field of view. 9. The system of claim 1, wherein th... | Precision EP GmbH | A61B005/00 A61B005/01
US10719933 | B2 | 2020-07-21 | OCT angiography calculation with optimized signal processing | 2018-04-27 | Methods and systems for angiographic imaging with optical coherence tomography (OCT) are described using ratio-based and angiographic deviation based calculations. In using these calculations to determine motion, arbitrary interframe permutations may be used, post-calculated, non-linear results for projection visualization may be averaged, poor matches may be eliminated on an A-line by A-line basis, windowing functions may be used to improve results, partial spectrums may be used when capturing data, and a minimum intensity threshold may be used for determining which pixels to use. | 1. An angiographic optical coherence tomography (OCT) method comprising: calculating values of OCT images generated from captured OCT data of a volume, the values representing dissimilarities between at least two of the OCT images;generating angiographic images based on the values; anddisplaying, rendering, and/or storing the OCT images and/or the angiographic images,wherein at least one of the values achieves a greater sensitivity to the dissimilarities between OCT images than an amplitude decorrelation method, andwherein a contrast of each pixel of one of the angiographic images corresponds to a different one of the values. 2. The angiographic OCT method of claim 1, wherein the values are calculated according to a function that receives, as a variable input, a value corresponding to a ratio between two pixel intensities of the at least two OCT images. 3. The angiographic OCT method of claim 2, wherein the value corresponding to the ratio is a difference between log-scale intensity information between two OCT images. 4. The angiographic OCT method of claim 1, wherein the values are represented by a ratio between two pixel intensities. 5. The angiographic OCT method according to claim 1, wherein the values are log-normal deviation values or geometric standard deviation values. 6. The angiographic OCT method of claim 1, wherein the values are calculated by dividing first pixel intensity values of an OCT image obtained at a first time by corresponding second pixel intensity values of an OCT image obtained at a second time. 7. The angiographic OCT method according to claim 1, further comprising: averaging the values for an X-Y position of the at least two OCT images. 8. The angiographic OCT method according to claim 7, further comprising: comparing the averaged values to a criterion; andexcluding values that do not meet the criterion. 9. The angiographic OCT method of claim 8, wherein the averaging, comparing, and excluding are performed A-line by A-line. 10. The angiographic OCT method of claim 1, further comprising filtering the captured OCT data before generating the OCT images, before generating the angiographic images, and/or before displaying the OCT images and/or angiographic images. 11. The method according to claim 1, wherein a partial spectrum of an OCT light source is used to generate at least one of the OCT images. 12. The method according to claim 1, wherein the OCT images are generated by applying a function to interferograms, or from interferograms having envelopes, and wherein an equivalent noise bandwidth (ENBW) for generating structural OCT images is greater than the ENBW for generating angiographic OCT images. 13. The method according to claim 12, wherein the envelopes are of interferograms immediately before a discrete Fourier transform is applied. 14. The method according to claim 1, wherein the values are one-sided ratio-based values. 15. An angiographic optical coherence tomography (OCT) method comprising: calcul... | Kabushiki Kaisha TOPCON | G06T007/00 G06T005/50 A61B005/00 A61B005/02 G06K009/52 G06K009/62
US10702196 | B2 | 2020-07-07 | Pulse photometer | 2017-11-21 | A pulse photometer includes a light emitter, a light detector, a first signal separator that separates a first signal component and a first noise component, a second signal separator that separates a second signal component and a second noise component, a first determining section, a second determining section, and a concentration calculating section that calculates a concentration of a light absorption material in blood of the subject based on a fundamental frequency of the first signal component or a fundamental frequency of the second signal component. | 1. A pulse photometer comprising: a memory storing instructions;a processor configured to execute the instructions;an emitter that emits first light beam of a first wavelength, a second light beam of a second wavelength, a third light beam of a third wavelength, and a fourth light beam of a fourth wavelength; anda detector that outputs a first intensity signal in accordance with a first intensity of the first light beam that transmitted through or reflected from a living tissue of a subject, a second intensity signal in accordance with a second intensity of the second light beam that is transmitted through or reflected from the living tissue of the subject, a third intensity signal in accordance with a third intensity of the third light beam that is transmitted through or reflected from the living tissue of the subject, and a fourth intensity signal in accordance with a fourth intensity of the fourth light beam that is transmitted through or reflected from the living tissue of the subject,wherein the processor executing the instructions stored in the memory is configured to implement functions of: a first signal separator that separates a first signal component of a first observed signal pair made of two of the first to fourth intensity signals from a first noise component of the first observed signal pair;a second signal separator that separates a second signal component of a second observed signal pair made of two of the first to fourth intensity signals excluding a combination of the first to fourth intensity signals in the first observed signal pair from a second noise component of the second observed signal pair;a third signal separator that that separates a third signal component of a third observed signal pair made of two of the first to fourth intensity signals excluding combinations of the first to fourth intensity signals in the first and second observed signal pairs from a third noise component of the third observed signal pair;a first determining section that determines whether a first difference between a first signal absorbance ratio of the first signal component and a first noise absorbance ratio of the first noise component is equal to or larger than a first predetermined value;a second determining section that determines whether a second difference between a second signal absorbance ratio of the second signal component and a second noise absorbance ratio of the second noise component is equal to or larger than a second predetermined value;a third determining section that, when it is determined that the first difference is smaller than the first predetermined value and the second difference smaller than the second predetermined value, determines whether a third difference between a third signal absorbance ratio of the third signal component and a third noise absorbance ratio of the third noise component is equal to or larger than a third predetermined value; anda concentration calculating section that, calculates a concentration o... | NIHON KOHDEN CORPORATION | A61B005/1455 A61B005/024 A61B005/00 G01N021/31 G01N033/49
US10716466 | B2 | 2020-07-21 | Device for in-vivo determination of eye moisture | 2017-10-29 | A medical diagnostic device and method, the device including: (a) an alternating current source, adapted to produce an alternating current; (b) an electrode arrangement having at least first and second electrodes, separated by an electrically insulating region, the arrangement having an at least semi-rigid region that fixes the electrodes in a spaced-apart manner, the arrangement adapted to contact the soft tissue on the inner surface of the eyelid; and (c) a processor, associated with the electrode arrangement, said electrodes electrically connected to the alternating current source; wherein, when the electrode arrangement is provided with the alternating current, and is disposed against the soft tissue, the soft tissue electrically bridges between the electrodes to form an electrical circuit, wherein an electrical signal is produced by the alternating current electrically passing between the electrodes via the soft tissue; wherein the processor is adapted to receive in-vivo based electrical information originating from the electrical signal, via the circuit, and to produce an output relating to, or derived from, the moisture parameter, based on the in-vivo electrical information; and wherein the processor is designed and configured to compute the moisture parameter in the soft tissue, at least partially based on the in-vivo electrical information, and based on an empirical correlation between the in-vivo electrical information and the moisture parameter. | 1. A device for evaluation of a moisture characterization parameter associated with moisture of soft tissue on an inner surface of an eyelid of a subject, the device comprising: (a) an alternating current source, adapted to connect to a power supply and to produce an alternating current;(b) an electrode arrangement having at least a first electrode and a second electrode, said first electrode electrically separated from said second electrode by an electrically insulating region, said electrode arrangement having an at least semi-rigid region fixing said first electrode and said second electrode in a spaced-apart manner,said electrode arrangement adapted to contact the soft tissue on the inner surface of the eyelid, said electrodes and said electrically insulating region composed of biocompatible materials; and(c) a processor, associated with said electrode arrangement,said first and second electrodes being electrically connected to said alternating current source,wherein, when said electrode arrangement is provided with said alternating current, and is disposed against the soft tissue, the soft tissue electrically bridges between said electrodes to form an electrical circuit, such that an electrical signal is produced by said alternating current passing from said first electrode to said second electrode via the soft tissue,said processor being adapted to receive in-vivo based electrical information originating from said electrical signal, via said electrical circuit, and to produce an output relating to, or derived from, the moisture characterization parameter associated with the moisture of the soft tissue on the inner surface of the eyelid, based on said in-vivo electrical information,said processor being designed and configured to compute the moisture characterization parameter associated with the moisture of the soft tissue on the inner surface of the eyelid, at least partially based on said in-vivo electrical information, and based on an empirical correlation between said in-vivo electrical information and said moisture characterization parameter associated with the moisture of the soft tissue on the inner surface of the eyelid,said empirical correlation including an inverse relationship between an electrical impedance derived from said electrical signal or from said in-vivo electrical information, and said moisture characterization parameter associated with the moisture of the soft tissue on the inner surface of the eyelid, such that an increasing level of moisture of soft tissue on the inner surface of the eyelid is correlated, by said processor, with a decreasing of said electrical impedance,wherein said moisture characterization parameter associated with the moisture of the soft tissue on the inner surface of the eyelid includes an eye-moisture characterization parameter selected from the group of parameters consisting of a calculated in-vitro osmolarity, a calculated Schirmer's Test absorption length, a calculated Meibomian Grading Scor... | 8679410 CANADA INC. | A61B003/10 A61B005/053
US10715021 | B2 | 2020-07-14 | Mobile capsule device and control method thereof | 2013-07-18 | A mobile capsule device 10 comprises a long capsule body 11 having a permanent magnet 13 movable in the lengthwise direction with respect to the long capsule body and a coil for driving the permanent magnet 13, while a propulsion force is generated entirely by applying an alternate current to the coil and performing back and forth movements of the permanent magnet 13. The coil has first and second coil parts 15 and 16 arranged circumferentially in front and back of the permanent magnet 13, and a frequency of an alternate current applied to the first and second coil parts 15 and 16 is made to accord with a resonance frequency of the capsule device 10 generated by a back and forth vibration of the permanent magnet 13. Thereby, a self-propelled, mobile capsule device 10 which is downsized, compact and efficient and a control method thereof can be provided. | 1. A mobile capsule device, comprising: a long capsule body provided with a permanent magnet and a coil, the permanent magnet being movable in the lengthwise direction with respect to the capsule body, the coil driving the permanent magnet, the coil being applied an alternate current through an alternate current generating means to perform a back and forth vibration of the permanent magnet and generate a propulsion force,wherein the coil has first and second coil parts arranged in a forward and a backward motion of the permanent magnet so as to surround the permanent magnet, and a frequency of an alternate current applied to the first and second coil parts corresponds to a resonance frequency of the capsule device generated by a back and forth vibration of the permanent magnet,the first and second coil parts are wound around a cylinder-shaped body having a slight interspace between the cylinder-shaped body and the permanent magnet to enable movement of the permanent magnet inside the cylinder-shaped body,the cylinder-shaped body wound by the first and second coil parts is inside the capsule body and the capsule device,stoppers are arranged on both sides of the cylinder-shaped body,an alternate current generated by the alternate current generating means is composed of a positive/negative symmetrical alternate current further added by a direct current,an excitation force generated by the first and second coil parts and the permanent magnet is intensified in one direction,the moving direction of the capsule device is determined based on a polar character of the direct current,by an electromagnetic force generated by the first and second coil parts, the permanent magnet performs a back and forth vibration while crashing into the stoppers, andthe resonance frequency is a frequency at which an electric current value becomes minimum when an alternate current is actually applied to the first and second coil parts. 2. The mobile capsule device as set forth in claim 1, wherein the capsule body is provided with a radio which liaises with an external control device. 3. The mobile capsule device as set forth in claim 2, wherein the coil is used as an antenna for the radio. 4. The mobile capsule device as set forth in claim 1, wherein the capsule body is provided with one or more of the following: an illuminated microcamera, instrument for procedures, dosing feature and position sensor. 5. A method for controlling a mobile capsule device having a long capsule body provided with a permanent magnet and a coil, the permanent magnet being movable in the lengthwise direction with respect to the capsule body, the coil driving the permanent magnet, the coil being applied an alternate current to perform a back and forth vibration of the permanent magnet and generate a propulsion force, the method comprising: determining the moving direction of the capsule body by making a frequency of an alternate current applied to the coil accord with a resonance frequency of a ... | KYUSHU INSTITUTE OF TECHNOLOGY | H02K033/14 A61B005/06 A61B001/00 A61B001/04 H02K033/16 A61B010/02 A61M031/00 A61B005/07
US10713791 | B2 | 2020-07-14 | Computational simulations of anatomical structures and body surface electrode positioning | 2020-02-07 | A method may include identifying a simulated three-dimensional representation corresponding to an internal anatomy of a subject based on a match between a computed two-dimensional image corresponding to the simulated three-dimensional representation and a two-dimensional image depicting the internal anatomy of the subject. Simulations of the electrical activities measured by a recording device with standard lead placement and nonstandard lead placement may be computed based on the simulated three-dimensional representation. A clinical electrogram and/or a clinical vectorgram for the subject may be corrected based on a difference between the simulations of electrical activities to account for deviations arising from patient-specific lead placement as well as variations in subject anatomy and pathophysiology. | 1. A system, comprising: at least one processor; andat least one memory including program code which when executed by the at least one processor provides operations comprising: identifying a three-dimensional representation of at least a portion of an anatomy of a subject including a target organ;identifying a non-standard lead placement of one or more electrogram leads on a body of the subject;generating, based at least on the three-dimensional representation, one or more simulated electrical activations of the target organ;generating, based at least on the one or more simulated electrical activations, a non-standard electrogram associated with the non-standard lead placement of the one or more electrogram leads on the body of the subject;generating, based at least on the one or more simulated electrical activations, a standard electrogram associated with a standard lead placement of the one or more electrogram leads on the body of the subject; andcorrecting, based at least on a difference between the nonstandard electrogram and the standard electrogram, an actual electrogram generated for the subject using the non-standard lead placement. 2. The system of claim 1, wherein the standard electrogram, the nonstandard electrogram, and the actual electrogram comprise electrocardiograms, electroencephalograms, or vectorcardiograms. 3. The system of claim 1, wherein the correcting includes generating a transformation matrix to transform the nonstandard electrogram to the standard electrogram and applying the transformation matrix to the actual electrogram. 4. The system of claim 1, wherein the identifying of the three-dimensional representation includes comparing a two-dimensional image of the portion of the anatomy of the subject to one or more two-dimensional images included in a library mapping the one or more two-dimensional images to one or more corresponding three-dimensional representations. 5. The system of claim 1, wherein the nonstandard lead placement is identified based at least on an analysis of a two-dimensional image of the portion of the anatomy. 6. The system of claim 1, wherein the operations further comprise: identifying a simulated electrogram matching the corrected electrogram by at least searching a library including a plurality of simulated electrograms, the library mapping the plurality of simulated electrograms to one or more characteristics of the target organ used to generate the plurality of simulated electrograms. 7. A computer implemented method, comprising: identifying a three-dimensional representation of at least a portion of an anatomy of a subject including a target organ;identifying a non-standard lead placement of one or more electrogram leads on a body of the subject;generating, based at least on the three-dimensional representation, one or more simulated electrical activations of the target organ;generating, based at least on the one or more simulated electrical activations, a non-standard electrogram associ... | The Regents of the University of California Vektor Medical, Inc. | G06T007/00 G16H030/40 G16H050/20 A61B005/055 A61B005/0402
US10716477 | B2 | 2020-07-21 | Medical device including manipulable portion with connected elongate members | 2019-01-28 | A medical system may include a shaft member and a structure physically coupled to the shaft member. A portion of the shaft member may be sized to be delivered through a bodily opening leading to a bodily cavity. The structure may include at least two flexible couplings, each flexible coupling extending transversely from an intermediate portion of a respective one of at least two elongate members of the structure. The flexible coupling extending transversely from the intermediate portion of a first one of the at least two elongate members forms at least a part of a first closed loop arranged to receive a portion of the flexible coupling of a second one of the at least two elongate members therethrough to limit a spacing between the intermediate portions of the first and the second ones of the at least two elongate members. | 1. A medical system comprising: a shaft member including a portion sized to be delivered through a bodily opening leading to a bodily cavity; anda structure physically coupled to the shaft member, the structure including a plurality of elongate members, the structure selectively moveable between a delivery configuration in which the structure is suitably sized to be deliverable through the bodily opening to the bodily cavity, and an expanded configuration in which the structure has a size too large to be deliverable through the bodily opening to the bodily cavity,wherein each of at least one elongate member of the plurality of elongate members comprises a respective plurality of material layers,wherein a first subset of the respective plurality of material layers of the one elongate member contacts a second subset of the respective plurality of material layers of the one elongate member, an external surface of the first subset of the respective plurality of material layers of the one elongate member forming part of a first surface of the one elongate member, an external surface of the second subset of the respective plurality of material layers of the one elongate member forming at least part of a second surface of the one elongate member, and the first surface of the one elongate member and the second surface of the one elongate member forming opposing surfaces on opposite sides of the one elongate member,wherein the first subset of the respective plurality of material layers of the one elongate member includes a notch that exposes a surface of the second subset of the respective plurality of material layers of the one elongate member, the surface of the second subset of the respective plurality of material layers of the one elongate member that is exposed by the notch forming a portion of the first surface of the one elongate member, andwherein an aperture is formed at least in part in the surface of the second subset of the respective plurality of material layers of the one elongate member. 2. The medical system of claim 1, wherein each elongate member of the plurality of elongate members comprises a proximal end, a distal end, and a respective intermediate portion positioned between the proximal end and the distal end, andwherein the notch is located within the respective intermediate portion of the one elongate member. 3. The medical system of claim 1, wherein the notch extends from an edge of the one elongate member, the edge between the first and second surfaces of the one elongate member. 4. The medical system of claim 1, wherein the notch has a trapezoidal shape. 5. The medical system of claim 1, wherein the surface of the second subset of the respective plurality of material layers of the one elongate member that is exposed by the notch surrounds the aperture. 6. The medical system of claim 1, wherein the first subset of the respective plurality of material layers of the one elongate member comprises an electrically conductive laye... | KARDIUM INC. | A61B018/18 A61B005/04 A61B005/01 A61B018/14 A61N001/05 A61B005/00 A61B005/042 A61B018/00
US10721574 | B2 | 2020-07-21 | Fitting unilateral electric acoustic stimulation for binaural hearing | 2015-08-03 | A system and method of fitting an Electric Acoustic Stimulation (EAS) system of a patient that has binaural hearing is provided. The patient has an ipsilateral ear and a contralateral ear opposite the ipsilateral ear, with the EAS system associated with the ipsilateral ear of the patient. Stimulation parameters of the EAS system are developed unilaterally while taking into account hearing abilities of the contralateral ear. The EAS system is configured based on the developed stimulation parameters. | 1. A method of fitting an Electric Acoustic Stimulation (EAS) system of a patient that has binaural hearing, the patient having an ipsilateral ear and a contralateral ear opposite the ipsilateral ear, the EAS system configured to be used on the ipsilateral ear of the patient whose contralateral ear is not associated with a hearing device, the method comprising: obtaining an unaided audiogram of the contralateral ear;determining hearing thresholds of the contralateral ear based on the unaided audiogram;developing stimulation parameters of the EAS system unilaterally to obtain equal loudness for both ears based, at least in part, on the hearing thresholds of the contralateral ear; andconfiguring the EAS system based on the developed stimulation parameters. 2. The method according to claim 1, wherein developing stimulation parameters of the EAS system includes developing stimulation parameters pertaining to acoustical and/or electrical coverage of the EAS system. 3. The method according to claim 2, further comprising: determining a frequency range covered by acoustical residual hearing in the ipsilateral ear that is associated with the EAS system; anddetermining a frequency range covered by electrical stimulation associated with the EAS system. 4. The method according to claim 3, wherein if there is a frequency gap that is not covered by either the acoustical residual hearing or the electrical stimulation of the EAS system, the method further includes developing stimulation parameters such that the frequency gap is covered by electrical stimulation. 5. The method according to claim 3, wherein if there is overlap in frequency coverage of the acoustical residual hearing and the electrical stimulation of the EAS system, the method further includes developing stimulation parameters to decrease electrical stimulation by at least one apical electrode of the EAS system, so as to allow only acoustical stimulation in regions of acoustical residual hearing. 6. The method according to claim 1, wherein developing stimulation parameters of the EAS system unilaterally includes providing binaural sound processing of the patient. 7. The method according to claim 6, wherein providing binaural sound processing of the patient includes optimizing interaural level differences (ILD). 8. The method according to claim 6, wherein providing binaural sound processing of the patient includes optimizing interaural time delays (ITD). 9. A system for modifying stimulation parameters of an Electric Acoustic Stimulation (EAS) system of a patient that has binaural hearing, the patient having an ipsilateral ear and a contralateral ear opposite the ipsilateral ear, the EAS system configured to be used on the ipsilateral ear of the patient whose contralateral ear is not associated with a hearing device, the system comprising: a controller configured to communicate with the EAS system, the controller further configured to: determine hearing thresholds of the contralateral ear ba... | MED-EL Elektromedizinische Geraete GmbH | A61N001/05 H04R025/00 A61B005/00 A61N001/36 A61B005/12
US10709373 | B2 | 2020-07-14 | Fluid analysis device and associated systems and methods | 2015-03-30 | A fluid analysis device including a fluid receipt chamber, a first capacitive element for measuring fluid flow into the fluid receipt chamber and a controller operatively coupled to the first capacitive element, wherein the controller is configured to measure a first capacitance of the first capacitive element. | 1. A fluid analysis device, comprising: a fluid receipt chamber having a first mechanical engagement feature and configured to receive fluid;a first capacitive element for measuring fluid flow of the fluid into the fluid receipt chamber; anda housing having a second mechanical engagement feature configured to engage with the first mechanical engagement feature of the fluid receipt chamber, the housing configured to directly contact the fluid receipt chamber when the first mechanical engagement feature and the second mechanical engagement feature are engaged, the housing including: a controller electrically coupled to the first capacitive element when the first mechanical engagement feature and the second mechanical engagement feature are engaged, wherein the controller is configured to: take a first measurement of a first capacitance of the first capacitive element based on signals received from the first capacitive element, wherein the first measurement of the first capacitance is indicative of a first volume of the fluid at a first point in time; andtake a second measurement of the first capacitance of the first capacitive element based on signals received from the first capacitive element after waiting a sampling time, wherein the second measurement of the first capacitance is indicative of a second volume of the fluid at a second point in time, wherein a flow rate of the fluid is determinable by (1) subtracting the first volume from the second volume to determine a difference in volume and (2) dividing the difference in volume by the sampling time. 2. The fluid analysis device of claim 1, wherein the first capacitive element comprises a pair of plates, each plate having a surface, wherein the surfaces of the plates are facing one another with a first gap therebetween, wherein the first capacitance is measured across the first gap. 3. The fluid analysis device of claim 2, wherein the fluid receipt chamber comprises a double-walled chamber having an inner wall and an outer wall, wherein one of the pair of first capacitive element plates is associated with the inner wall and the other of the pair of first capacitive element plates is associated with the outer wall. 4. The fluid analysis device of claim 3, wherein the inner wall defines an inner fluid compartment, wherein the outer wall surrounds the inner wall with a fluid space therebetween, and wherein the inner fluid compartment and fluid space are fluidly connected to one another by at least one fluid passage. 5. The fluid analysis device of claim 4, wherein fluid movement between the inner fluid compartment and the fluid space is dampened by the fluid space. 6. The fluid analysis device of claim 4, further comprising a second capacitive element for measuring a second capacitance, the second capacitive element being disposed within the fluid receipt chamber such that, upon entry of fluid into the fluid receipt chamber, the second capacitive element is completely covered by fluid before ... | University of Utah Research Foundation | A61B005/20 B01L003/00 G01F001/56 G01N027/22 G01N033/493 A61B010/00 A61B005/145 A61B005/00
US10709424 | B2 | 2020-07-14 | Method and system for cuff-less blood pressure (BP) measurement of a subject | 2015-06-18 | Embodiments herein disclose a method and system for cuff-less blood pressure (BP) measurement of a subject. The method includes measuring, by one or more sensors, a local pulse wave velocity (PWV) and/or blood pulse waveforms of an arterial wall of the subject. Further, the method includes measuring, by an ultrasound transducer, a change in arterial dimensions over a cardiac cycle of the arterial wall of the subject. The arterial dimensions include an arterial distension and an end-diastolic diameter. Furthermore, the method includes measuring, by a controller unit, BP of the subject based on the local PWV and the change in arterial dimensions. | 1. A system for cuff-less calibration free measurement of blood pressure (BP) of a subject, the system comprising: a probe comprising: a first sensor;a second sensor spaced at a distance from said first sensor; anda third sensor;an ultrasound transducer connected with said first sensor and said second sensor, wherein the ultrasound transducer is placed at a center of the third sensor in said probe and in between said first sensor and said second sensor; anda controller unit connected to said probe;wherein said first sensor and second sensor are configured to measure a local pulse wave velocity (PWV) of an arterial wall of said subject;wherein said ultrasound transducer configured to measure a change in a change in an arterial dimensions over a cardiac cycle of said arterial wall of said subject, wherein said arterial dimensions include an arterial distension and an end-diastolic diameter; andwherein said third sensor is configured to provide a feedback to said controller unit to measure BP of said subject and wherein said controller unit is configured to measure the blood pressure (BP) waveform at the artery measurement site of said subject based on said local PWV and said change in arterial dimensions in each heart beat. 2. The system of claim 1, wherein said ultrasound transducer is coupled to said first sensor, wherein said system is configured to simultaneously acquire a surface pressure waveform and an arterial distention waveform from same location. 3. The system of claim 2, wherein an output from said first sensor and said second sensor or said surface pressure waveform and said arterial distention waveform are used to measure said local PWV. 4. The system of claim 1, wherein said first sensor and said second sensor are configured to measure said local PWV of said arterial wall of said subject by: detecting a first pulse waveform at first site along said artery wall;detecting a second pulse waveform at a second site along said artery wall, wherein said first site and said second site is spaced at a distance; andmeasuring said local PWV by computing a time difference between said first pulse waveform and said second pulse waveform. 5. The system of claim 1, wherein said ultrasound transducer is configured to measure said arterial dimensions over said cardiac cycle and the said arterial diameter waveform of said arterial wall of said subject by: sending ultrasound signals through body of said subject;obtaining reflected ultrasound signals from a near arterial wall and a far arterial wall;converting said ultrasound signals to corresponding voltage pulses;splitting said voltage pulses into said plurality of frames;detecting a location of said arterial wall based on a motion profile signal identified from said plurality of frames in said plurality of voltage pulses;tracking motion of said arterial wall from a frame to a subsequent frame, wherein said frame and said subsequent frame is from said plurality of frames;measuring said arterial di... | HEALTHCARE TECHNOLOGY INNOVATION CENTRE | A61B008/08 G16H050/50 A61B005/021 G06F019/00 A61B008/04 A61B005/107 A61B005/00 A61B005/02
US10709599 | B2 | 2020-07-14 | Systems and methods for providing an automated titration for oral appliance therapy | 2016-03-31 | Systems and methods for setting one or more parameters for a multi-test-period titration for oral appliance therapy are described herein. A method can include performing a titration for oral appliance therapy during a first test period, identifying a variable associated with the titration performed during the first test period, and setting a parameter for a titration for oral appliance therapy to be performed during a second test period. The parameter can be dependent on the variable associated with the titration performed during the first test period. The method can also include performing the titration for oral appliance therapy during the second test period, and establishing an outcome of oral appliance therapy based on the titrations performed during the first and second test periods. | 1. A system for setting one or more parameters for a multi-test-period titration for oral appliance therapy, comprising: a mandibular displacement device configured to be positioned in an oral cavity of a subject;a monitoring unit configured to sense one or more physiological inputs from the subject; anda control unit comprising a processing unit and a memory operatively coupled to the processing unit, the memory having computer-executable instructions stored thereon that, when executed by the processing unit, cause the system to: perform a titration for oral appliance therapy during a first test period;identify a variable associated with the titration performed during the first test period;set a parameter for a titration for oral appliance therapy to be performed during a second test period, wherein the parameter is dependent on the variable associated with the titration performed during the first test period;perform the titration for oral appliance therapy during the second test period; andestablish an outcome of oral appliance therapy based on the titrations performed during the first and second test periods. 2. The system of claim 1, wherein the memory has further computer-executable instructions stored thereon that, when executed by the processing unit, cause the system to: receive a desired outcome criteria of oral appliance therapy; andselect a prediction protocol for establishing the outcome of oral appliance therapy based on the desired outcome criteria. 3. The system of claim 1, wherein the memory has further computer-executable instructions stored thereon that, when executed by the processing unit, cause the system to receive diagnostic information concerning the subject, wherein the diagnostic information is used to establish the outcome of oral appliance therapy or to set the parameter for a titration. 4. The system of claim 3, wherein the diagnostic information comprises at least one of a baseline measure of respiratory events experienced by the subject, a distribution of respiratory events, a level of desaturation, the subject's body position during a diagnostic sleep study, a sleep stage during a diagnostic sleep study, the subject's physical characteristics, the subject's medical history, or the subject's symptoms. 5. The system of claim 1, wherein the variable is at least one of a predicted outcome of oral appliance therapy, a predicted effective protrusion level for oral appliance therapy, a predicted optimal protrusion level for oral appliance therapy, a measure of respiratory events, a measure of respiratory events at or above a specific protrusion level, a percentage of time spent at or above a specific protrusion level, a sleep stage, or an occurrence of arousal. 6. The system of claim 1, wherein the parameter is at least one of a beginning protrusion level, a protrusion level adjustment rate, a protrusion level range, a criterion for adjusting protrusion level, a width and position of a correlation window, a type of p... | ZST Holdings, Inc. | A61F005/56 A61C019/04 G01N033/48 A61C007/08 A61B005/08 A61B005/11 A61B005/00
US10716539 | B2 | 2020-07-21 | Ultrasound probe | 2017-10-05 | An optical fiber guides light output from a light source to an ultrasound probe. The ultrasound probe includes a light guiding section that guides the light from a light input end, which is optically coupled with the optical fiber to a light output end provided in the vicinity of ultrasonic transducers. The light guiding section has a first light guiding portion that includes the light input end, and a second light guiding portion that includes the light output end. The first light guiding portion is formed by glass, and magnifies input light. The second light guiding portion is formed by resin, and emits the light toward a subject from the light output end. | 1. A probe, comprising: an optical fiber that guides light emitted from a light source to a probe main body;at least one light guiding section, comprising a first light guiding portion and a second light guiding portion, that guides light from a light input end, at the first light guiding portion and which is optically coupled with the optical fiber, to a light output end at the second light guiding portion; andan ultrasonic transducer that detects photoacoustic waves that are generated within a subject due to light guided by the light guiding section being emitted onto the subject,wherein the first light guiding portion that includes the light input end, that guides input light from the light input end toward the light output end, and that enlarges a cross sectional area of the input light, at an output end of the first light guiding portion,wherein the second light guiding portion that is connected h surface to the first light guiding portion, that includes the light output end, that guides light guided by the first light guiding portion to the light output end, and that emits light from the light output end toward the subject, andwherein the light diffusing surface is formed at a boundary between the second light guiding portion and the first light guiding portion. 2. The probe as defined in claim 1, wherein a detecting surface of the ultrasonic transducer and the light output end of the light guiding section are in the same plane. 3. The probe as defined in claim 1, wherein the first light guiding portion is formed from glass and the second light guiding portion is formed from resin. 4. The probe as defined in claim 1, wherein the first light guiding portion and the second light guiding portion are transparent. 5. The probe as defined in claim 1, wherein the light guiding section comprises a core having light transmitting properties, and low refractive index members that sandwich the core therebetween and have a lower refractive index than the core. 6. The probe as defined in claim 5, wherein the low refractive index members are formed from an organic material. 7. The probe as defined in claim 1, wherein the light guiding section comprises a core having light transmitting properties, and reflective members that sandwich the core therebetween. 8. The probe as defined in claim 1, wherein the ultrasonic transducer comprises a plurality of transducers arranged along one direction of the ultrasonic transducer. 9. The probe as defined in claim 8, further comprising a plurality of light guiding sections, including the at least one light guiding section, arranged along the one direction of the ultrasonic transducer. 10. The probe as defined in claim 8, further comprising a plurality of light guiding sections, including the at least one light guiding section, which face each other with the ultrasonic transducer interposed therebetween. 11. The probe as defined in claim 1, further comprising: the light diffusing surface,wherein the light diff... | FUJIFILM Corporation | A61B008/00 A61B005/00 G01N029/24
US10702177 | B2 | 2020-07-07 | Catheter with bipole electrode spacer and related methods | 2016-08-24 | An electrophysiology catheter constructed with assembly and wiring of ring electrodes using spacer rings of electrically-nonconductive material, whose length is predetermined/premeasured, includes a distal section with at least one spine having two electrodes, and a spacer member therebetween, wherein the spacer member configured to provide a separation gap between the two electrodes. The separation gap may span in an axial direction and/or in a circumferential direction. The spacer member may be configured generally as a ring with a center axial opening configured to receive the spine therethrough. The spacer member may include an axial extension providing a first separation gap between a first pair of electrodes in a circumferential direction, and a circumferential providing a second separation gap between a second pair of electrodes in an axial direction. | 1. An electrophysiology catheter comprising: an elongated body;a distal section distal of the elongated body, the distal section including at least one spine having at least two electrodes, and a generally cylindrical spacer member mounted in a generally surrounding relationship on the at least one spine and having at least a portion spanning between the at least two electrodes, the at least two electrodes spanning less than 360 degrees circumferentially around the at least one spine, wherein the spacer member is made of a nonconductive material and the portion of the spacer member spanning between the at least two electrodes is configured to provide a separation gap of smaller than 1.0 mm between the at least two electrodes. 2. The catheter of claim 1, wherein the separation gap spans in an axial direction. 3. The catheter of claim 1, wherein the separation gap spans in a circumferential direction. 4. The catheter of claim 1, wherein the separation gap is about 50 microns to about 200 microns. 5. The catheter of claim 1, wherein the spacer member includes an axial extension configured to provide a first separation gap between a first pair of the at least two electrodes in a circumferential direction and a circumferential extension configured to provide a second separation gap between a second pair of the at least two electrodes in an axial direction. 6. The catheter of claim 1, wherein the spacer member includes an axial extension spanning between a first pair of the at least two electrodes in a circumferential direction, and a circumferential extension spanning between a second pair of the at least two electrodes. 7. The catheter of claim 1, wherein the spacer member has a first axial extension spanning between a first pair of the at least two electrodes in a circumferential direction, a second axial extension spanning between a second pair of the at least two electrodes in the circumferential direction, and at least one circumferential extension extending between the first and second axial extensions and spanning between an electrode of the first pair and an electrode of the second pair in an axial direction. 8. The catheter of claim 1, wherein outer surfaces of the at least two electrodes and an outer surface of the spacer member together form a 360 degree circumferential surface that surrounds an outer surface of the spine. 9. A method of assembling the catheter of claim 1, comprising: connecting a distal end of a respective lead wire to each of the at least two electrodes;sliding a first one of the at least two electrodes on the spine;sliding the spacer member on the spine;sliding a second one of the at least two electrodes on the spine; andabutting the first one of the at least two electrodes and the second one of the at least two electrodes against the spacer member. 10. An electrophysiology catheter comprising: an elongated body;a distal electrode assembly distal of the elongated body and having a plurality of spines, at least o... | Biosense Webster (Israel) Ltd. | A61B005/042 A61B018/14 A61B005/00 A61M025/00 A61B017/00 A61B018/00
US10716497 | B1 | 2020-07-21 | Method and apparatus for non-invasively monitoring blood glucose level in domesticated animals | 2016-04-06 | A device and method for noninvasively monitoring biofluids of an animal provides measuring physical parameters of an analyte with a device located within an oral cavity of the animal. | 1. A device for measuring at least one analyte level in an animal, the device comprising: a body having a predetermined shape and size at least partially insertable into an oral cavity of the animal and configured to expose a surface of the body to biofluids in the oral cavity;an analyte measuring system housed with the body, including a biosensing unit for measuring at least one analyte in biofluids present on the surface of the body in the oral cavity of the animal, comprising: a at least one capillary conduit connected on a first end to a biofluid container and a second end disposed at or near the surface of the body and configured to enable collection and transport of biofluids from the oral cavity of the animal to the biofluid container, andat least one analyte sensor in fluid connection with biofluids transported to the biofluid container; anda signal output component operably coupled to the biosensing unit and configured to generate at least one output signal in response to the at least one analyte sensor. 2. The device of claim 1, wherein the analyte comprises glucose. 3. The device of claim 1, wherein the signal output component comprises a wireless transmitter. 4. The device of claim 2, further comprising a remote receiver wirelessly receiving the at least one output signal. 5. The device of claim 1, wherein the biofluids comprise at least one of saliva, mucus, blood or breath. 6. The device of claim 1, wherein the biosensing unit further comprises: at least one sensor indicating a desired positioning of the body at least partially within the oral cavity of the animal; anda computing means for activating the biosensing unit to acquire analyte sensing data while the body is in the desire position in the oral cavity. 7. The device of claim 1, wherein: the at least one analyte sensor comprises a plurality of sensors detecting a set of physical variables functionally dependent on glucose concentration in the biofluids and outputting physical variable signals;the biosensing unit further comprises a computing means in communication with the plurality of sensors and determining a glucose value in response to the physical variable signals. 8. The device of claim 7, wherein the plurality of sensors are selected from the group consisting of electrochemical sensors, photoelectric sensors, infrared sensors, acoustic sensors and electromagnetic sensors. 9. The device of claim 1, wherein the biosensing unit comprises: a sensory network including a plurality of sensors detecting a set of physical variables functionally dependent on at least one analyte level in the biofluids and outputting physical variable signals;a computing means for conditioning the physical variable signals to provide stable data and estimating variables for determining the at least one analyte level. 10. The device of claim 9, wherein the estimating variables possess unambiguous relationship with the at least one analyte level, such that the estimating variables are use... | Inesa, Inc. | A61B005/00 A61B005/145 A61B010/00
US10720032 | B2 | 2020-07-21 | Apparel with ultrasonic position sensing and haptic feedback for activities | 2018-07-19 | An article of apparel, a system, and methods include a fabric configured to conform to a body of a wearer. A plurality of ultrasonic positioning sensors are secured with respect to the fabric at a first set of predetermined locations, each of the ultrasonic positioning sensors configured to emit a sound wave configured to be detected by other ones of the plurality of ultrasonic positioning sensors and output an electronic signal indicative of having emitted or detected a sound wave. A plurality of feedback devices secured with respect to the fabric at a second set of predetermined locations, each of the feedback devices configured to output a feedback signal configured to be detectable by the wearer of the article of apparel. | 1. An article of apparel, comprising: a fabric configured to conform to a body of a wearer;a plurality of ultrasonic positioning sensors secured with respect to the fabric at a first set of predetermined locations, each of the ultrasonic positioning sensors configured to emit a sound wave configured to be detected by other ones of the plurality of ultrasonic positioning sensors and output an electronic signal indicative of having emitted or detected a sound wave; anda plurality of feedback devices secured with respect to the fabric at a second set of predetermined locations, each of the feedback devices configured to output a feedback signal configured to be detectable by the wearer of the article of apparel;wherein a processor is configured to:at a plurality of discrete sequential times, determine positional values of the plurality of ultrasonic positioning sensors based, at least in part, on electronic signals output by the plurality of ultrasonic positioning sensors; andcause at least some of the plurality of feedback devices to output the feedback signal based, at least in part, on a difference between the positional values as determined at the plurality of discrete sequential times and a parameter set of an activity program, wherein the activity program comprises a plurality of parameter sets, including the parameter set, sequentially organized to describe a transitional movement of the body of the wearer from a first position to a second position different than the first position, the plurality of parameter sets corresponding to ones of the plurality of discrete sequential times during the transitional movement of the body of the wearer from the first position to the second position. 2. The article of apparel of claim 1, wherein the parameter set comprises at least one target value indicative of a desired distance between two or more of the plurality of positioning sensors, and wherein the processor is configured to cause the individual ones of the plurality of feedback devices to output the feedbacksignal based on a variation between an associated positional value and the target value. 3. The article of apparel of claim 2, wherein the target value is associated with a tolerance value and wherein the processor is configured to cause the individual ones of the plurality of feedback devices to output the feedback signal based on the variation exceeding the tolerance value. 4. The article of apparel of claim 1, wherein the first set of locations comprise a left shoulder, a right shoulder, a left arm, a right arm, a left side, and a right side. 5. The article of apparel of claim 1, wherein the positional values include a distance between a pair of the plurality of ultrasonic positioning sensors. 6. The article of apparel of claim 1, wherein each one of the plurality of feedback devices comprises at least one haptic motor configured to cause individual ones of the plurality of feedback devices to output the feedback signal to induce the wea... | NIKE, Inc. | H04B003/36 G08B006/00 A61B005/11 A41D001/00 A61B005/08 A61B005/00 A41D013/00 A61B005/113
US10720240 | B2 | 2020-07-21 | Context detection for medical monitoring | 2016-02-21 | There is provided a context detection apparatus for use with a medical monitoring device arranged to measure a physiological characteristic of a subject. The context detection apparatus comprises at least one context sensor comprising a camera; and a processing unit. The processing unit is arranged to receive image data acquired by the camera from the at least one context sensor; analyze the received image data to generate a plurality of different types of context information; wherein each type of context information comprises information which relates to a factor capable of influencing measured values of the physiological characteristic and which is not measurable by the medical monitoring device; and output a signal to the medical monitoring device based on the generated context information. The processing unit is arranged to analyze the received image data by detecting features in the received image data; and classifying each detected feature as relating to a particular one of the plurality of types of context information. | 1. A method for use in monitoring a physiological characteristic of a subject using a medical monitoring device, the method comprising: receiving a single type of sensor data acquired by at least one context sensor;analyzing, by a processing unit, the received single type of sensor data by detecting features in the received single type of sensor data and classifying each detected feature as relating to a particular type of a plurality of different types of context information, wherein each type of context information comprises information which relates to a factor capable of influencing measured values of the physiological characteristic and which is not measurable by the medical monitoring device, and the generated context information comprises one or more of information relating to the movement of a body part of the subject, information relating to the posture of a subject, information relating to the activity of a subject, information relating to the location of a subject, information relating to one or more persons in the vicinity of the subject, and information relating to the environment of the subject;comparing each type of generated context information to at least one predefined criterion comprising a set of conditions, which must all be met by a given item of context information in order for a positive relevance value to be assigned to that context information:assigning a relevance value to each type of generated context information based on the comparing step; andoutputting via a communication functionality of the processing unit a context signal to the medical monitoring device based on the generated context information. 2. The method of claim 1, further comprising: measuring multiple physiological characteristics using the medical monitoring, wherein the at least one predefined criterion is specific to a particular type of context information and to a particular physiological characteristic, and wherein assigning the relevance value to each type of generated context information further comprises:assigning the relevance value in respect of each of the multiple physiological characteristics to each type of generated context information. 3. A method according to claim 1, further comprising: receiving, by the medical monitoring device, measured values of a physiological characteristic from a sensor;receiving, by the medical monitoring device, the context signal; andselecting, by the medical monitoring device, an action to perform, wherein the action is selected in dependence on the received context signal and on the received measured values. | KONINKLIJKE PHILIPS N.V. | G16H050/20 A61B005/00 A61B005/11 G16H040/63 A61B005/0205 G06F019/00 A61B005/024 A61B005/1455 A61B005/0402 A61B005/01 A61B005/0215
US10728434 | B2 | 2020-07-28 | Optical coherence tomography imaging system | 2018-11-09 | An imaging system, a method for imaging an object of interest, and an image of the object of interest. In one embodiment, an imaging system includes a channel configured to receive information from an object of interest, a 3DSS sub-system configured to capture information from the object of interest and generate at least one of 3D surface data or 3D position data based on the information captured by the 3DSS sub-system, and an OCT sub-system configured to perform a line scan of light reflected from the object of interest, generate OCT image data from the line scan, receive the at least one of the 3D surface data or the 3D position data, and generate one or more OCT images using the OCT image data and the at least one of the 3D surface data or the 3D position data. | 1. An imaging system comprising: a channel configured to receive information from an object of interest;a three-dimensional surface scanning (3DSS) sub-system configured to capture information from the object of interest, andgenerate at least one of three-dimensional surface data or three-dimensional position data based on the information captured by the 3DSS sub-system; andan optical coherence tomography (OCT) sub-system configured to perform a line scan of light reflected from the object of interest and captured by the OCT sub-system,generate optical coherence tomography (OCT) image data from the line scan,receive the at least one of the three-dimensional surface data or the three-dimensional position data from the 3DSS sub-system, andgenerate one or more optical coherence tomography (OCT) images using the OCT image data and the at least one of the three-dimensional surface data or the three-dimensional position data, wherein the at least one of the three-dimensional surface data or the three-dimensional position data is indicative of a common spatial frame of reference between the 3DSS sub-system and the OCT sub-system. 2. The imaging system of claim 1, further comprising a housing that encloses the channel, the 3DSS sub-system, and the OCT sub-system. 3. The imaging system of claim 1, wherein the 3DSS sub-system is configured to provide the at least one of the three-dimensional surface data or the three-dimensional position data at the same time that the OCT sub-system is configured to perform the line scan. 4. The imaging system of claim 1, wherein the OCT sub-system is further configured to determine whether the OCT sub-system is within a measurement range of the object of interest based on the at least one of the three-dimensional surface data or the three-dimensional position data, andgenerate the one or more OCT images using the OCT image data and the at least one of the three-dimensional surface data or the three-dimensional position data when the OCT sub-system is within the measurement range of the object of interest. 5. The imaging system of claim 1, wherein the OCT sub-system is further configured to determine whether the OCT sub-system is within a measurement range of the object of interest based on the at least one of the three-dimensional surface data or the three-dimensional position data, andadjust one or more imaging parameters associated with the OCT sub-system when the OCT sub-system is not within the measurement range of the object of interest. 6. The imaging system of claim 1, wherein the 3DSS sub-system is further configured to determine whether the OCT sub-system is within a measurement range of the object of interest based on the at least one of the three-dimensional surface data or the three-dimensional position data, andcontrol the OCT sub-system to generate the one or more OCT images using the OCT image data and the at least one of the three-dimensional surface data or the three-dimensional position data when th... | DENTAL IMAGING TECHNOLOGIES CORPORATION ORMCO CORPORATION | G06K009/00 H04N005/225 A61B005/00 A61C009/00
US10716631 | B2 | 2020-07-21 | Apparatus and methods for use with skeletal procedures | 2017-03-13 | Apparatus for performing a procedure using a tool configured to be advanced into a skeletal portion within a body of a subject along a longitudinal insertion path, and for use with: (a) a 3D imaging device configured to acquire 3D image data of the skeletal portion.(b) a 2D x-ray imaging device that is unregistered with respect to the subject's body and configured, while a portion of the tool is disposed at a first location along the longitudinal insertion path. | 1. A method for performing a procedure using a tool that has at least one radiopaque feature, the tool being configured to be advanced into a skeletal portion within a body of a subject along a longitudinal insertion path, the method comprising: acquiring 3D image data of the skeletal portion;while a portion of the tool is disposed at a first location along the longitudinal insertion path with respect to the skeletal portion, sequentially:acquiring a first 2D x-ray image of at least the portion of the tool and the skeletal portion from a first view, using a 2D x-ray imaging device that is unregistered with respect to the subject's body and that is disposed at a first pose with respect to the subject's body;moving the 2D x-ray imaging device to a second pose with respect to the subject's body; andwhile the 2D x-ray imaging device is at the second pose, acquiring a second 2D x-ray image of at least the portion of the tool and the skeletal portion from a second view;using at least one computer processor:registering the first and second x-ray images to the 3D image data, the registering comprising:generating a plurality of 2D projections from the 3D image data, andidentifying respective first and second 2D projections that match the first and second 2D x-ray images of the skeletal portion;identifying a location of the portion of the tool with respect to the skeletal portion within the first and second 2D x-ray images, using the at least one radiopaque feature of the tool, by means of image processing;based upon the identified location of the portion of the tool within the first and second 2D x-ray images, and the registration of the first and second 2D x-ray images to the 3D image data, determining the first location of the portion of the tool with respect to the 3D image data;subsequently, moving the portion of the tool to a second location along the longitudinal insertion path with respect to the skeletal portion;subsequent to moving the portion of the tool to the second location, acquiring one or more additional 2D x-ray images of at least the portion of the tool and the skeletal portion from a single image view;using the computer processor:identifying the second location of the portion of the tool within the one or more additional 2D x-ray images, using the at least one radiopaque feature of the tool, by means of image processing;deriving the second location of the portion of the tool with respect to the 3D image data, based upon the second location of the portion of the tool within the one or more additional 2D x-ray images, and the determined first location of the portion of the tool with respect to the 3D image data; andgenerating an output, at least partially in response thereto. 2. The method according to claim 1, wherein acquiring the one or more additional 2D x-ray images of at least the portion of the tool and the skeletal portion from the single image view comprises acquiring the one or more additional 2D x-ray image of at least the por... | Vuze Medical Ltd. | A61B005/05 G06K009/00 A61B034/20 A61B046/20 A61B006/12 A61B006/00 A61B090/00 A61B008/00 G06T015/06 G06T015/08 A61B034/30 G06T007/30
US10702670 | B2 | 2020-07-07 | Utilization of packaging paper as stabilization device | 2015-09-23 | A medical device packaging assembly includes a cover (10), a medical packaging paper (20), a medical device (50), and an adhesive dressing (30). The cover has a pocket (14) and a perimeter (12) extending from the perimeter of the pocket. The medical packaging paper forms a hermetically tight seal between the medical packaging paper and the perimeter of the cover around the pocket. The medical device is contained in the hermetically sealed pocket along with the adhesive dressing. The adhesive dressing is removably adhered to the medical packaging paper. | 1. A combination catheter assembly and packaging assembly comprising: a cover comprising an inside surface defining a pocket and a perimeter extending outwardly from a perimeter of the pocket;a packaging paper with a release liner surface peelably affixed to the cover, the packaging paper forming a hermetical seal with the perimeter of the cover around the pocket and capturing a catheter assembly at least partially inside the pocket such that the catheter assembly is directly exposed to the inside surface of the cover, the catheter assembly comprising a catheter hub having a catheter tube extending from the catheter hub and a needle hub having a needle extending from the needle hub; anda strip of adhesive dressing having an adhesive layer located within a space defined by the pocket and the packaging paper and separated from the catheter assembly, the strip of adhesive dressing is located between the cover and the catheter assembly and separable from the cover or between the packaging paper and the catheter assembly and separable from the packaging paper. 2. The combination of claim 1, wherein the catheter assembly comprises a needle guard having a surface located to a side of the needle in a ready to use position. 3. The combination of claim 2, wherein the strip of adhesive dressing has a narrow region located between two regions that are wider than the narrow region. 4. The combination of claim 3, wherein the strip of adhesive dressing is compatible with EtO and gamma sterilization. 5. The combination of claim 2, wherein the strip of adhesive dressing is made from a breathable material, a polyester fabric, a polyurethane fabric, or a foam material. 6. The combination of claim 1, wherein the strip of adhesive dressing is removably attached to the packaging paper. 7. The combination of claim 6, wherein the release liner surface contacts the strip of adhesive dressing. 8. The combination of claim 7, wherein the release liner surface is a surface coated with polyethylene, polypropylene, fluorocarbons, or silicone release agents. 9. The combination of claim 7, further comprising an adhesive dressing sheet attached to the release liner surface of the packaging paper, wherein the strip of adhesive dressing forms a portion of the adhesive dressing sheet and wherein the adhesive dressing sheet is removably adhered to the packaging paper. 10. The combination of claim 9, wherein the strip of adhesive dressing is attached to the adhesive dressing sheet along perforated lines, and the strip of adhesive dressing is configured to separate from the adhesive dressing sheet along the perforated lines. 11. The combination of claim 6, wherein the adhesive material is a medical grade, pressure-sensitive adhesive. 12. The combination of claim 11, wherein the adhesive material is an acrylic based, silicone, or hydrocolloid adhesive. 13. The combination of claim 6, wherein the strip of adhesive dressing has an opening matching a shape of a protrusion on th... | B. Braun Melsungen AG | A61M025/00 A61F013/00 A61M005/00 B65D075/32 A61B005/15 B65D085/20 B65D085/24 A61F013/02 A61M025/02 B65D025/10
US10709605 | B1 | 2020-07-14 | Eye monitoring and pressurization systems | 2017-08-28 | Assemblies and methods for modifying an intraocular pressure of a patient's one or both eyes are disclosed. The assemblies and methods can be used to treat, inhibit, or prevent ocular conditions such as glaucoma, high intraocular pressure, optic disc edema, idiopathic intracranial hypertension, zero-gravity induced papilledema, and other optic pressure related conditions. An assembly can include a goggle including at least one cavity, a pump in fluid communication with the at least one cavity, and a control mechanism. The control mechanism can be operatively coupled to the pump and can maintain a target pressure or target pressure range in the at least one cavity, which, when the assembly is worn by a patient, is the area between a patient's eye(s) and wall surfaces of the goggle. Controlling the pressure over the outer surfaces of the patient's eye(s) can drive a desired change in the intraocular pressure of the eye(s). | 1. A system to treat a patient eye comprising: a goggle, sized and shaped to fit over the patient eye, configured to define a cavity between the goggle and the patient eye;a cavity pressure sensor, configured to sense an indication of cavity pressure in the cavity; anda control circuit configured to receive the indication of cavity pressure and at least one of an indication of intraocular pressure (IOP) level in the patient eye or an indication of cerebrospinal fluid pressure (CSFP) level in the patient,wherein the control circuit is configured to adjust a pressure source to generate non-ambient pressure in the cavity based at least in part on one of the received indications of IOP level or CSFP level to adjust the indication of IOP level in the patient eye toward a target IOP level. 2. The system of claim 1; further comprising a biosensor configured to sense at least one of the indication of IOP level or the indication of CSFP level. 3. The system of claim 2, wherein the biosensor includes the biosensor configured to sense IOP level. 4. The system of claim 2, wherein the biosensor includes the biosensor configured to sense the indication of CSFP level. 5. The system of claim 1; further comprising a pressure source, in fluidic communication with the cavity, configured to generate non-ambient pressure in the cavity to adjust the indication of cavity pressure sensed by the cavity pressure sensor. 6. The system of claim 5, wherein the pressure source is configured to generate non-ambient pressure in the cavity greater than ambient pressure surrounding the cavity. 7. The system of claim 5, wherein the pressure source is configured to generate non-ambient pressure in the cavity less than ambient pressure surrounding the cavity. 8. The system of claim 5, wherein the control circuit is configured to adjust the pressure source to generate non-ambient pressure in the cavity to equalize an indication of-translaminar pressure (TLP) level associated with the patient eye, wherein equalizing the indication of TLP level includes reducing the indication of TLP level from a first TLP level to a lower second TLP level. 9. The system of claim 5, wherein the control circuit is configured to adjust the pressure source to generate non-ambient pressure in the cavity to enhance axonal transport in the optic nerve of the patient eye. 10. The system of claim 5, wherein the control circuit is configured to adjust the pressure source to generate non-ambient pressure in the cavity to help inhibit, treat, or prevent glaucoma or progression of glaucoma in the patient eye. 11. The system of claim 5, wherein the control circuit is configured to adjust the pressure source to generate non-ambient pressure in the cavity sufficient to achieve an IOP level in a range of about 10 mmHg to about 21 mmHg in the patient eye. 12. A method of using a system to treat a patient eye, the system comprising a goggle, sized and shaped to fit over the patient eye and configured to defin... | John Berdahl | A61F009/007 A61B005/00 A61F009/02 A61B003/16 A61B005/03
US10722185 | B2 | 2020-07-28 | Voice controlled assistance for monitoring adverse events of a user and/or coordinating emergency actions such as caregiver communication | 2017-11-03 | A user, such as an elderly person, may be assisted by an assistance device in his or her caregiving environment that operates in conjunction with one or more server computers. The assistance device may execute a schedule of assistance actions where each assistance action is associated with a time and is executed at that time to assist the user. An assistance action may present an input request to a user, process a voice input of the user, and analyze the voice input to determine that the voice input corresponds to a negative response event, a positive response event, or a non-response event. Based on the categorization of one or more voice inputs as negative response events, positive response events, or non-response events, it may be determined to notify a caregiver of the user, for example where the user has not responded to a number of assistance actions. | 1. A system for providing assistance to a user, the system comprising: at least one server computer comprising at least one processor and at least one memory; andan assistance device in a caregiving environment of the user, the assistance device comprising a processor, a memory, a speaker, and a microphone, wherein the at least one server computer and the assistance device are configured to: receive, from a caregiver, information about a schedule of assistance actions for interacting with the user, each assistance action corresponding to a predetermined time of day for implementing the assistance action, the schedule of assistance actions comprising a first assistance action corresponding to a first predetermined time of day for invoking the first assistance action,store, at the at least one server computer, the information about the schedule of assistance actions,transmit, by the at least one server computer to the assistance device at the first predetermined time of day, a first action instruction for implementing the first assistance action,in response to receiving the first action instruction from the at least one server computer, initiate, by the assistance device, the first assistance action by generating an invocation command and transmitting the invocation command to a speech processing server to obtain instructions for presenting audio of a first input request,receive, from the speech processing server, the instructions for presenting the audio of the first input request,present, by the assistance device, the audio of the first input request to the user via the speaker, receive a first voice input of the user via the microphone, and transmit the first voice input to the speech processing server,receive, from the speech processing server, an indication that the first voice input is a negative response event, a positive response event, or a non-response event,receive, from the caregiver, a safety profile of the user, wherein the safety profile specifies when a notification is to be sent to the caregiver based on at least one of a negative response event, a positive response event, or a non-response event, andselectively notify the caregiver by causing a notification to be sent to a device of the caregiver based on the safety profile and the indication that the first voice input corresponds to the negative response event, the positive response event, or the non-response event. 2. The system of claim 1, wherein the first assistance action assists the user in ordering dinner. 3. The system of claim 1, wherein the system comprises one or more sensing pods in the caregiving environment of the user, wherein each sensing pod comprises a microphone and a speaker. 4. The system of claim 1, wherein: the assistance device comprises a presence sensor and is configured to transmit a second invocation command in response to detecting a presence of the user with the presence sensor; andthe at least one server computer is configured to, in response to ... | LifePod Solutions, Inc. | A61B005/00 G08B021/04 G08B025/01 G08B017/10 G10L015/22 G10L015/18 G10L015/26 G10L025/63 G10L025/66 G06F003/01 G16H040/67 G06Q050/22 H04N021/2187 H04W004/90 G08B007/06 G06F019/00 A61B005/11 G06F003/16 G10L015/08
US10709568 | B2 | 2020-07-14 | Intervertebral body spacer assembly and kit | 2018-09-05 | An intervertebral spacer assembly having a plurality of articulated intervertebral spacers coupled in series, each intervertebral spacer releasably and connectable to an adjacent intervertebral spacer, each intervertebral spacer having a cannula extending from a first exterior surface of a tail portion through a nose portion to reach a second exterior surface of the nose portion. | 1. An intervertebral spacer kit, comprising: a plurality of candidate intervertebral spacers, each of the plurality of candidate intervertebral spacers releasably connectable to at least one other of the plurality of candidate intervertebral spacers, and wherein at least one of the plurality of candidate intervertebral spacers has different dimensions than others of the plurality of candidate intervertebral spacers; wherein each intervertebral spacer comprises a first end having a nose portion and an opposing second end having a tail portion, the nose portion sized and shaped to complementarily fit into the tail portion of an adjacent intervertebral spacer; and wherein each intervertebral spacer of the plurality of intervertebral spacers comprises a respective cannula extending through the intervertebral spacer, the cannula of each intervertebral spacer sized and shaped to receive a wire that extends through the cannula of each of the plurality of intervertebral spacers. 2. The intervertebral spacer kit of claim 1 wherein the plurality of candidate intervertebral spacers include a plurality of modular candidate intervertebral spacers, each modular candidate intervertebral spacer having a first end and a second end, the first end sized and shaped to nest with the second end of an adjacent modular candidate intervertebral spacer to form an articulable joint. 3. The intervertebral spacer kit of claim 1 wherein at least one of the plurality of candidate intervertebral spacers has asymmetric coronal dimensions. 4. The intervertebral spacer kit of claim 1 wherein at least one of the plurality of candidate intervertebral spacers has asymmetric sagittal dimensions. 5. An intervertebral spacer assembly, comprising: a plurality of articulated intervertebral spacers coupled in series, each intervertebral spacer releasably connectable to an adjacent intervertebral spacer, wherein each intervertebral spacer comprises a first end having a nose portion and an opposing second end having a tail portion, the nose portion sized and shaped to complementarily fit into the tail portion of an adjacent intervertebral spacer; and wherein each intervertebral spacer of the plurality of intervertebral spacers comprises a respective cannula extending through the intervertebral spacer, the cannula of each intervertebral spacer sized and shaped to receive a wire that extends through the cannula of each of the plurality of intervertebral spacers. 6. The intervertebral spacer assembly of claim 5 wherein the plurality of intervertebral spacers are selected from a plurality of candidate intervertebral spacers, wherein at least one of the plurality of candidate intervertebral spacers has different dimensions than remaining candidate intervertebral spacers. 7. The intervertebral spacer assembly of claim 5 wherein the nose portion has a rounded bullet-shape and the tail portion a concave shape sized and shaped to receive the nose portion of at least one other intervertebral spa... | - | A61F002/44 A61B017/16 A61M029/00 A61B005/00 A61F002/46 A61F002/30
US10718835 | B2 | 2020-07-21 | Electromagnetic shielding for magnetic resonance imaging methods and apparatus | 2018-09-05 | According to some aspects, a portable magnetic resonance imaging system is provided, comprising a B0 magnet configured to produce a B0 magnetic field for an imaging region of the magnetic resonance imaging system, a noise reduction system configured to detect and suppress at least some electromagnetic noise in an operating environment of the portable magnetic resonance imaging system, and electromagnetic shielding provided to attenuate at least some of the electromagnetic noise in the operating environment of the portable magnetic resonance imaging system, the electromagnetic shielding arranged to shield a fraction of the imaging region of the portable magnetic resonance imaging system. According to some aspects, the electromagnetic shield comprises at least one electromagnetic shield structure adjustably coupled to the housing to provide electromagnetic shielding for the imaging region in an amount that can be varied. | 1. A portable magnetic resonance imaging system comprising: a B0 magnet configured to produce a B0 field for an imaging region of the portable magnetic resonance imaging system, the B0 magnet comprising: a first magnet; anda second magnet positioned above the first magnet,wherein the imaging region is disposed between the first and second magnets;a housing for the B0 magnet, wherein at least a part of the housing is positioned between the B0 magnet and the imaging region; andat least one electromagnetic shield structure slidably coupled to the housing to provide electromagnetic shielding for the imaging region in an amount that can be varied by sliding the at least one electromagnetic shield structure about the imaging region. 2. The portable magnetic resonance imaging system of claim 1, further comprising a noise reduction system configured to detect and suppress at least some electromagnetic noise in and operating environment of the portable magnetic resonance imaging system. 3. The portable magnetic resonance imaging system of claim 2, wherein the at least one noise reduction system comprises at least one sensor to detect at least some of the electromagnetic noise in the operating environment of the portable magnetic resonance imaging system and at least one controller to suppress noise based on the at least some of the electromagnetic noise detected by the at least one sensor. 4. The portable magnetic resonance imaging system of claim 1, wherein the at least one electromagnetic shield structure is adjustable to shield between approximately seventy-five percent and approximately fifty percent of the imaging region. 5. The portable magnetic resonance imaging system of claim 1, wherein the at least one electromagnetic shield structure is adjustable to shield between approximately seventy-five percent and approximately twenty-five percent of the imaging region. 6. The portable magnetic resonance imaging system of claim 1, wherein the at least one electromagnetic shield structure is adjustable to shield between approximately fifty percent and approximately twenty-five percent of the imaging region. 7. The portable magnetic resonance imaging system of claim 1, wherein the at least one electromagnetic shield structure comprises a plurality of moveable shields coupled to the housing. 8. The portable magnetic resonance imaging system of claim 7, wherein the plurality of moveable shields are capable of being moved independently of one another to allow for the amount of shielding of the imaging region to be configured. 9. The portable magnetic resonance imaging system of claim 7, wherein each of the plurality of moveable shields comprises electrical gasketing disposed substantially about the periphery of the respective moveable shield. 10. The portable magnetic resonance imaging system of claim 1, further comprising: a plurality of radio frequency coils configured to, when operated, transmit radio frequency signals to the imaging region of the ... | Hyperfine Research, Inc. | G01R033/383 G01R033/385 G01R033/44 A61B050/13 G01R033/389 G01R033/421 G01R033/56 G01R033/38 A61B005/055 A61B006/00 A61G013/10 G01R033/34 G01R033/48 A61B090/00 G01R033/3873 G01R033/36 G01R033/422
US10716640 | B2 | 2020-07-21 | Electronic orientation monitor and an associated method | 2018-04-11 | The invention, in one aspect, provides an electronic orientation monitor. The monitor includes an orientation sensing electronics configured for calibration when in a reference orientation and being responsive to manipulation of the monitor so as to calculate first, second and third angles which together represent a difference between a current orientation of the monitor and the reference orientation. The monitor further includes a display being responsive to the first and second angles so as to display a point positioned relative to first and second axes, the display being further responsive to the third angle so as to display a line having a direction relative to the first and second axes such that a combination of the position of the point and the direction of the line is indicative to a user of the difference between the current orientation of the monitor and the reference orientation. | 1. A system for calibrating a surgical electronic orientation monitor, the system comprising: a surgical electronic orientation monitor detachably engageable to a brace and a surgical implement;the brace usable for taking measurements for calibrating the surgical electronic orientation monitor for subsequent use with the surgical implement, the brace comprising:a rigid frame including a first elongated portion extending in a first direction and a second elongated portion extending in a second direction crossing the first direction, wherein the first and second elongated portions are connected to each other at a first frame connectioin point;a docking station to which the surgical electronic orientation monitor can be attached;first and second arcuate members connected to the first elongated portion at respective first and second member connection points with the first frame connection point being disposed between the first and second arcuate members,wherein a position of at least one of the first and second member connection points is adjustable, allowing the at least one of the first and second arcuate members to be selectively positioned along the first elongated portion such that a separation distance between the first and second arcuate members can be selectively set to accommodate a patient's body between the first and second arcuate members, anda third elongated member connected to the second elongated portion;a first force sensor disposed on the first arcuate member, a second force sensor disposed on the second arcuate member and a third force sensor disposed on the third elongated member;wherein distal ends of the first and second arcuate members and a distal end of the third elongated member are respectively configured to be placed over the patient's skin, the first, second and third force sensors providing force measurements indicative of the thickness of soft tissue in the vicinity of each respective arcuate member when the brace is mounted with the first and second arcuate members and the third elongated member at respective anatomical sites on a patient, the brace being adjustable to vary the force readings to place the docking station in a reference orientation relative to the respective anatomical sites for calibration of the surgical electronic orientation monitor when attached to the docking station. 2. The system of claim 1, further comprising a first loop disposed at the distal end of the first arcuate member and a second loop disposed at the distal end of the second arcuate member. 3. The system of claim 2, wherein the first arcuate member is rotatable about a first axis of rotation, the first axis of rotation being orthogonal to an upper surface of the frame. 4. The system of claim 3, wherein a center of the first loop is in axial alignment with the first axis of rotation. 5. The system of claim 1, further comprising: a slot extending at least partially along the first elongated portion; anda threaded fastener configured ... | Gyder Surgical Pty Ltd. | A61B090/00 A61B005/00 A61B008/08 A61B005/107 A61B005/1495
US10716701 | B2 | 2020-07-21 | Oral and nasal devices for the treatment of sleep apnea and/or snoring with filter and sensors to provide remote digital monitoring and remote data analysis | 2015-10-07 | The invention relates to devices and methods for monitoring, diagnosing and/or treating sleep apnea. Devices of the invention comprise a mouthpiece having a filter assembly and a microchip. Devices of the invention may also comprise a nasal splint having a filter assembly and a microchip. The invention also relates to a system comprising devices of the invention. | 1. A system for treating sleep apnea in a patient, the system comprising: a mouthpiece;a bi-directional filter assembly; anda microchip comprising sensors,wherein the filter assembly and the microchip are positioned within the mouthpiece, andwherein the filter assembly comprises:a filter grid;a removable filter grid insert restrained by a slot of the filter assembly behind the filter grid, wherein the filter assembly is configured to bi-directionally direct a first portion of air through the filter grid insert and a second portion of air around the filter grid insert through a portion of the filter grid adjacent to the filter grid insert; anda filter flap that is configured to be closed against the filter grid insert during the patient's exhalation to create greater resistance during the patient's exhalation than inhalation to treat sleep apnea. 2. The system of claim 1, further comprising: a recharging cradle; anda non-transitory computer readable storage medium including instructions that, when executed by a processor, cause the system to analyze data obtained from the sensors and to generate reports. 3. The system of claim 1, further comprising a hydrating mesh surrounding the filter assembly. 4. The system of claim 1, wherein the slot of the filter assembly is positioned at a region of the filter assembly proximal to the patient's teeth. 5. The system of claim 1, wherein the slot of the filter assembly is positioned at a region of the filter assembly furthest away from the patient's teeth. 6. The system of claim 1, wherein the filter grid insert comprises a plurality of tubes with protuberances for generating laminar air resistance. 7. The system of claim 1, further comprising: a second removable filter grid insert configured to be placed in the slot of the filter assembly, replacing the removable filter grid insert restrained by the slot of the filter assembly. 8. The system of claim 1, further comprising: a face plate for the filter grid being positioned on a side of the [system] closer to the patient's mouth, wherein the filter flap is attached to the face plate. 9. The system of claim 8, wherein the filter flap is configured to be removable from the face plate and replaced with a different filter flap. 10. A device for treating sleep apnea in a patient, the device comprising: a mouthpiece;a bi-directional filter assembly positioned within the mouthpiece comprising:a filter grid;a face plate for the filter grid, the face plate being positioned on a side of the device closer to the patient's mouth;a removable filter grid insert restrained by a slot near the face plate, wherein the filter assembly is configured to bi-directionally direct a first portion of air through the filter grid insert and a second portion of air around the filter grid insert through a portion of the filter grid adjacent to the filter grid insert, wherein the filter grid insert comprises a plurality of tubes with protuberances for generating laminar air resista... | ALLREST TECHNOLOGIES | A62B023/02 A61F005/56 A61B005/00 A61F005/08 A62B023/06
US10716893 | B2 | 2020-07-21 | Adhesive and peripheral systems and methods for medical devices | 2018-02-15 | A repeater system may control a pump by using a repeater and a user interface. An adhesive patch system may be used for affixing a pump or other object to a human body. Such an adhesive patch system may include two sets of adhesive members, each member including an adhesive material on at least one side so as to attach to the body. The members of the first set are spaced to allow the members of the second set to attach to the body in spaces provided between the members of the first set, and the members of the second set are spaced to allow members of the first set to detach from the body without detaching the members of the second set. Also, fill stations and base stations are provided for personal pump systems. | 1. An inserter system for insertion of an infusion device in fluid-flow communication with the inlet or outlet port of a reservoir and for insertion of a sensor, the insertion system comprising: a bridge;an infusion device introduction needle and a sensor introduction needle supported by the bridge for movement between a retracted position and an extended position;an inserter housing for accommodating the bridge and for movement of the infusion device introduction needle and the sensor introduction needle between the retracted position and the extended position;an infusion device about the infusion device introduction needle and moveable with the infusion device introduction needle between the retracted position and the extended position;a sensor about the sensor introduction needle and moveable with the sensor introduction needle between the retracted position and the extended position;an infusion device release located on the bridge for releasing the infusion device in the extended position; anda sensor release located on the bridge for releasing the sensor in the extended position. 2. The inserter system of claim 1, the bridge further comprising a cover, the cover further comprising a needle passage through which the infusion device introduction needle and a sensor introduction needle extends when the infusion device introduction needle and a sensor introduction needle are in the retracted position and moving toward the extended position. 3. The inserter system of claim 1, the infusion device further comprising a central passage through which the infusion device introduction needle extends when the infusion device introduction needle is in the retracted position and moving toward the extended position. 4. The inserter system of claim 1, the sensor further comprising a central passage through which the sensor introduction needle extends when the sensor introduction needle is in the retracted position and moving toward the extended position. 5. The inserter system of claim 1, wherein the infusion device is located in a position adapted to extend into a user's skin when the infusion device introduction needle and infusion device are locked in the extended position. 6. The inserter system of claim 1, wherein the infusion device introduction needle is located in a position adapted to pierce a user's skin wherein the infusion device introduction needle is moved to the extended position, and wherein the infusion device introduction needle is located in a position external to a user's skin, when the infusion device introduction needle is in the retracted position. 7. The inserter system of claim 1, wherein the inserter housing comprises a spring. 8. The inserter system of claim 1, the bridge configured to be removably attached to the inserter housing, the inserter housing configured for imparting a force on the bridge in a direction for moving the infusion device introduction needle and the sensor introduction needle from thea retracted positio... | DEKA Products Limited Partnership | A61M005/145 A61M005/158 A61M005/168 F04B043/02 G05D007/06 F04B043/09 F04B043/12 A61M005/142 A61M005/172 G08C017/02 H04B007/26 G05B023/02 A61B005/00 A61M005/50 G01F022/00 A61M005/14 A61M005/162 A61M005/36 A61B005/155 A61J001/20
US10702214 | B2 | 2020-07-07 | Method and system for monitoring a patient's medical condition | 2017-11-17 | A method for monitoring and visualizing a patient's medical condition, wherein a graphical representation of the patient comprising a body having at least a torso and a head, as well as particularly two legs and two arms, is displayed using a display device, wherein said displayed graphical representation comprises at least one region which is allocated to at least one or several provided (e.g. measured and/or determined) patient monitoring quantities, and wherein the appearance of the at least one region is altered in real-time when the at least one patient monitoring quantity to which said at least one region is allocated changes. | 1. A method for monitoring a medical state of a patient through a dynamically rendered visualization of a synthetic two or three dimensional patient model, the monitoring involving monitoring one or more medical parameters of the patient, the patient model representing the medical state of the monitored patient and being rendered into a graphical representation of the patient over a range of values of a specific monitored medical parameter of the patient, the method comprising: a group of algorithms including: a first patient monitoring algorithm, wherein a volume of a rendered region corresponding to the patient's vena cava region is associated with a current value of the patient's mean central venous pressure (CVP), wherein the volume of the patient's vena cava region is re-rendered when the current value of the patient's mean central venous pressure (CVP) changes, and wherein the volume of the patient's vena cava region is proportional to the current value of the patient's mean central venous pressure (CVP);a second patient monitoring algorithm, wherein a volume of a rendered region corresponding to the patient's lung is altered between a first value and a second value derived from a tidal volume (TV), respectively, with a display refresh frequency associated with a respiratory rate (RR), and wherein the first value is smaller than the second value;a third patient monitoring algorithm, wherein a volume of a rendered region corresponding to the patient's expiratory CO2 cloud is altered between a first value and a second value derived from a capnographic value (CAP), respectively, with the display refresh frequency being associated with the patient's respiratory rate (RR), and wherein the first value is smaller than the second value;a fourth patient monitoring algorithm, wherein a rendered region corresponding to the patient's eye is altered when a bispectral index (BI) value changes, wherein when the index (BI) lies in a first range, the patient's eye region is rendered in an open state, and wherein when the index (BI) lies in a second range, the patient's eye region is rendered in a closed state, and wherein when the index (BI) is in a third range, the patient's the eye region is rendered in a partially open state; anda fifth patient monitoring algorithm, wherein a rendered region corresponding to the patient's hand is altered when a current neuromuscular transmission (NMT) value changes, wherein when the current neuromuscular transmission (NMT) value lies in a first range, the patient's hand region is rendered in a first state, in a relaxed state, and wherein when the current neuromuscular transmission (NMT) value lies in a second range, the patient's hand region is rendered in a different second state with an extended thumb and tightened fingers; andre-rendering the graphical representation of the monitored patient in response to changes in the monitored patient's medical state according to at least one algorithm selected from the group of a... | UNIVERSITÄT ZÜRICH | A61B005/00 G16H040/63 G16H010/60 A61B005/0205
US10722320 | B2 | 2020-07-28 | Thermo-electric generator | 2015-03-31 | Disclosed herein, a power supply may be recharged with a generally used system. In various embodiments, cleaning systems may be used to affect a generator to cause the generator to charge a power supply. Such charging systems may be used without the need for separate or specialized power charging systems. A charging source or system may, therefore, also charge a power supply without requiring additional steps. | 1. A system to power at least a portion of an instrument, comprising: at least one tracking device on the instrument configured to consume power to sense a condition for determination of a location of the at least one tracking device;a power source configured to power the at least one tracking device on the instrument; anda thermo-electric generator configured to generate a current to charge the power source;wherein the power source is configured to provide a current to the at least one tracking device for operation thereof on the instrument;wherein the thermo-electric generator is configured to generate the current during a heat cleaning cycle of the instrument. 2. The system of claim 1, wherein a temperature differential generated from an external heat source causes the thermo-electric generator to generate the current during the heat cleaning cycle of the instrument. 3. The system of claim 2, wherein the tracking device includes at least one coil of a conductive material configured to sense a field. 4. The system of claim 3, further comprising: a tracking system having a localizer configured to generate the field. 5. The system of claim 4, further comprising: a subject tracking device configured to be fixed to a subject for maintaining a registration of a subject space to a patient space;the subject tracking device having a second thermal electric generator to charge a second power source included with the subject tracking device. 6. The system of claim 1, further comprising: a tool bit, wherein the tool bit includes the at least one tracking device incorporated in the tool bit;wherein the thermal electric generator is separate from the tool bit. 7. The system of claim 1, further comprising: a navigation system having a display and a navigation processor;wherein the navigation processor is configured to determine a location of the at least tracking device based on the sensed condition and the display is configured to display a representation of the instrument based on the determined location of the at least one tracking device. 8. A system for navigating an instrument, comprising: an instrument for performing a procedure;a tracking system including a localizer configured to emit an electromagnetic field;a first tracking device on the instrument, the tracking device having: at least one coil of conductive material in communication with a communication system configured to consume power to sense the emitted electromagnetic field and transmit a property related to the sensed field for determination of a location of the at least one coil;a power source configured to provide the power to the at least one coil;a thermo-electric generator configured to generate a current to charge the power source;wherein the power source is configured to lose stored energy during use of the at least one coil and be recharged from the current generated with the thermo-electric generator;wherein the thermo-electric generator is configured to generate the curren... | Medtronic Navigation, Inc. | A61B005/00 A61B090/70 A61B034/20 A61B017/00 A61L002/04 H02J005/00 H04B005/00 A61B090/00 A61B017/16 H01L035/28
US10702207 | B2 | 2020-07-07 | System and method for determining spectral boundaries for sleep stage classification | 2015-12-04 | The present disclosure pertains to a system (10) configured to determine spectral boundaries (216, 218) for sleep stage classification in a subject (12). The spectral boundaries may be customized and used for sleep stage classification in an individual subject. Spectral boundaries determined by the system that are customized for the subject may facilitate sleep stage classification with higher accuracy relative to classifications made based on static, fixed spectral boundaries that are not unique to the subject. In some implementations, the system comprises one or more of a sensor (16), a processor (20), electronic storage (22), a user interface (24), and/or other components. | 1. A system configured to determine subject specific spectral boundaries for sleep stage classification in a subject, the system comprising: one or more sensors configured to generate output signals that convey information related to a respiratory wave amplitude metric for a sleep session of the subject; andone or more physical computer processors configured by computer readable instructions to: transform the information conveyed by the output signals in a first set of individual epochs of time into a frequency domain;determine individual frequencies of respiratory wave amplitude metric peaks within the first set of individual epochs of time;determine an aggregated frequency of the respiratory wave amplitude metric peaks by aggregating the determined individual frequencies of the respiratory wave amplitude metric peaks;determine subject specific upper spectral boundaries and lower spectral boundaries for sleep stage classification for the subject based on the aggregated frequency, wherein the upper spectral boundaries are a function of the aggregated frequency and upper coefficients, and wherein the lower spectral boundaries are a function of the aggregated frequency and lower coefficients; anddetermine sleep stages of the subject during a second set of individual epochs of time in a subsequent sleep session as a function of the aggregated frequency of the respiratory wave amplitude metric peaks using the determined subject specific upper and lower boundaries. 2. The system of claim 1, wherein the one or more sensors and the one or more physical computer processors are configured such that the respiratory wave amplitude metric is a power spectral density. 3. The system of claim 2, wherein the one or more physical computer processors are configured such that determining the aggregated frequency of the respiratory wave amplitude metric peaks comprises averaging frequencies of power spectral density peaks from the first set of individual epochs of time. 4. The system of claim 3, wherein the one or more physical computer processors are configured such that an average frequency of the power spectral density peaks from individual thirty second epochs of time during the sleep session is a mean respiratory frequency of the subject. 5. The system of claim 4, wherein the one or more physical computer processors are configured such that the subject specific upper spectral boundaries are determined based on the mean respiratory frequency using linear regression and upper regression coefficients, and the lower spectral boundaries are determined based on the mean respiratory frequency using linear regression and lower regression coefficients. 6. A method to determine subject specific spectral boundaries for sleep stage classification in a subject with a determination system, the determination system comprising one or more sensors and one or more physical computer processors, the method comprising: generating, with the one or more sensors, output signals t... | KONINKLIJKE PHILIPS N.V. | A61B005/00 A61B005/024 A61B005/08 A61B005/0402 A61B005/11 A61B005/113
US10702140 | B2 | 2020-07-07 | System and method for providing analysis of visual function using a mobile device with display | 2017-02-03 | A visual function evaluation is performed using a sequence of interactions with a mobile device. A patient user may perform a variety of visual tests using the mobile device. The mobile device transmits the test results to a remote server implementing analysis of the visual function results using network service. The network service receives the test results, processes the results, and provides the processed results to a healthcare provider. The processed results may include trends of the user's visual function test performance. The healthcare provider, such as a physician, may optimize and administer treatment based on the data. Early detection of changes in visual function can enable the healthcare provider to individualize treatment, helping to prevent vision loss while minimizing visits to the office, discomfort, and expense. | 1. A method for performing analysis of a visual function of a user with a retinal disease who is receiving recurrent treatment from a healthcare provider, comprising: receiving the user's visual function test data from a remote mobile device application;executing a test data analysis module stored in memory by a processor, the executed test data analysis module processing the visual function test data to generate a trend data of visual function and analyzing the trend data of visual function based on the retinal disease of the user to determine the user's response to treatment, need for treatment, or need for evaluation by a healthcare provider; andsending the visual function test data and the trend data to a healthcare provider. 2. The method of claim 1, further comprising: transmitting the predicted time for the next treatment to the remote computing device. 3. The method of claim 1, further comprising: transmitting the visual function test data and the trend data to a remote computing device. 4. The method of claim 1, further comprising: receiving a message from a physician for the user; andtransmitting the message to the remote computing device. 5. The method of claim 1, further comprising: analyzing the trend data of visual function and comparing the trend data to previous trend data for the user. 6. The method of claim 5, further comprising: transmitting an alert to the remote computing device if the trend data of visual function is outside of a range of acceptable variation in comparison to the previous trend data for the user. 7. The method of claim 1, further comprising: comparing the trend data of visual function to trend data aggregated for patients with a similar retinal disease to the user. 8. The method of claim 1, further comprising making visual function test data and the trend data available to the healthcare provider. 9. The method of claim 1, further comprising sending an alert to the healthcare provider if the change in visual function data is above a level specified by the healthcare provider. 10. The method of claim 1, wherein the visual function test data corresponds to tests that were pre-selected for the user by the user's healthcare provider and performed by the user using the remote mobile device application. 11. The method of claim 10, wherein the one or more visual function tests are pre-selected by a healthcare provider based on analysis of retinal imaging of the user. 12. The method of claim 10, wherein the test is pre-selected by the healthcare provider based on the retinal disease or the recurrent treatment. 13. The method of claim 12, wherein the recurrent treatment comprises monitoring disease progression or administration of a pharmacological agent. 14. The method of claim 1, wherein the healthcare provider is a doctor, nurse, medical professional, or medical technician. 15. The method of claim 1, wherein the recurrent treatment includes an intraocular injection of a pharmacological agent. 16. Th... | The Board of Trustees of the Leland Stanford Junior University | A61B003/00 A61B003/024 A61B005/00 A61B003/02 A61B003/06 G06Q050/22 A61B003/18 A61B003/032 G06F019/00
US10709978 | B2 | 2020-07-14 | Storage medium having stored thereon information processing program, and information processing device | 2015-09-03 | An example system includes a physiological sensor for generating physiological data associated with a user bearing the physiological sensor. The physiological sensor is configured to wirelessly communicate the physiological data. A portable device includes a touch input device; an accelerometer for generating accelerometer data indicative of activity of the user; wireless communication circuitry for receiving the physiological data; memory for storing the accelerometer data and the physiological sensor data; and a vibrator for providing tactile output to the user. The wireless communication circuitry transmits the physiological data and the accelerometer data to a computer device for use in a presentation application. | 1. An activity monitoring system comprising: a portable device comprising: a physiological sensor for generating physiological data associated with a user;a touch input device;an accelerometer for generating accelerometer data indicative of activity of the user;memory for storing the accelerometer data and the physiological data; andwireless communication circuitry for transmitting the physiological data and the accelerometer data; anda computer device comprising: wireless communication circuitry for receiving the accelerometer data and the physiological data communicated by the portable device; andprocessing circuitry configured to generate image data, the generated image data including first image data for a first object controlled based on one or more of the received accelerometer data and the physiological data, and second image data, different from the first image data, for information related to an assessment of control of the first object. 2. The activity monitoring system according to claim 1, wherein the physiological sensor comprises an optical sensor. 3. The activity monitoring system according to claim 1, wherein the physiological sensor generates physiological data associated with heart rate. 4. The activity monitoring system according to claim 1, wherein the accelerometer comprises a triaxial accelerometer. 5. The activity monitoring system according to claim 1, wherein the wireless communication circuitry of the portable device and the computer device are configured for Bluetooth communication. 6. The activity monitoring system according to claim 1, wherein the computer device further comprises a memory and the processing circuitry is configured to store the received physiological data as a history in the memory. 7. The activity monitoring system according to claim 1, wherein the first object comprises an animated character corresponding to the user. 8. The activity monitoring system according to claim 1, wherein the portable device further comprises a vibrator. 9. The activity monitoring system according to claim 1, wherein the portable device further comprises an operation button. 10. The activity monitoring system according to claim 1, wherein the portable device further comprises a connector for a wired connection. 11. The activity monitoring system according to claim 1, wherein the portable device further comprises: a vibrator;an operation button; anda connector for a wired connection. 12. The activity monitoring system according to claim 1, wherein aspects of movement of the first object are controlled based on the one or more of the received accelerometer data and the physiological data. 13. The activity monitoring system according to claim 1, wherein the processing circuitry is further configured to generate third image data for a second object and the information related to the assessment of the control of the first object is based on a positional relationship between the first object and the second object. 14... | NINTENDO CO., LTD. TOHOKU UNIVERSITY | A61B005/00 A63F013/428 A63F013/211 A61B005/024 A61M021/02 A63F013/40 A63F013/20 A61B005/0205 A61B005/026 A61B005/04 A61B005/11 A63F013/212 A61B005/021 A61B005/0402 A61B005/0488 A61B005/08 A61B005/1455 A61M021/00
US10722643 | B2 | 2020-07-28 | Patch-sized fluid delivery systems and methods | 2018-03-05 | A patch-sized fluid delivery device may include a reusable portion and a disposable portion. The disposable portion may include components that come into contact with the fluid, while the reusable portion may include only components that do not come into contact with the fluid. Redundant systems, such as redundant controllers, power sources, motor actuators, and alarms, may be provided. Alternatively or additionally, certain components can be multi-functional, such a microphones and loudspeakers that may be used for both acoustic volume sensing and for other functions and a coil that may be used as both an inductive coupler for a battery recharger and an antenna for a wireless transceiver. Various types of network interfaces may be provided in order to allow for remote control and monitoring of the device. | 1. A delivery device comprising: a first housing portion;a second housing portion configured to selectively engage with and disengage from the first housing portion;a rotatable member supporting a grouping of coupled projections for movement in an annular path with the rotation of the rotatable member;a resilient member supported by the first housing portion and having a flexible portion and arranged within at least a portion of the annular path of the grouping of coupled projections to be engaged at a single location along the annular path when the second housing portion and the first housing portion are engaged, the flexible portion of the resilient member being resiliently collapsible at the single location of engagement of the grouping of coupled projections, the resilient member being connectable in fluid flow communication with an injection site, wherein the flexible portion of the resilient member is supported on a flat support surface;a reservoir having an interior volume for containing a fluid, the interior volume of the reservoir being in fluid flow communication with the resilient member; anda motor supported by the second housing portion and operatively coupled to the rotatable member for selectively rotating the rotatable member. 2. A delivery device according to claim 1, wherein the rotatable member is supported for rotation by the second housing portion. 3. A delivery device according to claim 1, wherein the rotor comprises a rotary wheel and wherein the grouping of coupled projections is supported by the rotary wheel. 4. A delivery device according to claim 1, wherein the reservoir comprises a collapsible container structure. 5. A delivery device according to claim 1, wherein the flexible portion of the resilient member is supported on a backstop. 6. A delivery device according to claim 5, wherein the reservoir comprises a collapsible container structure. 7. A delivery device according to claim 1, wherein the flexible portion of the resilient member that is arranged within at least a portion of the annular path of the grouping of coupled projections comprises a portion of the resilient member arranged adjacent to the backstop. 8. A delivery device according to claim 1, wherein the grouping of coupled projections provide an intermittent pulse of force at the single location as the rotatable member rotates the grouping of coupled projections along the annular path while the second housing portion and first housing portion are engaged. 9. A delivery device comprising: a rotatable member supporting at least one coupled projection for movement in an annular path with the rotation of the member;a backstop arranged to abut a resilient member;a reservoir;wherein the resilient member is arranged between the grouping of coupled projections and the backstop, the resilient member arranged to flex in one direction to reduce the volume of the reservoir upon movement of the rotatable member associated with the resilient member in a first... | DEKA Products Limited Partnership | A61N001/30 A61M037/00 A61M005/142 A61B005/15 A61B005/151 A61B005/155 A61M005/158 A61M005/148
US10716547 | B2 | 2020-07-21 | Adapter for attaching devices to endoscopes | 2018-11-15 | An improved system for attaching devices to the distal end of endoscopic instruments is described. An attachment adapter comprises a distal stop, such as a complete or partial loop or cylinder, which limits the penetration of an endoscope into the adapter. The adapter has at least one flange protruding proximally along the endoscope. The adapter is secured on the endoscope by a securing means, which compresses the flange or flanges sufficiently to provide a friction fit. The flexibility of the inventive adapter system allows the adapter to work reliably with endoscopes that are used, damaged or repaired. The adapter may further provide a slit or window allowing viewing of the vessel wall by a lens on an endoscope, thereby improving the operator's ability to understand the location of the endoscope during operation. The adapter may carry any of a variety of devices, including devices for closing a surgical incision, or treating or manipulating tissue. Improvements are described in a number of areas to allow more flexible pairing of particular endoscopes or endoscopic instruments with devices to be carried on their exteriors, while maintaining a tightly locked position of the external device on the carrier. | 1. A system for deploying a tissue fastener, the system comprising: an adapter including a flange and a securing element, wherein the flange comprises a longitudinal opening and at least one attachment point protruding from the flange, wherein the flange comprises a membrane covering the longitudinal opening,the adapter being configured for insertion of an endoscopic instrument therein, the longitudinal opening and the membrane being expandable to allow the flange to conform to the endoscopic instrument,the securing element releasably attachable to the at least one attachment point to compress the flange against the endoscopic instrument for locking the adapter to the endoscopic instrument;a deployment assembly carried by the adapter, the deployment assembly including the tissue fastener for deployment from a distal end of the deployment assembly; anda pusher configured so that it can be pushed into the deployment assembly and push the tissue fastener out of the deployment assembly,wherein before the deployment, the tissue fastener is disposed in an annular space of the distal end of the deployment assembly defined by an inner shell, an outer shell, and a connecting wall,after the deployment, the tissue fastener is separated from the adapter. 2. The system of claim 1, wherein the deployment assembly defines an annular space to receive the tissue fastener. 3. The system of claim 1, wherein the deployment assembly and the pusher are carried by the adapter in a manner that does not obstruct visualization through the longitudinal opening. 4. The system of claim 3, wherein the pusher includes a longitudinal opening. 5. The system of claim 4, wherein the pusher is configured to allow visualization of tissue through the longitudinal opening of the pusher and the longitudinal opening of the adapter. 6. The system of claim 1, wherein the deployment assembly is activated by an actuating mechanism operable at a distance from the deployment assembly. 7. The system of claim 6, wherein the actuating mechanism includes an outer sheath connected to the pusher. 8. The system of claim 7, wherein the actuating mechanism further includes a wire extending through the outer sheath, the wire being connected to the deployment assembly. 9. The system of claim 1, further including a distal stop to limit penetration of the endoscopic instrument into the adapter. 10. The system of claim 9, wherein the distal stop comprises a closed circumference. 11. The system of claim 9, wherein the flange projects axially from the distal stop. 12. The system of claim 1, wherein the flange and the pusher are C-shaped in the cross-section. 13. The system of claim 1, wherein the flange is a C-shaped part-circle in profile. 14. A system for deploying a tissue fastener, the system comprising: an adapter configured for insertion of an endoscope instrument therein, the adapter including a flange and at least one protruding attachment point, wherein the adapter comprises a longitudi... | UNITED STATES ENDOSCOPY GROUP, INC. | A61B001/04 A61B017/00 A61B001/00 A61B017/064 A61B017/068 A61B005/00
US10713786 | B2 | 2020-07-14 | Detection of metal stent struts | 2018-06-07 | The disclosure relates to stent detection and shadow detection in the context of intravascular data sets obtained using a probe such as, for example, and optical coherence tomography probe or an intravascular ultrasound probe. | 1. A method of stent strut detection comprising: accessing a plurality of frames of image data obtained with respect to a blood vessel;detecting a lumen boundary in one or more frames;identifying a plurality of shadow regions in the plurality of frames;eliminating shadow regions corresponding to non-stent features from the plurality of identified shadow regions to identify shadow regions that are candidate stent strut shadow regions;scoring the candidate stent strut shadow regions to assign one or more scores to each candidate stent strut; andidentifying one or more stent struts based on the one or more scores of each candidate stent strut. 2. The method of claim 1 wherein the non-stent features are selected from the group consisting of: a guidewire, a side branch, and combinations thereof. 3. The method of claim 1, wherein the plurality of frames of image data are stored in an electronic memory device in electrical communication with an imaging system. 4. The method of claim 1, wherein eliminating shadow regions further comprises eliminating shadow regions that contain lumen pixels. 5. The method of claim 1, wherein the one or more scores are determined based on candidate stent strut's proximity to the lumen boundary, wherein candidate stent strut closest to the lumen boundary receives highest score. 6. The method of claim 1, wherein each candidate stent strut corresponds to a local maxima having a peak strength, wherein the one or more scores are determined based on peak strength. 7. The method of claim 6, further comprising ranking the local maxima based on peak strength to generate a peak score. 8. The method of claim 1, wherein the one or more scores is a malapposition score. 9. The method of claim 8 wherein candidate stent struts having a malapposition score that exceeds a threshold based on distance from lumen boundary are identified as false positive detections. 10. The method of claim 1 wherein the one or more scores are three scores, wherein the three scores are combined to generate a combined score, wherein the combined score is used to identify stent struts from the candidate stent struts. 11. The method of claim 1, wherein the one or more scores are selected from a group consisting of a peak score, a lumen boundary proximity score, and a malapposition score. 12. The method of claim 11 further comprising combining the peak score, the proximity score, and the malapposition score, wherein highest combined score is designated as a stent strut at a particular location in the blood vessel. 13. The method of claim 1 further comprising performing a cross-frame analysis to validate designated stent struts across multiple frames of the plurality of frames. 14. The method of claim 13 further comprising displaying on a graphical user interface the validated stent struts. 15. A computer readable medium comprising non-transitory instructions that when executed cause a processor to perform the steps of claim 1. 16. The method of clai... | LIGHTLAB IMAGING, INC. | G06T007/00 A61B005/00 A61B005/06 G06T007/70 G06K009/46
US10715759 | B2 | 2020-07-14 | Athletic activity heads up display systems and methods | 2017-05-12 | A method of using an athletic activity heads up display system during an athletic activity includes the steps of a heads up display unit receiving information about a sport ball and the heads up display unit displaying an image to an individual based on the information, where the image is overlaid on the individual's present field of view of an environment. | 1. A method of using an athletic activity heads up display system during an athletic activity, comprising: a body-mountable heads up display unit receiving information about a sport ball being used in the athletic activity, wherein the information comprises acceleration data about the sport ball received wirelessly from the sport ball;the heads up display unit taking a still photo or video clip in response to receiving the acceleration data from the sport ball;the heads up display unit further receiving information about a plurality of individuals engaged in the athletic activity; andthe heads up display unit displaying a plurality of images to a user wearing the heads up display unit based on the information about the sport ball and the information about the plurality of individuals,wherein the images are overlaid on the user's present field of view of an environment as viewed through the heads up display unit,wherein a first image includes a strike zone icon overlaid on top of the sport ball in the user's present field of view, andwherein a second image includes a point of impact icon overlaid on top of a ball icon, such that the point of impact icon is representative of a calculated point of impact based in part on the acceleration data received from the sport ball. 2. The method of claim 1, wherein the heads up display unit receiving information about the sport ball comprises the heads up display unit wirelessly receiving additional information transmitted from the sport ball. 3. The method of claim 2, wherein the heads up display unit displaying the image comprises the heads up display unit displaying one of an indication of the sport ball's speed, the sport ball's spin, and the sport ball's flight path. 4. The method of claim 1, further comprising: the heads up display unit receiving video information indicating information indicating the position of the sport ball, wherein the video information is used to trigger the taking of the still photo or the video clip by the heads up display unit. 5. The method of claim 1, wherein the heads up display unit receiving information about the sport ball comprises the heads up display unit receiving audio information indicating information about the contacts between the individuals and the sport ball, and wherein the audio information is used to determine one of the quality and the number of contacts. 6. The method of claim 1, wherein the heads up display unit receiving information about the sport ball comprises the heads up display unit wirelessly receiving information transmitted from the sport ball and the heads up display unit receiving video information providing information about the position of the sport ball. 7. The method of claim 1, wherein the sport ball comprises a soccer ball. 8. A method of using an athletic activity heads up display system during an athletic activity, comprising: a body-mountable heads up display unit identifying an objective for an individual wearing the heads up ... | adidas AG | H04N005/445 G06K009/00 G02B027/01 G06T019/00 A61B005/00 G09B019/00 A61B005/11 A63F013/213 A63F013/211 A63F013/25 A63F013/56 A63F013/212 A63F013/00 A63F013/5255 A63B024/00 A63B043/00 A63B069/00 A63B071/06
US10709379 | B2 | 2020-07-14 | Methods and devices for accurately classifying cardiac activity | 2017-09-14 | Methods, systems, and devices for signal analysis in an implanted cardiac monitoring and treatment device such as an implantable cardioverter defibrillator. In some examples, captured data including detected events is analyzed to identify likely overdetection of cardiac events. In some illustrative examples, when overdetection is identified, data may be modified to correct for overdetection, to reduce the impact of overdetection, or to ignore overdetected data. Several examples emphasize the use of morphology analysis using correlation to static templates and/or inter-event correlation analysis. | 1. An implantable medical device (IMD) comprising a plurality of electrodes coupled to operational circuitry configured to perform cardiac signal analysis, the IMD including a canister having the operational circuitry and at least one electrode associated therewith, the operational circuitry being configured to: detect at least first, second and third detected events by comparison of a cardiac signal from the plurality of electrodes to a threshold defined by a detection profile and declaring a detected event having a signal waveform when the threshold is crossed, wherein the first detected event occurs later in time than the second detected event, and the second detected event occurs later in time than the third detected event;compare the first detected event to the second detected event to yield a first comparison score;compare the first detected event to the third detected event to yield a second comparison score; andanalyze at least the first and second comparison scores to determine if one of the first, second and third detected events is likely due to overdetection. 2. The IMD of claim 1 wherein the operational circuitry is further configured to: detect a fourth detected event by comparison of a cardiac signal from the plurality of electrodes to a threshold defined by a detection profile and declaring a detected event having a signal waveform when the threshold is crossed, wherein the third detected event occurs later in time than the fourth detected event; andcompare the first detected event to the fourth detected event to yield a third comparison score;wherein the operational circuitry is configured such that when it analyzes the first and second comparison scores to determine if one of the first, second and third detected events is likely due to overdetection, the operational circuitry also analyzes the third comparison score. 3. The IMD of claim 1 wherein the operational circuitry is configured to analyze at least the first and second comparison scores to determine if one of the first, second and third detected events is likely due to overdetection by use of at least a first rule set configured to identify double detection including both R-wave and T-wave detection. 4. The IMD of claim 1 wherein the operational circuitry is configured to analyze at least the first and second comparison scores to determine if one of the first, second and third detected events is likely due to overdetection by use of at least a first rule set configured to identify triple detection of a cardiac cycle. 5. The IMD of claim 1 wherein the operational circuitry is configured to further determine, in response to determining that one of the first, second and third detected events is likely due to overdetection, whether the overdetection was of a T- wave. 6. The IMD of claim 1 wherein the operational circuitry is configured to calculate a cardiac rate for a patient using detected events, wherein when the operational circuitry identifies a detected event likel... | CAMERON HEALTH, INC. | A61B005/04 A61B005/00 A61B005/0468 A61B005/0452 A61N001/37 A61N001/365 A61B005/042 A61B005/046 A61B005/0464 A61N001/362 A61N001/39
US10709368 | B2 | 2020-07-14 | Regional oximetry sleeve for mobile device | 2018-01-16 | A system includes a mobile computing device, a sleeve, a contact surface, a sensor module, and a telemetry module. The mobile computing device has a first processor within a housing. The sleeve has an internal surface configured to fit the housing. The contact surface is disposed on an external surface of the sleeve. The sensor module is coupled to the contact surface and is configured to generate an electrical signal corresponding to a measured physiological parameter associated with tissue at the contact surface. The telemetry module is coupled to the sensor module and is configured to communicate data corresponding to the electrical signal to the mobile computing device. | 1. A method comprising: pairing a first communication module of a mobile computing device with a second communication module of a sleeve, the mobile computing device configured to fit an internal surface of the sleeve, the sleeve having a sleeve processor coupled to a sleeve memory and the sleeve having a contact surface, the contact surface having a curved contour;executing instructions using the sleeve processor, the instructions stored on the sleeve memory, the instructions configured to generate a measured signal using an optical sensor coupled to the contact surface, the measured signal corresponding to a tissue at the contact surface;communicating the measured signal from the sleeve to the mobile computing device using the first communication module and using the second communication module; andexecuting instructions using the mobile computing device processor to generate display data for a display of the mobile computing device. 2. The method of claim 1 wherein executing the instructions using the sleeve processor includes calculating a temperature using a temperature transducer. 3. The method of claim 1 wherein executing the instructions using the sleeve processor includes emitting light from the optical sensor. 4. The method of claim 1 wherein executing instructions using the sleeve processor includes calculating regional oximetry. 5. The method of claim 4 further including displaying regional oximetry data based on the display data using the display. | Nonin Medical, Inc. | A61B005/00 A61B005/1455 A61B005/0205 A61B005/145
US10716730 | B2 | 2020-07-21 | Gaze stabilization system and method | 2015-08-04 | A system and method are provided that allow a practitioner to rotate a human subject's head in three-dimensional space using a robotic arm with a head mount attached to one end, while the subject is focused on a visual target. The robotic arm is controlled by a computer system operatively coupled to the various sensors and drive motors embedded with the robotic arm. | 1. A method of treating at least one system of a human subject comprising: arranging the human subject within a human receiving area facing a visual target;securing a head mount to the human subject's head, wherein the head mount is affixed to and actuated by a distal end of a robotic arm, wherein the distal end is opposite of a base of the robotic arm;actuating the head mount to move the human subject's head in three dimensional space while the human subject focuses on the visual target, wherein the robotic arm comprises a sensor system that detects resistance to movement in each of three independent Cartesian axes during said actuating;measuring eye movements and adapting speed or amplitude of said actuating based on said measuring eye movements;compensating for increases in said resistance by decreasing speed or amplitude of said actuating; andengaging a safety cutoff mechanism if said resistance reaches or exceeds a threshold. 2. The method of claim 1 wherein the system is a vestibular system. 3. The method of claim 1 wherein the system is a visual system. 4. The method of claim 1 wherein the system is a proprioceptive system. 5. The method of claim 1 wherein the visual target is stationary during the actuating. 6. The method of claim 1 wherein the visual target is moving during the actuating. 7. The method of claim 1 further comprising the step of collecting data indicative of the position or movement of the human subject's eyes. 8. The method of claim 1 further comprising the step of collecting data indicative of the position or movement of the human subject's head. 9. The method of claim 1 wherein the system is at least one of a vestibular system, a visual system, and a proprioceptive system. 10. The method of claim 1 wherein the human subject is seated during the method. | OVARD, LLC | A61H005/00 A61B003/113 A61H001/02 A61B003/00 A61B005/00 A61B005/0496
US10722147 | B2 | 2020-07-28 | Activation by temperature sensor | 2017-06-05 | A sleep posture alerting method is provided, wherein the sleep posture of a person is detected. The person is automatically alerted by a stimulus generated by an alerting device when the sleep posture is out of a predetermined body posture range. A stimulus variation is effected by the alerting device. | 1. A sleep posture alerting apparatus, comprising: an electronic control unit;a body posture detecting sensor, operatively connected to the control unit, for detecting a posture of a body of a user;a temperature sensor, operatively connected to the control unit, for detecting a body temperature of the user; andan alarming device, operatively connected to the control unit, for generating a posture alarm responsive to the posture of the body detected by the body posture detecting sensor being outside a predefined posture range,wherein the temperature sensor is arranged to: activate the body posture detecting sensor during a normal operational mode responsive to the detected temperature being within a temperature range, the temperature range having a lower temperature and an upper temperature, such that the temperature sensor activates the body posture detecting sensor responsive to the detected temperature being higher than the lower temperature and the temperature sensor activates the body posture detecting sensor responsive to the detected temperature being lower than the upper temperature, andde-activate the body posture detecting sensor responsive to the detected temperature being outside the temperature range, such that the temperature sensor de-activates the body posture detecting sensor responsive to the detected temperature being lower than the lower temperature and the temperature sensor de-activates the body posture detecting sensor responsive to the detected temperature being higher than the upper temperature, and such that the body posture detecting sensor is active only when being worn by the user, and whereinthe alarming device is configured to generate a temperature alarm responsive to the detected temperature being different than a temperature value, the temperature value being within the temperature range. 2. The apparatus according to claim 1, wherein the control unit is programmed to send a posture alarm signal to the alarming device only if the detected temperature is within the temperature range. 3. The apparatus according to claim 1, wherein the temperature sensor is located in the apparatus such that a warm-up time is provided before the detected temperature is within the temperature range. 4. The apparatus according to claim 1, wherein the control unit is programmed not to send a posture alarm signal to the alarming device when the detected temperature is below the temperature range. 5. The apparatus according to claim 1, wherein the control unit is programmed not to send a posture alarm signal to the alarming device when the detected temperature is above the temperature range. 6. The apparatus according to claim 1, wherein the temperature sensor is an integrated temperature sensor. 7. A method for detecting a sleep posture of a user with a system comprising an electronic control unit, a body posture detecting sensor, a temperature sensor, and an alarming device, the method comprising: detecting, with the body posture ... | NIGHTBALANCE B.V. | G08B021/06 A61B005/11 A61F005/56 A61B005/00 A61M021/02 A61M021/00
US10716941 | B2 | 2020-07-21 | System and method for charge-balancing neurostimulator with neural recording | 2018-01-08 | Systems and methods are provided for an implantable neurostimulator system comprising an array of microelectrodes, a pair of current drivers, and a dynamic current allocation network (DCAN). Each current driver in the pair of current drivers is configured to deliver a respective portion of an electrical signal to the array of microelectrodes to deliver the desired electrical stimulation through the array of microelectrodes to the subject. The DCAN is coupled to the pair of current drivers and the array of microelectrodes to selectively electrically connect each of the pair of the current drivers to individual microelectrodes in the array of microelectrodes to deliver the electrical signal to selective collections of less than all microelectrodes in the array of microelectrodes to deliver the desired electrical stimulation. The pair of current drivers and the DCAN are arranged in respective or a common housing configured to be implanted into the subject. | 1. An implantable neurostimulator system comprising: an array of microelectrodes configured to deliver a desired electrical stimulation to a subject;a pair of current drivers electrically coupled to each microelectrode in the microelectrode array, wherein each current driver in the pair of current drivers is configured to deliver a respective portion of an electrical signal to the array of microelectrodes whereby the respective portions, together, form the electrical signal delivered to the array of microelectrodes to deliver the desired electrical stimulation through the array of microelectrodes to the subject;a dynamic current allocation network (DCAN) coupled to the pair of current drivers and the array of microelectrodes to selectively electrically connect each of the pair of the current drivers to individual microelectrodes in the array of microelectrodes to deliver the electrical signal to selective collections of less than all microelectrodes in the array of microelectrodes to deliver the desired electrical stimulation; andwherein the pair of current drivers and the DCAN are arranged in respective or a common housing configured to be implanted into the subject. 2. The system of claim 1 further comprising a controller configured to control the DCAN to selectively electrically connect each of the pair of the current drivers to individual microelectrodes in the array of microelectrodes to cause the desired electrical stimulation delivered to the subject to match a stored waveform. 3. The system of claim 2 wherein the stored waveform is stored in a memory accessible by the controller and rewritable to receive a user-selected waveform. 4. The system of claim 1 wherein the stored waveform has one of a symmetrical biphasic form, an asymmetrical biphasic form, or a pulse train form. 5. The system of claim 1 wherein the controller is configured to control the DCAN to selectively electrically connect each of the pair of the current drivers to individual microelectrodes in the array of microelectrodes to connect one current driver to multiple microelectrodes or couple the pair of current drivers to one microelectrode. 6. The system of claim 1 wherein an output of one or both of the current drivers in the pair of current drives is coupled to an output configured to receive a calibration capacitor. 7. The system of claim 6 wherein the DCAN is configured to couple the pair of current drivers to the calibration capacitor to calibrate each current driver independent of each other and the microelectrodes. 8. The system of claim 1 wherein DCAN is configured to electrically connect a given microelectrode to a residual voltage monitor (RVM) to sense a residual voltage on the given microelectrode upon delivering desired electrical stimulation and further comprising a controller configured to adjust parameters of the electrical signal to reduce the residual voltage. 9. The system of claim 1 wherein the controller is configured to adjust parameters of th... | Regents of the University of Minnesota | A61N001/36 A61N001/02 A61N001/05 A61B005/00 A61N001/372 A61B005/04 A61N001/08 A61B005/1495
US10716630 | B2 | 2020-07-21 | Systems and methods for a robotic-assisted revision procedure | 2017-07-13 | A method for performing a revision surgery using a robotic-assisted surgery system includes determining information related to an interface area between an implant component and a bone, and generating a planned virtual boundary associated with a portion of the interface area to be removed in a representation of the implant and the bone, based at least in part on the information related to the interface area. The method further includes tracking movement in the physical space of a cutting tool such that movement of the cutting tool is correlated with movement of a virtual tool, and providing a constraint on the cutting tool while the cutting tool removes the portion of the interface area. The constraint is based on a relationship between the virtual tool and the planned virtual boundary. The portion of the interface area is removed to remove the implant component from the bone. | 1. A method for performing a revision surgery using a robotic-assisted surgery system, the method comprising: determining, by a processing circuit associated with a computer, information related to an interface area between a primary implant component and a bone on which the primary implant component is implanted;generating, by the processing circuit, a planned virtual boundary corresponding to a portion of the interface area to be removed to detach the primary implant component from the bone based at least in part on the information related to the interface area;tracking, by a navigation system associated with the computer, movement in physical space of a cutting tool such that movement of the cutting tool is correlated with movement of a virtual tool;facilitating removal of the primary implant component from the bone by providing a constraint on the cutting tool while the cutting tool removes the portion of the interface area, the constraint based on a relationship between the virtual tool and the planned virtual boundary;determining, by the processing circuit, a second planned virtual boundary in a representation of the bone representing one or more additional cuts in the bone to prepare the bone to receive a replacement implant; andproviding a constraint on the cutting tool while the cutting tool performs the one or more additional cuts to prepare the bone, the constraint based on a relationship between the virtual tool and the second planned virtual boundary. 2. The method of claim 1, wherein determining information related to an interface area comprises generating models of the bone and the implant component implanted thereon. 3. The method of claim 2, wherein the bone models were created from images obtained in relation to a primary procedure during which the implant component was implanted on the bone. 4. The method of claim 2, wherein the models are generated by at least one modality from a group consisting of: CT, x-ray, fluoroscope, MRI, ultrasound, video camera, and tracked markers. 5. The method of claim 1, wherein determining information related to an interface area comprises digitizing the interface area with a tracked probe. 6. The method of claim 1, further comprising receiving an input for adjusting the virtual boundary relative to the representation of the implant and the bone. 7. The method of claim 1, wherein the virtual boundary is a haptic boundary and wherein providing the constraint comprises providing haptic feedback to the cutting tool. 8. The method of claim 1, wherein the virtual boundary is an autonomous control boundary and wherein providing the constraint comprises autonomously controlling the cutting tool to remain within the control boundary. 9. The method of claim 1, wherein the cutting tool is one or more tools from the group consisting of: a planar saw, a curved saw, and a burr. 10. The method of claim 1, further comprising determining, by the processing circuit, information related to at least one of... | MAKO Surgical Corp. | A61B034/20 A61B034/30 A61B090/00 A61B005/055 A61B006/03 A61B008/08 A61B017/15 A61B017/17 A61B034/10
US10702160 | B2 | 2020-07-07 | Skin assessment using image fusion | 2019-05-02 | Apparatuses and methods are disclosed for assessing the texture of skin using images thereof. In exemplary embodiments, a texture map of an area of skin is generated from a combination of a standard white light image, a parallel-polarized image, and a cross-polarized image of the area of skin. The texture map is then flattened to remove the underlying curvature of the skin. A texture roughness metric is then generated based on the flattened texture map. An image of the texture map and the metric can be displayed to provide visual and alphanumeric representations of the texture of skin, thereby facilitating the comparison of baseline and follow-up images of the skin, such as those taken before and after treatment. | 1. An apparatus comprising: a storage device configured to store instructions;a processor configured to execute instructions stored in the storage device to: obtain a standard image of a skin area;obtain a parallel-polarized image of the skin area;obtain a cross-polarized image of the skin area;generate a texture map from a combination of the standard, parallel-polarized and cross-polarized images;generate at least one of a texture map image or a texture metric from the texture map; andcause the at least one of the texture map, the texture map image, or the texture metric to be displayed, stored, communicated, or processed;an illumination source, wherein the illumination source is configured to illuminate the skin area with unpolarized light or polarized light of a first polarization orientation; andan image capture device configured to capture an image without polarization or with polarization of a second polarization orientation,wherein the first and second polarization orientations are the same for capturing the parallel-polarized image of the skin area and mutually orthogonal for capturing the cross-polarized image of the skin area. 2. The apparatus of claim 1, wherein the combination includes: an intensity of the parallel-polarized image divided by an intensity of the cross-polarized image,anda fraction of an intensity of the standard image. 3. The apparatus of claim 1, wherein the processor is configured to execute instructions to perform an image registration operation on the standard, parallel-polarized and cross-polarized images before generating the texture map. 4. The apparatus of claim 1, wherein the processor is configured to execute instructions to generate a flattened texture map. 5. The apparatus of claim 1, wherein generating the texture metric includes: applying a two-dimensional wavelet transform to the texture map to determine low- and high-frequency components, anddetermining a ratio of an energy of the high-frequency components to an energy of the low-frequency components. 6. The apparatus of claim 1, wherein generating the texture map image includes determining a gray scale value for each of a plurality of pixels of the texture map. 7. The apparatus of claim 1wherein at least one of the first or second polarization orientations is selectable. 8. The apparatus of claim 1, wherein the processor is configured to execute instructions to detect a skin feature by performing a segmentation operation using at least one of the texture map or the texture map image. 9. The apparatus of claim 1, wherein the processor is configured to execute instructions to generate a further image by modifying one or more of the standard, parallel-polarized and cross-polarized images in accordance with the texture map. 10. The apparatus of claim 1, wherein the processor is configured to execute instructions to generate at least one of a treatment evaluation or a treatment suggestion. 11. A method comprising: obtaining a standard image of a s... | Canfield Scientific, Incorporated | G06K009/00 H04N007/18 A61B005/00 G06T007/42 G06T007/30 G06T007/174 H04N005/225 G06T007/00
US10701987 | B2 | 2020-07-07 | Sense-enabled apparel | 2018-08-30 | Systems and methods for integrating and using sensor system in articles of apparel are provided. The system may include an apparel piece sized and dimensioned to be worn on a user and a sensor system integrated with the apparel piece. The sensor system may include at least one stretchable capacitive sensor positioned on the apparel piece so as to measure a biometric parameter of the user when the apparel piece is worn, an illumination system configured to provide illumination in a specified illumination area based on the measured biometric parameter of the user, and a control module to control one or more settings of the illumination system. | 1. A system comprising: an apparel piece sized and dimensioned to be worn on a user; anda sensor system integrated with the apparel piece, the sensor system including: at least one stretchable capacitive sensor positioned on the apparel piece so as to measure a biometric parameter of the user when the apparel piece is worn;an illumination system configured to provide illumination in a specified illumination area based on the measured biometric parameter of the user, wherein the specified illumination area is a projected area external to the system;a control module to control one or more settings of the illumination system;a time/intensity sensor configured to receive an input from the user specifying one or more settings of the illumination system; anda mode sensor configured to receive an input from the user specifying one or more modes of operation of the sensor system. 2. The system of claim 1, wherein the at least one stretchable capacitive sensor includes one or more sensors positioned proximate to a right deltoid when the apparel piece is worn on the user and one or more sensors positioned proximate to a left deltoid when the apparel piece is worn on the user, and the at least one stretchable capacitive sensor is configured to sense a swing of an arm of the user to sense a pace of the user. 3. The system of claim 1, wherein the at least one stretchable capacitive sensor is configured to sense a leg swing of the user to determine a pace of the user. 4. The system of claim 1, wherein the illumination system is configured to provide a pulsed illumination with a pulsation frequency set based on the measured biometric parameter of the user. 5. The system of claim 1, wherein the mode sensor includes a continuous mode component and a pace mode component, each having a respective input. 6. The system of claim 1, wherein the illumination system includes a plurality of LED modules. 7. The system of claim 1, wherein the control module is configured to adjust an intensity of the illumination system and a pulsation frequency of the illumination system. 8. The system of claim 1, wherein the control module is configured to be disconnectable from the sensor system and removable from the apparel piece. 9. The system of claim 1, wherein the control module comprises a hermetically sealed outer shell encasing electronic components therein. 10. The system of claim 1, wherein the illumination system is configured to provide a targeted and controllable illumination area based on a selected mode of operation. 11. The system of claim 1, wherein the illumination system is movable to adjust the specified illumination area. 12. The system of claim 1, wherein the at least one stretchable capacitive sensor is configured to detect a strain based on a type of biometric movement of the user. 13. The system of claim 1, wherein the control module is configured to connect to a computing device and to transfer data there between. 14. A system comprising: an appare... | NIKE, Inc. | A41D001/00 A41D001/04 A41D003/00 A41B001/08 A41D001/089 A41D001/06 F21V033/00 F21V023/04 F21V023/00 A61B005/11 A61B005/00 A41D027/08 G01L001/14 A41D013/01 H05B045/10 H05B047/105 A63B071/06 F21Y115/10
US10708550 | B2 | 2020-07-07 | Monitoring camera and mount | 2017-05-07 | A video monitoring system (20) includes a camera head (22, 62), including an infrared illumination source (76) and an image sensor (74). A mount (64, 102, 112) is configured to hold the camera head in a fixed location and orientation above a crib (24), so that the image sensor captures images of the crib and an intervention region (52) adjacent to the crib from a fixed perspective. | 1. A video monitoring system, comprising: a camera head, comprising an infrared illumination source and an image sensor, wherein the camera head is configured to transmit video signals generated by the image sensor over a network;a mount, which is configured to hold the camera head in a fixed location and orientation above a crib, so that the image sensor captures images of the crib and an intervention region adjacent to the crib from a fixed perspective; anda server, which is configured to receive and analyze the video signals so as to extract and provide behavioral information regarding sleep patterns of an infant in the crib and regarding actions taken by a caregiver in the intervention region. 2. The system according to claim 1, wherein the mount is selected from a set of mounts including a base having multiple legs for floor mounting and a flat base for at least one of shelf mounting and wall mounting, and the camera head is configured to be attached to and detached from any of the mounts in the set. 3. The system according to claim 2, wherein the camera head comprises a receptacle configured to accept a plug on the mount, and wherein the plug on each of the mounts has a distinguishing feature that fits into the receptacle and enables the camera head to identify the mount to which the camera head is attached. 4. The system according to claim 1, wherein the mount comprises a tripod base comprising three legs, wherein one of the legs is rotatable or removable in order to transform the mount from a freestanding configuration to a wall-supported configuration. 5. The system according to claim 1, wherein the mount is configured to stand between the crib and a wall adjacent to the crib and comprises at least one overmolded set screw, which is configured both to enable rotation and locking of the mount and to serve as a bumper against the wall. 6. The system according to claim 1, wherein the infrared illumination source directs infrared illumination toward the crib from a lower side of the camera head. 7. The system according to claim 6, wherein the camera head comprises a night light, which is configured to emit visible illumination from an upper side of the camera head. 8. The system according to claim 1, wherein the camera head is configured to transmit the video signals over a local network in data packets that are addressed so that the video signals are forwarded both to a local client device on the local network and via a public network to a remote server, from which video images are available to a remote client device. 9. Sleep monitoring apparatus, comprising: a memory, which is configured to receive and store a stream of images captured by a camera mounted over a crib in which an infant is put to sleep, wherein the images include an intervention region adjacent to the crib; anda processor, which is configured to analyze the stream of the images so as to detect and log events associated with sleep behavior of the infant, wherein at ... | UDISENSE INC. | H04N007/18 H04N005/225 G06K009/20 G06K009/62 A61B005/11 G06K009/00 H04N005/232 A61B005/00 H04N005/33
US10706541 | B2 | 2020-07-07 | System and method for estimating a quantity of a blood component in a fluid canister | 2019-04-23 | A variation of a method for estimating a quantity of a blood component in a fluid canister includes: within an image of a canister, identifying a reference marker on the canister; selecting an area of the image based on the reference marker; correlating a portion of the selected area with a fluid level within the canister; estimating a volume of fluid within the canister based on the fluid level; extracting a feature from the selected area; correlating the extracted featured with a concentration of a blood component within the canister; and estimating a quantity of the blood component within the canister based on the estimated volume and the concentration of the blood component within the canister. | 1. A method for assessment of blood within a canister, comprising: at a processor, selecting an area of an image of the canister upon detection of a reference marker on the canister;at the processor, determining a fluid level within the canister based upon the area within the image;at the processor, generating an estimated volume of blood within the canister based on the fluid level;at the processor, extracting a feature from at least one pixel of the area;at the processor, correlating the feature with a concentration, of a blood component within the canister;at the processor, estimating a quantity of the blood component within the canister based on the estimated volume and the concentration of the blood component within the canister; andby way of the processor, rendering an analysis indicative of the quantity of the blood component and an estimated total patient blood loss, based upon the estimated volume of blood within the canister. | Gauss Surgical, Inc. | A61B005/00 G06T007/00 G01N021/90 G06T007/70 G06T007/90 G01N021/25 G01N021/84 G01F023/292 G01F022/00 G01N033/49
US10722150 | B2 | 2020-07-28 | System for training spatial cognition, memory and spatiomotor coordination through drawing | 2019-06-02 | A method for improving subject's spatial cognition and spatial memory includes presenting spatial information to the person through non-visual means, parcellating the image to enhance understanding of the nature and relationships of the image components to augment image memory, drawing the image, detecting and encoding the errors in the drawing, and providing interactive feedback based on the individual's drawing performance. The methodology includes audio-haptic feedback to present the spatial image information non-visually using an audible signal that varies as a function of the distance of the finger/stylus from the regions of a display that express the image features. | 1. A method for improving a subject's spatial cognition and spatial memory, including: presenting spatial image information to the person through a non-visual medium; non-visual exploration and memorization of the spatial image to enhance the subject's understanding of the nature and relationships of components of the spatial image information to promote augmentation of the subject's spatial image memory;the subject thereafter drawing an image representation of the spatial image information;detecting and encoding errors in the image representation relative to the original spatial image; and,providing interactive feedback to the subject based on the subject's image representationwherein the step of providing interactive feedback includes identifying discrepancies between the explored spatial image information and the drawn image representation. including at least one of a group of factors including: the overall scale of the drawing, global shape distortions, the omission of notable features, the strategy of sequencing the various lines in the drawing the re-positioning of the drawing touch after being lifted from the drawing surface. the shapes of local image features, analyses in geometric or semantic nature. and strategies for global and local attention. 2. The method of claim 1, wherein the step of image exploration and memorization includes geometric parcellation of the spatial image information to direct attention to geometric relationships of components of the spatial image information. 3. The method of claim 1, wherein said step of image exploration and memorization includes semantic parcellation of the spatial image information to examine semantic relationships of components of the spatial image information. 4. The method of claim 3, wherein said step of semantic parcellation includes directing the subject to identify the morphology and image-wise identity of components of the spatial image information, including deconstructing the image-wise component into commonly identifiable elements. 5. The method of claim 4, further including geometric parcellation of the spatial image information, including directing the subject to analyze the spatial image information by at least one of the following tasks: identify line segments, identify in flection points, determine and describe spatial characteristics, both absolute and relative to adjacent line segments, segment-by-segment evaluation of line features, including segment orientation, length, angular relationships of segments, and curvature of line segments. 6. The method of claim 5, wherein said step of image parcellation includes integration of its subcomponents of geometric and semantic parcellation in an interwoven sequence of steps. 7. The method of claim 1, wherein the step of presenting spatial image information includes providing a tangible sheet of material with raised lines representing the lines in a line drawing, said raised lines being perceptible by finger touch. 8. The metho... | - | A61B005/11 G09B007/02 G09B023/28 A61B005/16 A61B005/0484 G16H050/30 G09B019/00 G16H020/70 G16H040/60 G09B021/00
US10702209 | B2 | 2020-07-07 | Graphical user interface for laser eye surgery system | 2013-10-24 | Methods and systems for planning and forming incisions in a cornea, lens capsule, and/or crystalline lens nucleus are disclosed. A method includes measuring spatial dispositions, relative to a laser surgery system, of at least portions of the corneal anterior and posterior surfaces. A spatial disposition of an incision of the cornea is generated based at least in part on the measured corneal anterior and posterior spatial dispositions and at least one corneal incision parameter. A composite image is displayed that includes an image representative of the measured corneal anterior and posterior surfaces and an image representing the corneal incision. | 1. A system for planning and performing a laser surgery procedure on an eye having a cornea, a pupil, and a lens, the cornea having an anterior surface and a posterior surface, the lens being disposed within a lens capsule having an anterior portion and a posterior portion, the system comprising: a treatment lase delivery system including a laser source configured to produce a treatment beam that includes a plurality of laser pulses;an integrated optical system that includes an imaging assembly containing at least one optical lens and operatively coupled to the treatment laser delivery system for the treatment beam such that the imaging assembly and the treatment laser delivery system share at least one common optical element, the integrated optical system being configured to locate at least a portion of the corneal anterior surface and at least a portion of the corneal posterior surface;a patient interface configured to couple the eye with the integrated optical system so as to constrain the position of the eye relative to the integrated optical system;a display device; anda controller operatively coupled with the display device, the laser source, and the integrated optical system; the controller being configured to: generate, relative to the integrated optical system, a spatial disposition of the corneal anterior surface and a spatial disposition of the corneal posterior surface by using the integrated optical system to locate at least portions of the corneal anterior and posterior surfaces;generate a spatial disposition of an incision to be formed in the cornea using the spatial dispositions of the corneal anterior and posterior surfaces and at least one corneal incision parameter, the at least one corneal incision parameter including a first numerical parameter that defines an amount of an uncut anterior region of the cornea and/or a second numerical parameter that defines an amount of an uncut posterior region of the cornea, the incision being a two-dimensional surface, wherein the corneal incision extends partially through the cornea so as to leave the uncut anterior or posterior region of the cornea aligned with one or more cut portions of the corneal incision, the corneal incision and the uncut anterior or posterior region of the cornea defining an access path for a cataract surgery instrument;display a composite image on the display device, the composite image including an image representative of the spatial dispositions of the corneal anterior and posterior surfaces and an image representing the corneal incision, and display the first and/or the second numerical parameter on the display device; andoperate the a treatment lase delivery system to deliver the treatment beam into the eye to form the incision in the cornea. 2. The system of claim 1, wherein the at least one corneal incision parameter includes a corneal incision line density parameter to control amount of overlap between adjacent lines of laser pulse focus points that will b... | AMO Development, LLC | A61N005/06 A61B005/00 A61B003/00 A61B003/10 A61F009/009 A61F009/008 A61F009/007 A61B034/00 A61B090/00 A61B034/10
US10722255 | B2 | 2020-07-28 | Systems and methods for removing obstructive matter from body lumens and treating vascular defects | 2015-06-22 | A catheter system can be used to remove obstructions, deliver implantable devices or substances, and/or restore flow through body lumens. The system can include an outer shaft having a lumen, sidewall, and a longitudinal window in the sidewall, an inner shaft disposed within the lumen, and an expandable member having a first end coupled to the outer shaft and a second end coupled to the inner shaft through the window. The expandable member can be positioned adjacent to a target region while in the collapsed configuration. The expandable member can be expanded to an expanded configuration by relative movement of the outer shaft and the inner shaft. | 1. A method for increasing flow of a body fluid through an obstructed body lumen, the method comprising: inserting a catheter system into the body lumen in a collapsed configuration, the system comprising (i) an outer shaft having a lumen, a sidewall surrounding the lumen, and a window extending longitudinally along a portion of a length of the sidewall and extending through an entire thickness of the sidewall, the window terminating proximal of a distal end of the sidewall, (ii) an inner shaft disposed within the lumen, and (iii) an expandable member having a first end coupled to the outer shaft and a second end coupled to the inner shaft through the window by a slidable collar via an inner shaft to collar bond;positioning the expandable member adjacent to a target region while in the collapsed configuration; andexpanding the expandable member from the collapsed configuration to an expanded configuration by relative movement of the outer shaft and the inner shaft, wherein, during relative movement of the outer shaft and inner shaft, the end of the expandable member remains at a radial position that is even with or outside of a radial position of the sidewall of the outer shaft. 2. The method of claim 1, wherein the expanding comprises proximally withdrawing the inner shaft while maintaining a longitudinal position of the outer shaft relative to the body lumen. 3. The method of claim 1, wherein the positioning comprises positioning the expandable member within a thrombus, and the expanding comprises expanding the expandable member within the thrombus to engage the thrombus. 4. The method of claim 3, further comprising proximally withdrawing the expandable member into a guide catheter. 5. The method of claim 4, wherein the guide catheter comprises a flared distal opening. 6. The method of claim 4, further comprising removing the catheter system along with the guide catheter and the thrombus from the body lumen. 7. The method of claim 1, wherein the expanding comprises compressing obstructive matter in the target region in a manner that improves flow of body fluid through the obstructive matter. 8. The method of claim 1, further comprising applying negative pressure through a second lumen of the inner shaft to aspirate obstructive matter into the second lumen. 9. The method of claim 1, wherein the inner shaft comprises a second lumen, the method further comprising advancing a working device though the second lumen and out of a distal opening of the catheter system. 10. The method of claim 9, further comprising using the working device to perform a therapeutic or diagnostic task. 11. The method of claim 9, wherein the working device comprises an embolectomy device, and the method further comprises using the embolectomy device to move obstructive matter through the distal opening and into the second lumen. 12. The method of claim 9, further comprising injecting a therapeutic substance to the body lumen via the second lumen of the inner sha... | COVIDIEN LP | A61B017/221 A61B017/12 A61M025/00 A61B017/22 A61B005/00 A61B017/34 A61B017/00 A61M027/00 A61B017/3207
US10716629 | B2 | 2020-07-21 | Methods and devices for facilitating visualization in a surgical environment | 2017-02-27 | Devices and methods for visually confirming the positioning of a distal end portion of an illuminating device placed within a patient include inserting a distal end portion of an illuminating device internally into a patient, emitting light from the distal end portion of the illuminating device, observing transillumination resulting from the light emitted from the distal end portion of the illuminating device that occurs on an external surface of the patient, and correlating the location of the observed transillumination on the external surface of the patient with an internal location of the patient that underlies the location of observed transillumination, to confirm positioning of the distal end portion of the illuminating device. | 1. A system comprising: (a) an illuminating guidewire having: (i) a proximal end,(ii) a distal end, and(iii) at least one illumination fiber extending between the proximal and distal ends,wherein the distal end is configured to be inserted through a nostril and into a paranasal sinus of a patient, wherein the at least one illumination fiber is configured to direct light distally through the distal end;(b) a dilation catheter having an expandable dilator, wherein the expandable dilator is movable over the illuminating guidewire for positioning within an opening to the paranasal sinus, wherein the expandable dilator is selectively transitionable between an unexpanded state and an expanded state for expanding the opening; and(c) a connector configured to releasably couple the proximal end of the illuminating guidewire with a light source, wherein the connector is configured to manipulate light transmitted distally therethrough from the light source to the illuminating guidewire, wherein the connector includes a tapering light channel configured to focus light received from the light source and communicate the focused light distally to the at least one illumination fiber of the illuminating guidewire. 2. The system of claim 1, wherein a distal end diameter of the tapering light channel is smaller than a proximal end diameter of the tapering light channel, wherein the distal end diameter is equal to or greater than an inner diameter of the at least one illumination fiber. 3. The system of claim 1, wherein a distal end diameter of the tapering light channel is smaller than a proximal end diameter of the tapering light channel, wherein the distal end diameter is equal to or greater than an outer diameter of the proximal end of the illumination guidewire. 4. The system of claim 1, wherein the connector includes a proximal portion and a distal portion, wherein the tapering light channel is arranged within the proximal portion and tapers distally toward the distal portion. 5. The system of claim 4, wherein the connector further includes a guidewire channel arranged within the distal portion, wherein the guidewire channel communicates proximally with the tapering light channel and is configured to releasably receive the proximal end of the illuminating guidewire. 6. The system of claim 5, wherein the connector further includes a locking mechanism configured to releasably retain the proximal end of the illuminating guidewire within the guidewire channel. 7. The system of claim 6, wherein the locking mechanism further comprises a locking arm, wherein the locking arm includes a first portion configured to extend into the guidewire channel and a second portion configured to extend outwardly from an exterior of the connector, wherein the second portion is retractable to transition the locking arm between a locked position and an unlocked position, wherein in the locked position the locking mechanism is operable to retain the proximal end of the illuminatin... | Acclarent, Inc. | A61B034/20 A61B001/07 G02B006/26 G02B006/24 A61B001/233 A61B005/00 A61M025/09 F21V008/00 A61B001/00 G02B006/42 A61B001/06 A61B005/06 A61B017/24 A61M029/02 G02B023/24 G02B006/122 G02B006/125 A61B018/22 A61M025/01 G02B006/04 A61B090/30
US10702190 | B2 | 2020-07-07 | Method for recognizing user activity and electronic device for the same | 2017-11-01 | Various exemplary embodiments provide an electronic device configured to include one or more sensor modules, at least one memory, a display, and a first processor or second processor operatively coupled to the one or more sensor modules, the at least one memory, and/or the display, wherein the first processor is configured to acquire sensor data from at least one sensor module among the one or more sensor modules, calculate a difference between at least two data points in the acquired sensor data, determine user's activity information based on at least one of a period of the sensor data, a magnitude of the difference, and a change amount of the difference, and deliver the user's activity information to the second processor, and wherein the second processor is configured to display, on the display, a user interface related to the user's activity information delivered from the first processor. Other exemplary embodiments are also possible. | 1. An electronic device comprising: one or more sensor modules;at least one memory;a display; anda first processor,wherein the first processor is operatively coupled to the one or more sensor modules, the at least one memory, and/or the display,wherein the first processor is configured to: acquire sensor data from a first sensor module among the one or more sensor modules,calculate a difference between at least two data points in the acquired sensor data,determine user's activity information based on at least one of a period of the acquired sensor data, a magnitude of the difference, and a change amount of the difference,if the period of the acquired sensor data is regular and if the magnitude of the difference is greater than or equal to a first threshold, drive a second sensor module different from the first sensor module based on the acquired sensor data, anddisplay, on the display, a user interface related to the user's activity information. 2. The electronic device of claim 1, wherein the first processor is further configured to determine whether the period of the acquired sensor data is regular if the magnitude of the difference is greater than or equal to the first threshold. 3. The electronic device of claim 1, wherein the first processor is further configured to determine the user's activity information as a user activity exercise if: the magnitude of the difference is greater than or equal to the first threshold,the period of the acquired sensor data is irregular, andthe change amount of the difference is greater than a second threshold. 4. The electronic device of claim 1, wherein the first processor is further configured to determine the user's activity information as a general exercise if the magnitude of the difference is greater than or equal to the first threshold and if the period of the acquired sensor data is regular. 5. The electronic device of claim 1, wherein the first processor is further configured to determine the user's activity information as a daily-life movement if the magnitude of the difference is less than or equal to the first threshold. 6. The electronic device of claim 1, wherein the first processor is further configured to determine, if the user's activity information is a user activity exercise, whether the user activity exercise is maintained based on an interval between a recent exercise time and a current time. 7. The electronic device of claim 1, wherein the first processor is further configured to determine an exercise intensity based on the change amount of the difference during a specific period of time, andwherein a second processor is configured to display information related to the exercise intensity. 8. The electronic device of claim 1, wherein the first processor is further configured to determine an exercise type for the user's activity information based on the acquired sensor data from the second sensor module. 9. The electronic device of claim 1, wherein the electronic device is an elect... | Samsung Electronics Co., Ltd. | A61B005/11 G16H040/60 G16H020/30 A61B005/00 G08B013/14 G08B021/04
US10720041 | B2 | 2020-07-21 | User reminding method and apparatus, and terminal device | 2019-02-20 | A user reminding method and apparatus, and a mobile terminal includes, obtaining, by a terminal device, a physical status of a user and a first blood pressure value of the user when a time difference between a current time and a preset reminding time is less than a first threshold; determining a first safety level based on the physical status and the first blood pressure value; and sending a first reminding signal in a manner corresponding to the first safety level. | 1. A user reminding method, comprising: obtaining, by a wearable device, a physical status of a user and a first blood pressure value of the user when a time length between a current time and a preset reminding time is less than a first threshold;determining, by the wearable device, a first safety level based on the physical status and the first blood pressure value; andsending, by the wearable device, a first reminder corresponding to the first safety level. 2. The user reminding method of claim 1, wherein the physical status is a sleeping state and the first reminder is an awakening reminder. 3. The user reminding method of claim 2, further comprising: obtaining, by the wearable device, a second blood pressure value of the user when the physical status of the user changes from the sleeping state to an awake state;determining, by the wearable device, a second safety level based on a difference between the first blood pressure value and the second blood pressure value; andsending, by the wearable device, a second reminder corresponding to the second safety level. 4. The user reminding method of claim 2, further comprising: obtaining, by the wearable device, a second blood pressure value of the user when the physical status changes from the sleeping state to an awake state;determining, by the wearable device, a second safety level based on a difference between the first blood pressure value and the second blood pressure value;determining, by the wearable device, a third safety level based on the second blood pressure value;selecting, by the wearable device, a higher safety level from the second safety level and the third safety level; andsending, by the wearable device, a second reminder corresponding to the higher safety level. 5. The user reminding method of claim 2, wherein the physical status is an awake state and the first reminder is a prompt reminder. 6. The user reminding method of claim 3, wherein the first reminder is different from the second reminder. 7. The user reminding method of claim 5, wherein the awakening reminder comprises normal ringing, soft ringing, or vibration, wherein the prompt reminder comprises either a first signal that indicates to the user to get-up normally, a second signal that indicates a time to the user to take medicine, a third signal that indicates to the user to get-up slowly, a fourth signal to prompt still lying that indicates to the user to lay still, or a fourth signal that indicates to the user to get help. 8. The user reminding method of claim 1, wherein the first threshold is greater than or equal to a time length needed to obtain the physical status and the first blood pressure value of the user. 9. A wearable device, comprising: at least one processor; andat least one memory, wherein the at least one memory comprises instructions that when executed by the at least one processor, causes the wearable device to perform, at least, the following:obtain a physical status of a user and a first blo... | HUAWEI TECHNOLOGIES CO., LTD. | G08B021/24 A61M021/00 A61B005/00 A61B005/021 G04G013/02 G16H020/10 A61B005/0205 A61B005/0402
US10702144 | B2 | 2020-07-07 | Systems for determining eye refraction | 2018-12-14 | Mobile computer devices and systems for refraction determination of an eye, for example for objective refraction determination and/or subjective refraction determination, are provided. Here, a display of the mobile computer device can be driven to display an image for refraction determination. | 1. A system for determining subjective refraction of the eyes of a person to facilitate selection of eye glass lenses or contact lenses by said person, the system comprising: a mount to be worn on a head or held thereagainst;a mobile computer device provided in said mount;said mobile computer device having a display, a processor, and a non-transitory computer-readable medium all contained within said device;said non-transitory computer-readable medium having a program code stored therein;said program code, when executed on the processor, causing said processor to drive said display to display structurally illuminated images thereon in a common plane defined by said display for viewing by said person to determine subjective refraction of the left and right eyes of said person;said mount including exchangeable viewing optics for viewing said display of the mobile computer device with a left eye and a right eye by said person;said viewing optics defining mutually adjacent optical viewing paths to said display for the left and right eyes, respectively, of said person;said exchangeable viewing optics including at least:first optical units having a first refractive power and being placeable in said mount for imaging the structurally illuminated images on said display to the left eye and the right eye, respectively, in corresponding ones of said beam paths; and,second optical units having a second refractive power different from said first refractive power and being exchangeable with said first optical units in said mount in corresponding ones of said beam paths for imaging said images concerning the left eye and the right eye, respectively, on said display to the left eye and the right eye to facilitate determination of said subjective refraction by receiving an indication from said person as to whether an impression of the images displayed on said display changes in sharpness to become sharper or less sharp in response to the exchange of said second optical units for said first optical units and so arrive at eye glass lenses or contact lenses having refractive properties required by said person to best view the images. 2. The system of claim 1, wherein said mount is configured for receiving changeable lenses, and/or changeable lenses are received in the changeable mount. 3. The system of claim 1 further comprising a combination of polarizer and analyzer for at least one eye. 4. The system of claim 1, wherein said viewing optics include a microlens arrangement. 5. The system of claim 1, wherein said viewing optics include a microlens film to be fitted on said display of said mobile computer device. 6. The system of claim 1, wherein said viewing optics include at least one of a color filter, a red filter, and an infrared filter. 7. The system of claim 1, further comprising a beam splitter arranged between the display of the mobile computer device and at least one eye, wherein the beam splitter is arranged to direct light emerging from the eye to a... | Carl Zeiss AG Carl Zeiss Vision International GmbH | A61B003/10 A61B003/14 A61B003/00 A61B003/032 G02C007/02 A61B003/103 A61B003/04 A61B003/06 A61B003/11 A61B005/00 A61B003/113
US10702191 | B2 | 2020-07-07 | Determination of a currently treated body portion of a user | 2016-08-24 | An apparatus for determining a body portion of a user treated by the user using a personal hygienic device. The apparatus includes a camera configured to capture the user to obtain a pictorial representation of the user while treating the body portion using the personal hygienic device; an interface configured to receive sensor data from at least one inertial sensor residing in the personal hygienic device; and an analyzer configured to analyze the pictorial representation and the sensor data to determine the body portion. | 1. An apparatus for determining a body portion of a user treated by the user using a personal hygienic device, the apparatus comprising: a camera configured to capture the user to obtain a pictorial representation of the user while treating the body portion using the personal hygienic device;an interface configured to receive sensor data from at least one inertial sensor residing in the personal hygienic device; andan analyzer configured to analyze the pictorial representation using a machine-learning method based on normalized face region samples with known or annotated sectors and to combine resulting pictorial data with the sensor data to determine the body portion. 2. The apparatus in accordance with claim 1, wherein the analyzer is configured to perform the determination by selecting the body portion out of a predetermined set of candidate body portions. 3. The apparatus in accordance with claim 2, wherein the set of candidate body portions at least comprises: a first candidate head portion lying at the user's left hand side;a second candidate head portion lying at the user's left hand side, but being displaced relative to the first candidate head portion along the user's vertical axis;a third candidate head portion lying at the user's right hand side;a fourth candidate head portion lying at the user's right hand side, but being displaced relative to the third candidate head portion along the user's vertical axis. 4. The apparatus in accordance with claim 2, wherein the personal hygienic device is a toothbrush and the set of candidate body portions at least comprises an upper jaw left side portion of the user's dentition,a lower jaw left side portion of the user's dentition,an upper jaw right side portion of the user's dentition, anda lower jaw right side portion of the user's dentition. 5. The apparatus in accordance with claim 1, wherein the pictorial representation comprises one or more pictures, and the analyzer is configured to associate a time-aligned portion of the sensor data to each of the one or more pictures to obtain a time-aligned mixed pictorial/acceleration data and determine the body portion based on the time-aligned mixed pictorial/acceleration data. 6. The apparatus in accordance with claim 1, wherein the analyzer is configured to subject the pictorial representation to a first evaluation analysis to obtain a first probability value for each of a first set of candidate body portions indicating how probable the body portion is the respective candidate body portion of the first set of candidate body portions,subject the sensor data to a second evaluation analysis to obtain a second probability value for each candidate body portion of a second set of candidate body portions indicating how probable the body portion is the respective candidate body portion of the second set of candidate body portions, andselect the body portion out of a third set of candidate body portions on the basis of the first probability values and th... | BRAUN GMBH | A61B005/11 A46B015/00 G06K009/62
US10716496 | B2 | 2020-07-21 | Determination of breathing signal from thermal images | 2015-04-28 | A data processing method performed by a computer for determining breathing signal data which represents a breathing cycle of a patient, comprising the steps of: —acquiring image data representing a sequence of training thermal images of at least a part of the surface of the patient's body over time, the sequence covering at least one half breathing cycle and being captured by a thermographic camera; and —tracking at least one tracked point in the image data over the sequence of training thermal images to find a trajectory of the tracked point as the breathing signal data, wherein the tracked point is a point on the surface of the patient's body. | 1. A method implemented by a computer having one or more processors, comprising: acquiring image data representing a sequence of training thermal images of at least a part of a surface of a patient's body over time, the sequence covering at least one half breathing cycle and being captured by a thermographic camera;tracking at least one tracked point in the image data over the sequence of training thermal images to find a trajectory of the at least one tracked point as breathing signal data, wherein the at least one tracked point is a point on the surface of the patient's body; anddefining a gating region in a live thermal image captured by the thermographic camera, wherein the gating region is an area of the live thermal image and wherein at least a part of the trajectory lies in the gating region;finding a position of the at least one tracked point in the live thermal image; andgenerating a gating signal indicating whether or not the at least one tracked point is within the gating region. 2. The method of claim 1, further comprising the steps of defining a corridor around the trajectory in a live thermal image captured by the thermographic camera, finding the position of the at least one tracked point in the live thermal image and outputting a movement warning signal if the position of the at least one tracked point is outside the corridor. 3. The method of claim 1, further comprising the steps of determining the speed of the at least one tracked point from at least two consecutive live thermal images captured by the thermographic camera, comparing the determined speed with the speed of the at least one tracked point at a corresponding position of the trajectory and outputting a speed warning if the difference of the speeds is above a predetermined threshold. 4. The method of claim 1, further comprising a step of a dimension reduction of the trajectory into one dimension. 5. The method of claim 4, wherein the dimension reduction step includes determining a main axis of the trajectory and projecting the trajectory onto the main axis. 6. The method of claim 1, wherein the at least one tracked point is selected automatically. 7. The method of claim 6, wherein the at least one tracked point is selected automatically within a user-defined area of the thermal image. 8. The method of claim 1, wherein a plurality of tracked points are tracked in the sequence of training thermal images to find trajectories of each of the tracked points. 9. The method of claim 8, wherein the plurality of tracked points comprises only tracked points with similar main axes of their trajectories. 10. The method of claim 8, further comprising the steps of defining a gating region for each of the plurality of tracked points in a live thermal image captured by the thermographic camera, finding the positions of each of the tracked points in the live thermal image and generating a gating signal indicating whether or not all or the majority of the tracked points are with... | BRAINLAB AG | A61B005/00 A61B005/113 A61B005/01 A61B005/11 H04N013/204 A61N005/10 G01J005/00 G03B035/02 G06T007/33 G06T007/70 A61B005/055 A61B006/03 A61B006/00 G06T007/00 H04N013/00 H04N013/239 G01J005/10 G01R033/48
US10722689 | B2 | 2020-07-28 | Guidewire with integral radiopaque markers | 2016-05-09 | Embodiments hereof relate to a guidewire formed from an elongated shaft, at least a portion of the shaft having an outer layer, a plurality of channels formed through the outer layer, and an inner core. The outer layer is formed from a material non susceptible to erosion by an erosion agent and the inner core is formed from a radiopaque material susceptible to erosion by the erosion agent. When exposed to the erosion agent, core material adjacent to the channels is removed to form a pattern of integral radiopaque segments or markers with a plurality of voids therebetween. By controlling the location of channels and the rate of erosion of the core material, the pattern of integral radiopaque segments and voids allow for in situ measurement when viewed under fluoroscopy. | 1. A guidewire comprising: an elongated shaft having an outer layer which is disposed over and encloses an innermost core, an inner surface of the outer layer abutting against an outer surface of the innermost core for an entire length of the innermost core, wherein the innermost core is formed from a radiopaque material susceptible to erosion by an erosion agent and the outer layer is formed from a material not susceptible to erosion by the erosion agent,wherein the innermost core is solid and continuous along a proximal portion of the shaft, andwherein the innermost core is non-continuous along a marker portion of the shaft and includes a plurality of segments of the radiopaque material with a plurality of voids therebetween such that each void extends between two adjacent segments of the radiopaque material, the marker portion being distal to the proximal portion of the shaft. 2. The guidewire of claim 1, wherein the marker portion includes a plurality of channels formed through the outer layer and the plurality of voids are adjacent to the plurality of channels. 3. The guidewire of claim 1, wherein the innermost core is solid and continuous along a distal tip portion of the shaft, the distal tip portion being distal to the marker portion. 4. The guidewire of claim 3, wherein an outer surface of the distal tip portion is tapered. 5. The guidewire of claim 3, wherein the proximal portion is the least flexible portion of the guidewire and the distal tip portion is the most flexible portion of the guidewire. 6. The guidewire of claim 1, wherein the radiopaque material is selected from the group consisting of tungsten, tantalum, platinum, platinum and iridium, rhenium, and gold. 7. The guidewire of claim 1, wherein each segment of the radiopaque material is between 20% and 150% of a length of each void. 8. A guidewire formed from an elongated shaft, at least a portion of the shaft comprising: an outer layer formed from a material non susceptible to erosion by an erosion agent; andan innermost core formed from a radiopaque material susceptible to erosion by the erosion agent, the outer layer being disposed over and enclosing the innermost core and an inner surface of the outer layer abutting against an outer surface of the innermost core for an entire length of the innermost core, wherein at least a portion of the innermost core is non-continuous and includes a series of a plurality of radiopaque markers with a plurality of voids therebetween such that each void extends between two adjacent radiopaque markers. 9. The guidewire of claim 8, wherein a portion of the innermost core is solid and continuous along a distal tip portion of the shaft, the distal tip portion being distal to the radiopaque markers. 10. The guidewire of claim 9, wherein the distal tip portion is the most flexible portion of the guidewire. 11. The guidewire of claim 8, wherein a portion of the innermost core is solid and continuous along a proximal portion of the shaft... | Medtronic Vascular, Inc. | A61B005/00 A61M025/09 A61L031/02 A61L031/14 A61L031/18 C23F004/00
US10702697 | B2 | 2020-07-07 | Device for deep electrical and optical brain stimulation | 2016-05-12 | A deep brain stimulation device includes a probe having a transparent distal portion, crossed by a channel open at a proximal end of the probe, and also includes electrodes affixed to the distal portion of the probe and electrode terminals affixed to a proximal portion of the probe. Further, the device includes an electrical housing including a channel open at its two ends and housing electrical terminals in contact with the terminals of the probe, and an electrical extension lead affixed to the housing and connected to the terminals of the housing. The channel of the housing is open at the other one of its ends and the device further includes a waveguide at least partly housed in the central channel of the probe, all the way to the distal portion of the probe, and a transmitter housing capable of being connected to the waveguide for the injection of a predetermined wavelength. | 1. A deep brain stimulation device comprising: a stimulation probe comprising: a longilineal body having a central channel void of material running therethrough, extending between a distal portion and a proximal end of the probe, the central channel being open at said proximal end and the distal portion of the body being capable of transmitting a predetermined wavelength;electrodes affixed to an outer surface of the body in the distal portion of the probe for electric brain stimulation; andelectrical terminals affixed to an outer surface of the body in the proximal end of the probe and connected to the electrodes for their electric power supply;an electrical housing comprising a channel having electrical terminals housed therein, said channel being open at one of its ends to receive the proximal portion of the probe to place into contact the electrical terminals of the probe and the electrical terminals of the housing; andan electrical extension lead affixed to the electrical housing and connected to the electrical terminals of the housing,wherein the channel of the electrical housing is open at the other one of its ends;andwherein the device further comprises: a waveguide capable of guiding the predetermined wavelength, housed at least partly in the central channel of the probe body, and extending all the way to the distal portion of the probe for the illumination of the brain therethrough; anda transmitter housing capable of being connected to the proximal end of the waveguide for the injection into it of the predetermined wavelength. 2. The deep brain stimulation device of claim 1, wherein the distal portion of the waveguide and the distal portion of the probe have a length in the range from 2 millimeters to 5 millimeters. 3. The deep brain stimulation device of claim 1, wherein the wavelength is in the range from 670 nanometers to 1,070 nanometers, and preferably from 670 to 810 nanometers, and wherein the optical power transmitted by the distal portion of the waveguide is smaller than or equal to 10 milliwatts. 4. The deep brain stimulation device of claim 1, wherein the waveguide is a HCS-type optical fiber having a silica core with a diameter in the range from 125 micrometers to 300 micrometers. 5. An electrical extension lead for a connection with a deep brain stimulation probe, said probe comprising a longilineal body and electrical terminals affixed to an outer surface of a proximal portion of the probe body, said electrical extension lead comprising: an electrical housing comprising: a channel open at a first and a second ends, the channel being capable of receiving the proximal portion of the body of the probe when the latter is introduced through the first end of the channel;electrical terminals housed in the channel of the electrical housing and capable of coming into contact with the probe terminals when the proximal portion thereof is introduced into the channel of the housing through the first end of the channel;and an electr... | COMMISSARIAT A L'ENERGIE ATOMIQUE ET AUX ENERGIES ALTERNATIVES | A61N001/36 A61N001/05 A61N005/06 A61N005/067 A61B005/04 A61B005/00
US10712879 | B2 | 2020-07-14 | Touch capacitance transduced energy harvesting system | 2017-06-15 | Aspects disclosed in the detailed description include a touch capacitance transduced energy harvesting system. The energy harvesting system includes a touch sensing electrode array and energy harvesting circuitry coupled to the touch sensing electrode array. When a movable conductive object (e.g., a human finger) moves toward or away from the touch sensing electrode array, capacitance of the touch sensing electrode array increases and decreases accordingly, thus transducing a direct current (DC) current in the touch sensing electrode array. As such, the energy harvesting circuitry can be configured to harvest electric energy from the DC current to generate and store a DC voltage. By harvesting the electric energy transduced from the kinetic energy of the movable conductive object, it is possible to power a low-power electronic device (e.g., a biosensor) with motions already used for interfacing with the low-power electronic device. | 1. An energy harvesting system comprising: a touch sensing electrode array having a capacitance and configured to: increase the capacitance in response to a movable conductive object being moved toward the touch sensing electrode array; anddecrease the capacitance in response to the movable conductive object being moved away from the touch sensing electrode array; andenergy harvesting circuitry coupled to the touch sensing electrode array and configured to harvest electric energy from a direct current (DC) current transduced by the touch sensing electrode array in response to the capacitance being increased and decreased alternately. 2. The energy harvesting system of claim 1 wherein the energy harvesting circuitry is further configured to generate and store a DC voltage based on the harvested electric energy. 3. The energy harvesting system of claim 1 wherein the touch sensing electrode array is configured to: increase the capacitance when the movable conductive object is in physical contact with the touch sensing electrode array; anddecrease the capacitance when the movable conductive object is out of physical contact with the touch sensing electrode array. 4. The energy harvesting system of claim 1 wherein the touch sensing electrode array is configured to: increase the capacitance when the movable conductive object is separated from the touch sensing electrode array by less than a defined distance; anddecrease the capacitance when the movable conductive object is separated from the touch sensing electrode array by more than the defined distance. 5. The energy harvesting system of claim 1 wherein the touch sensing electrode array is configured to: increase the capacitance in response to a movable human body part being moved toward the touch sensing electrode array; anddecrease the capacitance in response to the movable human body part being moved away from the touch sensing electrode array. 6. The energy harvesting system of claim 1 wherein the touch sensing electrode array is configured to: increase the capacitance in response to a human finger being moved toward the touch sensing electrode array; anddecrease the capacitance in response to the human finger being moved away from the touch sensing electrode array. 7. The energy harvesting system of claim 1 wherein the touch sensing electrode array is configured to: increase the capacitance in response to a conductive stylus held in a human hand being moved toward the touch sensing electrode array; anddecrease the capacitance in response to the conductive stylus being moved away from the touch sensing electrode array. 8. The energy harvesting system of claim 1 wherein the touch sensing electrode array comprises a projected capacitive touch (PCT) touchpad. 9. The energy harvesting system of claim 1 wherein the touch sensing electrode array comprises a plurality of electrodes organized into a plurality of rows and a plurality of columns. 10. The energy harvesting system of claim 1 provide... | The Regents of the University of California | G06F003/00 H02N001/08 G06F003/044 G06F003/041 B60L053/51 H01G005/16 G06F001/3234 A61B005/00
US10719574 | B2 | 2020-07-21 | Calibration of a chest-mounted wireless sensor device for posture and activity detection | 2018-01-08 | A method and system for calibrating a wireless sensor device are disclosed. In a first aspect, the method comprises determining a vertical calibration vector and determining a rotation matrix using the vertical calibration vector to line up native axes of the wireless sensor device with body axes. In a second aspect, a wireless sensor device comprises a processor and a memory device coupled to the processor, wherein the memory device includes an application that, when executed by the processor, causes the processor to determine a vertical calibration vector and to determine a rotation matrix using the vertical calibration vector to line up native axes of the wireless sensor device with body axes. | 1. A method to calibrate a wireless sensor device, the method comprising: transmitting a signal to a relay device;in response to a determination, by the relay device, that the wireless sensor device is within a predetermined proximity from the relay device based on the transmitted signal, communicatively connecting the wireless sensor device to the relay device, and receiving a manual calibration request from the relay device, wherein the determination, by the relay device, that the wireless sensor device is within a predetermined proximity from the relay device is based on: a first determination, by the relay device, of a signal index from the transmitted signal,a second determination, by the relay device, of a threshold value for the signal index,a third determination, by the relay device, that the signal index equals or exceeds the threshold value;determining a vertical calibration vector using a manual calibration; anddetermining a rotation matrix using the determined vertical calibration vector to line up native axes of the wireless sensor device with body axes,wherein the manual calibration request includes one of: an upright manual calibration request, a walking manual calibration request, or a bedridden user manual calibration request,wherein in response to the received manual calibration request is an upright manual calibration request, the determining the vertical calibration vector using the manual calibration includes: detecting a microelectromechanical systems (MEMS) based vertical calibration vector;filtering the MEMS based vertical calibration vector using a lowpass filter;replacing a current upright vertical calibration vector with the filtered MEMS based vertical calibration vector; andtransmitting a status message to the relay device,wherein in response to the received manual calibration request is the walking manual calibration request, the determining the vertical calibration vector using the manual calibration includes: calculating a microelectromechanical systems (MEMS) based vertical calibration vector during a user's walking period;determining whether the walking period meets a minimum footstep number threshold; andtransmitting a status message to the relay device, andwherein in response to the received manual calibration request is the bedridden user manual calibration request, the determining the vertical calibration vector using the manual calibration includes: determining both a supine calibration vector and a leaning calibration vector;calculating a microelectromechanical systems (MEMS) based vertical calibration vector based on the supine calibration vector and the leaning calibration vector; andtransmitting a status message to the relay device. 2. The method of claim 1, wherein the signal index is a received signal strength indicator (RSSI) or a received channel power indicator. 3. The method of claim 1, wherein the status message indicates one of a Manual Calibration Success status or a Manual Calibration Warning... | Vital Connect, Inc. | G06F017/00 G07C001/00 G01P021/00 G04F013/04 A61B005/11 A61B005/00 G06F017/40 G01P015/00 G01D021/00 G06F019/00
US10702128 | B2 | 2020-07-07 | Medical borescopes and related tip assemblies | 2019-07-02 | Medical borescopes and related methods. In some embodiments, a medical borescope may comprise a handle and a tube extending from the handle. The tube may be partially, or wholly, defined by an at least substantially non-conductive material, such as a suitable plastic, ceramic, or other non-metallic material. The borescope may further comprise a tip assembly positioned at a distal end of the tube, which may comprise an image sensor configured to take images through the distal end of the tube. | 1. A medical borescope, comprising: a handle;a tube extending from the handle, wherein the tube comprises a non-conductive material, and wherein the tube is opaque to visible light; anda tip assembly positioned at a distal end of the tube, wherein the tip assembly comprises an image sensor configured to take images through the distal end of the tube. 2. The medical borescope of claim 1, wherein the tip assembly further comprises: a housing;a light source; anda lens assembly. 3. The medical borescope of claim 2, wherein the tip assembly further comprises a printed circuit board, and wherein the printed circuit board is physically separated from the light source. 4. The medical borescope of claim 1, wherein the tube comprises at least one of a plastic and a ceramic material. 5. The medical borescope of claim 4, wherein the tube comprises at least one of a polycarbonate, a polyether ether ketone, and a ceramic material. 6. The medical borescope of claim 1, wherein the tube is wholly defined by a non-conductive material. 7. The medical borescope of claim 1, wherein the tube comprises a shielding tube portion comprising a non-conductive material. 8. The medical borescope of claim 7, wherein the tube further comprises an inner tube portion, and wherein the shielding tube portion extends externally of the inner tube portion. 9. The medical borescope of claim 8, wherein the inner tube portion comprises a conductive material, and wherein the shielding tube portion is configured to reduce electromagnetic interference with signals within the tube. 10. The medical borescope of claim 7, wherein the shielding tube portion extends along only a portion of a length of the tube. 11. The medical borescope of claim 1, further comprising a dongle comprising an image processor configured to receive image data from the image sensor, wherein the dongle is removably coupleable with the handle of the medical borescope such that the dongle can be coupled with a plurality of distinct medical borescopes. 12. A medical borescope, comprising: a handle;a tube extending from the handle and comprising a distal end, wherein the tube comprises a shielding portion configured to reduce electromagnetic interference with signals within the tube, GI wherein the shielding portion comprises a non-conductive material, and wherein the tube is opaque to visible light; andan image sensor configured to capture images through the distal end of the tube. 13. The medical borescope of claim 12, wherein the shielding portion comprises a plastic material. 14. The medical borescope of claim 13, wherein the non-conductive material comprises at least one of a polycarbonate, a polyether ether ketone, and a ceramic material. 15. The medical borescope of claim 12, wherein the shielding portion is positioned concentrically over an inner portion of the tube. 16. The medical borescope of claim 15, wherein the inner portion comprises a conductive material. 17. The medical borescope of claim 12... | Xenocor, Inc. | H04N005/225 A61B001/00 H04N007/18 H04N005/232 A61B005/01 G02B023/24 A61B001/045 A61B005/00
US10718709 | B2 | 2020-07-21 | Device for measuring radiation backscattered by a sample and measurement method using such a device | 2018-12-20 | A device for measuring radiation backscattered by a sample including: at least one light source that is configured to emit a light beam, along an axis of incidence, towards a surface of the sample so as to form, on said surface, an elementary illumination zone;an image sensor for forming an image of the radiation backscattered by the sample when the latter is illuminated by the light source, the image sensor lying in a detection plane;a bundle of optical fibres, extending, along an extension axis, between a proximal surface and a distal surface, the proximal surface being applied against the image sensor, the distal surface being configured to be applied against the surface of the sample; wherein the light source is arranged around the bundle of optical fibres, and wherein the distance between the light source and the bundle of optical fibres is less than 1 mm. | 1. A Device for measuring radiation backscattered by a sample including: at least one light source that is configured to emit a light beam, along an axis of incidence, towards a surface of the sample so as to form, on said surface, an elementary illumination zone;an image sensor for forming an image of the radiation backscattered by the sample when the latter is illuminated by the light source, the image sensor lying in a detection plane;a bundle of optical fibres, extending, along an extension axis, between a proximal surface and a distal surface, the proximal surface being applied against the image sensor, the distal surface being configured to be applied against the surface of the sample;wherein the light source is arranged around the bundle of optical fibres, and wherein the distance between the light source and the bundle of optical fibres is less than 1 mm. 2. The device according to claim 1, wherein the light source is arranged around the bundle of optical fibres, between the image sensor and the sample. 3. The device according to claim 1, wherein the light source is flush with the distal surface. 4. The device according to claim 1, wherein the extension axis is perpendicular to the detection plane. 5. The device according to claim 1, wherein the optical fibres of the bundle of optical fibres: are oriented in parallel to the extension axis;or form a bundle diverging or converging away from/toward the image sensor. 6. The device according to claim 1, including a plurality of light sources, each light source being arranged around the bundle of optical fibres. 7. The device according to claim 1, wherein at least two light sources emit light beams in different spectral bands, respectively. 8. The device according to claim 1, wherein at least one light source is such that the axis of incidence of the light beam produced by the latter intersects the extension axis of the bundle of optical fibres. 9. The device according to claim 1, wherein at least one light source is: a light-emitting diode;or a laser diode;or an end of an illuminating optical fibre, the illuminating optical fibre being coupled to an illuminating source. 10. The device according to claim 1, wherein the image sensor, the bundle of optical fibres and each light source are held securely to one another by means of a holding support. 11. The device according to claim 1, wherein the image sensor includes a matrix-array of pixels, such that the distance, parallel to the detection plane, between at least the light source and the matrix-array of pixels is less than 1 mm or less than 500 μm. 12. A method for estimating a scattering and/or absorption optical property of a sample using a device according to claim 1, the method including: a) illuminating the sample using the light source of the device so as to form, on a surface of the sample, an elementary illumination zone;b) forming, using the image sensor, an image representing radiation backscattered by the sample under the ... | Commissariat a I'energie atomique et aux energies alternatives | G01N021/47 A61B005/00 G02B006/08
US10702339 | B2 | 2020-07-07 | Method and system for patient-specific modeling of blood flow | 2014-05-13 | Embodiments include a system for determining cardiovascular information for a patient. The system may include at least one computer system configured to receive patient-specific data regarding a geometry of the patient's heart, and create a three-dimensional model representing at least a portion of the patient's heart based on the patient-specific data. The at least one computer system may be further configured to create a physics-based model relating to a blood flow characteristic of the patient's heart and determine a fractional flow reserve within the patient's heart based on the three-dimensional model and the physics-based model. | 1. A method for estimating a hemodynamic index, comprising: acquiring one or more vascular images depicting at least a narrowing of a patient's blood vessel;acquiring physiological information of the patient;generating, based on the one or more acquired vascular images and the physiological information of the patient, a combined model of the narrowing of the patient's blood vessel;acquiring, via simulating blood flow through the combined model of the narrowing of the patient's blood vessel, a first value of blood flow information comprising a value of blood velocity, blood pressure, or blood flow rate;computing, via simulating the blood flow through the combined model of the narrowing of the patient's blood vessel, a second value of blood flow information; andestimating a hemodynamic index indicative of vascular function at the narrowing of the patient's blood vessel based at least on the acquired first value of blood flow information, the computed second value of blood flow information, and the physiological information of the patient. 2. The method of claim 1, comprising: segmenting arteries from one or more CT images; andgenerating the combined model based at least on the segmented arteries. 3. The method of claim 2, comprising: identifying an ultrasound volume of interest;generating a set of velocity data for the ultrasound volume of interest; andcomputing the second value of blood flow information using the set of velocity data. 4. The method of claim 1, comprising: segmenting arteries from one or more reconstructed images; andgenerating the combined model based at least on the segmented arteries. 5. The method of claim 4, wherein the one or more reconstructed images are based on data acquired using an interventional imaging modality. 6. The method of claim 1, wherein the combined model further comprises a location of a vessel tree, a percentage of lumen reduction, or a stenosis composition. 7. The method of claim 1, wherein the narrowing of the blood vessel comprises a stenotic lesion. 8. The method of claim 1, wherein estimating the hemodynamic index comprises providing at least the combined model to a computational fluid dynamics model. 9. The method of claim 8, wherein computing the second value of blood flow information further comprises: determining a flow boundary condition used in the computational fluid dynamics model. 10. One or more non-transitory computer-readable media encoding one or more routines, wherein the one or more routines cause acts to be performed comprising: acquiring one or more vascular images depicting at least a narrowing of a patient's blood vessel;acquiring physiological information of the patient;generating, based on the one or more acquired vascular images and the physiological information of the patient, a combined model of the narrowing of the patient's blood vessel;acquiring, via simulating blood flow through the combined model of the narrowing of the patient's blood vessel, a first value of blood... | HeartFlow, Inc. | A61B034/10 G06F017/10 A61B034/00 G16B005/00 G16B045/00 G06T007/70 G06T007/73 G06T007/11 G06T007/12 G06T007/13 G06T007/149 G06T007/62 A61B005/02 A61B006/00 G16H010/60 G16H050/30 G16H050/50 G06F030/20 G06F030/23 A61B005/026 G06G007/60 A61B005/00 A61B005/029 A61B006/03 A61B008/06 A61B008/08 G01R033/563 A61B005/107 G06T007/00 G06T017/00 G06K009/00 G06K009/62 A61B005/021 A61B005/024 A61B005/22 G06K009/46 G06T007/20 G06T011/00 G06T011/20 A61M005/00 G06K009/52 A61B005/055 A61B005/11 G06T017/20 G06T015/10 G06T011/60 G06T007/60 A61B008/02 A61B008/04 G01R033/56 G16H030/20 G16H030/40 G16H010/40 G16H050/70 A61B090/00
US10712199 | B2 | 2020-07-14 | Spectrometric sensing apparatus for detecting one or more types of biometric information in electronic device and method thereof comprising a processor controlling a light emitting unit to selectively output light of a wavelength band corresponding to at least one mode | 2017-09-12 | Disclosed is an electronic device includes a housing, a display within the housing, and exposed through a surface of the housing, a light emitting unit and a light receiving unit within the housing, and a processor electrically coupled with the display, the light emitting unit and the light receiving unit. The light emitting unit includes at least one light source for outputting light of at least one wavelength band. The light receiving unit includes at least one region for receiving the light of the at least one wavelength band. The processor is configured to control a function of the light emitting unit and/or the light receiving unit based on at least one mode. Other embodiments are disclosed. | 1. An electronic device comprising: a housing;a display within the housing, and exposed through a surface of the housing;a light emitting unit and a light receiving unit within the housing; anda processor electrically coupled with the display, the light emitting unit and the light receiving unit,wherein the light emitting unit comprises at least one light source for outputting light of at least one wavelength band,the light receiving unit comprises at least one region for receiving the light of the at least one wavelength band, andthe processor is configured to: control at least one function of the light emitting unit and/or the light receiving unit based on at least one mode, andcontrol the light emitting unit to selectively output light of a wavelength band that corresponds to the at least one mode. 2. The electronic device of claim 1, wherein the light emitting unit comprises a plurality of light sources, each light source outputs light of mutually different wavelength bands, and the processor is further configured to selectively activate at least one of the plurality of light sources based on the at least one mode. 3. The electronic device of claim 1, wherein the processor is further configured to selectively activate at least one region of the light receiving unit based on the at least one mode. 4. The electronic device of claim 1, wherein the at least one mode comprises: a mode for detecting proximity of an object using light of a wavelength band including a wavelength of 950 nanometers (nm),a mode for detecting a user skin moisture using light of wavelength bands including a wavelength of 880 nm or 970 nm,a mode for detecting a user skin melanin using light of wavelength bands including a wavelength of 660 nm or 880 nm, ora mode for detecting a user skin erythema using light of wavelength bands including a wavelength of 568 nm or 880 nm. 5. The electronic device of claim 1, wherein the processor is further configured to: acquire information related with the at least one mode based on the light received by the light receiving unit, anddisplay the acquired information using the display. 6. The electronic device of claim 1, further comprising a communication module, and wherein the processor is further configured to: acquire information related with the at least one mode based on the light received by the light receiving unit, andtransmit the acquired information to another electronic device through the communication module. 7. The electronic device of claim 1, wherein the light receiving unit comprises: one or more first regions for receiving light of a proximity detecting wavelength band,one or more second regions for receiving light of an illuminance detecting wavelength band, and/orone or more third regions for receiving light of a biometric detecting wavelength band. 8. The electronic device of claim 1, wherein in a first mode: the light receiving unit detects light of a first wavelength band for the first mode and light of a seco... | Samsung Electronics Co., Ltd. | G01J003/02 G01J003/10 G01N021/31 G01J003/28 H04M001/725 A61B005/00 G01B011/14 A61B005/024 A61B005/021 A61B005/01 A61B005/145
US10722629 | B2 | 2020-07-28 | Device for emergency treatment of cardiac arrest | 2016-12-02 | A device for providing resuscitation or suspended state through redistribution of cardiac output to increase supply to the brain and heart for a patient. The device comprises an electrically controllable redistribution component in the form of a catheter, attachable to the patient and configured to interact with the patient to provide redistribution of the cardiac output to increase supply to the brain and heart, the redistribution component following a predefined reaction pattern based on an electrical signal, and computer configured to: receive a patient data which identifies physiological and/or anatomical characteristics of the patient; and provide the electrical signal for the redistribution component based on the patient data or a standard response. The device also comprises means for detection of blood vessels and motorized means for introduction of the catheter into an artery. | 1. A device for providing resuscitation or suspended state through redistribution of cardiac output to increase supply to the brain and heart for a patient, the device comprising: an electrically or manually controllable redistribution component in the form of a catheter configured to be attachable to the patient and to interact with the patient to provide redistribution of the cardiac output to increase supply to the brain and heart, the redistribution component following a predefined reaction pattern based on an electrical signal,computer configured to 1) receive a patient data which identifies physiological and/or anatomical characteristics of the patient; and 2) provide the electrical signal for controlling the redistribution component and/or for presenting the physiological and/or anatomical characteristics for a user based on the patient data or a standard response, anda component configured for detection of blood vessels and a motorized component configured for introduction of the catheter into an artery,wherein the component configured for introduction of the redistribution component comprises 1) a component configured for motorized sideways movement of a percutaneous cannula or needle, 2) a component configured for motorized introduction of the percutaneous cannula or needle into an artery via longitudinal axial movement of said cannula or needle, and 3) a component configured for motorized introduction of the catheter into an artery through or across said cannula or needle. 2. The device according to claim 1, wherein the component configured for detection of blood vessels comprises an ultrasound detection component, or an infrared or near-infrared imaging component. 3. A device for providing resuscitation or suspended state through redistribution of cardiac output to increase supply to the brain and heart for a patient, the device comprising: an electrically or manually controllable redistribution component in the form of a catheter configured to be attachable to the patient and to interact with the patient to provide redistribution of the cardiac output to increase supply to the brain and heart, the redistribution component following a predefined reaction pattern based on an electrical signal,computer configured to 1) receive a patient data which identifies physiological and/or anatomical characteristics of the patient; and 2) provide the electrical signal for controlling the redistribution component and/or for presenting the physiological and/or anatomical characteristics for a user based on the patient data or a standard response, anda component configured for detection of blood vessels and a motorized component configured for introduction of the catheter into an artery, wherein a cannula or needle can move sideways in a fixed or adjustable angle, preferably along a rail or cable. 4. A device for providing resuscitation or suspended state through redistribution of cardiac output to increase supply to the brain and heart for a pati... | Neurescue Aps | A61M001/10 A61B005/00 A61B017/12 A61M001/12 A61B005/0205 A61B005/06 A61B017/00 A61M005/32 A61M005/42
US10722184 | B2 | 2020-07-28 | Systems and methods for selecting, activating, or selecting and activating transducers | 2017-11-30 | A graphical representation may be displayed including at least a plurality of transducer graphical elements, each transducer graphical element of the plurality of transducer graphical elements representative of a respective transducer of a plurality of transducers of a transducer-based device. A set of user input may be received including an instruction set to reposition a first transducer graphical element in a state in which the first transducer graphical element is located at a first location in the graphical representation and a second transducer graphical element is located at a second location in the graphical representation, the second location closer to a predetermined location in the graphical representation than the first location. In response to conclusion of receipt of the set of user input, the first transducer graphical element may be repositioned from the first location in the graphical representation to the predetermined location in the graphical representation. | 1. A system comprising: a data processing device system;an input-output device system communicatively connected to the data processing device system; anda memory device system communicatively connected to the data processing device system and storing a program executable by the data processing device system, the program comprising:display instructions configured to cause the input-output device system to display a graphical representation including at least a plurality of transducer graphical elements, each transducer graphical element of the plurality of transducer graphical elements representative of a respective transducer of a plurality of transducers of a transducer-based device, the graphical representation including a first spatial relationship between the plurality of transducer graphical elements that is consistent with a second spatial relationship between the plurality of transducers of the transducer-based device;input-processing instructions configured to cause reception of a set of user input via the input-output device system, the set of user input including an instruction set to reposition a first transducer graphical element of the plurality of transducer graphical elements in a state in which the first transducer graphical element is located at a first location in the graphical representation and a second transducer graphical element of the plurality of transducer graphical elements is located at a second location in the graphical representation, the second location closer to a predetermined location in the graphical representation than the first location; andgraphical representation modification instructions configured to cause, in response to conclusion of receipt of the set of user input including the instruction set to reposition the first transducer graphical element, the input-output device system to reposition the first transducer graphical element from the first location in the graphical representation to the predetermined location in the graphical representation. 2. The system of claim 1, wherein the predetermined location is more centrally located in the graphical representation than the first location, and wherein the repositioning of the first transducer graphical element centralizes the first transducer graphical element in the graphical representation. 3. The system of claim 2, wherein the graphical representation modification instructions are configured to cause, in response to the conclusion of receipt of the set of user input including the instruction set to reposition the first transducer graphical element, the input-output device system to reposition the second transducer graphical element from the second location in the graphical representation to a third location in the graphical representation, and wherein the predetermined location is more centrally located in the graphical representation than the third location. 4. The system of claim 3, wherein the predetermined location is in a first direction extend... | KARDIUM INC. | A61B005/00 A61B005/01 A61B005/042 G06F003/0484 A61B005/02 G06F003/0482 G06F003/0481 A61B018/00 A61B005/15 A61B005/157 A61B018/14 A61B034/10 G16H040/63
US10722647 | B2 | 2020-07-28 | Patch-sized fluid delivery systems and methods | 2018-10-15 | A patch-sized fluid delivery device may include a reusable portion and a disposable portion. The disposable portion may include components that come into contact with the fluid, while the reusable portion may include only components that do not come into contact with the fluid. Redundant systems, such as redundant controllers, power sources, motor actuators, and alarms, may be provided. Alternatively or additionally, certain components can be multi-functional, such a microphones and loudspeakers that may be used for both acoustic volume sensing and for other functions and a coil that may be used as both an inductive coupler for a battery recharger and an antenna for a wireless transceiver. Various types of network interfaces may be provided in order to allow for remote control and monitoring of the device. | 1. A delivery device for delivering an infusion medium to a user, the device comprising: a first housing portion adapted to be secured to a user;a second housing portion configured to be selectively engaged with and disengaged from the first housing portion;a drive shaft supported by the second housing portion;a shape-memory alloy operatively connected to the drive shaft, to selectively move the drive shaft; anda pump device comprising: a housing provided with a barrel, a port extending to the barrel and configured for connection to a reservoir, the port in fluid flow communication with the barrel and configured for connection in fluid flow communication with an injection site;a plunger located in the barrel and having a surface for receiving a drive force from the drive shaft when the first housing portion and the second housing portion are engaged,wherein the plunger is moveable with movement of the drive shaft along the longitudinal axis of the barrel between fill and dispense positions, such that:in the fill position, the plunger is moved to a position to allow fluid flow communication through the port and into the barrel; andin the dispense position, the plunger is moved to a position wherein fluid flow is obstructed through the port and to reduce the volume of the barrel to force fluid in the piston chamber out of the port. 2. A delivery device according to claim 1, wherein the housing of the pump device further comprising an outlet chamber in fluid flow communication with the barrel and the port, and wherein the pump device further comprises a one-way valve between the barrel and the outlet chamber and arranged to allow fluid flow from the barrel to the outlet chamber and inhibit fluid flow from the outlet chamber to the barrel. 3. A delivery device according to claim 1, further comprising a reservoir having an interior volume for containing an infusion medium and a conduit connecting the interior volume of the reservoir in fluid flow communication with the port of the housing of the pump device. 4. A delivery device according to claim 3, wherein the reservoir is supported by the first housing portion. 5. A delivery device according to claim 1, further comprising electrical control circuitry contained in the second housing portion, wherein the electrical control circuitry controls the drive device for delivery of the infusion medium from the reservoir to the user when the first housing portion and the second housing portion are engaged. 6. A delivery device according to claim 1, wherein the first housing portion comprises a base portion having a bottom surface and an adhesive material on the bottom surface for securing the first housing portion to the skin of the user. 7. A delivery device according to claim 1, wherein the first housing portion comprises a base portion having a bottom surface securable to the skin of the user, the delivery device further comprising an injection site at which a hollow needle or cannula may be inserted... | DEKA Products Limited Partnership | A61M005/145 A61M005/168 F04B043/02 G05D007/06 F04B043/09 F04B043/12 A61M005/142 A61M005/172 G08C017/02 H04B007/26 G05B023/02 A61M005/158 A61B005/00 A61M005/50 G01F022/00 A61M005/14 A61M005/162 A61M005/36 A61B005/155 A61J001/20
US10722211 | B2 | 2020-07-28 | Acoustic-wave measuring apparatus and method | 2015-08-05 | An acoustic-wave measuring apparatus includes a holding unit configured to hold a subject and an irradiation unit configured to irradiate the subject with light. An acoustic-wave detecting unit detects acoustic waves generated in the subject due to irradiation with the light; an image pickup unit acquires an image of the subject; and a measurement position on the image acquired by the image pickup unit is designated. A coordinate transforming unit transforms the coordinates of the measurement position on the image to coordinates of a corresponding position on the holding unit; and a position control unit moves at least one of the irradiation unit and the acoustic-wave detecting unit to the corresponding position on the holding unit. | 1. An acoustic-wave measuring apparatus comprising: an irradiation unit configured to irradiate a subject with light;an acoustic-wave receiving unit configured to receive acoustic waves generated in the subject due to irradiation with the light;a holding unit configured to hold the subject;an image pickup unit configured to acquire an image of the subject held by the holding unit;a moving unit configured to move the acoustic-wave receiving unit; anda control unit configured to control driving of the moving unit;wherein the control unit is configured to control the driving of the moving unit in such a manner that the acoustic-wave receiving unit is moved to a position corresponding to a measurement position designated in the image acquired by the image pickup unit. 2. The acoustic-wave measuring apparatus according to claim 1, wherein the irradiation unit is configured to irradiate the subject with pulsed light, as the light. 3. The acoustic-wave measuring apparatus according to claim 1, wherein the moving unit is configured to move the acoustic-wave receiving unit and an irradiation region on which the irradiation with the light is performed, andwherein the control unit is configured to control the driving of the moving unit in such a manner that the irradiation region and the acoustic-wave receiving unit are moved to the position corresponding to the measurement position designated in the image acquired by the image pickup unit. 4. The acoustic-wave measuring apparatus according to claim 3, wherein the moving unit is configured to move the irradiation unit and the acoustic-wave receiving unit in a manner such that the irradiation unit and the acoustic-wave receiving unit move in synchronization with each other. 5. The acoustic-wave measuring apparatus according to claim 3, wherein the moving unit is configured to adjust the irradiation region being irradiated, by moving the irradiation unit. 6. The acoustic-wave measuring apparatus according to claim 5, wherein the moving unit is configured to move the irradiation unit so as to enable the image pickup unit to acquire the image of the subject in a state in which the irradiation unit and the subject do not overlap each other. 7. The acoustic-wave measuring apparatus according to claim 5, wherein the moving unit is configured to move the irradiation unit in a first direction parallel to the subject and a second direction perpendicular to the first direction. 8. The acoustic-wave measuring apparatus according to claim 1, further comprising a position designating unit configured to enable an observer to designate the measurement position in the image acquired by the image pickup unit. 9. The acoustic-wave measuring apparatus according to claim 8, wherein the position designating unit includes at least one of a mouse, a keyboard, and a touch-type position detecting sensor. 10. The acoustic-wave measuring apparatus according to claim 1, further comprising a housing configured to prevent the li... | CANON KABUSHIKI KAISHA | A61B005/00 A61B008/08 A61B008/13 A61B008/00 G01N021/17
US10714871 | B2 | 2020-07-14 | Anesthesia apparatus or ventilator with a hot wire sensor, hot wire sensor and hot wire sensor module for a hot wire sensor | 2017-09-19 | A medical device, for example, an anesthesia apparatus or ventilator, including a hot wire sensor (10); a hot wire sensor (10) and a hot wire module (14) for a hot wire sensor (10) are provided. A first hot wire and a second hot wire (26, 28), namely, a measuring wire (26) and a compensation wire (28), are connectable to the hot wire sensor (10), for example, in the form of a hot wire module (14), in an electrically conductive manner. A first contact pair (52, 54) is associated with the measuring wire (26) for contacting same and a second contact pair (56, 58) is associated with the compensation wire (28) for contacting same. The contacts of the second contact pair (56, 58) are configured as leading contacts in relation to at least one of the contacts of the first contact pair (52, 54). | 1. A hot wire sensor comprising: a hot wire sensor circuit comprising a measuring contact pair and a compensation contact pair, the measuring contact pair being disposed in a predetermined position relative to the compensation contact pair; anda hot wire module for electrically and mechanically detachably connecting to the hot wire sensor circuit, via the measuring contact pair and the compensation contact pair, wherein the hot wire module comprises:a first hot wire comprising a measuring wire;a second hot wire comprising a compensation wire; anda chronological contact means for chronologically electrically connecting the compensation wire to the hot wire sensor circuit prior to electrically connecting the measuring wire to hot wire sensor circuit, the chronological contact means comprising a contact pair arrangement comprising measuring wire contacts for an electrically conductive contact connection of the measuring wire to the hot wire sensor circuit in a state of a mechanical connection of the hot wire module to the hot wire sensor circuit and compensation wire contacts for an electrically conductive contact connection of the compensation wire to the hot wire sensor circuit in the state of the mechanical connection of the hot wire module to the hot wire sensor circuit, the chronological contact means chronologically electrically contact connecting the compensation wire contacts to the compensation contact pair prior to electrically contact connecting the measuring wire contacts to the measuring contact pair, as the hot wire module is being mechanically connected to the hot wire sensor circuit to change from a mechanically unconnected state to the state of the mechanical connection. 2. A hot wire sensor according to claim 1, wherein the chronological contact means is formed with the compensation wire contacts configured as leading contacts in relation to at least one of the measuring wire contacts for an electrically conductive connection of the compensation wire to the hot wire sensor circuit in advance of an electrically conductive connection of the measuring wire to the hot wire sensor circuit. 3. A hot wire module for an electrically and mechanically detachable connection to a measuring contact pair and a compensation contact pair of a circuit of a hot wire sensor, the hot wire module comprising: a first hot wire comprising a measuring wire;a second hot wire comprising a compensation wire; anda chronological contact means for chronologically electrically connecting the compensation wire to the circuit of the hot wire sensor prior to connecting the measuring wire to the circuit of the hot wire sensor, the chronological contact means comprising a contact pair arrangement comprising measuring wire contacts for an electrically conductive contact connection of the measuring wire to the circuit of the hot wire sensor, upon completing a mechanical connection of the hot wire module to the circuit of the hot wire sensor, and compensation wire cont... | DRÄGERWERK AG & CO. KGAA | H01R013/642 G01F001/69 G01F001/68 A61B005/087 A61M016/00 A61M016/01 G01F015/04 G01F001/699
US10702174 | B2 | 2020-07-07 | Medical monitor user interface | 2007-06-27 | Methods and devices useful for monitoring a patient and for monitoring and displaying the value of a physiological parameter are disclosed. In one embodiment, a user interface for a medical monitoring device is provided. The user interface can have a monitoring screen with a current value screen and a trend screen. The current value screen can display a graphical representation of a value of a physiological parameter over time, e.g., over a first time period, and the trend screen can display a graphical representation of a mean value of the parameter over time, e.g., over a second time period. In some embodiments, an out-of-limit condition for the parameter can be indicated with shading. In yet other embodiments, the user interface can provide an event marking screen, which can provide the ability to mark events on a historical trend screen that displays a value of the parameter over time. | 1. A method for monitoring a physiological parameter, comprising: receiving data representing a value of a physiological parameter over time with processing circuitry operatively connected to a display, the physiological parameter being measured from a patient and comprising intracranial pressure, the processing circuitry configured to perform the steps of: displaying a current value screen on the display in a first window on a monitoring screen on the display;displaying a trend screen on the display in a second, different window on the monitoring screen that is displaying the current value screen;drawing, on the current value screen, a graphical representation of the value of the physiological parameter over a first time period;determining a mean value of the physiological parameter over time;drawing, on the trend screen, a graphical representation of the mean value of the physiological parameter over a second time period, the first time period having a duration shorter than the second time period;drawing, superimposed on the graphical representation of the mean value on the trend screen, a threshold line indicating a predetermined value;drawing, on the current value screen, a graphical representation of the mean value of the physiological parameter over a third time period in response to a user input that changes the first time period to the third time period; anddrawing, on the trend screen, a graphical representation of the mean value of the physiological parameter over a fourth time period in response to a user input that changes the second time period to the fourth time period. 2. The method of claim 1, wherein the graphical representation drawn on the current value screen comprises a graph line, and further comprising shading a region between the graph line and a threshold line that corresponds to a predetermined threshold to indicate an out-of-limit condition for the physiological parameter. 3. The method of claim 1, wherein the graphical representation drawn on the trend screen comprises a trendline, and further comprising shading a region between the trendline and the threshold line to indicate an out-of-limit condition for the physiological parameter. 4. The method of claim 1, further comprising: comparing a value of the physiological parameter at a point in time to a threshold value to detect an out-of-limit condition, the out-of-limit condition occurring when the value of the physiological parameter exceeds the threshold value,wherein the display of the trend screen occurs only when the out-of-limit condition is detected. 5. The method of claim 1, further comprising: marking the graphical representation of the mean value of the physiological parameter with an event marker disposed at a time at which a medically-related event occurred, the time falling within the second time period. 6. A user interface for a medical monitoring device, comprising: a display;processing circuitry operatively connected to the display and configured t... | Integra Lifesciences Corporation | A61B005/03 A61B005/00 G16H040/63
US10709851 | B2 | 2020-07-14 | Systems, methods and devices for subcutaneous target location | 2016-10-31 | Embodiments of the invention provide devices, systems, and methods that precisely identify a minimum of one predetermined spot which is hidden under a skin. The system comprises a locator device and corresponding implanted target device. The port locator device preferably comprises one magnet with north and south magnetic pole, a body and a suspending component. The body may utilize specific geometry which improves accuracy. The implanted target device may include at least one magnet and at least one target or a plurality of targets and at least one magnet. Various configurations can be provided that precisely identify a single spot or a plurality of spots which are hidden under a skin. | 1. An assembly, comprising: a base including a first infusion port and a target magnet housing;a target magnet disposed in the target magnet housing, the target magnet including a magnetic pole proximal a first side of the target magnet housing and a magnetically opposite pole proximal a second side of the target magnet housing, wherein the target magnet is separated and spaced from the first infusion port;a locator assembly including a locator base and a locator housing, the locator base including a first port aperture proximal a first side of the locator base and a second port aperture proximal a second side of the locator base; andat least two locator magnets respectively disposed at least partially within the locator housing proximal respective sides of the locator housing, the two locator magnets being separated and spaced from the first and second port apertures. 2. The assembly of claim 1, wherein the base is elongated. 3. The assembly of claim 2, wherein the elongated base is ellipsoidal shaped, rectangular shaped, or oblong. 4. The assembly of claim 1, wherein the target magnet housing is disposed proximal a center of the base. 5. The assembly of claim 1, wherein the target magnet housing is elongated. 6. The assembly of claim 1, wherein the base includes a second infusion port disposed at a side of the base opposite to a side of the base including the first infusion port. 7. The assembly of claim 1, wherein the target magnet housing includes at least one attachment aperture. 8. The assembly of claim 7, wherein the attachment aperture includes a channel extending longitudinally along the target magnet housing. 9. The assembly of claim 1, further comprising a suspending guide extending from the locator housing away from the locator base. 10. The assembly of claim 9, further comprising a suspending line coupled to the suspending guide. 11. The assembly of claim 1, wherein the at least two locator magnets are disposed with respect to each other with reversed magnetic poles. 12. A locator system comprising: a target base including a first infusion port proximal a side of the base;a target magnet coupled to the base, the target magnet including a magnetic pole, and the target magnet being separated and spaced from the first infusion port;a locator assembly including a locator base and a locator housing, the locator base including a port aperture proximal a side of the locator base; andat least two locator magnets respectively disposed proximal opposite sides of the locator housing, the locator magnets disposed with respect to each other with reversed magnetic poles, and the locator magnets being separated and spaced from the port aperture. 13. The system of claim 12, further comprising a second infusion port disposed at a side of the locator base opposite to the side of the locator base including the first infusion port. 14. The system of claim 12, further comprising a second port aperture proximal a side of the locator base oppo... | SIENTRA, INC. | H01F007/20 A61M005/42 A61M039/02 A61B005/06 A61M005/142 A61B090/00
US10714212 | B2 | 2020-07-14 | Analyte meter | 2015-11-13 | A portable data-management system based on an analyte testing device which communicates wirelessly with a mobile device. The mobile device runs an application to manage and analyze data obtained by the analyte testing device. The mobile device may assist the user in displaying testing data, identifying patterns to assist healthy behavior or issue warnings based on the collected data. The mobile device may be connected to a network to store user health data for use by other parties. | 1. A system for managing health data, comprising: a testing device operable to measure an analyte in a sample fluid taken from a subject, the testing device including a first wireless transceiver, a testing sensor interface including a port for accepting the insertion of a test sensor, and an illumination panel surrounding the port; andan application stored in a storage device and downloadable via a network connection to a mobile device, the application, when executed by the mobile device, operable to: communicate via a second wireless transceiver on the mobile device to the first wireless transceiver of the testing device to pair the mobile device with the testing device;transmit data relating to the analyte measurement from the testing device via the first and second wireless transceivers to the mobile device; anddisplay the data relating to the analyte measurement on a display of the mobile device, wherein the illumination panel emits light in a first color to highlight the port prior to insertion of the test sensor in the port, emits light in a second color to indicate pairing of the testing device with the mobile device, and emits light in one of a plurality of other colors to communicate data relating to the measurement of the analyte. 2. The system of claim 1, wherein the mobile device includes a network transmitter to send the data to a cloud server. 3. The system of claim 1, wherein the mobile device is operable to receive data from the testing device while the mobile device is paired with the testing device. 4. The system of claim 3, wherein the pairing is activated by a tap between the testing device and the mobile device, wherein the tap is sensed by the testing device and the mobile device, and wherein the testing device and the mobile device register that the tap represents an intention to pair the testing device and the mobile device. 5. The system of claim 3, wherein the testing device is one of a plurality of testing devices paired with the mobile device. 6. The system of claim 5, wherein the application is operable to display information about the plurality of testing devices paired with the mobile device. 7. The system of claim 1, wherein the application is operable to provide an alert to perform a measurement based on an alert criterion. 8. The system of claim 1, wherein the application is operable to analyze the data and determine patterns of behavior of the subject. 9. The system of claim 1, wherein the application is operable to generate a warning when the data indicate that an analyte level falls outside selected analyte parameters for the subject. 10. The system of claim 1, wherein the application is operable to display different configurations available for the testing device and operable to configure the testing device while paired. 11. The system of claim 1, wherein the application is operable to allow entry of additional data to be associated with the data from the testing device. 12. The system of claim 1,... | Ascensia Diabetes Care Holding AG | G16H010/40 A61B005/00 G16H040/63 A61B005/145 G06F019/00 C12Q001/54 G01N033/66 A61B005/1486
US10722690 | B2 | 2020-07-28 | Medical devices with antithrombogenic coatings | 2016-12-22 | Pressure sensing guidewires are disclosed. The pressure sensing guidewires may include a tubular member having a proximal portion and a distal portion. The distal portion may have a plurality of slots formed therein. The distal portion may have a first wall thickness along a first region and a second wall thickness smaller than the first wall thickness along a second region. A pressure sensor may be disposed within the distal portion of the tubular member and housed within the second region. An anti-thrombogenic coating may be disposed on an inner surface, an outer surface, or both of the second region of the distal portion of the tubular member. | 1. A pressure sensing guidewire, comprising: a tubular member having a proximal portion and a distal portion;wherein the distal portion has a plurality of slots formed therein and two or more apertures formed therein, the two or more apertures positioned uniformly about a circumference of the distal portion and at similar axial locations distal to the plurality of slots;wherein the distal portion has a first wall thickness along a first region and a second wall thickness smaller than the first wall thickness along a second region;a pressure sensor disposed within the distal portion of the tubular member and housed within the second region distal to the plurality of slots; andan anti-thrombogenic coating disposed on an inner surface of the second region of the distal portion of the tubular member and at least partially covering one or more side walls of the two or more apertures;wherein the first region of the distal portion of the tubular member is free from the anti-thrombogenic coating. 2. The pressure sensing guidewire of claim 1, wherein the anti-thrombogenic coating comprises a hydrophilic coating. 3. The pressure sensing guidewire of claim 2, wherein the hydrophilic coating comprises a heparinized hydrophilic coating. 4. The pressure sensing guidewire of claim 1, wherein the anti-thrombogenic coating comprises a hydrophobic coating. 5. The pressure sensing guidewire of claim 4, wherein the hydrophobic coating comprises silicone. 6. The pressure sensing guidewire of claim 4, wherein the hydrophobic coating is disposed on the pressure sensor. 7. The pressure sensing guidewire of claim 1, wherein the first region of the distal portion is disposed proximal to the second region of the distal portion. 8. The pressure sensing guidewire of claim 1, wherein the second region of the distal portion is free from the plurality of slots. 9. The pressure sensing guidewire of claim 1, wherein the pressure sensor is positioned proximal to the one or more apertures. 10. A pressure sensing guidewire, comprising: a tubular member having a proximal portion and a distal portion;wherein the distal portion has a plurality of slots formed therein and two or more apertures formed therein, the two or more apertures having a length extending parallel to a longitudinal axis of the tubular member and positioned uniformly about a circumference of the distal portion and at similar axial locations distal to the plurality of slots, the length of the two or more apertures greater than a length of the plurality of slots;wherein the distal portion has a first wall thickness along a first region and a second wall thickness smaller than the first wall thickness along a second region;a pressure sensor disposed within the distal portion of the tubular member and housed within the second region proximal to the two or more apertures; andan anti-thrombogenic coating disposed on an outer surface and an inner surface of the second region of the distal portion of the tubular me... | BOSTON SCIENTIFIC SCIMED, INC. | A61M025/09 A61B005/0215 A61B005/00
US10722121 | B2 | 2020-07-28 | Chronic total occlusion crossing devices with imaging | 2017-12-26 | An imaging device includes a hollow flexible shaft having a central longitudinal axis and an imaging window therein. An optical fiber extends within the hollow flexible shaft substantially along the central axis. A distal tip of the optical fiber is attached to the hollow flexible shaft and aligned with the imaging window so as to transfer an optical coherence tomography signal through the imaging window. A handle is attached to the hollow flexible shaft configured rotate the hollow flexible shaft at speeds of greater than 1,000 rpm. | 1. An imaging device comprising: an outer sheath including an optically clear annular section at a distal end thereof;a hollow flexible shaft configured to be inserted within the outer sheath and rotated relative thereto, the hollow flexible shaft having an imaging window therein;an optical fiber extending within the hollow flexible shaft substantially along the central axis, a distal tip of the optical fiber attached to the hollow flexible shaft and aligned with the imaging window so as to transfer an optical coherence tomography signal through the imaging window. 2. The imaging device of claim 1, wherein the optical fiber extends substantially along the central axis for the entire length of the fiber. 3. The imaging device of claim 1, wherein the device is less than 0.1 inches in diameter. 4. The imaging device of claim 3, wherein the device is less than 0.08 inches in diameter. 5. The imaging device of claim 4, wherein the device is less than 0.05 inches in diameter. 6. The imaging device of claim 1, wherein the hollow flexible shaft is made of tungsten. 7. The imaging device of claim 1, wherein the hollow flexible shaft is made of multiple layers of wound filars. 8. The imaging device of claim 7, wherein the filars are counterwound. 9. The imaging device of claim 1, wherein the hollow flexible shaft further includes a mirror therein configured to reflect light from the optical fiber into adjacent tissue. 10. The imaging device of claim 1, wherein the optically clear annular section comprises tecothane or fluorinated ethylene propylene. 11. The imaging device of claim 1, wherein the optically clear annular section has a refractive index of between 1.35 and 1.45. 12. The imaging device of claim 1, wherein the imaging device further comprises a cutter thereon. 13. A method of imaging a body lumen, the method comprising: inserting an outer sheath into the body lumen, the outer sheath having an optically clear annular section at a distal end thereof;inserting an imaging device into a lumen of the other sheath, the imaging device including a hollow flexible shaft having a central longitudinal axis with an imaging window therein and an optical fiber extending within the hollow flexible shaft, wherein a distal tip of the optical fiber is attached to the hollow flexible shaft and aligned with the imaging window, the optical fiber extending substantially along the central longitudinal axis;rotating the hollow flexible shaft within the lumen of the outer sheath; andcollecting images of the body lumen through the imaging window and the optically clear annular section with the optical fiber as the hollow flexible shaft rotates. 14. The method of claim 13, wherein collecting images of the body lumen comprises collecting images of the body lumen at rates of greater than 10 frames per minute. 15. The method of claim 13, wherein the body lumen is a coronary artery. 16. The method of claim 13, wherein the body lumen is a peripheral artery. 17. ... | Avinger, Inc. | A61B005/00 A61B005/02 A61B017/3207 A61B001/005 A61B001/04 A61B001/07 A61B001/313 A61B017/32 A61B090/00 A61B017/22
US10726301 | B2 | 2020-07-28 | Method for treating a surface | 2018-06-22 | Method for treating a surface includes: automatically evaluating at least one digital image which includes the target surface; determining the nature of the target surface according to the evaluation of the at least one digital image; determining at least one available treatment implement according to the evaluation of the at least one image; determining the nature of the surface treatment according to the evaluation of the at least one image; automatically determining a use of the determined treatment implement in the determined treatment of the determined surface; and providing information analogous to the determined use of the treatment implement. | 1. A method for treating a target surface, the method comprising steps of: a. using machine learning in evaluating at least one digital image including the target surface;b. determining the nature of the target surface according to the evaluation of the at least one image;c. determining at least one available treatment implement according to the evaluation of the at least one image;d. determining the nature of the surface treatment according to the evaluation of the at least one image;e. using machine learning to determine a use of the determined implement in the determined treatment of the determined surface;f. altering performance characteristics of the implement or providing information analogous to the determined use. 2. The method according to claim 1 wherein the step of determining the nature of the target surface comprises determining one or more properties of the surface associated with surface treatments. 3. The method according to claim 1 wherein the step of determining at least one available treatment implements comprises determining the presence of a hand-held surface treatment appliance. 4. The method according to claim 1 wherein the step of determining the nature of the surface treatment comprises evaluating the nature of the determined surface and the nature of the determined implement. 5. The method of claim 1 further comprising the step of determining one or more properties of a treatment practitioner according to the evaluation of the at least one image. 6. The method according to claim 5, wherein the step of determining the nature of the surface treatment comprises using machine learning in evaluating at least one of: the nature of the surface, the implement, and the practitioner. 7. The method according to claim 1 further comprising the step of providing information regarding implements associated with a treatment of the determined target surface, such implements not determined to be present in the evaluation of the one or more images. 8. The method according to claim 1 further comprising the step of determining one or more environmental properties according to the evaluation of the one or more images together with at least one additional data source. 9. The method according to claim 1 wherein the step of providing information comprises providing cues analogous to the spatial interaction between the determined implement and the determined surface. 10. The method according to claim 1 wherein the step of providing information comprises providing information analogous to the temporal interaction of the determined implement and the determined surface. 11. The method according to claim 1 wherein the step of providing information comprises providing information through an alteration of a property of the determined implement. 12. The method according to claim 1 further comprising the steps of: g. tracking an initial determined treatment of a determined target surface;h. evaluating the initial determined treatment according... | The Procter & Gamble Company | G06K009/62 A45D044/00 A61B005/00 G06K009/46 G06N020/00 G06Q050/00
US10725129 | B2 | 2020-07-28 | Method for operating a local coil, local coil, and a magnetic resonance scanner | 2018-06-01 | The invention relates to a method for operating a local coil for a magnetic resonance scanner, which has a receiving antenna and a signal converter, and is coupled to a patient table by a signal line in terms of signal communication. The receiving antenna receives an analog magnetic resonance signal in a first signal frequency range, wherein the analog magnetic resonance signal is converted into a digital magnetic resonance signal by the signal converter and is frequency-shifted such that the digital magnetic resonance signal is shifted into a second signal frequency range that does not overlap with, and is preferably higher than, the first signal frequency range. | 1. A method for operating a local coil of a magnetic resonance scanner, said magnetic resonance scanner comprising a patient table and said local coil comprising a reception antenna and a signal converter connected thereto, with said signal converter being connected to a signal line that is coupled for signal communication to the patient table, said method comprising: during operation of the magnetic resonance scanner, receiving an analog magnetic resonance signal with said reception antenna in a first signal frequency range;in said signal converter, that is comprised within the local coil, converting the analog magnetic resonance signal into a digital magnetic resonance signal with frequency shifting that shifts the digital magnetic resonance signal into a second signal frequency range that does not overlap with said first signal frequency range; andtransmitting the frequency-shifted digital magnetic resonance signal via the signal line to the patient table. 2. A method as claimed in claim 1 wherein said second signal frequency range is higher than said first signal frequency range. 3. A method as claimed in claim 1 comprising frequency shifting the digital magnetic resonance signal in said signal converter by a frequency shift procedure selected from the group consisting of line coding and signal modulation. 4. A local coil for use in a magnetic resonance scanner comprising a patient table, said local coil comprising: a reception antennaa signal converter connected to said reception antenna, wherein said signal converter connected to a signal line that is coupled for signal communication to the patient table;during operation of the magnetic resonance scanner, said reception antenna receiving an analog magnetic resonance signal in a first signal frequency range; andsaid signal converter being configured to convert the analog magnetic resonance signal into a digital magnetic resonance signal with frequency shifting that shifts the digital magnetic resonance signal into a second signal frequency range that does not overlap with said first signal frequency range, and to transmit the frequency-shifted digital magnetic resonance signal via the signal line to the patient table. 5. A local coil as claimed in claim 4 wherein said signal converter comprises a logic module that implements said frequency shifting. 6. A local coil as claimed in claim 4 comprising a high-pass filter connected between said signal converter and said signal line, said high-pass filter attenuating signal frequencies of said digitized magnetic resonance signal that are outside of said second signal frequency range. 7. A local coil as claimed in claim 4 wherein said signal line is a symmetrical two-wire line comprising a plurality of standing wave traps, and thereby forming a symmetrical high-pass line. 8. A local coil as claimed in claim 7 wherein said signal terminal is a contact list Near Field Communication (NFC) connector. 9. A local coil as claimed in claim 4 wherein... | Siemens Healthcare GmbH | G01R033/34 G01R033/36 A61B005/055 G01R033/54 G01R019/252
US10725730 | B2 | 2020-07-28 | Physiological control based upon media content selection | 2019-01-07 | A media-playback device includes: a media-output device that plays media content items; a physiological measurement device programmed to measure at least one physiological measurement of a user of the media-output device; and a physiological control engine configured to: identify a current physiological measurement for the user; and cause the media-output device to modify playback of the media content items based upon the current physiological measurement. | 1. A method for playing media content items based on a physiological measurement, the method comprising: receiving at least one physiological measurement of a user, the physiological measurement associated with a level of exertion of the user during performance of a repetitive-motion activity;identifying a current physiological measurement of the user;estimating a future physiological measurement of the user from a workout plan that involves one or more changes in the physiological measurement of the user, the workout plan requiring the user to maintain a first heart rate for a first predetermined interval, and to change the first heart rate to a second heart rate for a second predetermined interval; andtransmitting a signal configured to cause a media-playback device to modify playback of media content items based on the estimated future physiological measurement. 2. The method of claim 1, further comprising comparing the current physiological measurement to the future physiological measurement to modify playback of media content items based on the comparison. 3. The method of claim 2, wherein the current physiological measurement is a current heart rate, and the future physiological measurement is a future heart rate. 4. The method of claim 1, wherein the first and second predetermined intervals are durations of time. 5. The method of claim 1, wherein the first and second predetermined intervals are distances. 6. The method of claim 1, wherein the signal is configured to cause the media-playback device to adjust a tempo used to select media content items for playback based on the estimated future physiological measurement. 7. The method of claim 6, wherein the signal is configured to cause the media-playback device to match transition beats between current media content and new media content to maintain a steady cadence transition. 8. The method of claim 1, wherein the signal is configured to cause the media-playback device to adjust a tempo of the media content that is currently being played by the media-playback device to a new tempo. 9. A media delivery system comprising: at least one processing device configured to control operation of the media delivery system; andat least one memory device storing software instructions that, when executed by the at least one processing device, cause the media delivery system to: receive at least one physiological measurement of a user, the physiological measurement associated with a level of exertion of the user during performance of a repetitive-motion activity;identify a current physiological measurement of the user;estimate a future physiological measurement of the user from a workout plan that involves one or more changes in the physiological measurement of the user, the workout plan requiring the user to maintain a first heart rate for a first predetermined interval, and to change the first heart rate to a second heart rate for a second predetermined interval; andtransmit a signal configured ... | SPOTIFY AB | G06F017/00 G06F003/16 G06F003/01 A61B005/024 A63B071/00 G01S019/19 A61B005/00 A61B005/08 A61B005/11 G16H040/67 A61B005/01 A61B005/083 A63B071/06 A61B005/0205
US10712357 | B2 | 2020-07-14 | Circuits and methods for artifact elimination | 2017-06-15 | Disclosed are apparatus and methods that provide the ability to electrical stimulate a physical system, and actively eliminate interference with signal acquisition (artifacts) that arises from the stimulation. The technique implemented in the circuits and methods for eliminating interference connects a discharge path to a physical interface to the system to remove charge that is built-up during stimulation. By placing the discharge path in a feedback loop that includes a recording preamplifier and AC-coupling circuitry, the physical interface is brought back to its pre-stimulation offset voltage. The disclosed apparatus and methods may be used with piezoelectric transducers, ultrasound devices, optical diodes, and polarizable and non-polarizable electrodes. The disclosed apparatus can be employed in implantable devices, in vitro or in vivo setups with vertebrate and invertebrate neural tissue, muscle fibers, pancreatic islet cells, osteoblasts, osteoclasts, bacteria, algae, fungi, protists, and plants. | 1. An apparatus comprising: an excitation circuitry that electrically excites a physical system by applying an excitation signal to a physical interface to the physical system; andsignal acquisition circuitry that actively eliminates interference arising from the excitation signal, said signal acquisition circuitry comprises a feedback loop that incorporates the physical interface, the feedback loop further providing an electrical discharge path for the physical interface, wherein said feedback loop eliminates interference arising from the excitation of the physical system by a continuous-time adjustment of an electrical characteristic of the physical interface wherein said adjustment of the electrical characteristic is a function of a present condition and a pre-excitation condition of the physical interface. 2. The apparatus of claim 1, wherein the signal acquisition circuitry further comprises an amplifier that is AC coupled to the physical interface. 3. The apparatus of claim 1, which is embodied in an integrated circuit. 4. The apparatus of claim 1, wherein the physical interface is selected from a group including piezoelectric transducers, ultrasound devices, optical diodes, polarizable and nonpolarizable electrodes. 5. The apparatus of claim 1, wherein the interference is actively eliminated by monitoring and discharging the physical interface to a value substantially equal to the pre-excitation condition of the physical interface. 6. The apparatus of claim 1, which is incorporated into an implantable medical device. 7. The apparatus of claim 1, wherein the signal acquisition circuitry further comprises discharge circuitry that actively eliminates interference by transiently increasing the tracking speed of a recording amplifier in the feedback loop. 8. The apparatus of claim 7, wherein the tracking speed of the recording amplifier is transiently increased by increasing a dominant high-pass pole of the recording amplifier. 9. The apparatus of claim 7, wherein increasing the tracking speed is accomplished by adjusting a bias current in the recording amplifier. 10. The apparatus of claim 9, wherein the bias current is adjusted by transiently adjusting a voltage level applied to a transconductance device that provides current that contributes to the bias current level. 11. The apparatus of claim 9, wherein the bias current is adjusted by selectively switching on or off a switch that provides current that contributes to the bias current level. 12. The apparatus of claim 9, wherein the bias current is adjusted by transiently adjusting a voltage level applied to a first transconductance device that provides current that contributes to the bias current level and by selectively switching on or off a second transconductance device that provides current that contributes to the bias current level. 13. The apparatus of claim 1, wherein the signal acquisition circuitry further comprises a discharge path comprising an additional feedback loo... | Georgia Tech Research Corporation | G01N037/00 A61B005/04 G01N033/483 A61B005/00 A61N001/36 A61N001/37
US10709412 | B2 | 2020-07-14 | System and method for replacing a rotational axis decoder in a scanning system | 2017-12-05 | A method for determining the position and/or orientation of at least one sensor system relative to the base structure of a scanning system during scanning of an object includes obtaining one or more tracking images using one or more cameras, where the cameras are in a fixed position with respect to the sensor system; and determining from the one or more tracking images the position and/or orientation of the sensor system relative to the base structure at a given time. | 1. A method for determining the position or orientation of at least one sensor system relative to the base structure of a cone beam computed tomography scanning system during scanning of an object, the method comprising: obtaining one or more tracking images using one or more cameras, where the cameras are in a fixed position with respect to the sensor system; anddetermining from the one or more tracking images the position or orientation of the sensor system relative to the base structure at a given time;wherein:a tracking element is attached in a fixed position with respect to the base structure of the scanning system, the tracking element comprising one or more fiducial markers in a known configuration;the tracking images are images taken of the tracking element;the tracking images are processed to identify the position of the one or more fiducial markers in a plurality of the tracking images;the identified position of the fiducial markers is used to determine the position of the tracking element at a given time. 2. The method according to claim 1, wherein the determined position or orientation of the one or more sensor systems relative to the base structure of the scanning system, is used to replace positional information from a physical motor encoder. 3. The method according to claim 1, wherein the object being scanned is a person. 4. The method according to claim 1, wherein the method is for determining the position and orientation of at least one sensor system relative to the base structure of a cone beam computed tomography scanning system during scanning of an object, wherein the determining step includes determining from the one or more tracking images the position and orientation of the sensor system relative to the base structure at a given time. 5. The method according to claim 4, wherein the determined position and orientation of the one or more sensor systems relative to the base structure of the scanning system, is used to replace positional information from a physical motor encoder. 6. The method according to claim 1, wherein the sensor system and the one or more cameras are rotated around the object while the tracking images are obtained from the one or more cameras and medical images are obtained from the sensor system. 7. The method according to claim 6, wherein the tracking element is fastened to the object, and the tracking element does not move relative to the base structure while the tracking images are obtained from the one or more cameras and the medical images are obtained from the sensor system. 8. The method according to claim 1, wherein the object is not moved while the tracking images are obtained from the one or more cameras and the medical images are obtained from the sensor system. 9. The method according to claim 8, wherein the sensor system and the one or more cameras are rotated around the object while the tracking images are obtained from the one or more cameras and medical images are obtained from th... | 3SHAPE A/S | G06K009/00 A61B005/05 A61B006/00 G06T007/292 A61B005/11 A61B005/117 A61B006/03
US10702175 | B2 | 2020-07-07 | Brain activity analysis method and apparatus thereof | 2017-07-06 | The present invention discloses a brain activity analysis method and apparatus, which is based on a nonlinear waveform decomposition technology, wherein the changes of the intrinsic features in brain waves are decomposed and demodulated to extract the modulation signals of the components, including the frequency-modulation signals and the amplitude-modulation signals. The present invention further uses a feature mask to determine whether to proceed further decomposition and demodulation of the extracted modulation signals. If not, the multidimensional changes of the intrinsic features are obtained according to the feature mask. Then, quantitation and identification is performed to obtain the status of brain function. The present invention not only effectively increases the accuracy of the identification but also uses the feature mask to obviously reduce the complexity and the load of computation. | 1. A brain activity analysis apparatus for processing at least one brain electrical signal to generate an analysis result, said brain activity analysis apparatus comprising: at least one brain activity sensor, said brain activity sensor configured to detect brain waves of a subject to acquire said at least one brain electrical signal;a signal processor in communication with said brain activity sensor, said signal processor configured to: receive said brain electrical signal,use a nonstationary decomposition method to decompose said brain electrical signal to acquire a plurality of sub-signals carrying intrinsic feature components,demodulate each of said sub-signals to generate modulation signals respectively corresponding to said sub-signals,recursively iterate said nonstationary decomposition method to further decompose and demodulate said modulation signals according to a preset feature mask, said feature mask formed of a word string, a sequence matrix, or a multidimensional matrix,select signals of interest from said modulation signals as feature modulation signals according to said feature mask, andundertake quantitation processes and identification processes of said feature modulation signals to obtain an analysis result corresponding to said brain electrical signal; anda display electrically connected with said signal processor and configured to present said analysis result. 2. The brain activity analysis apparatus according to claim 1, wherein said brain electrical signal is an electroencephalography (EEG) signal, an intracranial electroencephalogram (iEEG) signal, or an electrocorticography (ECoG) signal. 3. The brain activity analysis apparatus according to claim 1, wherein said nonstationary decomposition method is an empirical mode decomposition (EMD) method. 4. The brain activity analysis apparatus according to claim 1, wherein a normalization operation is used to demodulate each of said sub-signals to obtain said modulation signals. 5. The brain activity analysis apparatus according to claim 1, wherein said modulation signals include frequency-modulation parts and amplitude-modulation parts. 6. The brain activity analysis apparatus according to claim 1, wherein said analysis result is a status of brain activity. 7. The brain activity analysis apparatus according to claim 1, wherein said quantitation processes include calculation of power densities, instantaneous frequencies, or averaged periods. 8. The brain activity analysis apparatus according to claim 1, wherein a classification model is used in said identification processes, and wherein said classification model involves personal history and pre-trained parameters. 9. The brain activity analysis apparatus according to claim 1 further comprising a storage unit electrically connected with said signal processor and configured to store signals, data and results processed by said signal processor. | A-Neuron Electronic Corp. | A61B005/04 A61B005/0476 A61B005/00 G16H050/20 A61B005/048 G16H040/63
US10716391 | B2 | 2020-07-21 | Connected hairbrush | 2017-12-28 | A system that includes a styling tool configured to treat or shape a user's hair; and a client device configured to communicate with the styling tool and provide feedback to the user regarding a user's usage of the styling tool based at least in part on measured data captured by the styling tool. The styling tool may be a hairbrush that includes circuitry and/or sensors configured to detect one or more of sound, acceleration, force, rotation, hair humidity, and ambient temperature and humidity. | 1. A system, comprising: a hairbrush configured to treat or shape a user's hair and including a motion sensor; anda client device configured to communicate with the hairbrush and provide real-time feedback to the user regarding a user's usage of the hairbrush based at least in part on measured data captured by the hairbrush,wherein the client device is configured to output assessment information on the user's use of the hairbrush and to output guidance information on a method of using the hairbrush to achieve a target hair health,wherein the measured data captured by the hairbrush is a detected motion of the hairbrush detected by the motion sensor, and client device is configured to compare the detected motion with predetermined motion data and output instructions, based on the comparison, for the user to make a correction which includes at least an adjustment of a movement, speed, or force of the hairbrush. 2. The system according to claim 1, wherein the guidance information is image data, video data, or audio data on using the hairbrush. 3. The system according to claim 1, wherein the client device is configured to receive user information regarding physical characteristics of the user and outputs the guidance information based on the received user information. 4. The system according to claim 1, wherein the hairbrush includes circuitry configured to detect one or more of sound, acceleration, force, rotation, hair humidity, and ambient temperature and humidity as the measured data. 5. The system according to claim 1, wherein the outputted assessment information and guidance information is based on measured data received for a single continuous session of the user using the hairbrush. 6. The system according to claim 1, wherein the outputted assessment information is based on measured data received for a predetermined period of time that includes multiple sessions of the user using the hairbrush. | L'OREAL | A46B015/00 A46B009/02 A45D024/10 G09B019/00 A45D044/00 G09B019/24 A61B005/103 A61B005/00 G09B005/04 G09B005/02
US10726688 | B2 | 2020-07-28 | Facilitating a search of individuals in a building during an emergency event | 2019-08-14 | A computer-implemented method includes: determining, by a computing device, a location of an individual in a building relative to the location of an exit of the building; determining, by the computing device, the location of obstacles or hazards within the building; determining, by the computing device, a path from the individual to the exit based on determining the location of the individual relative to the exit and the location of the obstacles or hazards; generating, by the computing device, signaling instructions based on the path; and outputting, by the computing device, the signaling instructions to one or more signaling devices, wherein the outputting the signaling instructions control operations of the one or more signaling devices to guide the individual to the exit based on the path. | 1. A computer-implemented method comprising: determining, by a computing device, a location of an individual in a building relative to the location of an exit of the building;determining, by the computing device, locations of obstacles or hazards within the building;determining, by the computing device, distress level data representing a distress level associated with the individual;determining, by the computing device, a path from the individual to the exit based on determining the location of the individual relative to the exit and the locations of the obstacles or hazards;selecting, by the computing device, a type of signaling instructions based on the path and the distress level data; andgenerating, by the computing device, signaling instructions of the selected type of signaling instructions based on the path and the distress level data. 2. The method of claim 1, further comprising: monitoring, by the computing device, the location of the individual relative to the location of the exit of the building; andupdating, by the computing device, the signaling instructions based on the monitoring the location of the individual. 3. The method of claim 1, wherein the signaling instructions control at least one operation of the signaling devices selected from the group consisting of: light color;audio output rate;audio output tone;audio output speech or spoken words;audio output speech; anddisplayed content. 4. The method of claim, 1 wherein the one or more signaling devices include at least one from the group consisting of: lighting devices;audio output devices; andelectronic display devices. 5. The method of claim 1, wherein the location of the individual is determined based on at least one selected from the group consisting of: signal strength analysis from signals emitted by the user device of the individual;object detection sensor data;heat sensor data;audio sensor data; andvisual camera data. 6. The method of claim 1, wherein the location of the obstacles is determined based on object detection sensor data. 7. The method of claim 1, further comprising determining, by the computing device, a path from the individual to a responder, wherein the generating the signaling instructions is further based on the path between the individual to the responder. 8. The method of claim 1, further comprising: determining respective locations of a plurality of different individuals; andgenerating respective signaling instructions for each of the plurality of different individuals based on the respective locations. 9. The method of claim 8, further comprising: determining, by the computing device, distress level data that indicates a distress level associated with each of the plurality of different individuals, wherein the generating the signaling instructions is further based on the distress level data of the plurality of different individuals. 10. The method of claim 9, further comprising determining, by the computing device, a rescue priority based on... | INTERNATIONAL BUSINESS MACHINES CORPORATION | G01C021/20 G01C021/26 G08B007/06 G01C021/34 G06F003/0481 G06F001/16 A62B003/00 A62B005/00 H04W052/02 G08B025/01 G08B021/02 H04W004/90 G08B027/00 H04W004/33 A61B005/0402 A61B005/0205 A61B005/0295 G06F003/0484 H04W004/02
US10702700 | B2 | 2020-07-07 | Hybrid implantable connectors | 2018-06-01 | A hybrid connector for connecting a bulk conductor to a planar strand of a flexible circuit is described. The hybrid connector can include a connector body, a first connector end, and a second connector end. The first connector end can be attached to the connector body and can define a first opening having a first cross-sectional shape that corresponds to a first distal end of the strand. The second connector end can be attached to the connector body opposite the first connector end and can define a second opening having a second cross-sectional shape that corresponds to a second distal end of the bulk conductor. | 1. A system, comprising: a connector defining a cavity, with a first opening formed at a first end of the connector and a second opening formed at a second end of the connector, the first opening defining a first connector cross-sectional shape; anda flexible circuit defining a neural interface and comprising an elongate planar strand, the elongate planar strand comprising a conductive region disposed at a distal end of the elongate planar strand on a top surface of the elongate planar strand or a bottom surface of the elongate planar strand, the distal end comprising a strand cross-sectional shape,wherein at least a portion of the conductive region of the elongate planar strand is disposed within the first opening and permanently bonded to the connector. 2. The system of claim 1, further comprising a bulk conductor comprising a second distal end comprising a conductor cross-sectional shape, and wherein the second distal end of the bulk conductor is disposed within the second opening. 3. The system of claim 2, wherein the connector is permanently deformable to retain the portion of the conductive region of the elongate planar strand within the first opening and to retain the second distal end of the bulk conductor within the second opening. 4. The system of claim 2, wherein the portion of the conductive region is in physical contact with the second distal end of the bulk conductor. 5. The system of claim 2, wherein the strand cross-sectional shape corresponds to the first connector cross-sectional shape and is rectangular, and the conductor cross-sectional shape is circular. 6. The system of claim 2, further comprising a biased member disposed within the cavity of the connector, the biased member configured to electrically connect the bulk conductor and the conductive region. 7. The system of claim 6, wherein the bulk conductor and the elongate planar strand are disposed in different horizontal planes. 8. The system of claim 2, further comprising a tension member attached to the elongate planar strand and the bulk conductor, the tension member configured to support a tension force between the elongate planar strand and the bulk conductor. 9. The system of claim 1, wherein: the system further comprises a second connector; andthe flexible circuit comprises a second elongate planar strand comprising a second conductive region, with at least a second portion of the second conductive region being disposed within a second cavity of the second connector. 10. The system of claim 1, wherein: the elongate planar strand comprises a second conductive region;the connector defines a second cavity, with a third opening disposed at the first end of the connector and a fourth opening disposed at the second end of the connector; andwherein at least a second portion of the second conductive region of the elongate planar strand is disposed within the second cavity via the third opening. 11. The system of claim 1, wherein: the connector defines a second cav... | VERILY LIFE SCIENCES LLC | A61N001/372 A61B005/00 A61N001/05 A61N001/375 A61N001/37
US10709349 | B2 | 2020-07-14 | Annotation waveform | 2018-04-17 | Systems and methods for processing cardiac information include a processing unit configured to receive a set of cardiac electrical signals; receive an indication of a measurement location corresponding to each of the set of electrical signals; and identify, for each electrical signal of the set of electrical signals, a deflection. The deflection includes a deviation from a signal baseline. An activation waveform corresponding to the set of electrical signals is generated based on the identified deflections. | 1. A system for processing cardiac information, the system comprising: a processing unit configured to: receive a set of cardiac electrical signals;receive an indication of a measurement location corresponding to each of the set of electrical signals;identify, for each electrical signal of the set of electrical signals, a deflection, the deflection comprising a deviation from a signal baseline;generate, based on the identified deflections, an activation waveform corresponding to the set of electrical signals, wherein the activation waveform comprises values corresponding to probabilities that the identified deflections represent activations of cardiac tissue; andfacilitate presentation, on a display device and based on the activation waveform, of at least one of a cardiac map, a representation of the activation waveform, and a representation of an activation histogram. 2. The system of claim 1, wherein the processing unit is further configured to generate one or more cardiac map annotations based on the activation waveform. 3. The system of claim 1, wherein the processing unit is further configured to determine the signal baseline. 4. The system of claim 3, wherein the processing unit is configured to determine the signal baseline by referencing a memory in which the signal baseline is stored, the signal baseline comprising a pre-determined value. 5. The system of claim 3, wherein the processing unit is configured to determine the signal baseline based on an estimated noise floor associated with at least one electrical signal of the set of electrical signals. 6. The system of claim 1, wherein the processing unit is configured to identify, for each electrical signal of the set of electrical signals, a deflection by identifying, for each sample point of each electrical signal, whether the sample point represents a deviation from the signal baseline. 7. The system of claim 1, wherein the set of electrical signals comprises a plurality of electrical signals, wherein the processing unit is configured to adjust an activation waveform value corresponding to a sample point based on an evaluation of spatiotemporal consistency of corresponding deflections of the plurality of electrical signals. 8. The system of claim 1, wherein the processing unit is further configured to identify far-field signal components of each of the set of electrical signals. 9. The system of claim 1, wherein the processing unit is further configured to suppress far-field signal components of each of the set of electrical signals. 10. The system of claim 1, further comprising a catheter, wherein the processing unit is configured to receive the set of electrical signals from the catheter, the catheter comprising at least one of a mapping catheter, a diagnostic catheter, a CS catheter, and an ablation catheter. 11. A method of processing cardiac information, comprising: receiving a set of cardiac electrical signals;receiving an indication of a measurement location correspond... | Boston Scientific Scimed Inc. | A61B005/04 A61B005/044 A61B005/0452 A61B005/042 G06T017/00 A61B005/00 A61B005/06 G16H050/20
US10709627 | B2 | 2020-07-14 | Use of high pressure low volume air conversion to low pressure high volume air to power patient transport devices | 2016-01-14 | Systems for patient support, imaging, or transport include a modality with a support surface configured to support a patient thereon, a converter associated with the modality, the converter being configured to receive relatively low-volume high-pressure air from a source of the relatively low-volume high-pressure air and to convert the relatively low-volume high-pressure air into relatively high-volume low-pressure air, and an air flow device configured to receive the relatively high-volume low-pressure air from the converter and provide an air flow function to the modality using the relatively high-volume low-pressure air. | 1. A system for at least one of patient support, patient imaging, patient therapeutic procedures with or without the use of instruments, and patient transport, the system comprising: a modality with a support surface configured to support a patient thereon;a converter associated with the modality, the converter not comprising an electric blower, the converter being configured to receive relatively low-volume high-pressure air flow from a source of the relatively low-volume high-pressure air flow and to draw ambient air, and thereby to convert the relatively low-volume high-pressure air flow and the ambient air into relatively high-volume low-pressure air flow; andan air flow device configured to receive the relatively high-volume low-pressure air flow from the converter and provide an air flow function to the modality using the relatively high-volume low-pressure air flow. 2. The system of claim 1, wherein the air flow device is configured to direct the relatively high-volume low-pressure air flow toward the support surface of the modality, thereby facilitating transfer of the patient to or from the modality. 3. The system of claim 2, wherein the air flow device is configured to provide an air bearing positioned below the support surface of the modality. 4. The system of claim 1, wherein the air flow device is coupled to the modality. 5. The system of claim 1, wherein the converter is constructed of non-ferromagnetic material. 6. The system of claim 5, wherein the converter is constructed of a polymer-based material. 7. The system of claim 1, wherein the converter is integrated into the modality. 8. The system of claim 1, wherein the converter is removably connected to the modality. 9. The system of claim 1, wherein the modality, the converter, and the air flow device are constructed of materials such that the system is MRI safe. 10. A method of providing air flow to a modality, the method comprising: converting relatively low-volume high-pressure air flow from a relatively low-volume high-pressure air flow source and ambient air into relatively high-volume low-pressure air flow;receiving, at an air flow device, the relatively high-volume low-pressure air flow; andproviding, with the air flow device, an air flow function to the modality with the relatively high-volume low-pressure air flow,wherein the converting step does not utilize an electric blower. 11. The method of claim 10, wherein the providing step further comprises supplying air flow toward the support surface of the modality to facilitate transfer of the patient to or from the modality as the air flow function. 12. The method of claim 10, wherein the providing step further comprises inflating an air bladder positioned below the support surface of the modality to facilitate transfer of the patient to or from the modality as the air flow function. 13. The method of claim 10, further comprising imaging the patient with an MRI device, wherein the converter is constructed of a n... | Qfix Systems, LLC | A61G007/10 A61B006/04 A61B005/055 A61B017/225
US10722163 | B2 | 2020-07-28 | Methods, devices, systems, and kits for automated blood collection by fingerstick | 2016-03-31 | Methods, devices, systems and kits for obtaining blood samples are provided. Devices include a cuff, a pressure source, and a timing mechanism. Devices may further include one or more of a: warming mechanism, lancing mechanism; automated sample collection device, automated sample analysis device, and communication unit. Systems include such a device, and may include sample collection, sample analysis, or communication devices.;Methods include placing a cuff on a digit of a subject, inflating the cuff, and obtaining a small volume blood sample. Methods may further include warming a digit; lancing a digit; pulsing the cuff; and providing a signal indicating the end of the sample collection time period.;Kits may include a device, a sample collection vessel, and may include a disposable for use in sample collection.;These methods, devices, systems and kits for obtaining blood samples may be used to easily, reliably, and consistently obtain blood samples from subjects. | 1. A system for use in collecting a blood sample from a wound or incision on a subject using a blood sampling device, the system comprising: a cuff configured to at least partially encircle a digit on the subject;a pressure source configured to inflate said cuff effective to apply pressure to a digit disposed at least partially within the cuff, said pressure source comprising a housing having at least one button to toggle between starting pressurization and stopping pressurization, at least one visual indicator showing system status, and a flexible tubing coupling the cuff to the pressure source;a pressure-interruption switch for temporarily interrupting inflation of the cuff;an audio signal source;a control unit; anda timing mechanism configured to provide a timing mechanism signal at the beginning of a period of time and a timing mechanism signal at the end of said period of time, where the period of time is a desired period of time for collecting said blood sample from said digit, wherein said system is configured to end the inflation of said cuff at the end of the period of time;wherein said pressure-interruption switch interrupts said inflation without impacting a program timer of the timing mechanism;wherein said audio signal source provides a lancing audio cue in coordination with the timing mechanism for an operator to begin lancing, and an end-of-sample collection audio cue;wherein the pressure source and the button are controlled by the control unit to provide a series of inflations and deflations at a desired frequency between the beginning of the period of time and the end of the period of time. 2. The system of claim 1, wherein said cuff is configured to be inflated up to a maximum pressure of about 300 millimeters (mm) of mercury (Hg) of pressure. 3. The system of claim 1, wherein said cuff further comprises a warming mechanism configured to warm a said digit disposed within said cuff. 4. The system of claim 1, wherein said cuff comprises a plurality of lobes, wherein each lobe may be inflated independently of the other lobe or lobes. 5. The system of claim 1, wherein said cuff forms a tubular structure configured to completely encircle said digit placed within an internal portion of said tubular structure. 6. The system of claim 1, further comprising a lancing mechanism configured to puncture said digit effective to make said wound or incision in said digit, Wherein said wound or incision is suitable for providing blood for sample collection. 7. The system of claim 4, further comprising a lancing mechanism configured to puncture said digit effective to make said wound or incision in said digit, wherein said wound or incision is suitable for providing blood for sample collection. 8. A system, comprising a device of claim 1, and further comprising an automated sample collection device configured to receive a blood sample obtained from a puncture wound in said digit. 9. The system of claim 8, wherein said automated sample anal... | Labrador Diagnostics LLC | A61B005/00 A61B005/15 A61B005/151 G01N033/49
US10706533 | B2 | 2020-07-07 | Systems and methods for automatic detection and quantification of pathology using dynamic feature classification | 2017-05-15 | Methods, devices, and systems are provided for quantifying an extent of various pathology patterns in scanned subject images. The detection and quantification of pathology is performed automatically and unsupervised via a trained system. The methods, devices, and systems described herein generate unique dictionaries of elements based on actual image data scans to automatically identify pathology of new image data scans of subjects. The automatic detection and quantification system can detect a number of pathologies including a usual interstitial pneumonia pattern on computed tomography images, which is subject to high inter-observer variation, in the diagnosis of idiopathic pulmonary fibrosis. | 1. An automatic detection and quantification system, comprising: a server, comprising: a processor; anda computer readable medium, coupled to the processor and comprising instructions stored thereon that cause the processor to: receive an input image data scan of a subject;extract features from the input image data scan;determine, using an elements dictionary and a classification model, one or more normal and/or abnormal pathology areas of the input image data scan based on the extracted features, wherein each pathology area comprises a cluster of pixels;generate a marked image data scan by marking each of the determined one or more pathology areas on the input image data scan with a visual identifier;compute one or more quantitative scores for the input image data scan based on a comparison of a number of regions classified as various categories of normal and/or abnormal with a number of regions tested; andstore the marked image data scan and the one or more quantitative scores in a memory. 2. The automatic detection and quantification system of claim 1, wherein the elements dictionary is generated based on a plurality of sample image data scans received from a plurality of subjects. 3. The automatic detection and quantification system of claim 2, wherein the plurality of sample image data scans comprises healthy image data scans and pathology image data scans. 4. The automatic detection and quantification system of claim 3, wherein the classification model is generated in an unsupervised fashion using the elements dictionary and a clustering process. 5. The automatic detection and quantification system of claim 4, wherein the classification model is generated using training examples comprising delineated image data scans outlining regions of interest having pathology. 6. The automatic detection and quantification system of claim 5, further comprising: an image scan system configured to provide cross-sectional image data scans of the subject, wherein the server is communicatively coupled to the image scan system. 7. The automatic detection and quantification system of claim 6, wherein the marked image data scan is transmitted to a communication device across a communication network. 8. The automatic detection and quantification system of claim 7, wherein the marked image data scan is configured to be rendered by a display device of the communication device. 9. The automatic detection and quantification system of claim 8, wherein the visual identifier describes a pathology type and quantification value associated with the one or more pathology areas. 10. The automatic detection and quantification system of claim 1, wherein multiple elements dictionaries are generated based on a plurality of sample image data scans received from a plurality of subjects. 11. The automatic detection and quantification system of claim 10, wherein the multiple elements dictionaries comprise hierarchical dictionaries for features at multiple scales. 12. The ... | National Jewish Health | G06K009/00 G06T007/00 A61B005/00 G06N005/02 G16B099/00 A61B005/055 A61B006/00 G16H050/20 G06N020/10 G16H030/40 A61B006/03
US10718838 | B2 | 2020-07-21 | Systems and methods for calibrated multi-spectral magnetic resonance imaging | 2016-05-13 | Systems and methods are provided for performing a calibration “pre-scan” prior to acquiring data using a magnetic resonance imaging (“MRI”) system performing a multi-spectral imaging (“MSI”) acquisition. Information from the calibration scan is used to optimize the scanning and data collection during the MSI scan. As a result, scan times and motion artifacts are reduced. In addition, image resolution can also be increased, thereby improving image quality. As one example, the MSI acquisition can be a MAVRIC acquisition. In general, the calibration data is used to determine the minimum number of spectral bins required to achieve acceptable image quality near a specific metallic implant or device. | 1. A computer-implemented method for determining magnetic resonance imaging (“MRI”) scan parameters for a multi-spectral imaging (“MSI”) scan, the steps of the method comprising: (a) providing to a computer system, calibration data acquired with an MRI system from a field-of-view, wherein the calibration data contains data acquired at a plurality of different resonance frequency offsets;(b) computing with the computer system, a field map from the calibration data, the field map containing information about off-resonance effects in the field-of-view;(c) determining with the computer system, a spectral range from the computed field map;(d) setting with the computer system, a number of spectral bins based on the determined spectral range;(e) generating with the computer system, scan parameters based on the number of spectral bins and the determined spectral range, the scan parameters defining the spectral bins at which spins are to be excited and data are to be acquired with an MRI system performing an MSI scan; and(f) sending the scan parameters to the MRI system and performing the MSI scan in order to acquire data from a subject by operating the MRI system using the scan parameters, the data being acquired at the number of spectral bins set by the computer system. 2. The method as recited in claim 1, wherein step (a) includes acquiring the calibration data with the MRI system. 3. The method as recited in claim 1, wherein the calibration data provided in step (a) have a lower spatial resolution than the data to be acquired with the MRI system using the scan parameters generated in step (e). 4. The method as recited in claim 1, wherein step (c) includes calculating with the computer system, a frequency cutoff value from a frequency map generated by the computer system from the provided calibration data. 5. The method as recited in claim 4, wherein the computer system calculates the frequency cutoff value by: computing a histogram from the frequency map;calculating a forward cumulative distribution function (CDF) by integrating the histogram from low frequencies to high frequencies;calculating a backward CDF by integrating the histogram from high frequencies to low frequencies;determining a first frequency cutoff value using the forward CDF to identify a frequency value that retains a selected amount of signal intensities in the frequency map;determining a second frequency cutoff value using the backward CDF to identify a frequency value that retains the selected amount of signal intensities in the frequency map; andselecting the frequency cutoff value based on a value of the first frequency cutoff value and a value of the second frequency cutoff value. 6. The method as recited in claim 5, wherein the frequency cutoff value is selected as a signed value of the one of the first frequency cutoff value and second frequency cutoff value having a greater absolute value. 7. The method as recited in claim 4, wherein the computer system computes a magn... | The Medical College of Wisconsin, lnc. | G01R033/54 G01R033/24 G01R033/565 A61B005/055 A61B005/00 G01V013/00 G01V003/14 G01R033/58
US10718652 | B2 | 2020-07-21 | Detection system and method for aerosol delivery | 2018-03-23 | An apparatus comprises a detector, a pressure sensor and a processor. The detector is operable to detect light that is scattered by an aerosol that is associated with a pressure. The pressure sensor is operable to measure the pressure. The processor is coupled to the detector and to the pressure sensor, and is configured to receive at least a signal from the detector and the pressure sensor. The processor is further configured to use the received signals to calculate a volume of the first aerosol, and to output an output signal associated with the calculated volume. The various measurements can be repeated and compared, and the output signal can be a feedback signal for metering subsequent amounts of the aerosol, based on the comparison. | 1. An apparatus comprising: a sensor; anda processor coupled to the sensor, the processor configured to: receive a signal associated with the sensor that is associated with a first therapeutic agent introduced into a flow path;calculate a volume of the first therapeutic agent, the calculation based on the signal representing the sensor;receive a signal associated with the sensor from a second therapeutic agent introduced into the flow path;calculate a volume of the second therapeutic agent, the calculation based on the signal representing the sensor; andoutput a signal associated with a comparison of the volume of the first therapeutic agent and the volume of the second therapeutic agent. 2. The apparatus of claim 1, further comprising dose-selection means coupled to the processor, and wherein the output signal is received by the dose-selection means, and wherein the output signal includes information useful for metering a third therapeutic agent. 3. The apparatus of claim 1, wherein the output signal comprising information for metering a third therapeutic agent. 4. The apparatus of claim 1, further comprising a light source. 5. The apparatus of claim 4, further comprising a detector operable to detect light. 6. An apparatus comprising: a sensor;a light source located within a flow path; anda processor coupled to the sensor, the processor configured to: receive a signal associated with the sensor that is associated with a first therapeutic agent introduced into the flow path;calculate a volume of the first therapeutic agent, the calculation based on the signal representing the sensor;receive a signal associated with the sensor from a second therapeutic agent introduced into the flow path;calculate a volume of the second therapeutic agent, the calculation based on the signal representing the sensor; andoutput a signal associated with a comparison of the volume of the first therapeutic agent and the volume of the second therapeutic agent. 7. The apparatus of claim 6, further comprising a detector operable to detect light. 8. The apparatus of claim 6, further comprising dose-selection means coupled to the processor, and wherein the output signal is received by the dose-selection means, and wherein the output signal includes information useful for metering a third therapeutic agent. 9. The apparatus of claim 6, wherein the output signal comprising information for metering a third therapeutic agent. 10. An apparatus comprising: a sensor;a light source located within a flow path;a detector operable to detect light; anda processor coupled to the sensor, the processor configured to: receive a signal associated with the sensor that is associated with a first therapeutic agent introduced into the flow path;calculate a volume of the first therapeutic agent, the calculation based on the signal representing the sensor;receive a signal associated with the sensor from a second therapeutic agent introduced into the flow path;calculate a volume of the se... | DEKA Products Limited Partnership | G01F013/00 A61B005/08 A61M015/00 G01N015/02 G01N015/14 G01N021/53 G01F001/34 G01N001/22 A61M016/00 G01N021/47
US10713766 | B2 | 2020-07-14 | Automatic alignment of a contrast enhancement system | 2020-01-14 | An apparatus and method for insuring the proper alignment of a detected vein pattern and a projected vein pattern are disclosed. The apparatus enhances the visual appearance of veins so that an error that can lead to improper patient care or injury can be avoided. | 1. A vein imaging system comprising: a light source that illuminates a field of view with a first wavelength of light comprising an infrared wavelength, to create an image contrast of one or more veins formed by differential amounts of absorption and reflection of said first wavelength of light by the one or more veins and surrounding tissue in the field of view;an alignment card comprising a pattern formed of a material that emits a second wavelength of light when exposed to a third wavelength of light;a camera configured to capture said image contrast of the one or more veins at said first wavelength of light;a projector configured to receive and to project said captured image contrast onto the field of view using said third wavelength of light;wherein said camera is further configured to capture said second wavelength of light emitted by said pattern when positioned in the field of view, and to capture said third wavelength of light reflected from said pattern, said camera further configured to distinguish said third wavelength of light reflected by said pattern, from said second wavelength of light emitted by said pattern; andan image processor configured to align said projection of said image contrast with said image contrast formed by the differential absorption and reflection, using said distinction between said captured reflection of said pattern at said third wavelength of light, and said captured emission from said pattern as said second wavelength of light. 2. The vein imaging system according to claim 1 wherein said pattern comprises a fluorescent material. 3. The system according to claim 1 wherein said pattern comprises a known pattern. 4. The vein imaging system according to claim 1 wherein said third wavelength of light comprises a human visible wavelength of light. 5. The vein imaging system according to claim 1 wherein said third wavelength of light comprises a green wavelength of light. 6. The vein imaging system according to claim 5 wherein said second wavelength of light comprises a red wavelength of light. 7. The vein imaging system according to claim 6 wherein said green wavelength of light of said projected image contrast is projected onto the field of view outside of the one or more veins. 8. The vein imaging system according to claim 1 wherein said third wavelength of light and said infrared first wavelength of light are alternately projected and illuminated, respectively. 9. The vein imaging system according to claim 8 wherein said third wavelength of light is projected at a higher duty cycle than said illumination with said infrared wavelength. 10. A method of aligning a projection of an image of subcutaneous veins with the imaged veins, said method comprising: illuminating a field of view with a first wavelength of infrared light;capturing the first wavelength of infrared light reflected from the field of view as an image contrast formed by differential absorption and reflection by the veins and surrounding t... | AccuVein, Inc. | G06T005/00 H04N009/31 A61B005/00 H04N005/33
US10721569 | B2 | 2020-07-21 | Removable middle ear implant sensor | 2018-08-24 | A middle ear implant includes a first interface portion configured to interface with a first structure of a middle ear of a patient, a second interface portion configured to interface with a second structure of the middle ear of the patient, a shaft that connects the first and second interface portions, a carrier plate removably mounted in one of the first or second interface portions, and a removable sensor disposed at one end of the shaft, between the shaft and one of the first interface portion or the second interface portion. The removable sensor is configured to provide a DC signal output indicative of static pressure on the sensor based on placement of the sensor between the first and second structures, and provide an AC signal output indicative of a frequency response of the implant. The removable sensor is disposed at a portion of the carrier plate. | 1. A method of employing a removable sensor for providing feedback on implant placement during surgical procedures for a middle ear implant, the removable sensor comprising a sensing layer configured to have a first electrical lead contact a top surface of the sensing layer and a second electrical lead contact a bottom surface of the sensing layer, said method comprising: placing the removable sensor on a carrier plate that is insertable within a portion of the implant;installing the carrier plate into the portion of the implant with communication to a test set;placing the implant in the middle ear of a patient;generating electrical impulses based on mechanical forces imparted on the sensing layer of the implant;detecting, based on the generated electrical impulses, a DC component at a meter indicative of static pressure placed on the removable sensor based on its placement in the middle ear;detecting, based on the generated electrical impulses, an AC component at the meter indicative of frequency response of the implant;removing the carrier plate from the implant to enable removal of the removable sensor from the carrier plate; andreinstalling the carrier plate into the implant. 2. The method of claim 1, further comprising adjusting placement of the implant based on the detected DC and AC components. 3. The method of claim 2, further comprising: evaluating at least one of the detected DC component and the detected AC component relative to a respective predefined range to determine whether the placing of the implant results in the at least one of the detected DC component and the detected AC component being within the respective predefined range. | The Johns Hopkins University | H04R025/00 A61B005/00 A61F002/18
US10726547 | B2 | 2020-07-28 | Fundus image capturing | 2018-05-16 | An apparatus for producing a fundus image includes: a processor and a memory; an illumination component including a light source and operatively coupled to the processor; a camera including a lens and operatively coupled to the processor, wherein the memory stores instructions that, when executed by the processor, cause the apparatus to capture fundus images and provide controls for re-imaging the fundus. | 1. A method for imaging a fundus of a patient, comprising: receiving an eye input, the eye input including an indication of a given eye of the patient to image;receiving a focus input, the focus input indicating an image point focus of a portion of the given eye;capturing a first image of the given eye at the image point focus;presenting a control on a graphical user interface, the control configured to initiate an image acquisition sequence of the given eye;after receiving selection of the control, capturing a second image of the given eye at the image point focus;storing both the first image and the second image; anddisplaying an exam summary interface providing a summary of image captures, the summary of image captures providing an indication of a number of images captured for each given eye and the image point focus for each of the images captured. 2. The method according to claim 1, further comprising: displaying the first image on a review screen of the graphical user interface, the review screen additionally including the control. 3. The method according to claim 1, further comprising: analyzing the first image to determine a first image quality score; anddisplaying the first image quality score on the review screen. 4. The method according to claim 1, wherein the image acquisition sequence includes displaying an acquisition interface, the acquisition interface enabling a user to manually or automatically initiate image capture; and wherein capturing a second image occurs without deleting the first image. 5. The method according to claim 4, further comprising: setting a plurality of image capture parameters for capturing the first image, wherein the plurality of image capture parameters are also applied, without requesting user input, during capturing the second image. 6. The method according to claim 1, further comprising: after capturing either the first image or the second image, determining whether a preceding image has already been captured in a script for the given eye;if no preceding image has already been captured in the script for the given eye, enabling the control on the graphical user interface; andif the preceding image has already been captured in the script for the given eye, disabling the control on the graphical user interface. 7. The method according to claim 1, further comprising: determining if patient context is present; andif patient context is present, enabling a save control in the exam summary interface, the save control configured to initiate saving any captured images in a patient electronic medical record. 8. A method for capturing one or more images of a fundus of a patient, the method comprising: on a display, providing a summary interface, the summary interface including:an indication of the number of images captured for each of a right eye and a left eye of the patient;an indication of a focus type for each image captured for the right eye and the left eye of the patient; andfor each captured image, a... | WELCH ALLYN, INC. | G06K009/00 G06T007/00 G06F003/01 G16H010/60 G16H030/20 G16H040/60 A61B003/12 A61B003/00 G16H030/40 A61B005/1455
US10709330 | B2 | 2020-07-14 | Ambulatory medical telemetry device having an audio indicator | 2005-11-04 | An ambulatory medical telemetry device (10) is provided. This device includes at least one sensor (18) for detecting at least one physiological parameter of a patient and a housing that is securable to the patient. A circuit (50) is located in the housing for receiving and processing a signal representative of the physiological parameter from the sensor to generate recordable physiological data. An audible indicator (16) is located in the housing and operationally coupled to the circuit for generating an audio signal indicating at least one operational state of the device. | 1. An ambulatory medical telemetry device, comprising: at least one sensor configured for detecting at least one physiological parameter of a patient;a housing securable to the patient;a circuit located in said housing and configured for receiving and processing a signal representative of the physiological parameter from the at least one sensor to generate recordable physiological data;at least one processor located in said housing and connected to the circuit, said at least one processor programmed to perform a self-test of an overall operation and each of a plurality of operational states of the ambulatory medical telemetry device, the plurality of operational states including at least three operational states; andan audible indicator located in said housing and operationally coupled to said circuit and said at least one processor to generate an audio signal indicating at least one operational state of the device including a successful completion of the self-test indicating the operational state is operating correctly for each of the plurality of operational states of the ambulatory medical telemetry device;wherein said audible indicator is configured to generate a plurality of different tones for indicating each of the overall operation and each operational state of the ambulatory medical telemetry device, each different tone corresponding to a different operational state of the device, the plurality of different tones further including a tone indicating the successful completion of the self-test for each of the plurality of operational states of the ambulatory medial telemetry device. 2. The device of claim 1, wherein said audio signal generated by the audible indicator further indicates physiological-related information. 3. The device of claim 1, further comprising: a wireless receiver operationally coupled to the circuit and the audible indicator and located in said housing for receiving a wireless paging signal such that the audible indicator, in response to the wireless paging signal, generates an audio paging signal. 4. The device of claim 3, wherein said audible indicator generates a plurality of different tones indicative of each different operational state and the audio paging signal. 5. The device of claim 1, wherein one of the operational states of the device indicates a successful completion of a test result to obtain the physiological data. 6. The device of claim 1, wherein each of said plurality of tones differs in at least one characteristic selected from the group consisting of frequency, duration, intervals between tones, intensity, and tunes. 7. The device of claim 1, wherein at least one of the tones represents positive feedback, and at least one of the tones represents negative feedback in response to the self-test for the overall operation and each operational state of the device. 8. The device of claim 1, wherein the physiological parameter detected by the sensor includes at least one of: electrocardiogram/respirat... | KONINKLIJKE PHILIPS N.V. | A61B005/00 G06F019/00 G08B021/02 G08B021/18 A61B005/02 G16H040/63 G16H040/40 G08B021/00
US10709798 | B2 | 2020-07-14 | Crystallization process of tricyclic indole derivatives | 2013-11-28 | A composition comprising a tricyclic indole compound. The composition has a higher purity and better impurity profile than known compositions comprising said tricyclic indole compound and as a consequence has superior properties, particularly when said compound is destined for use in vivo as a therapeutic or diagnostic agent. | 1. A method of obtaining a composition comprising a compound of Formula Ia:wherein: R1 is a halogen;R2 is C1-3 alkoxy;R7 and R8 are C1-6 alkyl; andA1 is a bond or C1-10 alkylene, and R9 is a leaving group, or R9 is the group —O—R10 wherein R10 is hydrogen, C1-3 alkyl, C3-6 aryl, C7-10 arylalkyl, or a hydroxyl protecting group,wherein said composition comprises no more than 1% of a compound of Formula IIa:wherein each of R2, R7, R8, R9 and A1 are as defined for Formula Ia, the method comprising:crystallizing the composition from an organic solvent in the presence of the catalytic amount of N,N-diisopropyl ethylamine, triethyl amine, or diethyl amine to obtain said composition, wherein the composition comprises the compound of Formula Ia and no more than 1% of the compound of Formula IIa as a by-product. 2. The method of claim 1, wherein said organic base is selected from N,N-diisopropyl ethylamine or diethylamine. 3. The method of claim 1, wherein: R1 is halogen;R2 is methoxy;R7 and R8 are C1-3 alkyl;R9 is benzyloxy; andA1 is C1-3 alkylene. 4. The method of claim 1, wherein the composition comprises no more than 0.5% of the compound of Formula IIa. 5. The composition resulting from the method of claim 1, wherein R9 is a leaving group. | GE HEALTHCARE LIMITED | A61B005/00 A61B008/00 A61B010/00 A61K049/00 C07D209/88
US10709380 | B2 | 2020-07-14 | Method and apparatus for pulse signal analyzing | 2017-02-07 | A method and an apparatus for pulse signal analyzing are provided. The method includes: receiving a pulse signal; decomposing the pulse signal into a plurality of characteristics signal; generating a spectrum for each of the plurality of characteristics signal by using a spectrum projection; obtaining a plurality of quantized data for each of the plurality of characteristics signal by quantizing the spectrum corresponding to each of the plurality of characteristics signal; and determining a physiological condition corresponding to the pulse signal according to the plurality of quantized data of the plurality of characteristics signal. | 1. A method for pulse signal analysing, performed by a pulse signal analyzing device comprising a processing unit, the method comprising: detecting a pulse signal within a period of time by a detection unit;receiving, by the processing unit, the pulse signal;decomposing, by the processing unit, the pulse signal into a plurality of characteristics signals;performing, by the processing unit, a spectrum projection to each of the characteristics signals to generate a spectrum;obtaining, by the processing unit, a plurality of quantized data corresponding to each of the characteristics signals by quantizing the spectrum corresponding to each of the characteristics signals,wherein the plurality of quantized data comprise an energy density and an average period corresponding to each of the characteristics signals, and calculates an energy density and the average period of the spectrum corresponding to each of the characteristics signals, the energy density corresponding to a specific spectrum is obtained according to a total number of time points in the specific spectrum, an energy intensity corresponding to a time point within a time range of the specific spectrum,wherein an average period corresponding to the specific spectrum is obtained according to a time, a function of natural logarithm of the specific spectrum and the time,determining, by the processing unit, a physiological condition corresponding to the pulse signal according to the plurality of quantized data corresponding to the characteristics signals;comparing, by the processing unit, the quantized data corresponding to each of the characteristics signals with reference data to determine the physiological condition corresponding to the pulse signal and generates a comparison result;outputting, by the processing unit, the physiological condition corresponding to the pulse signal based on the comparison result. 2. The method for pulse signal analysing as claimed in claim 1, wherein the characteristics signals are combined to form the pulse signal. 3. The method for pulse signal analysing as claimed in claim 1, wherein the characteristics signals respectively correspond to a physiological significance. 4. The method for pulse signal analysing as claimed in claim 1, wherein the physiological condition at least comprises one of a plurality of traditional Chinese medicine pulse conditions. 5. A pulse signal analyzing device, comprising: a detection unit, configured to detect a pulse signal within a period of time; anda processing unit, whereinthe processing unit is configured to receive the pulse signal,the processing unit is configured to decompose the pulse signal into a plurality of characteristics signals,the processing unit is configured to perform a spectrum projection to each of the characteristics signals to generate a spectrum,the processing unit is configured to obtain a plurality of quantized data corresponding to each of the characteristics signals by quantizing the spectrum corres... | Winbond Electronics Corp. | A61B005/00 A61B005/024 A61B005/02
US10724943 | B2 | 2020-07-28 | Device for detecting an analyte in a bodily fluid | 2018-07-05 | A device is proposed for detecting at least one analyte in a bodily fluid. The device comprises at least one test element with at least one two-dimensional evaluation region. The device furthermore comprises at least one spatially resolving optical detector having a plurality of pixels. The detector is designed to image at least part of the test element onto an image region. In the process, at least part of the evaluation region is imaged onto an evaluation image region. The detector is matched to the test element such that a predetermined minimum number of pixels is provided for each dimension within the evaluation image region. The pixels are arranged in a two-dimensional matrix arrangement. The matrix arrangement has pixel rows and pixel columns, wherein the pixel rows are arranged substantially parallel to a longitudinal direction of the evaluation region and/or of the evaluation image region. | 1. A device for detecting an analyte in a bodily fluid, comprising: a microsampler;a test element associated with the microsampler and having a two-dimensional evaluation region; anda spatially resolving optical detector having a longitudinal side and a transverse side and a plurality of pixels, the detector being configured to image at least part of the microsampler onto an image region and to image at least part of the evaluation region onto an evaluation image region. 2. A device for detecting an analyte in a bodily fluid as claimed in claim 1, wherein the detector is matched to the test element such that a predetermined minimum number of pixels is provided for each dimension within the evaluation image region and the pixels are arranged in a two-dimensional matrix having pixel rows and pixel columns. 3. A device for detecting an analyte in a bodily fluid as claimed in claim 1, wherein the pixels are image-sensitive individual sensors. 4. The device of claim 1, further comprising a lancet element having a capillary, the capillary being configured to take up bodily fluid and the device being configured to cause the capillary to approach the test element to transfer bodily fluid onto the test element. 5. The device of claim 4, wherein the detector is configured to image the lancet and/or the capillary onto the image region. 6. The device of claim 4, wherein, the evaluation region is a region of the test element in which an optically detectable change occurs as a result of transferring the bodily fluid onto the test element. 7. The device of claim 1, wherein the device is configured to recognize the evaluation region automatically. 8. The device as claimed in claim 7, wherein the device is configured to recognize the evaluation region according to one or more of the following: a pattern recognition method, wherein there is at least one element selected from the group consisting of a lancet element and a capillary imaged in the image region, the element is recognized and an extrapolation of the element onto the test element is identified as the evaluation region;a signal-change method, wherein a region of the test element within which an optically detectable change occurs as a result of a transfer of the bodily fluid onto the test element is identified as the evaluation region. 9. The device in claim 8, wherein data from pixels outside the evaluation image region is discarded immediately after the evaluation image region is located within the image region. 10. The device as claimed in claim 1, wherein the evaluation region has a longitudinal side and a transverse side, wherein the detector has at least 3 pixel rows in the direction of the transverse side of the evaluation region and at least 1 pixel column in the direction of the longitudinal side of the evaluation region. 11. The device as claimed in claim 10, wherein the longitudinal side of the evaluation region is aligned parallel to a capillary of the device and the transverse side ... | Roche Diabetes Care, Inc. | G01N021/17 G01N021/84 A61B005/151 A61B010/00 H01L027/146 G01N021/78
US10722306 | B2 | 2020-07-28 | System for tracking guidewire with ray tracing capability | 2016-10-10 | A method, including registering a computerized tomography (CT) scan of a patient with a tracking system configured to track a location and an orientation of a sensor in the patient, and displaying the CT scan on a screen. The method further includes tracking coordinates of a probe having the sensor inside the patient using the tracking system. Based on the tracked coordinates, an icon representative of the location of the sensor is overlaid on the CT scan on the screen. The method also includes overlaying on the CT scan a ray, issuing from the icon in a direction corresponding to the orientation of the sensor, and traversing the screen to a termination point chosen responsively to a gray level distribution of the CT scan. | 1. A method, comprising: registering a computerized tomography (CT) scan of a patient with a tracking system, the tracking system comprising a processor configured to track a location and an orientation of a sensor in the patient and a field generator;displaying the CT scan on a screen;operating the field generator to produce a field configured to interact with the sensor and indicate a location and orientation of the sensor within the field;tracking coordinates of a probe comprising the sensor inside the patient using the tracking system;based on the tracked coordinates, overlaying on the CT scan on the screen an icon representative of the location of the sensor;overlaying on the CT scan a ray, issuing from the icon in a direction corresponding to the orientation of the sensor, and traversing the screen to a termination point chosen responsively to a gray level distribution of the CT scan at the termination point; andupdating the icon representative of the location of the sensor, the direction of the ray, and the termination point of the ray in response to a change in the location or orientation of the sensor. 2. The method according to claim 1, wherein the CT scan comprises at least one substantially black portion and at least one non-black portion, and wherein the gray level distribution is indicative of a local tissue density in the CT scan. 3. The method according to claim 1, further comprising setting a predefined threshold that indicates a set of gray levels that cannot be traversed by the ray and a set of gray levels that can be traversed by the ray, and wherein overlaying the ray comprises terminating the ray at a pixel where a gray level value of the displayed CT scan crosses the predefined threshold. 4. The method according to claim 1, and comprising directing laser radiation from the probe, in response to overlaying the ray to the termination point, to a location of a desired ablation target in the patient corresponding to the termination point. 5. The method according to claim 1, wherein the tracking system comprises a magnetic tracking system and the field generator comprises a magnetic field generator. 6. The method according to claim 1, wherein the CT scan comprises a CT scan of nasal sinuses of the patient, and wherein the probe comprises a guidewire configured and dimensioned to be inserted into the nasal sinuses. 7. Apparatus, comprising: a screen configured to display a computerized tomography (CT) scan of a patient;a probe comprising a sensor configured to be inserted into the patient;a tracking system comprising a processor configured to track a location and an orientation of the sensor and a field generator; andthe processor is configured to:receive the CT scan;register the CT scan with the tracking system,operate the field generator to produce a field configured to interact with the sensor and indicate a location and orientation of the sensor within the field,track coordinates of the probe using the tracking system,b... | Biosense Webster (Israel) Ltd. | A61B018/20 A61B006/03 A61B006/12 A61M025/09 A61B006/00 A61B005/00 A61B005/06 A61B034/20 A61B018/00 A61B034/10 A61B090/30 A61B034/00
US10702712 | B2 | 2020-07-07 | Tumor tracking during radiation treatment using ultrasound imaging | 2015-07-15 | Systems and methods for tracking a target volume, e.g., tumor, in real-time during radiation treatment are provided. Under one aspect, a system includes an ultrasound probe, an x-ray imager, a processor, and a computer-readable medium that stores a 3D image of the tumor in a first reference frame and instructions for causing the processor to: instruct the x-ray imager and ultrasound probe to substantially simultaneously obtain inherently registered x-ray and set-up ultrasound images of the tumor in a second reference frame; establish a transformation between the first and second reference frames by registering the 3D image and the x-ray image; instruct the ultrasound probe to obtain an intrafraction ultrasound image of the tumor; registering the intrafraction ultrasound image with the set-up ultrasound image; and track target volume motion based on the registered intrafraction ultrasound image. | 1. A method of tracking an anatomy of interest, the method comprising: receiving a first pre-acquired 3D image, of a first imaging modality, of the anatomy of interest of a patient in a first reference frame;receiving a second pre-acquired 3D image, of a second modality, of the anatomy of interest of the patient in the first reference frame;registering the first pre-acquired 3D image with the second pre-acquired 3D image;obtaining a set-up ultrasound image of the anatomy of interest of the patient in a second reference frame;establishing a transformation between the first and second reference frames by registering the second pre-acquired 3D image and the set-up ultrasound image;obtaining an intrafraction ultrasound image of the anatomy of interest of the patient;registering the intrafraction ultrasound image with the set-up ultrasound image; andtracking motion of the anatomy of interest based on the registered intrafraction ultrasound image. 2. The method of claim 1, wherein the first pre-acquired 3D image comprises a CT image and the second pre-acquired 3D image comprises an MR image. 3. The method of claim 1, further comprising: determining a position and orientation of the set-up ultrasound image in the second reference frame by obtaining from a tracking device positional data about markers affixed to an ultrasound probe; andidentifying a position and orientation of the ultrasound probe in the second reference frame based on the positional data. 4. The method of claim 1, further comprising adjusting a position and orientation of a radiation source based on the tracked motion of the anatomy of interest. 5. The method of claim 1, further comprising adjusting a position and orientation of a patient treatment couch relative to a radiation source based on the tracked motion of the anatomy of interest. 6. The method of claim 1, further comprising enabling or disabling a treatment beam based on the tracked motion of the anatomy of interest. 7. The method of claim 1, further comprising adjusting a multileaf collimator to adjust delivery of a radiation treatment beam based on the tracked motion of the anatomy of interest. 8. The method of claim 1, wherein at least one of the first or second pre-acquired 3D image is used to derive a treatment plan to irradiate a target volume associated with the anatomy of interest. 9. A system for tracking an anatomy of interest, the system comprising: a memory to store: a first pre-acquired 3D image of the anatomy of interest of a patient in a first reference frame;a second pre-acquired 3D image, of a second modality, of the anatomy of interest of the patient in the first reference frame; anda registration of the first pre-acquired 3D image with the second pre-acquired 3D image; anda processor, operative coupled with the memory, to: receive, from an ultrasound probe, a set-up ultrasound image of the anatomy of interest of the patient in a second reference frame;establish a transformation between the first and ... | ACCURAY INCORPORATED | A61N005/10 A61B008/00 A61B008/08 A61B006/03 A61B005/055 A61B006/00 A61B090/00 A61B006/02 A61B006/04
US10722148 | B2 | 2020-07-28 | Fall detection devices, systems, and methods | 2018-10-24 | Provided herein are systems and devices capable of detecting an event, such as the fall of a human. The devices may include a motion sensor, a heat sensor, and a vibration sensor. The devices and systems also may include an alarm and/or communication device configured to function when the event occurs. | 1. A method for detecting an occurrence of a fall of a human, the method comprising: providing a device comprising (i) a device body having a base configured for placement on a surface, (ii) a vibration sensor configured to detect vibrations of the surface, (iii) a motion sensor comprising a receiver configured to receive transmissions from one or more radio frequency transmitters, (iv) a heat sensor comprising a thermal camera, wherein the thermal camera comprises a plurality of pixels, and (v) a processing unit configured to (a) receive one or more signals from each of the motion sensor, the heat sensor, and the vibration sensor, and (b) determine an occurrence of an event based on the one or more signals received from the motion sensor, the heat sensor, and the vibration sensor, wherein the motion sensor, the heat sensor, the vibration sensor, and the processing unit are housed by the device body;detecting a motion-to-stationary transition of the human with the motion sensor;estimating a distance of the human from the heat sensor by determining a fraction of the plurality of pixels that detect a body temperature of the human;determining whether the human is standing after the motion-to-stationary transition based on a signal from the heat sensor; anddeclaring the occurrence of the fall if—[1] (i) the human is not standing after the motion-to-stationary transition, and (ii) a signal from the vibration sensor is greater than a threshold value based on the distance of the human from the heat sensor; or[2] (i) the signal from the heat sensor indicates that the human is standing after the motion-to-stationary transition, (ii) no movement is detected by the system for at least 10 seconds after the motion-to-stationary transition, and (iii) a signal from the vibration sensor is greater than a threshold value based on the distance of the human from the heat sensor. 2. A method for detecting an occurrence of a fall of a human, the method comprising: providing a device comprising (i) a device body having a base configured for placement on a surface, (ii) a vibration sensor configured to detect vibrations of the surface, (iii) a motion sensor comprising a receiver configured to receive transmissions from one or more radio frequency transmitters, (iv) a heat sensor comprising a thermal camera, wherein the thermal camera comprises a plurality of pixels, and (v) a processing unit configured to (a) receive one or more signals from each of the motion sensor, the heat sensor, and the vibration sensor, and (b) determine an occurrence of an event based on the one or more signals received from the motion sensor, the heat sensor, and the vibration sensor, wherein the motion sensor, the heat sensor, the vibration sensor, and the processing unit are housed by the device body;detecting a motion-to-stationary transition of the human with the motion sensor;estimating a distance of the human from the heat sensor by comparing a first signal and a second signal received ... | Florida State University Research Foundation, Inc. | A61B005/11 G08B021/04 A61B005/00 A61B005/1495 A61B005/01 G08B013/19 G01P013/00 G01P015/00
US10722129 | B2 | 2020-07-28 | Imaging system for intra-operative and post-operative blood perfusion monitoring | 2018-08-17 | Embodiments described herein generally relate to devices, methods and systems for determining blood oxygenation. By applying near infrared radiation of an appropriate wavelength to the tissue and determining the absorbance at a plurality of points where the distance between the source of the near infrared radiation and the detector are known, the oxygenation state of the hemoglobin can be determined based on position in a three dimensional space. | 1. A method for tissue imaging in a transplanted tissue, sequentially comprising: positioning a near infrared spectroscopy (NIRS) device in connection with a tissue portion located on a donor location of a first body, the NIRS device being positioned for a near infrared measurement;collecting a first NIRS measurement using the NIRS device, the first NIRS measurement providing information to create a pre-transfer volumetric map regarding blood oxygenation or tissue perfusion;creating the pre-transfer volumetric map regarding blood oxygenation or tissue perfusion from the first NIRS measurement;removing the tissue portion from the donor location of the first body;transplanting the tissue portion to a recipient location of the first body or a second body;collecting a second NIRS measurement, the second NIRS measurement providing information to create a post-transfer volumetric map regarding blood oxygenation or tissue perfusion;creating the post-transfer volumetric map regarding bloody oxygenation or tissue perfusion from the second NIRS measurement; andcomparing the pre-transfer volumetric map to the post-transfer volumetric map to determine a change in blood oxygenation or tissue perfusion. 2. The method of claim 1, wherein the NIRS device is connected with the tissue portion when the tissue portion is removed from the donor location and while the tissue portion is transplanted to the recipient location. 3. The method of claim 1, wherein the NIRS device comprises a plurality of first radiation sources, a plurality of second radiation sources and a plurality of detectors connected to a support. 4. The method of claim 3, wherein the first radiation sources and the second radiation sources are multiplexed. 5. The method of claim 4, wherein the first radiation sources each deliver a first radiation, and wherein the first radiation is a radiation with a wavelength between 650 nm and 800 nm. 6. The method of claim 4, wherein the second radiation sources each deliver a second radiation, wherein the second radiation is a radiation with a wavelength between 800 nm and 1000 nm. 7. The method of claim 4, wherein the first NIRS measurement and the second NIRS measurement are selected samples from a continuous NIRS measurement. 8. The method of claim 1 wherein: the pre-transfer volumetric map comprises dimensions corresponding to the tissue portion located on the donor location of the first body; andthe post-transfer volumetric map comprises dimensions corresponding to the tissue portion located on the recipient location of the first body or the second body. | University of Houston System The Methodist Hospital | A61B005/1455 A61B005/026 A61B005/00
US10725127 | B2 | 2020-07-28 | Magnetic field measurement method and magnetic field measurement apparatus | 2019-01-29 | In a magnetic field measurement apparatus, a light source irradiates a gas cell with linearly polarized light serving as pump light and probe light in a Z axis direction, and a magnetic field generator applies, to the gas cell, a magnetic field Ax which is a time function f(t) having the amplitude A0 taking n fixed values fi (where i=1, . . . , and n), and a magnetic field Ay which is a time function g(t) having the amplitude A0 taking m fixed values gj (where j=1, . . . , and m) in each of X axis and Y axis directions. A calculation controller calculates a magnetic field C (Cx, Cy, Cz) of a measurement region using the X axis and Y axis components Ax and Ay of an artificial magnetic field A, and a spin polarization degree Mx corresponding to a measurement value W− from a magnetic sensor. | 1. A magnetic field measurement apparatus in which a first direction, a second direction, and a third direction are perpendicular to each other, the apparatus comprising: a light source that emits light;a medium through which the light passes in the third direction and that changes optical characteristics depending on a magnetic field of the measurement region;a photodetector that detects the optical characteristics;a second magnetic field generator that applies a magnetic field in the second direction to the measurement region; anda calculation controller that causes the second magnetic field generator to generate a constant magnetic field with a first level on the second direction side, a constant magnetic field with a second level on the second direction side, and a constant magnetic field with a third level on the second direction side, as the magnetic field in the second direction, and calculates the magnetic field of the measurement region using a detection result from the photodetector and the magnetic field in the second direction. | SEIKO EPSON CORPORATION | G01R033/26 A61B005/04 G01R033/032
US10712411 | B2 | 2020-07-14 | Bulk magnet structure and bulk magnet system for NMR | 2017-07-27 | The purpose of the present invention is to make it possible to use a wider uniform magnetic field space and to achieve an additional use mode for a bulk magnet structure. A bulk magnet structure according to the present invention is provided with a plurality of oxide superconducting bulk bodies arranged so that the central axes thereof are on the same line and at least one outer circumferential reinforcing ring fitted to the bulk magnet structure so as to cover the outer circumferential surfaces of the plurality of oxide superconducting bulk bodies. The plurality of oxide superconducting bulk bodies includes a columnar oxide superconducting bulk body and/or a ring-shaped oxide superconducting bulk body. At least one set of adjacent oxide superconducting bulk bodies are spaced apart from each other in the direction of the central axes thereof. The interior of the bulk magnet structure includes a space through which the central axes pass. | 1. A bulk magnet structure comprising a plurality of oxide superconducting bulk bodies arranged such that their central axes are on the same line, andat least one outer circumferential reinforcing ring fitted to cover the outer circumferential surfaces of the plurality of oxide superconducting bulk bodies,wherein the plurality of oxide superconducting bulk bodies comprise at least a columnar oxide superconducting bulk body or a ring-shaped oxide superconducting body,wherein at least one pair of adjacent oxide superconducting bulk bodies are spaced apart from each other in the central axis direction and the bulk magnet structure has a space inside thereof through which the central axis passes, andwherein the space penetrates from a portion of a lateral circumferential portion of the bulk magnet structure to another portion of the lateral circumferential portion, in a direction perpendicular to the central axis. 2. A bulk magnet structure comprising a plurality of oxide superconducting bulk bodies, andat least one outer circumferential reinforcing ring fitted to cover the outer circumferential surfaces of the plurality of layered oxide superconductive bulk bodies,a spacer layered together with the oxide superconducting bulk bodies,wherein the plurality of oxide superconducting bulk bodies comprise at least a columnar oxide superconducting bulk body or a ring-shaped oxide superconducting bulk body,wherein the spacer has a space formed so as to penetrate from a portion of the lateral circumferential portion to another portion of the lateral circumferential portion, and the spacer is layered so that at least the central axis of the bulk magnet structure passes through the space, andwherein the space penetrates from a portion of the lateral circumferential portion of the bulk magnet structure to another portion of the lateral circumferential portion, in a direction perpendicular to the layering direction of the layered oxide superconducting bulk bodies. 3. A bulk magnet structure comprising a plurality of oxide superconducting bulk bodies,at least one outer circumferential reinforcing ring fitted to cover the outer circumferential surfaces of the plurality of layered oxide superconducting bulk bodies, anda space,wherein the plurality of oxide superconducting bulk bodies comprise a ring-shaped oxide superconducting bulk body, andwherein the space is capable of forming a magnetic field space by the bulk magnet structure, placing a sample and/or device and communicating with the outside of the bulk magnet structure, in two directions of a direction perpendicular to the layering direction and the layering direction. 4. The bulk magnet structure according to claim 1, wherein the plurality of oxide superconducting bulk bodies comprise a ring-shaped oxide superconducting body, and the space is capable of communicating with the outside of the bulk magnet structure, in a direction perpendicular to the central axis. 5. The bulk magnet structure according to ... | NIPPON STEEL CORPORATION | G01V003/00 G01R033/385 A61B005/055 H01F007/20 G01R033/38 H01F006/00 H01F006/04 H01F006/06
US10705603 | B2 | 2020-07-07 | System and method for detecting invisible human emotion in a retail environment | 2017-02-08 | A system for detecting invisible human emotion in a retail environment is provided. The system comprises a camera and an image processing unit. The camera is configured in a retail environment to capture an image sequence of a person before and during when a price of a product or service becomes visible. The image processing unit is trained to determine a set of bitplanes of a plurality of images in the captured image sequence that represent the hemoglobin concentration (HC) changes of the person, and to detect the person's invisible emotional states based on HC changes. The image processing unit is trained using a training set comprising a set of subjects for which emotional state is known. | 1. A system for detecting invisible human emotion in a retail environment within which a product is displayed in a product display to a person, the system comprising: a price display device for selectively displaying at least one price of the product, pursuant to a point of sale event;a camera configured to capture an image sequence of the person before and during the point of sale event; anda processing unit trained to determine a set of bitplanes of a plurality of images in the captured image sequence that represent the hemoglobin concentration (HC) changes of the person, to detect the person's invisible emotional states based on the HC changes, and to output the detected invisible emotional states, the processing unit being trained using a training set comprising HC changes of subjects with known emotional states. 2. The system of claim 1, wherein detecting the person's invisible emotional states based on the HC changes comprises generating an estimated statistical probability that the person's emotional state conforms to a known emotional state from the training set, and a normalized intensity measure of such determined emotional state. 3. The system of claim 1, wherein the point of sale event comprises the price display device displaying a price. 4. The system of claim 1, wherein the point of sale event comprises the price display device temporarily displaying a discounted price. 5. The system of claim 1, wherein the camera is integral to the price display device. 6. The system of claim 1, further comprising a motion sensor to detect motion in a region proximal the product display, and to, upon detecting motion in the region, trigger the camera to capture the image sequence and the price display device to display the price. 7. The system of claim 1, wherein the processing unit is configured to receive locations of the camera and the product, to perform gaze tracking to analyze the image sequence to determine whether the person is looking at the product during the point of sale event, and to discard the image sequence if the person is not looking at the product during the point of sale event. 8. The system of claim 1, further comprising a notification system to display a notification indicative of the person's detected invisible emotional states. 9. A method for detecting invisible human emotion in a retail environment within which a product is displayed in a product display to a person, the method comprising: selectively displaying, by a price display device, at least one price of the product, pursuant to a point of sale event;capturing, by a camera, an image sequence of the person before and during the point of sale event; anddetermining, by a processing unit, a set of bitplanes of a plurality of images in the captured image sequence that represent the hemoglobin concentration (HC) changes of the person, detecting the person's invisible emotional states based on the HC changes, and outputting the detected invisible emotional states,... | NuraLogix Corporation | G06F003/01 G06Q030/02 G06T007/20 H04N007/18 A61B005/16 A61B003/113 A61B005/00 A61B005/145 G06K009/78 A61B005/1455 G06K009/46 G06N003/08 G06N003/04 A61B005/08 A61B005/0402 G06K009/00
US10722180 | B2 | 2020-07-28 | Deep learning-based diagnosis and referral of ophthalmic diseases and disorders | 2018-10-15 | Disclosed herein are systems, methods, devices, and media for carrying out medical diagnosis of ophthalmic diseases and conditions. Deep learning algorithms enable the automated analysis of ophthalmic images to generate predictions of comparable accuracy to clinical experts. | 1. A computer-implemented method for detecting an ophthalmic disease, disorder, or condition, the method comprising: generating a machine learning classifier that classifies medical data, including image data, into one of a plurality of classifications, wherein the machine learning classifier is generated bytraining a convolutional neural network using non-domain images that are not labeled with any of the plurality of classifications, andretraining upper layers of the convolutional neural network using domain images that are ophthalmic images labeled with one or more of the plurality of classifications, while retaining lower layers of the convolutional neural network that were trained using the non-domain images;obtaining an ophthalmic image of an individual;evaluating the ophthalmic image using the machine learning classifier to generate a determination of the ophthalmic disease, disorder, or condition, the determination having a sensitivity of at least 90% and a specificity of at least 90% when tested against an independent data set of at least 200 samples; andproviding the determination to the individual or a third party. 2. The method of claim 1, wherein the ophthalmic image of the individual comprises a retinal image. 3. The method of claim 2, wherein the retinal image is an optical coherence tomography (OCT) image. 4. The method of claim 2, wherein obtaining the retinal image comprises capturing the retinal image of the individual with an electronic device comprising a camera and a portable device comprising an imaging component, wherein the portable device and the electronic device are positioned to align the imaging component with the camera. 5. The method of claim 1, wherein evaluating the ophthalmic image comprises uploading the ophthalmic image to a cloud network for remote analysis of the ophthalmic image using the machine learning classifier. 6. The method of claim 1, wherein the ophthalmic disease, disorder, or condition is a retinal disease or condition. 7. The method of claim 6, wherein the retinal disease or condition is age-related macular degeneration (AMD), diabetic macular edema (DME), choroidal neovascularization (CNV), or diabetic retinopathy. 8. The method of claim 1, further comprising providing a recommendation for treatment or further testing to the individual or the third party. 9. A computer-implemented system comprising: a server comprising at least one processor configured to generate a machine learning classifier that classifies medical data, including image data, into one of a plurality of classifications, wherein the machine learning classifier is generated by training a convolutional neural network using non-domain images that are not labeled with any of the plurality of classifications, andretraining upper layers of the convolutional neural network using domain images that are ophthalmic images labeled with one or more of the plurality of classifications, while retaining lower layers of the convolution... | - | A61B005/00 G16H030/40 G06N020/00 A61B003/00 A61B003/10 A61B003/12 A61B003/14 A61B005/053 A61B005/024 A61B005/021
US10716501 | B2 | 2020-07-21 | System and method for classification and quantitative estimation of cognitive stress | 2018-03-06 | This disclosure relates generally to stress classification and quantification, and more particularly to system and method for classification and quantitative estimation of cognitive stress from analysis of keystrokes and signals derived from physiological sensors. In one embodiment, a method includes obtaining, while a user is engaged in performance of a task, physiological signals from physiological sensors accessible to the user. Feature sets are identified from the physiological signals which correlate with cognitive stress experienced by the user. Using a regression model, a stress indicator metric comprising a quantitative estimate of the cognitive stress is predicted. The regression model is trained using the feature sets and independently determined quantitative estimates of cognitive stress used as a ground truth to output the value of the stress indicator metric. The ground truth is determined from keystroke data associated with the performance of keyboard-based tasks comprising navigation of moving objects to the target. | 1. A processor-implemented method for classification and quantitative estimation of cognitive stress in real-time, the method comprising: obtaining, while a user is engaged in performance of a task, a plurality of non-invasive physiological signals from a plurality of physiological sensors accessible to the user, via one or more hardware processors;identifying a plurality of feature sets from the plurality of non-invasive physiological signals, via the one or more hardware processors, the plurality of feature sets correlating with the cognitive stress experienced by the user while engaging in the performance of the task; andpredicting, using a regression model, a stress indicator metric comprising a quantitative estimate of the cognitive stress experienced by the user during the performance of the task, via the one or more hardware processors,wherein, the regression model is trained using the plurality of feature sets, and independently determined quantitative estimates of cognitive stress used as a ground truth to output value of the stress indicator metric, andwherein, the ground truth is determined from keystroke data associated with the performance of keyboard-based tasks comprising navigation of one or more moving objects to at least one target from a plurality of targets. 2. The method of claim 1, wherein the plurality of non-invasive physiological signals comprises Electroencephalogram (EEG) signal to estimate the neural responses of the user, Photoplethysmogram (PPG) to analyze cardiac parameters of the user, and Electrodermal or Galvanic Skin Response (GSR) to determine changes due to variations in skin conductance of the user, in real-time. 3. The method of claim 2, wherein identifying the plurality of feature sets from the plurality of non-invasive physiological signals comprises: extracting a feature set individually from each non-invasive physiological signal of the plurality of non-invasive signals to obtain the plurality of feature sets; andperforming feature reduction on each feature set of the plurality of feature sets to obtain an optimum feature set. 4. The method of claim 1, further comprising preprocessing the plurality of non-invasive physiological signals in order to reduce corruption of the plurality of non-invasive physiological by noise, reduce intra-user and inter-user variability and decrease computational effort for identifying the plurality of feature steps and predicting the stress indicator metric. 5. The method of claim 1, wherein determining the quantitative estimate of the cognitive stress from the keystrokes comprises: obtaining the keystroke data during performance of at least one keyboard-based task involving navigating a moving object to a target object out of a plurality of target objects;determining relative position of the moving object with respect to the ultimate target object during each epoch of the keyboard based task, an epoch associated with the moving object comprises a duration of time availa... | Tata Consultancy Services Limited | A61B005/16 A61B005/0476 A61B005/024 A61B005/053 G16H050/20 A61B005/11 A61B005/00 G06N005/02 G06N020/00 A61B005/0205 A61B005/048
US10702178 | B2 | 2020-07-07 | Catheter with high-density mapping electrodes | 2018-10-03 | High-density mapping catheters with an array of mapping electrodes are disclosed. These catheters can be used for diagnosing and treating cardiac arrhythmias, for example. The catheters are adapted to contact tissue and comprise a flexible framework including the electrode array. The array of electrodes may be formed from a plurality of columns of longitudinally-aligned and rows of laterally-aligned electrodes. | 1. A basket catheter comprising: an elongated catheter shaft including a proximal end and a distal end;a flexible basket catheter, including a plurality of splines, coupled to the distal end of the catheter shaft and configured to conform to tissue when extended into a deployed configuration;a plurality of electrodes mounted to the splines, the plurality of electrodes organized into triangular-shaped cliques along each of the splines; andwherein each of the plurality of splines include a plurality of pads that extend transverse to a length of the splines and are configured to host one or more of the plurality of electrodes, the pads on adjacent splines are longitudinally offset to facilitate interleaving of the pads when the flexible basket catheter is actuated into a contracted configuration. 2. The basket catheter of claim 1, wherein each of the splines are configured to nest with adjacent splines when the flexible basket catheter is actuated into a contracted configuration. 3. The basket catheter of claim 1, wherein the electrodes are spot electrodes and the splines include flexible electronic circuits that are communicatively and mechanically coupled to the plurality of electrodes. 4. The basket catheter of claim 1, wherein the distance between the electrodes in each clique is constant in contracted, deployed, and intermediate configurations of the basket catheter. 5. The basket catheter of claim 1, wherein the center-to-center distance between the electrodes in the cliques is between 0.5 and 4 millimeters. 6. The basket catheter of claim 1, wherein all the electrodes of a clique are coupled to the same spline, and the electrodes of the clique are configured to collect electrical signals emanating from the tissue in contact with the basket catheter, the electrical signals of the electrical characteristics of the tissue, the electrical characteristics of the tissue being independent of the orientation of the basket catheter relative to the tissue. 7. The basket catheter of claim 1, wherein the electrodes of the cliques form a right triangle. 8. The basket catheter of claim 1, wherein the electrodes of the cliques form an isosceles triangle with a vertex angle between 30°-140°. 9. The basket catheter of claim 1, wherein each of the splines is substantially saw-tooth shaped. 10. The basket catheter of claim 1, wherein the distance between each electrode of a clique is equal. | St. Jude Medical, Cardiology Division, Inc. | A61B005/042 A61B005/00 A61B018/14 A61B018/00 A61N001/362
US10722149 | B2 | 2020-07-28 | Real-time biofeedback rehabilitation tool guiding and illustrating foot placement for gait training | 2017-07-26 | A biofeedback rehabilitation display for gait training visible to the subject including a plurality of foot placement targets visible to the subject on the display and moving along the display at a speed proportional to the desired gait speed of the subject and in a direction along the display representative of the direction of locomotion of the subject; a pair of subject foot icons visible to the subject illustrating real-time foot positioning of the subject throughout the gait pattern of the subject, wherein the foot icons distinguish between stance phase and swing phase throughout the gait pattern of the subject; and at least one real-time proximity measurement visible to the subject for each stance phase providing the subject with an indication of the proximity of each stance phase of each foot with an intended foot placement target. | 1. A biofeedback rehabilitation display for gait training visible to the subject comprising: a plurality of foot placement targets configured to be visible to the subject on the display and configured for moving along the display at a speed proportional to the desired gait speed of the subject and in a direction along the display representative of the direction of locomotion of the subject;a pair of subject foot icons configured to be visible to the subject illustrating real-time foot positioning of the subject at all times throughout the gait pattern of the subject, wherein the foot icons distinguish between stance phase and swing phase throughout the gait pattern of the subject; andat least one real-time proximity measurement visible to the subject for each stance phase providing the subject with an indication of the proximity of the stance phase foot with an intended foot placement target. 2. The biofeedback rehabilitation display for gait training according to claim 1 wherein a stance phase is determined by the detection of a heel strike of the subject and the real-time proximity measurement is determined by the radial distance from a fixed point of the stance phase foot position to a fixed point on the nearest foot placement target. 3. The biofeedback rehabilitation display for gait training according to claim 1 wherein the real-time proximity measurement visible to the subject includes a visible indication when the proximity measurement is less than an acceptable threshold. 4. The biofeedback rehabilitation display for gait training according to claim 3 wherein the real-time proximity measurement visible to the subject includes a visible indication when the proximity measurement is greater than an unacceptable threshold. 5. The biofeedback rehabilitation display for gait training according to claim 4 wherein the real-time proximity measurement visible to the subject includes a visible indication when the proximity measurement is between the acceptable threshold and the unacceptable threshold. 6. The biofeedback rehabilitation display for gait training according to claim 3 wherein the foot placement target includes a visible representation of an associated proximity measurement. 7. The biofeedback rehabilitation display for gait training according to claim 1 wherein the real-time proximity measurement visible to the subject for each stance phase provide the subject with an indication of the proximity of each stance phase foot with an intended foot placement target in both the direction of locomotion of the subject and lateral to the direction of locomotion of the subject. 8. The biofeedback rehabilitation display for gait training according to claim 1 wherein a stance phase is determined by the detection of a heel strike of the subject and the stance phase illustrates foot orientation of the subject's foot at the time of the detected heel strike. 9. The biofeedback rehabilitation display for gait training according to claim 8 further... | VICTORIA UNIVERSITY | A61B005/11 A61B005/103 A61B005/00 G06F019/00 A61B005/22 A63B022/02
US10702183 | B2 | 2020-07-07 | Multi-channel reconfigurable systems and methods for sensing biopotential signals | 2016-05-12 | A biopotential monitoring device and use of such device. The device includes a configurable receiver circuit having a plurality of channels for receiving a plurality of biopotential signals from a biological tissue via a plurality of inputs coupled with the electrodes, and each channel substantially removes a DC (direct current) offset from a corresponding one of the biopotential signals and then band-pass amplifies such corresponding biopotential signal at a configurable gain and particular frequency range based on frequency control signals. The device further includes a controller circuit for receiving commands for configuring frequency characteristics of each biopotential signal. The controller automatically generates the frequency control signals based on such commands and outputs such frequency control signals to the configurable receiver circuit. The controller outputs a representation of each biopotential signal to an analyzer device that is configured to analyze such biopotential signal. | 1. A method of monitoring biopotential of biological tissue through a plurality of electrodes, comprising: receiving a plurality of commands into a biopotential sensing device, wherein the commands include commands for setting a high and low cutoff frequency and for selecting one or more channels from a plurality of channels for outputting biopotential measurements, wherein the commands further specify initiation of multiple impedance measurements of a particular electrode at a plurality of frequencies and specify an AC current or voltage waveform at the plurality of frequencies;the biopotential sensing device sensing a biopotential signal from biological tissue coupled with each of a plurality of electrodes coupled to the plurality of channels of the biopotential sensing device;the biopotential sensing device band-pass amplifying each biopotential signal based on the high and low cutoff frequency and outputting a digital representation of the band-pass amplified biopotential signal with substantially zero DC (direct current) offset for each selected channel by at least blocking any DC offset from the biological tissue that is coupled with the selected channel's electrode;the biopotential sensing device generating a capacitively-coupled AC (alternating current) current waveform, having an amplitude and frequency, that is applied to a particular input coupled to a particular channel to measure impedance of a corresponding particular one of the electrodes;determining the impedance by measuring a voltage signal received at the particular channel in response to the AC current waveform applied to the particular input;the biopotential sensing device generating the AC current waveform at the plurality of frequencies based on the specified AC current or voltage waveform at the plurality of frequencies; andthe biopotential sensing device sensing the AC voltage waveform at the plurality of frequencies for the particular electrode and outputting the sensed AC voltage waveform at the plurality of frequencies to an analyzer device configured for measuring the impedance of the particular electrode based on the AC voltage output for the particular electrode, which AC voltage is proportional to the particular electrode's impedance value for such frequency. 2. The method of claim 1, wherein the commands further include a command for setting a gain and wherein the biopotential sensing device amplifies each biopotential signal based on the gain. 3. The method of claim 1, wherein the biopotential sensing device is in the form of an integrated circuit device communicatively coupled to the electrodes and an analyzer device for issuing the commands and receiving the digital representation of the amplified and band-pass amplified biopotential signal with substantially zero DC offset for each selected channel. 4. The method of claim 3, wherein the biopotential sensing device is operable to receive commands from the analyzer device for reconfiguring a high and low cuto... | Intan Technologies, LLC | A61B005/053 A61B005/04 A61B005/0428 A61B005/00
US10709337 | B2 | 2020-07-14 | Photoacoustic measurement apparatus and photoacoustic measurement system | 2017-06-29 | In a photoacoustic measurement apparatus and a photoacoustic measurement system, blood flow information in a desired region can be generated using a photoacoustic image. A light source emits measurement light. A probe detects a photoacoustic wave generated in a subject after measurement light is emitted to the subject in each of the avascularized condition in which the subject is avascularized and the non-avascularized condition in which the subject is not avascularized. Photoacoustic image generation unit generates a photoacoustic image based on the detection signal of the photoacoustic wave. Blood flow information generation unit generates blood flow information based on the signal value of a photoacoustic image in a region of interest set in the photoacoustic image. | 1. A photoacoustic measurement apparatus, comprising: a receiving circuit that receives a detection signal of a photoacoustic wave, which is generated in a subject due to emission of measurement light to the subject, in each of at least an avascularized condition in which the subject is avascularized and a non-avascularized condition in which the subject is not avascularized;a processor configured to generate a photoacoustic image for each of the avascularized condition and the non-avascularized condition based on the detection signal of the photoacoustic wave; andgenerate blood flow information based on a signal value of each of the photoacoustic images in a region of interest set in the photoacoustic images,wherein the processor is configured to calculate a total value or an average value of the signal value in the region of interest for each time point or time frame, the each time point or the time frame corresponding to one of the photoacoustic images, andwherein the processor further generates, as the blood flow information, a score value based on a difference between the total value or the average value of the signal value in the region of interest in the avascularized condition and the total value or the average value of the signal value in the region of interest in the non-avascularized condition. 2. The photoacoustic measurement apparatus according to claim 1, further comprising: a tourniquet capable of changing a cuff pressure in order to bring the subject into the avascularized condition and the non-avascularized condition. 3. The photoacoustic measurement apparatus according to claim 2, the processor further configured to: control a cuff pressure of the tourniquet. 4. The photoacoustic measurement apparatus according to claim 1, wherein the processor binarizes the signal value by setting the signal value to a first value when the signal value is equal to or greater than a first threshold value and equal to or less than a second threshold value larger than the first threshold value and setting the signal value to a second value different from the first value when the signal value is less than the first threshold value or greater than the second threshold value, and generates the blood flow information based on the binarized signal value. 5. The photoacoustic measurement apparatus according to claim 2, wherein the processor binarizes the signal value by setting the signal value to a first value when the signal value is equal to or greater than a first threshold value and equal to or less than a second threshold value larger than the first threshold value and setting the signal value to a second value different from the first value when the signal value is less than the first threshold value or greater than the second threshold value, and generates the blood flow information based on the binarized signal value. 6. The photoacoustic measurement apparatus according to claim 1, wherein the processor further generates a graph showing a ... | FUJIFILM Corporation | A61B005/00 A61B008/13 A61B008/08 A61B008/00
US10719935 | B2 | 2020-07-21 | Image processing apparatus and image processing method thereof | 2016-11-14 | An image processing apparatus is disclosed. The image processing apparatus of the present invention comprises: an image receiving unit for receiving a first image and a second image of the same object taken at different times; a processor for obtaining transformation information by registering the first image on the basis of the second image, obtaining a first segment image corresponding to an area of the object from the first image, and generating a second segment image corresponding to an area of the object of the second image by transforming the obtained first segment image according to the transformation information; and an output unit for outputting the second segment image. | 1. An image processing apparatus, comprising: an image receiving unit for receiving a first image and a second image of the same object taken at different times;a processor configured for: obtaining transformation information by registering the first image on the basis of the second image,obtaining a first segment image corresponding to an area of the object from the first image, andgenerating a second segment image corresponding to an area of the object of the second image by transforming the obtained first segment image according to the transformation information; andan output unit for outputting the second segment image. 2. The image processing apparatus as claimed in claim 1, wherein the processor is further configured for controlling the output unit to output the first segment image obtained from the first image together with the second segment image. 3. The image processing apparatus as claimed in claim 1, wherein the processor is further configured for controlling the output unit to output a comparison result between the first segment image obtained from the first image and the second segment image. 4. The image processing apparatus as claimed in claim 1, wherein the object is a bronchus,wherein the first image is an image taken during inhalation,wherein the second image is an image taken during exhalation, andwherein the processor is further configured for comparing the first segment image and the second segment image and controls the output unit to output information relating to a change of the bronchus during inhalation and exhalation. 5. The image processing apparatus as claimed in claim 4, wherein the information relating to the change of the bronchus is information relating to a change of at least one of a thickness of a bronchus wall, a cross-sectional area of a bronchus wall, a diameter of a bronchus, a cross-sectional area of a bronchus, a ratio of a diameter of a bronchus to a thickness of a bronchus wall, and a ratio of a cross-sectional area of a bronchus to a cross-sectional area of a bronchus wall. 6. The image processing apparatus as claimed in claim 1, wherein the processor, in response to the first image including a plurality of images of the first image taken at different times, is further configured for: obtaining a plurality of first segment image from the plurality of images,generating a plurality of transformation images by transforming the plurality of first segment images according to a plurality of transformation information acquired in a registration process of the plurality of images, andgenerating the second segment images by combining the plurality of transformation images. 7. The image processing apparatus as claimed in claim 1, wherein at least one of the first image and the second image is an image which is acquired through a computed tomography (CT) apparatus, a magnetic resonance imaging (MRI) apparatus, an ultrasound imaging apparatus, a positron emission tomography (PET) apparatus, or an X-ray appa... | Samsung Electronics Co., Ltd. | G06T007/30 G06T007/10 G06T007/00 A61B005/055 A61B005/08 A61B006/03 A61B008/08 G06T007/62 G06T007/11 G06T007/168 G06T007/37 A61B006/00
US10702212 | B2 | 2020-07-07 | Magnetic resonance examination system with motion detection | 2013-06-17 | A magnetic resonance examination system includes an RF arrangement with an RF antenna to acquire magnetic resonance signals from an object to be examined. A motion sensing arrangement detects motion information of the object. The motion sensing arrangement is provided with one or more RF antenna motion sensors mounted on the RF antenna and one or more object motion sensors to be attached to the object. In an example the motion sensors are integrated devices having motion sensitivity along three independent axes. | 1. A magnetic resonance examination system comprising: a flexible RF antenna configured to receive magnetic resonance signals from a patient to be examined, the flexible RF antenna being configured to be movably mounted to the patient such that the flexible RF antenna and the object are movable relative to each other;a first plurality of motion sensors mounted to the flexible RF antenna, and configured to detect motion and deformation of the flexible RF antenna;a second plurality of motion sensors configured to be directly mounted to the patient to detect motion of the object, the second plurality of motion sensors being separate from the first plurality of motion sensors; anda data processor connected with the first plurality of motion sensors and the second plurality of motion sensors and configured to: determine relative motion between the patient and the flexible RF antenna from the signals indicative of flexible RF antenna motion and deformations from the first plurality of motion sensors mounted to the flexible RF antenna and signals indicative of the motion of the patient from the second pluralities of motion sensors mounted to the patient; andreconstruct a motion and RF antenna deformation corrected magnetic resonance image from the received magnetic resonance signals corrected for the relative motion between the flexible RF antenna and the patient. 2. The magnetic resonance examination system as claimed in claim 1, wherein the first and second plurality of motion sensors include three axis micro-electro-magnetic switches (MEMS). 3. The magnetic resonance examination system as claimed in claim 1, further including a transmit and receive switch connected to an RF antenna and configured to supply RF signals to the RF antenna to operate the RF antenna as a transmit coil and to supply the received magnetic resonance signals to the data processor for reconstruction. 4. The magnetic resonance examination system as claimed in claim 1, wherein the flexible RF antenna is a surface RF coil. 5. The magnetic resonance examination system as claimed in claim 1, wherein: the flexible RF antenna is configured to be movably mounted relative to the object such that the flexible RF antenna is deformable and movable relative to the object;the data processor is configured to determine the deformation and the motion of the flexible RF antenna relative to the object from signals from the first and second pluralities of motion sensors, and to reconstruct the received magnetic resonance signals into a magnetic resonance image corrected for the relative information and motion of the flexible RF antenna relative to the object. 6. A magnetic resonance examination system comprising: a deformable local, surface RF coil configured to be mounted to an object to be examined such that the coil and the object are movable relative to each other, the local, surface RF coil being configured to at least receive magnetic resonance signals from the object, the local, surfac... | Koninklijke Philips N.V. | A61B005/00 G01R033/565 A61B005/055 G01R033/30 G01R033/34 G01R033/385 G01R033/54 G01R033/563
US10709834 | B2 | 2020-07-14 | Medication fluid infusion set component with integrated physiological analyte sensor, and corresponding fluid infusion device | 2017-12-14 | Disclosed is a medical device component for delivering medication fluid to a patient. The medical device component includes a fluid infusion device to regulate delivery of medication fluid, a body-mountable base unit, and a top cover assembly that is removably couplable to the base unit and to the fluid infusion device. The base unit includes a cannula to deliver medication fluid under the control of the fluid infusion device, and a physiological analyte sensor to measure a physiological characteristic. The base unit also includes an electronics assembly electrically connected to sensor leads to obtain measurements in the analog domain, to convert measurements into digital sensor data, and to communicate conditioned digital sensor data to the fluid infusion device. The top cover assembly is configured to provide both fluid and electrical connections for the base unit, by way of an infusion tube having sensor conductors integrated therein or otherwise associated therewith. | 1. A medical device component for delivering medication fluid to a patient, the medical device component comprising: a body-mountable base unit comprising: a base structure;a lid assembly affixed and sealed to the base structure;a body-insertable cannula that accommodates delivery of medication fluid to the patient, the cannula having an upstream end securely attached to the base structure and having a downstream end extending from the base structure;a self-sealing septum coupled to the base structure to fluidly seal the upstream end of the cannula, wherein the self-sealing septum is maintained in position between the base structure and the lid assembly;a body-insertable physiological analyte sensor secured to and extending from the base structure, the sensor facilitating measurement of a physiological characteristic of the patient, and the sensor comprising a plurality of sensor leads; andan electronics assembly coupled to the base structure, the electronics assembly electrically connected to the sensor leads to obtain measurements of the physiological characteristic in an analog domain, and the electronics assembly comprising a digital processing circuit to convert measurements of the physiological characteristic from the analog domain into digital sensor data, to digitally process the digital sensor data into conditioned digital sensor data, and to communicate the conditioned digital sensor data to a fluid infusion device associated with the medical device component; anda top cover assembly that is removably couplable to the body-mountable base unit such that the top cover assembly covers the lid assembly of the body-mountable base unit, the top cover assembly comprising: a lid structure that releasably mates with the base structure, the lid structure comprising an interior space defined by an inner surface of the lid structure, and the lid structure comprising at least one structural support feature;an infusion tube secured to the inner surface of the lid structure and terminating within the interior space, wherein the at least one structural support feature of the lid structure receives, secures, and retains a downstream end of the infusion tube;a tubing connector fluidly coupled to the infusion tube, the tubing connector having a distal end supported by the at least one structural support feature of the lid structure, the distal end of the tubing connector penetrating the self-sealing septum to establish a fluid delivery flow path from inside the infusion tube, through the self-sealing septum, and into the cannula when the top cover assembly is coupled to the body-mountable base unit, wherein the self-sealing septum seals the upstream end of the cannula when the top cover assembly is removed from the body-mountable base unit and the tubing connector is withdrawn from the self-sealing septum;a plurality of sensor conductors carried by or integrated with the infusion tube, the sensor conductors terminating within the interior space; andan ele... | Medtronic MiniMed, Inc. | A61M005/14 A61M005/142 A61M005/172 A61B005/00 A61B005/145 A61B005/1473 A61M005/20 A61M005/00 A61M005/158 A61M005/145
US10702192 | B2 | 2020-07-07 | Dynamic analysis system | 2017-06-12 | A dynamic analysis system includes: an analytic value calculation unit configured to calculate an analytic value in a plurality of time phases on the basis of a dynamic image in the plurality of time phases obtained by performing dynamic photography on the chest of a subject; a ventilation state calculation unit configured to calculate, using a non-linear function, an index value representing a ventilation state of a lung field from the analytic value in the plurality of time phases; a display unit; and a control unit configured to display the index value calculated on the plurality of time phases, on the display unit. | 1. A radiographic dynamic analysis system comprising: an analytic value calculation unit configured to calculate an analytic value in a plurality of time phases on the basis of a radiographic dynamic image in the plurality of time phases obtained by performing radiographic dynamic photography on the chest of a subject, wherein the analytic value is one of transmitted X-ray intensity and a value calculated using the transmitted X-ray intensity;a ventilation state calculation unit configured to calculate, using a non-linear function that is predetermined, an index value representing a ventilation state of a lung field from the analytic value in the plurality of time phases, wherein the ventilation state represented by the index value is an air amount within the lung field or a change amount of the air within the lung field, wherein the non-linear function used in an inspiratory phase is a function having a downward curve at least in a portion of the region and the non-linear function used in an expiratory phase is a function having an upward curve at least in a portion of the region;a display unit; anda control unit configured to display information indicating the index value calculated on the plurality of time phases, on the display unit. 2. The radiographic dynamic analysis system according to claim 1, wherein the analytic value is one of transmitted X-ray intensity and a value calculated using the transmitted X-ray intensity. 3. The radiographic dynamic analysis system according to claim 1, wherein the index value represents the air amount within the lung field. 4. The radiographic dynamic analysis system according to claim 1, wherein the ventilation state calculation unit calculates the index value from the analytic value using different functions for each of an expiratory phase and an inspiratory phase. 5. The radiographic dynamic analysis system according to claim 4, wherein the ventilation state calculation unit uses a function in which an increase/decrease of the index value is reversed with respect to the increase/decrease of the analytic value, between the expiratory phase and the inspiratory phase. 6. The radiographic dynamic analysis system according to claim 1, the non-linear function used in the expiratory phase includes a region having a downward curve and a region having the upward curve. 7. The radiographic dynamic analysis system according to claim 1, further comprising: a storage unit configured to store the non-linear function; andan information acquisition unit configured to obtain the non-linear function stored in the storage unit,wherein the ventilation state calculation unit calculates the index value representing the ventilation state from the analytic value using the non-linear function obtained by the information acquisition unit. 8. The radiographic dynamic analysis system according to claim 1, wherein the analytic value calculation unit divides a lung field region included in the dynamic image in the plurality of... | KONICA MINOLTA, INC. | A61B005/11 A61B006/00 A61B005/113
US10716538 | B2 | 2020-07-21 | Hemodynamic ultrasound medical monitoring device | 2016-06-30 | An ultrasonic medical monitoring device may include: at least one ultrasonic probe attached to a body surface of a subject to scan the body surface and acquire an echo signal; a blood pressure measuring module to measure a blood pressure parameter of the subject; a processing module to receive the echo signal and process the echo signal into a blood flow parameter, the processing module being further configured to calculate a myocardial mechanic parameter according to the blood flow parameter and the blood pressure parameter; and a display module coupled to the processing module to display the blood pressure parameter, the blood flow parameter, and the myocardial mechanic parameter. The present device and method acquire the blood pressure parameter and the blood flow parameter using ultrasound and a blood pressure synchronous detection technique, and the myocardial mechanics parameters such as the systolic maximum elastance are calculated by a computer. | 1. An ultrasonic medical monitoring device, comprising: at least one ultrasonic probe that scans a body surface of a test subject and acquires an echo signal;a blood pressure measuring device that measures a blood pressure parameter of peripheral arteries;a general physiological monitoring device coupled to the monitoring host via a pluggable connector;a processor configured to receive the echo signal and processes the echo signal into a blood flow parameter, wherein the processor is integrated into the monitoring host, and the at least one ultrasonic probe and the blood pressure measuring device are coupled to the monitoring host via pluggable connectors; wherein the processor further calculates a myocardial mechanic parameter that specifically represents a synchronous coupling of a circulatory system of left ventricle-aorta-peripheral arteries circulation according to the blood flow parameter and the blood pressure parameter; anda display device that is integrated into the monitoring host and coupled to the processor and displays the blood pressure parameter, the blood flow parameter, and the myocardial mechanic parameter,wherein the myocardial mechanic parameter is selected from a group consisting of: slope of an ascending branch of an ABP waveform regressed by M-mode curve rising speed and slope of ascending branch of the ABP waveform regressed by aortic forward blood flow velocity. 2. The ultrasonic medical monitoring device according to claim 1, wherein the blood pressure parameter comprises peripheral arterial pulse wave information and blood pressure information. 3. The ultrasonic medical monitoring device according to claim 2, wherein the blood pressure information comprises a systolic pressure, a diastolic pressure, and a mean arterial pressure. 4. The ultrasonic medical monitoring device according to claim 1, further comprising a plurality of ultrasonic probes attached to different portions of the body surface of the tested subject. 5. The ultrasonic medical monitoring device according to claim 4, wherein the plurality of ultrasonic probes are attached to ventricle, aorta and peripheral arteries of the tested subject to perform synchronized real-time scanning. 6. The ultrasonic medical monitoring device according to claim 1, wherein the processor further digitally processes the echo signal to obtain a digital processing signal, and to obtain numerical parameter, waveform or trend graph calculation according to the digital processing signal, wherein the digital processing signal is selected from the group consisting of: RF signal, baseband signal, and envelope signal. 7. The ultrasonic medical monitoring device according to claim 1, wherein the processor further digitally processes the echo signal to obtain a digital processing signal, and to obtain an ultrasound image according to the digital processing signal and a chosen imaging mode, and to obtain the blood flow parameter according to the ultrasound image calculation, wherein ... | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD. | A61B008/00 A61B008/06 A61B008/04 A61B005/021 A61B005/0215 A61B005/022 A61B008/08 A61B008/14
US10709370 | B2 | 2020-07-14 | Biological fluid separation device | 2017-06-09 | A blood separation device adapted to receive a blood sample having a whole blood portion and a plasma portion is disclosed. The blood separation device includes a housing defining a first chamber, a second chamber, and a separation member disposed therebetween. The blood separation device also includes an actuator, wherein actuation of the actuator draws the blood sample into the first chamber and the separation member is adapted to allow the plasma portion to pass through the separation member to the second chamber. The blood separation device matches the plasma chamber volume, the blood chamber volume, and the applied single pressure source so that the correct trans-membrane pressure and shear rate is obtained. | 1. A blood separation device comprising: a housing defining a first chamber having a first chamber inlet and a first chamber outlet wherein the first chamber outlet is placed a lateral distance from the first chamber inlet, wherein the first chamber has a height of about 65 to about 95 μm,a second chamber having a second chamber outlet, and including a separation member disposed between the first chamber and the second chamber thereby dividing the housing into the first and second chambers and defining a tangential flow path from the first chamber inlet to the first chamber outlet, wherein the first chamber is adapted to receive a blood sample through the first chamber inlet; and a single vacuum source actuator for a single vacuum source,wherein the single vacuum source is applied adjacent the separation member by a first line to the first chamber outlet and draws the blood sample into the first chamber and along the tangential flow path defined by the separation member and wherein the single vacuum source is the only vacuum source in the blood separation device, wherein the separation member is adapted to allow a plasma portion of the blood sample to pass through the separation member into the second chamber. 2. The blood separation device of claim 1, wherein the separation member is adapted to trap a whole blood portion of the blood sample in the first chamber and allow the plasma portion to pass through the separation member and into the second chamber. 3. The blood separation device of claim 1, further comprising a plasma collection container in communication with the second chamber outlet. 4. The blood separation device of claim 1, wherein the single vacuum source actuator provides a first pressure to a portion of the first chamber via the first line. 5. The blood separation device of claim 4, wherein a second pressure in the second chamber is regulated by a porosity of the separation member and a volume of the second chamber. 6. The blood separation device of claim 1, wherein the single vacuum source actuator is an evacuated tube. 7. The blood separation device of claim 1, wherein the separation member comprises a track-etched membrane. 8. The blood separation device of claim 1, further comprising a vent transitionable between an open position and a closed position. 9. The blood separation device of claim 1 wherein the separation member has a pore size of 0.2 μm to 1 μm. 10. The blood separation device of claim 1 wherein the pressure difference between the outlet of the first chamber and the outlet of the second chamber is in the range of 8 psi to −8 psi. 11. The blood separation device of claim 10 wherein the separation member has a pore size of 0.2 μm to 1 μm. 12. The blood separation device of claim 1 wherein the first chamber height is 80 μm, the first chamber width is 10 mm and the first chamber length is 50 mm. 13. A blood separation device comprising: a housing defining a first chamber having a first chamber inlet and a f... | Becton, Dickinson and Company | A61B005/15 B01D061/18 G01N033/49 B01L003/00
US10716469 | B2 | 2020-07-21 | Ocular-performance-based head impact measurement applied to rotationally-centered impact mitigation systems and methods | 2019-01-31 | A system or method for measuring human ocular performance can be implemented using an eye sensor, a head orientation sensor, and an electronic circuit. The device is configured for measuring vestibulo-ocular reflex, pupillometry, saccades, visual pursuit tracking, vergence, eyelid closure, dynamic visual acuity, retinal image stability, foveal fixation stability, focused position of the eyes or visual fixation of the eyes at any given moment and nystagmus. The eye sensor comprises a video camera that senses vertical movement and horizontal movement of at least one eye. The head orientation sensor senses pitch and yaw in the range of frequencies between 0.01 Hertz and 15 Hertz. The system is implemented as part of an impact reduction helmet that comprises an inner frame having interior pads configured to rest against a person's head and one or more shock absorption elements attached between the inner frame and the spherical shell that couple the spherical shell to the inner frame. The spherical shell has a circular geometry, that when viewed horizontally at its horizontal midplane, includes a center point that is the rotational center of the spherical shell. The one or more shock absorption elements are sized to provide greater spacing between the inner frame and the spherical shell at the sides and rear of the spherical shell than at the front of the spherical shell. The one or more shock absorption elements are sized to configure the alignment of the rotational center of the spherical shell with the proximate rotational center of the wearer's head. | 1. A human ocular performance measuring device wherein: the device is configured for measuring an ocular performance characteristic selected from the group of: vestibulo-ocular reflex;ocular saccades;pupillometry;pursuit tracking during visual pursuit;vergence;eye closure;focused position of the eyes;dynamic visual acuity;kinetic visual acuity;virtual retinal stability;retinal image stability;foveal fixation stability; andnystagmus; andthe device comprises: an eye sensor wherein: the eye sensor comprises a video camera; andthe eye sensor senses eye information selected from the group of: horizontal eye movement;vertical eye movement;pupil size; andeyelid movement;a head orientation sensor wherein: the head orientation sensor senses a head movement selected from the group of pitch and yaw of a person's head wherein pitch represents a rotation about a first axis representing up and down movement of the person's face when the rear of the person's head moves in the opposite direction and yaw represents horizontal movement of the face when looked at from the front about a second axis substantially aligned with the spine and perpendicular to the first axis; andthe head orientation sensor senses the head movement in a range of frequencies between 0.01 Hertz and 15 Hertz;the head orientation sensor comprises a micro-electro-mechanical system integrated circuit comprising a module selected from the group consisting of an accelerometer, a magnetometer, and a gyroscope;an electronic circuit wherein: the electronic circuit comprises a central processing unit, and a memory unit;the electronic circuit is responsive to the eye movement information received from the eye sensor; andthe electronic circuit is responsive to head movement information received from the head orientation sensor; anda helmet, comprising an inner frame having an interior surface and one or more shock absorption elements attached between the inner frame and a spherical shell, wherein: the one or more shock absorption elements couple the spherical shell to the inner frame;the spherical shell is external to the inner frame;the spherical shell has a circular geometry that when viewed horizontally at its horizontal midplane, includes a center point that is the rotational center of the spherical shell;the one or more shock absorption elements are sized to provide greater spacing between the inner frame and the spherical shell at the sides and the rear of the spherical shell than at the front of the spherical shell;the one or more shock absorption elements are sized to configure the alignment of the rotational center of the spherical shell with a proximate rotational center of a wearer's head; andthe alignment between the rotational center of the spherical shell and the proximate rotational center of a wearer's head directly affects the effect of a tangential force caused by an impact on the spherical shell. 2. The device of claim 1 wherein: the device measures vestibulo-ocular reflex. 3. The ... | - | A61B003/14 A61B003/113 A61B003/11 A61B005/00 A61B005/11 A61B003/00 A61B003/032 A61B003/15
US10726551 | B2 | 2020-07-28 | Method for reconstructing joined part surfaces of two adjacent teeth of a 3D digital model of teeth | 2018-01-24 | A method for reconstructing joined part surfaces of two adjacent teeth of a 3D digital model of teeth is provided. The method includes: obtaining a 3D digital model of teeth which includes adjacent first tooth and second tooth having joined surfaces; identifying a set of intersection points of the joined surfaces; fitting a first plane using the set of intersection points; identifying a set of first control points from points on the 3D digital model of teeth near the set of intersection points basis curvature, where the set of first control points defines boundary of a first part surface of the first tooth to be reconstructed; reconstructing the first part surface using a linear equation system constructed basis the first plane and the set of first control points. | 1. A method for reconstructing joined part surfaces of two adjacent teeth of a 3D digital model of teeth, the method comprising: obtaining a 3D digital model of teeth which includes adjacent first tooth and second tooth having joined surfaces;identifying a set of intersection points of the joined surfaces;fitting a first plane using the set of intersection points;identifying a set of first control points from points near the set of intersection points based on curvature, where the set of first control points defines a boundary of a first part surface of the first tooth to be reconstructed;sampling on the first plane to obtain a first set of nodes;sampling on the first plane a first set of points, and moving the first set of points along a normal of the first plane toward the first tooth, to obtain a set of second control points;constructing a linear equation system based on the set of first control points, the set of second control points and the first set of nodes;solving the linear equation system to obtain a first set of Z coordinate values for the first set of nodes when a 3D coordinate system, whose Z axis is parallel with a normal of the first plane, is used; andreconstructing the first part surface based on the first set of nodes with the first set of Z coordinate values. 2. The method of claim 1, wherein the reconstructed first part surface is set to protrude along a normal of the first plane toward the second tooth. 3. The method of claim 1, wherein the set of first control points is identified based on one of principal curvature, Gaussian curvature, subordinate curvature, and mean curvature. 4. The method of claim 1, wherein the set of first control points is identified according to one of a predetermined curvature threshold and a predetermined curvature changing rate threshold. 5. The method of claim 4, wherein the predetermined curvature threshold and the predetermined curvature changing rate threshold are determined based on precision of the 3D digital model. 6. The method of claim 4, wherein the predetermined curvature threshold and the predetermined curvature changing rate threshold are determined based on type of the first tooth. 7. The method of claim 1, wherein the first set of point is sampled within an area defined by projection of the set of intersection points on the first plane. 8. The method of claim 1, wherein the first set of nodes is sampled within a maximum area defined by the greatest and the smallest X and Y coordinate values of projection of the set of intersection points on the first plane. 9. The method of claim 1, wherein the first set of points is moved by a first predetermined step to obtain the set of second control points. 10. The method of claim 9, wherein the first predetermined step is calculated by dividing the minimum distance between the set of first control points and the first plane by a first predetermined number. 11. The method of claim 1 further comprising: if the reconstructed first part... | Wuxi EA Medical Instruments Technologies Limited | G06T007/00 G06T017/20 G06T007/73 G06T007/50 G16H040/63 G16H050/50 A61B005/00 A61B005/107
US10709419 | B2 | 2020-07-14 | Dual modality imaging system for coregistered functional and anatomical mapping | 2016-09-27 | A real-time imaging system that provides ultrasonic imaging and optoacoustic imaging coregistered through application of the same hand-held probe to generate and detect ultrasonic and optoacoustic signals. These signals are digitized, processed and used to reconstruct anatomical maps superimposed with maps of two functional parameters of blood hemoglobin index and blood oxygenation index. The blood hemoglobin index represents blood hemoglobin concentration changes in the areas of diagnostic interest relative to the background blood concentration. The blood oxygenation index represents blood oxygenation changes in the areas of diagnostic interest relative to the background level of blood oxygenation. These coregistered maps can be used to noninvasively differentiate malignant tumors from benign lumps and cysts. | 1. An imaging system for visualization of slices into a depth of tissue of at least a portion of a body, comprising: a hand-held imaging probe comprising an array of ultrasonic transducers and a light emitting portion, wherein the light emitting portion comprises: a first fiber-optic bundle and a first light diffuser configured to form a first light beam and a second fiber-optic bundle and a second light diffuser configured to form a second light beam, wherein the first and second fiber-optic bundles are disposed on opposite sides of the array of ultrasonic transducers; anda processing system configured to receive data originating from the hand-held imaging probe and to process at least three independent images based at least in part upon said data, wherein the processing comprises deconvolution of an impulse response associated with the array of ultrasonic transducers, the impulse response based on a delta-function ultrasonic signal generated by laser pulses in an absorbing optical medium the three independent images together comprising: a first functional image reflecting distribution of total hemoglobin concentration;a second functional image reflecting distribution of blood oxygen saturation; anda morphological image of tissue structures. 2. The imaging system of claim 1, further comprising a signal filter that filters the delta-function ultrasonic signal after deconvolution of the impulse response associated with the array of ultrasonic transducers and communicates with the processing system. 3. The imaging system of claim 2, wherein the signal filter is a wavelet filter based on a wavelet transform that operates simultaneously in the frequency and time domain. 4. The imaging system of claim 3, wherein the wavelet filter is configured to filter frequency components of the delta-function ultrasonic signal that are noise at a first time; and preserve frequency components of the delta-function ultrasonic signal at a second time. 5. The imaging system of claim 3, wherein a frequency spectrum of the wavelet filter replicates a frequency band of an N-shaped optoacoustic signal. 6. The imaging system of claim 5, wherein the wavelet filter provides smooth window edges that do not cause signal distortions upon convolution. 7. The imaging system of claim 5, wherein the wavelet filter converts bipolar pressure signals into monopolar signals. 8. The imaging system of claim 1, wherein the deconvolution of the impulse response restores N-shaped pressure signals. 9. The imagining system of claim 1, wherein the hand-held imaging probe has a concave arc shape. | SENO MEDICAL INSTRUMENTS, INC. | A61B008/00 A61B005/00 A61B008/08 A61B005/145 A61B005/1455 A61B008/14 A61B090/30
US10716513 | B2 | 2020-07-21 | Systems and methods for cardiovascular blood flow and musculoskeletal modeling for predicting device failure or clinical events | 2015-12-09 | Computer-implemented methods are disclosed for assessing the effect of musculoskeletal activities on disease and/or clinical events, the method comprising: receiving a patient-specific vascular and musculoskeletal model of a patient's anatomy, including at least one vessel of the patient; receiving at least one characteristic of the patient's musculoskeletal activity; generating or updating a computational anatomic vascular and musculoskeletal model of the patient's anatomy based on the received at least one characteristic of musculoskeletal activity; performing at least one of a computational fluid dynamics analysis or a structural mechanics simulation on the computational anatomic vascular and musculoskeletal model; and estimating at least one of the patient's risk of disease or clinical events based on the performed computational fluid dynamics analysis and/or structural mechanics simulation. Systems and computer readable media for executing these methods are also disclosed. | 1. A computer-implemented method of non-invasively assessing a risk of a clinical event in a patient, the method comprising: obtaining, via an imaging device, one or more images of a patient's anatomy;generating a first three-dimensional patient-specific vascular model of the patient's anatomy based on the one or more images, including at least one vessel of the patient;measuring a duration and/or frequency of musculoskeletal behavior of the patient while the patient is at a given posture via a device associated with the patient;generating an updated three-dimensional patient-specific vascular model of the patient's anatomy by simulating or detecting the given posture corresponding to the measured duration and/or frequency of musculoskeletal behavior of the patient;computing mechanical and/or hemodynamic characteristics of the patient's vascular system at the given posture by performing at least one of a computational fluid dynamics analysis or a structural mechanics simulation using the updated three-dimensional patient-specific vascular model;using the at least one of the computational fluid dynamics analysis or structural mechanics simulation by creating a feature vector of the musculoskeletal behavior of the patient while the patient is at the given posture to estimate and output an effect of the computed mechanical and/or hemodynamic characteristics on at least one of the patient's risk of disease, a risk of failure of a device, or a performance characteristic of a device;assessing a risk of a clinical event in the patient at the given posture, using the estimated effect of the mechanical and/or hemodynamic characteristics of the vascular system of the patient on the at least one of the patient's risk of disease, the risk of failure of a device, or the performance characteristic of a device; andoutputting the risk of the clinical event to an electronic storage medium or a display for use by one or more healthcare providers. 2. The computer-implemented method of claim 1, wherein generating the updated three-dimensional patient-specific vascular model of the patient's anatomy comprises generating a computational anatomic vascular model of the patient's anatomy. 3. The computer-implemented method of claim 2, wherein the updated three-dimensional patient-specific vascular model of the patient's anatomy is generated using one or more of: a computational structural mechanics simulation; orimages of the patient's venous system in various postures taken by a magnetic resonance imaging (MRI) system. 4. The computer-implemented method of claim 2, wherein the musculoskeletal behavior of the patient includes one or more of sitting or lying down. 5. The computer-implemented method of claim 1, wherein estimating the effect of the mechanical and/or hemodynamic characteristics of the vascular system of the patient on at least one of the patient's risk of disease, the risk of failure of a device, or the performance characteristic of a device includes exe... | HeartFlow, Inc. | A61B005/00 A61B005/11 A61B005/026 G16H050/30 G16H050/50 A61B005/055 A61B005/021 A61B034/00 A61B034/10 A61B005/02
US10709312 | B2 | 2020-07-14 | Transitional region having cuts and a skive for an imaging catheter | 2019-04-05 | An imaging device for imaging a portion of a patient's vasculature with an imaging element may include a proximal portion having a relatively higher stiffness that provides rigidity for pushing the imaging device through a patient's vasculature, and may include a distal portion having a relatively lower stiffness that enables threading through a curved vasculature of the patient. The imaging device also may include a transition region disposed between the proximal portion and the distal portion. | 1. An intraluminal device, comprising: a flexible elongate member configured to be positioned within anatomy of a patient, the flexible elongate member comprising a proximal portion and a distal portion, the proximal portion comprising a relatively higher stiffness and the distal portion comprising a relatively lower stiffness;a transition region disposed between the proximal portion and the distal portion, the transition region comprising a graduated stiffness transitioning from the stiffness of the proximal portion to the stiffness of the distal portion, the transition region comprising: a hypotube comprising: a lumen; anda skive at a distal portion of the hypotube, the skive comprising a tapered portion that is relatively more stiff at a proximal region and less stiff at a distal region; anda cable extending from the proximal portion of the flexible elongate member over the skive to a functional element, wherein the cable is disposed within the lumen of the hypotube; anda polymer jacket extending over at least a portion of the hypotube and the skive, wherein the polymer jacket comprises an inner surface surrounding a lumen, and wherein the tapered portion of the skive and a portion of the cable are positioned within the lumen of the polymer jacket such that the tapered portion of the skive and the cable are fixed in place relative to one another by the inner surface of the polymer jacket and an inner surface of the tapered portion of the skive. 2. The intraluminal device of claim 1, wherein the polymer jacket comprises a circular cross section at a proximal region of the polymer jacket, and a non-circular cross section at a skive region distal of the proximal region, wherein the inner surface of the polymer jacket is disposed at the skive region of the polymer jacket. 3. The intraluminal device of claim 1, further comprising: an outer distal shaft extending over the distal region of the skive and a distal portion of the polymer jacket; andan inner member embedded within the outer distal shaft and configured to receive a guidewire, wherein the cable is disposed between the skive and the inner member. 4. The intraluminal device of claim 3, wherein the inner member has an opening end in the transition region. 5. The intraluminal device of claim 3, wherein the cable lies within the skive. 6. The intraluminal device of claim 1, comprising an outer distal shaft portion forming a part of the transition region, the outer distal shaft portion extending over the distal region of the skive to provide a graduated flexibility transition from the skive to the distal portion. 7. The intraluminal device of claim 1, wherein the skive has a length within a range of about 1-4 inches. 8. The intraluminal device of claim 1, wherein the hypotube includes a cylindrical portion proximal of the skive and a connection region proximal of the tapered portion, wherein the tapered portion comprises a curved edge to meet the cylindrical portion. 9. The intraluminal ... | PHILIPS IMAGE GUIDED THERAPY CORPORATION | A61B001/04 A61B001/00 A61B001/313 A61B008/12 A61B005/00 A61B008/00 A61M025/00
US10702705 | B2 | 2020-07-07 | Methods and systems for neural stimulation via visual, auditory and peripheral nerve stimulations | 2019-05-30 | Systems and methods of the present disclosure are directed to systems and methods for treating cognitive dysfunction in a subject in need thereof. The system can include a light source and a speaker. A visual neural stimulation system provides, via the light source, visual stimulation having a first value of a first parameter. An auditory neural stimulation system provides, via the speaker, audio stimulation having a second value of the second parameter. A stimuli orchestration component selects, for a first time interval, one of the visual stimulation or the audio stimulation to vary based on a policy, selects, for the first time interval, the other of the visual stimulation or the audio stimulation to keep constant based on the policy, and provides causes the one of the visual neural stimulation system or the auditory neural stimulation system to vary the one of the visual stimulation or the audio stimulation. | 1. A system, comprising: one or more processors configured to: select a light pattern having a first parameter and a second parameter;set a value of the first parameter of the light pattern based on a fixed predetermined stimulation frequency of neural oscillations, and set a first value of the second parameter of the light pattern based on a first default value;select an audio signal having a first parameter and a second parameter;set a value of the first parameter of the audio signal based on the fixed predetermined stimulation frequency of neural oscillations, and set a first value of the second parameter of the audio signal based on a second default value;construct a visual output signal based on the first parameter of the light pattern and the second parameter of the light pattern;construct an audio output signal based on the first parameter of the audio signal and the second parameter of the audio signal;provide the visual output signal to a light source to cause the light source to generate and direct light towards at least one eye of a subject in accordance with the constructed visual output signal;provide the audio output signal to a speaker to cause the speaker to provide sound to the subject in accordance with the constructed audio output signal;update, for a first time interval and responsive to a time-based or feedback-based function having one or more conditions, while maintaining the first parameter of the light pattern or the audio signal set based on the fixed predetermined stimulation frequency of neural oscillations, at least one parameter of one of the visual output signal or the audio output signal to generate an updated output signal; andprovide the updated output signal to cause one of the light source or the speaker to present the updated output signal. 2. The system of claim 1, wherein the one or more processors are configured to: select, for the first time interval, the first value of the second parameter of the light pattern to update; andupdate, for the first time interval, the first value of the second parameter of the light pattern to a second value to construct the updated output signal, without updating the first parameter and the second parameter of the audio output signal. 3. The system of claim 1, wherein the one or more processors are configured to: select one of the visual output signal or the audio output signal to sweep; andvary, for a plurality of time intervals, the at least one parameter of the selected one of the visual output signal or the audio output signal conduct a sweep through values of the at least one parameter, without varying the first parameter or the second parameter of the visual output signal or the audio output signal not selected for the sweep. 4. The system of claim 1, wherein the one or more processors are configured to: determine, during the first time interval, a level of attention of the subject based on input from the subject;vary, during a second time interval, the at least one... | Cognito Therapeutics, Inc. | A61N005/06 A61B005/00 A61N001/36 A61N001/04 H05B047/105 A61M021/00 G02C011/00 A61B005/0484 A61B005/16 A61B003/113 A61B005/024 A61B005/0482 A61B005/0496 A61B005/05 A61N001/05
US10722119 | B2 | 2020-07-28 | Patient monitor screen aggregation | 2019-05-06 | A remote interface system according to some embodiments includes one or more patient monitoring devices having one or more sensors for patient monitoring, one or more first processors to receive information from the sensors and generate patient data based thereon, a web server, a first communication system, and a first screen to display a representation of at least a first portion of the patient data, and a remote interface device including a second communication system, a second screen, one or more second processors, an application executed by the one or more second processors and configured to detect a presence of the first communication system, establish a communication link between the first and second communication systems, establish a reliable connection channel (e.g. secure websocket connection) with the web server, receive the patient data via the connection, and display at least a second portion of the patient data on the second screen. | 1. A remote interface system for medical applications, the remote interface system comprising: an external defibrillator monitor comprising: one or more sensors configured for physical attachment to a patient and to acquire electrocardiogram (ECG) information from the patient, a web server configured to support a secure communications connection and further configured to exchange information via the secure communications connection, and at least one first processor configured to: receive the acquired ECG information from the one or more sensors, transmit the ECG information for information exchange via the secure communications connection; and a remote interface device comprising: a web browser configured to receive information via the secure communications connection, and at least one second processor configured to: establish a communication link between the external defibrillator monitor and the remote interface device, establish the secure communications connection between the web server of the external defibrillator monitor and the web browser, send a request to the web server for the ECG information, receive, in response to the request via the secure communications connection, a first transmission representing at least a first portion of the ECG information, and receive, in response to the request via the secure communications connection, one or more second transmissions representing at least a second portion of the ECG information, the one or more second transmissions received by the at least one second processor continuously after the first transmission is received and without requiring a further request to the web server being sent by the at least one second processor. 2. The system of claim 1, wherein a latency of at least one of the first transmission and the one or more second transmissions of the ECG information from the external defibrillator monitor to the remote interface device is approximately 40 to 200 milliseconds. 3. The system of claim 1, wherein at least one of the first transmission and the one or more second transmissions represents between 40 to 200 milliseconds of the ECG information. 4. The system of claim 1, wherein the first transmission and the one or more second transmissions are sent between 40 to 200 milliseconds apart. 5. The system of claim 1, wherein the secure communications connection is configured to provide full-duplex communication between the external defibrillator monitor and the remote interface device. 6. The system of claim 1, wherein the first transmission comprises a first packet comprising a header and the first portion of the ECG information and the one or more second transmissions comprise one or more second packets comprising a header and the second portion of the ECG information. 7. The system of claim 6, wherein each of the first packet and the one or more second packets represent approximately 120 milliseconds of the ECG information. 8. The system of claim 6, wherein the at least one f... | Zoll Medical Corporation | A61B005/00 G06F019/00 A61N001/39 A61B034/10 G16H040/67
US10722678 | B2 | 2020-07-28 | Device and method for effective non-invasive two-stage neurostimulation | 2016-06-22 | A device for stimulating neurons that includes a non-invasive stimulation unit for generating stimuli in stimulation channels. A control unit controls the stimulation unit during first and second time intervals in different stimulation modes. The control unit controls the stimulation unit during 75% or more of the first time interval in the first stimulation mode to repeatedly generate sequences of stimuli and the order in which the stimulation channels generate stimuli is constant for not more than 5 successively generated sequences and then varied. The control unit controls the stimulation unit during 75% or more of the second time interval in a second stimulation mode such that the stimulation channels repeatedly generate sequences of stimuli and the order in which the stimulation channels generate stimuli is constant for at least 25 successively generated sequences and then varied. | 1. A device for stimulating neurons, comprising: a non-invasive stimulation unit configured to generate stimuli in a plurality of stimulation channels, such that the stimuli stimulate a neuron population in at least one of a brain and a spinal cord of a patient via the stimulation channels at respectively different points on the patient; anda control unit configured to operate the stimulation unit during a first time interval and a second time interval following the first time interval in different stimulation modes,wherein the control unit controls the stimulation unit during at least 75% of a duration of the first time interval in a first stimulation mode such that the stimulation unit repetitively generates sequences of stimuli and an order of the stimulation channels in which the stimuli are generated within a sequence is constant for not more than 5 successively generated sequences and then varied,wherein the control unit controls the stimulation unit during at least 75% of a duration of the second time interval in a second stimulation mode such that the stimulation unit repetitively generates sequences of stimuli and an order of the stimulation channels in which the stimuli are generated within a sequence is constant for at least 25 successively generated sequences and then varied, andwherein an intensity of the stimuli in the first stimulation mode is lower than or equal to a predetermined stimulus intensity and the intensity of the stimuli in the second stimulation mode is at least 1.3 times the predetermined stimulus intensity. 2. The device according to claim 1, wherein the sequences in at least one of the first and second stimulation mode are generated in a time pattern of consecutive cycles and at least in some of the cycles a sequence of stimuli is generated respectively. 3. The device according to claim 2, wherein, within a respective cycle, either exactly one sequence of stimuli is generated or no stimuli are generated. 4. The device according to claim 2, wherein the stimuli are generated during n consecutive cycles and no stimuli are generated during subsequent m cycles such that a pattern of the n consecutive cycles followed by the m cycles is periodically continued, wherein n and m are non-negative integers. 5. The device according to claim 1, wherein a pattern according to which an order in which the stimulation unit generates the stimuli in the plurality of stimulation channels within a sequence in the first stimulation mode is constant for at most 5 consecutively generated sequences and then varied is repeated multiple times. 6. The device according to claim 1, wherein a pattern according to which an order in which the stimulation unit generate the stimuli in the plurality of stimulation channels within a sequence in the second stimulation mode is constant for at least 25 consecutively generated sequences and then varied is repeated multiple times. 7. The device according to claim 1, wherein the stimuli are configured t... | Forschungszentrum Juelich GmbH | A61N005/10 G06N099/00 G06N007/00 G06N005/04 A61B006/03 A61M021/02 A61N002/00 A61B005/0476 A61B005/0488 A61N001/36 A61B005/00 A61B005/11 A61B005/053 A61M021/00 A61N002/02
US10724069 | B2 | 2020-07-28 | Methods and devices for cell detection | 2015-09-29 | The present disclosure relates to a cartridge, detection module, system, and kit for cell and particle detection and analysis. Devices disclosed herein may include at least an optical source, a fluidic chip, and a detection module, wherein the sample flows within the fluidic chip past a detection window, where the cells or particles are imaged by an image acquisition and analysis module that may include an optical detector. The image acquisition and analysis module may count the cells or particles of interest in real-time, or near real-time, or the module may capture images of the cells in order to analyze the sample from combined images at a later time. | 1. A cartridge for cell or particle detection and analysis comprising: a. one or more housing units;b. a fluidic chip incorporating at least one microfluidic channel comprising a plurality of meandering mixing microchannels that one or more cells or particles of a sample flows through within the fluidic chip, wherein the fluidic chip comprises at least a bellow actuator that drives either capillary or mechanical flow of the one or more cells or particles of the sample and together with the plurality of meandering mixing microchannels aid mixing of the sample with one or more reagents required for detection and analysis, wherein the one or more reagents are located downstream of at least one of the meandering mixing microchannels, and wherein the meandering mixing microchannels are non-symmetrical or non-uniform within the cartridge and generate turbulence within the microfluidic channel to enhance mixing of the reagents and sample;c. a detection window incorporated in one of the one or more housings, the detection window being operable to facilitate capture of one or more images of the one or more cells or particles of the sample flowing within the detection window. 2. The cartridge of claim 1, the fluidic chip further comprising: a. an inlet through which the sample is introduced to the microfluidic channel;b. an outlet through which the sample is removed from the fluidic chip; andc. a waste reservoir positioned near the outlet, the waste reservoir being operable to collect the sample after the sample has flowed through the microfluidic chip. 3. The cartridge of claim 1, wherein the detection window and a detection module are operable to apply multi-fluorescence detection. 4. The cartridge of claim 1, wherein sample preparation is integrated into the cartridge or wherein all reagents required for a test are supplied and sealed in the cartridge, and combinations thereof. 5. The cartridge of claim 4, wherein the reagents are dried or lyophilized or slow dried on the chip. 6. The cartridge of claim 4, wherein the reagents on the cartridge have a shelf life of at least 12 months at 0° C. to 40° C. 7. The cartridge of claim 4, wherein the reagents on the cartridge have a shelf life of at least 48 hours with temperature fluctuations between 0° C. and 50° C. 8. The cartridge of claim 1, wherein the cartridge holds a maximum volume of 200 microliters. 9. The cartridge of claim 1, wherein the cartridge requires a minimum of 2 microliters for analysis. 10. The cartridge of claim 1, wherein a sample size is 100 microliters or less. 11. The cartridge of claim 1, wherein the cartridge is marked with a unique identification number. | CHIPCARE CORPORATION | C12Q001/06 A61B005/151 A61B005/15 A61B005/157 G01N015/14 G01N015/10 G01N015/00 B01L009/00 B01L003/00 G06K009/00 G06T007/00
US10719962 | B2 | 2020-07-21 | Magnetic field distortion calculation apparatus, method, and program | 2018-07-26 | In a magnetic field distortion calculation apparatus, method, and program, information of magnetic field distortion inside a subject can be accurately acquired.;An image acquisition unit acquires a reference image and a three-dimensional image of the head of the subject. A feature point detection unit detects a plurality of feature points from the three-dimensional image, and a virtual feature point estimation unit estimates a plurality of virtual feature points, which are to be present in the brain in the three-dimensional image, using the plurality of feature points. A magnetic field distortion information acquisition unit acquires magnetic field distortion information, which indicates spatial magnetic field distortion caused by a three-dimensional image capturing apparatus included in the three-dimensional image, by performing registration between the plurality of feature points and the plurality of virtual feature points and a plurality of reference points. | 1. A magnetic field distortion calculation apparatus, comprising: a processor configured toacquire a medical image, which is acquired by arranging an auxiliary tool in which a plurality of marking points are arranged according to a predetermined arrangement rule around a target part of a subject and imaging the target part around which the auxiliary tool is disposed using an MRI apparatus and which includes a plurality of feature points represented by the plurality of marking points and the target part,acquire a reference image including a plurality of reference points arranged corresponding to the plurality of marking points according to the arrangement rule;detect the plurality of feature points from the medical image;estimate a plurality of virtual feature points, which are to be present in the target part in the medical image, using the plurality of feature points; andacquire magnetic field distortion information, which indicates spatial magnetic field distortion caused by the MRI apparatus included in the medical image, by performing registration between the plurality of feature points and the plurality of virtual feature points and the plurality of reference points. 2. The magnetic field distortion calculation apparatus according to claim 1, wherein the processor further configured to remove the magnetic field distortion from the medical image based on the magnetic field distortion information. 3. The magnetic field distortion calculation apparatus according to claim 1, wherein the processor configured to perform rigid registration between the plurality of feature points and the plurality of virtual feature points and the plurality of reference points included in the reference image, andacquire deformation vectors between the corresponding feature points and virtual feature points and the corresponding reference points, as the magnetic field distortion information, based on the plurality of feature points and the plurality of virtual feature points and the plurality of reference points included in the reference image after the rigid registration. 4. The magnetic field distortion calculation apparatus according to claim 2, wherein the processor configured to:perform rigid registration between the plurality of feature points and the plurality of virtual feature points and the plurality of reference points included in the reference image, andacquire deformation vectors between the corresponding feature points and virtual feature points and the corresponding reference points, as the magnetic field distortion information, based on the plurality of feature points and the plurality of virtual feature points and the plurality of reference points included in the reference image after the rigid registration. 5. The magnetic field distortion calculation apparatus according to claim 3, wherein the processor configured to acquire the magnetic field distortion information by performing an interpolation operation, which is based on the deformation vec... | FUJIFILM Corporation | G06K009/00 G06T011/00 G01R033/56 G06T007/33 G01R033/565 A61B005/00 A61B005/055 G01R033/58 G06T005/00
US10702199 | B2 | 2020-07-07 | Blood collecting device | 2016-04-27 | A blood collecting device collects blood in a short time after puncture without scattering the blood into the open air other than a container for collecting the blood. One of representative blood collecting devices has a container unit that includes a screw portion, a holder that holds the container unit serving as a system whose one end is closed, a puncture unit that can be attached to the container or the holder, a through-hole of the puncture unit, and a puncture unit protection member that protects the puncture unit and the container unit. | 1. A blood collecting device comprising: a container unit that has an opening portion and a closing portion, and that holds the opening portion in a state where the opening portion faces upward;a fixing unit that fixes a blood collecting target in a state where the blood collecting target faces the opening portion;a pressure changing unit that pressurizes the blood collecting target or applies negative pressure into a space surrounded by the blood collecting target and the container, or that performs both of these in combination with each other;a puncture unit that is positioned below the container unit, and that punctures the blood collecting target after pressure is changed by the pressure changing unit; anda lid that closes the opening portion of the container containing blood,wherein the container unit has a slide mechanism which slides with respect to the fixing unit in a horizontal direction. 2. The blood collecting device according to claim 1, wherein the pressure changing unit includes a sphygmomanometer which serves as a compression body for pressurizing the blood collecting target by using pneumatic pressure, and which measures blood pressure while pressurizing the blood collecting target by using the compression body. 3. The blood collecting device according to claim 2, further comprising: a pressurizing source that pressurizes the blood collecting target so that the blood pressure obtained by the sphygmomanometer is in a range of the lowest blood pressure to the highest blood pressure. 4. The blood collecting device according to claim 1, further comprising: a cover that protects the blood collecting target after blood collecting. 5. The blood collecting device according to claim 1, wherein all of the container unit, the fixing unit, and the puncture unit are configured to be separable from each other. 6. The blood collecting device according to claim 1, further comprising: a blood collecting amount confirmation mechanism which confirms an amount of the blood contained inside the container unit. 7. The blood collecting device according to claim 1, wherein at least a portion of the container unit has a transparent window. 8. A blood collecting device comprising: an upper disk that has at least one through-hole;a lower disk that is positioned below the upper disk, and that is rotatable independent of the upper disk;a negative pressure source that applies negative pressure to a surrounding environment of a blood collecting target pressed against the through-hole;a puncture unit that is fixed to the lower disk, and that punctures the blood collecting target after pressure is changed by a pressure changing unit; anda container unit that has an opening portion and a closing portion, and that is fixed to the lower disk in a state where the opening portion faces upward. 9. The blood collecting device according to claim 8, wherein the container unit and the puncture unit are detachably fixed to the lower disk. 10. The blood collecting ... | HITACHI HIGH-TECH CORPORATION | A61B005/15 A61B005/022 A61B005/151 A61B005/02 A61B005/154
US10722182 | B2 | 2020-07-28 | Method and apparatus for heart rate and respiration rate estimation using low power sensor | 2016-05-31 | Provided is a system for measuring biological signals of a user, which includes a sensor module configured to acquire ballistocardiogram (BCG) signals of a user via a present channel, where the present channel is at least one channel of the sensor module, a decomposition module configured to decompose the BCG signals to decomposed signals, a reconstruction module configured to reconstruct at least a portion of the decomposed signals to reconstructed signals, a processing module configured to process the reconstructed signals to at least one of a heart rate, respiration rate, phases of respiration, and blood pressure, and a display module configured to display at least one output corresponding to the at least one of the heart rate, the respiration rate, phases of respiration, and the blood pressure on a display device. | 1. A system for measuring biological signals of a user, comprising: a sensor module configured to acquire ballistocardiogram (BCG) signals of the user via a present channel, wherein the present channel is at least one channel of the sensor module;a decomposition module configured to decompose the BCG signals into decomposed signals;a reconstruction module configured to reconstruct at least a portion of the decomposed signals into reconstructed signals using a probability distribution of the decomposed signals;a processing module configured to process the reconstructed signals to at least one of a heart rate, a respiration rate, phases of respiration, and a blood pressure; anda display configured to display at least one output corresponding to the at least one of the heart rate, the respiration rate, the phases of respiration, and the blood pressure,wherein the processing module comprises a diagnostic processor configured to select a next channel when motion is detected by the present channel, wherein the next channel and the present channel are channels of a same accelerometer in the sensor module, andwherein: the reconstruction module is configured to reconstruct the decomposed signals into reconstructed signals using time domain smoothing based on one or both of moving averages and the maximum modulus principle, andduring reconstruction, the diagnostic processor is configured to convert the BCG signals using a pre-defined frequency transfer function between the BCG signals and photoplethysmogram (PPG) signals. 2. The system of claim 1, wherein the decomposition module is configured to decompose the BCG signals using a function corresponding to one or more of a wavelet, a Hilbert transform, one or more finite impulse response (FIR)/infinite impulse response (IIR) filters with respective cut-off frequencies and stop bands, a time-domain based moving average, and multi-order derivatives. 3. The system of claim 1, wherein the reconstruction module is configured to reconstruct the at least portion of the decomposed signals using statistical averaging. 4. The system of claim 1, wherein the diagnostic processor is configured to determine the heart rate based on Cepstrum analysis of at least one of the BCG signals and the reconstructed signals. 5. The system of claim 1, wherein channel evaluation criteria for selecting the next channel is based on at least one of a respiration rate estimation confidence indicator determined by peaks in a very low frequency (VLF) band of the channel and a heart rate estimation confidence indicator determined by peaks in a spectrum of the reconstructed signals. 6. The system of claim 1, wherein the probability distribution is determined based on an energy contained in each of the decomposed signals. 7. A non-transitory machine-readable medium storing machine executable instructions that when executed cause a computing system to control operations for measuring biological signals of a user, the operations comprising... | Samsung Electronics Co., Ltd. | A61B005/02 A61B005/00 A61B005/11 A61B005/0205 A61B005/113 A61B005/024
US10702201 | B2 | 2020-07-07 | Instrumented device for analysing and quantifying physiological parameters of the upper limbs of a sports person, and associated method | 2016-05-25 | An instrumented interface for quantifying at least one physiological parameter of the upper limbs of a person, capable of being secured to a gripping device, and including at least one gripping system which the person can hang from, the instrumented interface including at least one force sensor which measures a deformation of the interface when a person exerts a force on the interface via the one gripping device, and is capable of producing a value that translates a force value. | 1. A method for modelling an exercise capacity of an individual's upper limbs, implemented by an intelligent device, said method comprising: a step of configuration, consisting in receiving at least one message containing a weight value, a value representing a notion of difficulty and a value representing a gripping technique;a step of reception of measuring data consisting in receiving at least one measurement message containing at least one force value associated with a time value transmitted by an instrumented interface, said instrumented interface: securely attached to a gripping device, the gripping device having at least one gripping system allowing the individual to be suspended, said instrumented interface including at least one force sensor measuring a deformation of said instrumented interface when an individual exerts a force on said instrumented interface via said gripping device, and the at least one force sensor configured to produce a value representing a force value;a step of calculation of at least one physiological parameter representative of the performance of an individual's forearms or arms, from at least said difficulty value and a force value associated with a time value;a step of transmission consisting in transmitting at least one message containing a value representative of a physiological parameter representing the performance of an individual's forearms or arms, calculated in the previous step. 2. The method for modelling the exercise capacity of an individual's upper limbs according to claim 1, wherein said weight value and/or said value representing a notion of difficulty and/or value representing a gripping technique is a local value entered by a Man-Machine Interface or a value received from a remote source via a communication system. 3. The method for modelling the exercise capacity of an individual's upper limbs according to claim 1, wherein said value representing a notion of difficulty is a value relative to the gripping thickness of the at least one gripping system or a value relative to the inclination of the at least one gripping system. 4. The method for modelling the exercise capacity of an individual's upper limbs according to claim 1, further comprising a step of assistance consisting in transmitting at least one message containing a value relative to a start and an end of an exercise to be performed by said individual. 5. The method for modelling the exercise capacity of an individual's upper limbs according to claim 1, wherein the calculation step consists in calculating: an index representative of the force of the forearms, where the index is a function of said force value associated with a time value, and a function of an equation which is a function of said value representing a notion of difficulty and of the selected gripping technique and/or;an index representative of the resistance of the individual's forearms, where the index is a function of said force value associated with a time value and... | CENTRE NATIONAL DE LA RECHERCHE SCIENTIFIQUE (CNRS) UNIVERSITE D'AIX MARSEILLE | A61B005/22 G16H050/30 A61B005/00 A63B069/00 A63B023/12 A63B001/00 A63B021/00 A63B071/06 H04M001/725 G16H020/30
US10714208 | B2 | 2020-07-14 | Computer systems for treating diseases | 2016-01-12 | The present invention relates to computer-implemented methods and system for analysing a biomarker which cycles in a subject. In some other aspects, the present invention relates to analysing a biomarker which at least initially increases or decreases in amount in a subject following a treatment for a disease. In further aspects, the present invention relates to computer-implemented methods and systems for determining a preferred time to administer a therapy to treat a disease in a subject. The present invention also relates to computer program product to implement the methods. Further, the present invention relates to methods of determining the timing of treating a disease in a subject in which the immune system is cycling. | 1. A non-transitory computer-readable information storage medium having stored thereon instructions, that if executed by one or more processors, perform a method comprising: i) obtaining from a first memory measurements of an immune system marker in samples from a subject having a disease characterized by cycling immune system markers;ii) obtaining from a second memory a model of cycling of the immune system marker;iii) fitting the measurements to the model;iv) obtaining a best-fit curve of the measurements to the model to estimate a periodicity of the cycling of the immune system marker;v) projecting the obtained best-fit curve into the future; andvi) outputting a preferred time in the future, based on the projected best-fit curve, to administer a therapy to treat the disease. 2. The computer-readable information storage medium of claim 1, wherein the best-fit curve is a fit to a periodic model of the cycling of the immune system marker, and obtaining the best-fit curve further comprises fitting the periodic model to the measurements multiple times using different initial conditions for each fit. 3. The computer-readable information storage medium of claim 1, wherein the best-fit curve is a fit to a harmonic model of the cycling of the immune system marker, and obtaining the best-fit curve further comprises fitting the harmonic model to the measurements multiple times using different initial conditions for each fit. 4. The computer-readable information storage medium of claim 1, wherein the method further comprises generating a confidence measure by comparing a similarity in a fit result produced by each fitting. 5. The computer-readable information storage medium of claim 1, wherein obtaining the best-fit curve further comprises imposing a box constraint on at least some parameters of the model to guide optimization to biologically realistic regions. 6. The computer-readable information storage medium of claim 1, wherein the model has a heavy tail to allow for noisy nature of the measurements. 7. The computer-readable information storage medium of claim 1, wherein the method further comprises determining a variance-stabilizing transformation of the measurements prior to determining the periodicity based on the transformed measurements. 8. The computer-readable information storage medium of claim 7, wherein the variance-stabilizing transformation of measurements is logarithm of the measurements. 9. The computer-readable information storage medium of claim 1, wherein the measurements are for two or more different immune system markers which are cycling in the subject. 10. The computer-readable information storage medium of claim 1, wherein the disease is cancer, an infection, a degenerative disease, an autoimmune disease or transplant rejection. 11. The computer-readable information storage medium of claim 10, wherein the disease is cancer. 12. The computer-readable information storage medium of claim 1, wherein the immune system marke... | BIOTEMPUS PTY LTD | G16B005/00 G16B040/00 G01N033/564 G01N033/574 G16H050/50 A61B005/00 G06F019/00 A61K031/495 G01N033/68
US10716672 | B2 | 2020-07-21 | System and method for intraprocedural assessment of geometry and compliance of valve annulus for trans-catheter valve implantation | 2016-04-01 | A system for determining the dimensions of a native valve annulus may include a sensing catheter, a transmitter, and a computing device. The sensing catheter may include a shaft extending in a longitudinal direction; a plurality of splines attached to the shaft, the splines having an expanded condition spaced radially outward from the shaft; and at least one sensing coil located on each of the splines. The transmitter may generate a magnetic field to induce a potential in each of the sensing coils, and the computing device may identify the positions of the sensing coils based on the induced potentials. | 1. A system for determining the dimensions of a native valve annulus, the system comprising: a sensing catheter that includes: a shaft extending in a longitudinal direction;a plurality of splines positioned about the shaft, each spline having a first end fixedly attached to the shaft and a second end, the splines having an unconstrained expanded condition spaced radially outward from the shaft, the splines in the unconstrained expanded condition collectively defining a plurality of circles that are perpendicular to the longitudinal direction;a ring slidably mounted around the shaft, the second end of each of the splines being attached to the ring; andeach of the splines including a sensing coil located at an expansion center of the spline, the expansion center being a position on the spline such that the expansion centers of the plurality of splines collectively define a largest one of the plurality of circles;a transmitter configured to generate a magnetic field to induce a potential in each of the sensing coils; anda computing device configured to identify positions of the sensing coils based on the induced potentials. 2. The system of claim 1, wherein the sensing catheter includes an expandable balloon positioned between the shaft and the splines. 3. The system of claim 1, wherein the transmitter is an array of drive coils. 4. The system of claim 1, further comprising: a display configured to output the position of each of the sensing coils. 5. The system of claim 1, wherein the sensing catheter further comprises: a spline sheath slidable relative to the shaft between a first position overlying the splines and a second position exposing the splines for movement toward the expanded condition. 6. A method for determining the dimensions of a native valve annulus, the method comprising: delivering a sensing catheter to a native valve annulus, the sensing catheter including a shaft extending in a longitudinal direction and a plurality of splines, each spline having a first end fixedly attached to the shaft and a second end, the splines having a contracted condition and an unconstrained radially expanded condition, the splines in the unconstrained expanded condition collectively defining a plurality of circles that are perpendicular to the longitudinal direction, each of the splines including a sensing coil located at an expansion center of the spline, the expansion center being a position on the spline such that the expansion centers of the plurality of splines collectively define a largest one of the plurality of circles;deploying the plurality of splines from the contracted condition to a second radially expanded condition in which ones of the sensing coils contact the native valve annulus and each of the sensing coils lies in a first plane relative to the native valve annulus;generating at least one magnetic field that induces a first potential in each of the sensing coils;determining a first position of each of the sensing coils based on t... | St. Jude Medical, Cardiology Division, Inc. | A61F002/24 A61B005/107 A61B005/00 A61B017/00 A61B018/00
US10726550 | B2 | 2020-07-28 | Method for proposing personalized cosmetics | 2016-03-30 | The present invention provides a method for proposing personalized cosmetics and a system for implementing the same, the method comprising the steps of: photographing the visible spectrum of the skin of a subject to be measured; mapping a personalized visible spectrum to the measured visible spectrum of the skin so as to generate a spectral comparison table; matching a personalized composition to the spectral comparison table; and displaying the matched personalized composition on an image display unit. | 1. A method for proposing a personalized cosmetic composition, comprising the steps of: photographing a visible spectrum of a skin of a subject to be measured to prepare a measured visible spectrum of the skin;generating a spectral comparison table by mapping a personalized visible spectrum to the measured visible spectrum of the skin;matching a personalized cosmetic composition to the spectral comparison table; anddisplaying the matched personalized cosmetic composition on an image display unit,wherein the personalized visible spectrum is a preset obtained by performing a step of photographing a standard group of skin conditions preferred according to a fashion to calculate a mean intensity value for each wavelength,wherein the step of generating a spectral comparison table comprises storing a data pair comprising W and I in the spectral comparison table by comparing the measured visible spectrum of the skin with the personalized visible spectrum for each wavelength range thereof, wherein the W of the data pair refers to a wavelength range, andwherein the I of the data pair refers to a difference value between an intensity of the personalized visible spectrum and an intensity of the measured visible spectrum of the skin in the W,wherein the step of matching a personalized cosmetic composition to the spectral comparison table includes matching a component of the personalized cosmetic composition to at least one data pair stored in the spectral comparison table,wherein the step of matching a component of the personalized cosmetic composition includes matching a corresponding wavelength reinforcement material to the W of the at least one data pair stored in the spectral comparison table, andwherein the wavelength reinforcement material is one wavelength reinforcement powder or two or more wavelength reinforcement composite powders which selected from the group consisting of mica, synthetic mica, alumina, borosilicate powder, talc, and sericite. 2. The method according to claim 1, wherein the personalized visible spectrum is at least one preset spectrum selected from the group consisting of a moisture glow-mimicking spectrum, a shine-mimicking spectrum, a skin color-complementing spectrum, and a milky-color-mimicking spectrum, and obtained by performing a step of photographing a standard group of preferred skin conditions according to make-up trend to calculate a mean intensity value for each wavelength, in which:the moisture glow-mimicking spectrum refers to a patterned spectrum that appears in a state in which a skin is moisture-supersaturated,the shine-mimicking spectrum is a patterned spectrum that appears in a state in which a skin is oil-supersaturated,the skin color-complementing spectrum is a patterned spectrum that appears in a state in which the intensity in a red wavelength range is reinforced, andthe milky-color-mimicking spectrum is a patterned spectrum that appears when a yellow wavelength is reduced in a state in which the intensity ... | Amorepacific Corporation | G06K009/62 G06T007/00 A61B005/00 G06Q030/06 A61B005/103 A61K008/18
US10716949 | B2 | 2020-07-21 | Changing cardiac shock delivery parameters based on a transform value | 2018-03-19 | A medical device that includes a power source, a therapy delivery interface, therapy electrodes, electrocardiogram (ECG) sensing electrodes to sense ECG signal of a heart of a patient, a sensor interface to receive and digitize the ECG signal, and a processor. The processor is configured to analyze the ECG signal to determine a cardiac rhythm and a transform value representing a magnitude of a frequency component of the cardiac rhythm, analyze the cardiac rhythm and the transform value to detect a shockable cardiac arrhythmia by classifying the cardiac rhythm as a noise rhythm or a shockable cardiac arrhythmia rhythm based on the transform value, and causing the processor to detect the cardiac arrhythmia if classifying the cardiac rhythm as a shockable cardiac arrhythmia rhythm, initiate a treatment alarm sequence, adjust the shock delivery parameter for a defibrillation shock, and provide the defibrillation shock via the therapy electrodes. | 1. An ambulatory medical monitoring and treatment device comprising: a power source;a therapy delivery interface, coupled to the power source;a plurality of therapy electrodes coupled to the therapy delivery interface;a plurality of electrocardiogram (ECG) sensing electrodes to sense ECG signal of a heart of a patient;a sensor interface to receive and digitize the ECG signal; andat least one processor coupled to the sensor interface and the therapy delivery interface, the at least one processor configured to: analyze the ECG signal to determine a cardiac rhythm and a first transform value that represents a magnitude of at least one frequency component of the cardiac rhythm,analyze the cardiac rhythm and the first transform value to detect a shockable cardiac arrhythmia by classifying the cardiac rhythm as one of a noise rhythm and a shockable cardiac arrhythmia rhythm based at least in part on the first transform value, andcausing the at least one processor to detect the cardiac arrhythmia in response to classifying the cardiac rhythm as a shockable cardiac arrhythmia rhythm,initiate a treatment alarm sequence requesting a response from the patient to verify that the patient is unconscious in response to detecting the shockable cardiac arrhythmia,initiate a release of a conductive gel on the patient's chest,determine a second transform value that represents a current magnitude of the at least one frequency component of the cardiac rhythm after the release of the conductive gel,adjust, based on the first and second transform value, at least one shock delivery parameter for a defibrillation shock, andcause the therapy delivery interface to provide the defibrillation shock via the plurality of therapy electrodes based on the at least one shock delivery parameter on receiving no response from the patient. 2. The device of claim 1, wherein the at least one processor is configured to initiate the treatment alarm sequence to generate an alert in response to detecting the shockable cardiac arrhythmia, the alert indicating detection of the shockable cardiac arrhythmia. 3. The device of claim 2, wherein the alert is a first alert and prior to initiating the release of the conductive gel, the at least one processor is configured to generate a second alert indicating shock delivery. 4. The device of claim 1, wherein the at least one processor is configured to perform a comparison between the first transform value and the second transform value and adjust the at least one shock delivery parameter based on the comparison. 5. The device of claim 4, wherein the first transform value and the second transform value are determined over a window of time of 100 milliseconds to twenty seconds. 6. The device of claim 4, wherein the first alert is generated over a window of time of about twenty seconds and the second alert is generated over a window of time of about fifteen seconds. 7. The device of claim 1, wherein the at least one shock delivery parameter is a p... | ZOLL Medical Corporation | A61N001/39 A61N001/04 A61B005/0408
US10709365 | B2 | 2020-07-14 | System and method for noninvasive analysis of subcutaneous tissue | 2017-06-20 | A method for noninvasive analysis of subcutaneous tissue includes irradiating a surface of the tissue with short wave infrared (SWIR) radiation in a first spectral band that is strongly absorbed by water, and with SWIR radiation in a second spectral band such that an interaction of the radiation in both spectral bands with a component of the tissue other than water is substantially identical. An intensity of the radiation in each of the spectral bands that emerges from the tissue is measured. A relative absorption by the tissue of radiation in one of spectral bands relative to absorption by the tissue of radiation in the other of the spectral bands is calculated. A state of the tissue is determined in accordance with the calculated relative absorption. | 1. A method for noninvasive analysis of tissue, the method comprising: irradiating, with at least one source of infrared radiation, a surface of the tissue with short wave infrared (SWIR) radiation in a first spectral band that is strongly absorbed by water, and with SWIR radiation in a second spectral band such that an interaction of the radiation in both spectral bands with a component of the tissue other than water is substantially identical;measuring, with at least one radiation detector, an intensity of the radiation that emerges from the tissue in each of the spectral bands;determining change in at least one of shape and intensity of signals received by the at least one radiation detector;calculating a relative absorption by the tissue of radiation in one of the first and second spectral bands relative to absorption by the tissue of radiation in the other of the first and second spectral bands; anddetermining, by a processor, a subcutaneous deep tissue injury under intact skin in which liquids accumulate subcutaneously, in accordance with the calculated relative absorption and in accordance with determined change in the received signal,wherein the interaction of the radiation comprises at least one of absorption and scattering. 2. The method of claim 1, wherein the first spectral band is in the wavelength range of 1400 nm to 1500 nm, and wherein the second spectral band is in the wavelength range of 1000 nm to 1350 nm or 1500 nm to 2100 nm. 3. The method of claim 1, further comprising applying pressure on surface of the tissue in a periodical regime. 4. The method of claim 1, wherein a gap in wavelength between the first and second spectral bands is less than 200 nm. 5. The method of claim 1, wherein measuring the intensity comprises measuring the intensity of the radiation that is transmitted across the tissue. 6. The method of claim 5, wherein the tissue comprises tissue of a finger or an ear. 7. The method of claim 1, wherein measuring the intensity comprises measuring the intensity of the radiation that is reflected by the tissue. 8. The method of claim 1, wherein determining the subcutaneous deep tissue injury comprises determining a concentration of a substance in blood within the tissue. 9. The method of claim 8, wherein the substance is an introduced substance. 10. A system for noninvasive analysis of tissue, the system comprising: at least one source of infrared radiation to irradiate the tissue, the infrared radiation including SWIR radiation in a first spectral band that is strongly absorbed by water, and including radiation in a second spectral band such that an interaction of the radiation in both spectral bands with a component of the tissue other than water is substantially identical;at least one radiation detector to measure an intensity of radiation in each of the two spectral bands that emerges from the tissue; anda processor that is configured to calculate a relative absorption by the tissue of radiation in one o... | I. R. Med Ltd. | A61B005/1455 A61B005/00
US10726545 | B2 | 2020-07-28 | Systems and methods for administering treatment of atrial fibrillation | 2018-01-22 | Methods and systems are described for treatment of atrial fibrillation. Certain embodiments provide methods of assessing a risk of thromboembolic stroke in a patient. The method comprises acquiring image data and determining an indicator of a degree of fibrosis of a patient's left atrium (LA) based on the image data. A treatment modality is then determined and administered to the patient based on the analysis of the image data. | 1. A method of administering treatment of atrial fibrillation (AF) to a patient, the method comprising: capturing MR image data of left atrium tissue of the patient by performing a delayed-enhancement magnetic resonance imaging (DE-MRI) procedure using an MRI system, wherein the MR image data is captured before performing a first ablative treatment procedure;automatically determining, by an electronic processor, an amount of delayed intravascular contrast enhancement in the MR image data relative to a total LA tissue volume, wherein the amount of delayed intravascular contrast enhancement is positively proportional to an amount of unhealthy tissue in a wall of the left atrium of the patient, the amount being determined with the electronic processor by defining epicardial and endocardial borders of the left atrium in the MR image data to isolate the wall of the left atrium,determining the total tissue volume of the left atrium wall based on the isolated wall of the left atrium in the MR image data,identifying, in the MR image data, a number of voxels within the left atrium wall that exhibit delayed intravascular contrast enhancement indicative of unhealthy tissue based on a pixel intensity histogram, anddetermining the amount of delayed intravascular contrast enhancement relative to the total tissue volume of the left atrium wall based on the determined number of voxels that exhibit delayed intravascular contrast enhancement indicative of unhealthy tissue and the determined total tissue volume of the left atrium wall;determining, by the electronic processor, a risk of recurrent atrial fibrillation after the first ablative treatment procedure wherein the electronic processor is configured to determine that the risk of recurrent atrial fibrillation after the first ablative treatment procedure is significant in response to a determination by the electronic processor that the amount of delayed intravascular contrast enhancement in the MR image data relative to the total tissue volume of the left atrium wall is greater than a threshold amount, andwherein the electronic processor is configured to determine that the risk of recurrent atrial fibrillation after the first ablative treatment procedure is not significant in response to a determination by the electronic processor that the amount of delayed intravascular contrast enhancement in the MR image data relative to the total tissue volume of the left atrium wall is less than the threshold amount;performing the first ablative treatment procedure in response to a determination by the electronic processor that the risk of recurrent atrial fibrillation after the first ablative treatment procedure is not significant; andtreating the atrial fibrillation with a different treatment modality in response to a determination that the risk of recurrent atrial fibrillation after the first ablative treatment procedure is significant. 2. The method of claim 1, wherein the unhealthy tissue present in the LA wall compr... | UNIVERSITY OF UTAH RESEARCH FOUNDATION | G06T007/12 G06T007/00 A61B005/055 A61B005/00 A61B006/03 A61B006/00 G16H050/30 G16H050/20 G06F019/00
US10702165 | B2 | 2020-07-07 | Estimation of human core temperature based on heart rate system and method | 2017-03-04 | The invention in at least one embodiment includes a method for determining the core body temperature of a person by setting an initial core body temperature with a processor; receiving a heart rate of the person with the processor; calculating a predicted core body temperature with the processor using an extended Kalman filter based on the heart rate and the initial core body temperature; and providing the predicted core body temperature. In another embodiment, a system for performing the method. | 1. A method for indirectly determining a core body temperature of a person comprising: setting an initial core body temperature at a mobile device comprising a heart rate monitor having a plurality of electrodes, a processor, and a display;receiving at least two heart rates of the person with the processor, where the heart rates are detected by the heart rate monitor using the plurality of electrodes attached to the person;calculating a predicted core body temperature with the processor using an extended Kalman filter based solely on the second heart rate and the initial core body temperature; andproviding the predicted core body temperature to the display, andwherein the initial core body temperature is set by the processor based solely on the first heart rate of the person. 2. The method according to claim 1, further comprising: receiving an additional heart rate of the person with the processor, where the additional heart rate is detected by the heart rate monitor;calculating a new predicted core body temperature with the processor using the extended Kalman filter based solely on the additional heart rate and the last predicted core body temperature; andrepeating the receiving additional heart rate and calculating a new predicted core body temperature at predetermined intervals. 3. The method according to claim 2, further comprising: when the new predicted core body temperature exceeds a predetermined threshold, sending an alarm signal from said processor prior to proceeding to repeating, wherein the alarm signal triggers an alarm; andwhen the new predicted core body temperature does not exceed a predetermined threshold, proceeding to repeating. 4. The method according to claim 2, further comprising obtaining the predetermined interval. 5. The method according to claim 1, further comprising recording the predicted core body temperature in memory in the mobile device; and wherein the mobile device is a processing hub on the person being monitored, andthe extended Kalman filter includes a time update phase and an observation update phase. 6. The method according to claim 1, wherein providing the predicted core body temperature includes transmitting the predicted core body temperature with a transmitter in the mobile device to an external device. 7. The method according to claim 1, wherein calculating a predicted core body temperature includes calculating a preliminary core body temperature estimate;calculating a preliminary estimate of the variance of the preliminary core body temperature estimate;calculating an extended Kalman filter mapping function variance coefficient;calculating a Kalman gain weighting factor based on the preliminary estimate of variance and the extended Kalman filter variance coefficient; andcalculating the predicted core body temperature using a preliminary core body temperature estimate, an error between the new heart rate and an expected heart rate given the preliminary estimate of the core body temperature. 8. T... | The Government of the United States, as Represented by the Secretary of the Army | A61B005/01 A61B005/024 A61B005/00
US10713921 | B2 | 2020-07-14 | Control device for controlling an alarm output and method for controlling an alarm output as well as medical device | 2020-01-06 | A control device, controlling an alarm output, includes a data network interface, with a signaling interface for outputting a control signal indicating a request for an acoustic and/or optical alarm signal a memory unit and a processor. The data network interface is configured to receive a group message indicating a sender identity and a presence of an alarm state. The memory unit provides a first data set, which indicates a list with potential sender identities, as well as further a second data set, which indicates one or more alarm output time periods. The processor is configured to operate a timekeeping function, and to output the control signal via the signaling interface as a function of an agreement between the sender identity with one of potential sender identities, and as a function of a comparison of a current value of the timekeeping function with data of the second data set. | 1. A control device for controlling an alarm output, the control device being a part of a network unit, the control device comprising: a data network interface connected to a data network and being configured to receive a group message via the data network, which group message is addressed and sent to a plurality of network units connected to the data network, the plurality of network units to which the group message is addressed and sent including the network unit of which the control device is a part and which group message indicates a sender identity identifying a sender of the group message and which further indicates a presence of an alarm state;a signaling interface for outputting a control signal, which control signal indicates a request for an output of an acoustic and/or optical alarm signal;at least one memory unit to provide a first data set, which indicates a list with potential sender identities, as well as further a second data set, which indicates one or more time periods, during which an alarm signal shall be outputted;an input interface for receiving an input signal, which indicates an acknowledgment of the alarm signal by a user; andat least one processor configured to operate a timekeeping function, with which a start of timekeeping is started on arrival of the group message at the network unit of which the control device is a part, the timekeeping comprising monitoring a time duration from the start of the timekeeping to a current time duration value, and to output the control signal via the signaling interface as a function of an agreement of the sender identity indicated in the group message with one of the potential sender identities of the first data set as well as further as a function of a comparison of the current time duration value of the timekeeping of the timekeeping function with the one or more time periods data of the second data set, wherein the processor is further configured to suppress the output of the control signal and further to send a status message, which status message indicates the acknowledgment of the alarm signal by the user, to the sender of the group message via the data network interface in the presence of the input signal. 2. A control device in accordance with claim 1, wherein the processor is configured to suppress the output of the control signal no later than when suppression of the output of the control signal is indicated by the one or more time periods of the second data set. 3. A control device in accordance with claim 1, wherein: the group message has further a data element that indicates a type of an alarm state; the second data set further indicates an assignment of the time periods to types of alarm states; andthe processor is configured to take into account the type of alarm state indicated in the group message in the course of the comparison. 4. A control device in accordance with claim 1, wherein the processor is further configured: to check, on the basis of a status informati... | DRÄGERWERK AG & CO. KGAA | G08B021/18 A61B005/00 G08B029/12
US10702179 | B2 | 2020-07-07 | Method for sensing arrhythmia for providing user feedback | 2018-08-29 | A two-dimensional visual display that allows for the user and/or another individual, such as the user's spouse, to experience self-evident, abnormally complex elevated heart rate arrhythmia. The system of the present invention uses a heart rate monitor, along with data analysis, to provide the user with a visual display, seismic feel, and/or aural feedback that allow for the user to be quickly alerted as to any arrhythmia, thereby signaling to the user to change their activity in order to avoid sudden cardiac death. | 1. A method for sensing arrhythmia and providing arrhythmia feedback to a user, the method comprising: detecting arrhythmia via a heart sensor when the user is either asymptomatic or symptomatic;evaluating heart rate patterns based on data collected from the heart sensor, the data including beats per minute;converting the data collected from the heart sensor, including beats per minute to audio data;defining an arrhythmia control pattern based on the data collected from the heart sensor, the arrhythmia control pattern being associated with an asymptomatic arrhythmia of the user;defining one or more patterns from the data collected from the heart sensor that exceed the arrhythmia control pattern, thereby indicating arrhythmia outside of the control pattern;defining cardiac event triggers based on the arrhythmia control pattern and the one or more patterns and creating a database of cardiac event triggers;providing the user with a two-dimensional display; andproviding the user with physical feedback based on the one or more patterns, the arrhythmia control pattern, and the cardiac event triggers, the physical feedback being one of audio feedback, visual feedback, or vibrational feedback and the physical feedback being an indication of one of the cardiac event triggers such that the user is altered to a cardiac event;wherein the feedback provides the user with details to change their behavior to reduce stress, and further to reduce a risk of sudden cardiac arrest (SCA), silent cerebral ischemia (SCI), and/or sudden cardiac death (SCD). 2. The method of claim 1, wherein the heart sensor is a chest sensor. 3. The method of claim 1, wherein the heart sensor is an aural sensor. 4. The method of claim 1, further comprising: transmitting the audio data to an ear sensor to relay the audio to the user's inner ear. 5. The method of claim 1, further comprising: detecting SpO2 levels through an SpO2 sensor and logging SpO2 data collected;wherein the SpO2 data can be used to determine a change in SpO2 levels associated with arrhythmia. 6. The method of claim 1, further comprising: displaying, in real-time, a pulse-to-pulse display of heart rate and SpO2 levels to the user on a computing device. 7. The method of claim 1, further comprising: encouraging behavior via a computing device to the user to aid in reducing stress. 8. The method of claim 1, further comprising: corresponding the one or more patterns and the arrhythmia control pattern to seismic-feel to the user via a vibration sensor to be carried by the user. | - | A61B005/04 A61B005/044 A61B005/0205 A61B005/0472 A61B005/0468 A61B005/024 A61B005/0464 A61B005/1455 A61B005/0245
US10702155 | B2 | 2020-07-07 | Transmission handling of medical devices | 2017-04-27 | A method for handling an Implantable Medical Device, IMD, transmission from an IMD provider such that the transmission comprises one or more alerts and/or IMB data attributes indicative for a status of the IMD or the patient. The alert comprises a provider specific alert priority. The method comprises the following steps: retrieving the IMD transmission from the IMD provider; and assigning to the one or more alerts a generic alert severity based on a predetermined mapping between the provider specific alert priority and the generic alert severity; and/or calculating an IMD data severity based on the IMD data attributes; and calculating a transmission severity based on the generic alert severity assigned to the one or more alerts and/or the IMD data severity. | 1. A computer-implemented method for handling an Implantable Medical Device, IMD, transmission from an IMD provider; and wherein the IMD transmission is related to a patient and comprises one or more alerts and/or IMD data attributes indicative for a status of the IMD or the patient; andwherein an alert comprises a provider specific alert priority;wherein the method comprises the following steps:providing a non-transitory computer-readable medium comprising instructions executable by a computing system of an 1 MB alert-management system for automatically performing the method;retrieving the IMD transmission from the 1 MB provider; andassigning to the one or more alerts generic alert severity based on a predetermined mapping between the provider specific alert priority and the generic alert severity;and/or calculating an IMD data severity based on the IMD data attribute; andcalculating a transmission severity based on the generic alert severity assigned to the one or more alerts and/or the 1 MB data severity. 2. The method according to claim 1 wherein the method further comprises: overriding the generic alert severity of each of the one or more alerts by an alert specific override factor associated with a type of the respective alert or the patient. 3. The method according to claim 1 wherein the calculating the transmission severity comprises calculating a weighted sum between the IMD data severity and the generic alert severity of the one or more alerts of the IMD transmission. 4. The method according to claim 1 wherein the method further comprises calculating a patient follow-up priority based on the transmission severity. 5. The method according to claim 4 wherein the calculating the patient follow-up priority is further based on a patient specific weight factor. 6. The method according to claim 4 wherein the calculating the patient follow-up priority is further based on one or more weighted earlier transmission severities obtained from earlier IMD transmissions. 7. The method according to claim 6 further comprising classifying the patient into a priority group according to the calculated patient follow-up priority. 8. The method according to claim 7 wherein the priority group is selected from the group of: a first level indicative for a low medical priority, anda second level indicative for a medium medical priority, anda third level indicative for a high medical priority. 9. A computer program product comprising computer-executable instructions for performing the method according to claim 1 when the program is run on a computer. 10. A data processing system programmed for carrying out the method according to claim 1. 11. The computer-implemented method of claim 1, wherein the computing system comprises a bus, a processor, a memory, and a storage device. | LINDACARE NV | G08B001/08 A61B005/00 A61N001/372 G08B023/00 A61N001/365 G06F016/95
US10722711 | B2 | 2020-07-28 | Device for non-invasive neuro-stimulation by means of multichannel bursts | 2016-01-29 | A device is provided for the stimulation of neurons that includes a non-invasive stimulation unit to generate stimuli in multiple stimulation channels, where the stimulation unit stimulates a neuron population in the brain and/or spinal cord of a patient in different locations for each of the stimulation channels. Moreover, the device includes a control unit that controls the stimulation unit to generate repetitive bursts in each of the stimulation channels, where each of the bursts includes multiple stimuli and is designed so that they do not reset the phase of the neuronal activity of the respective stimulated neurons. | 1. An apparatus for stimulating neurons, the apparatus comprising: a non-invasive stimulator configured to generate stimuli in a plurality of stimulation channels that stimulate a neuron population in at least one of a brain and a spinal cord of a patient by the plurality of stimulation channels at respective different points in at least one of the brain and the spinal cord;a measurement unit configured to record measurement signals that represent a neuronal activity of the neuron population stimulated with the stimuli; anda control unit that controls the stimulator to generate repetitive bursts that each comprise a plurality of stimuli in each of the stimulation channels, wherein the bursts generated in different stimulation channels of the plurality of stimulation channels are delayed in time with respect to one another,wherein the stimulator is configured to generate the repetitive bursts such that the repetitive bursts do not reset a phase of the neuronal activity of the respective stimulated neurons,wherein the control unit is configured to determine whether the repetitive bursts generated by the stimulator do not reset the phase of the neuronal activity of the respectively stimulated neurons based on the measurement signals,wherein, when the control unit determines that the repetitive bursts reset the phase of the neuronal activity of the respectively stimulated neurons, the control unit is further configured to control the stimulator such that subsequently generated bursts have a reduced stimulation strength, andwherein the control unit is configured to continue to control the stimulator to subsequently generate bursts having a reduced stimulation strength until the control unit determines that the generated bursts do not reset the phase of the neuronal activity of the respectively stimulated neurons. 2. The apparatus in accordance with claim 1, wherein the stimulator is configured to generate stimuli that are selected from the group of members consisting of acoustic stimuli, visual stimuli, tactile stimuli, vibratory stimuli, olfactory stimuli, gustatory stimuli, transcutaneous electric stimuli, transcutaneous magnetic stimuli, transcranial electrical stimuli, transcranial magnetic stimuli, ultrasound stimuli and combinations thereof. 3. The apparatus in accordance with claim 1, wherein two respective bursts following one another in time and generated in different stimulation channels of the plurality of stimulation channels are delayed in time with respect to one another by Tstim/S, wherein Tstim=1/fstim and fstim is a frequency in the range of 1 to 30 Hz and S being a number of the plurality of stimulation channels. 4. The apparatus in accordance with claim 1, wherein the stimulator is configured to generate the repetitive bursts in a time pattern comprising consecutive cycles that respectively have a duration Tstim, wherein Tstim=1/fstim and fstim is a frequency in the range of 1 to 30 Hz. 5. The apparatus in accordance with claim ... | Peter Alexander Tass | A61N001/36 A61B005/00 A61N005/06 A61N002/00 A61N002/02 A61B005/048 A61H023/02 A61B005/04 A61H023/00 A61N007/00 A61B005/11 A61B005/0488 A61B005/053 A61N001/04
US10709343 | B2 | 2020-07-14 | Computer-based systems and methods for monitoring the heart muscle of a patient with contextual oversight | 2017-07-31 | A computer-based system for monitoring the heart muscle function of a patient, with contextual oversight, includes a sensor array for collecting physiological and environmental data that are pertinent to the patient. A context register is also included which contains periodically updated patient-specific data that establishes a relevant contextual oversight capability for the system. In operation, a computer identifies anomalies in the physical data and detects aberrations in the environmental data. These anomalies and aberrations are then interactively evaluated together, relative to the contextual oversight capability, to determine whether clinical intervention for the patient is warranted. | 1. A computer-implemented system for monitoring the heart muscle function of a patient which comprises: a sensor array positionable on the patient for collecting physiological and environmental data pertinent to the patient;a cardio-profile for receiving the physiological data from the sensor array to manifest the physiological aspects of the heart muscle function of the patient;a response matrix for receiving the environmental data from the sensor array to manifest the external factors influencing the patient;a context register for storing patient data, wherein the patient data includes medical records of the patient, routine behavioral patterns of the patient, a care plan for the patient, and notes regarding perceptions recorded by physicians and clinical staff pertinent to the patient, wherein the patient data is continuously updated;a first analyzer for receiving updated patient data from the context register, wherein the first analyzer determines whether a change order based on the updated patient data is required to alter the cardio-profile;a second analyzer for receiving updated patient data from the context register, wherein the second analyzer determines whether a change order based on the updated patient data is required to alter the response matrix; anda comparator for comparing anomalies identified by the cardio-profile with aberrations detected by the response matrix, wherein the cardio-profile and the response matrix include alterations required by updated patient data from the context register, to evaluate the heart muscle function, and for generatinga report based on the comparator evaluation, wherein the report is prepared for determining whether a clinical intervention is warranted. 2. The system of claim 1 wherein the sensor array comprises: a cardiac sensor for collecting cardiac input data; andat least one perturbation sensor for collecting perturbation input data. 3. The system of claim 2 wherein the cardio-profile includes measurable parameters from the cardiac input data and establishes acceptable ranges, R, for variations in individual parameters of the cardiac input data, and wherein the measurable parameters are based on a waveform of the heart muscle and are selected from the group consisting of waveform shape characteristics, amplitudes within the waveform, the repetition rate of heart function cycles in the waveform variability of the waveform shape, discontinuities in the waveform and variability of the repetition rate. 4. The system of claim 3 wherein a change order for the cardio-profile involves an alteration of at least one of the ranges R. 5. The system of claim 4 wherein an anomaly is identified when an individual parameter in the cardiac input data extends beyond an acceptable range R in the cardio-profile. 6. The system of claim 2 wherein the response matrix comprises: a plurality of measurable parameters taken from the perturbation input data, wherein each measurable parameter results from an external ... | National Cardiac, Inc. | A61B005/024 G06F019/00 G16H040/63 A61B005/00 G16H050/20 A61B005/0432
US10709372 | B2 | 2020-07-14 | System for monitoring incontinence | 2018-12-12 | According to one aspect of the inventive concept there is provided a system for monitoring incontinence comprising: a urine sensitive circuit arranged to present a changed electrical characteristic when exposed to urine; a measurement circuit arranged to perform a measurement on a urine bladder of a wearer to determine at least one parameter which varies with a fill level of the urine bladder; a sensor arranged to determine an orientation and/or a movement of the sensor; and a processing circuit arranged to: determine whether the urine sensitive circuit has been exposed to urine; estimate an amount of urine released on to the urine sensitive circuit; and in response to determining that the urine sensitive circuit has been exposed to urine, record data representing said at least one parameter determined by the measurement circuit, an estimated movement and/or posture of the wearer based on an orientation and/or a movement determined by the sensor, and an estimate of the amount of urine released on to the urine sensitive circuit. | 1. A system for monitoring incontinence for a user comprising: a urine sensitive circuit configured to present a changed electrical characteristic when exposed to urine, wherein the changed electrical characteristic of the urine sensitive circuit is indicative of an amount of urine released on to the urine sensitive circuit;a measurement circuit configured to perform a measurement on a urine bladder of the user to determine at least one parameter which varies with a fill level of the urine bladder;a sensor configured to be attached on the user and determine an orientation and/or a movement of the sensor; anda processing circuit configured to: determine whether the urine sensitive circuit has been exposed to urine,estimate the amount of urine released on to the urine sensitive circuit, andin response to determining that the urine sensitive circuit has been exposed to urine, record data representing said at least one parameter determined by the measurement circuit, an estimated movement and/or posture of the user based on an orientation and/or a movement determined by the sensor, and an estimate of the amount of urine released on to the urine sensitive circuit. 2. The system according to claim 1, wherein the changed electrical characteristic circuit includes one or more selected from the group comprising: a changed resistance, a changed capacitance, a changed inductance, a changed impedance, a changed resonance frequency, a changed voltage generated by the urine sensitive circuit, a changed current generated by the urine sensitive circuit, a changed resonance frequency of the urine sensitive circuit. 3. The system according to claim 2, wherein the processing circuit is configured to estimate the amount of urine released on to the urine sensitive circuit by determining a size of the change. 4. The system according to claim 1, wherein the processing circuit is configured to provide a signal which is indicative of a urine bladder fill level based on the at least one determined parameter. 5. The system according to claim 1, wherein the measurement circuit is configured to be positioned at or near a bladder region of the user and wherein the sensor is configured in a same unit as the measurement circuit. 6. The system according to claim 1, further comprising an underwear or undergarment wherein the urine sensitive circuit is configured to be positioned in connection to the underwear or undergarment. 7. The system according to claim 6, wherein the measurement circuit is configured to be secured to an edge of the underwear or undergarment. 8. The system according to claim 1, wherein the processing circuit is further configured to, subsequent to detecting that the urine sensitive circuit has been exposed to urine, perform a comparison between: the at least one parameter represented by the data recorded in response to said detecting, and a parameter determined by the measurement circuit at an instant subsequent to the detection of the urination event... | STICHTING IMEC NEDERLAND | A61B005/20 A61B005/11 A61B005/07 A61B005/00 A61B005/1473 A61B005/145
US10716506 | B2 | 2020-07-21 | Prosthetic implant and method for the production of such an implant | 2017-07-20 | A prosthetic implant includes a cavity that opens at an outer face of the prosthetic implant, the cavity forming a housing for receiving: a printed circuit including a radio tag; and a cover which closes the cavity at least partially when positioned on the implant in a so-called closed position, the radio tag and printed circuit being removably housed in the cavity. | 1. A prosthetic implant comprising: an upper surface, a lower bone-engaging surface, and a perimeter between said upper and lower surfaces,a cavity formed in said lower bone-engaging surface and opening onto an outer face of the perimeter of said prosthetic implant, said cavity forming a housing receiving: a printed circuit board including a radio-frequency identification (RFID) tag, said RFID tag comprising an antenna connected to an electronic chip containing an identifier and additional data;a force sensor arranged on the printed circuit board and connected to the printed circuit board; anda cover slidably received in said cavity, the cover closing at least in part said cavity when it is positioned on said implant in a position called closed position;wherein the RFID tag, force sensor, and the printed circuit board are removably housed in the cavity,wherein the cover has a window arranged on a surface separating the cavity from the outside of the prosthetic implant, said window being arranged in order to receive a deformable membrane and to transmit a force exerted on said membrane to said force sensor,wherein the cover comprises a shape and size complementary to shape and size of the cavity,wherein the cover is flush-mounted to the lower bone-engaging surface of the prosthetic implant when in said closed position, in which an exposed surface of said cover forms part of said lower bone-engaging surface,wherein the prosthetic implant is a tibial plate. 2. The prosthetic implant according to claim 1, comprising a capacitive proximity sensor arranged on the printed circuit board and connected to the printed circuit board. 3. The prosthetic implant according to claim 1, in which the window is produced in a single piece with the cover, a thickness of the cover being reduced at the level of said window. 4. The prosthetic implant according to claim 1, in which the force sensor is of the piezoresistive type and a force-transmission element is arranged between the membrane and said force sensor, the transmission element being arranged in order to transmit a force exerted on said membrane to the force sensor. 5. The prosthetic implant according to claim 1, in which the force sensor is of the capacitive type and a reaction element is arranged between the membrane and said force sensor, said reaction element being firmly fastened to the membrane on a face of the membrane turned towards the cavity, the reaction element being arranged in order to be displaced when a force is exerted on said membrane. 6. The prosthetic implant according to claim 1, in which the force sensor is of the piezoelectric type and a flexible piezoelectric body is arranged between the membrane and said force sensor, said flexible body being firmly fastened to the membrane on a face of the membrane turned towards the cavity, said flexible body being arranged in order to be deformed when a force is exerted on said membrane. 7. The prosthetic implant according to claim 1, in which... | CENTRE NATIONAL DE LA RECHERCHE SCIENTIFIQUE UNIVERSITE DE MONTPELLIER BONETAG | A61B005/00 A61F002/30 A61F002/38 A61B005/07 A61F002/46 A61B005/103
US10716487 | B2 | 2020-07-21 | Sub-dermally implanted electroencephalogram sensor | 2017-12-27 | A method for obtaining an electroencephalogram (EEG) of a user is disclosed. A reference sensor is attached to the user by connecting a first component of the reference sensor to a second component of the reference sensor, at least a portion of the first component being sub-dermally implanted on or adjacent to a mastoid process of the user. At least one active sensor is attached to the user. A first signal is detected from the reference sensor simultaneously as a second signal is detected from the at least one active sensor. The EEG is obtained based on the first signal and the second signal. | 1. A method for obtaining an electroencephalogram (EEG) of a user, the method comprising: attaching a reference sensor to the user by connecting a first component of the reference sensor to a second component of the reference sensor, wherein the second component comprises: an electrode sub-dermally implanted on or adjacent to a mastoid process of the user,a shaft extending outwards from a surface of the electrode and extending through the user's skin, anda semi-circular ring attached to a proximal end of the shaft, andwherein the first component comprises: an arc-shaped body comprising an electrical lead coupled thereto, anda spring loaded arm pivotably coupled at one end of the arc-shaped body, the spring loaded arm configured to pivot between a first position and a second position;attaching at least one active sensor to the user by passing a portion of the arc-shaped body through an aperture of the semi-circular ring;simultaneously detecting a first signal from the reference sensor and a second signal from the at least one active sensor; andobtaining the EEG based on the first signal and the second signal. 2. The method of claim 1, wherein the lead is electrically connected to the sensor to a bioamplifier. 3. The method of claim 1, wherein the first component is electrically connected to a wireless transceiver in communication with a bioamplifier. 4. The method of claim 1, wherein the electrode comprises titanium or gold. 5. The method of claim 1, wherein the second component is anchored to a skull bone of the user's mastoid process. 6. The method of claim 1, further comprising attaching a ground sensor to the user and detecting signals from the ground sensor simultaneously to detecting signals from the reference sensor and the at least one active sensor. 7. The method of claim 1, wherein the reference sensor forms an electrical path from the electrode to arc-shaped body to transmit EEG signals along the electrical lead to a bioamplifier during use. 8. An electroencephalogram (EEG) system, comprising: a bioamplifier; anda sensor having a first component and a second component, the sensor being in communication with the bioamplifier when the first component is coupled with the second component, wherein: the first component comprises: an electrode shaped for implantation under a user's skin,a shaft extending outwards from a surface of the electrode, the shaft sized to extend through the user's skin, anda semi-circular ring attached to a proximal end of the shaft; andthe second component comprises: an arc-shaped body comprising an electrical lead coupled thereto, anda spring loaded arm pivotably coupled at one end of the arc-shaped body, the spring loaded arm configured to pivot between a first position and a second position; andthe sensor forms an electrical path from the electrode to arc-shaped body to transmit EEG signals along the electrical lead to the bioamplifier during use of the EEG system. 9. The EEG system of claim 8, wherein th... | X Development LLC | A61B005/0478 A61B005/00 A61B005/04 A61B005/0484
US10716492 | B2 | 2020-07-21 | Quadruped lameness detection | 2018-03-30 | Systems and methods for detecting lameness in a limb of an animal are described herein. In an embodiment, a lameness detection system comprises a plurality of wearable footwear elements. Each of the footwear elements comprises a housing that is configured to fit snugly on a paw of the animal during ordinary gait. Each of the footwear elements comprise a sensor or matrix of sensors that is configured to detect and generate data representative of pressure or force exerted by the paw on a ground surface. Each footwear element additionally comprises a wireless transceiver that is configured for short-distance communication of the data captured by the sensors. The system further comprises a computing device with a wireless transceiver configured to communicate with the transceivers of the footwear elements and instructions which, when executed by the computing device, cause the computing device to analyze lameness in a limb of the animal based on data received from the sensors. | 1. A lameness measurement system for a quadruped animal, comprising: a plurality of wearable footwear elements, each of the footwear elements comprising a housing that is configured to fit snugly on a paw of the animal, and each of the footwear elements comprising a sensor that is configured to detect and generate first data representative of pressure or force exerted by the paw upon a ground surface, and a first wireless transceiver that is configured for short-distance communication of the first data captured by the sensor; anda computing device comprising a second wireless transceiver that is configured to communicate wirelessly with the first wireless transceiver, and comprising a non-transitory computer-readable data storage medium storing sequences of instructions of an application program which when executed by the computing device cause: continuously obtaining the first data from each sensor of the footwear elements; continuously transforming the first data to produce a real-time data set indicating pressure or force of a paw of the animal relative to one or more other paws of the animal; determining, in real-time from the real-time dataset indicating the pressure or force of the paw of the animal relative to the one or more other paws of the animal, that the animal is being affected by lameness; and causing display of an analysis of lameness of the animal on a client computing device. 2. The system of claim 1, further comprising a wireless data relay device that is configured to receive the first data from the first wireless transceiver of each of the footwear elements via a first wireless networking protocol, to format the first data into messages of a second wireless networking protocol, and to transmit the messages using the second wireless networking protocol to the computing device, wherein the second wireless transceiver communicates wirelessly with the first wireless transceiver through the wireless data relay. 3. The system of claim 2, wherein the wireless data relay device is affixed to a collar of the animal. 4. The system of claim 2, further comprising one or more proximity sensors that are programmed or configured to detect proximity of a food and/or water bowl to the wireless data relay. 5. The system of claim 4, wherein the food and/or water bowl comprises one or more sensors configured to measure food and/or water in the food and/or water bowl prior to detection of the proximity of the food and/or water bowl to the wireless data relay and after cessation of the detection of the proximity of the food and/or water bowl to the wireless data relay. 6. A data processing method comprising: continuously receiving, at a wireless data relay device, from a first wireless transceiver communicatively coupled to first one or more sensors attached to a footwear element worn on a first paw of an animal, first data representative of measurements of pressure or force exerted by the first paw upon a ground surface measured by the first... | Bay West Veterinary Surgery, Inc. | A61B005/103 G01L005/00 A01K029/00 A61B005/00 A61D099/00 A61B005/11 H04L029/06 H04W004/80 H04W084/12
US10712722 | B2 | 2020-07-14 | Systems and articles for enhancing wellness associated with habitable environments | 2015-02-25 | Environmental characteristics of habitable environments (e.g., hotel or motel rooms, spas, resorts, cruise boat cabins, offices, hospitals and/or homes, apartments or residences) are controlled to eliminate, reduce or ameliorate adverse or harmful aspects and introduce, increase or enhance beneficial aspects in order to improve a “wellness” or sense of “wellbeing” provided via the environments. Control of intensity and wavelength distribution of passive and active Illumination addresses various issues, symptoms or syndromes, for instance to maintain a circadian rhythm or cycle, adjust for “jet lag” or season affective disorder, etc. . Air quality and attributes are controlled. Scent(s) may be dispersed. Noise is reduced and sounds (e.g., masking, music, natural) may be provided. Environmental and biometric feedback is provided. Experimentation and machine learning are used to improve health outcomes and wellness standards. | 1. An environmental control system, comprising: a plurality of active subsystems operable to effect a condition in a habitable space;a first number of sensors that collect habitable space wellness data that includes, for at least the habitable space, at least one wellness parameter indicative of a wellness associated with the respective habitable space; anda control subsystem including at least one processor and at least one nontransitory processor-readable medium communicatively coupled to the at least one processor and which stores at least one of instructions or data including at least collected personal wellness data, the at least one of instructions or data executable by the at least one processor to control the active subsystems, the first number of sensors that collect habitable space wellness data communicatively coupled to the at least one processor, the control subsystem which: identifies at least one pattern in the collected personal wellness data via at least one machine learning technique, identifies at least one pattern in the collected habitable space wellness data via at least one machine learning technique, determines at least one adjustment to at least one operational parameter based on the identified at least one pattern in the collected personal wellness data and the collected habitable space wellness data, dynamically adjusts at least one operational parameter of at least one of the active subsystems based on both the collected habitable space wellness data and on collected personal wellness data, and adjusts a baseline set of operational parameters of the one or more active subsystems based on at least one of the collected habitable space wellness data or the collected personal wellness data. 2. The environmental control system of claim 1, further comprising: a second number of sensors that collect the collected personal wellness data from at least one individual. 3. The environmental control system of claim 2 wherein the second number of sensors includes one or more biometric sensors, and the at least one processor evaluates the collected personal wellness data from time-to-time. 4. The environmental control system of claim 3 wherein the one or more biometric sensors includes at least one of a temperature sensor operable to detect a temperature of the at least one individual, a scale operable to detect a weight of the at least one individual, a heart rate sensor operable to detect a heart rate of the at least one individual, a blood oxygen sensor operable to detect a level of blood oxygen of the at least one individual, a respiratory cycle sensor operable to detect at least one characteristic of a respiratory cycle of the at least one individual, or an electroencephalography (EEG) sensor operable to detect at least one brainwave pattern of the at least one individual. 5. The environmental control system of claim 3 wherein the one or more biometric sensors has at least one motion sensor operable to detect a motion of the... | Delos Living LLC | A61B005/024 F24F011/30 G06F019/00 G06Q050/10 A61B005/08 A61B005/145 G05B015/02 G06Q050/12 G06Q050/16 G06Q050/22 G16H010/60 G16H040/63 A61B005/0205 A61B005/0476 F24F110/00 F24F120/20 G05B019/048
US10716467 | B2 | 2020-07-21 | Image processing apparatus and image processing method | 2017-12-01 | An image processing apparatus includes a storage unit configured to store a plurality of names associated with depth ranges, a selection unit configured to select one name from the plurality of names, an acquisition unit configured to acquire a plurality of tomographic images of a subject's eye, a generation unit configured to generate a projection image on which a plural pieces of motion contrast data generated based on the plurality of tomographic images is projected in a depth range associated with the selected name, and a control unit configured to display the generated projection image on a display unit. | 1. An image processing apparatus comprising: one or more processors; andat least one memory storing executable instructions, which when executed by the one or more processors, cause the image processing apparatus to: display, on a display, a definition screen for defining a depth range, wherein the definition screen includes a name input portion for inputting a name, a boundary shape designation portion for designating two boundary shapes, a depth position designation portion for designating depth positions as offset values with respect to the boundary shapes, and a screen element for instructing addition of the name input into the name input portion;store, in the at least one memory storing a first name associated with a first depth range, a second name in association with a second depth range, the second name having been input into the name input portion, the second depth range based on the two boundary shapes having been designated in the boundary shape designation portion and the depth positions having been designated as offset values in the depth position designation portion, the first name having been previously stored in the at least one memory, the second name being newly added in the at least one memory based on an instruction for adding the second name, wherein the instruction for adding the second name is based on an input that indicates selection of the screen element included on the definition screen;select one name from a plurality of names including the first name and the second name, wherein each of the plurality of names is associated with a respective depth range;acquire a plurality of tomographic images of a constituent of a subject;generate a projection image on which a plurality of tomographic images or a plurality of pieces of motion contrast data generated based on the plurality of tomographic images is projected in the depth range associated with the selected name; anddisplay the generated projection image on the display. 2. The image processing apparatus according to claim 1, wherein the executable instructions, when executed by the one or more processors, further cause the image processing apparatus to: analyze the tomographic image to detect a plurality of layer boundary shapes,wherein the depth range is defined based on the layer boundary shapes detected by the analyzing. 3. The image processing apparatus according to claim 2, wherein the depth range is further defined based on a depth position as an offset value with respect to the layer boundary shape detected by the analyzing. 4. The image processing apparatus according to claim 3, wherein the depth range is defined based on at least one layer boundary shape and at least one depth position. 5. The image processing apparatus according to claim 1, wherein the name includes any of a superficial capillary plexus, a deep capillary plexus, an outer retina, a radial peripapillary capillaries, and a choriocapillaris, andwherein the each of two boundary shapes is selecte... | Canon Kabushiki Kaisha | A61B003/10 A61B003/00 A61B003/12 A61B005/02 G06T007/00 G06T007/60 A61B005/06
US10710594 | B2 | 2020-07-14 | Occupant-status prediction system | 2017-12-27 | An occupant support adapted for use in a vehicle includes a seat bottom, a seat back, and a sensory system. The seat bottom is coupled to a floor of the vehicle. The seat back extends upwardly away from the seat bottom. The sensor system is configured to monitor for fatigue of an occupant of the occupant support. | 1. An occupant support system for use in a vehicle, the occupant support comprising a sensor system configured to obtain biometrics input associated with physiological and behavioral characteristics of an occupant of the occupant support system, sleep input associated with sleep characteristics of the occupant, and vehicle input associated with characteristics of the vehicle and environment surrounding the vehicle, anda control system configured to receive the biometrics input, sleep input, and vehicle input, determine a biometrics score based on the biometrics input, determine a sleep score based on the sleep input, determine a task score based on the vehicle input, determine occupant fatigue based on the biometrics score, the sleep score, and the task score, and activate a vehicle system based on the occupant fatigue,wherein the control system is further configured to identify the occupant based on at least one of input from the occupant and the biometrics input and to associate the biometrics input, sleep input, and task input with a unique occupant data profile for the occupant,wherein the control system is further configured to receive supplemental biometrics input after activating the vehicle system, determine a supplemental biometrics score based on the supplemental biometrics input, and determine occupant fatigue based on the supplemental biometrics score, the sleep score, and the task score. 2. The occupant support system of claim 1, wherein the sleep input includes a number of hours that the occupant has been without sleep, a sleep history, and circadian information. 3. The occupant support system of claim 2, wherein the vehicle input includes traffic information, a level of carbon dioxide in a cabin of the vehicle, a light level in the cabin of the vehicle, a noise level in the cabin, and vibration of the occupant support. 4. The occupant support system of claim 3, wherein the biometrics input includes heart rate of the occupant, a respiration rate of the occupant, an eye blink rate of the occupant, a head position of the occupant, and a posture of the occupant. 5. The occupant support system of claim 1, wherein the control system is further configured to compare the biometrics score and the supplemental biometrics score and associate changes to the biometrics score in the unique occupant data profile. 6. The occupant support system of claim 1, wherein activating the vehicle system includes generating and displaying a suggested travel route to the occupant. 7. The occupant support system of claim 1, wherein activating the vehicle system includes changing the vehicle from manual operation to autonomous operation. 8. The occupant support system of claim 1, wherein the biometrics input includes information indicative of a respiration rate of the occupant included in signals received from a piezoelectric sensor and an electrode included in the sensor system, a posture of the occupant included in signals received from a load cell and... | Faurecia Automotive Seating, LLC | B60W040/08 G01C021/34 G05D001/00 B60W050/16 B60N002/90 G08B021/06 B60N002/00 B60N002/20 A61B005/11 A61B005/18 A61B005/16 A61B005/00 A61B005/0205 A61B005/024 A61B005/08
US10716485 | B2 | 2020-07-21 | Deep brain source imaging with M/EEG and anatomical MRI | 2015-11-09 | A method for non-invasively resolving electrophysiological activity in sub-cortical structures located deep in the brain by comparing amplitude-insensitive M/EEG field patterns arising from activity in subcortical and cortical sources under physiologically relevant sparse constraints is disclosed. The method includes a sparse inverse solution for M/EEG subcortical source modeling. Specifically, the method employs a subspace-pursuit algorithm rooted in compressive sampling theory, performs a hierarchical search for sparse subcortical and cortical sources underlying the measurement, and estimates millisecond-scale currents in these sources to explain the data. The method can be used to recover thalamic and brainstem contributions to non-invasive M/EEG data, and to enable non-invasive study of fast timescale dynamical and network phenomena involving widespread regions across the human brain. | 1. A method for non-invasively characterizing electrophysiological activity within deep brain regions based on data measurements acquired during a functional task of interest, the method comprising: acquiring measured data representative of brain activity of a subject during a functional task of interest,acquiring magnetic resonance images of brain anatomy of the subject,constructing superficial (cortical) source spaces and deep (subcortical) source spaces for regions of the brain from the magnetic resonance images, the cortical source spaces including a first set of cortical source spaces having patches with a first size, a second set of cortical source spaces having patches with a second size, the second size being smaller than the first size;determining a first relationship between the source spaces and the measured data;determining a minimum norm source current estimate (MNE) based on the source spaces, the first relationship, and the measured data, the MNE including a first subset of source spaces;refining the MNE a first time by using and determining a second relationship between the measured data, and cortical source spaces from the second set of cortical source spaces that at least partially overlap with the first subset of source spaces, the refined MNE including a second subset of source spaces, the second subset of source spaces including a cortical source space;further refining the MNE by using and determining a third relationship between the second subset of source spaces, the subcortical source spaces, and the measured data, the further refined MNE including a superficial source space and a deep source space that underlies the measured data. 2. The method of claim 1, wherein the measured data is obtained by at least one of an electroencephalography (EEG) and an magnetoencephalography (MEG). 3. The method of claim 1, wherein constructing source spaces from the magnetic resonance images of brain anatomy of the subject includes positioning dipoles within a given source space, and wherein at least one source space corresponds to at least one of: a neocortex, a hippocampus, a gray-white matter interface, and a subcortical volume. 4. The method of claim 1, wherein each cortical source space is a surface patch having dipoles, and wherein the subcortical source spaces includevolume subdivisions and surface patches having dipoles. 5. The method of claim 4, wherein the surface patches and volume subdivisions are sized to homogenize current strengths. 6. The method of claim 4, wherein the surface patches and the volume subdivisions having high current densities are assigned finer subdivisions, and wherein the surface patches and the volume subdivisions having low current densities are assigned larger subdivisions to homogenize current strengths thereby enabling comparable consideration of the surface patches and volume subdivisions in an inverse solution. 7. The method of claim 1, wherein determining the second relationship includes sear... | The General Hospital Corporation Massachusetts Institute of Technology | A61B005/0436 A61B005/0476 A61B005/055 A61B005/00
US10722179 | B2 | 2020-07-28 | Residual-based monitoring of human health | 2017-07-25 | Improved human health monitoring is provided in the context of sensor measurements of typical vital signs and other biological parameters, by a system and method using an empirical model of the parameters and disposed to estimate values of the parameters in response to actual measurements. Residuals resulting from the difference between the estimates and actual measurements are analyzed for robust indications of incipient health issues. Residual analysis is both more robust and more sensitive than conventional univariate range checking on vital signs. | 1. A computer-automated monitoring system for early detection of incipient health problems in a monitored patient, comprising: a computer-accessible memory for storing a set of exemplary data observations of multiple vital signs obtained from said monitored patient comprising a personalized model of said patient;a data feed for receiving new data observations of said multiple vital signs from said monitored patient;a computer configured by a kernel-based modeling software module to generate an estimate of the value of at least one of said multiple vital signs, as a weighted combination of a subset of said exemplary data observations where said weights are determined from a kernel function comparison of said subset of exemplary data observations and a said received observation;said computer being further configured by a localization software module to select said subset of exemplary data observations based on at least some of the vital signs in said received observation;said computer being further configured by a residual analysis software module to compare said estimate to the corresponding measured value of the at least one of said multiple vital signs in said received data observation to generate a residual; andsaid computer being further configured by an alerting module having a rules system configurable to provide a test, wherein the test applies rules to said generated residual of each of a plurality of said multiple vital signs in combination to selectively trigger an alert, the alert being triggered based upon results of the test, and present said alert in a web based interface usable by a medical practitioner for monitoring said patient for incipient health problems. 2. A system according to claim 1, wherein said kernel-based modeling software module implements a similarity-based model. 3. A computer automated monitoring system according to claim 1, wherein said test of a said generated residual of each of a plurality of said multiple vital signs selected from the set of: said residual exceeds a threshold: said residual exceeds said threshold for a preselected number of data observations in a preselected window: said residual exceeds a running statistic of a time series of said residuals: and a sequential probability ratio test. 4. A computer automated monitoring system according to claim 3, wherein said multiple vital signs includes heart rate. 5. A computer automated monitoring system according to claim 3, wherein said multiple vital signs includes respiration rate. 6. A computer automated monitoring system according to claim 3, wherein said multiple vital signs includes heart rate variability. 7. A computer automated monitoring system according to claim 3, wherein said multiple vital signs includes a measure of a time lag between a QRS event and a subsequent blood pressure event. 8. A computer automated monitoring system according to claim 3, wherein said multiple vital signs includes a measure of a time lag between a respirator... | PhysIQ, Inc. | A61B005/00 A61B005/0205 G06K009/62 G16H050/70 G16H050/20 G16H050/50 G06F009/54 A61B005/0402 A61B005/1455 G16B005/00 G16B099/00 A61B005/021 A61B005/0245 A61B005/0456 A61B005/08 A61B005/11 A61B005/145 G16H050/30
US10708574 | B2 | 2020-07-07 | Three dimensional imaging apparatus with color sensor | 2018-04-13 | An imaging apparatus includes a first and second light source, focusing optics, a probe, a detector, an optical transmission medium and a color sensor. The first light source is to generate light beams that travel through the focusing optics along an optical path to the probe. The probe directs the light beams toward a three dimensional object to be imaged. The detector detects returning light beams that are reflected off of the three dimensional object and directed back through the probe and the focusing optics. The second light source is to generate multi-chromatic light. The optical transmission medium is outside of the optical path and is to receive a ray of the multi-chromatic light reflected off of a spot on the three dimensional object and through the probe. The color sensor is to receive the ray from the optical transmission medium and determine a color of the spot on the three dimensional object. | 1. An imaging apparatus comprising: a first light source to generate first light that is to illuminate a three dimensional object;a probe to receive a plurality of rays of the first light that are reflected off of the three dimensional object and to direct the plurality of rays along an optical path associated with an imaging axis;a detector to detect the plurality of rays, wherein measurements of the plurality of rays indicate a shape of the three dimensional object;a second light source to generate multi-chromatic light that is to illuminate the three dimensional object;an optical transmission medium, outside of the optical path, to receive one or more rays of the multi-chromatic light reflected off of a spot on the three dimensional object and through the probe; anda color sensor, optically coupled to a first end of the optical transmission medium, to receive the one or more rays of the multi-chromatic light and determine a color of the spot on the three dimensional object based on an analysis of the one or more rays. 2. The imaging apparatus of claim 1, wherein the one or more rays of the multi-chromatic light have an oblique angle to an imaging axis associated with the optical path, and wherein the optical transmission medium is oriented at the oblique angle to the imaging axis. 3. The imaging apparatus of claim 1, wherein the one or more rays of the multi-chromatic light are parallel to the optical path, the imaging apparatus further comprising: a beam splitter disposed along the optical path, wherein the beam splitter directs the one or more rays of the multi-chromatic light into the optical transmission medium. 4. The imaging apparatus of claim 3, wherein the beam splitter passes a first portion of the one or more rays of the multi-chromatic light to the detector and directs a second portion of the one or more rays of the multi-chromatic light to the optical transmission medium, where the second portion is less than the first portion. 5. The imaging apparatus of claim 1, wherein the imaging apparatus is to: at a first time, activate the first light source, determine a first shape of a first portion of the three dimensional object, and determine a first relative position of the probe to the three dimensional object;at a second time, deactivate the first light source, activate the second light source, and determine the color of the spot on the three dimensional object;at a third time, deactivate the second light source, activate the first light source, determine a second shape of a second portion of the three dimensional object, and determine a second relative position of the probe to the three dimensional object; anddetermine a position of the spot on the three dimensional object based on an interpolation between the first relative position of the probe to the three dimensional object and the second relative position of the probe to the three dimensional object. 6. The imaging apparatus of claim 1, wherein the imaging apparatus is a c... | Align Technology, Inc. | H04N013/257 G02B027/10 G02B023/24 H04N005/235 H04N005/225 A61B001/06 A61B001/24 A61B001/00 A61B005/00 A61C009/00 A61B001/247 A61B001/07 G01J003/46 A61B001/04 H04N013/00 A61C001/08 H04N009/04
US10726116 | B2 | 2020-07-28 | Biometrics system, biologic information storage, and portable device | 2019-01-07 | In a biometrics system for a building entrance unlocking or a bank account authentication, reference information registered under administration by the system is transmitted to a room or mobile-phone for private storage, with the original reference information deleted from the system. Biologic information gotten upon authentication is transmitted through wireless system to the room or mobile-phone for comparison with the reference, the result being returned to the system. Or, the reference is tentatively sent back to the system for comparison with the gotten biologic information. The biologic information sent to mobile-phone also includes health control information for storage and display. Mobile-phone also can receive blood pressure information at a waiting lounge of medical institution though wireless local communication even if the main power shut down. The communication between the biometrics system and the mobile-phone is encrypted. The system includes sensor unit and protection unit, the abnormality thereof being separately checked. | 1. A structural sphere of activity for persons having a main entrance commonly utilized by the persons to go into an area in the structural sphere, and an authentication system for authorizing by authentication of one of the persons who is to access the main entrance for reaching one of sub areas within the area, the structural sphere of activity for the persons comprising: an information sensor, which is provided at the main entrance, arranged to sense information of the one of the persons to be authorized by the authentication;an administrator, which is provided in the authentication system, arranged to administrate the authentication system;a transmitter, which is provided in the authentication system, arranged to transmit in advance information of success in the authentication at the main entrance to the one of the sub areas to which the one of the persons is going to reach after going through the main entrance; anda lock control system for the main entrance commonly utilized by the persons, the lock control system being arranged to unlock the main entrance in accordance with the authentication, wherein the one of the sub areas is not unlocked by the information of success in the authentication at the main entrance in advance, but is to be unlocked in response to the one of the persons who actually reaches the one of the sub areas,wherein the transmitter is arranged not to transmit the information of the success in the authentication of the one of the persons at the main entrance to another one of the sub areas other than the one of the sub areas to which the one of the persons is going to reach, andwherein the one of the sub area is unique destination of the one of the persons and receives the information of success in the authentication at the main entrance before the one of the persons comes to the unique destination. 2. The structural sphere of activity for persons according to claim 1, wherein the transmitter includes a wireless local communication apparatus. 3. The structural sphere of activity for persons according to claim 1, wherein the information sensor includes a biologic information sensor arranged to sense biologic information of the one of the persons to be authorized by the authentication, and wherein the administrator is arranged to administrate the authentication system by means of a comparison between the biologic information gotten from the person by the biologic information sensor and preliminarily registered reference biologic information of the person. 4. The structural sphere of activity for persons according to claim 1, wherein the transmitter is arranged to inform a person in the one of sub areas of arrival of the person at the main entrance in advance by means of the transmission of the information of success in the authentication of the person at the main entrance. 5. The structural sphere of activity for persons according to claim 1, wherein the transmitter is arranged to control an equipment in the one of sub... | NL GIKEN INCORPORATED | G06F021/32 G06K009/00 G16H010/60 G07C009/25 G06F019/00 G06Q020/40 A61B005/117 A61B005/021 A61B005/1455 A61B005/00 G16H050/20
US10709935 | B2 | 2020-07-14 | System and method for measuring and adjusting physical resistance for athletic activities and fitness equipment | 2018-03-30 | Methods, apparatus, systems and articles of manufacture for measuring and adjusting physical resistance for athletic activities and fitness equipment are disclosed. An example system includes a user-wearable sensor to measure a first level of physical effort during a first athletic activity using fitness equipment and to generate sensor data indicative of the measured first level of physical effort. The example system also includes a controller to receive the sensor data to perform a comparison of the sensor data and data from second athletic activity indicative of a second level of physical effort, and automatically generate a control signal to adjust a setting of the fitness equipment based on the comparison. | 1. A system comprising: a user-wearable sensor including a housing having a wearable form factor, the user-wearable sensor to measure a first level of physical effort during a first athletic activity using fitness equipment and to generate sensor data indicative of the first level of physical effort; anda controller to: determine a first force-to-motion ratio based on the sensor data;conduct effort normalization based on the first force-to-motion ratio and a second force-to-motion ratio based on data from a second athletic activity indicative of a second level of physical effort; andautomatically generate a control signal to adjust a setting of the fitness equipment based on the effort normalization. 2. The system of claim 1, wherein the user-wearable sensor is to measure the first level of physical effort through a force adjacent the user. 3. The system of claim 2, wherein the force is air resistance. 4. The system of claim 2, wherein the force is a weight. 5. The system of claim 1, wherein the fitness equipment is first fitness equipment and the second athletic activity includes use of second fitness equipment different from the first equipment. 6. The system of claim 1, wherein the first athletic activity is at a first time and the second athletic activity includes use of the fitness equipment at a second time, the second time prior to the first time. 7. The system of claim 1, wherein the controller is to adjust the setting of the fitness equipment to decrease a physical resistance level when the effort normalization indicates a decrease level of physical effort. 8. The system of claim 1, wherein the user-wearable sensor is to measure the first level of physical effort through a physical signal generated by the user. 9. The system of claim 1, wherein the user-wearable sensor is to measure the first level of physical effort through a time of use of the fitness equipment. 10. The system of claim 1, wherein the controller is programmed to automatically generate the control signal to adjust a setting of the fitness equipment during use of the fitness equipment. 11. A system comprising: a first user-wearable sensor including a first housing having a wearable form factor, the first user-wearable sensor to measure a physical signal generated by the user during a first athletic activity using fitness equipment;a second user-wearable sensor including a second housing having a wearable form factor, the second user-wearable sensor to measure a force adjacent the user during the first athletic activity; anda controller to: determine a first level of physical effort exerted by the user based on the physical signal generated by the user and the force adjacent the user;determine a first force-to-motion ratio based on the first level of physical effort;determine a second force-to-motion ratio based on a second level of physical effort from a second athletic activity;conduct effort normalization based on the first force-to-motion ratio and the second... | Intel Corporation | A63B024/00 A63B069/16 A61B005/11 G01S019/19 G06F001/16 A63B069/06 G06F003/01 G16H020/30 A63B022/06 A61B005/01 A61B005/024 A63B071/06 A61B005/00 A63B022/00 A63B022/02
US10722686 | B2 | 2020-07-28 | Method and system of utilizing ECG signal for central venous catheter tip positioning | 2018-06-18 | Disclosed herein are a method and a medical system for utilizing of an intravascular ECG signal for central venous catheter placement. The medical system is capable of detecting the position of a catheter tip and assessing its location relative to the cavoatrial junction. The detection and assessment are performed by a multiscale analysis of the complexity of the intravascular signal data points. | 1. A medical system for placement of a central venous catheter comprising: a placement unit including an electrode usable to acquire intravascular ECG (ivECG) signal data;a multiscale complexity analysis module including software programmed for processing the ivECG signal data acquired by the electrode, wherein the processing uses Multiscale Complexity Analysis (MSCA) of dynamics of the ivECG signal data using a Complementary Probability Cumulative Distributive Function; anda visualization device configured to provide a visual indication of a location of the central venous catheter relative to a target location. 2. The medical system according to claim 1, wherein the target location is a cavoatrial junction of a body. 3. The medical system according to claim 1, wherein the software of the multiscale complexity analysis module is programmed to determine at least one proximity index (PxI) value based on the ivECG signal data acquired by the electrode. 4. The medical system according to claim 3, wherein the visualization device is configured to provide a user with the visual indication of the location of the central venous catheter relative to the target location based on a comparison of the at least one PxI value with a proximity threshold. 5. The medical system according to claim 1, wherein the electrode is positioned at a tip of the central venous catheter. 6. The medical system according to claim 5, wherein the electrode is attached to the central venous catheter. 7. The medical system according to claim 1, further comprising an audio device configured to provide an audio indication of the location of the central venous catheter relative to the target location. 8. A medical system for placement of an endovascular device comprising: an electrode usable to acquire intravascular ECG (ivECG) signal data;a multiscale complexity analysis module including software, the software comprising instructions to determine at least one proximity index (PxI) value calculated as a density of Tail Distribution Function (TDF) from the ivECG signal data acquired by the electrode; anda device configured to provide one or both of a visual indication and an audio indication of one or both of the at least one PxI value and a location of the endovascular device relative to a target location. 9. The medical system according to claim 8, wherein the endovascular device is a central venous catheter and the target location is a cavoatrial junction of a body. 10. The medical system according to claim 8, wherein the indication of one or both of the at least one PxI value and the location of the endovascular device relative to the target location is based on a comparison of the at least one PxI value with a proximity threshold. 11. A method of positioning an endovascular device in a body, comprising: advancing the endovascular device into the body with an electrode positioned near a tip of the endovascular device;using the electrode to acquire intravascular ECG (ivECG)... | C. R. Bard, Inc. | A61B005/00 A61M025/01 A61B005/06 A61B034/20 A61B005/042 G06F017/18
US10722301 | B2 | 2020-07-28 | Systems and methods for lesion assessment | 2015-11-03 | Ablation visualization and monitoring systems and methods are provided. In some embodiments, such methods comprise applying ablation energy to a tissue to form a lesion in the tissue, illuminating the tissue with a light to excite NADH in the tissue, wherein the tissue is illuminated in a radial direction, an axial direction, or both, monitoring a level of NADH fluorescence in the illuminated tissue to determine when the level of NADH fluorescence decreases from a base level in the beginning of the ablating to a predetermined lower level, and stopping ablation of the tissue when the level of NADH fluorescence reaches the predetermined lower level. | 1. A method for monitoring ablation of cardiac tissue of endocardium comprising: applying ablation energy to a cardiac tissue to form a lesion in the tissue;illuminating the cardiac tissue to excite nicotinamide adenine dinucleotide hydrogen (NADH) in the cardiac tissue and determining a base level of NADH fluorescence in the illuminated cardiac tissue;monitoring a level of NADH fluorescence in the illuminated cardiac tissue during the application of ablation energy to the illuminated cardiac tissue;reducing the NADH fluorescence in the illuminated cardiac tissue to between 60% and 80% of the base level and achieving a steady state NADH fluorescence; andupon achieving the steady state NADH fluorescence for between 5 and 10 seconds, stopping ablation of the illuminated cardiac tissue to limit an amount of ablation energy delivered to the cardiac tissue. 2. The method of claim 1 wherein the cardiac tissue is illuminated with light having a wavelength between about 300 nm and about 400 nm. 3. The method of claim 1 further comprising monitoring a level of the reflected light having a wavelength between about 450 nm and 470 nm. 4. The method of claim 1 wherein the ablation energy is selected from the group consisting of radiofrequency (RF) energy, microwave energy, electrical energy, electromagnetic energy, cryoenergy, laser energy, ultrasound energy, acoustic energy, chemical energy, thermal energy, electroporation energy, and combinations thereof. 5. The method of claim 1 further comprising advancing a catheter to the cardiac tissue, the catheter comprising: a catheter body;a distal tip positioned at a distal end of the catheter body for delivering ablation energy to the cardiac tissue, the distal tip defining an illumination cavity having one or more openings for exchange of light between the illumination cavity and the tissue; andone or more optical fibers extending through the catheter body into the illumination cavity of the distal tip, the one or more optical fibers being in communication with a light source and a light measuring instrument to illuminate the cardiac tissue and to relay light energy reflected from the cardiac tissue to the light measuring instrument. 6. The method of claim 1 further comprising illuminating the cardiac tissue in a radial direction and an axial direction with respect to a longitudinal axis of a catheter. 7. The method of claim 1 further comprising providing a real time visual feedback about the lesion formation by displaying the level of NADH fluorescence. 8. The method of claim 1, wherein the ablation energy is applied when a NADH fluorescence peak is detected. 9. The method of claim 1 further comprising performing an ultrasound evaluation of cardiac tissue in combination with monitoring the level of NADH fluorescence. 10. The method of claim 1 further comprising monitoring a level of the reflected light having a wavelength between about 375 nm and about 575 nm. 11. The method of claim 1, further compri... | The George Washington University 460Medical, Inc. | A61B018/14 A61B005/00 A61B018/12 A61B018/00 A61B090/00 A61B005/06
US10722784 | B2 | 2020-07-28 | Bimanual integrative virtual rehabilitation system and methods | 2017-08-06 | A system provides bi-manual game-based integrative therapy combining cognitive training with mild upper body physical exercising during game play. The system includes a pair of interfaces tracking a user's arms in 3D and detect trigger pressing. Custom therapeutic game controllers that detect grasping force, finger extension, 3D hand position, skin temperature and pulse may be used. Using one of these interfaces the patient plays a series of custom games displayed on a laptop, medical grade workstation, or other computer platform. The whole therapeutic system may be integrated on a medical cart, so to make the system mobile and easier to place in a clinical setting. Games were designed to improve motor control, shoulder strength, finger and arm range of motion, task sequencing, focusing, decision making (executive function), short term and long term visual/auditory memory, and were progressed in difficulty over 6 to 8 weeks of therapy. This therapy reduces depression. | 1. A method of providing therapy to a user having an arm and a hand comprising: providing a hand-held controller comprising a 3D movement sensor configured to detect movement of the hand-held controller, a grip sensor attached to the movement sensor and engageable by the user, and a finger extension sensor attached to the movement sensor and engageable by the user, at least a portion of the finger extension sensor being forced toward the grip sensor and one or more fingers of the user when the user grips the grip sensor, wherein the finger extension sensor comprises a rotating mechanical lever;executing a video game on a computer and portraying action from the video game on a display, the action being viewable by the user, the user holding the hand-held controller;sending, from the hand-held controller, one or more signals representative of data from one or more of the movement sensor, the grip sensor, and the finger extension sensor to the computer, and analyzing, at the computer, the one or more signals; andcontrolling, by the computer, action portrayed on the display based on the one or more signals. 2. The method of claim 1, wherein the hand-held controller further comprises a forearm support, a skin temperature sensor, and a pulse sensor, the one or more signals representative of data further comprising data from at least one of the skin temperature sensor and the pulse sensor. 3. A hand-held controller comprising: a movement sensor configured to detect movement of the hand-held controller;a grip sensor attached to the movement sensor and engageable by a user; anda finger extension sensor attached to the movement sensor and engageable by a user, at least a portion of the finger extension sensor being forced toward the grip sensor and one or more fingers of the user when the user grips the grip sensor,wherein the finger extension sensor comprises a rotating mechanical lever. 4. The hand-held controller of claim 3, further comprising: a forearm support attached to the movement sensor;a skin temperature sensor attached to the forearm support; anda pulse sensor attached to the forearm support. 5. The hand-held controller of claim 3, wherein the movement sensor comprises a 3D tracker configured to measure a position and an orientation of the movement sensor. 6. The hand-held controller of claim 5, wherein the position and orientation of the movement sensor is associated with one of an arm and a hand of a user. 7. The hand-held controller of claim 3, wherein the grip sensor comprises a sealed air chamber. 8. The hand-held controller of claim 7, wherein the sealed air chamber has a bulb shape. 9. The hand-held controller of claim 3, wherein the rotating mechanical lever comprises a rotary sensor configured to sense movement of the rotating mechanical lever. 10. The hand-held controller of claim 9, wherein the finger extension sensor comprises a spring having a spring constant, the spring constant having a value causing the rotating mechani... | Bright Cloud International Corporation | A63F013/218 A63F013/211 A61B005/0205 A63F013/85 A63F013/20 G16H050/30 G16H040/63 A61H001/00 G16H020/30 A61B005/024 A63F013/42
US10709388 | B2 | 2020-07-14 | Biometric, physiological or environmental monitoring using a closed chamber | 2016-05-09 | A monitoring device configured for insertion into a conduit of a subject includes a housing, at least one physiological sensor coupled to the housing, a transmitter coupled to the housing, an expandable element, inflatable element, stretched membrane or balloon coupled to the housing and configured to occlude at least a portion of the conduit, a power source attached to the housing, and a processor coupled to the housing and operatively coupled to memory containing computer instruction causing the monitoring device to obtain physiological information via the at least one physiological sensor where the physiological information includes one or more of pulse rate information, body temperature information, breathing rate information, blood pressure information, cardiac output information, or blood gas level information, and processing and analyzing the physiological information to provide a result. | 1. A monitoring device configured for placement in or insertion into a conduit of a subject, comprising: a stent, with an acoustic channel;a speaker, where the acoustic channel is configured to carry audio from the speaker;an ambient sound microphone configured to measure an ambient acoustic environment;a housing;at least one physiological sensor coupled to the housing;a transmitter coupled to the housing;an occlusion component coupled to the housing and configured to occlude at least a portion of the conduit wherein the conduit is an ear canal, where the occlusion component wraps circumferentially about a portion of the stent;a first light emitting diode within a sealed portion of the ear canal;a power source attached to the housing; anda processor coupled to the housing and operatively coupled to a memory containing computer instructions, the computer instructions when executed by the processor causes the monitoring device to perform operations comprising; obtaining physiological information from a subject via the at least one physiological sensor, wherein the physiological information comprises one or more of the following: pulse rate information, body temperature information, breathing rate information, blood pressure information, cardiac output information, and blood gas level information; andprocessing and analyzing the physiological information to provide a result that comprises an identification of a behavioral response of the subject. 2. The monitoring device of claim 1, further including an environmental sensor, and further comprising obtaining environmental condition information in a vicinity of the monitoring device via the environmental sensor attached to the housing, wherein the environmental condition information includes information about one or more of the following: volatile organic compounds (VOCs), pollution, noise, light, or temperature. 3. The monitoring device of claim 2, further comprising analyzing the processed physiological information, via the processor, to identify a correlation between the processed physiological information and the obtained environmental condition information. 4. The monitoring device of claim 3, further comprising generating, via the processor, a health and/or environmental exposure assessment that identifies and/or predicts one or more physiological or environmental issues in response to identifying a correlation between the obtained physiological information and environmental condition information. 5. The monitoring device of claim 4, further comprising analyzing the health and/or environmental exposure assessment, via the processor, to identify and/or predict psychological and/or physiological stress. 6. The monitoring device of claim 5, further comprising communicating and presenting identified and/or predicted psychological stress via the processor on a user interface. 7. The monitoring device of claim 2, further comprising transmitting the obtained environmental condition information to... | Staton Techiya, LLC | A61B005/00 A61B005/11 A61B005/145 A61B005/1172 H04R025/00 H04S007/00 A61B005/0205
US10721191 | B2 | 2020-07-21 | Fitness activity related messaging | 2018-10-15 | In one embodiment, a method for generating a message to a friend of a user is provided, comprising: processing activity data of a first user measured by an activity monitoring device to update a value of an activity metric for the first user; identifying a change in an inequality relationship between the value of the activity metric for the first user and a value of the activity metric for a second user; in response to identifying the change in the inequality relationship, prompting the first user to generate a message to the second user. | 1. A method comprising: receiving, over a network by at least one server computer, first activity data of a first user measured by a first activity monitoring device; the server computer configured to: receive, over a network by the at least one server computer, second activity data of one or more additional users measured by respective additional activity monitoring devices;define, based on the first and second activity data, and an input from the first user, at least one group event;cause an invitation to join the at least one group event to be sent to the one or more additional users; andgenerating, with the server computer, the at least one group event, wherein the group event defines one of a competitive race for the participating users, a common goal for each of the participating users, or a collective goal for the participating users. 2. The method of claim 1 wherein the group event is or includes a challenge or a race. 3. The method of claim 2, wherein the group event includes the participating users competing against one another to achieve a superior activity amount or activity-related outcome. 4. The method of claim 2, wherein the group event includes missions in which the participating users each have individual goals that are pursued together in a group context. 5. The method of claim 2, wherein the group event includes a group goal toward which the participating users collectively pursue a goal. 6. The method of claim 1, further comprising generating a competitive group based on the responses to the invitation. 7. The method of claim 6, further comprising determining a value for at least one activity metric for a given user in the competitive group based on activity data captured by an activity monitoring device associated with the given user. 8. The method of claim 7, wherein the at least one activity metric comprises one or more of a quantity of steps taken, a quantity of floors climbed, or a quantity of calories burned. 9. The method of claim 1, wherein the group event is defined for an activity metric for each of the participating users for a period of time. 10. The method of claim 1, wherein generating the group event includes one or more of displaying data pertaining to the group event, generating messages for the group event and displaying the messages. 11. The method of claim 1, further comprising generating customized messages to at least one user, the customization based on characteristics of the at least one user. 12. The method of claim 11, wherein generating messages includes sending messages to one or more client devices associated with the at least one user. 13. The method of claim 11, wherein generating messages includes sending messages to one or more activity monitoring devices associated with the at least one user. 14. The method of claim 1, wherein one or both of the common goal and the collective goal relates to at least one of: weight, food intake, and sleep data for one or more of the participating... | Fitbit, Inc. | H04L012/58 H04M001/725 G08B003/10 G08B005/36 G08B006/00 A61B005/22 G06F016/2457 A61B005/00 A61B005/11 G01C022/00 G06F019/00 G06F040/186 G06F003/041 G06F015/00 H04L029/06 H04L029/08 G08B021/18 A63B024/00 H04L012/18 A61B005/0205 A61B005/021 A61B005/024 A61B005/1455 G06F003/0482
US10716914 | B2 | 2020-07-21 | Catheter system | 2014-03-07 | In a catheter system and method for assembly thereof, one end of a catheter shaft is inserted into a connector for a handle, with the connector being generally tubular and having a channel extending longitudinally therethrough. At least a lengthwise segment of the channel is sized in transverse cross-section greater than an outer cross-section of the catheter shaft. The catheter shaft is inserted into the channel of the connector such that the catheter shaft is spaced from the connector within the lengthwise segment of the channel. Adhesive is directed through a port in the sidewall of the connector into the lengthwise segment of the channel to flow circumferentially about the catheter shaft within the channel. The adhesive is allowed to cure and the connector (with the catheter adhered thereto) is retained on the handle. | 1. A catheter system comprising: a catheter having an elongate flexible shaft, the shaft having an outer surface defining an outer transverse cross-sectional dimension; anda handle having a connector for connecting the catheter to the handle, the connector being generally tubular having a sidewall, an open front end, an open rear end and a channel extending longitudinally within the connector from the front end to the rear end thereof, the connector channel has an inner transverse cross-sectional dimension that receives the catheter shaft therein, at least a segment of the channel being configured such that the inner transverse cross-sectional dimension thereof is greater than the outer transverse cross-sectional dimension of the catheter shaft, said connector further having a port extending through the sidewall into fluid communication with said at least a segment of the channel, the port exposing the catheter shaft within said segment of the channel and permitting adhesive to enter the channel via the port and flow circumferentially about the outer surface of the catheter shaft within said segment of the channel, wherein, over a first segment of the connector channel, an inner surface of the sidewall is tapered such that the inner transverse cross-sectional dimension of the connector channel increases as the connector channel extends towards the front end. 2. The catheter system of claim 1, wherein the catheter shaft has an end and is connected to the handle connector at said end, the connector channel comprises a shaft receiving portion extending longitudinally rearward from the front end of the connector, and an end portion extending longitudinally forward from the rear end of the connector, the inner transverse cross-sectional dimension of the channel at said end portion being smaller than the outer transverse cross-sectional dimension of the catheter shaft to define a seat within the connector against which the end of the catheter shaft abuts upon insertion of the catheter shaft into the connector, wherein the seat includes an annular surface that is oriented substantially perpendicular to a longitudinal axis of the connector channel. 3. The catheter system of claim 2, wherein the catheter system further comprises a line extending within the catheter shaft along at least a portion of the length thereof, the line further extending outward from the end of the catheter shaft at the seat within the connector and outward of the connector through the open rear end of the catheter. 4. The catheter system of claim 1, wherein the connector is constructed at least in part of a material through which UV energy passes for curing adhesive in the channel of the connector, and wherein the port extends through the tapered inner surface. 5. The catheter system of claim 1, wherein the handle is configured to retain the connector on the handle. 6. The catheter system of claim 1, wherein the handle further comprises a flex relief member retained on the ha... | ST. JUDE MEDICAL, CARDIOLOGY DIVISION, INC. | A61M025/01 A61M025/00 A61B018/14 A61B018/00 A61B005/00
US10720964 | B2 | 2020-07-21 | Device for wireless transmission of data and/or power | 2017-12-24 | The present invention relates to a device (700a, 700b, 700c) for wireless transmission of data and/or power between the device and another device of a system, in particular of a patient monitoring system. To meet stringent relative time errors at low complexity the device comprises a connector (701) comprising a data transmission unit (703) and a magnetic coupling unit (702, 704) for transmitting power to and/or receiving power. A detection unit (705) detects coupling of a counterpart connector of another device of the system with the connector (701). A control unit (707) uses the detection that a counterpart connector of another device of the system has been coupled with the connector (701) as a trigger to determine and/or reset a relative time difference between a clock signal used by the device and a clock signal of the other device using i) the high frequency power signal of the magnetic coupling unit (702, 704) and of a magnetic coupling unit of the counterpart connector and/or ii) a received time calibration signal for determining and/or resetting the relative time difference. | 1. Device for wireless transmission of at least one of data and power between the device and another device of a system, comprising a patient monitoring system, said device comprising: a connector comprising a data transmission unit arranged for at least one of transmitting data to and receiving data from another device of the system having a counterpart connector, and a magnetic coupling unit for transmitting power to and receiving power from another device of the system having a counterpart connector by use of inductive coupling using a high frequency power signal,a detection unit for detecting coupling of a counterpart connector of another device of the system with the connector, anda control unit for using the detection that a counterpart connector of another device of the system has been coupled with the connector as a trigger to at least one of determine and reset a relative time difference between a clock signal used by the device and a clock signal of the other device using at least one of i) the high frequency power signal of the magnetic coupling unit and of a magnetic coupling unit of the counterpart connector and ii) a received time calibration signal for determining and/or resetting the relative time difference, anda data storage unit for storing the moment of detection that a counterpart connector of another device of the system has been coupled with the connector as a time stamp. 2. Device as claimed in claim 1, wherein the detection unit is configured to detect coupling of a counterpart connector in a magnetic, electric or optical manner. 3. Device as claimed in claim 1, wherein the detection unit is configured to detect coupling of a counterpart connector by at least one of (i) detecting if the strength of magnetic coupling between the magnetic coupling unit and a magnetic coupling unit of the counterpart connector exceeds a magnetic coupling threshold, in particular by detecting impedance, resonance frequency or induced voltage, and (ii) detecting if the intensity of data received by the data transmission unit from a data transmission unit of the other device exceeds an intensity threshold, in particular by detecting signal intensity or antenna impedance of an antenna of the data transmission unit. 4. Device as claimed in claim 1, wherein the data transmission unit is configured to transmit at least one of the time stamp and a determined relative difference between the clock signal used by the device and the clock signal of the other device to the other device for storage in a master device of the system. 5. Device as claimed in claim 1, wherein the other device comprises a battery and wherein the high frequency power signal used for transmitting power from the battery to the device is used for at least one of determining and resetting the relative time difference. 6. Device as claimed in claim 5, further comprising a battery, wherein the control unit is configured to switch off the battery or generate a control signal for ... | KONINKLIJKE PHILIPS N.V. | H04B005/00 H04B005/02 H02J050/10 A61B005/00
US10705519 | B2 | 2020-07-07 | Distributed vehicle system control system and method | 2017-03-16 | A distributed control system includes a remote control system configured to be communicatively coupled with plural separate vehicle systems. The remote control system is configured to remotely control operation of the vehicle systems and/or communicate with the local vehicle control system or operator. The remote control system also is configured to one or more of change how many of the vehicle systems are concurrently controlled by the remote control system or change how many remote operators of the remote control system concurrently control the same vehicle system of the vehicle systems. | 1. A remote control system comprising: a communication device configured to wirelessly communicate with one or more vehicle systems during movement of the one or more vehicle systems outside of vehicle yards and while the communication device is off-board the one or more vehicle systems, the communication device configured to remotely control the movement of the one or more vehicle systems by wirelessly communicating a first control signal to the one or more vehicle systems; anda controller configured to generate the first control signal based on manual input provided by a first remote operator that is located off-board the one or more vehicle systems and while the one or more vehicle systems are outside of a visible range of the first remote operator, wherein the first control signal that is generated by the controller and communicated to the one or more vehicle systems by the communication device designates an upper speed limit on the movement of the one or more vehicle systems while allowing the one or more vehicle systems to move at one or more slower speeds than the upper speed limit; andan alerter system configured to monitor an alertness of the remote operator and, responsive to determining that the first remote operator is not currently alert, switch remote control of the one or more vehicle systems to a different, second remote operator. 2. The remote control system of claim 1, wherein the controller is configured to generate different first control signals representative of different upper speed limits for different locations along one or more routes being traveled by the vehicle systems. 3. The remote control system of claim 1, wherein the controller is configured to generate the first control signal to direct an onboard operator of the at least one of the vehicle systems to change a speed of the at least one of the vehicle systems to match the upper speed limit while allowing the onboard operator to determine how quickly to change the speed of the at least one of the vehicle systems. 4. The remote control system of claim 1, wherein the communication device is configured to determine interference with wireless communication with the one or more vehicle systems via a communication link and, responsive to determining the interference, switch remote control of the movements of the one or more vehicle systems to another, different remote control system. 5. The remote control system of claim 1, wherein the controller is configured to switch remote control of the movements of at least one of the vehicle systems to a different, second remote operator responsive to the at least one of the vehicle systems leaving a first designated geographic area associated with the first remote operator or entering into a different, second designated geographic area associated with the second remote operator. 6. The remote control system of claim 1, wherein the controller is configured to generate and communicate one or more different, second control sig... | transportation ip holdings, llc | G05D001/00 B61L027/00 A61B005/00 A61B005/0205 A61B005/0402 A61B005/0476 A61B005/11 A61B005/18 H04L029/08 B61L003/12 A61B005/021 A61B005/024 A61B005/08 G08G009/00
US10716525 | B2 | 2020-07-21 | System and method for navigating to target and performing procedure on target utilizing fluoroscopic-based local three dimensional volume reconstruction | 2016-08-01 | A system and method for navigating to a target using fluoroscopic-based three dimensional volumetric data generated from two dimensional fluoroscopic images, including a catheter guide assembly including a sensor, an electromagnetic field generator, a fluoroscopic imaging device to acquire a fluoroscopic video of a target area about a plurality of angles relative to the target area, and a computing device. The computing device is configured to receive previously acquired CT data, determine the location of the sensor based on the electromagnetic field generated by the electromagnetic field generator, generate a three dimensional rendering of the target area based on the acquired fluoroscopic video, receive a selection of the catheter guide assembly in the generated three dimensional rendering, and register the generated three dimensional rendering of the target area with the previously acquired CT data to correct the position of the catheter guide assembly. | 1. A system for navigating to a target area within an organ using fluoroscopic-based three-dimensional volumetric data generated from two-dimensional fluoroscopic images, comprising: a medical device configured to be navigated through the organ to a target of the organ within the target area; anda computing device configured to: receive previously acquired CT data;generate a CT-based three-dimensional rendering of the target area of a patient based on the previously acquired CT data;determine a location of the medical device;receive a series of fluoroscopic images of the target area captured by a fluoroscopic imaging device;display the location of the medical device relative to the CT-based three-dimensional rendering of the patient;determine poses of the fluoroscopic imaging device corresponding to a plurality of fluoroscopic images of the series of fluoroscopic images;generate a fluoroscopic-based three-dimensional reconstruction of the target area based on the series of fluoroscopic images of the target area captured by the fluoroscopic imaging device and the poses determined;receive a selection of the medical device from at least one of the fluoroscopic-based three-dimensional reconstruction of the target area or at least two fluoroscopic images of the series of fluoroscopic images;receive a selection of the target from at least one of the fluoroscopic-based three-dimensional reconstruction of the target area or at least two fluoroscopic images of the series of fluoroscopic images;determine a position of the medical device relative to the target in the target area within the fluoroscopic-based three-dimensional reconstruction; andupdate the location of the medical device displayed relative to the CT-based three-dimensional rendering of the patient based on the position of the medical device relative to the target in the target area within the fluoroscopic-based three-dimensional reconstruction of the target area. 2. The system according to claim 1, wherein the computing device is further configured to receive the selection of the target located in the target area from the fluoroscopic-based three-dimensional reconstruction. 3. The system according to claim 1, wherein the computing device is configured to receive the selection of the target by: automatically detecting a portion of the target in at least one of the fluoroscopic-based three-dimensional reconstruction or at least two fluoroscopic images of the series of fluoroscopic images;presenting the detection to a user; andreceiving a user command either accepting or rejecting the detection. 4. The system according to claim 1, wherein the computing device is configured to receive the selection of the medical device by: automatically detecting a portion of the medical device in at least one of the fluoroscopic-based three-dimensional reconstruction or at least two fluoroscopic images of the series of fluoroscopic images;presenting the detection to a user; andreceiving a user command either... | Covidien LP | A61B005/05 A61B006/00 A61B005/06 A61B034/10 A61B034/20 A61B006/12 A61B090/00 A61B017/00
US10716480 | B2 | 2020-07-21 | Hearing aid earpiece covers | 2018-06-13 | A cover is configured to be removably attached to and at least partially cover a hearing aid earpiece. The removable cover includes light transmissive material in at least one portion thereof that is configured to guide light from at least one optical emitter associated with the hearing aid earpiece to a region of a body of a subject wearing the hearing aid earpiece and/or guide light from the body of the subject to at least one optical detector associated with the hearing aid earpiece. The removable cover may include at least one alignment member that facilitates stability of the hearing aid earpiece when inserted within an ear of the subject. The removable cover may include at least one lens region, at least one light diffusion region, and/or at least one luminescence region in optical communication with the light transmissive material. | 1. A cover configured to be removably attached to, and at least partially cover, a hearing aid earpiece, the cover comprising light transmissive material in at least one portion thereof configured to guide light from at least one optical emitter associated with the hearing aid earpiece to a region of an ear of a subject wearing the hearing aid earpiece and guide light from the ear of the subject to at least one optical detector associated with the hearing aid earpiece, wherein the light transmissive material comprises at least one lens region that facilitates inspection light coupling from the at least one optical emitter into the ear region and facilitates scattered light coupling to the at least one optical detector from the ear region. 2. The cover of claim 1, wherein the cover comprises a soft material which deforms when the hearing aid earpiece is inserted within an ear of the subject. 3. The cover of claim 1, wherein the cover comprises silicone. 4. The cover of claim 1, further comprising at least one alignment member that facilitates stability of the hearing aid earpiece when inserted within an ear of the subject. 5. The cover of claim 1, further comprising at least one light diffusion region in optical communication with the light transmissive material that diffuses light emitted by the at least one optical emitter. 6. The cover of claim 1, further comprising at least one luminescence region in optical communication with the light transmissive material, wherein the at least one luminescence region receives light emitted by the at least one optical emitter and converts at least a portion of the received light to light having a different wavelength from the received light. 7. The cover of claim 6, wherein the at least one luminescence region comprises phosphor. 8. The cover of claim 1, comprising an outer surface and an inner surface, and further comprising a light reflective material on at least a portion of one or both of the inner and outer surfaces. 9. The cover of claim 1, comprising an inner surface and an outer surface, and wherein a layer of cladding material is disposed on at least a portion of the inner surface and/or outer surface, wherein the cladding material has a lower index of refraction than the light transmissive material. 10. The cover of claim 1, further comprising at least one aperture through which sound from a speaker of the hearing aid earpiece can pass. 11. A cover configured to be removably attached to, and at least partially cover, a hearing aid earpiece, the cover comprising light transmissive material in at least one portion thereof configured to guide light from at least one optical emitter associated with the hearing aid earpiece to a non-ear canal region of an ear of a subject wearing the hearing aid earpiece and guide light from the non-ear canal region of the ear of the subject to at least one optical detector associated with the hearing aid earpiece, wherein the light transmissive material compr... | Valencell, Inc. | A61B005/00 A61B005/0205 A61B005/024 H04R001/10 A61B005/11 A61B005/1455 A61B005/0476 A61B005/16 F21V008/00 A61B005/021 A61B005/08 A61B005/01 A61B005/026 A61B005/145 H04R001/02 G16H040/63 G16H040/67 G16H050/20 G16H030/40 A61B005/0295 A61B005/091
US10722158 | B2 | 2020-07-28 | Handheld oximeter probe with replaceable probe tip | 2017-04-20 | A handheld oximeter probe has a replaceable probe tip. The probe tip can be removed or detached from the probe unit and discarded. A replacement probe tip can be attached to the probe tip. The replaceable probe tip allows the probe unit to be reused many times, each time with new sterile probe tip. | 1. An oximeter device comprising: a first housing portion comprising:a processor housed in the first housing portion;a display, accessible from an exterior of the first housing portion, electronically coupled to the processor; anda first electrical connector electronically coupled to the processor; anda second housing portion, which removably couples with the first housing portion, comprising:an optical sensor comprising a plurality of source structures, a plurality of detector structures, and a first plurality of optical waveguides that optically coupled to the plurality of source structures and the plurality of detector structures; anda second electrical connector, mating with the first electrical connector, wherein the first housing portion and second housing portion are coupled together to form a handheld oximeter such that the first and second electrical connectors mate, the handheld oximeter comprises a proximal end formed by the first housing portion and a distal end formed by the second housing portion, a joint where the first and second housing portions meet together is flush, an entirety of the joint is between the proximal and distal ends, and the first and second housing portions together are configured to be held by a hand of the user such that the thumb and index finger of the user will grip the second housing portion while the first housing portion is cradled by a webbing between the thumb and index finger of the same hand. 2. The device of claim 1 wherein the display and optical sensor are directed away from each other. 3. The device of claim 1 wherein the display is directed away from tissue being measured when the optical sensor is directed toward the tissue being measured. 4. The device of claim 1 wherein the first housing portion comprises: a memory, housed in the first housing portion, electronically coupled to the processor;a battery, housed in the first housing portion, coupled to and supplies power to the processor, the memory, and the display. 5. The device of claim 1 wherein the first housing portion comprises a plurality of light sources, a second plurality of waveguides optically coupled to the light sources, at least a first optical waveguide and a second optical waveguide of the second plurality of optical waveguides optically coupled to at least a first optical waveguide and a second optical waveguide of the first plurality of optical waveguides when the first and second housing portions are coupled. 6. The device of claim 5 wherein the first and second optical waveguides of the first and second pluralities of optical waveguides are optical fibers. 7. The device of claim 1 wherein the second housing portion comprises a finger rest surface, coupled to a back side surface of the second housing portion, wherein the finger rest surface is a concave surface that extends at an angle in a turn direction relative to the back side surface of the second housing portion. 8. The device of claim 1 wherein the first housin... | ViOptix, Inc. | A61B005/1455 A61B005/1495 A61B005/00
US10716628 | B2 | 2020-07-21 | Systems, methods and devices for calculating hip center of rotation, adjusting parameters of joint replacement for pelvic tilt and calculating leg length and offset | 2016-10-28 | Systems and methods are described herein to calculate implant orientation measurements of an acetabular cup implant in hip replacement surgery, such surgery performed with minimally invasive incisions. Surgeons may obtain real-time updated implant orientation measurements that compensate for pelvic tilt of a patient, such tilt measured pre-operatively. Implant orientation measurements may be measured from the patient's anterior pelvic plane, supine coronal plane, standing coronal plane or any other reference plane that the surgeon may find useful. Also disclosed are systems, methods and devices to measure hip center of rotation and provide leg length and offset measurements in hip replacement surgery. | 1. An electronic guidance system to provide intra-operative anatomic measurements to guide placement of an implant with respect to an anatomy of a patient, the electronic guidance system comprising: an optical sensor to generate and communicate image data of respective targets attached to a probe tool and a surgical tool to track the probe tool and the surgical tool during an implant placement procedure;an intra-operative computing unit communicatively coupled to the optical sensor, the intra-operative computing unit configured to: register a first reference plane for the anatomy of the patient in a defined coordinate system using the respective image data from the probe tool, where the probe tool points to at least three reference locations on the anatomy to define the first reference plane;provide a user interface to display the first reference plane;receive an inputted tilt angle of the anatomy of the patient relative to a second plane within the defined coordinate system;generate a second reference plane for the anatomy of the patient using the inputted tilt angle and the first reference plane;update the user interface to display the second reference plane, responsive to the inputted tilt angle;determine implant orientation measurements using the respective image data from the surgical tool as it is tracked during the implant placement procedure, referencing the implant orientation measurements relative to the second reference plane; anddisplay in the user interface the implant orientation measurements relative to the second reference plane. 2. The system of claim 1 wherein the intra-operative computing unit is further configured to: determine implant orientation measurements relative to the first reference plane using the respective image data from the surgical tool as it is tracked, referencing the implant orientation measurements to the first reference plane for the anatomy of the patient without using the inputted tilt angle; anddisplay in the user interface the implant orientation measurements relative to the first reference plane. 3. The system of claim 1 further comprising: an inclination sensor to generate and communicate inclination measurements to track the anatomy of the patient during the implant placement procedure where changes in inclination of the inclination sensor correspond to changes in inclination of the anatomy of the patient and the intra-operative computing unit is further configured to register a positional plane for the anatomy of the patient in the defined coordinate system using the inclination measurements, equating the positional plane to a horizontal plane within the defined coordinate system. 4. The system of claim 3 wherein the intra-operative computing unit is further configured to: calculate and display a tilt angle of the anatomy of the patient relative to the horizontal plane using the first reference plane and the positional plane. 5. The system of claim 3, wherein the inclination sensor comprises an ... | INTELLIJOINT SURGICAL INC. | A61B034/20 A61F002/34 A61F002/46 A61B005/026 A61B005/06 A61B005/00 A61B090/00 A61B034/00 A61B005/107
US10716493 | B2 | 2020-07-21 | Systems, methods, and apparatuses for peripheral arterial disease detection and mitigation thereof | 2017-02-27 | Embodiments described herein are directed to non-invasive detection of peripheral arterial disease. For example, a measuring apparatus is used to measure a patient's calf circumference. The measuring apparatus has text feature(s) or indicator(s) printed thereupon that indicate the likelihood that the patient has peripheral arterial disease based on the measured calf circumference. The assessment may be further refined by using a software application that assesses the likelihood of the patient having peripheral arterial disease using at least the calf circumference measurement, along with other information/data. Based on the assessment, a healthcare practitioner may prescribe a walking program for the patient to follow. A software application may track compliance of the walking program and provide escalating reminders to the patient if the patient continues to fail to comply with the prescribed walking program. | 1. A measuring apparatus for determining a calf circumference of a patient for diagnosing peripheral arterial disease, comprising: a strip of material, comprising: a first surface; anda second surface that opposes the first surface;wherein at least one of the first surface or the second surface comprise a plurality of color-coded segments arranged in series, a first segment of the plurality of color-coded segments having a first color and corresponding to a first range of calf circumference that represents a high likelihood that the patient has peripheral arterial disease, a second segment of the plurality of color-coded segments having a second color that is different than the first color and corresponding to a second range of calf circumference that represents a low likelihood that the patient has peripheral arterial disease, and a third segment having a third color that is different than the first color and second color and corresponding to a third range of calf circumference that represents a medium likelihood that the patient has peripheral arterial disease, andwherein the first range comprises values less than 34 centimeters, wherein the second range comprises values greater than 43 centimeters, and wherein the third range comprises values greater than 34 centimeters and less than 43 centimeters. 2. The measuring apparatus of claim 1, wherein the first surface further comprises a first indicator that represents a mean of patients that have peripheral arterial disease and a second indicator that represents a standard error of mean of patients that have peripheral arterial disease. 3. The measuring apparatus of claim 1, wherein the first surface further comprises a first indicator that represents a mean of patients that do not have peripheral arterial disease and a second indicator that represents a standard error of mean of patients that do not have peripheral arterial disease. 4. The measuring apparatus of claim 1, wherein at least one of the first surface or the second surface of the strip comprise a set of markings printed thereupon, said markings corresponding to a standard unit of length. 5. The measuring apparatus of claim 4, wherein the standard unit of length comprises centimeters. 6. The measuring apparatus of claim 4, wherein the standard unit of length comprises inches. 7. A measuring apparatus for determining a calf circumference of a patient for diagnosing peripheral arterial disease, comprising: a strip of material, comprising: a first surface; anda second surface that opposes the first surface;wherein at least one of the first surface or the second surface comprise a plurality of color-coded segments arranged in series, a first segment of the plurality of color-coded segments having a first color and corresponding to a first range of calf circumference that represents a high likelihood that the patient has peripheral arterial disease, and a second segment of the plurality of color-coded segments having a second color tha... | New Lifeware Inc. | A61B005/107 G01B003/1003 G16H020/30 A61B005/02 G01B003/1084 G16H050/20 G16H050/30 A61B005/00 G01C022/00
US10728642 | B2 | 2020-07-28 | Portable case for modular hearing assistance devices | 2018-09-19 | A portable case for storing hearing assistance devices is described that includes at least one retention structure configured to retain at least part of a hearing assistance device, one or more communication units configured to exchange information between the portable case and one or more external devices, and at least one processor. The processor is configured to: detect when the at least one retention structure shares an electrical connection with the at least part of the hearing assistance device that is retained by the at least one retention structure; and responsive to receiving data, from the hearing assistance device, via the electrical connection, while the at least part of the hearing assistance device is retained by the at least one retention structure, cause the one or more communication units to communicate with exchange the information data between the portable case and the one or more external devices. | 1. A portable case for storing hearing assistance devices, the portable case comprising: at least one retention structure configured to retain at least part of a hearing assistance device;one or more communication units configured to exchange information between the portable case and one or more external devices;one or more input components configured to detect user inputs as part of a user interface, wherein the one or more input components comprise an inertial measurement component configured to detect user gestures performed while holding the portable case; andat least one processor configured to: detect when the at least part of the hearing assistance device is retained by the at least one retention structure; andresponsive to receiving data, from the hearing assistance device, while the at least part of the hearing assistance device is retained by the at least one retention structure, cause the one or more communication units to communicate with the one or more external devices. 2. The portable case of claim 1, wherein the one or more external devices include: a computing device, the hearing assistance device, or at least one other hearing assistance device. 3. The portable case of claim 1, wherein the one or more communication units are configured for wireless communication between the portable case and the one or more external devices. 4. The portable case of claim 1, wherein the one or more communication units are configured for wired communication between the portable case and the one or more external devices. 5. The portable case of claim 1, wherein the one or more communication units are configured to communicate with the one or more external devices via at least one of: a near field communication radio, a Bluetooth® radio, a Wi-Fi® radio, a cellular radio, a software defined radio, a multimodal radio, a near field magnetic induction radio, or a radio configured to communicate via a wireless intra or inter body-network protocol. 6. The portable case of claim 1, wherein the at least one processor is further configured to: receive, via the one or more communication units, first data from a first external device of the one or more external devices, the first data being encoded according to a first communication protocol;after decoding the first data, generate second data, the second data being encoded according to a second communication protocol; andsend, via the one or more communication units, the second data to a second external device of the one or more external devices that is different than the first external device. 7. The portable case of claim 6, wherein: the hearing assistance device is a first hearing assistance device;the first external device is the first hearing assistance device; andthe second external device is a second hearing assistance device. 8. The portable case of claim 6, wherein: the hearing assistance device is a first hearing assistance device;the first external device is the first hearing assistance devic... | Starkey Laboratories, Inc. | H04R001/02 H04R001/10 H02J050/10 H02J007/00 H04R025/00 G16H040/63 A61B005/00 G16H040/67 G16H050/20 H02J007/10
US10722146 | B2 | 2020-07-28 | Patient movement and incontinence notification system | 2018-08-10 | A notification system is provided that provides notification of patient events such as movement and/or incontinence. The notification system provides for a plurality of different pressure sensor pads as well as an incontinence pad to be used in association with a single monitor. | 1. A patient movement notification system, comprising: a sensor pad comprising a pressure sensor for sensing pressure applied by a patient, and a transmitter circuit for transmitting a pressure signal; anda monitor having a receiver for receiving the pressure signal from the sensor pad, and a controller coupled to the receiver, the controller determining whether the pressure signal indicates that the patient is no longer applying pressure to the sensor pad and generating a notification of patient movement in response to the pressure signal if the pressure signal indicates that the patient is no longer applying pressure to the sensor pad; andat least one light for illuminating an area of the floor near the patient,wherein the controller controls the at least one light to selectively illuminate the area of the floor near the patient, wherein the monitor further comprises a housing in which the controller and the receiver are disposed, the housing including a mount for mounting the housing to a wall, and wherein the at least one light is disposed in the housing so as to project light from the housing to selectively illuminate the area of the floor near the patient, and wherein the monitor further includes a plurality of indicator lights for indicating a status of the sensor pad, the at least one light being separate from the plurality of indicator lights, andwherein the controller activates the at least one light in response to a determination that the pressure signal indicates that the patient is no longer applying pressure to the sensor pad. 2. The patient movement notification system of claim 1, wherein the receiver is part of a transceiver and the controller wirelessly controls the at least one light through the transceiver. 3. The patient movement notification system of claim 1 and further comprising a light strip for selectively illuminating an area where the patient is located, wherein the receiver is part of a transceiver and the controller wirelessly controls the light strip through the transceiver. 4. A patient movement notification system, comprising: a sensor pad comprising a pressure sensor for sensing pressure applied by a patient, and a transmitter circuit for transmitting a pressure signal;a light strip for selectively illuminating a path on a floor from a bed to a bathroom; anda monitor having a receiver for receiving the pressure signal from the sensor pad, the monitor generating a notification of patient movement in response to the pressure signal if the pressure signal indicates that the patient is no longer applying pressure to the sensor pad, the monitor having a mount for mounting the monitor to a wall,wherein the monitor is communicatively coupled to the light strip for controlling the light strip to illuminate the path in response to the pressure signal if the pressure signal indicates that the patient is no longer applying pressure to the sensor pad, wherein the monitor wirelessly transmits control signals to the light s... | Parasol Medical LLC | A61B005/00 A61B005/11 G08B021/04 G08B025/10 A61F013/42 G08B021/20 G08B021/22 G08B005/36 G08B029/18 A61B005/20
US10702213 | B2 | 2020-07-07 | Differentiation of heart failure risk scores for heart failure monitoring | 2018-04-26 | A method and device for differentiating heart failure risk scores that includes determining receipt of a current data transmission and acquiring patient metrics from a remote monitoring device, determining a daily heart failure risk score for each day occurring during a time period from a previous received data transmission to the current received data transmission based on the acquired patient metrics, and determining a maximum daily heart failure risk score of the determined daily heart failure risk scores during a lookback window prior to the current received data transmission. A heart failure risk score is determined for the received data transmission based on the determined maximum daily heart failure risk score, and a heart failure risk score alert is determined for the received data transmission based on the proximity of the determined maximum heart failure risk score and the current received data transmission. A display of at least one of the determined heart failure risk score and the determined heart failure risk score alert is generated, and the determined daily heart failure risk score for each day occurring during the time period and the generated display are stored. | 1. A method of differentiating heart failure risk scores in a heart failure monitoring device, comprising: receiving, at the monitoring device, a current data transmission from a remote monitoring device;acquiring patient metrics from the remote monitoring device;determining a daily heart failure risk score for each day occurring during a time period from a previous received data transmission to the current data transmission based on the acquired patient metrics;determining a maximum daily heart failure risk score of the determined daily heart failure risk scores during a lookback window prior to receipt of the current data transmission;determining a heart failure risk score for the current data transmission based on the determined maximum daily heart failure risk score;determining a heart failure risk score alert for the current data transmission based on whether the heart failure risk score for the current data transmission matches a prior heart failure risk score determined during the lookback window prior to receipt of the current data transmission; andgenerating a display of at least one of the determined heart failure risk score for the current data transmission and the determined heart failure risk score alert. 2. The method of claim 1, wherein the determined heart failure risk score for the current data transmission comprises one of a high heart failure risk score, a medium heart failure risk score, and a low heart failure risk score. 3. The method of claim 2, further comprising: determining, in response to the determined heart failure risk score for the current data transmission being high, whether a high alert or a high ongoing alert was determined within the lookback window prior to receipt of the current data transmission; anddetermining that the heart failure risk score alert for the current data transmission is a high alert in response to the high alert or the high ongoing alert not being determined within the lookback window prior to receipt of the current data transmission. 4. The method of claim 3, further comprising waiting for a next received data transmission in response to at least one of the determined heart failure risk score for the current data transmission not being high, the high ongoing alert not being enabled, a high alert or a high ongoing alert was not determined to occur within the lookback window prior to receipt of the current data transmission, or the most recent day with a daily heart failure risk score of high in a prior transmission is not included within the lookback window prior to receipt of the current data transmission. 5. The method of claim 3, further comprising determining a high heart failure risk score alert in response to at least one of the determined heart failure risk score for the current data transmission not being high, the high ongoing alert not being enabled, a high alert or a high ongoing alert not being determined to occur within the lookback window prior to receipt of the current dat... | Medtronics, Inc. | A61B005/00 A61B005/053 A61B005/0402 A61N001/37 A61N001/372 G16H050/30 A61B005/046
US10726546 | B2 | 2020-07-28 | Tissue-to-flow image generation in medical imaging | 2018-04-26 | A network is machine trained to estimate flow by spatial location based on input of anatomy information. A medical scan of tissue may be used to generate flow information without the delay or difficulty of performing a medical scan configured for flow imaging or CFD. Anatomy imaging is used to provide flow estimates with the speed provided by the machine-learned network. | 1. A method for flow estimation based on anatomical information from a medical imaging system, the method comprising: scanning a region of a patient by the medical imaging system in three dimensions, the region having tissue and fluid, and the scanning being configured for detecting tissue responses, the tissue responses representing the region in three dimensions;generating flow estimates for the fluid of the region from the tissue responses in three dimensions, the flow estimates generated by a machine-trained image-to-image network for tissue-to-flow conversion; anddisplaying a three-dimensional rendering from the flow estimates. 2. The method of claim 1 wherein scanning comprises scanning with the medical imaging system comprising a magnetic resonance or ultrasound scanner, and wherein generating the flow estimates comprises generating phase contrast or Doppler estimates by the machine-trained image-to-image network. 3. The method of claim 1 wherein generating comprises inputting the tissue responses to the machine-trained image-to-image network. 4. The method of claim 1 wherein generating comprises inputting features derived from the tissue responses to the machine-trained image-to-image network. 5. The method of claim 1 wherein the medical imaging system comprises a first modality; wherein scanning comprises detecting the tissue responses with a physics of the first modality;further comprising:determining tissue responses for physics of a second modality different than the physics of the first modality, the determining being by another machine-trained image-to -image network for tissue-to-tissue conversion between the first and second modalities;wherein generating the flow estimates comprises generating from the tissue responses for the physics of the second modality. 6. The method of claim 5 wherein the first modality comprises computed tomography, and wherein the second modality comprises magnetic resonance, and wherein generating the flow estimates comprises generating phase-contrast estimates. 7. The method of claim 1 wherein generating the flow estimates comprises generating the flow estimates as magnitude of flow. 8. The method of claim 1 wherein generating the flow estimates comprises generating the flow estimates as directional flow estimates. 9. The method of claim 1 further comprising determining a value of a hemodynamic parameter from the flow estimates, and wherein displaying the three -dimensional rendering comprises displaying the value of the hemodynamic parameter for the region of the patient. 10. The method of claim 1 wherein displaying the three -dimensional rendering comprises displaying a spatial representation of the flow estimates for the region of the patient. 11. The method of claim 1 wherein scanning comprises scanning for the tissue response without scanning for fluid response. 12. A medical imaging system for flow estimation based on anatomical information, the medical imaging system comprising: a compu... | Siemens Healthcare GmbH | G06T007/00 A61B008/06 A61B008/08 A61B005/00 A61B005/055 A61B006/03
US10709369 | B2 | 2020-07-14 | Needle assembly with flashback chamber for collecting blood or other liquid samples | 2015-10-23 | A needle assembly comprising a housing, a first needle hub provided by the housing, a first hollow needle portion mounted to the first needle hub and extending from the first needle hub in a distal direction. A second needle hub is provided by the housing. A second hollow needle portion is mounted to the second needle hub and extends from the second needle hub in a proximal direction, opposite to the distal direction, wherein a hollow flashback chamber is arranged within the housing between the first needle portion and the second needle portion. The flashback chamber communicates with the inner lumen provided within the first and second hollow needle portion and is connected to the atmosphere by a passage and wherein a filter is provided within the passage, which is designed such that it is permeable for air to pour out of the flashback chamber into the atmosphere. | 1. A needle assembly for collecting blood or other liquid samples for discharge into a container, the needle assembly comprising: a housing;a first needle hub provided by said housing;a first hollow needle portion mounted to said first needle hub and extending from said first needle hub in a distal direction;a second needle hub provided by said housing;a second hollow needle portion mounted to said second needle hub and extending from said second needle hub in a proximal direction, opposite to said distal direction;wherein a hollow flashback chamber is arranged within the housing and extends longitudinally from a proximal end of the first needle portion to a distal end of the second needle portion, wherein the flashback chamber communicates with an inner lumen provided within the first and second hollow needle portion, wherein the housing is formed from transparent material at least in a region close to the flashback chamber, wherein the flashback chamber is connected to the atmosphere by a passage and wherein a filter is provided within the passage, wherein the filter is configured to be permeable for air to pour out of the flashback chamber into the atmosphere but prevents blood or other liquid samples contained in the flashback chamber from pouring out through an opening into the atmosphere; andwherein a longitudinal length of the flashback chamber approximately equals an inner diameter of an entry opening of a cylindrical channel of the passage extending transversely from a longitudinal axis of the flashback chamber. 2. The assembly according to claim 1, wherein the housing is formed from transparent plastics. 3. The assembly according to claim 1, wherein the housing includes an outer shell, a distal body and an inner tubular portion, wherein the outer shell surrounds a hollow space, wherein the outer shell is connected to the inner tubular portion by the distal body. 4. The assembly according to claim 3, wherein the flashback chamber is arranged at least partially within the inner tubular portion. 5. The assembly according to claim 1, wherein the first needle hub is formed at least partially by a distal body. 6. The assembly according to claim 1, wherein the first needle portion is formed by a first needle and wherein the second needle portion is formed by a second needle, wherein the first needle is formed separately from the second needle, wherein the first needle comprises an end and the second needle comprises an end and the ends of the first and second needle are facing one another so as to define facing needle ends, wherein facing needle ends limit the flashback chamber and wherein a port opens into the flashback chamber between the facing needle ends of the first and second needle portions. 7. The assembly according to claim 1, wherein the first needle portion and the second needle portion are formed by a single needle extending through the housing, wherein the needle has an opening arranged within or close to the flashback cham... | POLY MEDICURE LIMITED | A61B005/15 A61B005/154 A61M025/06
US10716540 | B2 | 2020-07-21 | Resin composition for acoustic wave probe, and acoustic lens using the same, acoustic wave probe, acoustic wave measurement apparatus, ultrasound diagnostic apparatus, photoacoustic wave measurement apparatus, and ultrasound endoscope | 2018-09-27 | Provided are a resin composition for an acoustic wave probe including, a polymer which has a structural unit having a siloxane bond and a structural unit having a urea bond, an acoustic lens using the same, the acoustic wave probe, acoustic wave measurement apparatus, ultrasound diagnostic apparatus, photoacoustic wave measurement apparatus, and ultrasound endoscope. | 1. An acoustic wave probe comprising: an acoustic lens comprising a resin composition, the resin composition comprising a polymer which has a structural unit having a siloxane bond and a structural unit having a urea bond. 2. An acoustic wave measurement apparatus comprising: the acoustic wave probe according to claim 1. 3. An ultrasound diagnostic apparatus comprising: the acoustic wave probe according to claim 1. 4. A photoacoustic wave measurement apparatus comprising: an acoustic lens comprising a resin composition, the resin composition comprising a polymer which has a structural unit having a siloxane bond and a structural unit having a urea bond. 5. An ultrasound endoscope comprising: an acoustic lens comprising a resin composition, the resin composition comprising a polymer which has a structural unit having a siloxane bond and a structural unit having a urea bond. 6. The acoustic wave probe according to claim 1, wherein the structural unit having a siloxane bond is represented by General Formula (1) and the structural unit having a urea bond is represented by General Formula (2)in the formulae, R1 and R2 each independently represent a monovalent organic group, L1 and L2 each independently represent a single bond or a divalent linking group, X represents a divalent linking group, and m represents an integer of 1 to 10,000. 7. The acoustic wave probe according to claim 1, wherein an amount of urea in the polymer is 0.2 to 3.0 mmol/g. 8. The acoustic wave probe according to claim 1, wherein a mass average molecular weight of the polymer is greater than or equal to 100,000. 9. The acoustic wave probe according to claim 1, wherein a proportion of the structural unit having a siloxane bond in the polymer is greater than or equal to 70 mass %. 10. The acoustic wave probe according to claim 1, wherein a density of the polymer is greater than or equal to 1.1 g/cm3. 11. The acoustic wave probe according to claim 1, wherein the structural unit having a urea bond has an aromatic ring. 12. The acoustic wave probe according to claim 1, wherein the structural unit having a siloxane bond has an aromatic ring. | FUJIFILM Corporation | A61B008/00 C08G018/61 C08G018/76 C08G018/48 C08G018/73 C08G077/452 A61B005/00 A61B008/12 G10K011/30
US10709394 | B2 | 2020-07-14 | Method and system for 3D reconstruction of X-ray CT volume and segmentation mask from a few X-ray radiographs | 2018-01-15 | A method and apparatus for automated reconstruction of a 3D computed tomography (CT) volume from a small number of X-ray images is disclosed. A sparse 3D volume is generated from a small number of x-ray images using a tomographic reconstruction algorithm. A final reconstructed 3D CT volume is generated from the sparse 3D volume using a trained deep neural network. A 3D segmentation mask can also be generated from the sparse 3D volume using the trained deep neural network. | 1. A method for automated reconstruction of a 3D computed tomography (CT) volume one or more X-ray images, comprising: generating a sparse 3D volume from one or more X-ray images of a patient; andgenerating a final reconstructed 3D CT volume from the sparse 3D volume using a trained deep neural network. 2. The method of claim 1, wherein the one or more X-ray images of the patient comprise a first x-ray image and a second x-ray image, and generating the sparse 3D volume from the one or more X-ray images of the patient comprises: generating the sparse 3D volume from the first X-ray image and the second X-ray image using a tomographic reconstruction algorithm. 3. The method of claim 2, wherein the one or more x-ray images of the patient comprise only the first and second x-ray images, and generating the sparse 3D volume from the first X-ray image and the second X-ray image using a tomographic reconstruction algorithm comprises: generating the sparse 3D volume from the first X-ray image and the second X-ray image without any additional x-ray images using a tomographic reconstruction algorithm. 4. The method of claim 1, further comprising: generating a 3D segmentation mask of a target object from the sparse 3D volume using the trained deep neural network. 5. The method of claim 4, wherein the trained deep neural network is a multi-output deep image-to-image network having encoder layers that code the sparse 3D volume into a code whose size is smaller than the spare 3D volume and decoder layers that decode the code into the final reconstructed 3D volume and the 3D segmentation mask of the target object. 6. The method of claim 1, wherein the trained deep neural network is a deep image-to-image network that is trained in a generative adversarial network together with a discriminator network for distinguishing between synthesized reconstructed 3D CT volumes generated by the deep image-to-image network from input sparse 3D volume training samples and real reconstructed 3D CT volume training samples. 7. The method of claim 1, wherein the trained deep neural network is a deep image-to-image network that is trained in a conditional-generative adversarial network together with a discriminator network for distinguishing between synthesized reconstructed 3D CT volumes generated by the deep image-to-image network from input sparse 3D volume training samples and real reconstructed 3D CT volume training samples, conditioned on the input sparse 3D volume training samples. 8. The method of claim 7, wherein the conditional-generative adversarial network is integrated with a voxel-wise cost function that computes a voxel-wise error between the synthesized reconstructed 3D CT volumes generated by the deep image-to-image network from input sparse 3D volume training samples and corresponding ground-truth reconstructed 3D CT volume training samples, and the deep image-to-image network and the discriminator network are trained together to optimize, over a plurality o... | Siemens Healthcare GmbH | G06T007/00 G06T007/30 A61B005/00 A61B006/03 G06T007/11 G06T011/00 A61B006/00 G06T005/00
US10722169 | B2 | 2020-07-28 | Physiological condition determination based on pressure wave produced by an implantable medical device housing | 2012-01-24 | A physiological state of a patient is detected by at least producing and detecting pressure waves with a free wall of an implantable medical device (IMD) housing. An actuator element may contact the free wall, e.g., a portion of the IMD housing, and cause movement of the free wall that produces a pressure wave within the fluid and tissue of the patient. A detector element contacting the free wall may in turn detect reflected pressure waves received by the free wall. An acoustic module within the IMD may then determine a physiological condition of the patient, e.g., a bladder fullness state, based on the time delay between the transmitted and reflected pressure waves. In some examples in which the IMD also delivers stimulation therapy to the patient, e.g., incontinence therapy, the IMD may also automatically adjust stimulation therapy based on the determined physiological condition. | 1. A method comprising: producing a pressure wave within a patient with a portion of a device housing of an implantable medical device using an actuator element configured to contact the portion of the device housing, wherein the device housing encloses a processor and the actuator element, and wherein the implantable medical device is configured to couple to an implantable medical lead;detecting at least one reflected pressure wave with the portion of the device housing; anddetermining, by the processor and based on the at least one reflected pressure wave, a physiological condition of the patient. 2. The method of claim 1, wherein the portion of the device housing is configured to enclose operational circuitry of the implantable medical device, the operational circuitry comprising the processor, and wherein the portion of the device housing is tuned to one or more frequencies at least one of the produced pressure wave or the at least one reflected pressure wave. 3. The method of claim 1, wherein producing the pressure wave further comprises: producing the pressure wave from an internal location inferior to a bladder and lateral to a labia; anddirecting the pressure wave towards the bladder. 4. The method of claim 1, wherein the physiological condition includes a bladder fullness state. 5. The method of claim 1, wherein determining the physiological condition further comprises: determining a time delay between the produced pressure wave and the at least one reflected pressure wave; anddetermining, based on the time delay, a distance between the portion of the device housing and an anatomical structure of the patient from which the at least one reflected pressure wave originated. 6. The method of claim 1, wherein the at least one reflected pressure wave comprises a first reflected pressure wave and a second reflected pressure wave, and wherein determining the physiological condition further comprises: determining a first time delay between the produced pressure wave and the first reflected pressure wave;determining a second time delay between the produced pressure wave and the second reflected pressure wave; anddetermining a distance between first and second anatomical structures of the patient from which the first and second reflected pressure waves, respectively, originated based on a difference between the first time delay and the second time delay. 7. The method of claim 1, further comprising: receiving an indication of a therapy adjustment input from the patient, the therapy adjustment input being received by one of the implantable medical device or an external programmer;determining the physiological condition after an input delay period immediately following receipt of the therapy adjustment input to set the determined physiological condition as a first physiological state of the patient;retrieving, from a memory, a last physiological condition determination before the therapy adjustment input was received; andsetting the last physio... | Medtronic, Inc. | A61B005/00 A61B005/20 A61N001/36 A61B008/08 A61B008/12 A61B008/00 A61B005/11 A61B005/0488
US10719695 | B2 | 2020-07-21 | Method for pushing picture, mobile terminal, and storage medium | 2018-04-19 | Provided are a method for pushing a picture, a terminal and a storage medium. The method includes the follows. An image containing a facial expression of a user is acquired, when a screen is lit; the image is analyzed to determine an emotion type of the user; and a target picture is pushed according to the emotion type of the user. | 1. A method for pushing a picture, comprising: acquiring an image containing a facial expression of a user when a screen is lit;comparing the image with a template to determine locations of multiple facial features;analyzing shapes of the facial features to determine multiple candidate emotion types;determining an emotion type with the highest proportion among the candidate emotion types as the emotion type of the user; andpushing a target picture according to the emotion type of the user. 2. The method of claim 1, wherein the pushing the target picture according to the emotion type of the user comprises: determining a type of a picture to be pushed according to the emotion type of the user;obtaining the target picture that matches with the type of the picture to be pushed from a picture gallery; andpushing the target picture to the user. 3. A mobile terminal, comprising: a memory configured to store program codes;at least one processor coupled with the memory, wherein the program codes, when executed by the at least one processor, are operable with the at least one processor to perform: an acquiring unit, configured to acquire an image containing a facial expression of a user when a screen is lit;an analyzing unit, configured to analyze the image to determine an emotion type of the user;a determining unit, configured to determine a type of a picture to be pushed according to the emotion type of the user;a learning unit, configured to learn the expression images according to a preset machine learning scheme to obtain the user emotion model;an obtaining unit, configured to obtain a target picture that matches with the type of the picture to be pushed from a picture gallery; anda pushing unit, configured to push the target picture to the user;wherein the analyzing unit comprises: a second comparing sub-unit, configured to compare the image with a user emotion model to determine a target emotion picture from the user emotion model that matches with the image; anda determining sub-unit, configured to determine the emotion type of the user according to the target emotion picture;wherein the acquiring unit being configured to acquire a preset number of expression images of the user. 4. The mobile terminal of claim 3, wherein the program codes, when executed by the at least one processor, are operable with the at least one processor to perform: a recording unit, configured to record a choosing result of the user; andan updating unit, configured to update the user emotion model according to the choosing result of the user. 5. A mobile terminal, comprising; a camera, configured to acquire an image containing a facial expression of a user when a screen is lit; anda general-purpose processor, configured to: analyze the image to determine an emotion type of the user;determine a type of a picture to be pushed according to the emotion type of the user;obtain a target picture that matches with the type of the picture to be pushed from a picture gallery; and... | GUANGDONG OPPO MOBILE TELECOMMUNICATIONS CORP., LTD. | G06K009/62 G06K009/00 A61B005/16 G06F003/01 A61B005/00
US10716519 | B2 | 2020-07-21 | Mapping and quantifying blood stasis and thrombus risk in the heart | 2016-11-23 | The present disclosure provides methods for in-vivo assessment of the location and extent of blood flow stasis regions inside a cardiac chamber or blood vessel and systems for performing the methods. The disclosure provides methods for assessing risk of intracardiac or intravascular thrombus or of embolism originating in a cardiac chamber or vessel, and methods for assessing the need for and/or optimization of cardiac resynchronization therapy. | 1. A method for identifying regions of blood flow stasis inside a cardiac chamber or blood vessel of a subject comprising: obtaining flow-velocity images of blood particles inside a cardiac chamber or blood vessel of the subject;calculating one or more of residence time (TR), standard deviation of residence time (σR), kinetic energy, or rate of distortion of blood particles inside the cardiac chamber or blood vessel using the flow-velocity images to generate numerical metrics of blood flow; andgenerating a residence time (TR) map, a kinetic energy map, a rate of distortion map, or combinations thereof, using the numerical metrics to identify and characterize regions of blood flow stasis;wherein generating numerical metrics of blood flow comprises calculating a blood stasis timescale index. 2. The method of claim 1, wherein generating numerical metrics of blood flow comprises calculating the blood flow's residence time (TR) inside the cardiac chamber or blood vessel. 3. The method of claim 1, wherein generating numerical metrics of blood flow comprises calculating the standard deviation of the blood flow's residence time (σR) inside the cardiac chamber or blood vessel. 4. The method of claim 3, comprising calculating the standard deviation of the blood flow's residence time in regions with blood flow residence time greater than a threshold value inside the cardiac chamber or blood vessel. 5. The method of claim 1, wherein generating numerical metrics of blood flow comprises calculating the blood flow's kinetic energy inside the cardiac chamber or blood vessel. 6. The method of claim 5, comprising calculating the blood flow's kinetic energy in regions with blood flow residence time greater than a threshold value inside the cardiac chamber or blood vessel. 7. The method of claim 6, comprising calculating the blood flow's rate of distortion in regions with blood flow residence time greater than a threshold value inside the cardiac chamber or blood vessel. 8. The method of claim 1, wherein generating numerical metrics of blood flow comprises calculating the rate of distortion of blood flow inside any cardiac chamber or blood vessel. 9. The method of claim 1, wherein generating numerical metrics of blood flow comprises calculating the size, shape, mobility, distance to the chamber wall, and perimeter in contact with the chamber wall of regions with blood flow residence time greater than a threshold value. 10. The method of claim 1, wherein obtaining flow-velocity images of blood particles inside a cardiac chamber or blood vessel of the subject comprises obtaining flow-velocity images of blood particles inside the left ventricular chamber, left atrium chamber, left atrial appendage, right-ventricular chamber, or right atrium chamber. 11. The method of claim 1, wherein obtaining flow-velocity images of blood particles inside a cardiac chamber or blood vessel is performed using a medical image-based apparatus able to determine blood flow velocit... | The Regents of the University of California Fundación para la Investigación Biomédica del Hospital Gregorio Marañó | A61B005/00 A61B005/029 A61B005/02 A61B005/026 A61B008/08 G16H050/30 G16H030/40 A61B005/0285 A61B006/00 A61B005/055 A61B008/06
US10709389 | B2 | 2020-07-14 | System and method for heart rate detection with motion artifact reduction | 2018-03-28 | Example embodiments relate to systems and methods for heart rate detection with motion artifact reduction. One embodiment includes an electronic system for heart rate detection. The electronic system includes a random sampling sensor module. The random sampling sensor module includes a first sensor circuit configured to provide nonuniform random samples below a Nyquist rate of a photoplethysmographic signal. The random sample sensor module also includes a second sensor circuit configured to provided nonuniform random samples below a Nyquist rate of a motion signal. The motion signal and the photoplethysmographic signals are sampled with an equivalent pattern. The electronic system also includes a heart rate detection module. The heart rate detection module is configured to calculate a heart rave value based on frequencies corresponding to peak powers of calculated power spectral density value sets corresponding to the photoplethysmographic signals in a frequency range of interest. | 1. An electronic system for heart rate detection comprising: a random sampling sensor module comprising: a first sensor circuit configured to provide nonuniform random samples below a Nyquist rate of a photoplethysmographic signal; anda second sensor circuit configured to provide nonuniform random samples below a Nyquist rate of a motion signal, wherein the motion signal is sampled with an equivalent nonuniform pattern as the photoplethysmographic signal; anda heart rate detection module configured to: receive a plurality of the nonuniform random samples below the Nyquist rate of the photoplethysmographic signal;calculate a power spectral density value set based on a Lomb-Scargle periodogram of the nonuniform random samples below the Nyquist rate of the photoplethysmographic signal;receive a plurality of the nonuniform random samples below the Nyquist rate of the motion signal;calculate a power spectral density value set based on a Lomb-Scargle periodogram of the nonuniform random samples below the Nyquist rate of the motion signal;normalize the calculated power spectral density value sets corresponding to the photoplethysmographic signal and the motion signal;subtract the normalized power spectral density value set corresponding to the motion signal from the normalized power spectral density value set corresponding to the photoplethysmographic signal;renormalize the power spectral density value set corresponding to the photoplethysmographic signal; andcalculate a heart rate value based on a frequency corresponding to a highest power peak value of the calculated power spectral density value set corresponding to the photoplethysmographic signal in a frequency range of interest. 2. The electronic system for heart rate detection according to claim 1, wherein the heart rate detection module is further configured to: detect a motion signal power spectral density peak value and a frequency corresponding to the motion signal power spectral density peak value;normalize the power spectral density value set corresponding to the photoplethysmographic signal based on the motion signal power spectral density peak value; andnormalize the power spectral density value set corresponding to the motion signal based on the motion signal power spectral density peak value. 3. The electronic system for heart rate detection according to claim 2, wherein the motion signal power spectral density peak value is a maximum value in a frequency range of interest. 4. The electronic system for heart rate detection according to claim 2, wherein the power spectral density value set corresponding to the photoplethysmographic signal and the power spectral density value set corresponding to the motion signal are normalized by setting a power spectral density value corresponding to the frequency of the motion signal power spectral density peak value to an equivalent normalization value and adapting a remainder of the power spectral density values accordingly. 5. The electronic sys... | IMEC VZW Katholieke Universiteit Leuven, KU LEUVEN R&D | A61B005/00 A61B005/024 A61B005/11
US10722671 | B2 | 2020-07-28 | Method and apparatus for determining optimal endotracheal tube size | 2015-05-14 | A method and apparatus for determining the optimal endotracheal tube size for the intubation of a patient based upon both optimal tube length and optimal tube diameter which are determined by using either direct or indirect measurement devices or techniques to measure the patient's glottis aperture and tracheal length. The apparatus may include a maneuverable arm, a fiber optic video guidance system and an introducer guide to maneuver an end of the apparatus around obstructions in a patient who has a difficult airway to facilitate insertion of either an optimally sized endotracheal tube or a standard endotracheal tube into the patient's airway. | 1. An apparatus for intubating a patient with an endotracheal tube, the endotracheal tube being adapted to be inserted in a trachea of a patient requiring intubation, the patient having a head, a face, a mouth, an oral cavity, a thoracic area, a trachea having a length and forming an airway in the patient, vocal cords, a cricoid arch cooperating with the patient's vocal cords in forming a glottic aperture, and a carina defining a point at which the trachea separates into a left and a right bronchial tube, the apparatus comprising: a hand-held intubating laryngoscope operatively connected to a handle and a detachable sterile cartridge, the laryngoscope including a curvilinear blade adapted to be inserted into the patient's oral cavity, the handle having a plurality of controls positioned thereon; the detachable sterile cartridge operatively connected to the handle, the detachable sterile cartridge being adapted to store a semi-flexible, hollow guide tube or introducer, wherein the curvilinear blade contains an extendable flexible intubation arm therein; the extendable flexible intubation arm including a hollow core extending coaxially along a length thereof, the extendable flexible intubation arm comprising a plurality of sections, each one of the plurality of sections being sequentially connected to an adjacent one of the plurality of sections by one of a plurality of extensible, rotatable joints, and a tip; anda fiber optic video and ultrasound bundle having a fiber optic video portion and an ultrasound portion, wherein the fiber optic video portion of the bundle is adapted to aid in visually guiding the flexible extendable intubation arm into the patient's oral cavity; and wherein the ultrasound portion of the bundle is adapted to measure the length of the patient's trachea, wherein the fiber optic video and ultrasound bundle positioned within the extendable flexible intubation arm; andthe semi-flexible, hollow guide tube or introducer having measuring capabilites and being adapted to be deployed into the patient's trachea by sliding through a channel of the extendable flexible intubation arm and configured to be removed from the extendable flexible intubation arm, the curvilinear blade, and the handle to facilitate intubation of the patient with an endotracheal tube by guiding the endotracheal tube into the patient by passing the endotracheal tube over the introducer. 2. The apparatus of claim 1, wherein the apparatus further includes a device for measuring a diameter of the glottic aperture. 3. The apparatus of claim 2, wherein the device for measuring the diameter of the glottic aperture comprises a ring mandrel. 4. The apparatus of claim 2, wherein the device for measuring the diameter of the glottic aperture comprises an optical scanning device. 5. The apparatus of claim 2, wherein the device for measuring the diameter of the glottic aperture comprises an ultrasound scanning device. 6. The apparatus of claim 1, wherein the endotrachea... | Securisyn Medical, LLC | A61M016/04 A61B008/12 A61B005/107 A61B001/267 A61B005/00 A61B005/0205 A61B005/08 A61B001/00
US10716867 | B2 | 2020-07-21 | High-resolution optical molecular imaging systems, compositions, and methods | 2016-02-05 | A composition includes a plurality of gold nanoparticles each having at least one surface. The gold nanoparticles have an average length of at least about 90 nm and an average width of at least about 25 nm. | 1. A composition comprising: a plurality of gold nanorods within an aqueous solution, the plurality of gold nanorods prepared for in vivo optical coherence tomography imaging of a target tissue,each of the plurality of gold nanorods having at least one surface,the plurality of gold nanorods having an average length of between about 90 nm and about 150 nm and an average width of between about 25 nm and about 50 nm,each of the plurality of gold nanorods being coated with polystyrene sulfonate (PSS),the polystyrene sulfonate coating configured to maintain nonaggregation of the plurality of gold nanorods during the in vivo optical coherence tomography imaging of the target tissue. 2. The composition of claim 1, wherein the average length is about 90 nm and the average width is about 30 nm. 3. The composition of claim 1, wherein the plurality of gold nanorods is within the aqueous solution as a colloidal suspension. 4. The composition of claim 1, wherein the PSS is electrostatically adsorbed on the at least one surface of the plurality of gold nanorods. 5. The composition of claim 4, further comprising: one or more functional groups associated with the PSS. 6. The composition of claim 5, wherein the one or more functional groups are covalently associated with the PSS. 7. The composition of claim 5, wherein the one or more functional groups are electrostatically associated with the PSS. 8. The composition of claim 5, wherein the one or more functional groups are associated with the PSS via one or more thiol groups. 9. The composition of claim 5, wherein the one or more functional group includes a biotin. 10. The composition of any one of claim 5, wherein the one or more functional groups includes any of an avidin, a poly(ethylene) glycol (PEG), an antibody, or an antigen binding portion of the antibody. 11. The composition of claim 9, wherein the biotin is PEGylated biotin. 12. The composition of claim 5, wherein the one or more functional group comprises a biotinylated ligand. 13. The composition of claim 12, wherein the biotinylated ligand is at least one of a biotinylated antibody, a fragment of the biotinylated antibody, a biotinylated antibody-drug conjugate, a biotinylated peptide, a biotinylated enzyme, a biotinylated G protein coupled receptor (GPCR) ligand, or a biotinylated transmembrane receptor ligand. 14. The composition of claim 4, further comprising: a plurality of polystyrene beads conjugated to at least a portion of the plurality of gold nanorods. 15. A method of imaging of a target tissue, comprising: administering, to the target tissue, a composition including a plurality of gold nanorods each having at least one surface, at least one of the gold nanorods of the plurality of gold nanorods including an anionic polymer electrostatically adsorbed on the at least one surface of the gold nanorod, the anionic polymer being polystyrene sulfonate (PSS), the plurality of gold nanorods having an average length of between about 90... | The Board of Trustees of the Leland Stanford Junior University | A61K049/00 A61B005/00 A61K049/04 C01G007/00 C08L025/06 B82Y005/00
US10720942 | B2 | 2020-07-21 | Apparatus and method for data compression in a wearable device | 2016-06-01 | Described is an apparatus and method for data compression using compressive sensing in a wearable device. Described is also a machine-readable storage media having instruction stored thereon, that when executed, cause one or more processors to perform an operation comprising: receive an input signal from a sensor; convert the input signal to a digital stream; and symmetrically pad on either ends of the digital stream with a portion of the digital stream to form a padded digital stream. | 1. Non-transitory machine-readable storage media having instruction stored thereon, that when executed, cause one or more processors to perform an operation comprising: receive an input signal from a sensor;convert the input signal to a digital stream; andsymmetrically pad on either ends of the digital stream with a portion of the digital stream to form a padded digital stream, wherein the portion of the digital stream is non-zero, and wherein the portion of the digital stream is a sequence of zeros and ones which is identical to a sequence of zeros of ones in the digital stream. 2. The non-transitory machine-readable storage media of claim 1 having further instruction stored thereon, that when executed, cause the one or more processors to perform a further operation comprising: select a part of the digital stream;transmit the selected part of the digital stream to another device; andreceive one or more configuration parameters from the other device in response to the transmitted selected part of the digital stream. 3. The non-transitory machine-readable storage media of claim 2 having further instruction stored thereon, that when executed, cause the one or more processors to perform a further operation comprising: prior to symmetrically padding on either ends of the digital stream, compute a number of elements to be padded using the one or more configuration parameters. 4. The non-transitory machine-readable storage media of claim 3, wherein number of elements determines the portion of the digital stream. 5. The non-transitory machine-readable storage media of claim 2, wherein at least one of the one or more configuration parameters is dynamically received from the other device. 6. The non-transitory machine-readable storage media of claim 2, wherein at least one of the one or more configuration parameters is periodically received from the other device. 7. The non-transitory machine-readable storage media of claim 2, wherein at least one of the one or more configuration parameters is a compression ratio. 8. The non-transitory machine-readable storage media of claim 3 having further instruction stored thereon, that when executed, cause the one or more processors to perform a further operation comprising: construct an encoding matrix using the one or more configuration parameters; andencode the padded digital stream using the encoding matrix to generate compressed padded data. 9. The non-transitory machine-readable storage media of claim 8, wherein the encoding matrix is a Binary Permuted Block Diagonal (BPBD) matrix. 10. The non-transitory machine-readable storage media of claim 9 having further instruction stored thereon, that when executed, cause the one or more processors to perform a further operation comprising: transmit the compressed padded data to the other device for processing. 11. The non-transitory machine-readable storage media of claim 1 having further instruction stored thereon, that when executed, cause the one or more p... | Intel Corporation | H03M007/34 H03M007/30 H04L001/00 A61B005/00 A61B005/024 A61B005/0402 A61B005/0476
US10722263 | B2 | 2020-07-28 | Biopsy needle and photoacoustic measurement apparatus | 2017-09-21 | In a biopsy needle having a hollow tubular outer needle, an inner needle that is disposed in a hollow portion of the outer needle so as to be movable in the tube axis direction relative to the outer needle, and a recessed sample collection portion that is cut inward from a circumferential surface of the inner needle, an outer groove extending in the tube axis direction of the outer needle is provided on at least the distal end side and the rear end side of the sample collection portion in the inner needle, and a light guide member is disposed in the outer groove. At least a part of the light guide member is fixed by fillers filled in the outer groove. | 1. A biopsy needle, comprising: a hollow tubular outer needle;an inner needle that is movable in a tube axis direction relative to the outer needle and that is disposed in a hollow portion of the outer needle;a sample collection portion that has a recessed shape and that is provided on a circumferential surface of the inner needle;an outer groove that is provided on each of at least a distal end side and a rear end side of the sample collection portion in the inner needle and that is opened to the circumferential surface of the inner needle and extends in the tube axis direction; wherein each of the outer grooves is open to the sample collection portion respectively,a first light guide member that is disposed in each of the outer grooves and in the sample collection portion, and that is disposed from a rear end side of the inner needle to vicinity of a distal end of the inner needle;a first light absorber that absorbs light emitted from a distal end of the first light guide member and emits a photoacoustic wave; anda filler that is filled in the outer groove on at least one of on an inner needle rear end side or on an inner needle distal end side to fix at least a part of the first light guide member. 2. The biopsy needle according to claim 1, wherein the first light absorber is disposed in the outer groove provided on the distal end side of the sample collection portion in the inner needle. 3. The biopsy needle according to claim 2, wherein the first light guide member is fixed to a bottom surface of the sample collection portion by a fixing member that covers at least a part of the first light guide member. 4. The biopsy needle according to claim 3, wherein a bottom surface of the outer grooves and the bottom surface of the sample collection portion are located at the same height position. 5. The biopsy needle according to claim 2, wherein a bottom surface of the outer grooves and a bottom surface of the sample collection portion are located at the same height position. 6. The biopsy needle according to claim 2, wherein a bottom surface of the sample collection portion is located at a position higher than a bottom surface of the outer grooves. 7. The biopsy needle according to claim 1, wherein the first light guide member is fixed to a bottom surface of the sample collection portion by a fixing member that covers at least a part of the first light guide member. 8. The biopsy needle according to claim 7, wherein a bottom surface of the outer grooves and the bottom surface of the sample collection portion are located at the same height position. 9. The biopsy needle according to claim 7, wherein the bottom surface of the sample collection portion is located at a position higher than a bottom surface of the outer grooves. 10. The biopsy needle according to claim 1, wherein a bottom surface of the outer grooves and a bottom surface of the sample collection portion are located at the same height position. 11. The biopsy needle according to... | FUJIFILM Corporation | A61B017/34 A61B010/02 A61B005/00 A61B008/00
US10709376 | B2 | 2020-07-14 | System for supporting an elderly, frail and/or diseased person | 2016-05-26 | The present invention relates to a system (10) for supporting an elderly, frail and/or diseased person (12), in particular a person suffering from Parkinson's disease, wherein the system (10) comprises: a detection unit (14) including (i) a brain activity sensor (20) for detecting a brain activity signal relating to the brain activity of the person (12) and (ii) a motion detection unit (22) for detecting a motion signal relating to a motion of one or more body parts of the person (12); an analysis unit (16) for determining, based on the detected brain activity signal and motion signal, an activity level of the person (12) which is indicative of the motoric and cognitive activity of the person (12); and a feedback unit (18) for providing a feedback to the person (12) if the activity level of the person (12) exceeds a predetermined threshold. | 1. A system for supporting an elderly, frail and/or diseased person, wherein the system comprises: a detection unit including (i) a brain activity sensor configured to detect a brain activity signal relating to brain activity of the person and(ii) a motion detection unit configured to detect a motion signal relating to a motion of one or more body parts of the person;an analysis unit configured to determine, based on the detected brain activity signal and the detected motion signal, an activity level of the person which is indicative of how many different motoric and cognitive tasks the person is performing simultaneously, anddetermine, based on the detected brain activity signal and the detected motion signal, whether the person is walking; anda feedback unit configured to provide a feedback to the person the person is walking and the activity level of the person exceeds a predetermined threshold. 2. The system of claim 1, wherein the brain activity signal includes a plurality of different signal components each relating to the brain activity in a different region of the brain of the person, and the analysis unit is further configured to determine the activity level of the person by analyzing the different signal components in order to estimate how many different motoric and cognitive tasks the person is performing simultaneously. 3. The system of claim 2, wherein the analysis unit is further configured to consider the activity level of the person as exceeding the predetermined threshold andsteer the feedback unit to provide the feedback after a first one of the different signal components which relates to the brain activity in a first region of the brain of the person indicates a brain activity in said first region above a first predetermined brain activity level. 4. The system of claim 2, wherein the analysis unit is further configured to consider the activity level of the person as exceeding the predetermined threshold and steer the feedback unit to provide the feedback after a first one of the different signal components which relates to the brain activity in a first region of the brain of the person indicates a brain activity in said first region above a first predetermined brain activity level and after simultaneously a second one of the different signal components which relates to the brain activity in a second region of the brain of the person indicates a brain activity in said second region above a second predetermined brain activity level. 5. The system of claim 1, wherein the brain activity sensor is an EEG sensor and the brain activity signal is an EEG signal, and wherein the analysis unit is further configured to determine the activity level of the person by analyzing the EEG signal in one or more ranges of the frequency spectrum of the EEG signal in the frequency domain. 6. The system of claim 5, wherein the analysis unit is further configured to analyze at least one of (i) a signal power in the whole frequency spectrum of the... | KONINKLIJKE PHILIPS N.V. | A61B005/00 A61B005/048 A61B005/11 A61B005/16 G06F019/00 G16H020/30 A61B005/0488
US10709350 | B2 | 2020-07-14 | Body impedance measuring device | 2016-03-10 | A device for impedance measurements on body segments. Outputs of a current source are connected to inputs of current changeover switches. Each current changeover switch has outputs that are interconnectable with the input of the respective current changeover switch. The outputs of the current changeover switches are connected to electric feed lines of current electrodes. Measuring electrodes detect voltage signals, which are used with the current supplied by the current source for the impedance measurements of the body segments. Switching devices are incorporated into the electric feed lines near the current electrodes. By switching the current changeover switches and switching devices, the device respectively connects only two current electrodes to the current source and disconnects all other current electrodes from the electric feed line by switching the switching device and the electric feed line from the current source by switching the current changeover switches. | 1. Device for impedance measurements on segments of a human or animal body, the impedance measurements being multi-frequency impedance measurements, wherein the device comprises a current source, current changeover switches, current electrodes placeable on segments of the human or animal body, voltage changeover switches and measuring electrodes, wherein each output of the current source is connected to an input of a current changeover switch, wherein each current changeover switch has a plurality of outputs and the input of the respective current changeover switch is switchably interconnectable with one of its outputs, wherein each of the outputs of the current changeover switches is connected by an electric feed line to one of the current electrodes, wherein each voltage changeover switch has a plurality of inputs and an output and the output of the respective voltage changeover switch is switchably interconnectable with one of its inputs, wherein each of the inputs of the voltage changeover switches is connected by a signal line to one of the measuring electrodes, wherein the measuring electrodes detect voltage signals and the device determines the impedances of the segments of the human or animal body from the current supplied by the current source and the detected voltage signals, wherein switching devices are incorporated into each of the electric feed lines at the current electrodes, with which switching devices each of the current electrodes may be connected to or disconnected from its electric feed line, wherein the device is configured to always open up only one measuring section via the voltage changeover switches and respectively connect only two current electrodes to the current source by switching the current changeover switches and the switching devices, while all other current electrodes are disconnected from their electric feed lines by switching the switching devices, and to disconnect the electric feed lines from the current source by switching the current changeover switches. 2. Device according to claim 1, wherein voltage amplifiers or voltage followers with a high input resistance are integrated into the signal lines at the measuring electrodes. 3. Device according to claim 1, wherein the switching devices are switchable by the device. 4. Device according to claim 1, wherein the outputs of the voltage changeover switches are connected to inputs of a differential amplifier, which, based on the voltage signals applied to its inputs, determines a differential voltage (Udiff), which the device uses for the determination of impedance. 5. Device according to claim 1, wherein a compensating current may be applied to at least one of the current electrodes or an additional electrode via a resistor, wherein the connection of the resistor to the current electrode is positioned between the switching device and the current electrode. 6. Device according to claim 1, wherein a compensating current electrode is provided, to which a com... | Falko Skrabal | A61B005/053 A61B005/0428 A61B005/0205 A61B005/0408 A61B005/00 A61B005/024
US10722117 | B2 | 2020-07-28 | Systems and methods for storing data on medical sensors | 2018-12-17 | A system includes a monitor configured to be coupled to multiple sensors over time to monitor one or more physiological characteristics of one or more patients. A first sensor of the multiple sensors includes a sensor memory, and the monitor is configured to transfer information to the first sensor for storage in the sensor memory of the first sensor. The information includes information related to one or more sensors of the multiple sensors that were coupled to the monitor prior to the first sensor. | 1. A system, comprising: a monitor configured to be coupled to multiple sensors over time to monitor one or more physiological characteristics of one or more patients; anda first sensor of the multiple sensors, wherein the first sensor comprises a sensor memory, wherein the monitor is configured to transfer information to the first sensor for storage in the sensor memory of the first sensor, and the information comprises information related to one or more sensors of the multiple sensors that were coupled to the monitor prior to the first sensor. 2. The system of claim 1, wherein the information comprises a duration of operation of each sensor of the one or more sensors, an indication of compatibility between each sensor of the one or more sensors and the monitor, a number of alarms generated during operation of each sensor of the one or more sensors, a number of recorded reprocessing events for each sensor of the one or more sensors, a number of prior uses for each sensor of the one or more sensors, a level of ambient light present during operation of each sensor of the one or more sensors, or the one or more physiological characteristics of one or more patients obtained during operation of each sensor of the one or more sensors, or any combination thereof. 3. The system of claim 1, wherein the monitor is configured to automatically transfer the information to the first sensor in response to the first sensor being coupled to the monitor. 4. The system of claim 1, wherein the first sensor is configured to transfer first sensor information from the sensor memory to the monitor in response to the first sensor being coupled to the monitor, wherein the first sensor information comprises a serial number of the first sensor, calibration coefficients or information, a number of recorded reprocessing events for the first sensor, a number of prior uses of the first sensor, or a combination thereof. 5. The system of claim 4, wherein the first sensor is configured to erase the first sensor information from the sensor memory in response to transferring the first sensor information to the monitor. 6. The system of claim 5, wherein the monitor is configured to transfer the information to the sensor memory in response to the first sensor information being erased from the sensor memory. 7. The system of claim 1, comprising a sensor processing system configured to access the information in the sensor memory of the first sensor, wherein the sensor processing system comprises a processor configured to process the information to determine a clinical area of care in which the monitor was located during operation of the one or more sensors. 8. The system of claim 1, wherein the information comprises information related to the monitor. 9. The system of claim 1, wherein the monitor is configured to transfer first sensor information to the first sensor for storage in the sensor memory of the first sensor in response to an input indicating that the first sensor wil... | COVIDIEN LP | G08B021/00 A61B005/00 H04Q009/00 A61B005/1455
US10716473 | B2 | 2020-07-21 | Automated procedure-determination and decision-generation | 2017-03-22 | Methods and systems for automated monitoring are described herein. In various embodiments, image data of an enclosed area captured by an imaging device (102), audio data captured from the enclosed area by a sound sensor (104), and odor data captured from the enclosed area by an odor sensor (106), may be obtained by a computing device (HO). Based on the image, audio, and odor data, a determination (400) of an activity taking place within the enclosed area may be made. In various embodiments, one or more objects or persons in the enclosed area may be identified by performing image processing on the image data. The determined activity may be compared with corresponding data and parameters previously stored in a rules and schedules database (120). Based on the comparison and a decision generated by the rules and schedules database, an action may be selected to be performed. | 1. A method for automated monitoring comprising: obtaining, by one or more processors, image data captured of an enclosed area by an imaging device;obtaining, by the one or more processors, audio data captured from the enclosed area by a sound sensor and storing the captured sound data in a sounds database;obtaining, by the one or more processors, odor data captured from the enclosed area by an odor sensor and storing the captured odor data in an odor database;determining, by the one or more processors based on the image data, the audio data stored in the sounds database, and the odor data stored in the odor database, an activity taking place within the enclosed area, wherein one or more objects or persons are identified by performing image processing on the image data;comparing, by the one or more processors, the determined activity with corresponding data and parameters previously stored in a rules and schedules database; andselecting, by the one or more processors, an action to be performed based on the compared activity and a decision generated by the rules and schedules database. 2. The method of claim 1, further comprising: obtaining, by the one or more processors, motion data captured by a motion sensor, wherein the determining is further based on the motion data. 3. The method of claim 1, further comprising: identifying, by the one or more processors, motion of one or more persons in the enclosed area based on the image data captured by the imaging device. 4. The method of claim 3, wherein the determining is further based on the identified motion. 5. The method of claim 1, wherein after the activity violates one or more rules in the rules and schedules database, the action to be performed comprises: raising an alert to a caregiver that the patient has violated the one or more rules. 6. The method of claim 5, wherein the one or more rules comprise a rule that the patient is supposed to fast for a predetermined time interval. 7. The method of claim 5, wherein the one or more rules comprise a rule that the patient is supposed to be asleep during a predetermined time interval. 8. The method of claim 1, wherein the odor data comprises the smell of urine, feces, or vomit in the enclosed area. 9. The method of claim 8, wherein the action to be performed comprises: raising an alert to prompt a caregiver to change sheets or a diaper worn by a patient in the enclosed area. 10. The method of claim 1, wherein the odor data comprises the smell of smoke in the enclosed area. 11. The method of claim 1, further comprising: performing, by the one or more processors, image processing on the captured image data to identify one or more persons in the enclosed area. 12. The method of claim 11, wherein the image processing comprises identification of visual indicia worn by the one or more persons. 13. The method of claim 11, wherein the image processing comprises facial recognition of the one or more persons. 14. The method of claim 1, further com... | KONINKLIJKE PHILIPS N.V. | G16H050/20 A61B005/00 G16H040/20 G16H040/67 G06K009/00 G08B021/04
US10706717 | B2 | 2020-07-07 | Electronic device and control method thereof | 2019-06-27 | A device includes a first sensor to generate a movement signal of a user movement; a second sensor to generate a PPG signal of the user; and a processor configured to set each of time periods as a predetermined amount of time to determine one sleeping state of the user, set a first part of the predetermined amount of time in each time period to emit the light, and set a remaining part of the predetermined amount of time in each time period to not emit the light, control the light emitter of the second sensor to emit the light in the first part of each time period and to not emit the light in the remaining part of each time period, and determine, for each time period, the sleeping state of the user based on the movement signal and the PPG signal that are generated for each time period. | 1. An electronic device comprising: a first sensor for generating a movement signal corresponding to a user movement;a second sensor configured to physically contact a user to generate a photoplethysmography (PPG) signal, the second sensor comprising:a light emitter for emitting light for generating the PPG signal, anda light receiver for receiving light reflected from the user;a communicator; anda processor configured to: set a plurality of preset time periods which are consecutive time periods for analyzing a sleep of the user, each of the plurality of preset time periods being a predetermined amount of time to determine one sleeping state of the user,set a first part of the predetermined amount of time in each of the plurality of preset time periods to emit the light, and set a remaining part of the predetermined amount of time in each of the plurality of preset time periods to not emit the light,control the light emitter to emit the light in the first part of each of the plurality of preset time periods and to not emit the light in the remaining part of each of the plurality of preset time periods, anddetermine, for each of the plurality of preset time periods, the sleeping state of the user based on the movement signal and the PPG signal that are generated for each of the plurality of preset time periods, respectively,wherein the first part of the predetermined amount of time and the remaining part of the predetermined amount of time are variably set based on a charge state of a battery of the electronic device. 2. The electronic device as claimed in claim 1, wherein the processor is further configured to determine the sleeping state of the user as any one of an awakening stage, a non-rapid eye movement (NREM) sleeping stage, and a rapid eye movement (REM) sleeping stage. 3. The electronic device as claimed in claim 1, wherein the processor is further configured to smooth the PPG signal generated for one of the plurality of preset time periods using the PPG signal generated prior to the one of the plurality of preset time periods, and compare the smoothed PPG signal with a certain value to determine the sleeping state of the user in the one of the plurality of preset time periods. 4. The electronic device as claimed in claim 1, further comprising a touch screen for receiving a sleep analysis start command, wherein the processor is further configured to control the first sensor and the second sensor to generate the movement signal and the PPG signal, respectively, based on the sleep analysis start command being received. 5. The electronic device as claimed in claim 4, wherein the touch screen is further configured to receive a sleep analysis termination command, and calculate a sleep efficiency from a first time point at which the sleep analysis start command is received, to a second time point at which the sleep analysis termination command is received, based on the sleep analysis termination command being input. 6. The electronic devic... | SAMSUNG ELECTRONICS CO., LTD. | G08C019/16 G08C017/02 A61B005/024 A61B005/00
US10704973 | B2 | 2020-07-07 | Data acquisition device and method of timing data sampling | 2017-12-22 | A method of timing data sampling includes, in a data acquisition device, generating sampling intervals from a system clock of the data acquisition device, sampling data at the generated sampling intervals, and receiving start of frame (SOF) signals from a port, such as a USB port. For a selected number of SOF signals received, an actual number of system clock cycles is determined for a time interval corresponding to the selected number of SOF signals. The actual number of system clock cycles is compared to a nominal number of system clock cycles and a sampling interval is adjusted based on the comparison. The actual number of system clock cycles may be determined using a count-and-capture counter. The nominal number of system clock cycles may be calculated based on a nominal system clock rate and on a known SOF interval. Generating sampling intervals may include using a divide-by-N counter. | 1. A method of timing data sampling comprising, in a data acquisition device: generating sampling intervals from a system clock of the data acquisition device;sampling data at the generated sampling intervals;receiving start of frame signals from a port; andfor a selected number of start of frame signals received:determining an actual number of system clock cycles for a time interval corresponding to the selected number of start of frame signals;comparing the actual number of system clock cycles to a nominal number of system clock cycles; andadjusting a sampling interval based on the comparison. 2. The method of timing data sampling as claimed in claim 1, wherein determining the actual number of system clock cycles comprises using a count-and-capture counter. 3. The method of timing data sampling as claimed in claim 1, wherein the nominal number of system clock cycles is calculated based on a nominal system clock rate and on a known start of frame interval. 4. The method of timing data sampling as claimed in claim 1, wherein generating sampling intervals comprises using a divide-by-N counter. 5. The method of timing data sampling as claimed in claim 1, wherein the port is a USB port. 6. A data acquisition device comprising: a system clock;an interval generator that generates sampling intervals from the system clock;a data sampler that samples data at the generated sampling intervals;an input port that receives start of frame signals from a port; anda sampling interval adjuster configured to:determine an actual number of system clock cycles for a time interval corresponding to a selected number of start of frame signals;compare the actual number of system clock cycles to a nominal number of system clock cycles; andadjust length of the sampling intervals based on the comparison. 7. The device as claimed in claim 6, wherein the sampling interval adjuster includes a count-and-capture counter and the actual number of system clock cycles is determined using the count-and-capture counter. 8. The device as claimed in claim 6, wherein the nominal number of system clock cycles is calculated based on a nominal system clock rate and a known start of frame interval. 9. The device as claimed in claim 6, wherein the interval generator includes a divide-by-N counter and the sampling intervals are generated using the divide-by-N counter. 10. The device as claimed in claim 6, wherein the port is a USB port. | Advanced Mechanical Technology, Inc. | G01L005/16 A61B005/00 A61B005/103 A61B005/11 G01L005/161 G01L025/00 G01G019/44 G01G021/23
US10709333 | B2 | 2020-07-14 | Instrument for acquiring co-registered orthogonal fluorescence and photoacoustic volumetric projections of tissue and methods of its use | 2016-07-25 | Disclosed are instruments and methods for acquiring co-registered orthogonal fluorescence and photoacoustic volumetric projections of an interrogated object. In an embodiment, an instrument includes an imaging tank filled with a liquid coupling medium. An object positioning mechanism is configured to position the interrogated object inside the coupling medium and to rotate the interrogated object. An optical excitation unit that is fixed with respect to the tank is configured to induce both fluorescence and photoacoustic responses inside the interrogated object using the same optical excitation spectrum and the same irradiation pattern at the surface of the interrogated object. An array of unfocused photoacoustic transducers is fixed with respect to the tank, and each element of the array is configured to detect photoacoustic signals generated inside the interrogated object. The array is arranged such that a plane connecting the rotation axis and a central portion of the array cuts through the interrogated object separating the two sides irradiated by the optical excitation unit. An optical detector that is fixed with respect to the tank is configured to register planar projections of sources of fluorescence generated inside the object. A data acquisition unit is configured to synchronize acquisition of photoacoustic data, acquisition of fluorescence data, optical excitation, and rotational position of the interrogated object. | 1. An instrument for acquiring co-registered orthogonal fluorescence and photoacoustic volumetric projections, comprising: an imaging tank filled with liquid, the liquid comprising a coupling medium selected from the group consisting of: a single chemical, mixture, solution, suspension, emulsion or gel;an object positioning mechanism configured to position an interrogated object, or part thereof, inside the coupling medium in the tank and to rotate the interrogated object about a rotation axis relative to the tank;an optical excitation unit fixed with respect to the tank and configured to induce both fluorescence and photoacoustic responses inside the interrogated object using the same optical excitation spectrum and the same irradiation pattern at the surface of the interrogated object, with light-emitting outputs arranged so as to illuminate two opposite sides of the interrogated object at each of a plurality of rotational positions of the interrogated object;an array of unfocused photoacoustic transducers fixed with respect to the tank, each element of the array configured to detect photoacoustic signals generated inside the interrogated object, the array being arranged such that a plane connecting the rotation axis and a central portion of the array cuts through the interrogated object separating the two sides irradiated by the optical excitation unit;at least one optical detector fixed with respect to the tank, configured to register planar projections of sources of fluorescence generated inside the object, and arranged opposite to the array of photoacoustic transducers with respect to the interrogated object; and,a data acquisition unit configured to synchronize acquisition of photoacoustic data, acquisition of fluorescence data, optical excitation, and rotational position of the interrogated object;the instrument being further configured to amplify, filter, digitize, and transfer acquired data to a peripheral system for subsequent storage, processing, or visualization. 2. The instrument of claim 1 wherein the coupling medium has (a) specific acoustic impedance of 1.1×10{circumflex over ( )}6 to 1.8×10{circumflex over ( )}6 kg/(m{circumflex over ( )}2*s), and (b) optical absorption coefficient and effective optical attenuation coefficient of 0 to 0.2 cm{circumflex over ( )}−1 each for wavelengths employed by the optical excitation unit. 3. The instrument of claim 1 wherein the object positioning mechanism comprises a rotary stage that rotates the interrogated object, the rotary stage being operated in a continuous mode or a stepper mode with an encoder measuring angular position of the interrogated object at each of a plurality of photoacoustic and fluorescence data acquisition events. 4. The instrument of claim 1 wherein the object positioning mechanism incorporates an animal restrainer and a life support unit with a breathing unit, a gas anesthesia unit, or an intravenous delivery unit. 5. The instrument of claim 4 wherein the breathi... | - | A61B005/00 A61D003/00 A61D007/04 A61B005/145 A61B005/1455 A61B090/30 A61B090/00
US10720242 | B2 | 2020-07-21 | Systems and methods for modeling and predicting effective treatments for schizophrenia and other disorders | 2015-12-07 | A system and method for evaluating an effectiveness of a therapy for a psychological condition includes selecting a therapy to be analyzed relative to psychological pathology. The selected therapy is applied to a model of the psychological condition that includes hyperdopaminergia as a function. A response is determined using an output of the model of the psychological condition. The response is compared to a control to determine a wellness metric and a report is generated indicating an effectiveness of the therapy based on the wellness metric. | 1. A computer system comprising: a non-transitory memory having stored thereon a computational model for a psychological condition and a computational control model and having instructions for evaluating an effectiveness of a therapy for a psychological condition;a processor configured to access the memory and carry out steps comprising: (i) selecting a therapy to be analyzed for effectiveness for the psychological condition, the therapy represented by a virtual therapy;(ii) applying the virtual therapy selected in step (i) to the computational model of the psychological condition by introducing at least one simulated therapy effect to the model to simulate at least one physiological parameter, wherein the model includes hyperdopaminergia as a function;(iii) determining a response using the at least one simulated physiological parameter;(iv) comparing the response to the computational control model to determine a wellness metric; and(v) generating a report indicating an effectiveness of the therapy based on the wellness metric;wherein the wellness metric is calculated as:W⁢⁢M=1-(1DX×Δ⁢⁢PX);wherein WM is the wellness metric, Dx represents a deficit in response associated with the psychological condition at a selected biomarker condition, X, at baseline, compared with the control model, and ΔPX is given by:Δ⁢⁢PX=PX,s-PX,cPX,c;andwherein PX,s is the response from the computational model of the psychological condition at the selected biomarker condition, X, and PX,c is the response from the computational control model at the selected biomarker condition, X. 2. The computer system of claim 1 wherein the therapy is a drug and wherein the processor is further configured to apply a virtual drug representing the drug to the computational model of the psychological condition by introducing at least one simulated medication effect to the model to carry out step (ii). 3. The computer system of claim 1 wherein the computational model of the psychological condition is configured to cause the processor to model increases in dopamine levels as a decrement in gamma band response. 4. The computer system of claim 1 wherein the computational control model is configured to cause the processor to model increases in dopamine levels as an enhancement of gamma band activity. 5. The computer system of claim 1 wherein the processor is further configured to: use auditory steady state response (ASSR) tasks as a biomarker for the psychological condition;compare the ASSR tasks when the virtual therapy is applied to the computational model of the psychological condition to ASSR tasks of the computational control model; andindicate a positive effectiveness of the therapy if the ASSR tasks when the therapy is applied to the model of the psychological condition approach ASSR tasks of the control model. 6. The computer system of claim 5 wherein the response includes a power response and is determined for ASSR tasks at 20 Hz (P20,s), 30 Hz (P30,s), and 40 Hz (P40,s). 7. The ... | MCLEAN HOSPITAL CORPORATION | G16H050/50 A61B005/00 A61B005/0484 G16H050/20 A61B005/04
US10712421 | B2 | 2020-07-14 | MRI Apparatus with RF corrected each TR interval | 2014-12-10 | An MRI apparatus includes, a generating unit configured to generate radio frequency pulses applied in a pulse sequence; a sequence control unit configured to apply a radio frequency pulse related to acquisition of an image signal and a corrective radio frequency pulse during execution of one TR of a pulse sequence; and a calculation unit configured to measure the corrective radio frequency pulse and calculate a correction value for the radio frequency pulse. Based on the correction value, the generating unit corrects a radio frequency pulse related to acquisition of an image signal to be applied during a following TR later than a TR during which the corrective radio frequency pulse is measured. | 1. An MRI apparatus comprising: a generating unit configured to generate radio frequency (RF) pulses according to a pulse sequence, the pulse sequence including a repeating series for each of plural repetition time (TR) intervals;a sequence control unit configured to apply an imaging RF pulse acquiring an imaging signal and a corrective RF pulse for correction following the imaging RF pulse during each TR interval of the pulse sequence thereby applying plural sets of RF pulses, each set including an imaging RF pulse and a corrective RF pulse, during each of the plural TR intervals of the pulse sequence;a wave detector configured to detect the corrective RF pulse during each of the plural TR intervals; anda calculation unit configured to calculate a correction value based on the detected corrective RF pulse during each TR interval,wherein the generating unit is configured to correct an imaging RF pulse to be applied during a second and subsequent TR intervals of the plural TR intervals based on the correction value detected during earlier TR intervals in the pulse sequence. 2. The MRI apparatus according to claim 1, wherein the calculation unit calculates the correction value based on a corrective RF pulse applied when a spoiler gradient magnetic field is applied. 3. The MRI apparatus according to claim 1, wherein the calculation unit calculates the correction value based on a corrective RF pulse applied at a site outside an imaging region. 4. The MRI apparatus according to claim 1, wherein the corrective RF pulse applied to calculate the correction value is a part of an imaging sequence. 5. The MRI apparatus according to claim 1, wherein the corrective RF pulse is measured by a pickup coil, the pickup coil being a receiving coil configured to receive imaging RF pulses. 6. The MRI apparatus according to claim 1, wherein the corrective RF pulse is measured via a directional coupler configured to input the imaging RF pulses outputted from the generating unit. 7. The MRI apparatus according to claim 4 wherein a triangular wave is used as the corrective RF pulse applied to calculate the correction value. 8. The MRI apparatus according to claim 4, wherein a sinc wave is used as the corrective RF pulse applied to calculate the correction value. 9. The MRI apparatus according to claim 1, wherein the corrective RF pulse applied to calculate the correction value is outputted in an amplitude that includes a maximum value and minimum value of the imaging RF pulse. 10. The MRI apparatus according to claim 1, wherein the calculating unit calculates the correction value based on the corrective RF pulses which are different in amplitude among different TRs. 11. The MRI apparatus according to claim 1, wherein: the calculation unit calculates a correction value each time the corrective RF pulse is outputted; andaccording to the correction value, the generating unit generates a corrected imaging RF pulse by correcting the imaging RF pulse outputted in a T... | TOSHIBA MEDICAL SYSTEMS CORPORATION | G01R033/54 G01R033/48 G01R033/565 A61B005/055 G01R033/28 G01R033/58
US10702198 | B2 | 2020-07-07 | Wire-assembly apparatus for invasive biosensors | 2016-11-29 | In some aspects, an apparatus for a biosensor includes a sensor wire and a rigid member. The rigid member may be coupled to the sensor wire and include a contact surface. The contact surface may be sized to enable a suction head of a robotic placement device to create a vacuum seal on the contact surface for lifting the sensor wire and rigid member. | 1. A method of assembling a biosensor, comprising: connecting a sensor wire and a rigid member by connecting a coupling surface of the rigid member to a proximal end of the sensor wire, the coupling surface having an adhesive disposed thereon, and the sensor wire having a distal end for insertion into a patient's skin;creating, by a suction head of a robotic placement device, a vacuum seal on a smooth surface of the rigid member for grasping the rigid member, the smooth surface being on an opposing side of the coupling surface;positioning, by the robotic placement device and using the vacuum seal, the sensor wire and the rigid member on a printed circuit board by connecting the coupling surface of the rigid member to the printed circuit board; andapplying, by the robotic placement device, a force on the rigid member in a direction of the printed circuit board to cause a pressure-sensitive adhesive disposed on the rigid member to adhere the rigid member to the printed circuit board. 2. The method of claim 1, further comprising: coupling, at a connection point on the printed circuit board, the sensor wire to circuitry disposed on the printed circuit board,wherein positioning the sensor wire and the rigid member on the printed circuit board includes:positioning the proximal end of the sensor wire on the printed circuit board at the connection point; andpositioning the rigid member on the printed circuit board such that the connection point is between a section of the printed circuit board and the rigid member. 3. The method of claim 1, further comprising: applying a potting material proximate to the rigid member to create a moisture barrier at a connection point on the printed circuit board for the sensor wire and circuitry disposed on the printed circuit board. 4. The method of claim 1, further comprising disposing the printed circuit board within a housing of a wearable monitoring device, the housing having a substantially planar surface for positioning the wearable monitoring device on the patient's skin. 5. The method of claim 1, further comprising disposing a reactive material on the distal end of the sensor wire to react with glucose present in interstitial fluid beneath the patient's skin. 6. A device for a biosensor, comprising: a sensor wire; anda rigid member including a coupling surface having an adhesive disposed thereon for coupling the rigid member to a printed circuit board of the biosensor and for coupling the rigid member to a proximal end of the sensor wire, the rigid member further including a contact surface on an opposing side of the coupling surface, the contact surface sized to enable a suction head of a robotic placement device to create a vacuum seal on the contact surface. 7. The device of claim 6, wherein the adhesive is a pressure-sensitive adhesive. 8. The device of claim 6, wherein the sensor wire includes one or more electrodes, wherein the rigid member is oriented on the printed circuit board to electrically co... | DexCom, Inc. | A61B005/00 A61B005/1473 A61B005/145
US10725136 | B2 | 2020-07-28 | Cerebrospinal diffusion phantom | 2017-12-21 | A cerebrospinal diffusion phantom including a housing having a shape and size configured for insertion into a magnetic resonance coil in one or more preselected poses. A scaffold support structure is mounted on an interior of said housing and a plurality of elongated diffusion mimicking members supported on the support array. The elongated diffusion mimicking members are affixed to the scaffold support structure such that elongated diffusion mimicking members extend in directions needed to substantially emulate a 3 dimensional arrangement of cerebrospinal diffusion fiber tracts in a living organism; modules for elimination of resolution based-bias, angular accuracy evaluation, diffusion rate calibration, and quality assurance image referencing. Each elongated diffusion mimicking member includes an aqueous component which can undergo diffusion along the elongated diffusion mimicking member. The phantom includes a cerebrospinal tissue mimic matrix material contained in the housing enveloping the array of elongated diffusion mimicking members. | 1. A cerebrospinal diffusion phantom, comprising: a) a housing having a shape and size configured for insertion into a magnetic resonance coil in one or more preselected poses;b) a plurality of elongated diffusion mimicking members contained within the housing, the plurality of elongated diffusion mimicking members positioned such that the elongated diffusion mimicking members extend in directions needed to substantially emulate a three-dimensional (3D) arrangement of cerebrospinal diffusion fiber tracts in a living organism, each elongated diffusion mimicking member including a liquid component which can undergo diffusion along the elongated diffusion mimicking member; andc) a cerebrospinal tissue mimic matrix material contained and sealed in the housing enveloping the plurality of elongated diffusion mimicking members, the housing being made of a material whose magnetic susceptibility substantially matches that of the cerebrospinal tissue mimic matrix material. 2. The phantom according to claim 1, wherein the elongated diffusion mimicking members are comprised of a flexible fiber. 3. The phantom according to claim 2, wherein the flexible fiber includes a micro-rod bundle having an assembly of micro-rods with uniform and varying cross-sectional planform so as to generate micro-lumen structures that are filled with liquid solution. 4. The phantom according to claim 2, wherein the flexible fiber includes a sheet of fibers organized in a fanned or planer arrangement. 5. The phantom according to claim 1, wherein the elongated diffusion mimicking members are flexible pulled micro-rods having two or more channels formed in an outer surface thereof and extending along a length of the micro-rods to give a surface area that is higher compared to the flexible pulled micro-rods not having the two or more channels. 6. The phantom according to claim 5, wherein the two or more channels having a size selected to be of a size to be narrow in comparison to the distance liquid can diffuse on the timescale of a diffusion tensor imaging (DTI) protocol, thereby restricting the possible diffusion in all directions except for along the direction of the fiber. 7. The phantom according to claim 6, wherein the two or more channels range in size from about 0.5 to about 10 micrometers. 8. The phantom according to claim 5, wherein the one or more elongated diffusion mimicking members is a plurality of said flexible pulled micro-rods bundled together to form flexible micro-rod bundles. 9. The phantom according to claim 1, wherein said cerebrospinal mimic matrix material is any one or combination of a polyvinyl alcohol cryogel (PVA-C), solution, water, or a cross-linked polymer gel. 10. The phantom according to claim 9, wherein each elongated diffusion mimicking member is comprised of a plurality of flexible micro-rod elements, and wherein said aqueous component is a thin layer of liquid enveloping said flexible micro-rod elements such that when said phantom is place... | Synaptive Medical (Barbados) Inc. | G01R033/58 G09B023/30 G01R033/48 G01R033/563 A61B005/055
US10702354 | B2 | 2020-07-07 | RF tag marker, emission marker, and the detector | 2016-10-11 | An object of the present invention is to provide an emission marker for being indwelled in a lumen, the emission marker having: a base part; a coil wound around the base part; an emission part electrically connected to the coil; and a cover part for covering at least the base part and the coil, wherein the emission part has a light source that emits visible light. Additionally, the present invention relates to a detector for detecting, from outside of a lumen, an emission marker indwelled in the lumen, the detector having a detection unit having an antenna that acts with a coil of the emission marker. | 1. A detection system comprising an emission marker and a detector, wherein the emission marker is indwelled in an inner surface of a lumen of a digestive tract which is selected from an esophagus, a stomach, a duodenum, a small intestine, a large intestine, a colon, or a rectum; the emission marker comprising: a base part;a coil wound around the base part, wherein the coil receives an induced current from the detector outside the lumen of the digestive tract that is separate from the emission marker;an emission part electrically connected to the coil; anda cover part for covering at least the base part and the coil,wherein the emission part has a light source that emits visible light, andwherein the emission marker has a size enabling insertion into a forceps channel of an endoscope, andwherein the detector comprises a detection unit having an antenna that acts with the coil of the emission marker. 2. The detection system according to claim 1, wherein the light source of the emission marker is an LED. 3. The detection system according to claim 1, wherein the base part of the emission marker further has a perforation part. 4. The detection system according to claim 1, wherein the cover part of the emission marker is a polymer base agent. 5. A detector for detecting, from outside of a lumen, an emission marker of claim 1 indwelled in the lumen, the detector comprising a detection unit having an antenna that acts with a coil of the emission marker. 6. The detector according to claim 5, wherein the detector is capable of being inserted into a laparoscope port. 7. The detector according to claim 5, wherein at least the detection unit is formed to be high-temperature and high-pressure steam sterilizable. 8. The detection system according to claim 1, wherein the base part of the emission marker further has a memory. 9. An emission marker for being indwelled in an inner surface of a lumen of a digestive tract which is selected from an esophagus, a stomach, a duodenum, a small intestine, a large intestine, a colon, or a rectum, comprising: a base part;a coil wound around the base part, wherein the coil receives an induced current from outside the lumen of the digestive tract;an emission part electrically connected to the coil; anda cover part for covering at least the base part and the coil,wherein the emission part has a light source that emits visible light, andwherein the emission marker has a size enabling insertion into a forceps channel of an endoscope. | Osaka Prefectural Hospital Organization | A61B005/00 A61B090/00 A61B090/98
US10709351 | B2 | 2020-07-14 | Magnetic resonance imaging apparatus, magnetic resonance imaging method and magnetic resonance imaging system | 2018-03-23 | A magnetic resonance imaging apparatus according to an embodiment includes a display, an input interface, and processing circuitry. The display displays at least a locator image and a reference image. The input interface sets a region of interest on the locator image displayed on the display. The processing circuitry scans a subject to obtain three dimensional data, generates a locator image from the three dimensional data and displaying the locator image on the display, generates a reference image corresponding to the location of the region of interest and displaying the reference image on the display, and makes, when a size or position of the region of interest on one of the locator image and the reference image is changed by the input interface, adjustments to correspondingly change the display magnification or position of the other one of the locator image and the reference image. | 1. A magnetic resonance imaging apparatus, comprising: a display configured to display at least a locator image and a reference image;an input interface configured to set a region of interest on the locator image displayed on the display; andprocessing circuitry configured toobtain three dimensional data of a subject,generate the locator image from the three dimensional data and display the locator image on the display,generate the reference image corresponding to the location of the region of interest and display the reference image on the display,when the reference image is zoomed in or zoomed out by the input interface, adjust the size of the region of interest being set on the locator image so as to correspond to an area of the reference image, the area being displayed by the display after the reference image has been zoomed in or zoomed out, and,when the region of interest being set on the locator image is increased or reduced by the input interface, zoom in or zoom out the reference image so as to correspond to the region of interest having been increased or reduced. 2. The magnetic resonance imaging apparatus according to claim 1, wherein the processing circuitry further scans a part of the subject corresponding to the region of interest to obtain image data. 3. The magnetic resonance imaging apparatus according to claim 1, wherein the processing circuitry generates a multi-plane reconstruction (MPR) image at a predetermined position and orientation from the three dimensional data, as the locator image. 4. The magnetic resonance imaging apparatus according to claim 1, wherein the processing circuitry generates a volumetric rendering image rendered from a predetermined orientation from the three dimensional data as the locator image. 5. The magnetic resonance imaging apparatus according to claim 2, wherein the processing circuitry applies a local excitation scan to obtain two dimensional or three dimensional data of the part corresponding to the region of interest from the subject. 6. The magnetic resonance imaging apparatus according to claim 1, wherein the processing circuitry obtains volumetric data as the three dimensional data. 7. The magnetic resonance imaging apparatus according to claim 1, wherein the processing circuitry obtains a multi-slice image as the three dimensional data. 8. A magnetic resonance imaging apparatus, comprising: a display configured to display at least a locator image and a reference image;an input interface configured to set a region of interest on the locator image displayed on the display; andprocessing circuitry configured toobtain three dimensional data of a subject,generate the locator image from the three dimensional data and display the locator image on the display,generate the reference image corresponding to a location of the region of interest and display the reference image on the display,make, when a size or position of one of the region of interest on the locator image and the reference ima... | Canon Medical Systems Corporation | G01R033/20 A61B005/00 A61B005/055 G01R033/54 G01R033/483
US10716479 | B2 | 2020-07-21 | Signal synchronization device, as well as stethoscope, auscultation information output system and symptom diagnosis system capable of signal synchronization | 2017-11-29 | The invention provides a signal synchronization device, which obtains electrocardiographic, heart sound signal calibration factors by performing time domain and frequency domain transformations for an electrocardiographic signal and a heart sound signal, and thereby performs calibration for the electrocardiographic signal and the heart sound signal to synchronize the electrocardiographic signal and the heart sound signal in time domain, such that a diagnosis rate for cardiovascular disease is increased. Moreover, the invention further provides a stethoscope with signal synchronization processing, an auscultation information output system and a symptom diagnosis system capable of signal synchronization. | 1. A stethoscope using a signal synchronization device for auscultating a user, the stethoscope including: a pickup member for collecting a heart sound signal of the user;a sound guiding pipe for conducting a heart sound signal;a slave member sheathing the pickup member for contacting the user due to a motion of the pickup member;a heart sound transducer arranged on the sound guiding pipe, the heart sound transducer having a heart sound transduction portion, which is oriented to the pickup member for acquiring the heart sound signal;an electrocardio transducer arranged on the slave member, the heart sound transducer having an electrocardio transduction portion, which is exposed to a portion of the slave member contacting the user for acquiring the electrocardiographic signal;wherein the signal synchronization device is used to perform synchronous calibration for the electrocardiographic signal and the heart sound signal on the stethoscope, the signal synchronization device including:a signal receiving module for receiving the electrocardiographic signal acquired by the electrocardiographic signal transducer and the heart sound signal acquired by the heart sound transducer separately to generate a first electrocardiographic time domain information and a first heart sound time domain information separately, the first electrocardiographic time domain information comprising a first electrocardio generation time point of the electrocardiographic signal, the first heart sound time domain information comprising a first heart sound generation time point of the heart sound signal;a synchronization factor generation module for receiving the electrocardiographic signal and the heart sound signal synchronously to generate a mixed time domain information, and performing time domain and frequency domain transformations for the mixed time domain information to obtain a second electrocardiographic time domain information representing the electrocardiographic signal in the mixed time domain information, the second electrocardiographic time domain information comprising a second electrocardio generation time point of the electrocardiographic signal, and to obtain a second heart sound time domain information representing the heart sound signal in the mixed time domain information, the second heart sound time domain information comprising a second heart sound generation time point of the heart sound signal, and allowing the second electrocardio generation time point and the second heart sound generation time point to be synchronous in time domain by using the second electrocardiographic time domain information as an electrocardiographic signal calibration factor, and using the second heart sound time domain information as a heart sound signal calibration factor, wherein the time domain synchronization is performed by adjusting time axis; anda synchronization module used to calibrate the electrocardiographic signal and the heart sound signal received by the signal re... | NATIONAL CHENG KUNG UNIVERSITY | A61B005/0205 A61B005/04 A61B007/04 A61B005/044 A61B005/00 A61B005/0402 A61B005/0452
US10726956 | B2 | 2020-07-28 | Fever epidemic detection system and method thereof | 2016-04-15 | The fever epidemic detection system comprises a detection module, a control module and a communication module. The detection module measures and obtains a body-temperature measured value. The control module comprises a first operation unit, a determination unit and an alert unit. The first operation unit receives and calibrates the body-temperature measured value with a calibration factor, and generates a body-temperature calibrated value. The determination unit receives and determines whether the body-temperature calibrated value is within a preset normal body-temperature range and generates a determination result. The alert unit receives the determination result. The communication module transmits data to an external device. The alert unit generates a first alerting message, if the determination result shows that the body-temperature calibrated value is not within a preset normal body-temperature range. The control module transmits the first alerting message and the corresponding body-temperature to the external device via the communication module. | 1. A fever epidemic detection method, used in a fever epidemic detection system for automatically and widely detecting fever epidemic in a region via Internet of Things (“IoT”), comprising: calculating a number of people passing through the region during a predetermined time interval;detecting temperature of the region in the predetermined time interval to obtain an environment average temperature and calculating an environment average temperature parameter according to the environment average temperature;using a thermal imager to measure temperature of each person in the region to obtain a measured temperature;calibrating the measured temperature with a calibration factor to calculate a calibrated temperature, wherein the calibration factor is with respect to a type of the thermal imager;determining whether the calibrated temperature is within a predetermined normal temperature range so as to generate a determination result;generating a first alerting message and outputting the first alerting message and a corresponding calibrated temperature if the determination result indicates that the calibrated temperature is not within the normal temperature range, and transmitting the first alerting message and the corresponding calibrated temperature to an external device;counting a number of the first alerting messages accumulated during the predetermined time interval and calculating a ratio of the number of the first alerting messages to the number of people passing through the region during the predetermined time intervalgenerating a third alerting message once the ratio is larger than a predetermined ratio; andtransmitting the third alerting message, the ratio, the corresponding number of the first alerting messages, and the number of people to the external device;wherein the third alerting message indicates that among people who have passed through the region during the predetermined time interval, the people whose temperatures are abnormal are equal to or more than a certain percentage of the people who have passed through the region. 2. The fever epidemic detection method according to claim 1, further comprising: calculating a group average temperature in every predetermined time interval according to the measured temperature, and calculating a group average temperature parameter according to the group average temperature; andadjusting the normal temperature range according to the environment average temperature parameter and the group average temperature parameter. 3. The fever epidemic detection method according to claim 1, further comprising: wherein a second alerting message is generated once the number of the first alerting messages is larger than a predetermined message number, and the second alerting message and the number of the first alerting messages are outputted. | NATIONAL TAIWAN UNIVERSITY | G16H050/80 G16H040/40 G06Q050/22 A61B005/01 A61B005/00
US10722170 | B2 | 2020-07-28 | Closed-loop glucose and/or insulin control system | 2017-06-01 | A method and/or system may be provided for determining a suggested change in a recommended therapy for a patient based, at least in part, on sensor measurements, and generating an alert to an attendant in a hospital environment upon detection of the suggested change. A method and/or system may also be provided to automatically determine a maximum interval to alert an attendant following receipt of a measurement at an operator interface. A method and/or system may also be provided to blood-glucose sensor calibration. | 1. A method comprising: receiving a signal representative of a measurement value entered at an operator interface;receiving signals representative of blood glucose sensor measurements obtained subsequent to receipt of the signal representative of the measurement value; andexecuting instructions on a special purpose computing apparatus to:compute one or more commands to a delivery system based, at least in part, on the received signals representative of the blood glucose sensor measurements; anddetermine a maximum duration following entry of the measurement value until initiation of an alert to an operator based, at least in part, on at least one of the one or more computed commands. 2. The method of claim 1, wherein the maximum duration to be based, at least in part, on the measurement value. 3. The method of claim 1, and further comprising executing the instructions on the special purpose computing apparatus to determine the maximum duration based, at least in part, on a signal representative of a measured rate of change in blood glucose of a patient. 4. The method of claim 1, wherein the one or more computed command to the delivery system comprise one or more proportional, plus integral, plus derivative (PID) commands based, at least in part, on the signals representative of the blood glucose sensor measurements. 5. The method of claim 1, and further comprising executing the instructions on the special purpose computing apparatus to: determine the maximum duration based, at least in part, on whether a glucose bolus was infused to a patient contemporaneously with receipt of the signal representative of the measurement value. 6. The method of claim 1, and further comprising further executing the instructions on the special purpose computing apparatus to determine the maximum duration based, at least in part, on one or more signals representative of a measured rate of change in blood glucose of a patient. 7. The method of claim 1, wherein the entered measurement value to comprise a blood glucose sample measurement value. 8. The method of claim 1, and further comprising generating the alert as a callback alert message to an attendant. 9. The method of claim 8, and further comprising shortening the maximum duration based, at least in part, on a forecasted blood glucose level, proximity of a blood glucose level to limits of a target blood glucose range or proximity of the blood glucose level to hypoglycemic or hyperglycemic warning limits, or a combination thereof. 10. An apparatus comprising: an operator interface to receive entry of an input value representative of a measurement value;a blood glucose sensor to provide signals representative of blood glucose sensor measurements obtained from a patient;andone or more processors to:compute one or more commands to a delivery system based, at least in part, on signals representative of blood glucose sensor measurements obtained from the blood glucose sensor subsequent to entry of the input valu... | Medtronic Minimed, Inc. | A61B005/00 A61M005/172 A61B005/145 A61M005/142 A61M005/158
US10716518 | B2 | 2020-07-21 | Blood pressure estimation by wearable computing device | 2017-02-27 | According to one aspect of the present disclosure, a method for estimating blood pressure is provided, comprising training a machine learning model on a cohort data set. The cohort data set may include subject-specific contextual data, time-varying features, and blood pressure measurements for a plurality of subjects. The method may include receiving contextual data for a specific subject, wherein the contextual data includes medical history data of the subject. The method may further include personalizing the machine learning model to the subject based on the contextual data. The method may include calibrating the machine learning model to the subject based on a set of time-varying features and blood pressure measurements of the subject. In addition, the method may include using the machine learning model and the time-varying features for the subject to generate a blood pressure estimate. | 1. A method for estimating blood pressure, comprising: training a machine learning model on a cohort data set, wherein the cohort data set includes subject-specific contextual data, time-varying features, and blood pressure measurements for a plurality of subjects;receiving contextual data for a specific subject, wherein the contextual data includes medical history data of the subject;personalizing the machine learning model to the subject based on the contextual data;detecting a pulse pressure wave signal of the subject over time using a pulse pressure sensor of a wearable sensing device;calibrating the machine learning model to the subject based on a set of time-varying features for the subject including the pulse pressure wave signal detected during an initial period of time, and blood pressure measurements of the subject; andusing the machine learning model following calibration and the time-varying features for the subject including the pulse pressure wave signal detected during a subsequent period of time to generate a blood pressure estimate. 2. The method of claim 1, wherein the contextual data of the subject includes information from the subject's medical history data selected from the set consisting of demographic information, comorbidities, past and present medications, vital signs, laboratory test results, recent weight change, echocardiogram results, cardiovascular disease history, smoking history, and past and present pregnancy. 3. The method of claim 1, wherein the subject-specific contextual data is used to determine parameters of the machine learning model. 4. The method of claim 1, wherein the machine learning model is calibrated at least in part using a brachial cuff measurement. 5. The method of claim 1, wherein the machine learning model is calibrated at least in part by the subject going through a series of perturbations based on posture. 6. The method of claim 1, wherein the machine learning model is calibrated at least in part by the subject exercising at varying levels of intensity. 7. The method of claim 1, wherein the machine learning model is calibrated at least in part by the subject taking medication which changes the blood pressure of the subject. 8. The method of claim 1, wherein the blood pressure estimate is conveyed for display on a display device. 9. The method of claim 1, wherein the wearable sensing device includes an electrocardiograph (EKG). 10. The method of claim 1, wherein the pulse pressure sensor is configured to detect a pressure wave morphology of the pulse pressure wave signal, including pressure wave morphology metrics selected from a group consisting of augmentation index, maximum systolic slope, systolic rise time, ejection time, dicrotic notch height, dicrotic notch time, pulse pressure, reflected wave arrival time, and heart rate. 11. The method of claim 1, wherein the wearable sensing device includes both the pulse pressure sensor and an EKG, wherein: the wearable sensing device is c... | MICROSOFT TECHNOLOGY LICENSING, LLC | A61B005/00 A61B005/022 A61B005/021 A61B005/0245 A61B005/0402 A61B005/024
US10709383 | B2 | 2020-07-14 | Wrist-worn pulse transit time sensor | 2015-06-25 | A wrist-worn device heart-monitoring device is presented. The wrist-worn heart-monitoring device includes a radial tonometer configured to output a pressure signal indicating a pulse pressure wave at a user's wrist, two or more electrodes configured to output an electrical signal indicating a user's heart has been commanded to contract, and a microphone configured to output an audio signal indicating a closing of a user's aortic valve. The wrist-worn heart-monitoring device further includes a pulse transit time monitor configured to calculate a pre-ejection period of the user's heart based on at least the pressure, electrical, and audio signals, and calculate a pulse transit time based on at least the pre-ejection period, the pressure signal, and the electrical signal. | 1. A wrist-worn heart monitoring device, comprising: a primary device;an adjustable strap, the primary device and adjustable strap cooperatively forming a loop configured to be worn around a wearer's wrist, the loop having an inner surface facing the wearer's wrist and an outer surface facing away from the wearer's wrist;a first electrode on the inner surface of the loop and configured to contact a first measurement location on the wearer's skin on a first effective side of the wearer's heart;a second electrode on the outer surface of the loop at the adjustable strap and configured to contact a second measurement location on the wearer's skin on a second, opposite, effective side of the wearer's heart;a microphone positioned on the outer surface of the loop and configured to output an audio signal indicating a closing of the wearer's aortic valve;a radial tonometer configured to output a pressure signal indicating a pulse pressure wave at the wearer's wrist;a controller configured to measure an electrical potential between the first measurement location and the second measurement location; anda pulse transit time monitor configured to: receive a first electrical signal from the first and second electrodes, wherein the first electrical signal is an electrocardiograph (ECG);receive a first audio signal from the microphone;receive a first pressure signal from the radial tonometer;calculate a pre-ejection period (PEP) of the wearer's heart, wherein the PEP is calculated as an elapsed time between a first identified timepoint in the ECG indicating a command to contract a left ventricle and a second identified time point indicating a timing of an aortic valve opening,wherein a third identified timepoint in the first audio signal indicates a timing of an aortic valve closing,wherein a total aortic valve opening period is determined by subtracting a fourth identified timepoint in the first pressure signal indicating the aortic valve opening from a fifth identified timepoint in the first pressure signal indicating the aortic valve closing, andwherein the second identified time point is determined by subtracting the total aortic valve opening period from the third identified timepoint,calculate a first pulse transit time based on at least the pre-ejection period, the first pressure signal, and the first electrical signal;store the calculated pre-ejection period of the wearer's heart;calculate a radial pulse arrival time responsive to receiving a second electrical signal indicating the wearer's heart has been commanded to contract, and a second pressure signal indicating the pulse pressure wave at the wearer's wrist; andcalculate a second pulse transit time based on at least the radial pulse arrival time and the stored pre-ejection period without receiving a second audio signal indicating a closing of the wearer's aortic valve. 2. The wrist-worn heart monitoring device of claim 1, wherein the first electrode is configured to be in contact with an underside... | MICROSOFT TECHNOLOGY LICNESING, LLC | A61B005/00 A61B005/024 A61B005/0245 A61B005/02 A61B005/0205 A61B005/021 A61B005/022 A61B005/029 A61B005/0408 A61B005/107 A61B005/11 A61B007/04
US10716488 | B2 | 2020-07-21 | Imaging using gated elements | 2014-12-29 | Exemplary methods and systems may image an object of interest using one or more waveguide assemblies including gated elements. The gated elements may be configurable in a transmission state, a reception state, or a passive state. The exemplary methods and systems may deliver electromagnetic energy (e.g., microwave energy) to an object of interest using a gated element of a waveguide assembly configured in a transmission state and sample scattered field using a gated element of another waveguide assembly configured in the reception state while the remainder of gated elements are configured in the passive state. | 1. A method of imaging an object using microwave imaging, wherein the method comprises: providing a plurality of waveguide assemblies positioned about an object, wherein each waveguide assembly of the plurality of waveguide assemblies comprises: a waveguide structure configured to deliver and sample electromagnetic energy, andone or more gated elements, wherein each gated element of the one or more gated elements defines an aperture extending into the waveguide structure, wherein each gated element of the one or more gated elements is configurable in at least a transmission state, a reception state, and a passive state, wherein each gated element of the one or more gated elements is configured to allow electromagnetic energy to radiate from the waveguide structure through the aperture when in the transmission state, to allow electromagnetic energy to be received through the aperture into the waveguide structure when in the reception state, and to neither allow electromagnetic energy to radiate from the waveguide structure through the aperture nor allow electromagnetic energy to pass through the aperture into the waveguide structure when in the passive state;delivering electromagnetic energy using at least one waveguide assembly of the plurality of waveguide assemblies to irradiate the object resulting in scattered electromagnetic energy;sampling the scattered electromagnetic energy using at least one waveguide assembly of the plurality of waveguide assemblies; andreconstructing an image of the object based on the sampled scattered electromagnetic energy. 2. The method of claim 1, wherein, to sample the scattered electromagnetic energy using at least one waveguide assembly of the plurality of waveguide assemblies, a modulation signal is applied to an individual gated element of the one or more gated elements to configure the individual gated element into the reception state. 3. The method of claim 2, wherein the modulation signal has a lower frequency than the electromagnetic energy delivered using at least one waveguide assembly of the plurality of waveguide assemblies. 4. The method of claim 1, wherein sampling the scattered electromagnetic energy using at least one waveguide assembly of the plurality of waveguide assemblies comprises sampling the scattered electromagnetic energy using each gated element individually of the one or more gated elements of each waveguide assembly of the plurality of waveguide assemblies until each gated element of the one or more gated elements of the plurality of waveguide assemblies has been individually used to sample the scattered electromagnetic energy, wherein, to sample the scattered electromagnetic energy using each gated element individually, a selected gated element is configured in the reception state, the remaining gated elements of the one or more gated elements of the plurality of waveguide assemblies are configured in the passive state, and the waveguide assembly comprising the selected gated elem... | The University of Manitoba | A61B005/05 G01S013/89 G01N022/00 G01S013/00
US10716512 | B2 | 2020-07-21 | Bed with user occupancy tracking | 2018-09-25 | If a user enters a bed, a sensor in the pump of the bed can detect the user's presence in the bed. | 1. A method performed by data processing apparatuses, the method comprising: determining, by a computing device, that a mattress of a bed is unoccupied;receiving, from a mattress pump and at a computing device, an empty pump pressure reading, the empty pump pressure reading recording air pressure of the mattress when the mattress is not subject to pressure from a person;receiving, from the mattress pump and at a computing device, a stream of pump pressure readings, the pump pressure readings recording air pressure of the mattress when the mattress is subject to pressure from an external body;for each received pump pressure readings: calculating, by the computing device, a test value that includes the most recent pump pressure reading in the stream of pump pressure readings; anddetermining, by the computing device, if the test value is greater than a threshold value that is greater than the empty pump pressure reading; andresponsive to the computing device determining that the test value is greater than the threshold value, engaging a peripheral device. 2. The method of claim 1, wherein the test value is a trailing average pressure that represents the average of the N most recent pump pressure readings in the stream of pump pressure readings. 3. The method of claim 1, the method further comprising: identifying, based on the stream of pump pressure readings, a supplemental indication of presence in the bed; andwherein, responsive to the computing device determining that the test value is greater than the threshold value, engaging a peripheral device further comprises: combining the supplemental indication of presence in the bed with the determining that the difference is significant compared to the threshold value. 4. The method of claim 3, wherein the supplemental indication of presence in the bed comprises identifying biological activity by an occupant of the bed. 5. The method of claim 1, the method further comprising modifying the threshold value. 6. The method of claim 5, wherein the threshold value is modified to account for seasonality. 7. The method of claim 5, wherein the threshold value is modified to account for growth of the user. 8. The method of claim 1, the method further comprising updating a user-facing computer interface to show that a user is detected in the bed. 9. The method of claim 1, wherein engaging a peripheral device comprises engaging an under-bed lighting system. 10. The method of claim 1, wherein: the computing device comprises a controller having a computing processor and computing memory, the controller being physically couplable to the bed; andthe pump comprises: a motor configured to be energized in response to a command to increase pressure of the mattress; anda pressure transducer configured to sense air-pressure of the mattress;wherein the pump is in fluid communication with at least one air chamber of the mattress. 11. A controller configured to: determine that a mattress is unoccupied;receive, from a... | Sleep Number Corporation | A47C027/08 A61B005/00 F04D015/00 F04D027/00 G05B015/02
US10720044 | B2 | 2020-07-21 | Headwear device for detection and communication of information received from an ingestible device | 2019-01-24 | A headwear device for detecting an electrical signal generated by an ingestible event marker is disclosed. The headwear device includes a detection subsystem to receive an electrical signal generated by an ingestible event marker from a detection arrangement. A processing subsystem is coupled to the detection subsystem to decode the electrical signal. A radio subsystem is configured to transmit the decoded electrical signal to a wireless node. A system includes the headwear device and the detection arrangement. A method includes receiving the electrical signal generated by the ingestible event marker at the headwear device, decoding the electrical signal to extract information associated with the ingestible event marker, and transmitting the information to a wireless node. | 1. A headwear device for detecting an electrical signal generated by an ingestible event marker, the headwear device comprising: a detection subsystem to receive the electrical signal generated by the ingestible event marker, the detection subsystem comprising: a pair of electrodes arranged at opposite sides of the headwear device and configured to contact opposing skin surfaces of a head of a user when the headwear device is worn by the user, the pair of electrodes being configured to detect the electrical signal generated by the ingestible event marker; anda detection circuit module coupled to the pair of electrodes; andan antenna coupled to the detection subsystem and wirelessly couplable to a mobile device, the antenna configured to wirelessly transmit information associated with the ingestible event marker to the mobile device. 2. The headwear device of claim 1, wherein the detection circuit module is configured to detect information associated with the ingestible event marker encoded in the electrical signal generated by the ingestible event marker. 3. The headwear device of claim 1, wherein the detection circuit module is configured to be wirelessly powered by inductive coupling. 4. The headwear device of claim 1, wherein the detection circuit module is configured to be wirelessly powered in response to an interrogation signal received from the mobile device. 5. The headwear device of claim 4, wherein the detection subsystem is configured to initiate receipt of the electrical signal via the pair of electrodes and transmission of information associated with the ingestible event marker to the mobile device when powered in response to the interrogation signal. 6. The headwear device of claim 1, further comprising a battery configured to power the detection circuit module. 7. The headwear device of claim 1, wherein the antenna is configured to transmit the information associated with the ingestible event marker to the mobile device using Bluetooth communication. 8. The headwear device of claim 1, wherein the antenna is configured to transmit the information associated with the ingestible event marker to the mobile device using Wi-Fi communication. 9. The headwear device of claim 1, wherein the headwear device is selected from the group consisting of a helmet, a hat, and a visor. 10. The headwear device of claim 1, further comprising an annular support configured to be supported upon the head of the user, wherein the pair of electrodes are positioned at opposite sides of the annular support. 11. A headwear device for detecting an electrical signal generated by an ingestible event marker after ingestion by a user, the headwear device comprising: a first electrode;a second electrode;wherein the first electrode and the second electrode are arranged at opposite sides of the headwear device and configured to contact opposing skin surfaces of a head of the user to receive the electrical signal generated by the ingestible event marker;a dete... | PROTEUS DIGITAL HEALTH, INC. | G08C017/02 G16H040/67 A61B005/07 A61B005/00 G06F019/00 G16H030/20 G16H020/10
US10702224 | B1 | 2020-07-07 | Positioning guidance system For X-ray exams | 2020-02-14 | A system for assisting an x-ray operator with properly positioning a patient's body part to be x-rayed. The system uses a range sensor supported on an x-ray emitter to collect data about the patient's body part to be x-rayed. The collected data is transmitted to a processor and used to create a mapped envelope of the patient's body part. The processor compares the mapped envelope to a selected reference envelope. If the processor determines that the mapped envelope is not fully contained within the reference envelope, the processor will provide the x-ray operator with a negative notification, indicating that the patient's body part needs to be adjusted. If the processor determines that the mapped envelope is fully contained within the reference envelope, the processor will provide the x-ray operator with a positive notification, indicating the patient's body part is ready to be x-rayed. | 1. A system, comprising: an x-ray emitter having a field of view;a three-dimensional range sensor having a field of view that at least partially overlaps the field of view of the x-ray emitter; anda processor having a memory within which a plurality of reference envelopes are stored, the processor is in communication with the range sensor and is configured to use data from the range sensor to generate a virtual map that includes a body part within the field of view of the range sensor, to generate within the virtual map an envelope that closely surrounds the body part, and to compare the mapped envelope to a selected reference envelope. 2. The system of claim 1 in which the reference envelope is selected in response to input from a human user of the system. 3. The system of claim 1 in which the comparison performed by the processor determines whether the mapped envelope is fully contained within the selected reference envelope. 4. The system of claim 1, further comprising: a monitor in communication with the processor, the monitor configured to display, in real time and in superposition, representations of the body part within the field of view of the sensor and the selected reference envelope. 5. The system of claim 1, further comprising: an optical camera having a field of view that at least partially overlaps the field of view of the x-ray emitter. 6. The system of claim 5, further comprising: a monitor in communication with the processor, the monitor configured to display, in real time and in superposition, representations of the body part within the field of view of the optical camera and the selected reference envelope. 7. The system of claim 3, further comprising: a notification system controlled by the processor that sends a signal to a human user of the system if the mapped envelope is fully contained within the selected reference envelope. 8. The system of claim 7 in which the signal sent by the notification system is an audible signal. 9. The system of claim 7 in which the signal sent by the notification system is a visual signal. 10. The system of claim 9 in which the visual signal comprises a flashing light. 11. The system of claim 3, further comprising: a notification system controlled by the processor that sends a signal to a human user of the system if the mapped envelope is not fully contained within the selected reference envelope. 12. The system of claim 11 further comprising: a monitor in communication with the processor, the monitor configured to display, in real time and in superposition, representations of the body part within the field of view of the sensor and the selected reference envelope; andin which the signal sent by the notification system is displayed on the monitor. 13. The system of claim 1 further comprising: an interface in communication with the processor and comprising a menu listing a plurality of different body parts and a plurality of different x-ray views. 14. The system of claim 13 in which... | - | A61B006/00 A61B006/04 H04N005/32 G06F003/0482 A61B005/00
US10702166 | B1 | 2020-07-07 | Devices and methods for respiratory variation monitoring by measurement of respiratory volumes, motion and variability | 2014-04-07 | This invention is directed to devices and methods for assessing a patient. The devices have at least one impedance measuring element functionally connected to a programmable element, programmed to analyze an impedance measurement, and to provide an assessment of at least one respiratory parameter of the patient. Preferably the device includes electronics which aid in calibration, signal acquisition, conditioning, and filtering. | 1. A device for assessing a patient, the device comprising: a sensor array comprised of a set of impedance measurement electrodes adapted to be placed in one or more tetrapolar configurations on the chest of the patient for acquiring a bioimpedance signal from the patient;a programmable element coupled to the sensor array and having acquisition circuitry;wherein the programmable element: determines a Respiratory Variation Monitoring (RVM) calibration coefficient based on the patient's bioimpedance signal and demographic information of the patient;adjusts the acquisition circuitry based on the calibration coefficient;determines signal variation in the bioimpedance signal;adjusts the acquisition circuitry for the signal variation; andanalyzes the bioimpedance signal using the RVM calibration coefficient to determine a respiratory volume and/or respiratory flow parameter of the patient; anda visual and/or audible alarm coupled to the programmable element or a medical therapy system coupled to the programmable element, wherein the alarm or medical therapy system is activated upon the respiratory volume and/or respiratory flow parameter exceeding a predetermined threshold. 2. The device of claim 1, wherein the calibration coefficient is further based on at least one of: physiological measurements of the patient, measured ECG signals of the patient, baseline impedance levels of the patient, measurements from a spirometer, and measurements from a ventilator. 3. The device of claim 1, wherein the respiratory volume and/or respiratory flow parameter is associated with a measurement of one or more of the patient's respiratory rate, the patient's respiratory pressure, the patient's end tidal CO2, the patient's sublingual CO2, the patient's intensity of respiration, a measurement that assesses variability, variation, or complexity in at least one of the patient's respiratory rate, the patient's respiratory pressure, the patient's respiratory flow, a patient's end tidal CO2, the patient's sublingual CO2, the patient's intensity of respiration, a measurement of at least one of the shape of the patient's respiratory curve, change in the shape of the patient's respiratory curve, a respiratory curve based on the patient's inhaled volume, a respiratory curve based on the patient's exhaled volume, a respiratory curve based on the patient's inhaled pressure, a respiratory curve based on the patient's exhaled pressure, a respiratory curve based on the patient's inhaled flow, a respiratory curve based on the patient's exhaled flow, a respiratory curve based on motion of the patient's chest as measured by imaging, a respiratory curve based on motion of the patient's chest as measured by contact sensors placed on the chest, a respiratory curve based on motion of the patient's abdomen as measured by imaging, a respiratory curve based on motion of the patient's abdomen as measured by contact sensors placed on the abdomen, a respiratory curve based on motion of both the ... | Respiratory Motion, Inc. | A61B005/0205 A61B005/0402 A61B005/00 A61B005/113 A61B005/11 A61M016/00 A61M016/04 A61B005/1455 A61B005/087 A61B005/053 A61B005/08 A61B005/083 A61B005/091 A61B005/085 A61B005/024
US10716483 | B2 | 2020-07-21 | Method and system for vector analysis of electrocardiograms | 2016-09-16 | A method for vector analysis of an electrocardiogram for assessment of risk of sudden cardiac death includes receiving data about electrical activity of heart of a subject recorded on electrocardiogram device, generating a vector cardiogram based on the data, analyzing the vector cardiogram to determine arrhythmogenic right ventricular dys-plasia/cardiomyopathy to identify a presence of a micro-scar in a three-dimensional vector loop of the vector cardiogram, determining a risk of SCD for the subject based on the identification of the presence of a micro-scar, and storing the risk in a database. | 1. A method stored on a non-transitory medium and performed by a processor for three-dimensional vector analysis of an electrocardiogram in search for micro-scar (“bite”) in a three-dimensional vector loop for assessment of sudden cardiac death (SCD) risk in arrhythmogenic right ventricular dysplasia/cardiomyopathy based on presence of micro-scar (“bite”), the method comprising: receiving electrocardiogram (EKG) waveform data from twelve leads of EKG recording of a patient's heart activity as one of a real time recording, or from an EKG print out, and an EKG image;generating a three-dimensional vectorcardiogram (VCG) based on the received EKG waveform data by transformation of the twelve EKG leads by application of inverse orthogonal projection that converts a two-dimensional VCG into a three-dimensional VCG;analyzing the three-dimensional VCG to identify the presence of a micro-scar (“bite”) in the three-dimensional vector loop without time delay which is achieved by setting borders for a beginning of a Q part, an R part, and an S part, and an end of the S part, of the three-dimensional vector loop from the twelve EKG leads;determining the SCD risk in arrhythmogenic right ventricular dysplasia/cardiomyopathy for the patient based on the identification of the presence of micro-scar (“bite”) in the three-dimensional vector loop; andstoring, in a database, the SCD risk in arrhythmogenic right ventricular dysplasia/cardiomyopathy based on the presence of micro-scar (“bite”) in the three-dimensional vector loop. 2. The method according to claim 1, wherein analyzing the three-dimensional VCG comprises: searching for a plurality of micro-scars (“bites”) where each micro-scar from the plurality of micro-scars (“bites”) represents deviation from an ideal curve as observed in the three-dimensional vector loop in a pre-defined region of one of the S part and the R part, of the three-dimensional vector loop;when analyzing the S part of the three-dimensional vector loop to determine a deviation from expected trajectory, if a trajectory of the S part is out of range of preset limit of epsilon environment a deviation representing the presence of micro-scar (“bite”) is determined and the SCD risk in arrhythmogenic right ventricular dysplasia/cardiomyopathy is established and the search terminates, if not, the search continues for the R part of the three-dimensional vector loop for detection of the presence of micro-scar (“bite”);when deviation of the trajectory of the S part is determined to be within a defined epsilon environment, which is a preset limit stored in software code, analyzing process for the R part of the three-dimensional vector loop continues in search for the presence of micro-scar (“bite”) indicative of the SCD risk in arrhythmogenic right ventricular dysplasia/cardiomyopathy, which is done by searching for an rR′ of the R part as an abrupt change of direction in one of a negative rotation and a clockwise rotation of the three-dimensional vec... | Ivana I. Vranic | A61B005/04 A61B005/0456 A61B005/0472 A61B005/00
US10709378 | B2 | 2020-07-14 | Closed loop neural activity triggered rehabilitation device and method | 2016-10-10 | A closed loop, neural activity triggered rehabilitation device and method are provided for facilitating recovery of a patient from the effects of a sensory motor disability. The device includes a sensor system positionable adjacent the brain of the patient. The sensor system detects neural signals. A functional stimulation component is operatively connectable to at least one body part, such as a muscle or a nerve. The functional stimulation component stimulates the at least one body part in response to the neural signals detected. A sensory stimulation module is operatively connected to the patient to provide sensory feedback thereto. | 1. A closed loop, neural activity triggered rehabilitation device for facilitating recovery of a patient from the effects of a sensory motor disability, the patient including a brain and at least one muscle, the device comprising: a sensor system positionable adjacent the brain at a predetermined location, the sensor system configured to detect real-time neural signals at the predetermined location corresponding to an area of the brain used by the patient to move a limb having plurality of muscle sites contemporaneously with volitional commands of the brain by the patient;a control system operatively connected to the sensor system for receiving the real-time neural signals from the sensor system, the control system generating a plurality of unique, synchronized control signals which are based solely on the real-time neural signals detected by the sensor system and a predetermined set of patient parameters, without utilizing any neural signals of the patient obtained previous to and non-contemporaneously with the real-time neural signals, wherein the predetermined set of patient parameters is entirely independent of real time patient feedback; anda functional stimulation component operatively connectable to the control system and including a plurality of channels, each channel of the plurality of channels operatively connectable to a different muscle site of the limb and being adapted to receive a corresponding one of the plurality of unique, synchronized control signals and to stimulate a corresponding different muscle site in response to the corresponding one of the plurality of unique, synchronized control signals received from the control system. 2. The device of claim 1 further comprising a sensory stimulation module operatively connectable to the control system and to the patient to provide sensory feedback thereto. 3. The device of claim 2 wherein the sensory stimulation module includes a tongue stimulator operatively connectable to a tongue of the patient. 4. The device of claim 1 wherein the control system includes a central processing unit interconnecting the sensor system and the functional stimulation component, the central processing unit generating the control signals. 5. The device of claim 1 wherein the sensor system includes an electroencephalography sensor. 6. The device of claim 1 wherein: the control system includes a central processing unit;the sensor system includes a plurality of sensors;each channel of the functional stimulation component includes a muscle stimulator connectable to one of the different muscle sites of the limb; andthe central processing unit is operatively connected to each channel of the functional stimulation component; andwherein the device further comprises: a multichannel amplifier interconnecting each of the plurality of sensors to the central processing unit. 7. The device of claim 6 further comprising a sensory stimulation module operatively connected to the central processing unit and operati... | Wisconsin Alumni Research Foundation | A61B005/04 A61N001/36 A61B005/00 A61B005/0482 A61B005/0476
US10716727 | B2 | 2020-07-21 | Garment system including at least one muscle or joint activity sensor and at least one actuator responsive to the sensor and related methods | 2017-05-31 | Embodiments disclosed herein relate to a garment system including at least one muscle or at least one joint activity sensor, and at least one actuator that operates responsive to sensing feedback from the at least one muscle or the at least one joint activity sensor to cause a flexible compression garment to selectively compress against or selectively relieve compression against at least one body part of a subject. Embodiments disclosed herein also relate to methods of using such garment systems. | 1. A garment system, comprising: at least one flexible compression garment adapted to be worn on at least one body part of a subject, the at least one flexible compression garment defining an interior space adapted to receive the at least one body part;one or more sensors positioned and configured to sense at least one of a blood flow or a temperature of the at least one body part, the one or more sensors are further configured to output one or more sensing signals indicative of the blood flow or temperature;one or more actuators positioned relative to the at least one flexible compression garment and configured to cause the at least one flexible compression garment to selectively compress against or selectively relieve compression against the at least one body part; anda control system operably coupled to the one or more actuators and further operably coupled to the one or more sensors to receive the one or more sensing signals therefrom, the control system including control electrical circuitry configured to direct the one or more actuators to cause the at least one flexible compression garment to selectively compress against or selectively relieve compression against the at least one body part responsive to the one or more sensing signals from the one or more sensors. 2. The garment system of claim 1, wherein the one or more sensors include at least one of an electromyography sensor, a thermal sensor, an oxygenation sensor, a chemical sensor, or an acoustic sensor. 3. The garment system of claim 1, wherein the one or more sensors include one or more near infrared radiation sensors positioned and configured to detect a level of blood flow in the at least one body part. 4. The garment system of claim 3, wherein the one or more near infrared radiation sensors include one or more near infrared source-detector pairs. 5. The garment system of claim 3, wherein the control electrical circuitry is configured to direct the one or more actuators to selectively compress or selectively relieve compression against the at least one body part responsive to the level of blood flow being above or below a threshold level. 6. The garment system of claim 3, wherein the control electrical circuitry is configured to direct the one or more actuators to selectively compress or selectively relieve compression against the at least one body part intermittently, responsive to the level of blood flow being above or below a threshold level. 7. The garment system of claim 1, wherein the one or more sensors include a thermal sensor. 8. The garment system of claim 7, wherein the thermal sensor includes a passive infrared sensor positioned and configured to sense infrared radiation from the at least one body part indicative of a temperature of the at least one body part. 9. The garment system of claim 7, wherein the thermal sensor is positioned and configured to sense a temperature of the at least one body part, and wherein the control electrical circuitry is configured... | ELWHA LLC | A61H001/00 A61F005/01 A61B005/00 A61B005/11 A61H007/00 A61H011/00
US10709825 | B2 | 2020-07-14 | Wound exudate monitor accessory | 2016-11-16 | A stand-alone system for assessing wound exudates from the wound of a patient is described. The system contains functionality to detect, process and report various wound parameters. The system also may make treatment determinations based on these findings. The system may detect one or more physiological values of the wound exudates from the wound of the patient. The system may also compare detected physiological values to predetermined physiological values, in order to obtain a comparison result in real time. The system may include a processor (15) which provides an electronic signal based on the comparison result in which the electronic signal may corresponds to guidelines for treating the wound (3). The system described may be an accessory, which may be used on its own, or in conjunction with other wound treatment devices (9). | 1. A stand-alone system for assessing wound exudate from a wound of a patient, said system comprising: a sensor or detector for detecting one or more physiological values of the wound exudate from the wound of the patient;a processor for comparing the one or more detected physiological values to one or more predetermined physiological values to obtain a comparison result in real time, and for providing an electronic signal based on the comparison result, in which the electronic signal corresponds to guidelines for treating the wound;a compartment for receiving wound exudate from the wound of the patient;an exit tube, wherein the compartment is connected to the exit tube;a valve positioned relative to the exit tube to control removal of wound exudate from the compartment via the exit tube when the wound exudate reaches a set level in the compartment;a reservoir for receiving the removed wound exudate from the exit tube; anda transmitter to output the electronic signal for reception by a negative pressure wound therapy device which controls a flow of the wound exudate from the wound of the patient;wherein the sensor or detector is positioned on an exterior of the exit tube. 2. The system of claim 1, further comprising a flow disruption element for controlling a flow of wound exudate from the wound of the patient using the electronic signal, in which the flow disruption element is integrally arranged in the system. 3. The system of claim 1, wherein the exit tube comprises an obstruction. 4. The system of claim 1, further comprising an attachment element for attaching to a negative pressure wound therapy device and for enabling the one or more physiological values of the wound exudate to be detected while the negative pressure wound therapy device is in use. 5. The system of claim 1, in which the one or more physiological values include values of tone, volume, thermal mass, flow rate, or viscosity. 6. The system of claim 1, further comprising a memory containing the one or more predetermined physiological values. 7. The system of claim 6, in which the one or more predetermined physiological values are representative of previously obtained physiological values. 8. The system of claim 1, wherein the sensor or detector is positioned to detect wound exudate within the compartment and/or as the wound exudate is removed from the compartment. 9. The system of claim 1, wherein the sensor or detector is positioned on an exterior of the compartment. 10. The system of claim 1, wherein the valve is structured to empty the compartment of wound exudate in response to the wound exudate reaching the set level in the compartment. 11. A stand-alone system for assessing wound exudate from a wound of a patient, said system comprising: a wound treatment device comprising a compartment for receiving wound exudate from the wound of the patient, wherein the compartment comprises or is connected to a port or exit tube, anda valve positioned relative to the port or ... | ConvaTec Technologies Inc. | A61M001/00 A61B005/145 A61B005/1455 A61B005/00 A61B005/01
US10716516 | B2 | 2020-07-21 | Monitor recorder-implemented method for electrocardiography data compression | 2017-08-14 | A monitor recorder-implemented method for electrocardiography data compression and an electrocardiography monitor recorder with integral data compression are provided. A series of data items are obtained, each of the data items associated with a magnitude of an ECG signal sensed by a monitor recorder. A range is set for an initial one of the data items in the series. Each of the data items remaining in the series is processed, including: obtaining an estimation of probabilities of the data items appearing next to that data item in the series; dividing the further range into sub-ranges, each sub-range representing a fraction of the further range proportional to the probabilities of the next data items; selecting the sub-range corresponding to the data item next to that data item in the series; and representing the next data item by the selected sub-range in a non-volatile memory. | 1. A monitor recorder-implemented method for electrocardiography data compression, comprising the steps of: obtaining by a monitor recorder a series of data items, each of the data items associated with an electrocardiographic signal magnitude, the monitor recorder comprising a housing configured to fit within a receptacle on a a patch applied to a patient, an electrocardiographic front end circuit within the housing that is operable to sense electrocardiographic signals through electrocardiographic electrodes each positioned on one end of the electrode patch, a non-volatile memory within the housing, and a micro-controller within the housing that is interfaced to the electrocardiographic front end circuit and the non-volatile memory, wherein the data items are derived by the micro-controller from the sensed electrocardiographic signals;performing via the microcontroller of the monitor recorder a lossy compression of the data items by encoding each data item with a value, wherein one or more of the values comprise multiple numbers;setting by the micro-controller a range for one of the compressed data items in the series; andfurther compressing by the micro-controller each of the data items remaining in the series by converting each of the multi-number values to a single number value, comprising: obtaining an estimation of probabilities of the data items appearing next to that data item in the series;dividing the range into sub-ranges, each sub-range representing a fraction of the range proportional to the probabilities of the next data items;selecting the sub-range corresponding to the data item next to that data item in the series; andrepresenting the next data item by the selected sub-range in the non-volatile memory,wherein the lossy compression followed by the further compression increases data storage capacity of the non-volatile memory and reduces power consumption of the monitor recorder to extend wear time of the monitor recorder. 2. A method according to claim 1, further comprising: using a statistical model to obtain the estimations of the probabilities. 3. A method according to claim 2, wherein the statistical model uses an electrocardiography data sample to obtain the estimations of the probabilities. 4. A method according to claim 2, further comprising: for each of the data items, generating at least one table of the estimations of the probabilities. 5. A method according to claim 4, wherein the estimations of the probabilities of the next data items for one of the data items are calculated using at least two of the data items preceding the one data item in the series. 6. A method according to claim 5, wherein a number of the tables generated for each of the data items depends on a number of the preceding data items used in calculation. 7. A method according to claim 1, wherein the data items are represented in the non-volatile memory periodically, continually, or en masse upon a completion of the processing. 8. A method accor... | Bardy Diagnostics, Inc. | A61B005/0402 A61B005/00 A61B005/04 A61B005/0432 A61B005/0408 A61B005/0452 A61B005/11
US10722750 | B2 | 2020-07-28 | Method and apparatus for determining effect of training on improving fitness | 2013-03-15 | The present disclosure concerns determining physiological training effect of a physiological performance of a person by monitoring the performance using one or more performance-monitoring means in order to obtain performance data, and, according to one aspect of the invention, determining, using computing means capable of utilizing the performance data, a third training effect parameter describing a third physiological effect of the performance using a third determination method, the third physiological effect being a combination effect of the first and second physiological effects which are different from each other and are descriptive of different physiological effects of training, such as homeostatic disturbance and cumulative physiological load, respectively. | 1. A method of determining physiological training effect of a physiological performance of a person, comprising: monitoring the physiological performance of the person during an exercise session using one or more sensors including a heart rate sensor in order to obtain performance data;determining, using a computing unit capable of utilizing the physiological performance data, a first training effect parameter describing a first physiological peak training effect relating to an experienced maximum stress of the physiological performance using a first determination method, anda second training effect parameter describing a second physiological effect relating to cumulative physiological load of the physiological performance using a second determination method, and/ora third training effect parameter describing a third physiological effect of the physiological performance using a third determination method, the third physiological effect being a combination effect of the first physiological effect and second physiological effect and relating to recovery time; andstoring the first and second training effect parameters on a storage unit and displaying at least one of said first, second and third training effects indicating the first, second and third effect of the exercise session on improvement of at least one of: cardiorespiratory fitness and fatigue resistance. 2. The method according to claim 1, wherein the third training effect parameter is determined and the third training effect parameter describes a cumulative physiological training effect of the physiological performance. 3. The method according to claim 1, wherein the third training effect parameter is calculated at least partly directly on the basis of the physiological performance data using the third determination method. 4. The method according to claim 1, wherein the first physiological peak training effect of the physiological performance correlates with homeostatic disturbance experienced by the person during the physiological performance. 5. The method according to claim 1, wherein the first determination method comprises estimating oxygen intake during or after the physiological performance for determining the first training effect parameter. 6. The method according to claim 5, wherein the physiological performance data and at least one sensor used for estimating the oxygen intake comprise interbeat interval data and a heartbeat sensor. 7. The method according to claim 1, wherein the second determination method comprises estimating the cumulative physiological load of the performance. 8. The method according to claim 1, wherein the second determination method comprises using at least one of the following as the physiological performance data and the one or more sensors: heart beat frequency data measured using the heart rate sensor, energy consumption data measured using an energy consumption sensor, position or velocity data measured using a positioning sensor and/or veloci... | Amer Sports Digital Services Oy | A63B024/00 A61B005/22 G06Q050/22 A61B005/11 A61B005/024 A61B005/00
US10712851 | B2 | 2020-07-14 | Electronic device and method for controlling the same | 2018-08-06 | An electronic device may include an edge touch screen including a main display region and an edge display region extending from the main display region each including one or more of red pixels, near infrared ray pixels, and sensor pixels for detecting light with different wavelengths; and a controller configured to, drive the edge touch screen in response to a touch input for the edge display region being maintained for a set time by instructing at least one selected red pixel of the red pixels and at least one selected near infrared ray pixel of the near infrared ray pixels corresponding to a position of the touch input to emit light, and measure biometrics based on light amounts of light of different wavelengths received from at least one selected sensor pixel of the sensor pixels corresponding to the position of the touch input. | 1. An electronic device comprising: an edge touch screen including a main display region and an edge display region extending from the main display region, each of the main display region and the edge display region including one or more of red pixels, near infrared ray pixels, and sensor pixels for detecting light with different wavelengths, the edge display region including a left edge display region and a right edge display region extending from respective sides of the main display region; anda controller configured to, drive the edge touch screen to measure biometrics in the left edge display region and the right edge display region in response to maintaining a multi-touch input in both the left edge display region and the right edge display region for a set time by instructing at least one selected red pixel of the red pixels and at least one selected near infrared ray pixel of the near infrared ray pixels corresponding to a position of the multi-touch input to emit light, andmeasure the biometrics in the left edge display region and the right edge display region based on light amounts of light of different wavelengths received from at least one selected sensor pixel of the sensor pixels corresponding to the position of the multi-touch input. 2. The electronic device of claim 1, wherein the edge touch screen includes: a display panel including the red pixels and the near infrared ray pixels;a touch panel on the display panel, the touch panel including at least one touch sensor;a sensor panel including the sensor pixels; anda light blocking member configured to block part of the display panel from being visible to an outside of the electronic device. 3. The electronic device of claim 2, wherein the at least one selected near infrared ray pixel is in a region of the edge touch screen corresponding to the light blocking member. 4. The electronic device of claim 2, wherein the at least one selected near infrared ray pixel is in the edge display region. 5. The electronic device of claim 2, wherein the at least one selected near infrared ray pixel is in a portion of the main display region adjacent to the edge display region. 6. The electronic device of claim 2, wherein the sensor pixels are in a region of the edge touch screen corresponding to the light blocking member. 7. The electronic device of claim 2, wherein the sensor pixels are in the edge display region. 8. The electronic device of claim 2, wherein at least one of the sensor pixels extends to both the edge display region and a portion of the main display region adjacent to the edge display region. 9. The electronic device of claim 2, wherein the sensor pixels include: a first sensor pixel configured to detect light of a near infrared ray wavelength, anda second sensor pixel configured to detect light of a red wavelength, whereina size of the first sensor pixel is different from a size of the second sensor pixel. 10. The electronic device of claim 2, wherein the red pixels and th... | Samsung Electronics Co., Ltd. | G06F003/044 G06F003/041 G06F001/16 G06F003/0488 G06F021/32 H01L027/32 H04M001/02 A61B005/00 A61B005/024 A61B005/1455
US10702150 | B2 | 2020-07-07 | Systems and methods for detection of cognitive and developmental conditions | 2018-06-14 | Systems, devices, and methods are described for the assessment, screening, monitoring, or diagnosis of developmental or cognitive conditions, including autism spectrum disorders (ASD) by analysis of eye tracking data generated from feedback received as a result of display of specific predetermined visual stimuli to a subject or patient. Subsequent to a calibration phase, a testing procedure is performed by presenting predetermined stimuli (e.g., videos) to a subject via a display device. Eye tracking data (from the subject moving his or her eyes in response to predetermined movies or other visual stimuli) are collected. During the data collection period, the system periodically presents targets to reflexively elicit the subject's gaze. These data are used to later verify accuracy. Analysis of the subject's viewing patterns during these stimuli is used for the assessment, screening, monitoring, or diagnosis of developmental or cognitive conditions such as ASD. | 1. A system comprising: a sensor configured to detect ocular behavior of a subject in response to a stimulus, wherein the ocular behavior of the subject is with respect to a fixation target, the stimulus comprising: a display of a first visual image followed by a display of the fixation target; anda display of a second visual image following the display of the fixation target;at least one processor configured to determine whether the ocular behavior indicates any of a blink, saccade, or smooth pursuit of the subject with respect to the displayed fixation target; andwherein the ocular behavior of the subject is rejected for calibration purposes, at least in part, if the ocular behavior indicates any of a blink, saccade, or smooth pursuit during display of the fixation target. 2. The system of claim 1, wherein the fixation target is a still visual image. 3. The system of claim 1, wherein the fixation target is a dynamic video. 4. The system of claim 1, wherein the fixation target triggers reflexive ocular behavior of the subject. 5. The system of claim 1, wherein the ocular behavior of the subject is elicited without verbal instruction to the subject. 6. The system of claim 1, further comprising at least one processor that is configured to: receive data associated with ocular behavior from the sensor;identify a fixation from the data indicating ocular behavior of the subject;calculate a fixation location for the fixation; anddetermine whether the fixation location is proximal to a known target location for the displayed fixation target. 7. The system of claim 6, wherein the data indicating ocular behavior of the subject is rejected for calibration purposes if the fixation location is not proximal to the known target location. 8. The system of claim 6, wherein the at least one processor is further configured to receive an indication from an operator that identifies the fixation from an observation by the operator. 9. The system of claim 1, wherein data indicating ocular behavior of the subject is rejected for calibration purposes if the data does not indicate a fixation by the subject. 10. A method comprising: detecting ocular behavior of a subject in response to a stimulus, wherein the ocular behavior of the subject is with respect to a fixation target, the stimulus comprising: a display of a first visual image followed by a display of the fixation target; anda display of a second visual image following the display of the fixation target;determining whether the ocular behavior indicates any of a blink, saccade, or smooth pursuit of the subject with respect to the displayed fixation target; andrejecting the ocular behavior of the subject for calibration purposes, at least in part, if the ocular behavior indicates any of a blink, saccade, or smooth pursuit during display of the fixation target. 11. The method of claim 10, wherein the ocular behavior is detected by a sensor, the method further comprising: receiving data associated with ocula... | Children's Healthcare of Atlanta, Inc. | A61B003/14 A61B003/113 A61B003/00 A61B003/032 A61B005/16
US10722164 | B2 | 2020-07-28 | Method and system for detection and analysis of cognitive flow | 2017-03-22 | A system and method for detection and analysis of learner's/performers cognitive flow have been provided. The system is configured to measure the cognitive state of the performer, while they are performing tasks of various complexity levels, using physiological responses like brain activation, heart rate variability and galvanic skin response. The system uses a Bayesian network based framework to probabilistically evaluate the cognitive state of the learner from the difficulty levels of the tasks, IQ level of the learner and observations made using the physiological sensing. The system also measures the actual cognitive state using a questionnaire. The predicted cognitive state and the actual cognitive state is compared and based on the outcome of comparison a relevant step is taken. | 1. A method for detection and analysis of a learner's cognitive flow, the method comprising: asking the learner to perform a task, wherein the task is performed at a difficulty level;recording EEG, GSR and HRV parameters related to the task using a data capture tool;measuring brain activity of the learner by sensing an electroencephalogram (EEG) signal using an EEG device;measuring galvanic skin response (GSR) of the learner using a GSR sensing device;measuring stress level of the learner by sensing heart rate variability (HRV);measuring an actual cognitive state of the learner after the task using a questionnaire filled by the learner after performing the task;creating, by a processor, a Bayesian network framework, wherein the Bayesian network framework includes a plurality of evidence nodes, a plurality of sensor nodes and a plurality of state nodes and wherein the plurality of sensor nodes include the EEG, GSR and HRV parameters;providing intelligence quotient (IQ) level of the learner and the difficulty level of the task as inputs to the plurality of evidence nodes;predicting, by the processor, a cognitive state of the learner using results obtained from the plurality of sensor nodes and the inputs provided at the plurality of evidence nodes;comparing, by the processor, the predicted cognitive state with the actual cognitive state, wherein if the predicted cognitive state is same as the actual cognitive state then the processor validates the Bayesian network framework, elsethe processor resolves a conflict generated due to a difference between the predicted cognitive state and the actual cognitive state by reasoning between the nodes using probabilistic queries. 2. The method of claim 1, wherein the difficulty level is at least one of a high, medium and low level. 3. The method of claim 1 wherein the HRV is measured by sensing a photo-plethysmogram (PPG) signal using a pulse oximeter. 4. The method of claim 1, wherein the task include playing a modified Tower of London game on a console. 5. The method of claim 4, wherein the difficulty level of the modified Tower of London game is dependent on number of moves made by the learner, total number of disks and amount of empty space available. 6. The method of claim 1 further include analyzing mid-beta and theta frequency of the EEG signal of the learner. 7. The method of claim 1, wherein the questionnaire includes a set of questions related to flow and boredom states in the task. 8. A system for detection and analysis of a learner's cognitive flow, the system comprising: a console configured to be used by the learner to perform a task at a difficulty level;a data capture tool to capture EEG, GSR and HRV a plurality of parameters related to the task;an electroencephalogram (EEG) device for measuring brain activity of the learner, wherein the brain activity indicates a skill-challenge balance of the learner;a GSR sensing device for measuring galvanic skin response (GSR) of the learner, wherei... | Tata Consultancy Services Limited | A61B005/16 A61B005/00 A61B005/0205 A61B005/024 A61B005/053 A61B034/10 A61B005/048 G16H050/30 G16H010/20 G16H050/50 A61B005/0476
US10722138 | B2 | 2020-07-28 | Method and system for electron paramagnetic resonance | 2015-06-11 | Electron paramagnetic resonance (EPR) systems and methods for transcutaneous oxygen monitoring (TCOM) and subcutaneous oxygen monitoring (SCOM) are provided herein. Optionally, the EPR systems provided herein can be portable and/or handheld to facilitate EPR oximetry in clinical environments. | 1. A system for detecting an electron paramagnetic resonance (EPR) signal from an oxygen-sensitive EPR probe material, comprising: a spectrometer having a radio frequency (RF) source and being configured for generating, via said RF source, an RF pulse sequence selected for measurement of T2 relaxation times; anda housing containing a magnetic assembly, a resonator assembly and a coupling loop, wherein the magnet assembly is configured to apply a magnetic field to a region-of-interest comprising the EPR probe and wherein the resonator assembly is in operative communication with the RF source through the coupling loop and is configured to resonate at a predetermined radio frequency, the housing having a distal end positionable proximal to the EPR probe;said spectrometer being in operative communication with the resonator assembly through the coupling loop, wherein the spectrometer is configured to receive a signal from the coupling loop and to process the received signal to measure said T2 relaxation times. 2. The system of claim 1, wherein the EPR probe comprises an oxygen-sensitive EPR material embedded in an oxygen-permeable polymer. 3. The system of claim 2, wherein the EPR probe is adherable to a surface of a tissue of a subject. 4. The system of claim 3, wherein the spectrometer is further configured to obtain data corresponding to a transcutaneous oxygen concentration of the tissue from the signal. 5. The system of claim 3, wherein the distal end of the housing includes an indentation for accommodating a portion of the EPR probe. 6. The system of claim 1, wherein the EPR probe is implanted in a tissue of a subject. 7. The system of claim 6, wherein the spectrometer is further configured to obtain data corresponding to a subcutaneous oxygen concentration of the tissue from the signal. 8. The system of claim 1, wherein the housing is a handheld wand. 9. The system of claim 8, wherein the region-of-interest is located at least partially external to the handheld wand. 10. The system of claim 8, wherein the region-of-interest is located entirely external to the handheld wand. 11. The system of claim 10, wherein the region-of-interest is located approximately 3 mm from the handheld wand. 12. The system of claim 1, wherein a maximum size of the region-of-interest is approximately 0.5 mm×0.5 mm×0.5 mm. 13. The system of claim 1, wherein the magnetic assembly comprises a plurality of magnets, the magnets being arranged around the resonator assembly. 14. The system of claim 13, wherein the magnets are arranged in a ring around the resonator assembly. 15. The system of claim 1, wherein the magnet assembly is configured to generate a substantially uniform magnetic field in the region-of-interest. 16. The system of claim 15, wherein the substantially uniform magnetic field has a field inhomogeneity of approximately 1000 ppm or less across the entire region-of-interest. 17. The system of claim 16, wherein the predetermined radio frequency ... | Technion Research & Development Foundation Limited | A61B005/055 G01R033/38 G01R033/383 G01R033/60 A61B005/145 A61B005/00 G01R033/28 G01R033/44 G01R033/343
US10716481 | B2 | 2020-07-21 | Apparatus, systems and methods for monitoring and evaluating cardiopulmonary functioning | 2013-12-10 | Wearable apparatus for monitoring various physiological and environmental factors are provided. Real-time, noninvasive health and environmental monitors include a plurality of compact sensors integrated within small, low-profile devices, such as earpiece modules. Physiological and environmental data is collected and wirelessly transmitted into a wireless network, where the data is stored and/or processed. | 1. A system for evaluating the cardiopulmonary functioning of a person, the system comprising: a portable telecommunication device; anda sensor module configured to be worn by the person, the sensor module comprising: at least one auscultatory sensor that is pointed in a direction of the person when the sensor module is being worn, wherein the at least one auscultatory sensor is configured to sense physiological sounds from the person that are related to cardiopulmonary functioning;at least one external acoustical sensor pointing away from the person and towards an environment around the person, wherein the at least one external acoustical sensor is configured to sense external sounds from the environment;a signal processor configured to process signals from the at least one auscultatory sensor and the at least one external acoustical sensor to at least partially remove external sounds from the physiological sounds and generate new signals containing cleaner physiological information about the cardiopulmonary functioning of the person; anda transmitter in communication with the signal processor and configured to transmit the new signals from the signal processor to the portable telecommunication device;wherein the at least one auscultatory sensor, the at least one external acoustical sensor, the signal processor, and the transmitter combined operate with less than 100 mW of total operating power;wherein the portable telecommunication device comprises a user interface configured to display the cleaner physiological information about the cardiopulmonary functioning information for the person. 2. The system of claim 1, wherein the portable telecommunication device comprises at least one physiological sensor configured to sense one or more of the following types of physiological information generated by the person: pulse rate, blood flow, blood gas levels, body temperature, and blood pressure. 3. The system of claim 1, wherein the sensor module comprises an earpiece module and wherein the at least one auscultatory sensor is configured to sense physiological sounds generated from an ear region of the person, and wherein the physiological sounds comprise pulse rate and breathing rate. 4. The system of claim 1, wherein the sensor module further comprises a blood pressure sensor comprising an actuator and a pressure transducer. 5. The system of claim 1, wherein the sensor module further comprises a blood pressure sensor comprising a pulse rate sensor and a blood volume sensor. 6. The system of claim 1, further comprising a personal database and an anonymous database, wherein the personal database contains cardiopulmonary information for the person, and wherein the anonymous database contains cardiopulmonary information for multiple other persons, and wherein the anonymous database is publicly accessible through the Internet. 7. The system of claim 6, wherein the portable telecommunication device comprises at least one signal processor configured to... | Valencell, Inc. | A61B005/0205 A61B005/22 A61B005/00 A61B005/021 A61B005/11 A61B005/145 A61B005/029 A61B007/00 A61B005/12 A61B007/04 A61B005/03 A61B005/0476 A61B005/0496 A61B005/1455 A61B005/01 A61B005/04 A61B005/103 A61B005/1477 A61M037/00 A61N001/32 A61B005/024 G01D011/00 G01N021/47 G01N021/59 A61B005/026 A61B005/08 A61B005/087 A61B005/053
US10716515 | B2 | 2020-07-21 | Systems, devices, and methods for tracking and compensating for patient motion during a medical imaging scan | 2016-11-18 | In some embodiments, a motion detection and correction system and/or device for tracking and compensating for patient motion during a medical imaging scan can be adapted to be integrated into a medical imaging scanner, such as an MRI scanner, or be adapted to retrofit a pre-existing medical imaging scanner. In certain embodiments, the motion detection system and/or device can comprise one or more mounting brackets and a motion correction device housing, which can further comprise one or more detector modules and a power unit. In further embodiments, the collected motion data can be further analyzed by an image processing unit of the motion tracking and/or correction system. | 1. A motion correction device for a medical imaging scanner, the device comprising: a device housing, wherein the device housing comprises an arcuate surface, and wherein the device housing comprises: one or more optics openings on the arcuate surface;one or more detectors configured to detect motion of a subject of the medical imaging scanner through the one or more optics openings, wherein each of the one or more detectors further comprises: a detector housing; anda sensor module placed within the detector housing, wherein the sensor module is configured to be removably coupled to the detector module housing;a power unit configured to regulate power to the one or more detectors; andone or more wires configured to connect the one or more detectors to the power unit,wherein the device housing is configured to be removably coupled to a top inner surface of a bore of the medical imaging scanner. 2. The motion correction device of claim 1, wherein the device is configured to be removably coupled to a plurality of medical imaging scanners, wherein each of the plurality of medical imaging scanners comprises a bore of a different size. 3. The motion correction device of claim 1, wherein the device is configured to detect motion of the subject of the medical imaging scanner and transmit the detected motion to a motion tracking system for processing the detected motion. 4. The motion correction device of claim 1, wherein the device is configured to be removed and reattached to the medical imaging scanner without losing alignment of the one or more detectors. 5. The motion correction device of claim 1, wherein the one or more optics openings comprises indium tin oxide coated glass. 6. The motion correction device of claim 1, wherein the one or more optics openings protrude from the arcuate surface at an angle. 7. The motion correction device of claim 1, wherein the device housing further comprises one or more radiofrequency chokes. 8. The motion correction device of claim 1, wherein the device housing further comprises one or more mounting clips, wherein the one or more mounting clips are configured to be removably attached to a mounting bracket, wherein the mounting bracket is attached to the top inner surface of the bore. 9. The motion correction device of claim 1, wherein the detector housing is flash plated with a material configured to delay oxidation. 10. The motion correction device of claim 1, wherein the detector housing comprises a top cover and a bottom cover, wherein the top cover comprises one or more non-parallel walls to eliminate standing waves. 11. The motion correction device of claim 10, wherein the top cover comprises nickel. 12. The motion correction device of claim 1, wherein the sensor module comprises an imaging sensor, an optics module, and a light source. 13. The motion correction device of claim 12, wherein the optics module comprises a sensor and an optics. 14. The motion correction device of claim 12, wherein the ... | Kineticor, Inc. | A61B005/05 A61B005/00 A61B005/11 A61B005/055 A61N005/10 G01R033/565 G01R033/28 A61B005/042
US10722157 | B2 | 2020-07-28 | Apparatus and method for monitoring biometric parameters using an ambient light collector | 2015-06-12 | An apparatus and method, the apparatus comprising; a light collector configured to collect ambient light and provide the collected light to a portion of skin of a user; a photodetector configured to enable a biometric parameter to be monitored by detecting changes in the light absorbed by the portion of skin; and an amplifier configured to amplify an output signal provided by the photodetector; wherein the light collector is configured to filter the collected ambient light to increase the proportion of light within a selected range of wavelengths which is provided to the portion of skin of the user relative to the collected ambient light. | 1. An apparatus comprising; a light collector configured to collect ambient light, increase an intensity of the ambient light collected, and provide the collected light to a portion of skin of a user;a photodetector configured to enable a biometric parameter to be monitored by detecting changes in the light absorbed by the portion of skin; andan amplifier configured to amplify an output signal provided by the photodetector;wherein the light collector comprises at least one waveguide patterned with a coupler microstructure configured to scatter the collected ambient light in the waveguide, the waveguide being configured to focus the collected ambient light onto the portion of the skin of the user and to filter the collected ambient light to increase the proportion of light within a selected range of wavelengths which is provided to the portion of skin of the user relative to the collected ambient light. 2. An apparatus as claimed in claim 1 wherein the waveguide comprises photoluminescent material configured to absorb light of a first wavelength and reemit light within the selected range of wavelengths. 3. An apparatus as claimed in claim 1 wherein the waveguide is configured to filter for different selected ranges of wavelengths to enable different biometric parameters to be monitored. 4. An apparatus as claimed in claim 1 wherein the waveguide is configured to provide light in one range of wavelengths selected to monitor heart rate. 5. An apparatus as claimed in claim 1 wherein the waveguide is configured to provide light in two different ranges of wavelengths selected to monitor blood oxygenation levels. 6. An apparatus as claimed in claim 1 wherein the apparatus comprises an array of photodetectors configured to enable a biometric parameter to be monitored by detecting changes in the light absorbed by the portion of skin. 7. An apparatus as claimed in claim 1 wherein the amplifier is configured to provide a variable gain. 8. An apparatus as claimed in claim 7 wherein the gain of the amplifier is dependent upon the intensity of the ambient light. 9. An apparatus as claimed in claim 1 further comprising a further photodetector configured to detect the brightness of the ambient light. 10. An apparatus as claimed in claim 1 wherein the apparatus comprises a supplementary light source. 11. An apparatus as claimed in claim 10 wherein the supplementary light source comprises one or more light emitting diodes. 12. An apparatus as claimed in claim 1 wherein the apparatus is configured so that the light collector and photodetector are positioned adjacent to the skin of the user. 13. An apparatus as claimed in claim 1 wherein the apparatus is configured to be worn by a user. 14. An apparatus as claimed in claim 1, wherein the waveguide is configured to filter using a filter that allows light within a selected range of wavelengths to pass through. 15. A method comprising; providing a light collector configured to collect ambient light, incre... | Nokia Technologies Oy | A61B005/1455 A61B005/024 A61B005/00
US10712337 | B2 | 2020-07-14 | Detecting gases and respiration by the conductivity of water within a porous substrate sensor | 2015-10-22 | A method of and system for detecting a gas or vapor includes providing a sensor comprising an electrode pair in electrical contact with a layer of porous material, the porous material layer having water adsorbed on its surface; contacting the sensor with a gas or vapor sample to be analysed; applying a voltage across the electrode pair of the sensor; and measuring a response, the response correlating to the presence of a target gas or vapor. | 1. A method of detecting respiration of a patient comprising: providing a sensor comprising an electrode pair in electrical contact with a layer of hygroscopic porous cellulose paper material capable of adsorbing and desorbing water as a function of relative humidity;positioning the sensor in the proximity of a mouth and/or nose of the patient; andmonitoring the change in water content of the sensor over time by application of a voltage across the electrode pair, such that a current is induced that is proportional to the ionic content of water in the hygroscopic porous cellulose material, and wherein the increase and decrease in water content is correlated to exhalation and inhalation, respectively. 2. The method of claim 1, wherein the method detects a respiratory condition such as apnea, hypopnea and respiratory arrhythmia. 3. The method of claim 2, wherein the respiratory condition is one or more of apnea, hypopnea and respiratory arrhythmia. 4. The method of claim 1, wherein the response is an increase in current, indicating an increase in the adsorbed water in the sensor. 5. The method of claim 1, wherein the response comprises one or more of current, electrical resistance, voltage, or conductance. 6. The method of claim 1, wherein said response is wirelessly transmitted from the sensor to a remote location. 7. The method of claim 1 wherein a microcontroller is used to process the response of the sensor to generate data representative of said change of water content over time in an exhaled gas stream. 8. The method of claim 1, wherein the increase and decrease in water content is used to generate data representative of breathing rate. 9. The method of claim 1, wherein the increase and decrease in water content is used to generate data representative of breathing depth. 10. A respiration monitor, comprising: a support capable of being positioned in the proximity of a mouth and/or nose of a patient; anda sensor comprising an electrode pair in electrical contact with a layer of hygroscopic porous cellulose paper material capable of adsorbing and desorbing water as a function of relative humidity, said sensor attached to the support and located so as to be in the proximity of patient inhalation and exhalation, wherein upon application of a voltage across the electrode pair, current is induced that is proportional to the ionic content of water in the hygroscopic porous cellulose paper material. 11. The respiration monitor of claim 10, further comprising an amplification agent in or on the hygroscopic porous cellulose paper material. 12. The respiration monitor of claim 10, further comprising a damping agent in or on the hygroscopic porous cellulose paper material. 13. The respiration monitor of claim 10, wherein the hygroscopic porous cellulose paper material comprises woven textiles, non-woven textiles, non-woven meshes, fabrics, or paper. 14. The respiration monitor of claim 10, further comprising a voltage source configured to app... | President and Fellows of Harvard College | A61B005/02 G01N033/497 A61B005/08 G01N027/12 A61B005/00
US10709862 | B2 | 2020-07-14 | Device for withdrawing and for transporting a breathing gas stream | 2015-10-05 | A device 10 withdraws a breathing gas stream (A) from a ventilation system (B) and transports the breathing gas stream (A) to a gas analysis system (G). The device 10 has a tubular configuration with an inner side (41) and with an outer side (42) and includes two tube sections (11, 11′) and a drying stage (12, 14, 22) with an inner side (43, 43′) and with an outer side (44, 44′), and at least one liquid storage device (13, 21). The drying stage (12, 14, 22) includes a gas-tight and moisture-permeable material that transports moisture from the inner side (43, 43′) of the drying stage (12, 14, 22) through the gas-tight and moisture-permeable material to the outer side (42) of the tubular device (10). The drying stage (12, 14, 22) and/or the liquid storage device (13, 21) is arranged at least partially between the two tube sections (11, 11′). | 1. A device for withdrawing a breathing gas stream from a ventilation system and for transporting the breathing gas stream to a gas analysis system, the device comprising: an upstream tube section comprising a gas-tight and moisture-permeable material defining a gas passage for the breathing gas stream and forming an upstream drying stage with an inner side and with an outer side and configured such that moisture is transported from the inner side of the upstream drying stage through the gas-tight and moisture-permeable material of the upstream drying stage to the outer side of the upstream tube section;a downstream tube section comprising a gas-tight and moisture-permeable material defining a gas passage for the breathing gas stream and forming a downstream drying stage with an inner side and with an outer side and configured such that moisture is transported from the inner side of the downstream drying stage through the gas-tight and moisture-permeable material of the downstream drying stage to the outer side of the downstream tube section; andat least one liquid storage device disposed downstream of the upstream drying stage and disposed upstream of the downstream drying stage and in fluid communication with the breathing gas stream for absorbing moisture from the breathing gas stream and storing absorbed moisture at the liquid storage and for releasing stored moisture from the liquid storage back into the breathing gas stream, wherein the breathing gas stream passes through the device in the sequence of the upstream drying stage of the first tube section, followed by the liquid storage device and then followed by the downstream drying stage of the second tube section, wherein the upstream tube section, the downstream tube section and the at least one liquid storage device together form a tubular configuration. 2. A device in accordance with claim 1, wherein the at least one liquid storage device is selected from among a hydrophilic porous sintered plastic material, a nonwoven, a plasma-treated granular polyethylene or a matrix of microstructured plastic particles. 3. A device in accordance with claim 1, wherein the at least one liquid storage device has a tubular configuration. 4. A device in accordance with claim 1, wherein a desiccant, comprising zeolite, silica gel or both zeolite and silica gel, is applied to the outer side of at least one drying stage. 5. A device in accordance with claim 1, further comprising a further drying stage comprising a drying element connected to the tubular configuration, wherein the at least one drying stage is arranged at least partially in the drying element. 6. A device in accordance with claim 5, wherein the at least one liquid storage device has a wick-shaped configuration. 7. A device in accordance with claim 1, wherein the at least one drying stage has a membrane based on polyether imides, polyether block amides or polyurethanes. 8. A device in accordance with claim 7, wherein the material of wh... | Drägerwerk AG & Co. KGaA | A61M016/08 A61B005/08 A61B005/097 A61M016/10 B01D053/26 B01D053/22
US10702188 | B2 | 2020-07-07 | System and methods for video-based monitoring of vital signs | 2019-06-24 | The present invention relates to the field of medical monitoring, and in particular non-contact, video-based monitoring of pulse rate, respiration rate, motion, and oxygen saturation. Systems and methods are described for capturing images of a patient, producing intensity signals from the images, filtering those signals to focus on a physiologic component, and measuring a vital sign from the filtered signals. | 1. A video-based method of measuring a patient's vital sign, comprising: generating a video signal from a video camera having a field of view that includes exposed skin of a patient;extracting from the video signal a plurality of time-varying intensity signals;combining two or more of the time-varying intensity signals within a region of the field of view to produce a regional intensity signal;transforming the regional intensity signal into the frequency domain to produce a regional frequency signal;accumulating peaks in the regional frequency signal over a sliding time window;selecting a median frequency from the accumulated peaks;converting the median frequency into a measured vital sign value and adding the measured vital sign value to a running average; andoutputting the averaged vital sign value for further processing or display. 2. The method of claim 1, wherein the region is selected based on a strength of modulations of intensity signals in the region. 3. The method of claim 1, wherein the region includes non-adjacent areas or pixels. 4. The method of claim 1, further comprising removing peaks from the accumulated peaks when they reach an age limit. 5. The method of claim 1, further comprising discarding peaks outside of a physiologic limit. 6. The method of claim 1, further comprising discarding the measured vital sign value when it differs from the running average by more than a defined amount. 7. The method of claim 1, further comprising discarding peaks if they are sub-harmonics of already identified peaks. 8. The method of claim 1, further comprising identifying a frequency content of each of said plurality of time-varying intensity signals, averaging the color signals from said plurality of time-varying intensity signals, and de-noising the signal. 9. The method of claim 8, further comprising performing a fast Fourier transform operation over a sliding time window to identify frequency components of said plurality of time-varying intensity signals. 10. The method of claim 1, wherein the vital sign comprises respiration rate or heart rate. 11. The method of claim 10, wherein the vital sign comprises heart rate. 12. A video-based method of measuring a patient's vital sign, comprising: generating a video signal from a video camera having a field of view that includes exposed skin of a patient;extracting from the video signal a time-varying color signal for each of a plurality of pixels in the field of view;identifying a frequency content of each color signal;combining color signals within a region of the field of view;measuring said vital sign from the combined color signal, including over a period of time, accumulating frequency peaks from the combined signal, selecting a median frequency and updating a running average vital sign value of a patient by converting the median frequency into a measured vital sign value and adding the measured vital sign value to the running average, wherein accumulating frequency peaks comprise... | Covidien LP | A61B005/103 A61B005/00 A61B005/024 A61B005/1171 A61B005/1455 A61B005/0205 A61B005/145 A61B005/08
US10702445 | B2 | 2020-07-07 | Apparatus and methods for assisting breathing | 2018-03-16 | The present invention provides, among other things, apparatus and methods of use for treating a subject in need of assistance with breathing. In some embodiments the subject suffers from airflow obstruction. In some embodiments, the subject suffers from chronic obstructive pulmonary disease. | 1. A method of treating a subject comprising: (a) sensing an onset of the subject's expiratory phase of breathing; and (b) delivering a stimulus to the subject's thorax or abdomen during at least part of the expiratory phase of breathing in response to said sensing, said stimulus effective to cause a contraction in the subject's abdominal muscle so as to assist the subject with ventilation, wherein at least one of the sensing and delivering is carried out by one or more implanted sensors and/or implanted device adapted to deliver a stimulus. 2. The method of claim 1, wherein performing the method decreases the subject's end expiratory lung volume (EELV) as compared with the subject's EELV when unassisted. 3. The method of claim 1, wherein the subject suffers from a disorder that impairs exhalation. 4. The method of claim 1, wherein the subject suffers from chronic obstructive pulmonary disease (COPD). 5. The method of claim 1, wherein the subject has an abnormally high EELV in the absence of the stimulus. 6. The method of claim 1, wherein the subject suffers from dynamic hyperinflation in the absence of the stimulus. 7. The method of claim 1, wherein step (a) comprises detecting one or more characteristics intracorporeally via one or more implanted sensors. 8. The method of claim 7, wherein step (a) comprises detecting intra-abdominal pressure. 9. The method of claim 7, wherein step (a) comprises detecting respiratory movement of the subject's thorax or abdomen using one or more sensors selected from the group consisting of a strain gauge, a respiratory plethysmography sensor, a piezoelectric sensor, a pressure sensor for measuring air flow caused by breathing of the subject, and combinations thereof. 10. The method of claim 7, wherein the one or more implanted sensors is in wireless communication with the device adapted to deliver the stimulus. 11. The method of claim 7, wherein the one or more implanted sensors is in direct electrical communication with the device adapted to deliver the stimulus. 12. The method of claim 7 wherein step (b) is carried out by a device implanted in the subject to deliver a stimulus internally to nerves and/or muscles of a subject. 13. The method of claim 12, wherein the one or more sensors and the device are in communication with an external computer. 14. The method of claim 12, wherein physiological data of the subject is communicated wirelessly to an external computer. 15. The method of claim 1, wherein step (a) comprises directly detecting action potentials or muscle contractions via one or more implanted sensors. 16. The method of claim 1, wherein step (b) is carried out by a device implanted in the subject to deliver a stimulus internally to nerves and/or muscles of a subject. 17. The method of claim 16, wherein step (b) comprises delivering electrical stimulation to one or more abdominal muscles. 18. The method of claim 16, wherein step (b) comprises delivering electrical stimulation to one o... | Apellis Holdings, LLC | A61N001/36 A61H009/00 A61H031/00 A61B005/08 A61B005/00 A61N002/00 A61B005/113 A61H011/00
US10722123 | B2 | 2020-07-28 | Method for detecting at least one of a heart rate and a respiratory rate of a subject | 2016-12-08 | A method and system for detecting at least one of a heart rate and a respiratory rate of a subject are disclosed. In one aspect, the method includes transmitting a signal towards the subject and receiving a reflected signal from the subject being Doppler-shifted due to at least one of the heart rate and the respiratory rate. The method also includes providing the reflected signal to a first input of a phase comparator, providing an adjustible reference signal to a second input of the phase comparator, and generating an output signal by the phase comparator based on the reflected signal and the reference signal. The method includes varying, by the reference signal generator, at least one of a phase and a frequency of the adjustable reference signal based on the output signal of the phase comparator to track a phase of the reflected signal. | 1. A method for detecting at least one of a heart rate and a respiratory rate of a subject, comprising: transmitting, via a transmitter antenna, a radio frequency signal towards the subject;receiving, via a receiver antenna, a reflected signal from the subject, the reflected signal being Doppler-shifted relative to the radio frequency signal due to at least one of the heart rate and the respiratory rate;providing the reflected signal to a first input of a phase comparator;generating an adjustable reference signal by a signal generator and providing the adjustable reference signal to a second input of the phase comparator;generating an output signal by the phase comparator based on the reflected signal and the adjustable reference signal;generating a filtered output signal by filtering the output signal with a filter and providing the filtered output signal to an integrator and a processing system;integrating the filtered output signal by the integrator to generate an integrated filtered output signal;varying, by the signal generator, at least one of a phase and a frequency of the adjustable reference signal based on the integrated filtered output signal or the filtered output signal to track a phase of the reflected signal; anddetermining, by the processor, the at least one of the heart rate and the respiratory rate of the subject based on at least one of the integrated filtered output signal and the filtered output signal. 2. The method according to claim 1, wherein determining the at least one of the heart rate and the respiratory rate comprises performing a frequency analysis of the filtered output signal. 3. The method according to claim 1, wherein determining the at least one of the heart rate and the respiratory rate comprises performing a frequency analysis of the integrated filtered output signal. 4. The method according to claim 1, further comprising estimating a magnitude of motion of tissue of the subject. 5. The method according to claim 1, wherein the phase comparator includes a mixer. 6. The method according to claim 1, wherein the signal generator includes a voltage controlled oscillator. 7. The method according to claim 1, wherein the reference signal generator includes an oscillator and a phase modulator. 8. The method according to claim 7, wherein providing the integrated filtered output signal to the signal generator comprises providing the integrated filtered output signal to the phase modulator. 9. The method according to claim 1, wherein the adjustable reference signal is generated with a phase which is correlated to a phase of the radio frequency signal. 10. The method according to claim 1, wherein the radio frequency signal transmitted towards the subject is a fixed-frequency signal. 11. A system for detecting at least one of a heart rate and a respiratory rate of a subject, the system comprising: a transmitter configured to transmit a radio frequency signal towards the subject;a receiver configured to receive a ... | IMEC vzw Stiching IMEC Nederland | A61B005/0205 A61B005/05 A61B005/113 A61B005/08 A61B005/00 A61B005/024
US10703223 | B2 | 2020-07-07 | Seat occupancy recognition | 2017-05-03 | The invention relates to a seat occupancy recognition device for recognising the occupancy of a seat, characterised by at least one volume detection device with at least one first sensor which is provided and designed to detect a volume on the vehicle seat, and at least one body function detection device with at least one second sensor, which is provided and designed to detect body functions, and at least one body function recognition device for recognising the detected body functions. | 1. A method for recognizing the seat occupancy of a seat, comprising:a. Performing a volume detection process for detecting a volume on a vehicle seat by a volume detection device, wherein the seat comprises the volume detection device,b. Performing a body function detection process by a body function detection device only after positive volume detection, wherein the seat comprises the body function detection device,c. Performing a body function recognition process by a body function recognition device only after positive body function detection, wherein the body function recognition process comprises at least one process of detecting a cardiac function and detecting a pulmonary function, wherein the seat comprises the body function recognition device,wherein after a negative body function recognition, either: (i) the body function detection process is performed again,(ii) one continues with method step a., or(iii) at least one additional body function is detected by the body function detection device. 2. The method according to claim 1, further comprising:the at least one additional body function is selected from a group which comprises detecting a temperature, perspiration, gastric acoustics, body odor and brain waves and/or a combination of these. 3. The method according to claim 1, whereindata that is recorded by the body function detection device is shown by a display. 4. The method according to claim 1, whereina health status is detected by the body function recognition device after positive body function recognition. 5. A vehicle seat comprising: a seat occupancy recognition device for recognizing seat occupancy of a seat,having:at least one volume detection device having a first sensor, which is provided and designed for detecting a volume on the vehicle seat, and at least one body function detection device having a second sensor, which is provided and designed for detecting body functions, and at least one body function recognition device for recognizing the detected body functions, the volume detection device and the body function detection device being connected to a control device, the control device operating the body function detection device only in response to receiving signal from the volume detection device indicating that volume on the vehicle seat has been detected. 6. The vehicle seat according to claim 5, whereinby the first sensor of the volume detection device, at least one process of contour detection and mass detection can be performed. 7. The vehicle seat according to claim 5, whereinby the second sensor of the body function detection device, at least one of the functions of the cardiac function and pulmonary function can be detected. 8. The vehicle seat according to claim 5, whereinthe seat occupancy recognition device has a storage device by which the sensor data, which can be recorded by the body function detection device, can be stored. 9. The vehicle seat according to claim 5, whereinthe second sensor is de... | GRAMMER AG | B60N002/00 B60N002/75 A61B005/024 G01G019/00 G06K009/00 A61B005/0205 G01V011/00 A61B005/08 A61B005/11
US10708553 | B2 | 2020-07-07 | Measurement support device, endoscope system, processor for endoscope system | 2019-06-19 | The invention aims at providing a measurement support device, an endoscope system, and a processor for an endoscope system capable of displaying an accurate indicator with a simple configuration. In a measurement support device related to one aspect of the invention, the coordinates of a spot, and coordinates of points indicating an actual size of a measurement target in a subject and indicating a circular marker distorted in accordance with distortion aberration of an imaging optical system are stored in association with each other in a storage unit, the coordinates of the points indicating the circular marker are acquired with reference to the storage unit on the basis of the measured coordinates of the spot, and the circular marker is displayed on the basis of the acquired coordinates. Thus, the distance measurement is unnecessary, the configuration is simple, and the processing load is low. Additionally, since the circular marker is displayed in the vicinity of the spot (for example, centering on a spot position), there is little deviation between the spot position and a marker position, the circular marker is accurate as an indicator. Additionally, since the indicator is not widely displayed, there is little processing load. | 1. A measurement support device comprising: a head that emits measurement auxiliary light;an imaging unit that captures an image of a subject on which a spot is formed with the measurement auxiliary light via an imaging optical system and an imaging element;a measurement unit that measures coordinates of the spot in the image;a storage unit that stores the coordinates of the spot and coordinates of points indicating an actual size of a measurement target in the subject and indicating a circular marker distorted in accordance with distortion aberration of the imaging optical system in association with each other and that stores the coordinates of the points indicating the circular marker with respect to a plurality of points in a trajectory along which the spot moves on the image in a case where an imaging distance of the image is changed;a coordinate acquisition unit that refers the storage unit on the basis of the measured coordinates of the spot and acquires the coordinates of the points indicating the circular marker corresponding to the coordinates of the spot; anda display control unit that causes the circular marker to be displayed in the vicinity of the spot in the image on the basis of the acquired coordinates,wherein the head emits the measurement auxiliary light that has an inclination angle that is not 0 degrees with respect to an optical axis of the imaging optical system and crosses an angle of view of the imaging optical system, in a case where an optical axis of the measurement auxiliary light is projected on a plane including the optical axis of the imaging optical system. 2. The measurement support device according to claim 1, wherein the coordinate acquisition unit acquires the coordinates of the points indicating the circular marker corresponding to a point of which a distance from the spot is equal to or smaller than a threshold value, among the plurality of points. 3. The measurement support device according to claim 1, wherein the coordinate acquisition unit acquires the coordinates of the points indicating the circular marker by interpolating coordinates corresponding to two or more points sandwiching the spot, among the plurality of points. 4. The measurement support device according to claim 1, wherein the coordinate acquisition unit acquires the coordinates of the points indicating the circular marker by extrapolating coordinates corresponding to two or more points that do not sandwich the spot, among the plurality of points. 5. The measurement support device according to claim 1, wherein the storage unit stores the coordinates of the points indicating the circular marker in correspondence with a range where size measurement of the measurement target by the circular marker is effective, in the image. 6. The measurement support device according to claim 1, wherein the storage unit stores coordinates of a plurality of first points, which are actually measured in correspondence with a circle centering on the spot, and ... | FUJIFILM Corporation | H04N007/18 A61B001/00 A61B001/005 A61B001/04 A61B001/06 A61B005/107 G06T003/00 G06T007/00 G02B023/24 A61B001/07
US10724066 | B2 | 2020-07-28 | Saliva glucose measurement devices and methods | 2017-01-27 | Devices and methods capable of detecting glucose in saliva (FIG. 12). The devices feature a sensor having a substrate containing electrodes and one or more reagents on the electrodes. A detection device is operably coupled with the sensor to detect glucose based on measurement of an electrical parameter when electricity is applied to the electrode. | 1. A method of detecting a blood glucose level using saliva, the method comprising: contacting a sensor containing electrodes and one or more glucose-detection reagents on said electrodes with saliva; andenergizing said electrodes such that a measurement of an electrical parameter indicative of an amount of glucose in said saliva is produced and measured,wherein said one or more glucose-detection reagents comprise glucose dehydrogenase (GDH) and ferricyanide in 1X phosphate buffered saline (PBS) buffer. 2. A method of detecting a blood glucose level using saliva, the method comprising: contacting a sensor containing electrodes and one or more glucose-detection reagents on said electrodes with saliva; andenergizing said electrodes such that a measurement of an electrical parameter indicative of an amount of glucose in said saliva is produced and measured, wherein said one or more reagents comprise glucose dehydrogenase flavine-adenine dinucleotide (GDH-FAD) and ferricyanide in 1×PBS buffer. 3. A method of monitoring a glucose level against a desired range in a patient, comprising: contacting a sensor containing electrodes and one or more glucose-detection reagents on said electrodes with saliva; andenergizing said electrodes such that a measurement of an electrical parameter indicative of an amount of glucose in said saliva is produced and measured against said desired range,wherein said one or more glucose-detection reagents comprise glucose dehydrogenase (GDH) and ferricyanide in 1×PBS buffer. 4. The method of claim 3, further including a sulfonated tetrafluoroethylene-based fluoropolymer-copolymer coating or mesoporous carbon coating on one or more of the electrodes. 5. The method of claim 3, wherein said glucose-detection reagents are dried prior to contacting with the saliva. 6. A method of monitoring a glucose level against a desired range in a patient, comprising: contacting a sensor containing electrodes and one or more glucose-detection reagents on said electrodes with saliva; andenergizing said electrodes such that a measurement of an electrical parameter indicative of an amount of glucose in said saliva is produced and measured against said desired range, wherein said one or more reagents comprise glucose dehydrogenase flavine-adenine dinucleotide (GDH-FAD) and ferricyanide in 1×PBS buffer. | Arizona Board of Regents on behalf of Arizona State University | G01N027/327 C12Q001/00 G01N027/403 G01N027/70 G01N027/31 A61B005/145
US10709393 | B2 | 2020-07-14 | Methods and apparatus for dynamic displays | 2018-12-03 | Apparatus and methods for providing dynamic displays. In one embodiment, a method is provided wherein health parameters having goals associated therewith are displayed; data relating to the health parameters is received and the display is incremented according to a progress of the user toward the goal. When collection of the data is completed, it is determined whether a current value of a first one of the health parameters within a predetermined range of the goal associated therewith. When the current value of the first one of the health parameters is within the range, a portion of the display relating thereto is upgraded. When the current value of all of the health parameters is within the range, the entire display is upgraded. The display upgrades may comprise display of an animation and/or using color, light, and/or flashing to simulate a shining or shimmering of the display. | 1. A method of providing a dynamic display to a user, said method comprising: displaying to the user a display comprising a plurality of portions, each of said plurality of portions of said display relating to a respective one of a plurality health parameters having goal associated therewith and comprising at least shading representative of a progress of said user toward said goal;receiving health parameter data relating to one or more of said plurality of health parameters;updating said shading representative of said progress of said user toward said goal associated to said one or more of said plurality of health parameters based on said received health parameter data;when said update to said shading representative of said progress of said user toward said goal associated to said one or more of said plurality of health parameters comprises an update to complete said shading of said portion of said display relating to at least one of said one or more of said plurality of health parameters, upgrading said portion of said display; andwhen said update to said shading representative of said user's progress toward said goal associated to said one or more of said plurality of health parameters comprises an update to complete said shading of all of said portions of said display, upgrading an entirety of said display. 2. The method of claim 1, wherein said upgrading said portion of said display comprises changing a color used to depict said shading of said portion of said display. 3. The method of claim 1, wherein said upgrading said portion of said display comprises using color, light, or flashing to simulate a shining or shimmering of said portion of said display. 4. The method of claim 1, wherein said upgrading said entirety of said display comprises a change to all colors used to depict said shading. 5. The method of claim 1, wherein said upgrading said entirety of said display comprises causing an animation to be displayed, said animation comprising using color, lights, or flashing to simulate a shining or shimmering of said display. 6. The method of claim 1 wherein said plurality of health parameters comprise at least one of: steps, nutrition, weight, fitness, and/or sleep. 7. An apparatus for management of progress with regard to a plurality of health parameters, said apparatus comprising: at least one first interface configured to enable communication with a plurality of health devices;at least one second interface configured to provide a display to a user;a storage entity; anda processor configured to communicate with said storage entity, said at least one first interface, and said at least one second interface, said processor configured to execute at least one computer program thereon, said computer program comprising a plurality of instructions which are configured to, when executed by said processor, cause said apparatus to: display a plurality of portions at said second interface, each of said plurality of portions of said display rela... | Under Armour, Inc. | A61B005/02 A61B005/11 G06T011/00 G06T013/80 A61B005/00 A61B005/024 G06T011/20
US10709375 | B2 | 2020-07-14 | Devices, systems and methods for testing for degenerative disorders of the central nervous system | 2018-07-10 | The present system is directed in various embodiments to devices, systems and methods for detection, evaluation and/or monitoring olfactory dysfunction by measuring and determining the patient's olfactory detection threshold for the left and the right nostril. In some cases, the present invention relates to devices, systems and methods for detecting an asymmetric differential in a patient's relative olfactory detection and/or identification threshold (left vs right nostril) which, when present, may be used as a device to detect, diagnose and/or monitor olfactory deterioration resulting from Alzheimer's disease, and/or the efficacy of treatment regimens therefor over time. Alternatively, general olfactory dysfunction may be evaluated and/or monitored without regard to the individual nostril's performance. | 1. An inhaler device for presenting an aroma to one nostril of a test subject who inhales and may exhale during the aroma presentation, the device comprising: a lower inhalation body defining:an aroma storage compartment,at least one inhalation port in fluid communication with the ambient atmosphere;an aroma source disposed within the aroma storage compartment;an upper cover connected with the lower inhalation body and comprising a tip operatively connected with an upper housing, the tip having a nasal aperture therethrough, wherein the tip and upper housing define a check ball compartment therein, and wherein the check ball compartment is immediately adjacent and contiguous with the aroma storage compartment;an aperture connecting the check ball compartment in fluid communication with the aroma source compartment; anda check ball disposed within the check ball compartment, wherein the check ball is configured to engage a check ball seat defined within the check ball compartment and cover the aperture connecting the check ball compartment and the aroma source compartment when the test subject exhales and to disengage from the check ball seat when the test subject inhales. 2. The inhaler device of claim 1, wherein the check ball is configured to allow air flow from the aroma storage compartment to the check ball compartment when the test subject inhales and prevent air flow from the check ball compartment to the aroma storage compartment when the test subject either stops inhaling or exhales. 3. The inhaler device of claim 2, wherein the aroma storage compartment is further defined by a bottom section of the lower inhalation body, wherein the bottom section is removable. 4. The inhaler device of claim 1, further comprising keyed indicia on an external surface of the lower inhalation body, the keyed indicia correspondent with the name of the aroma source disposed within the aroma storage compartment. 5. The inhaler device of claim 1, wherein at least one of the group consisting of the upper cover, the tip and the upper housing is at least partially transparent. 6. The inhaler device of claim 1, wherein the aroma comprises either a pure odorant or an odorant. 7. The inhaler device of claim 1, further comprising the upper cover removably attached to the lower inhalation body. 8. The inhaler device of claim 2, further comprising: the upper housing defining a photo detection circuit therein, the photo detection circuit comprising:a light source incorporated into an inner surface of the check ball compartment,a photo detector incorporated into the inner surface of the check ball compartment at a location across the check ball compartment from the light source, the light source aligned with the photo detector and adapted to actuate the photo detector when light from the light source is sensed by the photo detector,a transceiver in operative communication with the photodetector, anda battery in operative communication with the light source, the pho... | Inspired Life Medical, Inc. | A61B005/00 A61B005/087
US10708478 | B2 | 2020-07-07 | Image transformation and display for fluorescent and visible imaging | 2016-03-23 | Improved fluoresced imaging (FI) and other sensor data imaging processes, devices, and systems are provided to enhance display of different secondary types of images and reflected light images together. Reflected light images are converted to a larger color space in a manner that preserves the color information of the reflected light image. FI or secondary images are transformed to a color range within the larger color space, but outside the color area of the reflected light images, allowing the FI or secondary images to be combined with them in a manner with improved distinguishability of color. Hardware designs are provide to enable real-time processing of image streams from medical scopes. The combined images are encoded for an electronic display capable of displaying the larger color space. | 1. An imaging scope system, comprising: a light source configured to direct visible light and fluorescent excitation light toward a subject scene;an optical assembly configured to direct received light from the subject scene toward an image sensor assembly;the image sensor assembly including at least one image sensor, the image sensor assembly capable of detecting visible light and first fluoresced light from the subject scene and producing at least one output signal response;image forming circuitry configured to receive the at least one output signal and produce a first image stream including visible light images having a first color space and a second image stream representing the first fluoresced light;image processing circuitry configured to:(a) convert a format of the first image stream into a second data format having a second color space larger, in terms of the range of colors, than the first color space, while preserving color space content of the first image stream;(b) format the second image stream to the second data format and transform the second image stream to a first color range inside the second color space and outside the first color space;(c) combine the converted first image stream and the transformed second image stream to produce a combined image stream; andvideo encoder circuitry configured to encode the combined image stream to a video encoding format configured for display on an electronic display capable of displaying the second color space. 2. The imaging scope system of claim 1 in which the first image stream expresses the first color space with a first bit depth, and the converted image stream expresses the second color space with a second bit depth larger than the first bit depth. 3. The imaging scope system of claim 1 in which transforming the second image stream to a first color range comprises mapping an intensity range in the second image stream to an intensity range of a selected color. 4. The imaging scope system of claim 1 in which the second image stream has a bit depth expressing intensity variation of the detected first fluoresced light, said bit depth smaller than the bit depth of the second color space, and in which the image processing circuitry further comprises circuitry operable to transforming from second image stream by adding an intensity variation in the second image stream to a base value for a single selected color in the second color space. 5. The imaging scope system of claim 1 in which transforming the second image stream to a first color range further comprises mapping an intensity range in the second image stream to the first color range. 6. The imaging scope system of claim 1 in which the first fluoresced light is from an indocyanine green (ICG) fluorescing agent and the first color range consists essentially of a green or multiple greens. 7. The imaging scope system of claim 1 in which: the image sensor assembly is further capable of detecting second fluoresced light at a spectrum di... | KARL STORZ Imaging, Inc. | H04N005/232 H04N005/225 A61B001/04 A61B001/00 A61B008/08 H04N007/18 G06T011/00 A61B005/00 H04N009/64 G02B023/24 A61B008/12
US10714987 | B2 | 2020-07-14 | Systems and methods for wirelessly powering or communicating with sterile-packed devices | 2019-07-17 | Systems and methods are disclosed herein that can allow for wirelessly powering and/or communicating with a sterile-packed electronic device without removing the electronic device from its sterile packaging and while maintaining the sterility of the electronic device. In some embodiments, a base station with a power transmitter wirelessly transfers power to a power receiver of the electronic device, for example using inductive, capacitive, or ultrasonic coupling. The base station or another external device can also be used to wirelessly program or interrogate the electronic device. Battery charging circuits and switching circuits for use with said systems and methods are also disclosed. | 1. A surgical method, comprising: aligning a power receiver of a surgical electronic device with a power transmitter of a base station;wirelessly transferring power from the power transmitter to the power receiver through a sterile container comprising an inner blister pack in which the electronic device is disposed, an outer blister pack in which the inner blister pack is disposed, and an outer box in which the outer blister pack is disposed;upon receipt of wireless power, automatically initiating communications with the base station; andautomatically switching the device to battery power upon activation of a switch, the switch being configured to detect when the device is physically removed from the inner blister pack. 2. The method of claim 1, wherein the surgical electronic device comprises at least one of an instrument or an implant. 3. The method of claim 1, wherein aligning the power receiver comprises aligning a graphical mark formed on the base station with a corresponding graphical mark formed on the sterile container. 4. The method of claim 1, wherein aligning the power receiver with the power transmitter comprises positioning the sterile container in a recess defined in the base station. 5. The method of claim 1, wherein the sterile container is configured to maintain the power receiver in a fixed position with respect to the sterile container. 6. The method of claim 1, further comprising communicating device data from the electronic device to the base station. 7. The method of claim 6, wherein the device data comprises at least one of device identification information, device diagnostic information, or device charge level information. 8. The method of claim 1, wherein transferring the power is effective to charge a battery of the electronic device. 9. The method of claim 1, wherein transferring the power comprises driving an ultrasonic transducer of the base station to generate ultrasonic waves received at an ultrasonic transducer of the electronic device. 10. The method of claim 1, wherein transferring the power comprises generating a magnetic field between a primary coil of the power transmitter and a secondary coil of the power receiver. 11. The method of claim 1, wherein transferring the power comprises generating an electric field between an electrode of the power transmitter and an electrode of the power receiver. 12. The method of claim 1, wherein the power transmitter comprises a light source and wherein transferring the power comprises directing light from the light source through the sterile container to a photovoltaic cell of the electronic device. 13. The method of claim 1, further comprising, after said transferring and communicating steps, opening the sterile container and using the electronic device in a surgical procedure. 14. The method of claim 1, wherein the electronic device comprises a sensor configured to sense a position or orientation of the electronic device. 15. The method of claim 1, further c... | DePuy Synthes Products, Inc. | H02J050/90 A61B005/00 H02J050/10 H02J007/02 H02J050/15 H02J050/30 H02J050/80 H02J050/05 H02J050/12 H02J050/20 G16H040/67 H04B005/00
US10722167 | B2 | 2020-07-28 | Method, an optical probe and a confocal microscopy system for inspecting a solid organ | 2010-09-17 | A method to inspect a solid organ in a subject includes introducing a needle in a predetermined area of the solid organ, inserting an optical probe through a lumen of the needle, and imaging the predetermined area using the optical probe. An optical probe to inspect a solid organ in a subject, the optical probe being intended to be positioned in the solid organ through a needle, the optical probe includes an optical fiber bundle, a ferule to protect the distal tip of the optical fiber bundle, the ferule comprising a shank and a head, and a sheath wrapping the fiber bundle and the shank, wherein the head of the ferule has a length adapted for the optical probe to image the solid organ while keeping the sheath inside the needle. | 1. A method to image a solid organ in a subject comprising: introducing a needle in a predetermined area of the solid organ, wherein the solid organ is one selected from the group consisting of a pancreas, a liver, a spleen, a lymph node, a prostate, a kidney, breast and ovaries, wherein the needle comprises a beveled distal tip;inserting an optical probe through a lumen of the needle, such that the optical probe, after insertion, protrudes outside the lumen and beyond the beveled distal tip of the needle by at least 2 mm; andimaging the predetermined area using the optical probe, wherein the optical probe comprises:an optical fiber bundle;a ferule that protects a distal tip of the optical fiber bundle, the ferule comprising a shank and a head, wherein the shank and the head of the ferule have a tubular shape with a same internal diameter and are coaxially mounted;a sheath that wraps the fiber bundle and the shank, wherein the head of the ferule has a length adapted for the optical probe to image the solid organ while keeping the sheath inside the needle, the solid organ being at least one selected from a group consisting of a pancreas, a liver, a spleen, a lymph node, a prostate, a kidney, a breast, and an ovary,wherein an external diameter of the head of the ferule is equal to an external diameter of the sheath; anda single objective directly connected coaxially at a distal tip of the optical fiber bundle, wherein the ferule connects the single objective to the distal tip of the optical fiber bundle, and wherein the ferule completely surrounds the single objective up to a distal tip of the single objective at a distal tip of the optical probe,wherein the single objective is forward facing and captures an image in a coaxial direction of the optical probe, andthe optical fiber bundle includes:a flexible portion where the optical fiber bundle is surrounded by the sheath and not surrounded by the ferule, anda rigid portion where the optical fiber bundle is surrounded by both the sheath and the ferule,wherein the external diameter of the flexible portion is the same as the external diameter of the rigid portion. 2. The method according to claim 1, wherein introducing the needle in the solid organ comprises puncturing the solid organ with a tip of the needle. 3. The method according to claim 2, further comprising percutaneously inserting the needle in the organ. 4. The method according to claim 1, wherein introducing the needle in the solid organ comprises puncturing the solid organ using a stylet preliminary inserted in the lumen of the needle, the stylet being driven to protrude out of the needle and being removed from the lumen before inserting the optical probe through the lumen of the needle. 5. The method according to claim 1, further comprising passing the needle through a working channel of an endoscope inserted in the subject through a natural orifice to approach the solid organ. 6. The method according to claim 5, further comprising pa... | Mauna Kea Technologies | A61B005/00
US10702167 | B2 | 2020-07-07 | Method to trigger storage of onset of physiologic condition in an implantable device | 2017-02-08 | An apparatus includes a sensing circuit configured to generate a sensed physiological signal that includes physiological information of a subject, a detection circuit, and a control circuit. The detection circuit detects a physiological condition of a subject using the physiological signal. The control circuit stores sampled values of a segment of the physiological signal in temporary memory storage; and stores the sampled values in non-temporary storage in response to receiving an indication of continued detection of the physiological condition. | 1. An apparatus comprising: a sensing circuit configured to generate a sensed physiological signal, wherein the physiological signal includes ventricular depolarization (V-V) interval information of a subject;a detection circuit electrically coupled to the sensing circuit and configured to: during a detection window, detect a cardiac arrhythmia of a subject using a V-V interval stability detection criterion; andat the end of the detection window, confirm the cardiac arrhythmia using a second detection criterion more specific than the V-V interval stability detection criterion, wherein the V-V interval stability detection criterion is more sensitive than the second detection criterion; anda control circuit configured to: store, in response to detection of cardiac arrhythmia using the V-V interval stability, sampled values of a segment of the physiological signal during the detection window in temporary memory storage;receive, after detection of the cardiac arrhythmia using the V-V interval stability, an indication of confirmation of the detected cardiac arrhythmia using the second detection criterion; andstore, in response to receiving the indication of confirmation of the detected cardiac arrhythmia, the sampled values of the segment of the physiological signal during the detection window in non-temporary storage. 2. The apparatus of claim 1, wherein the control circuit is configured to confirm the cardiac arrhythmia using the sampled values of the segment of the physiological signal stored in the temporary memory storage in response to detection of onset of the cardiac arrhythmia; and to store the sampled values in the non-temporary storage in response to receiving the indication of confirmation of the detected cardiac arrhythmia. 3. The apparatus of claim 2, wherein the detection circuit is configured to detect syncope of the subject, and wherein the control circuit is configured to trigger storage of the sampled values in temporary memory storage in response to detection of syncope of the subject; and to store the sampled values in non-temporary storage in response to receiving an indication of confirmation of syncope of the subject. 4. The apparatus of claim 3, wherein the sensing circuit is configured to generate a sensed hemodynamic signal representative of hemodynamic information of the subject, and wherein the detection circuit is configured to detect syncope of the subject using the hemodynamic signal. 5. The apparatus of claim 4, wherein the sensing circuit is configured to generate a sensed hemodynamic signal representative of at least one of respiration of the subject, blood pressure of the subject, one or more heart sounds of the subject. 6. The apparatus of claim 3, wherein the sensing circuit is configured to generate one or both of a sensed posture signal representative of posture of the subject and a sensed activity signal representative of patient activity, and wherein the detection circuit is configured to detect syncope o... | Cardiac Pacemakers, Inc. | A61B005/0205 A61B005/046 A61B005/0215 A61B005/08 A61B005/0456 A61B005/00 G16H040/67 A61B005/021 A61B005/0245 A61B005/11 A61B007/02 A61N001/362 A61N001/365 A61N001/39 A61B005/042
US10702342 | B2 | 2020-07-07 | Navigation guidance method for complex catheters | 2016-09-08 | A method and system for determining a target location for a medical device having complex geometry relative to an anatomical feature, and for navigating and positioning the medical device at the target location. The system may include a medical device including a treatment element having a centroid, one or more navigation electrodes, and a longitudinal axis and a navigation system in communication with the one or more navigation electrodes, the navigation system including a processing unit. The processing unit may be programmed to define a plane that approximates a surface of the anatomical feature, define a centroid of the anatomical feature, define a vector that is normal to the plane and extends away from the centroid of the anatomical feature, and determine a target location for the treatment element of the medical device based on the vector to assist the user in placing the device for treatment. | 1. A system for recommending a target location for a medical device relative to a target anatomical feature, the system comprising: a medical device including a balloon having complex geometry, a centroid, a first navigation electrode, a second navigation electrode, and a longitudinal axis, the balloon being disposed between the first navigation electrode and the second navigation electrode;a navigation system in communication with the first navigation electrode and the second navigation electrode, the navigation system including a plurality of surface electrodes in communication with the first navigation electrode and the second navigation electrode, a display, and a processing unit having processing circuitry configured to, during use: define a plane that approximates a surface of the target anatomical feature;define a centroid of the target anatomical feature and a boundary of the target anatomical feature;define a vector that is normal to the plane and extends away from the centroid of the target anatomical feature;determine a first target location for the balloon;recommend a second target location for the first navigation electrode and recommend a third target location for the second navigation electrode; anddisplay on the display a first graphical indicator for the second target location for the first navigation electrode, a second graphical indicator for the third target location for the second navigation electrode, and a line corresponding to the vector,the recommended target location for each of the first and second navigation electrode being displayed on the display before beginning at least one of a treatment phase and a mapping phase,the processing circuitry being configured to determine the first target location in three dimensions according to a first equation:[xbybzb]=[x0y0z0]+d⁡[abc]where d is a distance between the centroid of the balloon and the centroid of the target anatomical feature calculated according to a second equation:d=rb⁢cos⁡(sin-1⁡(ros,maxrb))where ros,max is a distance between the centroid of the target anatomical feature and the boundary of the anatomical feature and rb is a distance between the centroid of the balloon and the boundary of the anatomical feature, the processing circuitry being configured to recommend the second target location in three dimensions according to a third equation:ed=[xdydzd]=[xbybzb]-ld⁡[abc]and recommend the third target location in three dimensions according to a fourth equation:ep=[xpypzp]=[xbybzb]+lp⁡[abc]where ld is a distance between the centroid of the balloon and the first navigation electrode and lp is a distance between the centroid of the balloon and the second navigation electrode. 2. The system of claim 1, wherein the balloon has more than one layer of material. 3. The system of claim 1, wherein the display displays the first and second graphical indicators superimposed on a three-dimensional image of the anatomical feature. 4. The system of claim 3, wherein the anatomi... | Medtronic, Inc. | A61B034/20 A61B005/06 A61B090/00 A61B018/14 A61B005/00 A61M025/01 A61B018/00
US10716510 | B2 | 2020-07-21 | Smart clothing with converging/diverging bend or stretch sensors for measuring body motion or configuration | 2017-09-12 | This invention is smart clothing with stretch and/or bend sensors for measuring changes in a person's body configuration. This smart clothing has an elastic nonconductive layer onto which helical converging/diverging conductive pathways are printed using ink which is a mixture of elastic nonconductive material and conductive material. Changes in the transmission of electromagnetic energy through the pathway are analyzed to measure changes in the configuration of the person's body. | 1. Smart clothing with stretch sensors for measuring changes in body configuration comprising: a wearable layer which is configured to be worn by a person, wherein the wearable layer comprises a first electromagnetically-nonconductive material; andan electromagnetically-conductive pathway, wherein a distal portion of the electromagnetically-conductive pathway is more elastic or stretchable than a proximal portion of the electromagnetically-conductive pathway, or vice versa, wherein the electromagnetically-conductive pathway is created by printing electromagnetically-conductive ink onto the wearable layer, wherein the ink comprises a mixture of a second electromagnetically-nonconductive material and an electromagnetically-conductive material, wherein stretching the electromagnetically-conductive pathway causes changes in the transmission of electromagnetic energy through the electromagnetically-conductive pathway, and wherein the changes in the transmission of electromagnetic energy through the electromagnetically-conductive pathway are analyzed to measure changes in the configuration of the person's body. 2. The smart clothing in claim 1 wherein the first electromagnetically-nonconductive material, the second electromagnetically-nonconductive material, or both comprise a polymer. 3. The smart clothing in claim 1 wherein the first electromagnetically-nonconductive material, the second electromagnetically-nonconductive material, or both comprise polyurethane. 4. The smart clothing in claim 1 wherein the first electromagnetically-nonconductive material, the second electromagnetically-nonconductive material, or both comprise elastane. 5. The smart clothing in claim 1 wherein the electromagnetically-conductive material comprises aluminum, carbon, copper, gold, silver, or steel. 6. The smart clothing in claim 1 wherein the first electromagnetically-nonconductive polymer comprises polyurethane, the second electromagnetically-nonconductive polymer comprises polyurethane, and the electromagnetically-conductive material comprises aluminum. 7. The smart clothing in claim 1 wherein the first electromagnetically-nonconductive polymer comprises polyurethane, the second electromagnetically-nonconductive polymer comprises polyurethane, and the electromagnetically-conductive material comprises carbon. 8. The smart clothing in claim 1 wherein the first electromagnetically-nonconductive polymer comprises polyurethane, the second electromagnetically-nonconductive polymer comprises polyurethane, and the electromagnetically-conductive material comprises copper. 9. The smart clothing in claim 1 wherein the first electromagnetically-nonconductive polymer comprises polyurethane, the second electromagnetically-nonconductive polymer comprises polyurethane, and the electromagnetically-conductive material comprises gold. 10. The smart clothing in claim 1 wherein the first electromagnetically-nonconductive polymer comprises polyurethane, the second electromagneticall... | Medibotics | A61B005/00 A61B005/0488 A61B005/11 G06F003/01
US10702208 | B2 | 2020-07-07 | Apparatus and method for electroencephalographic examination | 2015-03-31 | An apparatus comprises a data processing unit (100) which receives electroencephalographic data based on a measurement of a person (110) exposed to anesthetic drug substance which has one or more estimated or measured non-zero levels of concentration in the body (108) of the person (110) as a function of time; and determines and present activity of slow waves of the electroencephalographic data at the one or more estimated or measured levels of the non-zero concentration of the anesthetic drug substance. | 1. An apparatus comprising: an administration device configured to cause estimated or measured non-zero amounts of an anesthetic drug substance to be introduced into a body of a person and to adjust input of the anesthetic drug substance into the body; andat least one processor and associated memory configured to: receive electroencephalographic data for a brain based on at least one measurement of the person exposed to the estimated or measured non-zero amounts of the anesthetic drug substance, the electroencephalographic data comprising slow waves having a frequency at or below 1 Hz;control the administration device to introduce continuously into the body the anesthetic drug substance at a rate of from 2 mg/(kg*h) to 4 mg/(kg*h) and to adjust the estimated or measured non-zero amounts of the anesthetic drug substance based on at least one of the following: the electroencephalographic data, activity of the slow waves, and data relating to the estimated or measured non-zero amounts of the anesthetic drug substance;determine changes in the activity of the slow waves of the electroencephalographic data with respect to the estimated or measured non-zero amounts of the anesthetic drug substance;predict a good neurological recovery of the person if the activity of the slow waves increases by more than 50% relative to the activity of the slow waves prior to the introduction of the anesthetic drug substance and predict a poor neurological recovery of the person if the activity of the slow waves does not increase by more than 50% relative to the activity of the slow waves prior to the introduction of the anesthetic drug substance; andpresent result data that represent the changes and that show (i) whether the activity of the slow waves has increased by more than 50% with the introduction of the anesthetic drug substance, thereby indicating the good neurological recovery, or (ii) whether the activity of the slow waves has not increased by more than 50% with the introduction of the anesthetic drug substance, thereby indicating the poor neurological recovery. 2. The apparatus of claim 1, wherein the at least one processor is further configured to form the result data based on the activity of the slow waves and present the result data for a prediction of a neurological recovery level of the person. 3. The apparatus of claim 1, wherein the at least one processor is further configured to form different pieces of the result data for the increase in the activity of the slow waves being more than 50% and the increase in the activity of the slow waves not being more than 50%. 4. The apparatus of claim 1, wherein the at least one processor is further configured to present the prediction of the good neurological recovery or the poor neurological recovery. 5. The apparatus of claim 1, wherein the at least one processor is further configured to measure the activity of the slow waves based on at least one of the following: amplitude spectrum, frequency spectrum, ph... | CERENION OY | A61B005/00 A61M016/01 A61B005/04 A61B005/0476 A61M016/10 A61B005/048 G16H050/30 A61M005/172
US10704890 | B2 | 2020-07-07 | Dynamic motion detection system | 2017-02-23 | A dynamic motion detection system including a plurality of active independent emitting elements, a signal capturing apparatus, and a computing apparatus is provided. The active independent emitting elements are suitable for being respectively affixed to different parts of a person to-be-tested and actively emitting positioning signals having preset wavelengths. The signal capturing apparatus is configured to capture the positioning signal of each of the active independent emitting elements and calculate a plurality of emitting coordinates according to the positioning signals. The computing apparatus is configured to control operation of the signal capturing apparatus and receive the emitting coordinates from the signal capturing apparatus. The computing apparatus defines a corresponding relationship between the emitting coordinates and the active independent emitting elements by comparing a geometry relationship among the emitting coordinates. | 1. A dynamic motion detection system, comprising: a plurality of active independent emitting elements, respectively affixed to different parts of a person to-be-tested, and each of the active independent emitting elements being configured to actively emit a positioning signal having a preset wavelength;a signal capturing apparatus, configured to capture the positioning signal of each of the active independent emitting elements and obtain a plurality of emitting coordinates according to the positioning signals; anda computing apparatus, configured to control operation of the signal capturing apparatus and receive the emitting coordinates from the signal capturing apparatus,wherein the computing apparatus defines a corresponding relationship between the emitting coordinates and the active independent emitting elements by comparing a geometry relationship among the emitting coordinates,wherein the signal capturing apparatus comprises:a first image capturing module, configured to capture a left-eye image;a second image capturing module, disposed at a position with a predetermined interval from the first image capturing module, and configured to capture a right-eye image; andan image processing module, coupled to the first image capturing module and the second image capturing module, and configured to calculate a plurality of left emitting coordinates of the left-eye image according to the positioning signals and calculate a plurality of right emitting coordinates of the right-eye image according to the positioning signals, wherein the emitting coordinates comprise the right emitting coordinates and the left emitting coordinates,wherein the computing apparatus defines a corresponding relationship between the left emitting coordinates and the active independent emitting elements by comparing a geometry relationship among the left emitting coordinates,wherein the computing apparatus determines whether the left emitting coordinates pass a complete verification according to the corresponding relationship between the left emitting coordinates and the active independent emitting elements. 2. The dynamic motion detection system according to claim 1, wherein when a dynamic motion of the person to-be-tested is detected, the active independent emitting elements are controlled to simultaneously emit the positioning signals. 3. The dynamic motion detection system according to claim 1, wherein each of the active independent emitting elements is a light-emitting diode (LED). 4. The dynamic motion detection system according to claim 1, wherein the preset wavelengths of the positioning signals ranges from 760 nm to 1000 nm. 5. The dynamic motion detection system according to claim 4, wherein each of the active independent emitting elements is an infrared (IR) emitter. 6. The dynamic motion detection system according to claim 1, wherein each of the active independent emitting elements is an active radio frequency identification (RFID) tag. 7. The dynamic motion ... | Giant Manufacturing Co., Ltd. | G01B011/00 G06K009/00 A61B005/11 B62J099/00 G01P013/00 G09B019/00 G06K009/32
US10709404 | B2 | 2020-07-14 | Radiography support device, radiation detection device, radiography apparatus, radiography support method, and radiography support program | 2018-11-21 | A radiography support device that supports imaging of an object with a radiation detection panel in which a plurality of pixels are two-dimensionally arranged and which detects radiation emitted from a radiation source, includes: a first reading control unit that reads first pixel signals from at least some of the plurality of pixels in a state in which the radiation is not emitted from the radiation source and a state in which the object is placed on the radiation detection panel; a first arithmetic unit that calculates a position of the object with respect to the radiation detection panel based on the first pixel signals; and a notification processing unit that performs a notification process in a case in which the position of the object is out of a predetermined range. | 1. A radiography support device that supports imaging of an object with a radiation detection panel in which a plurality of pixels are two-dimensionally arranged and which detects radiation emitted from a radiation source, the device comprising: a first reading control unit that reads first pixel signals from at least some of the plurality of pixels in a state in which the radiation is not emitted from the radiation source and a state in which the object is placed on the radiation detection panel;a first arithmetic unit that calculates a position of the object with respect to the radiation detection panel based on the first pixel signals; anda notification processing unit that performs a notification process in a case in which the position of the object is out of a predetermined range. 2. The radiography support device according to claim 1, further comprising: a second reading control unit that reads second pixel signals from at least some of the plurality of pixels in the state in which the radiation is not emitted from the radiation source and a state in which the object is not placed on the radiation detection panel,wherein the first arithmetic unit specifies a pixel of which a value obtained by processing the first pixel signal using a value obtained based on the second pixel signals is equal to or greater than a threshold value and calculates the position of the object with respect to the radiation detection panel based on a position of the specified pixel. 3. The radiography support device according to claim 2, wherein, in a case in which there are three or more separated regions including the specified pixel, the first arithmetic unit calculates a gravity center position of a range surrounded by the three or more separated regions as the position of the object. 4. The radiography support device according to claim 2, wherein, in a case in which there are two separated regions including the specified pixel, the first arithmetic unit calculates an intermediate position between gravity centers of the two separated regions as the position of the object. 5. The radiography support device according to claim 3, wherein, in a case in which there are two separated regions including the specified pixel, the first arithmetic unit calculates an intermediate position between gravity centers of the two separated regions as the position of the object. 6. The radiography support device according to claim 1, wherein, in a case in which information for using a chest or an abdomen of a living body as the object is set in a radiography apparatus, functions of the first reading control unit, the first arithmetic unit, and the notification processing unit are activated. 7. The radiography support device according to claim 2, wherein, in a case in which information for using a chest or an abdomen of a living body as the object is set in a radiography apparatus, functions of the first reading control unit, the first arithmetic unit, and the notification proc... | FUJIFILM Corporation | A61B006/00 A61B005/00
US10722160 | B2 | 2020-07-28 | Non-invasive and wearable chemical sensors and biosensors | 2015-12-03 | A non-invasive epidermal electrochemical sensor device includes an adhesive membrane; a flexible or stretchable substrate disposed over the adhesive membrane; and an anodic electrode assembly disposed over the flexible or stretchable substrate including an iontophoretic electrode. The device includes a cathodic electrode assembly disposed adjacent to the anodic electrode assembly over the flexible or stretchable substrate and includes an iontophoretic electrode. Either the cathodic electrode assembly or the anodic electrode assembly also includes a sensing electrode that includes a working electrode and at least one of a counter electrode or a reference electrode. The iontophoretic electrode in either the anodic electrode assembly or the cathodic electrode assembly that includes the sensing electrode is disposed on the substrate to at least partially encompass the working electrode and the at least one of the counter electrode or the reference electrode. The device includes an electrode interface assembly including independent electrically conductive contacts. | 1. A non-invasive epidermal electrochemical sensor device, comprising: an adhesive membrane;a flexible or stretchable substrate disposed over the adhesive membrane;an anodic electrode assembly disposed over the flexible or stretchable substrate, the anodic electrode assembly including an iontophoretic electrode;a cathodic electrode assembly disposed over the flexible or stretchable substrate and adjacent to the anodic electrode assembly, wherein the cathodic electrode assembly includes an iontophoretic electrode;wherein at least one of the anodic electrode assembly or the cathodic electrode assembly include a sensing electrode that includes a working electrode and at least one of a counter electrode or a reference electrode, and wherein the iontophoretic electrode in the anodic electrode assembly or the cathodic electrode assembly that includes the sensing electrode is disposed on the flexible or stretchable substrate to at least partially encompass the working electrode and the at least one of the counter electrode or the reference electrode,an electrode interface assembly including independent electrically conductive contacts disposed on the flexible or stretchable substrate and electrically coupled to each electrode of the anodic electrode assembly and the cathodic electrode assembly; andwherein the iontophoretic electrodes in the anodic electrode assembly and the cathodic electrode assembly are operable to pass an electrical signal between the iontophoretic electrodes to apply an electric field to drive ion flow from interstitial fluid (ISF) toward the working electrode and the at least one of the counter electrode or the reference electrode,wherein the working electrode includes an electrochemical transducer layer including a catalyst to selectively catalyze a corresponding analyte in the ISF to cause a reaction detectable at the anodic and cathodic electrode assemblies, andwherein the electrically conductive contacts of the electrode interface assembly are configured to electrically couple to one or more electrical circuitry to transmit a sensor signal indicative of the reaction to detect the analyte,wherein the adhesive membrane includes a silicone membrane, a urethane membrane, or an acrylic adhesive membrane,wherein the adhesive membrane includes a gel layer. 2. The non-invasive epidermal electrochemical sensor device of claim 1, wherein the gel layer includes hydrogel or cryogel. 3. The non-invasive epidermal electrochemical sensor device of claim 1, wherein the gel layer includes a charged chemical agent. 4. The non-invasive epidermal electrochemical sensor device of claim 3, wherein the charged chemical agent includes a chemical agent that enhances extraction of ISF. 5. The non-invasive epidermal electrochemical sensor device of claim 4, wherein the ISF includes biofluid and the chemical agent includes skin permeation enhancers. 6. The non-invasive epidermal electrochemical sensor device of claim 5, wherein the iontophoretic elec... | THE REGENTS OF THE UNIVERSITY OF CALIFORNIA | A61B005/1486 A61B005/00 A61B005/145 G01N033/66
US10702184 | B2 | 2020-07-07 | Low power measurement of skin electrical properties | 2017-06-07 | Various examples are directed to a measurement system for measuring an electrical property of skin comprising an excitation circuit, a receiver circuit, and a sequencer circuit. The excitation circuit may generate a periodic excitation signal that, when provided to the skin, generates a response signal in the skin indicative of the electrical property. The sequencer circuit may be configured to activate the excitation circuit to provide the excitation signal to the skin. While the excitation circuit is activated to provide the excitation signal to the skin, the sequencer circuit may activate the receiver circuit to execute a first sample cycle to generate a first plurality of samples of the response signal. A first value for the electrical property of the skin may be determined based at least in part on the first plurality of samples of the response signal. | 1. A measurement system for measuring an electrical property of skin, comprising: an excitation circuit comprising a digital waveform generator to generate a periodic excitation signal that, when provided to the skin, generates a response signal in the skin indicative of the electrical property;a receiver circuit comprising an analog-to-digital converter (ADC); anda sequencer circuit configured to perform operations comprising: activate the excitation circuit to provide the excitation signal to the skin;while the excitation circuit is activated to provide the excitation signal to the skin, activating the receiver circuit to execute a first sample cycle to generate a first plurality of samples of the response signal, wherein the receiver circuit oversamples the response signal during the first sample cycle;after the receiver circuit executes the first sample cycle, deactivating the receiver circuit;after the deactivating of the receiver circuit and while the excitation circuit is activated to provide the excitation signal to the skin, reactivating the receiver circuit to execute a second sample cycle to generate a second plurality of samples of the response signal;determining a first value for the electrical property of the skin based at least in part on the first plurality of samples of the response signal; anddetermining a second value for the electrical property of the skin based at least in part on the second plurality of samples of the response signal. 2. The measurement system of claim 1, wherein the receiver circuit is configured to perform operations comprising: capturing a first sample of the first plurality of samples when the excitation signal is at a first position of a response signal period;capturing a second sample of the first plurality of samples when the excitation signal, is at a second position of the response signal period;capturing a first sample of the second plurality of samples when the excitation signal is at the first position of the response signal period; andcapturing a second sample of the second plurality of samples when the excitation signal is at the second position of the response signal period. 3. The measurement system of claim 1, wherein the first sample cycle is executed after an excitation setup time has elapsed since the activating of the excitation circuit. 4. The measurement system of claim 1, wherein the sequencer circuit is further configured to perform operations comprising: deactivating the excitation circuit after the receiver circuit executes the first sample cycle; andreactivating the excitation circuit, wherein the reactivating the receiver circuit is after the reactivating of the excitation circuit by an excitation setup time. 5. The measurement system of claim 1, further comprising a discrete Fourier transform (DFT) circuit configured to determine a DFT of the response signal based at least in part on the first plurality of samples. 6. The measurement system of claim 5, wherein the DFT circu... | Analog Devices International Unlimited Company | A61B005/053 A61B005/00 A61B005/01
US10724897 | B2 | 2020-07-28 | Cascaded gapped cantilever for low-frequency vibration sensing | 2016-06-16 | This disclosure describes a sensor for detecting vibrations, particularly low-frequency vibrations. The sensor is of a cascaded gapped cantilever construction and may include a single sensing beam extending from a base to a proof mass across the plurality of gaps, a mechanical beam opposite the piezoelectric beam. The sensor may be included in a system including a housing, a support, and electronic components for converting and measuring signal. The sensor may be mounted to an item of furniture, such as a bed or a sofa, and measure physiological data including ballistocardiogram and respiratory data of a subject positioned on the furniture. | 1. A sensor device to be connected to a support and for detecting vibrations, the sensor device comprising: a base for connecting to the support;a proof mass movable relative to the base;a mechanical beam connecting the base to the proof mass, the mechanical beam comprising an inner surface;a piezo beam having a first end and extending to a second end, the piezo beam comprising a sensing material, the first end being directly attached to the base, the second end being directly attached to the proof mass; andat least one ridge extending from the inner surface of the mechanical beam to the piezo beam, such that the at least one ridge is in direct contact with the piezo beam and forms a plurality of gaps between the piezo beam and the mechanical beam, the mechanical beam and the piezo beam being in opposing relationship relative to the at least one ridge such that the movement of the piezo beam is consistent with movement of the proof mass. 2. The sensor device of claim 1 comprising a plurality of ridges. 3. The sensor device of claim 2, wherein the plurality of ridges are spaced at substantially equal distances from one another, such that each of the plurality of gaps has substantially the same stage length. 4. The sensor device of claim 2, wherein each of the plurality of gaps has a different stage length. 5. The sensor device of claim 4, wherein the stage length of the gap closest the proof mass is less than the stage length of the gap closest the base. 6. The sensor device of claim 4, wherein the stage length of each gap decreases in a direction from the base to the proof mass. 7. The sensor device of claim 4, wherein the stage length of each gap in a device having n gaps in which each of the plurality of ridges have substantially the same width, is defined by the equation:lpml1=lpm+2⁢(l1+h)l2=lpm+2⁢(l1+l2+2⁢h)l3=…=lpm+2⁢(l1+l2+…+ln-1+(n-1)⁢h)lnwherein lpm is the length of the proof mass, l1 is the stage length of the gap closest the proof mass, ln is the stage length of the gap closest the base, and h is the width of the ridges. 8. The sensor device of claim 2, wherein each ridge has substantially the same width. 9. The sensor device of claim 1, wherein the mechanical beam is of uniform thickness adjacent each gap. 10. The sensor device of claim 1, wherein the mechanical beam has a different thickness adjacent each gap, the thickness of the mechanical beam increasing adjacent each gap from the proof mass to the base. 11. The sensor device of claim 1, wherein the sensing material comprises lead zirconate titanate. 12. The sensor device of claim 1, wherein the sensing material comprises silicon. 13. The sensor device of claim 1, wherein the base, the proof mass, and the mechanical beam are of monolithic construction. 14. The sensor device of claim 1, wherein the piezo beam is adhered to each of the plurality of ridges. 15. A system for sensing vibrations comprising: a support;a sensor device having a gapped cantilever arrangement, t... | National Science Foundation | G01H011/08 A61B005/11 A61B005/00 G01P015/09
US10722173 | B2 | 2020-07-28 | Methods, systems and computer program products for visualizing anatomical structures and blood flow and perfusion physiology using imaging techniques | 2015-10-13 | Methods for combining anatomical data and physiological data on a single image are provided. The methods include obtaining an image, for example, a raw near-infrared (NIR) image or a visible image, of a sample. The image of the sample includes anatomical structure of the sample. A physiologic map of blood flow and perfusion of the sample is obtained. The anatomical structure of the image and the physiologic map of the sample are combined into a single image of the sample. The single image of the sample displays anatomy and physiology of the sample in the single image in real time. Related systems and computer program products are also provided. | 1. A method for combining anatomical data and physiological data on a single image, the method comprising: obtaining an image of a sample using a single camera, the image of the sample being at least one of a raw near-infrared (NIR) image of the sample or a visible light image of the sample including anatomical structure of the sample;obtaining a physiologic map of blood flow and perfusion of the sample; andcombining the anatomical structure of the sample and the physiologic map of the sample into a single image of the sample, the single image of the sample displaying anatomy and physiology of the sample in the single image in real time,wherein the obtaining a physiologic map of blood flow and perfusion of the sample is carried out using laser speckle imaging (LSI); laser Doppler imaging (LDI); or a blood flow and perfusion angiography resemblance from a fluorescence image; andwherein at least one of the obtaining an image, obtaining a physiologic map and combining is performed by at least one processor. 2. The method of claim 1, wherein obtaining the image comprises obtaining the raw NIR image using a wavelength of from 780 nm to 2500 nm and/or obtaining the visible light image using a wavelength of from 400 nm to 700 nm. 3. The method of claim 1, wherein combining the anatomical structure of the image and the physiologic map of the sample into a single image comprises adjusting one or more properties of the raw NIR image and/or the visible light image and/or the physiologic map. 4. The method of claim 3: wherein the one or more properties comprise at least one of colorization, transparency and a weight function; andwherein the physiologic map illustrates one of blood flow and perfusion, flow distribution, velocity, and/or volume rate of blood flow (cc/min) quantification in primary vessels based on fluid dynamic modeling. 5. The method of claim 3, wherein combining further comprises creating an 8 bit RGB color image represented by the following equation: Img(i,j)=ImgA(i,j)×RGB(i,j)wherein ImgA(i,j) is an 8 bit gray scale visible image of the target tissue/organ, wherein i and j are pixel indexes along horizontal and vertical directions, respectively, and ImgA(i,j) for each color channel is adjusted separately to achieve a desired visualization effect. 6. The method of claim 1: wherein the sample comprises one of tissue and an organ; andwherein the obtained image comprises anatomical structure of a vasculature of at least one of the tissue and the organ. 7. The method of claim 1, wherein obtaining the image is preceded by illuminating the sample with at least one light source, the method farther comprising reflecting a portion of light from the at least one light source to obtain the image and the physiologic map during a single data acquisition. 8. The method of claim 1, further comprising combining a plurality of images with a corresponding plurality of physiologic maps to provide a video displaying anatomy and physiology of the sample ... | East Carolina University | A61B005/00 A61B005/026 G06T005/50 A61B005/0275 G06T007/00 A61B090/00
US10718831 | B2 | 2020-07-21 | Magnetic resonance imaging receive coil with reduced radiation attenuation | 2015-08-07 | A magnetic resonance antenna includes a surface coil and a receive coil. The magnetic resonance antenna includes one or more antenna elements. The magnetic resonance antenna further includes a preamplifier for the antenna element and a coil former for supporting the antenna element. The coil former is formed from a porous material. The antenna is divided into an irradiation zone and at least one reduced radiation zone. The preamplifier for each of multiple antenna elements is located within the at least one reduced radiation zone. The multiple antenna elements are located at least partially within the irradiation zone. The coil former has a perimeter. The irradiation zone extends continuously from a first edge of the perimeter to a second edge of the perimeter. The first edge and the second edge are opposing edges. | 1. A medical instrument, comprising: a linear accelerator (LINAC) with an X-ray source for directing X-ray radiation at a target zone, wherein the LINAC is adapted for rotating the X-ray source about a rotational axis; anda magnetic resonance imaging system for acquiring magnetic resonance data from an imaging zone, wherein the target zone is within the imaging zone, the magnetic resonance imaging system comprising a surface receive coil and a magnet for generating a magnetic field within the imaging zone, wherein the X-ray source is adapted for rotating at least partially about the magnet, wherein the surface receive coil comprises: a plurality of antenna elements;a plurality of preamplifiers for the antenna elements; anda rigid coil former for supporting the antenna elements, the rigid coil former having a perimeter and comprising a porous material, wherein: the rigid coil former is divided in tandem into a first reduced X-ray radiation zone, an X-ray radiation zone and a second reduced X-ray radiation zone, wherein no detuning circuitry is located within the X-ray radiation zone; the preamplifiers for the antenna elements are located within the first reduced X-ray radiation zone; the antenna elements are located at least partially within the X-ray radiation zone; the X-ray radiation zone extends continuously from a first edge of the perimeter to a second edge of the perimeter, with the antenna elements adjacent to and not overlapping one another; the antenna elements are disposed between the first edge and the second edge; the first edge and the second edge are opposing edges; and a cross section of the rigid coil former from the first edge to the second edge is substantially concave when observed from a direction from the rigid coil former to the antenna elements. 2. The medical instrument of claim 1, wherein the porous material is any one of the following: a foam, expanded polypropylene, Polyurethane foam, Polyimide foam, polyether ether ketone (PEEK) foam, a corrugated structure, corrugated cardboard, and a honeycomb structure and/or has an attenuation of less than 2 percent for X-ray radiation between 1.8 MeV and 8 MeV. 3. The medical instrument of claim 1, wherein the cross section is any one of the following: a semi-circle; substantially flat with rounded portions near the first edge and the second edge; and a series of connected straight segments. 4. The medical instrument of claim 1, wherein the rigid coil former comprises one or more mounting fixtures for attaching the surface receive coil to a subject support. 5. The medical instrument of claim 1, wherein the antenna elements are disposed on a flexible printed circuit board, wherein the flexible printed circuit board is attached to the rigid coil former. 6. The medical instrument of claim 1, wherein the first reduced X-ray radiation zone and the second reduced X-ray radiation zone each comprise PIN diodes for controlling detuning of the antenna elements, respectively, wherein th... | KONINKLIJKE PHILIPS N.V. | A61B005/055 A61N005/10 G01R033/34 G01R033/341 G01R033/3415 G01R033/48
US10709341 | B2 | 2020-07-14 | Devices, systems, and methods for pulmonary arterial hypertension (PAH) assessment and treatment | 2013-11-21 | Provided herein are devices, systems, and methods for assessing, treating, and for developing new treatments for pulmonary arterial hypertension (PAH) using pulmonary artery pressure (PAP) values and/or cardiac output (CO) estimates. | 1. A method of evaluating progress of pulmonary arterial hypertension (PAH) in a subject, or an outcome in the subject having PAH, comprising: administering a pharmaceutical agent to the subject;implanting a wireless pressure sensor in a pulmonary artery of the subject;obtaining a pulmonary artery pressure waveform over a cardiac cycle from the wireless pressure sensor;determining one or more pulmonary artery pressure (PAP) values experienced within the pulmonary artery from the pulmonary artery pressure waveform;estimating cardiac output (CO) using one or more of the PAP values experienced within the pulmonary artery over the cardiac cycle;comparing one or more of the PAP values and one or more of the CO estimates to standard PAP and CO values respectively, wherein the standard CO value is based on one or more previously estimated CO values from the subject;using the comparisons of PAP and CO to monitor a PAH progression or outcome in the subject; andchanging at least one of the pharmaceutical agent, dosage or dosing regimen based on the comparison of the PAP and CO. 2. The method of claim 1, wherein the standard PAP value is one or more previously determined PAP values from the subject. 3. The method of claim 1, wherein: a. a decrease in the one or more PAP values compared to the standard PAP value combined with one or more increased or stable CO estimates compared to the standard CO value, orb. one or more stable PAP values compared to the standard PAP value combined with one or more increased CO estimates compared to the CO standard value indicates or predicts improvement of the PAH in the subject. 4. The method of claim 1, wherein: a. a decrease in one or more CO estimates compared to the standard CO value combined with one or more increased or stable PAP values compared to the standard PAP value, orb. an increase in one or more PAP values compared to the standard PAP values combined with a stable or decreased CO value compared to the standard CO value indicates or predicts worsening of the PAH in the subject. 5. The method of claim 1, wherein the pressure sensor lacks percutaneous connections. 6. The method of claim 5, wherein the pressure sensor is energized from an external source. 7. The method of claim 6, wherein the pressure sensor is a passive sensor energized to return pressure readings by an electromagnetic field. 8. The method of claim 1, wherein the pulmonary artery pressure waveform is obtained while the subject is exercising. 9. The method of claim 1, wherein the pulmonary artery pressure waveform is obtained while the subject is asleep. 10. The method of claim 1, further comprising modifying or maintaining the administration of the pharmaceutical agent. 11. The method of claim 10, further comprising maintaining administration of the pharmaceutical agent at its current dosage and or dosing regimen when improvement is indicated. 12. A method for evaluating treatment of a subject having pulmonary arterial hypertension (... | ST. JUDE MEDICAL LUXEMBOURG HOLDINGS II S.A.R.L. | A61B005/02 A61B005/0215 G16H050/20 A61B005/029 A61B005/00 G06F019/00
US10716490 | B2 | 2020-07-21 | Wound dressing with impedance sensor | 2010-07-09 | A wound sensor comprising at least one electrode (10a, 10b) and a non-adherent porous layer (25) proximate at least part of the electrode. | 1. A wound sensor comprising: at least one electrode and an external measurement device coupled or configured to couple with the at least one electrode, the external measurement device being configured to receive electrical measurement signals from the at least one electrode and to determine impedance of a wound or wound fluid using the received electrical measurement signals, the at least one electrode having a sensing surface and a signal carrying portion, the signal carrying portion being configured for carrying the electrical measurement signals from the sensing surface to the external measurement device, the signal carrying portion and the sensing surface both being disposed on an electrically insulating substrate having insulating properties during use of the wound sensor;an insulating layer that covers said signal carrying portion, such that said signal carrying portion is positioned between the substrate and the insulating layer; and a non-adherent porous layer in direct contact with the sensing surface, wherein:said sensing surface is positioned between the substrate and the non-adherent porous layer that is in direct contact with the sensing surface, such that the non-adherent porous layer that is in direct contact with the sensing surface and the substrate at least partly define a bounded volume in which the sensing surface is provided; andthe non-adherent porous layer directly overlies and contacts the sensing surface, forms a wound contact surface, and collects a volume of wound fluid within the bounded volume that is measurable by the external measurement device using the at least one electrode to determine the impedance. 2. A wound sensor according to claim 1, wherein the non-adherent porous layer is a biocompatible and/or non-irritant surface, material or coating. 3. A wound sensor according to claim 1, wherein the sensor comprises two or more electrodes. 4. A wound sensor according to claim 1, wherein the at least one electrode is printed. 5. A wound sensor according to claim 1, wherein the non-adherent porous layer is separate from the at least one electrode. 6. A wound sensor according to claim 1, wherein the non-adherent porous layer fully covers the at least one electrode. 7. A wound sensor according to claim 1, wherein the non-adherent porous layer is non-conducting. 8. A wound sensor according to claim 1, wherein the non-adherent porous layer is adjacent to at least part of the at least one electrode. 9. A wound sensor according to claim 1, wherein the sensor is arranged such that the non-adherent porous layer is provided on an opposite side of the at least one electrode to the substrate. 10. A wound sensor according to claim 1, wherein the non-adherent porous layer comprises silicone. 11. A wound sensor according to claim 1, wherein the non-adherent porous layer comprises a cellulosic material. 12. A wound sensor according to claim 1, wherein the non-adherent porous layer is between 50 500 pm thick. 13. A wound... | University of Strathclyde | A61B005/04 A61B008/00 A61B005/053 A61B005/00 A61F013/02 A61F013/00
US10712329 | B2 | 2020-07-14 | Methods of measuring metal pollutants on skin | 2018-07-03 | A method for measuring an airborne metal pollutant on skin can include the taking of a skin sample and measuring of a target metal utilizing inductively coupled plasma tandem mass spectrometry. | 1. A method of identifying a skin cleanser which can reduce the amount of airborne pollution on skin, comprising: a) identifying a subject with pollution on at least a portion of the skin;b) applying a first tape strip to a skin collection site on the subject;c) removing the first tape strip with a first skin sample from the first skin collection site;d) placing the first tape strip in a container;e) washing a comparable second skin collection site with a skin cleanser;f) allowing the second skin collection site to dry;g) applying a second tape strip to the second skin collection site;h) removing the second tape strip with a second skin sample from the second skin collection site;i) placing the second tape strip in a container;j) digesting the first skin sample to form a first digestion solution and the second skin sample to form a second digestion solution; andk) measuring the level of a target metal from the first digestion solution and second digestion solution using inductively coupled plasma tandem mass spectrometry;wherein a decrease of 10% or more from the first strip to the second strip indicates the ability of the skin cleanser to remove at least a portion of the target metal from the surface of the skin. 2. The method of claim 1, further comprising comparing the level of the target metal in the skin samples to the geological record to determine if the level of target metal is anthropogenically enriched. 3. The method of claim 1, wherein the tape strips have a peel force of 3 g/cm or more. 4. The method of claim 1, wherein the tape strips have a surface area of about 1 cm2 to about 12 cm2. 5. The method of claim 1, further comprising correcting for the tape strip in the metal measurement by measuring the level of the target metal from a blank tape strip in the same manner as the tape strips placed on the skin and subtracting the value of the target metal in the blank tape strip from the value of the target metal in the skin samples. 6. The method of claim 1, wherein the target metal comprises lead, sodium, magnesium, potassium, calcium, vanadium, chromium, manganese, iron, nickel, copper, zinc, arsenic, rubidium, strontium, cadmium, cesium, lanthanum, uranium, or a combination thereof. 7. The method of claim 1, wherein after the tape strips are placed on the skin, the skin is marked near one or more edges of the tape strips, so that subsequent tape strips can be placed in the same areas. 8. The method of claim 1, wherein the tape strips are removed from the skin using a tool. 9. The method of claim 8, wherein the tool is a tweezer or a forceps. 10. The method of claim 1, wherein a-k are repeated about 20 times or until the glistening layer is reached. | The Procter & Gamble Company | G01N033/20 G01N001/02 G01N001/28 A61B010/02 A61B005/00 G01N001/40 C12M001/34
US10725652 | B2 | 2020-07-28 | Wearable apparatus for continuous blood glucose monitoring | 2016-12-27 | A system, a method, and a computer program product for providing wearable continuous blood glucose monitoring. In some embodiments, there is provided a method that includes receiving, at a smartwatch, an alert representative of a glucose state of a host-patient coupled to a glucose sensor; detecting, at the smartwatch, a predetermined action indicative of a request to generate a glance view providing an indication of the glucose state of the host-patient; and presenting, at the smartwatch and in response to the detecting, the glance view providing the indication of the glucose state of the host-patient. | 1. A method comprising: receiving, at a smartwatch, an alert representative of a glucose state of a host-patient coupled to a glucose sensor;detecting, at the smartwatch, a predetermined action indicative of a request to generate a glance view providing an indication of the glucose state of the host-patient;presenting, at the smartwatch and in response to the detecting, the glance view providing the indication of the glucose state of the host-patient;selecting, at the smartwatch, a handoff icon;sending, from the smartwatch, a message to a second device, the message indicating a current view at the smartwatch; andpresenting, at the second device, a user interface view that is either the glance view or is different from the glance view. 2. The method of claim 1, wherein the alert is received via a low power radio access transceiver at the smartwatch. 3. The method of claim 2, wherein the low power radio access transceiver is configured in accordance with at least one of Bluetooth, Bluetooth Low Energy, or NFC. 4. The method of claim 2, wherein the alert is received from a receiver wirelessly coupled to the glucose sensor via transmitter sensor electronics. 5. The method of claim 4, wherein the receiver comprises at least one of a smartphone and/or a tablet. 6. The method of claim 4, wherein the receiver includes a continuous blood glucose application configured to interact with the smartwatch. 7. The method of claim 1, wherein the alert is received from a remote server. 8. The method of claim 1, wherein the detecting of the predetermined action triggers the smartwatch to replace a home screen and/or a default user interface view with the glance view. 9. The method of claim 1, further comprising: triggering, at the smartwatch and in response to the detecting, a feedback comprising a haptic indicator, an audio indicator, and/or a visual indicator. 10. The method of claim 1, wherein the predetermined action comprises a selection of a physical button. 11. The method of claim 1, wherein the predetermined action comprises a detection of a certain wrist movement. 12. The method of claim 1, wherein the predetermined action comprises a detection of a certain eve movement of a wearer of the smartwatch. 13. The method of claim 1, further comprising: detecting a swipe pattern;mapping the detected swipe pattern to at least one of a plurality of glance views, each of which is mapped to a different swipe pattern; andpresenting, based on the detected swipe pattern, the at least one glance view. 14. The method of claim 1, further comprising: mapping the glucose state to at least one of a plurality of glance views, each of which is mapped to a different glucose state; andpresenting, based on the glucose state, the at least one glance view. 15. The method of claim 1, further comprising: determining a role of a wearer of the smartwatch;mapping the role of the wearer to at least one of a plurality of glance views, each of which is mapped to a different ro... | DexCom, Inc. | A61B005/00 G06F003/0488 G08B021/04 A61B005/145 G06F001/16 H04W004/80
US10709371 | B2 | 2020-07-14 | System and methods for serving a custom meditation program to a patient | 2016-09-09 | One variation of a method for addressing psychological conditions of a patient through guided meditation includes: serving a first psychological test to the patient through a virtual patient portal; transforming results of the first psychological test into a set of psychological conditions of the patient; matching a first predefined meditation path to the patient based on a severity of each psychological condition in the set of psychological conditions; serving a first guided meditation track in the first meditation path to the patient through the virtual patient portal; serving a second psychological test to the patient through the virtual patient portal; calculating new intensities of each psychological condition in the set of psychological conditions of the patient based on results of the second psychological test; and matching a second predefined meditation path to the patient based on the new intensities of each psychological condition in the set of psychological conditions. | 1. A method for automatically prescribing guided meditation based on biometric signals of a patient, the method comprising: serving a first psychological test to the patient through a virtual patient portal executing on a computing device;transforming results of the first psychological test into a set of psychological conditions of the patient;matching intensities of psychological conditions in the set of psychological conditions to a particular template psychological test result in a set of template psychological test results;assigning a first meditation path, from a set of predefined meditation paths, to the patient based on a severity of each psychological condition in the set of psychological conditions, the first mediation path associated with the particular template psychological test result;at a first time, serving a first guided meditation track in the first meditation path to the patient through the virtual patient portal;calculating a meditation score comprising a quantitative value inversely proportional to a magnitude of alpha waves read from an electroencephalography headset worn by the patient during completion of the first meditation path by the patient;at a second time succeeding completion of the first meditation path by the patient, serving a second psychological test to the patient through the virtual patient portal;calculating new intensities of each psychological condition in the set of psychological conditions of the patient based on results of the second psychological test;matching a second meditation path, from the set of predefined meditation paths, to the patient based on the new intensities of each psychological condition in the set of psychological conditions and the meditation score; andat a third time succeeding the second time, serving a second guided meditation track in the second meditation path to the patient through the virtual patient portal. 2. The method of claim 1: wherein assigning the first meditation path to the patient comprises selecting the first meditation path associated with a first difficulty level;wherein matching the second meditation path to the patient comprises selecting the second meditation path associated with a second difficulty level greater than the first difficulty level in response to the meditation score of the patient trending upwardly through the first meditation path. 3. The method of claim 1: wherein calculating the meditation score comprises, at the computing device, downloading electroencephalography data from the electroencephalography headset during completion of the first guided meditation track in the first meditation path by the patient at a medical facility; andfurther comprising, at a personal mobile computing device carried by the patient, serving a survey prompting the patient to indicate severity of a psychological condition in the set of psychological conditions following completion of a second guided meditation track in the first meditation path by the user outside... | WellBrain, Inc. | A61B005/16 G16H010/20 G16H050/20 A61B005/00
US10722137 | B2 | 2020-07-28 | Systems and methods for accelerated MR thermometry | 2015-04-02 | Aspects of the present disclosure relate to magnetic resonance thermometry. In one embodiment, a method includes acquiring undersampled magnetic resonance data associated with an area of interest of a subject receiving focused ultrasound treatment, and reconstructing images corresponding to the area of interest based on the acquired magnetic resonance data, where the reconstructing uses Kalman filtering. | 1. A method for accelerated magnetic resonance thermometry, comprising: acquiring, by applying a magnetic; resonance imaging sequence generated by a magnetic resonance imaging device, undersampled k-space magnetic resonance data associated with an area of interest of a subject receiving focused ultrasound treatment, wherein the magnetic resonance imaging sequence comprises a three-dimensional interleaved stack-of-spirals spoiled gradient echo sequence; andreconstructing images corresponding to the area of interest based on the acquired magnetic resonance data, wherein the reconstructing of the imams comprises:reconstructing the images from the undersampled k-space data using Kalman filtering, wherein the Kalman filtering comprises using a dynamic state-space model for the Kalman filtering,wherein the center of k-space is fully sampled and provides low resolution images used for initial values and training data for noise covariance for the Kalman filtering, andwherein the dynamic state-space model is according to: xk=f(xk-1)+wk andyk=UxFxk+vk wherein xk is target image at kth frame, yk is corresponding acquired data, f(xk-1) is state transition function, F is Fourier transform operator, Uk is undersampling scheme at acquisition k, and w and v are system and measurement noise. 2. The method of claim 1, further comprising: measuring phase changes from proton-resonance frequency (PRF) shift associated with heating in the area of interest caused by the focused ultrasound treatment. 3. The method of claim 2, further comprising: based on at least one of the measured phase changes and images generated from the reconstruction, determining one or more characteristics associated with a physiological activity in the area of interest. 4. The method of claim 1, wherein reconstructing the images includes reconstruction by 2D gridding. 5. A system for accelerated magnetic resonance thermometry, comprising, a magnetic resonance imaging device configured to apply a magnetic resonance imaging sequence configured to acquire undersampled k-space magnetic resonance data associated with an area at interest of a subject receiving focused ultrasound treatment, wherein the magnetic resonance imaging sequence comprises a three-dimensional interleaved stack-of-spirals spoiled gradient echo sequence; andan image reconstruction device configured to reconstruct images corresponding to the area at interest based on the acquired magnetic resonance data, wherein the reconstructing of the images comprises:reconstructing the images from the undersampled k-space data using Kalman filtering, wherein the Kalman filtering comprises using a dynamic state-space model for the Kalman filtering,wherein the center of k-space is fully sampled and provides low resolution images used for initial values and training data for noise covariance for the Kalman filtering, andwherein the dynamic; state-space model is according to: xkf(xk-1)+wk andyk=UxFxk+vk wherein xk is target image at kth frame... | University of Virginia Patent Foundation | A61B005/055 A61N007/02 A61B005/01 A61B005/00 A61B018/00 A61B090/00
US10722131 | B2 | 2020-07-28 | Body-worn system for continuous, noninvasive measurement of cardiac output, stroke volume, cardiac power, and blood pressure | 2011-12-27 | The invention provides a system for measuring stroke volume (SV), cardiac output (CO), and cardiac power (CP) from a patient that features: 1) an impedance sensor connected to at least two body-worn electrodes and including an impedance circuit that processes analog signals from the electrodes to measure an impedance signal (e.g. TBEV waveform); 2) an ECG sensor connected to at least two chest-worn electrodes and including an ECG circuit that processes analog signals from the electrodes to measure and ECG signal; 3) an optical sensor connected to a body-worn optical probe and including an optical circuit that processes signals from the probe to measure at least one optical signal (e.g. a PPG waveform) from the patient; 4) a processing system, typically worn on the patient's wrist and connected through a wired interface to the optical sensor, and through either a wired or wireless interface to the TBEV and ECG sensors. | 1. A system for simultaneously measuring blood pressure and stroke volume of a patient, comprising: an impedance sensor comprising at least two electrodes and an impedance circuit, the at least two electrodes configured to be positioned on the patient in the clavicle/brachial region, and the impedance circuit configured to receive and process signals from the at least two electrodes to measure a transbrachial electro-velocimetry (TBEV) signal from the patient;an optical sensor configured to be positioned on the patient and comprising an optical probe configured to be positioned on a digit on a hand of the patient and an optical circuit, the optical circuit configured to receive and process signals from the optical probe to measure at least one optical signal from the patient;an ECG sensor configured to be positioned on the patient and comprising at least two electrodes configured to be positioned on the patient's torso and an ECG circuit, the ECG circuit configured to receive and process signals from the at least two electrodes to measure an ECG signal from the patient; and,a processing system configured to: (i) be worn on the patient's body; (ii) be operably connected to the impedance sensor, optical sensor, and the ECG sensor; (iii) receive the TBEV signal from the impedance sensor; (iv) receive the at least one optical signal from the optical sensor; (v) receive the ECG signal from the ECG sensor and collectively process the ECG signal and the at least one optical signal to determine a blood pressure value; (vi) process the TBEV signal to determine a time dependent derivative of the TBEV signal (d(TBEV)dZ/dt), a first zero value in d(TBEV)dZ/dt indicating onset of a pulse in d(TBEV)dZ/dt representing opening of the patient's aortic valve, a second zero value in d(TBEV)dZ/dt after onset of the pulse indicating closing of the patient's aortic valve, a d(TBEV)dZ/dtmax value; (vii) determine a first systolic flow time (SFT) calculated using the first zero value in d(TBEV)dZ/dt and the second zero value in d(TBEV)dZ/dt; (viii) collectively process d(TBEV)dZ/dtmax and the first SFT value to determine a first stroke volume (SV) value; (ix) process the at least one optical signal to determine a second SFT value; (x) collectively process d(TBEV)dZ/dtmax and the second SFT value to determine a second SV value; (xi) collectively process the first and second SV values to determine a third SV value; and (xii) display the third SV value on a display device. 2. The system of claim 1, wherein the processing system is further configured to determine a time difference between the ECG signal and the at least one optical signal. 3. The system of claim 2, wherein the processing system is further configured to determine a pulse arrival time separating a QRS complex in the ECG signal and an onset of a pulse in the at least one optical signal. 4. The system of claim 3, wherein the processing system is further configured to process the pulse arrival time and a bloo... | SOTERA WIRELESS, INC. | A61B005/0295 A61B005/029 A61B005/053 A61B005/08 A61B005/11 A61B005/1455 A61B005/00 A61B005/01 A61B005/021 A61B005/02 A61B005/0205 A61B005/145 G06F019/00 A61B005/024 A61B005/0245
US10726740 | B2 | 2020-07-28 | Image processing device and image processing method | 2015-10-09 | Provided is an image processing device including an infrared image acquisition unit that acquires an infrared image of an imaged object, a visible light image acquisition unit that acquires a visible light image of the imaged object, a generation unit that generates cutaneous sensation control parameters on the basis of the infrared image acquired by the infrared image acquisition unit, and a data processing unit that associates the visible light image acquired by the visible light image acquisition unit with the cutaneous sensation control parameters generated by the generation unit. | 1. An image processing device, comprising: an infrared image acquisition unit configured to acquire a plurality of infrared images at a plurality of infrared ray irradiation angles with respect to an imaged object;a visible light image acquisition unit configured to acquire a visible light image of the imaged object, wherein each pixel region of a plurality of pixel regions of an infrared image of the plurality of infrared images corresponds to a set of pixels of the visible light image;a generation unit configured to: estimate a surface roughness of the imaged object in each pixel region of the plurality of pixel regions based on a distribution of pixel values corresponding to each pixel region of the plurality of pixel regions at the plurality of infrared ray irradiation angles;generate a first parameter of a plurality of cutaneous sensation control parameters based on the estimated surface roughness of the imaged object and pixel values of the plurality of infrared images;estimate a density of texture boundaries on a surface of the imaged object based on a dispersion of the estimated surface roughness of the imaged object; andgenerate a second parameter of the plurality of cutaneous sensation control parameters based on the estimated density of the texture boundaries on the surface of the imaged object and the pixel values of the plurality of infrared images, wherein the plurality of cutaneous sensation control parameters includes tactile sensation control parameters, andthe tactile sensation control parameters include the first parameter to control a first tactile sensation corresponding to the surface roughness of the imaged object, and the second parameter to control a second tactile sensation corresponding to the density of the texture boundaries on the surface of the imaged object; anda data processing unit configured to assign the plurality of cutaneous sensation control parameters to the set of pixels of the visible light image. 2. The image processing device according to claim 1, wherein the first parameter includes a value that indicates a frequency of an oscillation of a target, andthe second parameter includes a value that indicates a duration time of the oscillation. 3. The image processing device according to claim 1, wherein the first parameter includes a value of a voltage applied to targets, andthe second parameter includes a value that indicates an interval of the targets. 4. The image processing device according to claim 1, wherein the plurality of cutaneous sensation control parameters further includes temperature sensation control parameters, andthe generation unit is further configured to generate the temperature sensation control parameters based on the pixel values of the plurality of infrared images. 5. The image processing device according to claim 4, wherein the temperature sensation control parameters include an absolute value of a current applied to a target and a value indicating a direction of the current. 6.... | SONY CORPORATION | G09B021/00 A61B005/01 A61B005/00 H04N005/33 G06F003/03 G06F003/01 H04N005/247 G06T007/13 G06T007/40
US10702156 | B2 | 2020-07-07 | Systems and methods for improved tractography images | 2016-10-12 | The present disclosure discusses systems and methods for identifying biomarkers that can help with the diagnosis, prognosis, and treatment choices of patients with neurodegenerative diseases. Diffusion based magnetic resonance imaging can often fail for patients with a neurodegenerative disease because parameters fractional anisotropy, mean diffusivity, and radial diffusivity are based on simple models that can fail in the presence of neurodegeneration, such as demyelination. The present disclosure discusses systems and methods that enhance dMRI images and enable tractography to be performed on images of a damaged nervous system. The damaged tracks identified by the present system can be used as a biomarker for the assessment of patients. In some implementations, the biomarkers are converted into clinical scales that can be used to compare patients to one another or over time. | 1. A method comprising: receiving anatomical image data and diffusion-weighted data of a patient's nervous system;generating a segment from the anatomical image data identifying a first volume of voxels;generating enhanced diffusion-weighted data in the segment responsive to receiving the diffusion-weighted data;generating a track comprising a second volume of voxels in the enhanced diffusion-weight data;registering the segment identifying the first volume of voxels with the track in the enhanced diffusion-weight data, the segment located in the same space as the track, the segment configured for comparison with the track comprising the second volume of voxels; andcalculating a damage score to the track based on the first volume of voxels, the second volume of voxels, and the segment registered with the track. 2. The method of claim 1, wherein each of the one or more segments correspond to damage identified in the anatomical image data. 3. The method of claim 2, wherein the damage comprises at least one of a brain lesion and a plaque. 4. The method of claim 1, wherein generating the enhanced diffusion-weighted data comprises convolving the diffusion-weighted data with a kernel representing Brownian motion. 5. The method of claim 1, wherein the track is one of a corpus callosum, an optical radiation, and a corticospinal tract. 6. The method of claim 1, further comprising generating the track responsive to a track template. 7. The method of claim 1, further comprising determining a number of voxels of the second volume of voxels contained within the segment. 8. The method of claim 7, wherein calculating the damage is calculated responsive to the number of voxels of the second volume of voxels contained within the segment and the second volume of voxels. 9. The method of claim 1, further comprising calculating a clinical score responsive to the determined damage to the track. 10. The method of claim 1, wherein the anatomical image data comprises one or more of T1, T2, HARDI, and Mill data. 11. A system comprising: an assessment engine comprising one or more processors and a processor readable medium storing instructions that when executed by the one or more processors cause the one or more processors to: receive anatomical image data and diffusion-weighted data of a patient's nervous system;generate a segment from the anatomical image data identifying a first volume of voxels;generate an enhanced diffusion-weighted data responsive to the diffusion-weighted data;generate a track comprising a second volume of voxels in the enhanced diffusion-weight data;register the segment identified as the first volume of voxels with the track in the enhanced diffusion-weight data, the segment located in the same space as the track, the segment configured for comparison with the track comprising the second volume of voxels; andcalculate a damage score to the track based on the first volume of voxels, the second volume of voxels, and the segment registered ... | Mint Labs Inc. | A61B005/00 G06T007/00 G16H050/20 A61B005/055 G06F019/00 G01R033/48 G01R033/50 G01R033/563
US10722302 | B2 | 2020-07-28 | Methods and systems to combine RF ablation therapy with device navigation | 2016-09-27 | Methods and systems for combining ablation therapy with navigation of the ablation device. An ablation system may be configured for use with one of two methods to prevent loss of navigation signals during ablation energy delivery. In the first method, ablation energy signals are filtered from the navigation signal. In the second method, the delivery of ablation energy is sequenced with the delivery of navigation energy such that ablation energy and navigation energy are not delivered at the same time and navigation signals received by the system are time-division multiplexed to reconstruct the navigation signals and determine a location of the device within the patient. | 1. A system for determining a location of a medical device during a procedure in which ablation energy is delivered to a tissue area within a patient, the system comprising: a medical device including at least one energy delivery electrode and at least one mapping electrode;an ablation energy source in electrical communication with the at least one energy delivery electrode, the ablation energy source being configured to deliver ablation energy to the at least one energy delivery electrode during an on period and stop the delivery of ablation energy during an off period;an energy detector in communication with the ablation energy source, the energy detector being configured to determine when energy is being delivered from the ablation energy source to the at least one energy delivery electrode;a navigation system being in communication with the at least one mapping electrode of the medical device;a switching system in communication with the ablation energy source and the navigation system, the switching system being configured to selectively place the at least one energy delivery electrode in communication with the energy source or the navigation system; anda logic apparatus in communication with the energy detector and the switching system, the logic apparatus being configured to selectively open the switching system when energy is being delivered from the ablation energy source to the at least one energy delivery electrode and close the switching system when no energy is being delivered from the ablation energy source to the at least one energy delivery electrode, when energy is first detected by the energy detector, the switching system recognizes the delivery of energy, the switching system remains closed, and the medical device is in communication with the navigation system, then after a period where no energy is detected by the energy detector, the switching system opens and disconnects the navigation system from the medical device. 2. The system of claim 1, further comprising a switch between the ablation energy source and the navigation system, and the switching system further includes a shunt between the switch and the navigation system. 3. The system of claim 1, wherein the navigation system further includes a plurality of external electrodes and an alternating current source that is configured to deliver alternating current from a plurality of sources to the plurality of external electrodes. 4. The system of claim 3, wherein the alternating current source is configured to deliver alternating current electricity to each of the plurality of sources at a same frequency. 5. The system of claim 4, wherein the navigation system includes a processing unit having processing circuitry that is configured to time-division multiplex the frequency for each of an X plane, a Y plane, and a Z plane. 6. The system of claim 5, wherein the processing circuitry is further configured to time-division multiplex voltage signals received from the at leas... | Medtronic Ablation Frontiers LLC | A61B018/14 A61B005/00 A61B005/042 A61B005/0428 A61B018/12 A61B018/00 A61B034/20 A61B090/00 A61B017/00 A61B018/02
US10704809 | B2 | 2020-07-07 | System for warming-up and cooling-down a superconducting magnet | 2016-12-20 | A cryogenic refrigerator system with heaters is constructed in modular form to serve as a portable servicing system to warm up and then cool down a target object by circulating a gaseous cryogen through a target object cryostat without moving the target object or breaking its vacuum. The main module is a refrigerator cryostat containing a fan that circulates gas through one or more heat exchangers which can warm or cool the gas by heaters and by one or more GM or Brayton cycle expanders. Additional components including one or more compressors, a gas charge and vent assembly, a control system, gas lines, power lines, and vacuum jacketed transfer lines can be assembled in the main module or additional modules. An example is a system that can be wheeled through a hospital to service a MRI cryostat. | 1. A servicing system for servicing a target object disposed in a target object cryostat, the servicing system comprising: one or more expanders and one or more compressors for cooling-down the target object to a cryogenic temperature;one or more heaters for warming-up the target object from the cryogenic temperature;a fan and a fan circuit, the fan circulating gas through the fan circuit; anda refrigerator cryostat housing the one or more expanders and the fan which are surrounded by a vacuum;a plurality of heat exchangers for transferring heat from the heater to the gas in the fan circuit during warm-up of the target object and for transferring heat from the target object to the gas in the fan circuit during cool-down of the target object;a controller and a plurality of sensors, the plurality of sensors for measuring a temperature and a pressure in the fan circuit, the controller responsive to the plurality of sensors controlling at least a speed of the fan, a speed of the expanders, and a power of the heaters;a removable supply transfer line for bringing gas from the fan circuit to the target object cryostat;a removable return transfer line for returning gas from the target object to the fan circuit, the transfer lines being interchangeably connected to the target object cryostat;a pressure relief valve for relieving excess pressure in the fan circuit; anda vent valve that can be actuated to open to one of atmospheric pressure and a gas recovery system. 2. A system in accordance with claim 1, further comprising a cart for moving the servicing system, the cart and servicing system fitting through a 78 cm wide and 160 cm high doorway. 3. A system in accordance with claim 1, wherein each expander is a GM type expander and each heat exchanger is an integral part of a respective cold end of the expander. 4. A system in accordance with claim 1, wherein each expander is a Brayton type expander and each heat exchanger is separate from a respective cold end of the expander. 5. A system in accordance with claim 1, wherein the heat exchanger for cooling the gas is different from the heat exchanger that warms the gas. 6. A system in accordance with claim 1, wherein the pressure relief valves are set to relieve excess pressure in the fan circuit of less than 200 kPa above atmospheric pressure. 7. A method of servicing a target object disposed in a target object cryostat using a service system, the target object cryostat having a supply port for receiving cold gas during cool down and a return port, the service system comprising: one or more expanders and one or more compressors for cooling-down the target object to a cryogenic temperature;one or more heaters for warming-up the target object from the cryogenic temperature;a fan and a fan circuit, the fan circulating gas through the fan circuit; anda refrigerator cryostat housing the one or more expanders and the fan which are surrounded by a vacuum;a plurality of heat exchangers for transferring heat ... | SUMITOMO (SHI) CRYOGENICS OF AMERICA, INC. | F25B009/14 F17C003/08 A61B005/055
US10716475 | B2 | 2020-07-21 | Localized monitoring | 2013-09-25 | This document relates to receiving and processing signals during cardio-pulmonary resuscitation (CPR) treatment. This document further relates to establishing a localized, patient-specific network for communication of information from one or more sensors to a computing device. | 1. A medical device system for providing a localized, patient-specific wireless network, the system comprising: an external defibrillator configured to provide a wireless network according to a wireless transmission protocol, wherein the external defibrillator includes a first wireless module that comprises a first transmitter and a first receiver configured to communicate via the wireless network;at least one sensor configured to be placed in contact with a skin of a patient and obtain patient-specific data, wherein the at least one sensor includes a second wireless module that comprises a second transmitter and a second receiver configured to communicate with the external defibrillator via the wireless network; andat least one defibrillator pad configured to be placed in contact with the skin of the patient, and configured to communicate with the external defibrillator over a wired electrical connection,wherein the external defibrillator is configured to communicate with the at least one sensor via a first signal transmission path and a second signal transmission path, the first signal transmission path comprising the wired electrical connection between the external defibrillator and the at least one defibrillator pad, and the second signal transmission path comprising a body of the patient as a conductive medium between the at least one defibrillator pad and the at least one sensor, and the external defibrillator is further configured to: establish the localized, patient-specific wireless network, wherein the localized, patient-specific wireless network includes the at least one sensor and the external defibrillator,perform a secure authentication process comprising an exchange of multiple authentication signals between the at least one sensor and the external defibrillator, the multiple authentication signals comprising: an initial signal transmitted from the second wireless module of the at least one sensor to the first wireless module of the external defibrillator via the wireless network according to the wireless transmission protocol, wherein the initial signal establishes communication between the at least one sensor and the external defibrillator, andan electrical reply signal from the external defibrillator to the at least one sensor, in response to the initial signal, transmitted over the first signal transmission path and the second signal transmission path; andsubsequent to the secure authentication process, enable the at least one sensor to send and receive signals via the localized, patient-specific wireless network; andwherein the secure authentication process ensures that the at least one sensor is correctly paired with the external defibrillator and the localized, patient-specific wireless network associated with the patient so that the patient-specific data is only received by recipients authorized to view and use the patient-specific data. 2. The system of claim 1, wherein the external defibrillator is configured to initiate... | Zoll Medical Corporation | A61B005/00 A61B005/0402 A61N001/39 A61B005/0205 A61B005/145 A61B005/021 A61B005/08
US10726948 | B2 | 2020-07-28 | Medical imaging device- and display-invariant segmentation and measurement | 2017-12-08 | Medical imaging device- and display-invariant segmentation and measurement is provided. In various embodiments, a plurality of medical images is read from a data store. Metadata of each of the plurality of medical images is read. The metadata identifies an image acquisition device associated with each of the plurality of medical images. Based on the plurality of medical images and the metadata of each of the plurality of images, a learning system is trained to determine one or more image correction parameters. The one or more image correction parameters optimize segmentation of the plurality of medical images. | 1. A method comprising: reading a plurality of medical images from a data store;reading metadata of each of the plurality of medical images, the metadata identifying an image acquisition device associated with each of the plurality of medical images;based on the plurality of medical images and the metadata of each of the plurality of images, training a learning system to determine one or more image correction parameters, the one or more image correction parameters adapted to optimize the plurality of medical images for segmentation. 2. The method of claim 1, further comprising: after training, applying the learning system to determine one or more image correction parameters of an additional medical image. 3. The method of claim 2, further comprising: applying the determined one or more image correction parameters to the additional medical image to thereby generate a corrected medical image; andsegmenting the corrected medical image. 4. The method of claim 1, wherein training the learning system comprises: determining a segmentation of each of the plurality of medical images;comparing the segmentation to ground truth. 5. The method of claim 1, further comprising: training the learning system to determine one or more segmentation parameters, the one or more segmentation parameters optimizing segmentation of the plurality of medical images. 6. The method of claim 5, further comprising: determining an expected error range of the segmentation. 7. The method of claim 1, wherein the one or more image correction parameters includes window, level, noise, or smoothing. 8. The method of claim 1, wherein the metadata further identifies one or more image acquisition parameter associated with each of the plurality of medical images. 9. The method of claim 1, wherein the one or more image correction parameters includes exposure, intensity, image reconstruction kernels, acquisition pixel and slice spacing, contrast agent dose amount, light power, noise, or contrast level. 10. The method of claim 1, wherein the learning system comprises a genetic algorithm. 11. The method of claim 1, further comprising: further training the learning system based on user-generated segmentation. 12. The method of claim 1, wherein the metadata further comprises an indication of a subject anatomy of the plurality of medical images. 13. A system comprising: a datastore;a computing node comprising a computer readable storage medium having program instructions embodied therewith, the program instructions executable by a processor of the computing node to cause the processor to perform a method comprising: reading a plurality of medical images from the data store;reading metadata of each of the plurality of medical images, the metadata identifying an image acquisition device associated with each of the plurality of medical images;based on the plurality of medical images and the metadata of each of the plurality of images, training a learning system to determine one or more im... | INTERNATIONAL BUSINESS MACHINES CORPORATION | G16H030/20 A61B005/00 G06T007/11 G16H030/40 G06T007/136 G06T007/00
US10702245 | B2 | 2020-07-07 | Method of detecting microbubbles in a vessel | 2017-05-04 | The present invention provides a method of detecting microbubbles in a vessel of an affected part, comprising aggregates the microbubbles, acquiring phase-contrast magnetic resonance images and analyzing the phase-contrast magnetic resonance images. Thus, we can detect or monitor the size and location of the microbubbles in vessels of any part of body. | 1. A method of detecting microbubbles in a vessel of an affected part, comprising: delivering an ultrasonic energy within the affected part to aggregate the microbubbles to form a plurality of aggregated microbubbles;acquiring a phase-contrast magnetic resonance image by a magnetic resonance device; andlocating the plurality of aggregated microbubbles by analyzing the phase-contrast magnetic resonance image to obtain a velocity value and a vorticity value of fluid flow in a region of interest in the phase-contrast magnetic resonance image;wherein the plurality of aggregated microbubbles is located at wherein the velocity value is within the lowest 10%, and the vorticity value is −0.18 to 0.18;wherein the ultrasonic energy has a frequency of 0.83-1.25 MHz and an acoustic pressure of 0.025-0.1 MPa; andwherein the region of interest is an area of the vessel. 2. The method as claimed in claim 1, wherein the region of interest is an area of the vessel. 3. The method as claimed in claim 1, wherein the microbubbles have a diameter of 1-1.5 μm. 4. The method as claimed in claim 1, wherein the microbubbles were substantially composed of C3F8 gas or C5F12 droplet core encapsulated by a lipid shell. 5. The method as claimed in claim 4, wherein the microbubbles further comprise a drug for treating the affected part. 6. The method as claimed in claim 1, wherein the phase-contrast magnetic resonance image is a plurality of phase-contrast magnetic resonance images in a time sequence. 7. The method as claimed in claim 6, further comprising: calculating a velocity change of each pixel of the region of interest in the plurality of phase-contrast magnetic resonance images, the velocity change is defined as follows:Velocity⁢⁢change⁢⁢(%)=velocity-average⁢⁢velocity⁢⁢⁢during⁢⁢preFUSaverage⁢⁢velocity⁢⁢during⁢⁢preFUS×100⁢%,wherein preFUS is a time duration before the ultrasonic energy delivered, and a concentration of the microbubbles is equal to (the velocity change a first y-intercept value)/a first slope value, wherein the first y-intercept value is 1 to 1.5 and B the first slope value is 7.56 to 11.34. 8. The method as claimed in claim 7, wherein the first y-intercept value is 1.25 and the first slope value is 9.45. 9. The method as claimed in claim 6, further comprising: calculating a temporal velocity standard deviation of each pixel of the region of interest in the plurality of phase-contrast magnetic resonance images,wherein a concentration of the microbubbles is equal to the temporal velocity standard deviation/a second slope value, wherein the second slope value is 0.04 to 0.06. 10. The method as claimed in claim 9, wherein the second slope value is 0.05. 11. The method as claimed in claim 6, further comprising: calculating a temporal velocity standard deviation of each pixel of the region of interest in the plurality of phase-contrast magnetic resonance images; andcalculating a range of temporal velocity standard deviation which is 90th percentile o... | NATIONAL TSING HUA UNIVERSITY | A61B008/08 A61B005/055 A61B008/06 G01R033/563 G01R033/48 G06T007/00 A61K009/50 A61B008/00 A61K049/18
US10706592 | B2 | 2020-07-07 | Systems and methods for myocardial perfusion MRI without the need for ECG gating and additional systems and methods for improved cardiac imaging | 2015-01-06 | In some embodiments, the present application discloses systems and methods for cardiac MRI that allow for continuous un-interrupted acquisition without any ECG/cardiac gating or synchronization that achieves the required image contrast for imaging perfusion defects. The invention also teaches an accelerated image reconstruction technique that is tailored to the data acquisition scheme and minimizes or eliminates dark-rim image artifacts. The invention further enables concurrent imaging of perfusion and myocardial wall motion (cardiac function), which can eliminate the need for separate assessment of cardiac function (hence shortening exam time), and/or provide complementary diagnostic information in CAD patients. | 1. A method for performing first-pass myocardial perfusion magnetic resonance imaging (MRI) on a subject, comprising: using an MRI machine to apply a pulse sequence with continuous radial k-space sampling to a volume of interest (VOI) comprising a region of a heart of the subject, T1 weighting of the volume of interest being provided by a steady state magnetization of the pulse sequence resulting from the continuous radial k-space sampling, the pulse sequence having a flip angle selected so as to maximize a ratio of (i) a contrast-to-noise ratio of hypoperfused tissue to (ii) a contrast-to-noise ratio of non-perfused tissue;introducing a contrast agent into a vascular system of the subject's prior to or during imaging; andusing a model-based, iterative image reconstruction technique to generate one or more perfusion images of one or more slices or VOIs within the heart region, the image reconstruction technique including utilizing a sliding window scheme to generate the one or more perfusion images, the one or more perfusion images being indicative of passage of the contrast agent through at least the region of the heart of the subject. 2. The method of claim 1, wherein no electrocardiogram (ECG) signal acquisition is required during the MRI. 3. The method of claim 1, wherein an apodization scheme is employed in the image reconstruction technique to reduce or eliminate a dark-rim artifact. 4. The method of claim 1, wherein the one or more perfusion images depict the heart or a portion thereof within 30 to 60 heartbeats after the introduction of the contrast agent. 5. The method of claim 4, wherein an apodization scheme is employed in the image reconstruction technique to reduce or eliminate a dark-rim artifact. 6. The method of claim 1, wherein the one or more perfusion images include multiple image frames generated per each heartbeat at a rate of at least 8 image frames per second to depict the heart motion. 7. The method of claim 1, wherein the Mill machine is a 3.0T scanner. 8. The method of claim 1, wherein the subject has an arrhythmia. 9. The method of claim 1, further comprising diagnosing the subject with the presence or absence of a condition associated with a perfusion defect or a wall motion abnormality, based upon the one or more images. 10. A magnetic resonance imaging system, comprising: a magnet operable to provide a magnetic field;a transmitter operable to transmit to a region within the magnetic field;a receiver operable to receive a magnetic resonance signal from the region; anda processor operable to control the transmitter and the receiver; wherein the processor is configured to direct the transmitter and receiver to execute a sequence, comprising:applying a pulse sequence with continuous radial k-space sampling to a volume of interest (VOI) comprising a region of a heart of a subject, T1 weighting of the VOI being provided by a steady state magnetization of the pulse sequence resulting from the continuous radial k-sp... | Cedars-Sinai Medical Center | G06T011/00 A61B005/055 A61B005/00 G01R033/48 G01R033/563 G16H030/40 G16H040/60 G16H050/50 G01R033/56
US10719698 | B2 | 2020-07-21 | System, method and apparatus for occupancy detection | 2017-04-28 | Systems and methods are directed toward occupancy detection in a predefined space including one or more sub-zones. Occupancy of the predefined space can be determined by determining the occupancy of an outer zone comprising the predefined space, and determining the occupancy of individual sub-zones to the outer zone. Occupancy values of the individual sub-zones can be rescaled such that the sum of the individual sub-zones occupancies sum to the already determined outer zone occupancy. The occupancy of the predefined space then may be determined using the rescaled occupancies of the various sub-zones included in the predefined space. Such processes can be repeated for individual sub-zones, which may be treated as a separate outer zone including one or more sub-zones. | 1. An occupancy system comprising: a plurality of counting sensors; anda central processing function operable to hierarchically determine figures for occupancy of a predefined space by: determining a total occupancy of an outer zone comprising the predefined space, wherein the outer zone comprises a plurality of individual sub-zones;determining occupancies of each individual sub-zone in the plurality of individual sub-zones that make up the outer zone;responsive to determining that a sum of the occupancies for each individual sub-zone is not equal to the total occupancy of the outer zone, resealing the determined occupancies of each individual sub-zone such that the determined occupancies sum to the total occupancy of the outer zone,wherein a first counting sensor of the plurality of counting sensors is positioned to count traffic entering a first sub-zone from a location outside of the outer zone, and wherein a second counting sensor of the plurality of counting sensors is configured to count traffic entering a second sub-zone from the first sub-zone. 2. The system of claim 1, wherein the central processing function is further operable to iteratively determine and rescale occupancies to further zones, or sub-zones, of the predefined space which are fully enclosed. 3. The system of claim 2, wherein a sub-zone is treated as another outer zone, and the sub-zone total occupancy is scaled such that it sums to the already determined sub-zone total. 4. The system of claim 1, further comprising at least one counting sensor operable to sense any one of more of: occupants in a zone, occupants entering a zone or occupants exiting a zone. 5. The system of claim 4, wherein the at least one counting sensor is operable to sense two or more zones in which the occupancy is to be determined, which together make up a larger space. 6. The system of claim 4, wherein the central processing function comprises a central server, and the at least one counting sensor have their counts fed over a network into the central server. 7. The system of claim 1, wherein determining the occupancy of an individual sub-zone comprises any one or more of: determining a first estimate of occupancy as sum(in)−sum(out) for each time point in a day;determining a scaling factor for the day as the ratio of total(in) to total(out), or a weighted sum of the two;multiplying the ‘in’ and/or ‘out’ counts by the ratio such that the scaled totals are equal; and/ordetermining a second estimate for the occupancy for each time point using sum(scaled_in)−sum(scaled_out) for each time point. 8. The system of claim 7, wherein, for each time point, if the second estimate for occupancy is negative, then the central processing function is configured to divide the day at the respective time point into two time segments, and the occupancy determination process re-run for each of the two time segments. 9. The system of claim 7, wherein the result will be an occupancy estimate for the day, such that occ... | Fluke Corporation | G05B019/048 G06K009/00 A61B005/103 F24F120/10 G07C009/00
US10722701 | B2 | 2020-07-28 | Nerve stimulation apparatus and biomagnetic field measurement system | 2016-11-07 | A nerve stimulation apparatus is provided. The nerve stimulation apparatus detects nerve activities from a body surface and applies stimulation. The nerve stimulation apparatus includes a stimulation apparatus with multiple electrodes which are arranged on skin and a current supply unit which supplies a current to the electrodes, which stimulation apparatus provides the current to a living body percutaneously to stimulate a target nerve; a measurement apparatus which measures activities of muscles governed by the nerve according to the stimulation from the stimulation apparatus; and an information processing apparatus which determines, based on a measurement result of the nerve activities obtained from the measurement apparatus, which electrode is capable of providing the target nerve activities equal to or greater than a desired value. | 1. A nerve stimulation apparatus for detecting nerve activities from a body surface and applying stimulation, the nerve stimulation apparatus comprising: a stimulation apparatus, including a plurality of electrodes, a three-dimensional structure that has a convex portion on at least one side, and a current supply unit for supplying a current to the electrodes, which are configured to provide the current percutaneously to a living body to stimulate a target nerve;a measurement apparatus configured to measure activities of muscles governed by the target nerve stimulated by the stimulation apparatus, or the nerve activities of the target nerve itself; andan information processing apparatus configured to determine, based on a measurement result of the nerve activities obtained from the measurement apparatus, which electrode is capable of providing the target nerve activities equal to or greater than a desired value,wherein the electrodes include a plurality of stimulation negative electrodes, a center-to-center distance of adjacent ones of the stimulation negative electrodes is less than 3 mm, an area of each of the stimulation negative electrodes is equal to or greater than 12 mm2, and the stimulation negative electrodes are arranged on the convex portion of the three-dimensional structure. 2. The nerve stimulation apparatus according to claim 1, wherein the stimulation apparatus includes a selection circuit, and the current supply unit supplies the current to the electrode selected by the selection circuit. 3. The nerve stimulation apparatus according to claim 1, wherein the three-dimensional structure has a semicircular column shape. 4. The nerve stimulation apparatus according to claim 1, wherein the three-dimensional structure has a shape of a frustum of a cone. 5. The nerve stimulation apparatus according to claim 1, further comprising: a band configured to attach and fix the electrodes to the living body and include a movable unit disposed on a surface of a side of the band which side is in contact with the living body, whereinthe stimulation negative electrodes are arranged on the movable unit. 6. The nerve stimulation apparatus according to claim 5, wherein the electrodes include a stimulation positive electrode, and the movable unit causes the stimulation negative electrodes to slide on the band and causes a positional relationship between the stimulation negative electrodes and the stimulation positive electrode to be changed. 7. A biomagnetic field measurement system comprising: a nerve stimulation apparatus according to claim 1; anda magnetism measurement apparatus configured to measure magnetic fields evoked in the living body by electrical stimulation of the nerve stimulation apparatus. 8. A nerve stimulation apparatus for detecting nerve activities from a body surface and applying stimulation, the nerve stimulation apparatus comprising: a stimulation apparatus, including a plurality of electrodes and a current supply unit for s... | Ricoh Company, Ltd. National University Corporation Tokyo Medical and Dental University | A61N001/04 A61B005/0488 A61N001/36 A61B005/04 A61B005/00 A61B005/05 A61N001/02
US10702322 | B2 | 2020-07-07 | Medical insertion apparatus | 2014-08-25 | Provided are a medical insertion apparatus includes a screw body to be inserted into a body, and an electrode provided in the screw body and including an externally exposed portion, wherein a through-hole may be provided in the screw body or the electrode, and a medical insertion apparatus includes a screw body to be inserted into a body, and an electrode provided in the screw body and exposed to an outside of the screw body, wherein the screw body and the electrode may be provided in a connectable or separable structure. | 1. A medical insertion apparatus comprising: a screw body to be inserted into a body, wherein the screw body is a first material; andan electrode positioned in the screw body, wherein the electrode is formed to include an externally exposed portion in which the electrode is externally exposed through the screw body, and wherein the electrode is a second material having a greater electrical conductivity than the first material,wherein a through-hole is formed in the screw body and the electrode in a first direction, the first direction corresponding to a longitudinal direction of the screw body, the through-hole extending in the first direction from one end of the screw body to a terminal portion of the screw body,wherein the externally exposed portion of the electrode is provided in a second direction different from the first direction at a position spaced apart from the terminal portion of the screw body on an outer circumference of the screw body,wherein a first portion of the through-hole is formed to extend in a central portion of the screw body and the electrode and a second portion of the through-hole is formed to extend through the central portion of the screw body toward the terminal portion of the screw body without extending through the electrode,wherein the externally exposed portion of the electrode is oriented in the second direction and an end of the through-hole is oriented in the first direction,wherein the screw body that is the first material and the electrode that is the second material are in direct contact with one anotherwherein the first material is titanium, andwherein the second material is selected from the group consisting of platinum, gold, silver and tungsten. 2. The medical insertion apparatus of claim 1, wherein the electrode is formed using a method of filling with a melted electrode material in the screw body and hardening the electrode material by cooling the electrode material at room temperature. 3. The medical insertion apparatus of claim 1, wherein a guide element is disposed in the through-hole to guide an insertion of the screw body into the body. 4. A medical insertion apparatus comprising: a screw body to be inserted into a body, wherein the screw body is titanium; andan electrode positioned in the screw body, wherein the electrode is formed to include an externally exposed portion in which the electrode is externally exposed through the screw body, and wherein the electrode is a material having a greater electrical conductivity than the screw body,wherein a through-hole is formed in the screw body and the electrode in a first direction, the first direction corresponding to a longitudinal direction of the screw body, the through-hole extending in the first direction from one end of the screw body to a terminal portion of the screw body,wherein the externally exposed portion of the electrode is provided in a second direction different from the first direction at a position spaced apart from the terminal... | Kyungpook National University Industry-Academic Cooperation Foundation | A61B017/86 A61N001/05 A61B005/00 A61B017/70 A61B005/0492 A61C008/00 A61N001/36 A61C019/04 A61B017/00
US10702648 | B2 | 2020-07-07 | Fully-automatic regional citrate anticoagulation machine | 2015-09-15 | A fully-automatic regional citrate anticoagulation machine is disclosed. A sampling portion extracts a blood sample on line and in real time and sends it to a storage purifier for storage and purification. A sample injection portion conveys the blood sample in the blood sample storage purifier to a sample loading slot through a sample injection tube. An assay and analysis portion conveys, through a sample suction tube, the blood sample to an analyzer for blood sample assay and analysis. The sampling portion, the sample injection portion, the assay and analysis portion and an administration portion are connected to the program control portion. The program control portion controls, according to an internally-set program, operation of the sampling portion, the sample injection portion and the assay and analysis portion, and sends an instruction to the administration portion according to the data result obtained by the assay and analysis portion. | 1. An automatic regional citrate anticoagulation machine comprising: a sampling portion having a connecting catheter, an air detector, a pressure sensor, a first peristaltic pump, and a blood sample storage purifier successively connected in series, said sampling portion adapted to extract a blood sample and to send the blood sample to the blood sample storage purifier for storage and purification;a sample injection portion having a second peristaltic pump, a sample injection tube and a sample loading slot, one end of the second peristaltic pump communicating with the first peristaltic pump, another end of the second peristaltic pump communicating with the sample injection tube, the sample injection tube being insertable into the sample loading slot, said sample injection portion adapted to convey the blood sample in the blood sample storage purifier to the sample loading slot through a sample injection portion channel;an assay and analysis portion having a sample suction tube and an analyzer, one end of the sample suction tube being inserted in the sample loading slot, and other end of the sample suction tube being connected to the analyzer, the sample suction tube adapted to convey the blood sample in the sample loading slot to the analyzer for blood sample analysis;a program control portion connected to said sample injection portion and said assay and analysis portion, the program control portion controlling said sampling portion and said sample injection portion and said assay and analysis portion so as to compare data results by the analyzer; andan administration portion connected to said program control portion so as to receive instructions relative to the compared data results. 2. The automatic regional citrate anticoagulation machine of claim 1, wherein said administration portion has a first connecting tube, one end of the first connecting tube being connected to an arterial-end line inlet, another end of the first connecting tube being connected to a program-controlled infusion pump, the program-controlled infusion pump adapted to pump a citrate solution into the arterial-end line inlet via the first connecting tube, said administration portion having a second connecting tube, one end of the second connecting tube being connected to a venous-end line inlet, another end of the second connecting tube being connected to the program-controlled infusion pump. 3. The automatic regional citrate anticoagulation machine of claim 1, wherein the connecting catheter having an arterial line connecting catheter and a venous line connecting catheter, the arterial line connecting catheter and the venous line connecting catheter being connected to the passage of said sampling portion through a two-position three-way valve and communicate with the air detector, a front end of the air detector having a first cut-off clamp thereon, the first pressure sensor and the second peristaltic pump having a second cut-off clamp therebetween, the blood sample stora... | - | A61M001/36 A61M001/14 A61M001/16 G01N001/14 G01N001/34 G01N033/49 A61B005/00
US10716505 | B2 | 2020-07-21 | Systems and methods for estimating clinical effects of electrical stimulation | 2018-07-10 | A method for generating a clinical effects map for electrical stimulation using an electrical stimulation lead includes receiving at least one clinical response for stimulation using each of multiple first sets of stimulation parameters where each clinical response is associated with at least one therapeutic effect or side effect. The method also includes determining, using the clinical responses and the first sets of stimulation parameters, a spatial relationship between the electrical stimulation lead and at least one effect region, where each effect region is associated with at least one of the therapeutic effects or side effects; estimating, based on the spatial relationship, at least one clinical response for each of multiple second sets of stimulation parameters; and generating and presenting a clinical effects map illustrating at least the estimated clinical response for one or more of the second sets of stimulation parameters. | 1. A system for generating a clinical effects map for electrical stimulation, the system comprising: a display; anda processor coupled to the display and configured to perform actions comprising: receiving, for each of a plurality of first sets of stimulation parameters, at least one clinical response resulting from stimulation of tissue of a patient using the first set of stimulation parameters, wherein each first set of stimulation parameters comprises a selection of at least one electrode of an electrical stimulation lead to generate the stimulation, wherein each of the at least one clinical response is associated with at least one therapeutic effect or side effect;determining, using the clinical responses and the first sets of stimulation parameters, a spatial relationship between the electrical stimulation lead and at least one effect region in the tissue, wherein each effect region is associated with at least one of the at least one therapeutic effect or side effect;estimating, based on the spatial relationship, at least one clinical response for each of a plurality of second sets of stimulation parameters which are different from the first sets of stimulation parameters; andgenerating and presenting, on the display, a clinical effects map illustrating at least the estimated at least one clinical response for each of a plurality of the second sets of stimulation parameters. 2. The system of claim 1, wherein the determining the spatial relationship comprises estimating, for at least one of the first sets of stimulation parameters, a region of tissue stimulated by the electrical stimulation lead using the at least one of first sets of stimulation parameters. 3. The system of claim 2, wherein the estimating the region comprising determining a stimulation field model (SFM) for the at least one of the first sets of stimulation parameters. 4. The system of claim 2, wherein the determining the spatial relationship further comprises estimating a degree of overlap between the estimated region of tissue stimulated by the at least one of the first sets of stimulation parameters and a one of the at least one effect region based on the the at least clinical response resulting from stimulation using the one of the first sets of stimulation parameters. 5. The system of claim 1, wherein the estimating at least one clinical response comprises estimating, for at least one of the second sets of stimulation parameters, a region of tissue stimulated by the electrical stimulation lead using the at least one of the second sets of stimulation parameters. 6. The system of claim 5, wherein the estimating the region comprising determining a stimulation field model (SEM) for the at least one of the second sets of stimulation parameters. 7. The system of claim 5, wherein the estimating at least one clinical response further comprises determining a degree of overlap between the estimated region of tissue stimulated by one of the second sets of stimulation paramete... | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | A61B005/00 A61N001/36 G06T007/30 G06T007/70 A61N001/05
US10722155 | B2 | 2020-07-28 | Method and device for generating image showing concentration distribution of biological substances in biological tissue | 2014-05-28 | Generating an image of a molar concentration ratio between a first and second biological substance, an absorption spectrum of the first and second biological substance within a predetermined wavelength range including a first, second, third, and fourth isosbestic point, acquiring first imaging data of the tissue using light extracted from white light using a first filter to collectively selectively extract light in a first wavelength range demarcated by the first and second isosbestic point, light in a second wavelength range demarcated by the second and third isosbestic point, and light in a third wavelength range demarcated by the third and fourth isosbestic point; acquiring second imaging data by taking an image of the tissue using light from the white light using a second optical filter to selectively extract light in the second wavelength range; and generating the distribution image on the basis of the first and second imaging data. | 1. A device that generates a distribution image showing a molar concentration ratio between a first biological substance and a second biological substance included in a biological tissue of which an absorption spectrum within a predetermined wavelength range has a first isosbestic point, a second isosbestic point, a third isosbestic point and a fourth isosbestic point, wherein the wavelengths of the first, second, third and fourth isosbestic points are about 528 nm, about 546 nm, about 570 nm, and about 584 nm respectively, the device comprising: a light source which emits white light;a first optical filter configured to selectively transmit, from the white light, light within a wavelength range from the first isosbestic point to the fourth isosbestic point, and not to transmit light outside the wavelength range, the wavelength range including a first wavelength range from the first isosbestic point to the second isosbestic point, a second wavelength range from the second isosbestic point to the third isosbestic point, and a third wavelength range from the third isosbestic point to the fourth isosbestic point;a second optical filter configured to selectively transmit light within the second wavelength range from the white light, and not to transmit light outside the second wavelength range;a controller configured to sequentially insert the first optical filter and the second optical filter one after another, into a light path of the white light emitted by the light source;an image pick-up device configured to take a first image of the biological tissue using the light transmitted through the first optical filter and a second image of the biological tissue using the light transmitted through the second optical filter; andan image processor configured to generate the distribution image on the basis of the first image and the second image;wherein the first and second biological substances are oxyhemoglobin and deoxyhemoglobin,wherein the molar concentration ratio of the first biological substance and the second biological substance included in the biological tissue is oxygen saturation,wherein the first optical filter has a transmission wavelength range of about 528 nm-584 nm and the second optical filter has a transmission wavelength range of about 546 nm-570 nm, andwherein the image processor generates the distribution image on the basis of the first image and the second image by acquiring a first absorbance of the biological tissue in a transmission wavelength range of the first optical filter on the basis of the first image, acquiring a second absorbance of the biological tissue in a transmission wavelength range of the second optical filter on the basis of the second image, and calculating an index X, as X=A1−2kA2 where k is a constant, A1 is the first absorbance, A2 is the second absorbance; and generating the distribution image on the basis of the index X. 2. The device according to claim 1 comprising an endoscope device including an endosco... | HOYA CORPORATION | A61B005/00 A61B005/1455 A61B005/1459 A61B001/06 G06T007/00 A61B001/00 A61B001/04 A61B005/145
US10716500 | B2 | 2020-07-21 | Systems and methods for normalization of chemical sensor data based on fluid state changes | 2016-06-22 | Embodiments herein include implantable medical systems, devices and methods including chemical sensors. In an embodiment, an implantable medical device system includes a chemical sensor; a fluid state sensor such as a posture sensor; an activity sensor; and/or a respiration sensor. The implantable medical device system can further include normalization circuitry receiving data from the chemical sensor and the fluid state sensor and normalizing the chemical sensor data based on data from the fluid state sensor. In an embodiment, a method of operating an implantable medical device system is included. The method can include measuring the amount of a chemical analyte using a chemical sensor, measuring the fluid status in a patient using a fluid state sensor, and normalizing the measured amount of the chemical analyte as indicated by the chemical sensor using normalization circuitry based on data from the fluid state sensor. Other embodiments are also included herein. | 1. An implantable medical device system comprising: a chemical sensor configured to generate chemical sensor data;a fluid state sensor, the fluid state sensor comprising a posture sensor configured to generate fluid state sensor data to differentiate between patient postures, the fluid state sensor data corresponding to patient postures comprising at least an upright posture and a recumbent posture; andnormalization circuitry configured to receive the chemical sensor data and the fluid state sensor data and normalize the chemical sensor data based on the patient postures comprising at least an upright posture and a recumbent posture corresponding to the fluid state sensor data; andwherein the normalization circuitry is further configured to categorize the chemical sensor data based on the patient postures comprising at least an upright posture and a recumbent posture corresponding to the fluid state sensor data. 2. The implantable medical device system of claim 1, the normalization circuitry discarding chemical sensor data, ignoring chemical sensor data, or suspending chemical sensor data collection when the fluid state sensor data indicates a fluid state exceeding a threshold value. 3. The implantable medical device system of claim 1, the normalization circuitry labeling data coming from the chemical sensor based on the patient postures corresponding to the fluid state sensor data. 4. The implantable medical device system of claim 1, wherein the categorization of chemical sensor data based on the patient postures comprising at least an upright posture and a recumbent posture corresponding to the fluid state sensor data further comprises categorizing the chemical sensor data into one or more categories comprising a fully recumbent posture or a partially recumbent posture. 5. The implantable medical device system of claim 4, the normalization circuitry adjusting alert and/or therapy thresholds according to the category. 6. The implantable medical device system of claim 4, the normalization circuitry adjusting the chemical sensor data according to the category. 7. The implantable medical device system of claim 1, wherein the normalization circuitry matches fluid state sensor data with chemical sensor data that is offset in time from one another. 8. The implantable medical device system of claim 1, wherein the normalization circuitry uses fluid state sensor data from an earlier time to process chemical sensor data from a later time. 9. The implantable medical device system of claim 1, the normalization circuitry matching a measured fluid state as determined by the fluid state sensor data to a template for that fluid state. 10. The implantable medical device system of claim 9, the normalization circuitry further adjusting the chemical sensor values according to the template. 11. The implantable medical device system of claim 1, the normalization circuitry calculating a weighted average for the chemical sensor data with the weight based on t... | Cardiac Pacemakers, Inc. | A61B005/1473 A61B005/145 A61B005/00 A61N001/372 A61B005/01 A61B005/05 A61B005/08 A61B005/1459 A61B005/11 A61N001/365 A61N001/36
US10709331 | B2 | 2020-07-14 | Automatic transfer and identification of monitored data with hierarchical key management infrastructure | 2007-07-27 | A measurement device (14) includes a measuring unit (42) for obtaining health related parameters of a patient (12), and a body-coupled communication unit (40) for sending at least measurement results. An identification device (20), associated with the patient, includes a body-coupled communication unit (26) for receiving and sending out the measurement results. A gateway device (72) includes a body-coupled communication unit (78) for receiving patient's measurement results. Additionally, a hierarchical relational deployment model (100) facilitates grouping wireless devices in a healthcare environment into subgroups based on relationships between the devices, and a hierarchical key pre-distribution scheme (110) permits distribution of unique keying material for security domains of respective groups of devices, prior to deploying the devices in a healthcare network. Devices can then execute a handshake protocol by which they identify their deepest common node in the deployment model (100), and employ keying material assigned to that nodes group in the pre-distribution scheme (110) to facilitate granular communication security. | 1. A method for measuring health related parameters of least two patients, comprising: sending, by a patient identification device, an identification for each of the at least two patients to a measurement device;receiving and reading a valid identification for at least one of the at least two patients and measuring, with the measurement device, a health related parameter of the at least one of the at least two patients having a valid identification;communicating the health related parameter measurement results of the patient having a valid identification from the measurement device to the patient identification device carried by the patient via body-coupled communication;communicating the measurement results separately and concurrently with a patient identification code from the identification device to a gateway device via body-coupled communication;associating, via the gateway device, the health related parameter measurements with the separately and concurrently received patient identification code as the health related parameter measurements and patient identification code are received at the gateway device, the association between the measurement results and the patient identification code by the gateway device being a first electronic association between the measurement results and the patient identification code;receiving and reading an invalid identification for the other patient of the at least two patients and measuring, with the measurement device, a health related parameter of the other patient of the at two patients having an invalid identification; anddisplaying the measurement results for the other patient having an invalid identification locally on the measurement device. 2. The method as set forth in claim 1, wherein the patient identification device includes a measurement memory and further including: storing the measurement results in the measurement memory. 3. The method as set forth in claim 1, wherein the measuring unit includes a sensor and further including: measuring a physiological function of the patient. 4. The method as set forth in claim 1, further including: assigning measurement devices to one or more patients, each measurement device being encoded with an electronic device identification code that uniquely identifies the selected patient which it is assigned to;communicating the electronic identification code of the measurement device to a measurement memory contemporaneously with communicating the measurement results; andassociating the communicated measurement results with the patient via the electronic identification code of the measurement device. 5. The method as set forth in claim 1, further including: assigning a patient identification device to the patient which identification device is encoded with an electronic patient identification code that uniquely identifies the patient which it is assigned to;communicating the patient identification code to the gateway device contemporaneously with the communica... | KONINKLIJKE PHILIPS N.V. | A61B005/00 H04L029/06 H04L009/08 H04W012/04 H04B013/00 H04L029/08 H04W084/18 H04W012/00
US10702172 | B2 | 2020-07-07 | Blood pressure cuff with tapered bladder | 2015-09-16 | A blood pressure cuff may comprise an alignment component, a sleeve, and a tapered inflatable bladder disposed within the sleeve. The tapered inflatable bladder may have a width that increases as the distance from the alignment component increases, and the tapered inflatable bladder and the alignment component may be configured to position the blood pressure cuff around a limb having a blood vessel and a circumference, such that at a position coincident with the blood vessel, the width of the tapered inflatable bladder is about 40% of the circumference of the limb. A method of using such a blood pressure cuff may comprise placing the alignment component at the position coincident with the blood vessel of the limb, wrapping the cuff around the limb such that the width of the tapered inflatable bladder overlaying the blood vessel is about 40% of the circumference of the limb, and inflating the bladder. | 1. A blood pressure cuff comprising: an alignment component having a front straight edge;a sleeve; anda tapered inflatable bladder disposed within the sleeve, wherein the tapered inflatable bladder has a width that increases as the distance from the front straight edge increases, and wherein the tapered inflatable bladder and the alignment component are configured to position the blood pressure cuff around a limb having a blood vessel and a circumference such that at a position coincident with the blood vessel, the width of the tapered inflatable bladder is about 40% of the circumference of the limb for any circumference within a range of circumferences;wherein the alignment component has a length between the front straight edge and the tapered inflatable bladder;wherein the tapered inflatable bladder does not extend into the alignment component; andwherein the front straight edge of the alignment component is configured to be placed at the position coincident with the blood vessel. 2. The blood pressure cuff of claim 1, further comprising a rectangular section having a first end and a second end opposite the first end. 3. The blood pressure cuff of claim 2, wherein a length from an end of the alignment component opposite the edge to the second end of the rectangular section is from about 19 cm to about 35 cm. 4. The blood pressure cuff of claim 2, wherein the circumference of the limb is from about 27 cm to about 34 cm, and a length from an end of the alignment component opposite the front straight edge to the second end of the rectangular section is about 27 cm. 5. The blood pressure cuff of claim 1, wherein the tapered inflatable bladder is logarithmically tapered. 6. The blood pressure cuff of claim 1, wherein the tapered inflatable bladder is curvedly tapered. 7. The blood pressure cuff of claim 1, wherein the tapered inflatable bladder is linearly tapered. 8. The blood pressure cuff of claim 1, further comprising at least one tube configured to transport air to and from the tapered inflatable bladder. 9. The blood pressure cuff of claim 1, further comprising at least one tube configured to transmit sound from the tapered inflatable bladder. 10. The blood pressure cuff of claim 1, further comprising at least one sensor configured to detect sound. 11. The blood pressure cuff of claim 1, further comprising at least one sensor configured to detect pressure. 12. The blood pressure cuff of claim 1, further comprising a ring disposed at an end of the sleeve. 13. The blood pressure cuff of claim 1, wherein the range of circumferences is a continuous range of circumferences. 14. A method of using a blood pressure cuff, the method comprising: providing a blood pressure cuff comprising: an alignment component having a front straight edge;a sleeve; anda tapered inflatable bladder disposed within the sleeve, wherein the tapered inflatable bladder has a width that increases as the distance from the front straight edge increases, and wherein ... | ST. LUKE MEDICAL, INC. | A61B005/022 A61B005/025 A61B017/135 A61B007/04
US10709361 | B2 | 2020-07-14 | Methods and systems for correcting blood analyte measurements | 2018-03-29 | A system, transceiver, and method for calculating and correcting levels (e.g., analyte levels) in a first medium (e.g., blood) using measurements from a second medium (e.g., interstitial fluid). In some embodiments, a transceiver may calculate an initial second medium level. The transceiver may calculate an initial second medium level rate of change (“ROC”) using at least the initial second medium level and past second medium level(s). The transceiver may calculate a first medium level using at least the initial second medium level and the initial second medium level ROC. The transceiver may calculate a subsequent second medium level. The transceiver may calculate an updated second medium level ROC using at least the initial second medium level, the subsequent second medium level, and past second medium level(s). The transceiver may calculate a corrected first medium level using at least the initial second medium level and the updated second medium level ROC. | 1. An analyte monitoring system comprising: an analyte sensor including an indicator element that exhibits one or more detectable properties based on an amount or concentration of an analyte in proximity to the indicator element; anda transceiver configured to: receive first sensor data from the analyte sensor;calculate a first interstitial fluid analyte level using at least the first sensor data;calculate a first interstitial fluid analyte level rate of change using at least the first interstitial fluid analyte level and one or more past interstitial fluid analyte levels;calculate a first blood analyte level using at least the first interstitial fluid analyte level and the first interstitial fluid analyte level rate of change;after receiving the first sensor data, receive second sensor data from the analyte sensor;calculate a second interstitial fluid analyte level using at least the second sensor data;calculate an updated first interstitial fluid analyte level rate of change using at least the first interstitial fluid analyte level, the second interstitial fluid analyte level, and the one or more past interstitial fluid analyte levels; andcalculate a corrected first blood analyte level using at least the first interstitial fluid analyte level and the first interstitial fluid analyte level rate of change. 2. The system of claim 1, wherein the transceiver is further configured to: calculate a second interstitial fluid analyte level rate of change using at least the first and second interstitial fluid analyte levels; andcalculate a second blood analyte level using at least the second interstitial fluid analyte level and the second interstitial fluid analyte level rate of change. 3. The system of claim 2, further comprising a display device configured to: receive and display the first blood analyte level;receive and display the second blood analyte level; andreceive and display the corrected first blood analyte level. 4. The system of claim 1, further comprising a display device configured to: receive and display the first blood analyte level; andreceive and display the corrected first blood analyte level. 5. The system of claim 4, wherein the display device is configured to: display the first blood analyte level until the corrected first blood analyte level is received; andafter receiving the corrected first blood analyte level, display the corrected first blood analyte level instead of the first blood analyte level. 6. The system of claim 1, wherein the transceiver is further configured to: after receiving the second sensor data, receive third sensor data from the analyte sensor;calculate a third interstitial fluid analyte level using at least the third sensor data; andcalculate the updated first interstitial fluid analyte level rate of change using at least the first interstitial fluid analyte level, the second interstitial fluid analyte level, the third interstitial fluid analyte level, and the one or more past interstitial fluid analyte l... | Senseonics, Incorporated | A61B005/1455 A61B005/145 A61B005/1459 A61B005/00 A61B005/1486 A61B005/07 A61B005/1473
US10722143 | B2 | 2020-07-28 | Exhaled gas measurement compensation during high flow respiratory therapy | 2015-09-04 | The present disclosure relates to determining a corrected exhaled gas measurement during high flow respiratory therapy. Measuring exhaled gas concentration during high flow respiratory therapy is difficult and inaccurate due to a phenomenon known as flushing. The high flows delivered to the patient flush the dead space in the conducting airways, which causes a dilution effect that results in underestimated or overestimated exhaled gas measurement depending on the gas composition delivered by the high flow system. This can lead to incorrect clinical measurements and diagnoses. Various algorithms are disclosed herein to account for the dilution effect caused by flushing, allowing for the method of measuring gas concentrations to still be used accurately for clinical measurements. | 1. A system for estimating exhaled gas during high flow respiratory therapy, the system comprising: a high flow respiratory therapy system;an unsealed patient interface;an exhaled gas monitoring system; anda hardware processor which is configured to receive exhaled gas measurement data from the exhaled gas monitoring system and flow data from the respiratory flow therapy system and determine a compensated exhaled gas measurement by estimating a flow rate dilution. 2. The system of claim 1, wherein the exhaled gas is CO2. 3. The system of claim 1, wherein the exhaled gas monitoring system is a capnograph. 4. The system of claim 1, wherein the exhaled gas monitoring system is a capnometer. 5. The system of claim 1, wherein the unsealed patient interface is a nasal cannula and wherein the system comprises a display arranged to display the compensated exhaled gas measurement. 6. The system of claim 1, wherein the hardware processor determines the compensated exhaled gas by estimating expiratory volume dilution. 7. The system of claim 1, wherein the high flow respiratory therapy system is configured to deliver gas at a preset flow rate. 8. The system of claim 7, wherein the preset flow rate comprises a constant flow rate. 9. The system of claim 7, wherein the preset flow rate comprises an oscillatory flow rate. 10. The system of claim 1, wherein the high flow respiratory therapy system is configured to deliver humidified gas to the unsealed patient interface. 11. The system of claim 1, wherein the hardware processor is configured to determine the compensated exhaled gas measurement at least in part based on determining a concentration or volume of a gas component in a flow of gases delivered via the unsealed patient interface. 12. The system of claim 1, wherein the high flow respiratory therapy system is configured to deliver gas at a flow rate of greater than 10 L/min. 13. A method for estimating exhaled gas during flow respiratory therapy, the method comprising: receiving an exhaled gas measurement;receiving flow data; anddetermining a compensated exhaled gas measurement using a hardware processor based on the exhaled gas measurement and the flow data, wherein the determining comprises determining an estimate of flow rate dilution;wherein the flow respiratory therapy is a non-sealed high flow respiratory therapy that involves an unsealed patient interface. 14. The method of claim 13, wherein receiving the exhaled gas measurement comprises receiving an exhaled gas measurement from a capnograph. 15. The method of claim 13, wherein receiving the exhaled gas measurement comprises receiving an exhaled gas measurement from a capnometer. 16. The method of claim 13, wherein the high flow respiratory therapy is nasal high flow respiratory therapy. 17. The method of claim 13, wherein determining the compensated exhaled gas measurement comprises determining an estimate of expiratory volume dilution. 18. The method of claim 13, wherein the non-... | Fisher & Paykel Healthcare Limited | A61B005/083 A61M016/00 A61M016/08 A61B005/091 A61M016/06 A61M016/10
US10702193 | B2 | 2020-07-07 | Analyte sensing biointerface | 2020-01-14 | Disclosed herein is an analyte sensing biointerface that comprises a sensing electrode incorporated within a non-conductive matrix comprising a plurality of passageways extending through the matrix to the sensing electrode. Also disclosed herein are methods of manufacturing a sensing biointerface and methods of detecting an analyte within tissue of a host using an analyte sensing biointerface. | 1. A transcutaneous continuous glucose sensor system comprising: a substantially planar sensor, the sensor comprising: a first conductive layer associated with a first electrode;a first non-conductive layer located at least in part over the first conductive layer;a second conductive layer associated with a second electrode, wherein the second conductive layer is located at least in part over the first non-conductive layer;a second non-conductive layer located at least in part over the second conductive layer;a third conductive layer associated with a third electrode, wherein the third conductive layer is located at least in part over the second non-conductive layer; anda membrane located over at least a portion of a working electrode;wherein at least one of the first electrode, the second electrode, or the third electrode is the working electrode, and wherein the working electrode is configured to measure a signal indicative of a glucose concentration. 2. The transcutaneous continuous glucose sensor system of claim 1, wherein the membrane comprises an enzyme. 3. The transcutaneous continuous glucose sensor system of claim 2, wherein the membrane further comprises a resistance domain located over the enzyme. 4. The transcutaneous continuous glucose sensor system of claim 1, wherein the membrane comprises a bioactive agent. 5. The transcutaneous continuous glucose sensor system of claim 1, wherein the membrane is substantially planar. 6. The transcutaneous continuous glucose sensor system of claim 1, wherein the membrane covers at least a portion of the first conductive layer and at least a portion of the third conductive layer. 7. The transcutaneous continuous glucose sensor system of claim 1, wherein at least a portion of the first non-conductive layer is in physical contact with at least a portion of the first conductive layer. 8. The transcutaneous continuous glucose sensor system of claim 1, further comprising sensor electronics configured to electrically couple to the working electrode. 9. The transcutaneous continuous glucose sensor system of claim 1, wherein at least a portion of the first non-conductive layer is physically separate from the first conductive layer. 10. The transcutaneous continuous glucose sensor system of claim 1, further comprising sensor electronics having one or more contacts for operably connecting to the working electrode. 11. A transcutaneous continuous glucose sensor system, comprising: a substantially planar sensor, the sensor comprising: a first conductive layer associated with a first electrode, wherein the first electrode is a counter electrode;a first non-conductive layer located over at least a portion of the first conductive layer;a second conductive layer associated with a second electrode, wherein the second conductive layer is located over at least a portion of the first non-conductive layer, wherein the second electrode is a working electrode;a second non-conductive layer located over at least a... | DexCom, Inc. | A61B005/05 A61B005/145 A61B005/1486 C12Q001/00 B33Y080/00 A61B005/00
US10702187 | B2 | 2020-07-07 | Apparatus and method for controlling radiation source variability for optical gas measurement systems | 2012-04-04 | An infrared detector includes a radiation source configured to emit electromagnetic radiation. The detector includes a source monitoring apparatus configured to generate output signals (62) related to electrical resistance through the radiation source. The output signals are used to account for radiation source temperature in measurements by the detector. The output signals may be used (72) in the provision of power to the radiation source to maintain the temperature of the detector at a beneficial level. The detector is configured to monitor a level of a gaseous molecular species within a flow of breathable gas. | 1. A method of monitoring a level of a gaseous molecular species within a flow of breathable gas with a detector, the method comprising: estimating irradiance of a radiation beam before the beam passes through a flow path of a flow of breathable gas that communicates with an airway of a subject, the radiation beam including a first wavelength absorbed by a selected species in the breathable gas in the flow path and a second reference wavelength;after the radiation beam has passed through the flow path, separating the first and second wavelengths into a first wavelength beam and a second wavelength beam, respectively;with a first radiation detector, receiving the first wavelength beam and outputting an electronic signal indicative of the irradiance of the first beam;comparing irradiance of the first wavelength beam with the estimated irradiance of the beam before passing through the flow path to determine a difference indicative of a level of the gaseous molecular species in the breathable gas in the flow path;with a second radiation detector, receiving the second wavelength beam and outputting an electronic signal indicative of the irradiance of the second beam;comparing the estimated irradiance of the beam before passing through the flow path with an irradiance of the second wavelength beam to determine a difference; andresponsive to the difference between the estimated irradiance and the second wavelength beam irradiance exceeding a threshold, generating a notification to initiate a recalibration of the detector. 2. The method of claim 1, wherein estimating the irradiance of the radiation beam before passing through the flow path includes determining a resistance of a radiation source which generates the radiation beam and further including: controlling a provision of power to the radiation source dynamically based on the determined electrical resistance. 3. The method of claim 1, wherein estimating the irradiance of the radiation beam before it passes through the flow path includes determining an electrical resistance of a radiation source that generates the radiation beam. 4. The method of claim 1, wherein separating the first and second wavelength beams includes: passing the radiation beam, after passing through the flow path, through a beam splitter which separates the radiation beam into the first wavelength beam and the second wavelength beam. 5. The method of claim 4, further including: generating an electronic output signal indicative of the irradiance of the first wavelength beam and generating a reference output signal indicative of the irradiance of the second wavelength beam; andcomparing the first and second output signals using a processor to determine a difference indicative of a level of a molecular species in the flow path. 6. The method of claim 3, further including: converting the irradiance of the second wavelength beam into an electrical output signal indicative of the irradiance of the second wavelength beam;with a pr... | Koninklijke Philips N.V. | A61B005/00 A61B005/08 G01N021/3504 G01J003/10 G01N021/27 G01N021/85
US10716560 | B2 | 2020-07-21 | Implant unit delivery tool | 2018-07-25 | An implant unit delivery tool is provided. The implant delivery tool may include a body, a holder disposed at a distal end of the body and adapted to hold an implant unit, and an implant activator associated with the body, the implant activator configured to receive power from a power source. The implant activator may be configured to selectively and wirelessly transfer power from the power source to the implant unit during implantation of the implant unit into the body of a subject to cause modulation of at least one nerve in the body of the subject, and determine a degree of nerve modulation response resulting from the selective and wireless transfer of power from the power source to the implant unit claims. | 1. A method of testing the functionality of an implant unit, the method comprising: moving an implant activator from a first position to a second position, prior to final fixation of the implant unit, the implant unit being configured for implantation within a body of a subject; andtransferring power from the implant activator to the implant unit when the implant activator is in the second position,wherein transferring power to the implant unit provides for one or more patient signals of the subject. 2. The method of claim 1, wherein power is wirelessly transferred from the implant activator to the implant unit. 3. The method of claim 1, wherein the implant activator includes a power source configured to supply the power to the implant unit. 4. The method of claim 1, wherein the implant activator includes a power source configured to supply the power to the implant unit, and wherein the power source is a battery. 5. The method of claim 1, wherein the second position is closer to the implant unit than the first position. 6. The method of claim 1, further comprising selectively activating the implant activator and causing the implant activator to move toward and away from the implant unit via a slidable engagement portion. 7. The method of claim 1, further comprising: determining that an implant location is suitable based on the one or more patient signals of the subject; andimplanting the implant unit at the implant location. 8. The method of claim 1, further comprising: transmitting power between an antenna on the activator and a secondary antenna of the implant unit,receiving, at the implant activator, feedback from an implant unit once the implant unit is activated by implant activator, in order to test the implant unit; anddetermining, by a processor that, implant unit is not in a sufficient location from a desired nerve to properly cause modulation of the nerve. | NYXOAH SA | A61B017/04 A61B005/04 A61B005/0488 H04B005/00 A61B017/30 A61N001/378 A61N001/05 A61N001/36 A61N001/372
US10722130 | B2 | 2020-07-28 | Body-worn system for continuous, noninvasive measurement of cardiac output, stroke volume, cardiac power, and blood pressure | 2011-12-27 | The invention provides a system for measuring stroke volume (SV), cardiac output (CO), and cardiac power (CP) from a patient that features: 1) an impedance sensor connected to at least two body-worn electrodes and including an impedance circuit that processes analog signals from the electrodes to measure an impedance signal (e.g. a TBEV waveform); 2) an ECG sensor connected to at least two chest-worn electrodes and including an ECG circuit that processes analog signals from the electrodes to measure and ECG signal; 3) an optical sensor connected to a body-worn optical probe and including an optical circuit that processes signals from the probe to measure at least one optical signal (e.g. a PPG waveform) from the patient; 4) a processing system, typically worn on patient's wrist and connected through a wired interface to the optical sensor, and through either a wired or wireless interface to the TBEV and ECG sensors. | 1. A system for measuring a stroke volume of a patient, comprising: an impedance sensor comprising at least two electrodes and an impedance circuit, the at least two electrodes configured to be positioned on the patient in the clavicle/brachial region, and the impedance circuit configured to receive and process signals from the at least two electrodes to measure a transbrachial electro-velocimetry (TBEV) signal from the patient;an optical sensor configured to be positioned on a hand of the patient and comprising an optical probe and an optical circuit, the optical circuit configured to receive and process signals from the optical probe to measure at least one optical signal from the patient; anda processing system configured to: (i) be worn on the patient's body; (ii) operably connect to both the impedance sensor and the optical sensor; (iii) receive the TBEV signal from the impedance sensor; iv) receive the at least one optical signal from the optical sensor; (v) process the TBEV signal to determine a time dependent derivative of the TBEV signal (d(TBEV)dZ/dt), a first zero value in d(TBEV)dZ/dt indicating onset of a pulse in d(TBEV)dZ/dt representing opening of the patient's aortic valve, a second zero value in d(TBEV)dZ/dt after onset of the pulse indicating closing of the patient's aortic valve, a d(TBEV)dZ/dtmax value; (vi) determine a first systolic flow time (SFT) value calculated using the first zero value in d(TBEV)dZ/dt and the second zero value in d(TBEV)dZ/dt; (vii) collectively process d(TBEV)dZ/dtmax and the first SFT value to determine a first stroke volume (SV) value; (viii) process the optical signal to determine a second SFT value; (ix) collectively process d(TBEV)dZ/dtmax and the second SFT value to determine a second SV value; (x) collectively process the first and second SV values to determine a third SV value; and (xi) display the third SV value on a display device. 2. The system of claim 1, wherein the impedance sensor comprises a housing configured to be worn on the patient's body. 3. The system of claim 1, wherein the impedance sensor comprises an analog-to-digital converter configured to convert an impedance signal into a digital impedance signal. 4. The system of claim 3, wherein the impedance sensor comprises a transceiver configured to transmit the digital impedance signal through a cable to the processing system. 5. The system of claim 3, wherein the impedance sensor comprises a transceiver configured to transmit the digital impedance signal through a wireless interface to the processing system. 6. The system of claim 5, wherein the transceiver comprises a short-range radio. 7. The system of claim 1, wherein the optical probe is configured to be worn on the patient's thumb. 8. The system of claim 7, wherein the processing system comprises a wrist strap configured to wrap around the patient's wrist so that the processing system can be worn on the wrist. 9. The system of claim 8, wherein the optical circuit and... | SOTERA WIRELESS, INC. | A61B005/0295 A61B005/029 A61B005/053 A61B005/08 A61B005/11 A61B005/1455 A61B005/00 A61B005/01 A61B005/021 A61B005/02 A61B005/0205 A61B005/145 G06F019/00 A61B005/024 A61B005/0245
US10706971 | B2 | 2020-07-07 | System for management and intervention of neurocognitive related conditions and diseases | 2018-08-02 | A system and apparatus for the management and intervention of neurocognitive-related conditions and diseases. The intervention may be implemented in the form of a wearable device providing one or more features of medication adherence, voice, data, SMS reminders, alerts, location via SMS, and 911 emergency. The device may function in combination with an application software accessible to multiple clients (users) executable on a remote server to provide patient support of memory, social contact, daily activities, safety as well as support for caregivers, and feedback for healthcare providers in the management of neurocognitive and neurodegenerative conditions, syndromes, and diseases including dementia and AD. Alternative embodiments implementing monitoring and intervention include using mobile apps or voice-controlled speech interface devices to access cloud control services capable of processing automated voice recognition-response and natural language understanding-processing to perform functions and fulfill user requests. | 1. A system for management of neurocognitive and neurodegenerative conditions, syndromes, and diseases comprising: a speech interface device operably engaged with a communications network, the speech interface device being configured to receive a voice input from a patient user of the speech interface device, process a voice transmission from the voice input, and communicate the voice transmission over the communications network via at least one communications protocol, wherein the speech interface device has at least one audio output means, and the patient user of the speech interface device is a patient having a neurocognitive or neurodegenerative condition, syndrome, or disease; and,a remote server operably engaged with the speech interface device via the communications network to receive the voice transmission, the remote server executing a control service comprising an automated speech recognition function, a natural-language processing function, and one or more application protocols, the control service processing the voice transmission to execute one or more system commands, the one or more system commands comprising communicating an interaction of the patient user to the speech interface device, determining a health status of the patient user, fulfilling a request of the patient user, and communicating with one or more third-party servers, wherein: the control service executing on the remote server serves as a relational agent that provides conversational interactions with the patient user based on the automated speech recognition function, the natural-language processing function, and the one or more application protocols, the conversational interactions being directed to the management of neurocognitive and neurodegenerative conditions, syndromes, and diseases of the patient user; andthe remote server comprises an application software comprising instructions for a user management function and a patient treatment function, the user management function being operable to define roles and permissions between the patient user and one or more non-patient users, and the patient treatment function being operable to perform one or more patient care protocols associated with the control service, the one or more patient care protocols comprising functions to perform non-pharmacologic interventions, wherein, in response to a command by the remote server, the relational agent further provides suggestions for non-pharmacologic interventions as corrective actions in the management of neurocognitive and neurodegenerative conditions, syndromes, and diseases, andwherein in response to a command by the remote server, the speech interface device communicates the suggested non-pharmacologic interventions to the patient user. 2. The system of claim 1 wherein the one or more patient care protocols further comprise instructions for tracking a location of the patient user, monitoring medication adherence of the patient user, and storing and tracking health dat... | Elements of Genius, Inc. | G08B023/00 G16H080/00 A61B005/00 G06F003/0482 G16H020/30 G06F016/9032 A61B005/11 G16H050/20 G16H040/63 G16H040/67 G16H040/20 G16H020/10 G06F001/16 G08B021/04 A61B005/0205 G10L015/18 G10L015/22 A61B005/145 A61B005/021
US10726283 | B2 | 2020-07-28 | Finger vein authentication device | 2016-10-31 | Provided is a device which performs vein authentication by using a downward irradiation-type thin module and selecting an image of a proper angle by varying the irradiation direction of near-infrared illumination. Realized are a photographing method and control method of a finger vein image suitable for thin devices such as a smartphone. Adopted is a finger vein authentication device comprising an imaging unit, an illumination unit which is disposed on a substantially same plane as the imaging unit, and irradiates a finger to be captured by the imaging unit with light in which an irradiation angle is variable, an image selection unit which selects an image according to the irradiation angle of the illumination unit, and an authentication processing unit which performs authentication processing using the image selected by the image selection unit. | 1. A finger vein authentication device which takes a photograph of a blood vessel pattern of a finger of an individual, and authenticates the individual by matching the photographed blood vessel pattern with a registered blood vessel pattern, comprising: a plurality of imaging sensors;an illumination device which is disposed on a substantially same plane as the plurality of imaging sensors, the illumination device comprising a light source and a movable reflecting mirror in which an angle of reflection is variable, the illumination device irradiates a finger to be captured by the plurality of imaging sensors by directing light from the light source to the movable reflecting mirror in a direction parallel to the plane and reflecting light from the movable reflecting mirror at a variable angle;a processor in communication with the plurality of imaging sensors, the processor comprising: an image selection circuit which selects an image according to the angle of reflection of the light; andan authentication processing circuit which performs authentication processing using the image selected by the image selection circuit. 2. The finger vein authentication device according to claim 1, wherein the illumination device is disposed at a position shifted from a position where the finger is placed to a position where the finger vein authentication device is projected, and irradiates a back part of the placed finger with light. 3. The biometric authentication device according to claim 1, wherein the image selection circuit performs correction processing related to brightness or distortion of a partial image for each imaging sensor. 4. The finger vein authentication device according to claim 1, further comprising: a guide imaging sensor; anda display which superimposes and displays, on an image captured by the guide imaging sensor, a guide for adjusting placement of a subject's finger to a proper position and direction. 5. The finger vein authentication device according to claim 1, wherein the the direction of light from the light source is fixed. 6. A finger vein authentication device according to claim 1, wherein light irradiated by the illumination device is infrared light, andwherein the plurality of imaging sensors are capable of detecting infrared light. 7. A finger vein authentication device which takes a photograph of a blood vessel pattern of a finger of an individual, and authenticates the individual by matching the photographed blood vessel pattern with a registered blood vessel pattern, comprising: a plurality of imaging sensors and an optical lens assembly which capture images of different parts of the finger;an illumination device which is disposed on a substantially same plane as the plurality of imaging sensors, and irradiates a finger to be captured by the plurality of imaging sensors with light in which an irradiation angle is variable;a processor in communication with the plurality of imaging sensors, the processor comprising: an imag... | HITACHI, LTD. | G06K009/00 A61B005/1171 G01N021/01 G06F021/32 G06T001/00 G01N021/47
US10716626 | B2 | 2020-07-21 | Method and system for interactive 3D scope placement and measurements for kidney stone removal procedure | 2016-06-23 | The present teaching relates to interactive medical image processing for surgical procedure planning. In one example, a three dimensional (3D) image of a kidney is obtained. The three dimensional image is rendered on a display screen. An input is received from a user specifying a location with respect to a representation of the kidney in the rendered three dimensional image. A representation of an instrument is rendered on the display screen based on the location. The instrument is automatically aligned with an infundibulum pathway of calyx at the location with respect to the kidney. A graphical line extension is rendered on the display screen to visualize the alignment of the instrument. One or more measurements related to the kidney are determined based on the location and an anatomical structure of the kidney. | 1. A method, implemented on a computing device having at least one processor, storage, and a communication platform capable of connecting to a network for determining one or more measurements related to a kidney, the method comprising: obtaining a three dimensional (3D) image of the kidney;rendering the three dimensional image in a 3D virtual space;receiving an input from a user, the input corresponding to a direct interaction of the user within the 3D virtual space and specifying a location with respect to a representation of the kidney in the rendered three dimensional image;rendering a representation of an instrument in the 3D virtual space based on the location;automatically aligning in the 3D virtual space, the representation of the instrument with an infundibulum pathway of calyx at the location with respect to the kidney based on the received input from the user;rendering a graphical line extension from the representation of the instrument in the 3D virtual space to visualize the alignment of the representation of the instrument; anddetermining the one or more measurements related to the kidney based on the location and an anatomical structure of the kidney. 2. The method of claim 1, wherein the one or more measurements are used for at least one of: surgical planning, research of the kidney, and educational material preparation. 3. The method of claim 1, wherein the instrument is represented by a probe. 4. The method of claim 3, further comprising adjusting the alignment of the probe based on an interaction between the user and the three dimensional image. 5. The method of claim 3, further comprising adjusting the location based on an interaction between the user and the three dimensional image, wherein a tip of the probe is placed at the location. 6. The method of claim 3, further comprising computing a central curve that connects a tip of the probe and a point in the kidney, wherein the central curve goes through an urological structure of the kidney. 7. The method of claim 1, wherein the one or more measurements include at least one of the following: an infundibular width of the kidney;an infundibular length of the kidney;an infundibular height of the kidney;an infundibulo-pelvic angle of the kidney; andan inter-calyx angle of the kidney. 8. The method of claim 1, further comprising: computing a central curve that connects a tip of the instrument and a point in the kidney, wherein the central curve goes through an urological structure of the kidney. 9. A system, having a processor, a storage comprising instructions, and a hardware communication platform capable of connecting to a network for determining one or more measurements related to a kidney, wherein the instructions when executed by the system cause the system to: obtain a three dimensional (3D) image of the kidney;render the three dimensional image in a 3D virtual space;receive an input from a user, the input corresponding to a direct interaction of the user within the 3... | EDDA TECHNOLOGY, INC. | A61B034/10 A61B017/22 A61B005/00
US10702442 | B2 | 2020-07-07 | Walking aid device | 2016-05-19 | Provided is a walking aid device capable of easily offering walking training by aiding the oscillation of the user's pelvis. The walking aid device comprises a drive part which supports each of at least two or more locations of the user's lumbar region, and applies external force independently to each of the supported locations; a gait recognition unit which recognizes the user's gait; and a control unit which sets the user's ideal walking motion pattern based on a recognition result of the gait recognition unit, and controls the external force applied by the drive part to each of the locations of the lumbar region so that the user's pelvis position within the lumbar region will oscillate to match the walking motion pattern. | 1. A walking aid device which aids a user's gait, comprising: a drive part which includes a first tension adjusting part configured to adjust tension of a first wire in which one end is configured to be fixed to a location corresponding to a left side of a user's lumbar region, a second tension adjusting part configured to adjust tension of a second wire in which one end is configured to be fixed to a location corresponding to a right side of the user's lumbar region, and a frame erected relative to a walking surface of a walking belt and supports the first tension adjusting part and the second tension adjusting part so that they are configured to attain a mutually right-and-left paired relation at a position that is higher than the user's lumbar region, and the drive part is configured to apply external force independently via the first wire and the second wire while being configured to support each of the locations corresponding to the left side and the right side of the user's lumbar region;a gait recognition unit configured to measure a user's stride, walking cycle, walking speed and centroid position during a user's walking motion and recognizes a user's gait; anda control unit configured to set a user's ideal walking motion pattern by setting each of a user's pelvis inclination and rotation angle and speed thereof as optimal values, in which a time-oriented trunk balance is optimal, of the ideal walking motion pattern with a user's supported state as a reference based on a recognition result of the gait recognition unit,wherein the control unit is configured to drive the first tension adjusting part and the second tension adjusting part in synchronization with the walking motion pattern to respectively control a level of increase/decrease of tension of the first wire and tension of the second wire such that (i) a user's overall pelvis may swing alternately from side to side based on a walking direction to match changes in a time-orientated centroid position of a user's left leg and right leg and (ii) the user's pelvis itself engages in three-dimensional inclination oscillation and rotation oscillation based on a user's trunk. 2. The walking aid device according to claim 1, further comprising: a storage unit configured to store pattern data representing the ideal walking motion pattern which is measured in advance for each of a plurality of users,wherein the control unit is configured to read the pattern data corresponding to a designated user from the storage unit as the recognition result of the gait recognition unit, and thereafter controls the first tension adjusting part and the second tension adjusting part to match the ideal walking motion pattern based on the pattern data read by the control unit. 3. The walking aid device according to claim 1, wherein the drive part, among each of the locations within the user's the lumbar region, is configured to assign the first tension adjusting part and the second tension adjusting part to eac... | CYBERDYNE INC. UNIVERSITY OF TSUKUBA | A61H003/00 A61B005/11
US10707692 | B2 | 2020-07-07 | Recharge of implanted medical devices | 2017-01-26 | Systems, devices and methods are disclosed that allow recharging a power source located in an implanted medical device implanted in a patient, the recharging device comprising first and second pairs of electrical coils configured to generate first and second uniform magnetic fields in overlapping first and second cylindrical regions located between the respective pairs of electrical coils. | 1. A method for recharging a power source located in an implantable medical device, the method comprising: generating, by a signal generator, an electrical signal having a varying voltage waveform;applying, by a first driver circuit, the electrical signal to a first pair of electrical coils, the first pair of electrical coils having a first central axis, the first pair of coils configured to generate a first uniform magnetic field in a first cylindrical region located between the first pair of coils when the first pair of coils is electrically energized;applying, by a second driver circuit, the electrical signal to a second pair of electrical coils, the second pair of electrical coils having a second central axis, the second pair of coils configured to generate a second uniform magnetic field in a second cylindrical region located between the second pair of coils when the second pair of coils is electrically energized,wherein the first uniform magnetic field and the second uniform magnetic field form a resultant magnetic field within an area common to both the first cylindrical region and the second cylindrical region within an area common to both the first cylindrical region and the second cylindrical region, the resultant magnetic field operable to provide a charging energy to at least one receiving antenna of the implantable medical device located within the resultant magnetic field when the first pair of coils and the second pair of coils are electrically energized; androtating, by a position controller and a motor controlled by the position controller, the second pair of coils around the first central axis so that when rotated, the second central axis remains perpendicular to the first central axis. 2. The method of claim 1, further comprising: receiving, by communication circuitry, a feedback signal indicating a level of coupling between the resultant magnetic field and a receiving antenna of the implantable medical device located within the resultant magnetic field;determining, by processing circuitry coupled to the communication circuitry, a final position for the second pair of coils as the second pair of coils is being rotated around the first central axis based on the feedback signal; andpositioning, by the positioning circuitry and the motor, the second pair of coils at the final position. 3. The method of claim 1, wherein the first pair of coils are separated from each other by a first distance and the second pair of coils are separated from each other by a second distance such that at least a portion of a patient including the implantable medical device may be disposed in the area common to both the first cylindrical region and the second cylindrical region between the first pair of coils and the second pair of coils. 4. A system for recharging an implantable medical device, the system comprising: a first pair of electrical coils having a first central axis, the first pair of coils configured to provide a first uniform magnetic f... | Medtronic, Inc. | A61N001/372 H02J007/02 H02J050/12 H02J050/80 H02J050/40 A61N001/378 A61N001/362 A61N001/39 H01F038/14 A61B005/00 H02J050/10 A61B005/024 A61N001/375 A61B005/11 A61B005/08 A61B005/01 H02J007/00
US10718840 | B2 | 2020-07-21 | Method for preparation of highly polarized nuclear spins containing samples and uses thereof for NMR and MRI | 2016-06-26 | A method for the preparation of a highly polarized nuclear spins containing sample of an organic or inorganic material, containing H or OH groups or adsorbed water molecules. Such highly polarized nuclear spins containing samples can be subjected to nuclear magnetic resonance (NMR) measurement and/or can be thawed and immediately administered to an individual undergoing a magnetic resonance imaging (MRI) scan. The method is based on generating unstable radicals on the surface of the sample in the presence of ionized environment followed by cooling the sample to cryogenic temperatures. A device for carrying out a particular step of said method is also discloses. | 1. A method for the preparation of a sample comprising highly polarized nuclear spins, said method comprising: (i) providing a sample in the form of a powdery solid, in a sealed vessel;(ii) setting a controlled pressure and composition gas environment in said vessel by evacuating gas from said vessel maintaining a moderate vacuum level of 1×10−6 to 1×10−3 bar,(iii) generating unstable radicals in said sample by electrical ionization of the gas environment for 1-3 minutes, wherein said unstable radicals are fixed to the solid and thus do not decay for tens of minutes and up to several days under moderate conditions;(iv) cooling said sample to a cryogenic temperature in the range of 0° K to 120° K; and(v) performing a dynamic nuclear polarization (DNP) process on the cooled sample so as to transfer spin polarization from the electron spins to the nuclear spins, thereby obtaining said sample comprising highly polarized nuclear spins. 2. The method of claim 1, wherein the DNP process performed in step (v) involves microwave irradiation at a frequency of either the sum or difference of the electron and nuclear Larmor frequencies, under an applied magnetic field. 3. The method of claim 1, wherein the sample provided in step (i) comprises a material having no carbonyl groups. 4. The method of claim 1, further comprising the step of subjecting the sample obtained in step (v) to a nuclear magnetic resonance (NMR) measurement. 5. The method of claim 4, wherein said NMR measurement is solid state NMR, magic-angle spinning (MAS)-NMR, a liquid state NMR measurement, or an NMR measurement for the investigation of a surface. 6. The method of claim 1, further comprising the step or thawing the sample obtained in step (v) and immediately administering said sample, in liquid or gaseous form, to a subject undergoing a magnetic resonance imaging (MRI) scan. 7. The method of claim 6, wherein said MRI scan is metabolic MRI or functional MRI. | Technion Research & Development Foundation Limited Yeda Research and Development Co. Ltd. | G01R033/56 G01R033/30 G01R033/62 A61K049/20 G01R033/28 A61B005/055 A61M005/00 G01R033/48
US10722166 | B2 | 2020-07-28 | Method and system for biomechanical analysis of the posture of a cyclist and automatic customized manufacture of bicycle parts | 2015-12-30 | A system for biomechanical analysis of user posture and automatic customized manufacture of bicycle parts includes a servo-assisted simulator having a handlebar, a saddle, pedal cranks, and actuators, a device detecting input data that includes a 3D scanner for automatically detecting the position of body segments of the user and the angular ranges therebetween and generating three-dimensional physical data units, an electronic platform detecting pressure data of the user, a pair of insoles detecting plantar pressure, a computer connected to the actuators and to the detection device, a memory unit storing optimized initial data and instantaneous data, software comparing the optimized initial data and the instantaneous data and generating final data of the characteristics of the main parts, a spatial representation device spatially representing the final data, and a device for immediate manufacture of the parts using 3D printers. A method of biomechanical analysis and custom manufacture of bicycle parts. | 1. A method of biomechanical analysis of a posture of a user, and of automatic customized manufacture of bicycle parts, comprising the steps of: a) determining a matrix of optimal angular ranges (βREF) between body segments (B) of a user (U) at his/her joints (J);b) collecting historical data (DA) of the user (U) and converting the historical data into numeric strings;c) performing a functional morphological analysis of the user (U) by detecting positions of the body segments (B) of the user (U) with a 3D scanner (20);d) analyzing dysmetria and postural defects of the user (U) using pressure data (DW) detected in standing and bent positions by a platform (25);e) providing a computer (14) with software (SW) installed therein for processing said optimal angular ranges (βREF), said numeric strings of the historical data (DA), and said pressure data (DW) to determine initial data (DIN) concerning parts of a bicycle;f) providing pressure sensors (28) associated with insoles (27) designed to be inserted into shoes (S) of the user (U) and a servo-assisted pedaling simulator (2) assisted by said computer (14), having additional pressure sensors (29) in saddle (4) and handlebar (3) and actuators (6) for automatically changing an optimized three-dimensional position of the saddle (4) and the handlebar (3) according to said initial data (DIN);g) detecting the pressure data (DW) at a specific moment using said sensors (28, 29) and angles (β) between body segments (B) using said 3D scanner (20) during a dynamic test of the user (U) on said servo-assisted pedaling simulator (2) by setting a standard pedaling resistance thereon, according to heartbeat rate, to obtain instantaneous data (DIS);h) comparing said instantaneous data (DIS) with said initial data (DIN) by said computer (14) to obtain optimized final data (DOUT) of the bicycle, to be transmitted to said servo-assisted pedaling simulator (2) for automatic adaptation of a position of the saddle (4) and for configuration of shapes and sizes of frame (7) and pedal cranks (5);i) providing a holographic spatial representation of an image (0) of the user (U) on the servo-assisted pedaling simulator (2) and an optimized configuration of the bicycle by a 3D holographic projector (31) controlled by said computer (14), to check proper positioning of the user (U) and make any corrections to configuration; andj) transmitting said optimized final data (DOUT) to one or more 3D printers for instantaneous manufacture of the parts of the bicycle using this three-dimensional technology or, alternatively, for computerized selection of existing parts in stock in warehouses and stores from databases. 2. The method as claimed in claim 1, wherein said matrix of optimal angular ranges (βREF) of the user (U) is determined according to cycling discipline and difficulty level selected by the user (U). 3. The method as claimed in claim 1, wherein, before said dynamic test, the user (U) undergoes an analysis of sacral plane incli... | - | A61B005/22 A61B005/11 A61B005/107 B62K003/02 A61B005/00 A61B005/0205 A61B005/103 B62J001/00 B62K021/12 B62M003/08 G05B019/4099 A61B005/024
US10722729 | B2 | 2020-07-28 | Probe for localized neural optogenetics stimulation and neurochemistry recordings | 2017-01-11 | A neural probe is presented for local neural optogenetics stimulation and neurochemistry recordings. The neural probe includes a probe body, a shank extending from the probe body to a tip, a plurality of micro light-emitting diodes (LEDs) positioned across a length of a first surface of the shank for providing neuron-affecting light, a plurality of carbon devices, and a plurality of carbon electrodes positioned across a length of a second surface of the shank, the second surface in opposed relation to the first surface. The plurality of carbon electrodes can be vertically aligned carbon nanotubes or vertically aligned carbon nanofibers. The plurality of carbon electrodes can also be horizontally aligned carbon nanotubes. The plurality of micro LEDs activate neurons and the plurality of vertically aligned carbon electrodes electrochemically record neurotransmitters. | 1. A flexible neural probe comprising: a probe body;a shank extending from the probe body to a tip;a plurality of micro light-emitting diodes (LEDs) extending sequentially along an entire length of a first surface of the shank for providing neuron-affecting light; anda plurality of electrode sensor structures positioned across a length of a second surface of the shank, wherein each of the plurality of electrode sensor structures is integrated with a respective micro-LED of the plurality of micro LEDs to form single unitary components each defining a stacked configuration;wherein each of the plurality of electrode sensor structures includes a carbon nanotube or carbon nanofiber layer, carbon devices, and wiring levels such that vias extend from the electrode devices to the carbon nanotube or carbon nanofiber layer; andwherein the plurality of micro LEDs optogenetically activate neurons and the plurality of electrode sensor structures electrochemically record neurotransmitters and resulting neurochemistry, simultaneously, as neuronal data is being collected. 2. The neural probe of claim 1, wherein the wiring levels extend a horizontal length of each of the plurality of electrode sensor structures. 3. The neural probe of claim 1, wherein the carbon devices are positioned directly between the carbon nanotube or carbon nanofiber layer and the wiring levels. 4. The neural probe of claim 1, wherein each of the plurality of micro LEDs are defined on an LED die on a blanket film wafer. 5. The neural probe of claim 4, wherein the LED die is transferred to a first handler, a sapphire substrate of the wafer is removed, and a first bondpad is formed on an exposed semiconductor layer. 6. The neural probe of claim 5, wherein each of the plurality of electrode sensor structures is formed by patterning segmentation channels. 7. The neural probe of claim 6, wherein a release layer is applied and a second handler is bonded to each of the plurality of electrode sensor structures. 8. The neural probe of claim 7, wherein a polyimide (PI) layer is disposed directly between the carbon nanotube or carbon nanofiber layer and the carbon devices. 9. The neural probe of claim 1, wherein each respective micro LED is attached to each respective electrode sensor structure via first and second bondpads. 10. The neural probe of claim 1, wherein a portion of the carbon devices of an electrode structure are vertically misaligned with respect to a respective micro LED. 11. A flexible neural probe comprising: a probe body;a shank extending from the probe body to a tip;a plurality of micro light-emitting diodes (LEDs) extending sequentially along an entire length of a first surface of the shank for providing neuron-affecting light; anda plurality of electrode sensor structures positioned across a length of a second surface of the shank, wherein each of the plurality of electrode sensor structures is integrated with a respective micro-LED of the plurality of micro LEDs to for... | International Business Machines Corporation | A61N005/06 A61B005/04
US10721392 | B2 | 2020-07-21 | Method and a system for eye tracking | 2017-07-17 | According to an aspect of the present inventive concept there is provided a method for eye tracking, comprising: capturing a sequence of digital images of an eye of a user; outputting data including said sequence of images to an image processing unit; processing said data by the image processing unit to determine a sequence of positions of the eye, each position being indicative of a gaze direction, acquiring biosignal data representing an activity of the eye; and in response to detecting closing of the eye based on the acquired biosignal data, pausing at least one of said capturing, said outputting and said processing. A system for implementing the method is also disclosed. | 1. A method for eye tracking, the method comprising: capturing a sequence of digital images of an eye of a user;outputting data including said sequence of images to an image processing unit;processing said data by the image processing unit to determine a sequence of positions of the eye, each position being indicative of a gaze direction,acquiring biosignal data representing an activity of the eye;in response to detecting closing of the eye based on the acquired biosignal data, pausing said capturing; andreturning of said capturing to an active state from the paused state,wherein returning of said capturing to the active state includes transitioning said capturing from the paused state to a waiting state and transitioning said capturing from the waiting state to the active state. 2. The method according to claim 1, wherein said act of transitioning from the paused state to the waiting state is performed after a preset time delay from said act of pausing. 3. The method according to claim 1, wherein said transitioning from the waiting state to the active state is performed in response to detecting opening of the eye based on the acquired biosignal data. 4. The method according to claim 1, wherein said returning to an active state is performed in response to detecting opening of the eye based on the acquired biosignal data. 5. The method according to claim 1, wherein said returning to an active state is performed after a preset time delay from said pausing. 6. A system for eye tracking, the system comprising: an image capturing device configured to capture a sequence of digital images of an eye;an image processing unit configured to receive data including said sequence of images from the image capturing device, and determine a sequence of positions of the eye, each position being indicative of a gaze direction;a biosignal acquisition unit configured to acquire biosignal data representing an activity of the eye;a biosignal processing unit configured to receive the biosignal data and detect closing of the eye based on the acquired biosignal data; anda control unit configured to, in response to the biosignal processing unit detecting closing of the eye, outputa control signal for pausing an image capture operation of the image capturing device,wherein the control unit is further configured to output at least one control signal, wherein the image capture operation is returned to an active state from the paused state,wherein the control unit is further configured to output: a first control signal for transitioning the image capture operation in the paused state, to a waiting state; anda second control signal for transitioning the image capture operation in the waiting state, to the active state. 7. The system according to claim 6, wherein the control unit is configured to output the first control signal after a preset time delay from said output of the control signal for pausing. 8. The system according to claim 6, wherein the control unit is conf... | Stichting IMEC Nederland | H04N005/232 A61B003/113 A61B005/0496 G06F003/01 G06T007/73 G06K009/00 G06T007/20 A61B003/11
US10702073 | B2 | 2020-07-07 | Portable rocker for newborn baby or infant | 2018-02-21 | A rocker for supporting an infant or other item includes a carrier assembly and a support assembly. The carrier assembly has a support platform. The support assembly is disposed below the carrier assembly is operable to selectively impart motion to the carrier assembly in a plurality of different motion patterns. The motion patterns can include one or more of translation along a first linear direction, translation along a second linear direction perpendicular to the first linear direction, yaw relative to a first axis, pitch relative to a second axis that is perpendicular to the first axis, roll about a third axis that is perpendicular to the first and second axes, and elevation changes relative to the support assembly. | 1. An infant rocking system, comprising: a carrier assembly configured to support an infant thereon, the carrier assembly comprising a support platform having a generally planar top surface and an opening on a bottom surface of the support platform;a support assembly disposed below the carrier assembly and configured for placement on a planar surface, the support assembly configured to impart motion to the carrier assembly in one or more of rotation about a first axis that extends through the opening, translation along a first linear direction, translation along a second linear direction perpendicular to the first linear direction, yaw relative to the axis, pitch relative to a second axis that is perpendicular to the first axis and roll about a third axis that is perpendicular to the first and second axes, elevation above the support assembly, the support assembly comprising: a base platform,a frame movably disposed on the base platform and operable to provide said rotation about the first axis,a universal joint having a lower yoke fixedly mounted on the frame and an upper yoke slideably coupled within the opening, the upper yoke configured to engage the opening to impart said rotation about the first axis to the carrier assembly based on rotation of the frame,a plurality of motor assemblies mounted on the frame and arranged circumferentially about the universal joint, each of the motor assemblies configured to engage a different portion of the bottom surface of the support platform and being actuatable independently of each other to impart pitch, roll, and vertical motion to the carrier assembly, each of the motor assemblies comprising:an electric motor operable to rotate a rotor extending therefrom,a drive arm coupled with the rotor, anda drive shaft coupled with the drive arm, the drive shaft configured to translate vertically between the frame and the support platform, a vertical position of the drive shaft controlled by the position of the rotor of the motor, an upper end of the drive shaft configured to engage the bottom surface of the support platform; andan electronic controller unit operable to control an operation of the plurality of motor assemblies to impart the pitch, roll, and vertical motion on the support platform. 2. The baby rocker of claim 1, further comprising a motor assembly mounted on the base platform and configured to operatively engage the frame and to rotate about the first axis; the motor assembly comprising a pinion that engages a rack on the frame to effect rotation;the electronic controller unit operable to control operation of the motor assembly to rotate the frame. 3. The baby rocker of claim 1, wherein the base platform is mounted on a second platform, a motor assembly mounted on the second platform and configured to operatively engage the base platform to translate in the first linear direction relative to the second platform; the motor assembly comprising a pinion that engages a rack on the base to effect tra... | Steven Paperno | A47D009/02 G08B007/06 G05D003/10 G05D003/12 A61B005/11 A61B005/0205 A61B005/024 A61B005/08 A61B005/00
US10718837 | B2 | 2020-07-21 | 3D UTE imaging using variable-TE stack-of-spirals acquisition | 2017-04-21 | Some aspects of the present disclosure relate to ultrashort-echo-time (UTE) imaging. In one embodiment, a method includes acquiring UTE imaging data associated with an area of interest of a subject. The acquiring comprises applying an imaging pulse sequence with a three-dimensional (3D) spiral acquisition and a nonselective excitation pulse. The method also includes reconstructing at least one image of the area of interest from the acquired UTE imaging data. | 1. A method for ultrashort-echo-time (UTE) imaging, comprising: acquiring UTE imaging data associated with an area of interest of a subject, the acquiring comprising applying an imaging pulse sequence with a three-dimensional (3D) spiral acquisition and a nonselective excitation pulse; andreconstructing at least one image of the area of interest from the acquired UTE imaging data. 2. The method of claim 1, wherein the 3D spiral acquisition comprises a variable echo time (TE) 3D stack-of-spirals acquisition. 3. The method of claim 2, wherein the 3D stack-of-spirals acquisition uses a 3D spoiled gradient-echo sequence. 4. The method of claim 2, wherein the minimum TE is about 50 μs. 5. The method of claim 1, wherein the nonselective excitation pulse is a nonselective hard RF excitation pulse. 6. The method of claim 1, wherein the area of interest of the subject comprises an area of the musculoskeletal system of the subject. 7. The method of claim 1, wherein the area of interest of the subject comprises the lungs of the subject. 8. The method of claim 1, further comprising: based on the reconstructed at least one image, determining at least one physiological characteristic of the area of interest of the subject. 9. A system for ultrashort-echo-time (UTE) imaging, comprising: a data acquisition device configured to acquire UTE imaging data associated with an area of interest of a subject, the acquiring comprising applying an imaging pulse sequence with a three-dimensional (3D) spiral acquisition and a nonselective excitation pulse; andone or more processors coupled to the data acquisition device and configured to cause the system to perform functions comprising: reconstructing at least one image of the area of interest from the acquired UTE imaging data. 10. The system of claim 9, wherein the data acquisition device comprises a magnetic resonance imaging device. 11. The system of claim 9, wherein the 3D spiral acquisition comprises a variable echo time (TE) 3D stack-of-spirals acquisition. 12. The system of claim 11, wherein the 3D stack-of-spirals acquisition uses a 3D spoiled gradient-echo sequence. 13. The system of claim 11, wherein the minimum TE is about 50 μs. 14. The system of claim 9, wherein the nonselective excitation pulse is a nonselective hard RF excitation pulse. 15. The system of claim 9, wherein the area of interest of the subject comprises an area of the musculoskeletal system of the subject. 16. The system of claim 9, wherein the area of interest of the subject comprises the lungs of the subject. 17. The system of claim 9, wherein the one or more processors coupled to the data acquisition device are configured to cause the system to perform functions that further comprise: based on the reconstructed at least one image, determining at least one physiological characteristic of the area of interest of the subject. 18. A non-transitory computer-readable medium having stored instructions that, when executed by one or mo... | UNIVERSITY OF VIRGINIA PATENT FOUNDATION SIEMENS HEALTHCARE GMBH | G01R033/48 A61B005/00 A61B005/055
US10722361 | B2 | 2020-07-28 | Heart valve prosthesis | 2018-07-17 | A heart valve prosthesis is proposed, including a housing, at least one leaflet installed within the housing, at least one sensor and/or valve performance monitoring device. The sensor and/or monitoring device is built into the housing and/or leaflet. In an embodiment, the housing and/or leaflet has an oscillating circuit including an inductor coil connected to the sensor and/or monitoring device. One such coil can be built into the housing, and another coil can be built into the leaflet. In another embodiment, one coil can be built into a first leaflet, and another coil can be built into a second leaflet, wherein the sensor and/or monitoring device determines a degree of opening/closing the leaflets by measuring inductive interconnection between the coils, and operation timing characteristics of the valve. In an embodiment, the valve may include an emission sensor disposed within the leaflet and/or housing capable of detecting the valve's open/closed positions. | 1. A heart valve prosthesis comprising: a housing defining a clearance therein; at least one leaflet installed within the housing and capable of opening and closing the clearance; at least one sensor A built into the housing, and at least one valve performance monitoring device built into the housing; said housing includes at least one segment; and at least one sensor B built into said at least one leaflet; wherein:the at least one sensor A includes a first inductor built into the respective at least one segment;the at least one sensor B includes a second inductor built into the respective at least one leaflet;the first inductor is substantially connected to the at least one valve performance monitoring device;each pair of the first inductor and the second inductor is positioned along opposite edges of the respective at least one segment and the respective at least one leaflet forming inductive interconnection between the first inductor and the second inductor;and wherein:the at least one sensor A, and/or the at least one sensor B, and/or the at least one valve performance monitoring device determine a degree of opening and closing of said at least one leaflet and timing characteristics of operation of the heart valve prosthesis by measuring the inductive interconnection for said each pair of the first inductor and the second inductor that indicates a current spatial position of the at least one leaflet. 2. The valve of claim 1 wherein: said at least one valve performance monitoring device is further equipped with at least one antenna arranged according to one of the following configurations: the at least one antenna is a single open or a single closed circuit wiring turn placed in a segment of said housing, along one side of said at least one sensor and/or said at least one valve performance monitoring device; the at least one antenna is a single open or a single closed circuit wiring turn placed in a segment of said housing, around one side of said at least one sensor and/or said at least one valve performance monitoring device; the at least one antenna is a closed-loop multi-turn wiring around the housing; and the at least one antenna is a closed-loop multi-turn wiring placed in the at least one leaflet. 3. The valve of claim 1 wherein: said at least one sensor additionally includes a temperature sensor. 4. The valve of claim 1 wherein: said at least one sensor and/or said at least one valve performance monitoring device is equipped with a number of wireless data transmission devices. 5. The valve of claim 1, wherein said at least one leaflet includes at least a leaflet CL and a leaflet DL being adjacent to the leaflet CL; said valve further comprising at least one additional sensor C built into the leaflet CL and at least one additional sensor D built into the leaflet DL, and at least one additional valve performance monitoring device built into the leaflet CL and/or into the leaflet DL; wherein: the leaflet CL defines a third edge, and t... | Alexander V. Samkov | A61F002/24 A61B005/00 A61B005/02 A61B005/021
US10724846 | B2 | 2020-07-28 | System and method for use in depth characterization of objects | 2017-05-23 | A system is described, for use in optical measurement of a sample. The system comprising: an illumination unit configured for providing coherent illumination of one or more selected wavelength ranges and directing it onto one or more selected inspection regions of the sample, a collection unit configured for collecting light returning from the inspection region and generating output data comprising a sequence of image data pieces indicative of secondary speckle patterns formed at an intermediate plane in optical path of light collection, a depth resolving module configured for affecting at least one of the illumination unit and the collection unit for determining an association between collected secondary speckle patterns and depth layers of the sample; and a control unit being connectable to said depth resolving module and configured for operating said depth resolving module and for receiving said sequence of image data pieces from the collection unit and processing said sequence of image data pieces by determining correlation functions between at least portions of said secondary speckle patterns associated with corresponding depth layers of the sample, thereby determining one or more parameter variations along depth of the sample. | 1. A system for use in optical measurement of a sample, the system comprising: an illumination unit configured for providing coherent illumination of one or more selected wavelength ranges and directing the coherent illumination onto one or more selected inspection regions of the sample,a collection unit configured for collecting light returning from the inspection region and generating output data comprising a sequence of image data pieces indicative of secondary speckle patterns formed at an intermediate plane located between the inspection region and the collection unit,a depth resolving module comprising an interferometric unit having at least one sample arm and at least one reference arm, the reference arm is configured for receiving a reference illumination beam from the illumination unit, and configured for temporally modulating optical path of said reference arm and combining light propagating in said reference arm with collected light returning from the inspection region, thereby causing the collection unit for collecting image data piece associated with interfering speckle patterns generated by interference between light returning from the sample through a sample arm and a reference beam travelling thorough said reference arm, said coding forms an association between data in the collected secondary speckle patterns and depth layers of the sample at the inspection region; anda control unit being connectable to said depth resolving module and configured and operable operating said depth resolving module and for receiving said sequence of image data pieces from the collection unit and processing and analyzing said sequence of image data pieces by determining correlation functions between at least portions of said secondary speckle patterns associated with corresponding depth layers of the sample, and for determining one or more parameter variations along depth of the sample at said one or more inspection region, wherein said control unit comprises a flickering detection module configured and operable for receiving said sequence of image data pieces and identifying at least one portion of image data piece having repetitive flickering and for marking speckle pattern associated with said at least one portion of the frames as relating to depth layer in accordance with frequency of said repetitive flickering, thereby enabling separate processing of speckle pattern portions associated with separate depth layers of the sample. 2. The system of claim 1, further comprising a stimulating unit configured for applying a predetermined stimulation onto the sample to thereby enable detection of sample response to said predetermined stimulation. 3. The system of claim 1, wherein optical path modulation of the reference arm is provided by varying location of a mirror of the reference arm within a predetermined axial range. 4. The system of claim 3, wherein said location of said mirror being moved at constant acceleration along one direction of the axial... | CONTINUSE BIOMETRICS LTD. | G01N021/21 G01B009/02 G01B011/16 G01N021/47 G06T007/514 A61B005/00 G01N021/17 G01H009/00
US10702173 | B2 | 2020-07-07 | System and method for camera-based heart rate tracking | 2019-07-09 | A system and method for camera-based heart rate tracking. The method includes: determining bit values from a set of bitplanes in a captured image sequence that represent the HC changes; determining a facial blood flow data signal for each of a plurality of predetermined regions of interest (ROIs) of the subject captured by the images based on the HC changes; applying a band-pass filter of a passband approximating the heart rate to each of the blood flow data signals; applying a Hilbert transform to each of the blood flow data signals; adjusting the blood flow data signals from revolving phase-angles into linear phase segments; determining an instantaneous heart rate for each the blood flow data signals; applying a weighting to each of the instantaneous heart rates; and averaging the weighted instantaneous heart rates. | 1. A method for heart rate tracking of a subject, the method comprising: receiving a plurality of images captured of the subject;determining hemoglobin concentration (HC) changes from the plurality of images;determining one or more blood flow data signals based on the HC changes;determining a heart rate as an average of the one or more blood flow data signals; andoutputting the determined heart rate. 2. The method of claim 1, wherein determining the heart rate comprises determining an instantaneous heart rate for each the blood flow data signals and averaging the instantaneous heart rates. 3. The method of claim 2, wherein determining the heart rate further comprises assigning a weighting to each of the instantaneous heart rates, and averaging the instantaneous heart rates comprises averaging the weightings. 4. The method of claim 1, wherein each of the blood flow data signals is associated with a separate one of a plurality of predetermined regions of interest (ROIs) of the subject captured by the images. 5. The method of claim 4, wherein the ROIs are captured from one or more of: the face, the wrist, or the neck of the subject. 6. The method of claim 1, wherein the HC changes of the subject are represented by bit values from a set of bitplanes in the captured image sequence, the set of bitplanes being those that are determined to approximately maximize a signal-to-noise ratio (SNR), and wherein determining the HC changes comprises using a machine learning model trained with an HC changes training set, the HC changes training set comprising bit values from each bitplane of images captured from a set of subjects for which heart rate is known. 7. The method of claim 6, wherein determining the set of bitplanes that maximize SNR comprises: performing pixelwise image subtraction and addition of bitplane vectors to maximize signal differences in all ROIs over a predetermined time period;identifying bit values from bitplanes that increase the signal differentiation and bit values from bitplanes that decrease the signal differentiation or do not contribute to signal differentiation; anddiscarding the bit values from the bitplanes that decrease the signal differentiation or do not contribute to signal differentiation. 8. The method of claim 6, wherein the machine learning model comprises a Long Short Term Memory (LSTM) neural network or a non-linear Support Vector Machine. 9. The method of claim 1, wherein the method further comprises applying a band-pass filter of a passband approximating the heart rate to each of the blood flow data signals, the passband being in a range of approximately 0.6 hertz to 1.2 hertz, where 60 heartbeats-per-minute is equivalent to 1 hertz. 10. The method of claim 1, further comprising adjusting the blood flow data signals from revolving phase-angles into linear phase segments, and applying a differential filter to the linear phase segments to convert phase-angle data into frequency units representing a count value, ... | NuraLogix Corporation | G06K009/00 A61B005/024 G06N020/00 A61B005/1455 G16H030/40 G16H050/20 A61B005/026 A61B005/0255 A61B005/00 G06T007/00 A61B005/145
US10722323 | B2 | 2020-07-28 | Method and apparatus to account for transponder tagged objects used during medical procedures | 2019-05-31 | Medical procedure related objects (e.g., instruments, supplies) tagged with transponders are accounted for in a medical procedure environment via a medical object accounting system using a number of antennas, and optional readers. A first set of antennas is configured to interrogate sterile fields and/or non-sterile fields which do not encompass a body of a patient to account for the objects proximate a start and an end of a medical procedure. Readers (e.g., symbol, RFID) may be employed. A database is maintained with information including a current status of each instrument or supply, for instance as checked in or checked out. On notification of a discrepancy, a handheld antenna and/or second set of antennas interrogate a volume encompassing a body of a patient for retained instruments or supplies. The system is automatically configured (e.g., loading appropriate software) on communicative coupling of a device (e.g., antenna, reader, peripheral device). | 1. A system to account for medical procedure objects, the system comprising: a plurality of sterile fields, each of the sterile fields encompassing an area that excludes any portion of a body of a patient on which a medical procedure is performed;a first plurality of antennas communicatively coupled to at least one receiver, the first plurality of antennas and the at least one receiver providing a range that is about coextensive with the plurality of sterile fields and which range does not extend into any portion of the body of the patient when the patient is supported by a patient support structure during the medical procedure;a second plurality of antennas, also coupled to the at least one receiver, the second plurality of antennas and the at least one receiver configured to monitor a volume of medical procedure objects including portions of the body of the patient; anda control subsystem including at least one processor and at least one processor-readable storage medium, the control subsystem configured to: at least proximate a start of a medical procedure on a patient, automatically detect each of a number of medical procedure objects for performing the medical procedure within any of the sterile fields; andat least proximate an end of the medical procedure on the patient, automatically determine whether each of the number of medical procedure objects detected within any of the sterile fields at least proximate the start of the medical procedure are present in at least one sterile field of the plurality of sterile fields or at least one non-sterile field. 2. The system of claim 1, further comprising: at least one transmitter communicatively coupled to the first plurality of antennas to transmit an interrogation signal to any transponders in the plurality of sterile fields. 3. The system of claim 1, wherein the plurality of sterile fields are coextensive with at least a portion of a back table, a supply table, and an instrument table. 4. The system of claim 3 wherein one of the plurality of sterile fields is coextensive with a portable instrument tray and at least one of the plurality of antennas is physically coupled to the portable instrument tray. 5. The system of claim 1, further comprising: a plurality of substrates, each of the substrates carrying at least a respective one of the first plurality of antennas, the plurality of substrates positionable with respect to a respective one of each of the plurality of sterile fields such that a range of the at least respective one of the plurality of antennas covers the respective sterile field. 6. The system of claim 1, wherein the at least one receiver includes a plurality of receivers, each of at least some of the plurality of receivers physically coupled to a respective one of the substrates. 7. The system of claim 1, wherein the at least one receiver includes a plurality of receivers, each of at least some of the plurality of receivers physically housed in a controller console, at least... | COVIDIEN LP | A61B090/98 A61B005/00 A61B005/06 A61B050/15 A61B050/33 A61G013/10 G06K019/077 A61B005/0205 A61B090/00
US10722140 | B2 | 2020-07-28 | Localized magnetic field generator | 2015-07-01 | An apparatus for generating a magnetic field for tracking of an object (12) can include a localized magnetic field generator (76) that is configured to generate a magnetic field (621) and to control the magnetic field (621) in an area of interest (38) and configured to control the magnetic field (621) in a separate area. The separate area can be displaced from the area of interest (38) and can include a magnetic field-disrupting component (86). The object (12) can be located in the area of interest (38). | 1. An apparatus for generating a magnetic field for tracking of an object, comprising the following: a localized magnetic field generator that is configured to generate a magnetic field and to control the magnetic field in an area of interest and configured to control the magnetic field in a separate area that includes a magnetic field disrupting component, wherein: the separate area is displaced from the area of interest, andthe object is located in the area of interest; andthe localized magnetic field generator includes a magnetic field shield disposed beneath a patient examination table, spaced apart from a bottom surface of the patient examination table, and a plurality of magnetic transmitting elements disposed on a first side of the magnetic field shield, the first side of the magnetic field shield facing the patient examination table, wherein the magnetic field shield is configured to be attached to the patient examination table. 2. The apparatus of claim 1, wherein the magnetic field-disrupting component moves with respect to the localized magnetic field generator. 3. The apparatus of claim 2, wherein the magnetic field-disrupting component includes at least one of an x-ray source associated with an x-ray imaging device and a c-arm associated with the x-ray imaging device. 4. The apparatus of claim 1, wherein the localized magnetic field generator is configured to decrease a strength of the magnetic field in the separate area. 5. The apparatus of claim 1, wherein the localized magnetic field generator is configured to deflect the magnetic field away from the separate area. 6. The apparatus of claim 1, wherein the plurality of magnetic transmitting elements are configured to generate the magnetic field. 7. The apparatus of claim 6, wherein: the magnetic field shield is located between the plurality of magnetic transmitting elements and the magnetic field-disrupting component;the plurality of magnetic transmitting elements are mounted in different locations; andthe plurality of magnetic transmitting elements are mounted in different orientations. 8. The apparatus of claim 1, wherein the localized magnetic field generator is configured to increase a ratio of a strength of a magnetic field located in the area of interest versus a strength of a magnetic field located in the separate area. 9. The apparatus of claim 1, wherein: the localized magnetic field generator is configured to be placed beneath a patient examination table; andthe area of interest is located on an opposite side of the patient examination table. 10. The apparatus of claim 1, wherein the localized magnetic field generator is a mobile localized magnetic field generator configured to be placed on a patient. 11. An apparatus for generating a magnetic field for tracking of an object, comprising the following: a localized magnetic field generator that is configured to generate a magnetic field and to control the magnetic field in an area of interest and configured to con... | St. Jude Medical International Holding S.à r.l. | A61B005/00 A61B005/06 A61B034/20 A61B006/00 A61B005/042 A61B090/00
US10726846 | B2 | 2020-07-28 | Virtual health assistant for promotion of well-being and independent living | 2017-06-03 | An electronic device for providing health information or assistance includes an input configured to receive at least one type of signal selected from sound signals, verbal signals, non-verbal signals, and combinations thereof, a communication module configured to send information related to the at least one user and his/her environment to a remote device, including the sound signals, non-verbal signals, and verbal signals, the remote device being configured to analyze a condition of the at least one user and communicate condition signals to the electronic device, a processing module configured to receive the condition signals and to cause the electronic device to engage in a passive monitoring mode or an active engagement and monitoring mode, the active engagement and monitoring mode including, but not limited to, verbal communication with the at least one user, and an output configured to engage the at least one user in verbal communication. | 1. A system for providing health information or assistance, the system comprising: at least one electronic device configured to engage in verbal communication with at least one user based upon at least one type of signal selected from the group consisting of 1) sound signals from the at least one user's environment, 2) non-verbal signals from the at least one user, 3) verbal signals from the user; and 4) any combination of (1), (2), and (3), the electronic device comprising: an input module configured to receive the sound signals from the at least one user's environment, the non-verbal signals from the at least one user, and the verbal signals for the at least one user;a communication module configured to send information related to the at least one user and the at least one user's environment, including the sound signals, the non-verbal signals, and the verbal signals, to a remote device, and to subsequently receive condition signals from the remote device;a processing module configured to receive the condition signals and to cause the electronic device to operate in a passive monitoring mode or an active engagement and monitoring mode, the active engagement and monitoring mode including, but not limited to, verbal communication with the at least one user;an output configured to engage the at least one user in specific types of verbal communication selected from a group consisting of: 1) asking the at least one user a question; 2) providing related suggestions to the user; and 3) provide information to the user; and the remote device comprising:a remote communication module configured to receive the information related to the sound signals, the non-verbal signals, and the verbal signals from the at least one electronic device, and to subsequently send the condition signals to the at least one electronic device; anda rules module configured to determine, based on the sound signals, the non-verbal signals, and the verbal signals, a responsive action;wherein the responsive action is communicated from the remote device to the at least one electronic device in the form of the condition signals. 2. The system of claim 1, wherein at least one of the electronic device and the remote device further comprises: a speaker identification module configured to identify and authenticate the at least one user. 3. The system of claim 1, wherein the remote device further comprises: a temporal database configured to receive and store the information related to the at least one user and the at least one user's environment; anda decisions database configured to receive and store decisions from at least one health care provider. 4. The system of claim 3, wherein the remote device further comprises: a training module configured to selectively retrieve the information related to the at least one user and the at least one user's environment from the temporal database, selectively retrieve the decisions from the at least one health care provider from the decisions data... | SRI INTERNATIONAL | G10L017/00 A61B005/00 A61B005/117 G10L017/06 G10L025/78
US10716507 | B2 | 2020-07-21 | Palpation assisting apparatus, palpation assisting system, and palpation assisting method | 2016-08-09 | A palpation assisting apparatus includes a synchronization controller configured to synchronize a timing of acquiring pressure data output from a pressure sensor attached to a position in contact with a lesion part with a timing of acquiring positional information of a marker attached to a position corresponding to the pressure sensor, and an entering part configured to store, in a storage part, the pressure data and the positional information of the marker that are acquired at the synchronized timing in association with time information indicating a time at which the pressure data and the positional information are acquired. | 1. A palpation data processing apparatus configured to be coupled to a pressure sensor and a marker, the palpation data processing apparatus comprising: a memory;an input part; anda processor coupled to the memory and the input part, the processor configured to perform a process including while the palpation data processing apparatus is coupled to the pressure sensor and the marker, while the pressure sensor is attached to a first portion on a palpation examiner who palpates a lesion part, and while the marker is attached to a second portion on the palpation examiner, synchronizing a timing of acquiring pressure data output from the pressure sensor with a timing of acquiring positional information of the marker;storing, in the memory, the pressure data and the positional information of the marker that are acquired at the synchronized timing in association with time information indicating a time at which the pressure data and the positional information are acquired; andupon receiving an input of chief complaint data through the input part of a patient receiving palpation, storing the chief complaint data in association with time information indicating time of the input and the positional information of the marker in the memory. 2. The palpation data processing apparatus as claimed in claim 1, wherein the palpation data processing apparatus is further configured to be coupled to measurement cameras, and wherein the process further includes, while the palpation data processing apparatus is coupled to the pressure sensor, the marker, and the measurement cameras, capturing images by the measurement cameras,computing coordinates of the marker in a world coordinates system based on the positional information of the marker acquired from the images captured by the measurement cameras, the images captured by the measurement cameras are used to measure a position of the marker, andthe storing stores in the memory the pressure data and the coordinates of the world coordinates system in association with the time information. 3. The palpation data processing apparatus as claimed in claim 2, wherein the process further includes computing elastic modulus of the lesion part based on the pressure data, the coordinates of the world coordinates system, and the time information, andwherein the storing stores in the memory the elastic modulus in association with the pressure data, the coordinates of the world coordinates system, and the time information. 4. The palpation data processing apparatus as claimed in claim 3, wherein the process further includes acquiring an image of a lesion part by a lesion part imaging camera configured to image the lesion part; anddisplaying, on a display part, distribution of the pressure data corresponding to the positional information of the marker onto the acquired image of the lesion part. 5. The palpation data processing apparatus as claimed in claim 4, wherein the measurement cameras include at least two or more measurement ca... | FUJITSU LIMITED | A61B005/00 A61B005/06 A61B005/11
US10702216 | B2 | 2020-07-07 | Method and system for body scanning and display of biometric data | 2016-05-04 | A system for scanning a body to create scan data comprises a processor, a range camera capable of capturing at least a first set of depth images of the body rotated to 0 degrees, and at least a second set of depth images of the body rotated to x degrees, wherein x is >0 degrees, and x | 1. A system for scanning a body to create scan data, comprising: a processor;a range camera capable of capturing at least a first set of depth images of the body rotated to 0 degrees, and at least a second set of depth images of the body rotated to x degrees, wherein x is >0 degrees, and x &lt;360 degrees, including any clothing worn, the range camera including an infrared camera comprising an infrared light emitter that emits infrared light and at least an infrared sensor that reads reflected infrared light;a rotation device capable of rotating the body;a first set of computer instructions executable on the processor capable of calculating a first set of three dimensional points from the first set of depth images and a second set of three dimensional points from the second set of depth images;a second set of computer instructions executable on the processor capable of applying rotational and translational matrix transformations on the first set of three dimensional points to align said first set of three dimensional points with the second set of three dimensional points and then calculating and storing a square of the sum of the differences between the first set of three dimensional points with the second set of three dimensional points as an error;a third set of computer instructions executable on a processor capable of reconstructing the surface of the body by minimizing an error between the first and second set of three dimensional points by repeatedly making iterative and small changes to the matrix transformations and then re-calculating the error until the error is minimized;a fourth set of computer instructions executable on the processor capable of applying poisson surface reconstruction to construct a mesh;a fifth set of computer instructions executable on the processor capable of rigging the mesh with a biped bone structure, scaled and sized to appropriately fit the mesh;a sixth set of computer instructions executable on the processor capable of utilizing the bone structure from within the mesh to identify one or more anthropometric points of measure; anda seventh set of computer instructions executable on the processor capable of extracting anthropometric body circumferences by calculating the convex hull associated with points around a specific region of the body's surface; andan eighth set of computer instructions executable on a processor capable of creating biometrics based on the anthropometric body circumferences. 2. The system of claim 1, comprising avatar processing software capable of creating an avatar from the mesh. 3. The system of claim 2, further comprising a display capable of displaying the avatar with the anthropometric body circumferences. 4. The system of claim 1, wherein the scan data includes data of the type selected from the group consisting of: scan data collected by the range camera, 3D geometry data created from raw scan data, body measurement data determined by analyzing raw scan data and 3D mesh data, 3D... | Styku, LLC | A61B005/00 H04N005/232 G06T007/60 A61B005/107 G06K009/00 G06T019/00
US10722174 | B2 | 2020-07-28 | Skin-conformal sensors | 2015-07-10 | Various aspects as described herein are directed to skin-conformal sensor devices and methods of using the same. As consistent with one or more embodiments, a sensor device includes an upper portion and lower portion. The upper portion includes a plurality of layers including at least one sensor. The lower portion includes a layer of microstructures configured and arranged to interface with skin of a subject and to interlock the skin with the at least one sensor. | 1. A sensor device, comprising: an upper portion including a plurality of layers, the plurality of layers including a layer with sensor circuitry and another layer having a plurality of upper-portion microstructures; anda lower portion including an electrically conductive layer configured and arranged to contact the plurality of upper-portion microstructures and including an array of lower-portion microstructures configured and arranged to interface with skin of a subject and to interlock or secure the skin with the sensor circuitry, the lower-portion microstructures characterized in that: each is flexible in response to engagement of the skin, and the array of lower-portion microstructures including microstructures characterized by geometrical and material properties and having an aspect ratio not greater than ten and not less than three, in terms of height and diameter, that manifests an ability of the lower portion microstructures to prevent buckling in response to an average radial artery force of 4 kPa being applied; wherein during operation of the sensor device the buckling is prevented in the presence of said average radial artery force corresponding to artery pulse waves of the subject, while the sensor circuitry senses the artery pulse waves. 2. The sensor device of claim 1, wherein the upper-portion microstructures and the sensor circuitry are further configured and arranged to measure data selected from: pulsations of venous pulses or arterial pulses of the subject, pulse waveforms of the subject, radial artery pulses of the subject, respiration of the subject, and carotid pulses of the subject, and a combination thereof. 3. The sensor device of claim 1, wherein: the plurality of upper-portion microstructures are further configured and arranged to enhance measurement of data selected from the group consisting of one of more of the following types of pressure: pressure pulses or waveforms that arise from venous pulses or arterial pulses of the subject, pulsations of venous pulses or arterial pulses of the subject, breathing, and a combination thereof, and wherein while the flexible lower-portion microstructures are in engagement with the skin, each of the lower-portion microstructures is asymmetrically shaped and configured to respond to said one of more of said types of pressure without buckling. 4. The sensor device of claim 1, further including circuitry configured and arranged to collect pressure data from the sensor circuitry, and a wireless transmitter configured and arranged to transmit the pressure data to a remote device, and the array of lower-portion microstructures include an adhesive portion to stick to the subject, wherein the lower-portion microstructures include microstructures having an aspect ratio of ten and a spacing ratio of three, thereby being configured and arranged to mitigate voids between the sensor device and the skin of the subject and to enhance a signal-to-noise ratio (SNR) via contact between the senso... | The Board of Trustees of the Leland Stanford Junior University | A61B005/00 A61B005/0205 A61B005/024 A61B005/0402 A61B005/0476 A61B005/01
US10713800 | B2 | 2020-07-14 | Ultrasound-based geometry determination for electrophysiological mapping | 2018-05-29 | Systems and methods can be used for ultrasound-based geometry determination for cardiac mapping. A patient can be scanned with an ultrasound while wearing body surface electrodes. While the scanning takes place, the location of the ultrasound transducer can be tracked in three-dimensional space. The electrodes can be tracked and located in the same coordinate system as the image volume. Therefore, each electrode's location can be determined relative to the acquired image volume such that corresponding geometry data is generated for the heart and the electrodes. | 1. A method comprising: receiving, by a system comprising a processor, ultrasound image data based on an ultrasound transducer scanning a patient's body that includes a plurality of body surface electrodes at distributed positions across a surface of the patient's body, the ultrasound image data including at least one ultrasound image of the patient's heart;receiving, by the system, tracking data from at least one ultrasound tracking sensor associated with the ultrasound transducer and from at least one electrode tracking sensor having a predetermined position with respect to at least some of the plurality of body surface electrodes;determining locations of the ultrasound transducer in three-dimensional space based on the tracking data received as the ultrasound transducer scans the patient's body;determining locations of the plurality of body surface electrodes in the three-dimensional space based on the tracking data received from the at least one electrode tracking sensor;determining, by the system, cardiac surface data to represent locations along a surface of the patient's heart based on the tracking data and the ultrasound image data; andderiving, by the system, geometry data representing a geometric relationship between the locations of the plurality of body surface electrodes and locations along the surface of the patient's heart represented by the cardiac surface data. 2. The method of claim 1, further comprising segmenting, by the system, the at least one ultrasound image to determine the locations along the surface of the patient's heart, the locations corresponding to a plurality of points distributed across the surface of the patient's heart. 3. The method of claim 2, wherein determining the cardiac surface data comprises generating an epicardial surface mesh of the patient's heart based on the segmenting, wherein the plurality of points distributed across the surface of the patient's heart correspond to nodes of the surface mesh. 4. The method of claim 2, wherein the patient's heart is segmented based on long and short axes of the patient's heart. 5. The method of claim 2, wherein the patient's heart is segmented based on an identification of one or more anatomical landmarks of the patient's heart. 6. The method of claim 5, wherein the one or more anatomical landmarks comprise at least one of the atrioventricular groove, the pulmonary artery, the aortic arch, and the apex. 7. The method of claim 1, wherein the at least one tracking sensor comprises a magnetic sensor, an electromagnetic sensor, or an optical sensor. 8. The method of claim 1, wherein the relationship between the locations of the ultrasound transducer reflected by the tracking data and the positions of the plurality of body surface electrodes expresses spatial locations of the plurality of body surface electrodes in a common three-dimensional coordinate system with the location of the ultrasound transducer. 9. The method of claim 8, wherein the positions of the... | CARDIOINSIGHT TECHNOLOGIES, INC. | G06T007/20 G06T007/292 A61B005/06 A61B005/00 A61B008/00 A61B008/08 A61B005/0408 G06T007/12 G06T007/62 G06T007/10 G06T007/30 A61N001/18 A61N001/00 A61N001/05
US10722183 | B2 | 2020-07-28 | Sleep signal conditioning device and method | 2017-04-06 | A method and device are provided for conditioning a sleep signal from a sleep sensor for a sleep monitor. A sleep sensor is configured to sense a physiological signal, such as a cardiac or respiratory signal from a sleeping mammal, such as a human, e.g. an infant or a baby. If the mammal exhibits also (gross) body movement during his sleep, the sensor signal may be clipped. The method and device provide a conditioning of the sleep signal by providing an estimation of the sleep signal instead of the conditioned sleep signal itself when the sensor signal is clipping. | 1. A sleep signal conditioning device for a sleep monitor, the device comprising: a first processing unit for providing a first parameter indicative of the magnitude of a first sensor signal from a first sleep sensor sensing a physiological signal from a sleeping mammal;a second processing unit for providing a second parameter indicative of the magnitude of a second sensor signal from a second sleep sensor sensing a physiological signal from the sleeping mammal, wherein the first sensor signal is different from the second sensor signal;a clipping detector for detecting when the first sensor signal is clipping;a divider for calculating a ratio based on dividing the first parameter by the second parameter when the first sensor signal is not clipping;a multiplier for generating an estimated first sensor signal based on multiplying a second output signal based on the second sensor signal by the ratio; anda switch for providing a first output sleep signal based on the estimated first sensor signal when detecting that the first sensor signal is clipping and a first output signal based on the first sensor signal when detecting that the first sensor signal is not clipping. 2. The device according to claim 1, wherein the first and second processing unit are arranged to provide an amplitude or an energy as the parameter indicative of the magnitude of the signal. 3. The device according to claim 1, wherein the first and second processing units comprise an IIR filter and/or FIR filter. 4. The device according to claim 1, wherein at least one of the processing units is arranged to determine the parameter indicative of the magnitude of the signal based also on a signal portion after a preceding signal portion where the first sensor signal was clipping, while the divider is arranged to use the determined parameter in the calculation of the ratio used during said preceding signal portion. 5. The device according to claim 1, wherein an ability to extract respiratory and cardiac features from a current signal portion is determined and the divider is arranged in dependence thereon to control whether or not to include the current signal portion in the calculation of the ratio. 6. The device according to claim 1, wherein the device comprises a first conditioning unit for conditioning the first sensor signal to the first output signal and a second conditioning unit for conditioning the second sensor signal to the second output signal, or the first parameter as provided by the first processing circuit constitutes the first output signal and the second parameter as provided by the second processing circuit constitutes the second output signal. 7. The device according to claim 1, wherein the device comprises a second clipping detector for detecting when the second sensor signal is clipping; and the divider is arranged for calculating the ratio when the second sensor signal is not clipping. 8. The device according to claim 1, comprising: a plurality of processing u... | KONINKLIJKE PHILIPS N.V. | A61B005/00 A61B005/0205 A61B005/024 A61B005/08 A61B005/11
US10722175 | B2 | 2020-07-28 | System and apparatus comprising a multisensor guidewire for use in interventional cardiology | 2015-07-10 | A system (1) and apparatus comprising a multisensor guidewire (100/200/300) for use in interventional cardiology, e.g., Transcatheter Valve Therapies (TVT), comprises a plurality of optical sensors (10/20) for direct measurement of cardiovascular parameters, e.g. transvalvular blood pressure gradients and flow. A conventional outer coil wire (35) contains a shaped core wire (31) having a cross-section defining a channel surface (132), e.g. grooves (32), extending along its length, to position optical fibers (11) and optical sensors (10/20) in a channel (33). Advantageously, the core wire has a diameter that provides sufficient stiffness to the guidewire for use as a support guidewire for TVT, e.g. Transcatheter Aortic Valve Implantation (TAVI), while accommodating multiple sensors and fibers within a guidewire of outside diameter ≤0.89 mm. An optical connector (112) couples the guidewire to a control system (150). Optionally, the guidewire includes a contact force sensor (60), a pre-formed tip (400-1/400-2) and a separable micro-connector (140). | 1. A multisensor guidewire assembly for measuring blood pressure concurrently at multiple locations during transcatheter heart valve therapies (TVT), the multisensor guidewire assembly comprising distal and proximal parts connected by a micro optical coupler that is separable to enable over the guidewire delivery of TVT components;the distal part of the multisensor guidewire assembly comprising: a guidewire comprising a tubular covering layer comprising a flexible coil (coil) of a first stiffness, the coil having a length extending between a proximal end and a distal end, an inner member having a second stiffness extending within the coil from the proximal end to the distal end; the distal end comprising a flexible distal tip; and the proximal end comprising a male part of the optical coupler;a plurality of optical sensors and a plurality of optical fibers; a sensor end of each optical fiber being attached and optically coupled to an individual one of the optical sensors; the plurality of optical sensors comprising at least two optical pressure sensors;the plurality of optical fibers and the inner member forming an assembly wherein sensor ends of each optical fiber are arranged to form a sensor arrangement wherein said plurality of optical sensors are positioned at respective sensor locations spaced apart lengthwise within a distal end portion of the guidewire;a proximal end of each of the plurality of optical fibers being coupled to an optical input/output comprising the male part of the optical coupler;the proximal part of the guidewire assembly comprising:a flexible optical cable containing a plurality of optical fibers and having at its distal end a female part of the optical coupler and at its proximal end an optical input/output for connection to an optical controller;the male and female parts of the optical coupler being configured to mutually align and optically couple the plurality of optical fibers of the distal and proximal parts of the guidewire assembly for operation of the optical pressure sensors; andthe male and female parts of the optical coupler being separable and the male part of the optical coupler having an outside diameter that accepts TVT components for over the guidewire delivery. 2. The multisensor guidewire assembly of claim 1, wherein the male part of the optical coupler has an outside diameter no greater than a maximum outside diameter of the coil of the guidewire. 3. The multisensor guidewire assembly of claim 2, wherein the inner member comprises a core wire and a proximal end of the core wire comprises a tapered portion that extends to form a core of the male part and the plurality of optical fibers extend around the tapered portion, around the core, and through a surrounding body of the male part. 4. The multisensor guidewire assembly of claim 3, wherein the optical coupler further comprises at least one of: alignment means comprising corresponding facets of the male and female parts of the optical coupler for ... | HemoCath Ltd. | A61B005/00 A61B005/042 A61B005/0215 A61F002/24
US10721696 | B2 | 2020-07-21 | Data synchronization across multiple sensors | 2019-03-20 | Various embodiments of the present invention provide methods, systems, apparatus, and computer program products for synchronizing data across at least two sensors. In one embodiment, first sensor data, which includes a universal time stamp from a universal time source, is recorded from at least one first sensor and second sensor data, which does not include a universal time stamp from a universal time source, is recorded from at least one second sensor. The first sensor data is correlated with the second sensor data, and a universal time that corresponds with the data from the second sensor is determined at least in part based on the correlation between the first sensor data and the second sensor data. | 1. A method of synchronizing data of a first sensor of a vehicle with data of a second sensor of the vehicle comprising: recording, by a vehicle telematics device positioned in the vehicle, first sensor data provided by the first sensor of the vehicle, wherein the first sensor data comprises (a) a first local time stamp based at least in part on a local time source indicating a time at which the first sensor data was recorded by the vehicle telematics device, and (b) a first universal time stamp based at least in part on a universal time source indicating a time at which the first sensor data was measured by the first sensor;recording, by the vehicle telematics device positioned in the vehicle, second sensor data provided by the second sensor of the vehicle, wherein the second sensor data (a) comprises a second local time stamp based at least in part on the local time source indicating a time at which the second sensor data was recorded by the vehicle telematics device but (b) does not comprise a second universal time stamp based at least in part on the universal time source indicating a time at which the second sensor data was measured by the second sensor;determining a time shift, by the vehicle telematics device, between the first sensor data and the second sensor data, wherein the determining the time shift comprises performing cross-correlation between the second sensor data and the first sensor data;determining a universal time that corresponds to the second sensor data based at least in part on the time shift between the first sensor data and the second sensor data; andsynchronizing the second sensor data with the first sensor data based at least in part on the universal time. 2. The method of claim 1, wherein the universal time source comprises a global navigation satellite system. 3. The method of claim 1 further comprising: interpolating the first sensor data to generate additional first sensor data such that a time step of the first sensor data equals a time step of the second sensor data. 4. The method of claim 3, wherein interpolating is selected from the group consisting of spline interpolation, linear interpolation, and polynomial interpolation. 5. The method of claim 1 further comprising: determining an orientation of the second sensor; andadjusting the second sensor data based at least in part on the determined orientation. 6. The method of claim 1 further comprising: converting the first sensor data from a first coordinate system associated with the first sensor data to a second coordinate system associated with the second sensor data. 7. The method of claim 1 further comprising: transforming the first sensor data to convert the first sensor data from a first value measured by the first sensor to a second value corresponding to a type of data measured by the second sensor. 8. The method of claim 1 further comprising: associating a universal time stamp corresponding to the determined universal time with the second sensor d... | SAFECO INSURANCE COMPANY OF AMERICA | H04W056/00 H04L007/00 A61B005/00
US10709360 | B2 | 2020-07-14 | Medical device with integrated biosensor | 2018-09-04 | The disclosure concerns medical devices configured to utilize nitric oxide (NO) donor materials to release NO gas upon contact with moisture in order to mitigate bacterial proliferation, coagulation, or a combination thereof, at a site of a biosensor integrated within the medical device. A novel multi-lumen medical device is described, the device having a first lumen for containing the NO donor material and producing the NO gas, and having a second lumen for housing biosensor componentry. A channel between the first and second lumens serves to communicate NO gas from the first lumen to a sensor-volume of the biosensor. | 1. A medical device (100) with integrated bio-sensor, comprising: a tubular body (101), the tubular body comprising a first lumen (250) and a second lumen (350) therein;the first lumen comprising a nitric oxide donor material (201) disposed therein, wherein said nitric oxide donor material is configured to release nitric oxide gas;the second lumen being longitudinally disposed parallel with respect to the first lumen and configured to house componentry of at least a first biosensor (300), the first biosensor having a first sensor-volume (375) associated therewith;wherein a first nitric oxide channel (275) is disposed between the nitric oxide donor material and the first sensor-volume, the first nitric oxide channel being configured to direct a flow of the nitric oxide gas into the first sensor-volume for preventing bacterial-proliferation, coagulation, or a combination thereof,wherein the first biosensor and the nitric oxide donor material are each disposed in a distinct lumen of the first and second lumens. 2. The medical device of claim 1, wherein the first bio-sensor comprises: a working electrode (301) and a reference electrode (321),the working electrode comprising: an insulation layer (311) covering at least a portion of the working electrode and configured to restrict diffusion of electroactive interference compounds through the insulation layer;an analyte reaction layer (312) comprising one or more immobilized or chemically linked antibodies, redox enzymes, or chelators for reacting with one or more corresponding analytes to achieve a detectable signal; andan analyte diffusion-control layer (313) for regulating diffusion of analytes of interest into the first sensor-volume of the first biosensor;the reference electrode being positioned in proximity to the working electrode. 3. The medical device of claim 2, wherein each of the working and reference electrodes are coupled to an electrode signal monitoring system. 4. The medical device of claim 2, wherein the insulation layer comprises: one or more perflourinated sulfonic acid resin-containing sub-layers, and a plurality of electropolymerized sub-layers disposed on an outer surface of the one or more perflourinated sulfonic acid resin-containing sub-layers. 5. The medical device of claim 4, wherein the one or more perflourinated sulfonic acid resin-containing sub-layers comprise sulfonated tetrafluoroethylene and one or more of: perfluoro (alkyl vinyl ether), sulfonyl acid fluoride, perfluorocycloalkene, vinyl fluoride, vinylidene fluoride, chlorotrifluoroethylene, propylene, hexafluoropropylene, perfluoropropylvinylether, perfluoromethylvinylether, polychlorotrifluoroethylene, perfluoroalkoxy polymer, fluorinated ethylene-propylene, polyethylenetetrafluoroethylene, polyethylenechlorotrifluoroethylene, perfluoroelastomer, tetrafluoroethylene-propylene, polyurethane, polyetheline, silicone or any combination thereof. 6. The medical device of claim 4, wherein the one or more perflourinat... | BIOCREDE INC. | A61B005/145 A61B005/00 A61B005/1486 C12Q001/00
US10709348 | B2 | 2020-07-14 | Standing motion assist device, standing motion assist method, and recording medium | 2017-08-24 | A standing motion assist device includes a support mechanism attached to a leg of a user to assist the user with motion, a sensor including first and/or second sensors that measure and output posture information and myoelectric potentials of the user, respectively, the sensor outputting measurement data including first measurement data measured after a sitting motion of the user sitting in a chair starts and second measurement data measured after the sitting motion ends, a memory that stores the measurement data, and a processor that controls the support mechanism. The processor detects the sitting motion and identifies the chair type based on the first measurement data, detects the start of a motion of the user standing from the chair based on the second measurement data, and outputs assist information corresponding to the identified chair type and used by the support mechanism to assist the user with standing. | 1. A standing motion assist device comprising: a support mechanism attached to a leg of a user, the support mechanism supporting the user with motion;a sensor including at least one of a first sensor and a second sensor, the first sensor measuring posture information in accordance with postures of the user and outputting the posture information, the second sensor measuring myoelectric potential information including myoelectric potentials of the user and outputting the myoelectric potential information, the sensor outputting measurement data including at least one of the posture information and the myoelectric potential information;a memory that stores the measurement data; anda processor that controls the support mechanism by using the measurement data stored in the memory,wherein the processor detects a sitting motion of the user sitting in a chair on the basis of the measurement data,wherein the detection of the sitting motion includes a detection of a start of the sitting motion and a detection of an end of the sitting motion,wherein the processor identifies a type of the chair on the basis a first measurement data included in the measurement data and measured after the start of the sitting motion,wherein the processor detects a start of a standing motion of the user standing from the chair on the basis of second measurement data included in the measurement data and measured after the end of the sitting motion, andwherein the processor outputs assist information used to cause the support mechanism to assist the user with the standing motion in accordance with the identified type of the chair. 2. The standing motion assist device according to claim 1, wherein the myoelectric potential information includes first myoelectric potentials of a muscle included in muscles in the leg of the user. 3. The standing motion assist device according to claim 2, wherein the measurement data includes the first myoelectric potentials, andwherein the processor detects the start of the sitting motion if a myoelectric potential, included in the first myoelectric potentials, is greater than or equal to a threshold value corresponding to the muscle. 4. The standing motion assist device according to claim 1, wherein the posture information is at least one of accelerations, angular velocities, and geomagnetisms of an upper body of the user. 5. The standing motion assist device according to claim 4, wherein the measurement data includes the accelerations, wherein the processor calculates movement distances by which the upper body of the user moves in a vertical direction on the basis of the accelerations, andwherein if a movement distance, included in the movement distances, is greater than or equal to a threshold value, the processor detects the sitting motion. 6. The standing motion assist device according to claim 4, wherein the measurement data includes the accelerations, wherein the measurement data includes third measurement data and fourth measurement data ... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | A61G005/14 A61B005/04 A61B005/00 A61B005/11 A61H003/02 A61H003/00 A61H001/02 B25J009/00 A61B005/0488
US10722134 | B2 | 2020-07-28 | Gel-assisted electroencephalogram sensor electrode | 2017-12-27 | An electroencephalogram (EEG) system is disclosed. The EEG system includes an EEG controller and an EEG sensor that includes a contact surface, the contact surface including an electrically-conductive portion in communication with the EEG controller and a sensor release element in communication with the EEG controller, the sensor release element being configured to perform, in response to a signal from the EEG controller, a release action to reduce adhesion of an electrically-conductive gel between the contact of the sensor and the user's skin. | 1. An electroencephalogram (EEG) system, comprising: an EEG controller; andan EEG sensor, comprising:a contact surface, the contact surface comprising an electrically-conductive portion in communication with the EEG controller; anda heating elementpositioned proximate to an aperture in the EEG sensor, the aperture being arranged to serve as a thermal conduit between the heating element and an electrically conductive gel on the contact surface of the EEG sensor, the heating element controllable by a signal from the EEG controller, to reduce adhesion of the electrically-conductive gel between the contact surface of the sensor and a user's skin,wherein the EEG controller is configured to control an output temperature of the heating element. 2. The EEG system of claim 1, wherein the EEG controller is configured to adjust the output temperature of the heating element to heat the electrically-conductive gel to a temperature that reduces adhesion of an adhesive in the electrically-conductive gel. 3. The EEG system of claim 1, wherein the EEG controller is configured to adjust the output temperature of the heating element to heat the electrically-conductive gel to a temperature within the range of 100° F. to 120° F. 4. The EEG system of claim 1, wherein the EEG sensor comprises a hygroscopic material applied to the contact surface. 5. The EEG system of claim 4, wherein, when heated by the heating element, the hygroscopic surface absorbs the electrically-conductive gel. 6. The EEG system of claim 4, wherein, when heated by the heating element, the hygroscopic surface releases to or absorbs water from the electrically-conductive gel to reduce adhesive forces between the EEG sensor and the user's skin. 7. The EEG system of claim 1, wherein an adhesiveness of the electrically-conductive gel is reduced by sweat from the user's skin, and wherein the EEG controller is configured reduce the adhesiveness of the electrically-conductive gel by adjusting the output temperature of the heating element to increase user perspiration. 8. The EEG system of claim 1, wherein the EEG controller is configured to adjust the output temperature of the heating element to reduce an amount of heat applied to the electrically-conductive gel to a temperature that reduces adhesion of an adhesive in the electrically-conductive gel. 9. The EEG system of claim 1, wherein the EEG sensor further comprises a gel dispenser, the gel dispenser comprising a gel reservoir; anda gel dispense mechanism in fluidic communication with the gel reservoir. 10. The EEG system of claim 9, wherein the gel dispense mechanism comprises a contact-driven release mechanism. 11. The EEG system of claim 10, wherein the gel dispense mechanism is disposed within the gel reservoir and the gel dispense mechanism comprises: a hollow shaft housing a spring and a probe, wherein the spring is arranged to exert a force on the probe that biases the probe to seal a dispense pathway through which gel is dispensed f... | X Development LLC | A61B005/0478 A61B005/00 A61B005/0484
US10722153 | B2 | 2020-07-28 | Medical photometer and medical photometer control method | 2017-10-25 | A medical photometer includes a processor and a memory that stores an instruction readable by a computer. When the instruction is executed by the processor, the medical photometer acquires an intensity of one of a first intensity signal and a second intensity signal, the first intensity signal corresponding to an intensity of a first light beam that has a first wavelength and that is transmitted through or reflected from a tissue of a living body, and a second intensity signal corresponding to an intensity of a second light beam that has a second wavelength and that is transmitted through or reflected from the tissue of the living body acquires an intensity ratio of the first intensity signal and the second intensity signal, and detects an artifact of the living body based on the acquired intensity and the acquired intensity ratio. | 1. A medical photometer comprising: a first light emitter that emits a first light beam having a first wavelength;a second light emitter that emits a second light beam having a second wavelength;a detector that outputs a first intensity signal corresponding to an intensity of the first light beam that is transmitted through or reflected from a tissue of a living body, and a second intensity signal corresponding to an intensity of the second light beam that is transmitted through or reflected from the tissue;a processor; anda memory that stores an instruction readable by a computer,wherein, when the instruction is executed by the processor, the medical photometer acquires an intensity of one of the first intensity signal and the second intensity signal, acquires an intensity ratio of the first intensity signal and the second intensity signal, detects a variation rate of the intensity ratio, cancels a temporal variation of the intensity ratio corresponding to pulsation of blood in the living body from the variation rate of the intensity ratio, detects an artifact in a variation rate of the intensity ratio, and calculates at least one of arterial oxygen saturation, a fraction of carboxyhemoglobin, a fraction of methemoglobin, and a concentration of total hemoglobin based on the artifact, both the first light beam and the second light beam are infrared light beams or red light beams, a ratio of a first extinction coefficient of a first blood light absorber at the first wavelength and a second extinction coefficient of the first blood light absorber at the second wavelength is different from 1, a third extinction coefficient of a second blood light absorber at the first wavelength is higher than the first extinction coefficient, a fourth extinction coefficient of the second blood light absorber at the second wavelength is higher than the second extinction coefficient, and a ratio of the third extinction coefficient and the fourth extinction coefficient is approximable by 1. 2. The medical photometer according to claim 1, wherein the artifact is detected based on a temporal variation of the acquired intensity ratio in a first coordinate plane in which a first axis represents a variation of the intensity of the one of the first intensity signal and the second intensity signal, and a second axis represents a variation of the intensity ratio. 3. The medical photometer according to claim 2, wherein, when the instruction is executed by the processor, the medical photometer reduces noises that is superimposed on the first intensity signal and the second intensity signal, and that is caused by the artifact, through a coordinate transformation from the first coordinate plane to a second coordinate plane in which a third axis represents a signal component, and a fourth axis represents an artifact component. 4. The medical photometer according to claim 1, further comprising a notifying section that performs notification when the detected artifact exceeds a th... | Nihon Kohden Corporation | A61B005/1455 A61B005/145 A61B005/00 G01N021/59 G01N021/27 G01N021/31 G01J001/04
US10709855 | B2 | 2020-07-14 | Apparatus and method for improved assisted ventilation | 2017-10-30 | Devices and methods for allowing for improved assisted ventilation of a patient. The methods and devices provide a number of benefits over conventional approaches for assisted ventilation. For example, the methods and devices described herein permit blind insertion of a device that can allow ventilation regardless of whether the device is positioned within a trachea or an esophagus. In addition, the methods and device allow for timed delivery of ventilations based on a condition of a thoracic cavity to increase the amount and efficiency of blood flow during a resuscitation procedure. | 1. A method of artificially ventilating an individual, the method comprising: inserting a working end of a ventilation device within a respiratory opening of the individual and positioning the working end of the ventilation device within a body passageway of the individual, where the working end includes a first opening fluidly coupled to a first lumen and a second opening fluidly coupled to a second lumen, where the second opening is located along the ventilation device proximal to the first opening, the ventilation device also having a pressure sensor configured to detect pressure changes within the body passageway where the ventilation device is configured for an automatic delivery of a bolus of air into an airway of the individual at a pre-determined rate;drawing suction through the first opening to induce collapse of the body passageway and maintaining the suction for a period of time;delaying delivery of the bolus of air when the sensor detects a pressure change in the air within the body passageway, where the pressure change within the body passageway results from a chest compression; anddelivering the bolus of air through the first lumen upon failure to detect collapse of the body passageway and delivering the bolus of air through the second lumen while maintaining the suction upon detecting collapse of the body passageway. 2. The method of claim 1, where delivering the bolus of air into the airway of the individual occurs after a pre-determined delay. 3. The method of claim 1, where the sensor is configured to intermittently detect pressure changes within the body passageway. 4. The method of claim 1, further comprising, after delaying the delivery of the bolus of air, the ventilation device resumes delivering the bolus of air at the predetermined rate if the pressure sensor fails to detect a pressure range within a first period of time. 5. The method of claim 1, further comprising measuring a condition of a thoracic cavity using the ventilation device. 6. The method of claim 5, where measuring the condition of the thoracic cavity comprises measuring a change in a compression of a chest of the patient. 7. The method of claim 6, where measuring the change of the compression of the chest of the patient comprises observing a force applied to the ventilation device by the body passageway. 8. The method of claim 6, where measuring the change of the compression of the chest of the patient comprises observing a deflection of the chest of the patient using one or more deflection sensors on the chest. 9. The method of claim 5, where measuring the condition of the thoracic cavity comprises measuring a state of air flow within the thoracic cavity. 10. The method of claim 9, where the ventilation device further comprises a second sensor and where measuring the state of air flow within the thoracic cavity comprises detecting airflow using the second sensor. 11. The method of claim 9, where delaying the delivery of the bolus of air comprises... | CoLabs Medical, Inc. | A61M016/04 A61M016/00 A61M016/20 A61M016/06 A61B005/087 A61B005/00 A61B005/08 A61B005/091 A61M025/00 A61N001/05 A61N001/362 A61N001/39
US10722141 | B2 | 2020-07-28 | Active voltage location (AVL) resolution | 2018-04-30 | A method includes transmitting into a heart of a patient multiple voltages using multiple respective surface-electrodes attached externally to the patient. Repeatedly per each region of a plurality of regions of the heart (i) multiple respective calibration-voltages are measured in the region; (ii) a respective partial sub-set of the surface-electrodes is selected for the region, whose corresponding calibration-voltages minimize a cost-function, wherein the cost-function includes a first term that depends on magnitudes of gradients of the measured calibration-voltages, and a second term that depends on correlations among the gradients; and (iii) the partial sub-set of surface-electrodes selected for the region, are recorded for use in a subsequent session; and, in the subsequent session, a position of a probe inserted into a given region of the heart is calculated based only on the sub-set selected for the given region. | 1. A method, comprising: transmitting into a heart of a patient multiple voltages using multiple respective surface-electrodes attached externally to the patient;repeating per each region of a plurality of regions of the heart: (i) measuring multiple respective calibration-voltages in the region;(ii) selecting for the region a respective partial sub-set of the surface-electrodes whose corresponding calibration-voltages minimize a cost-function, wherein the cost-function comprises a first term that depends on magnitudes of gradients of the measured calibration-voltages, and a second term that depends on correlations among the gradients; and(iii) recording the partial sub-set of surface-electrodes, selected for the region, for use in a subsequent session; andin the subsequent session, calculating a position of a probe inserted into a given region of the heart, based only on the sub-set selected for the given region. 2. The method according to claim 1, wherein measuring the calibration-voltages comprises receiving location measurements from a location-measuring system that tracks a calibration-tool at multiple locations in the heart, and deriving respective relations between the multiple location measurements and respective sets of measured calibration-voltages. 3. The method according to claim 2, wherein the location-measuring system comprises a magnetic location-tracking system. 4. The method according to claim 2, wherein the location-measuring system comprises a medical imaging system. 5. The method according to claim 2, wherein calculating the position of the probe comprises: measuring multiple investigation-voltages using the probe;deriving an identity of the given region from the measured investigation-voltages and the derived relations; andcalculating the position of the probe, based only on the measured investigation-voltages that originate from the surface-electrode belonging to the sub-set selected for the given region. 6. The method according to claim 5, wherein calculating the position of the probe comprises calculating the position of the probe based on the selected sub-set of measured investigation-voltages and based on a respective sub-set of the derived relations. 7. The method according to claim 1, wherein transmitting the multiple voltages comprises selecting one surface-electrode to serve as common ground, and applying the voltages between one or more remaining surface-electrodes and the common ground. 8. The method according to claim 3, wherein measuring the investigation-voltages comprises measuring the investigation-voltages using one or more investigation-electrodes fitted at a distal end of the probe. 9. A system, comprising: multiple surface-electrodes, which are attached externally to a patient and are configured to transmit multiple respective voltages into a heart of the patient;a processor, which is configured to: repeat per each region of a plurality of regions of the heart: (i) receiving multiple calibration-vo... | Biosense Webster (Israel) Ltd. | A61B005/06 A61B005/0428 A61B005/05 A61N001/08 G01R033/02
US10709332 | B2 | 2020-07-14 | Transcutaneous analyte sensor | 2019-08-16 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A system for measuring an analyte concentration in a host, the system comprising: a sensor system comprising:a first component comprising a sensor configured to measure a signal indicative of an analyte concentration in the host, and a second component comprising sensor electronics configured to operably connect to the sensor, wherein the sensor electronics comprise: one or more memories configured to store data indicative of analyte concentrations derived at least in part from measurements from the sensor;a radio frequency module configured to wireles sly transmit at least some of the data indicative of analyte concentrations to a receiver that comprises a display configured to display one or more estimated analyte concentration values, wherein the radio frequency module is configured to transmit at least some of the data indicative of analyte concentrations to the receiver in response to the receiver interrogating the sensor electronics;wherein the first component is unconnected to the second component during a first state, wherein during a second state the first component is configured to be connected to the second component by the host to form an assembly before the assembly is affixed to the host for sensor insertion. 2. The system of claim 1, wherein the second component comprises a housing, wherein the sensor comprises an in vivo portion and an ex vivo portion, and wherein the housing is configured to receive the ex vivo portion of the sensor. 3. The system of claim 1, wherein the sensor electronics are configured to receive inductive power from the receiver. 4. The system of claim 1, wherein the system is configured to have a sensor lifetime of about 14 days. 5. The system of claim 1, wherein both the first component and the second component are labeled with an alphanumeric code. 6. The system of claim 5, wherein the alphanumeric code is unique to the system. 7. The system of claim 5, wherein the alphanumeric code is associated with sensor calibration information. 8. The system of claim 5, wherein the alphanumeric code is stored in the sensor electronics. 9. The system of claim 8, wherein the alphanumeric code is associated with sensor calibration information. 10. The system of claim 8, wherein the sensor electronics are configured to transmit the alphanumeric code upon activation of the sensor system. 11. The system of claim 9, wherein the sensor electronics are configured to transmit the alphanumeric code upon activation of the sensor system. 12. The system of claim 11, wherein the sensor calibration information is configured to enable use of the system without requiring a reference analyte measurement during sensor use. 13. The system of claim 1, wherein during a third state the first component is connected to the second component. 14. The system of claim 13, wherein during a fourth state the sensor is transcutaneously implanted in the host. 15. The system of claim 1, wherein the radio frequency module comprises an ant... | DexCom, Inc. | A61B005/00 A61B005/15 A61B005/145 A61B005/1486 G16Z099/00 A61B005/1473 G06F019/00
US10712797 | B2 | 2020-07-14 | Two-phase deployment-initiated wakeup mechanism for body-mountable electronic device | 2018-10-23 | The technology described herein is related to a two-phase deployment-initiated wakeup mechanism for a body-mountable electronic device. During a first phase of the two-phase wakeup mechanism, a motion sensor detects an acceleration event indicative of deployment of the device onto the body of the user. During a second phase of the two-phase mechanism, control circuitry can be adapted to be enabled by the acceleration event. Once enabled, the control circuitry can verify that the device has been launched onto the body of a user via a deployment applicator in which the device is retained until deployment. Once verified, the control circuitry can wake up the body-mountable electronic device by transitioning the device from a sleep state to a functional (or operational) state. | 1. A mountable device deployment system comprising: a body-mountable electronic device including: a motion sensor adapted to detect a first deployment indicator, the first deployment indicator comprising an acceleration event indicative of deployment of the body-mountable electronic device, andcontrol circuitry adapted to transition the body-mountable electronic device from a sleep state to a wakeup verification state responsive to the first deployment indicator, and to transition the body-mountable electronic device from the wakeup verification state to an operational state responsive to detecting a second deployment indicator; anda deployment applicator having the body-mountable electronic device retained therein, the deployment applicator including a deployment mechanism adapted to launch the body-mountable electronic device. 2. The system of claim 1, wherein the body-mountable electronic device further includes a biosensor adapted to detect an analyte or interstitial fluid. 3. The system of claim 2, wherein the deployment mechanism comprises an extendable spring-loaded needle adapted to insert the biosensor into a body of a user upon deployment, and wherein the acceleration event is indicative of the launching of the extendable spring-loaded needle. 4. The system of claim 3, wherein the deployment applicator is adapted to retain the body-mountable electronic device until the extendable spring-loaded needle retracts and thereby disposing the body-mountable electronic device on the body of the user and the biosensor beneath a surface of skin of the user. 5. The system of claim 3, wherein the body-mountable electronic device further includes: a base having a bio-compatible adhesive disposed on a proximal surface for attaching the body-mountable electronic device to skin of the user. 6. The system of claim 2, wherein the control circuitry is further adapted to detect the second deployment indicator when an electrical current measured by the biosensor matches a preset pattern. 7. The system of claim 6, wherein the electrical current measured by the biosensor matches the preset pattern when the electrical current exceeds a predetermined threshold value. 8. The system of claim 1, wherein the body-mountable electronic device further includes: a wireless transmitter operably coupled with the control circuitry and adapted to transmit information to a communication device,wherein the control circuitry is further adapted to detect the second deployment indicator responsive to establishment of a wireless connection between the wireless transmitter and the communication device prior to expiry of a timeout. 9. The system of claim 1, wherein the control circuitry is further adapted to detect the second deployment indicator responsive to detecting a gesture pattern in motion signals outputted by the motion sensor. 10. The system of claim 9, wherein the gesture pattern comprises a pattern of body mountable electronic device movements, a pattern of rot... | Verily Life Sciences LLC | G06F001/32 G06F001/16 A61B005/145 G06F001/3206 A61B005/00
US10702232 | B2 | 2020-07-07 | Systems and methods for detecting complex networks in MRI image data | 2019-03-28 | Systems and methods for detecting complex networks in MRI image data in accordance with embodiments of the invention are illustrated. One embodiment includes an image processing system, including a processor, a display device connected to the processor, an image capture device connected to the processor, and a memory connected to the processor, the memory containing an image processing application, wherein the image processing application directs the processor to obtain a time-series sequence of image data from the image capture device, identify complex networks within the time-series sequence of image data, and provide the identified complex networks using the display device. | 1. A method for predicting and assigning effective treatment protocols, comprising: obtaining a time-series sequence of image data of a patient's brain using a medical imaging device;generating at least one neurological model, where the at least one neural model characterizes at least one neural circuit in the patient's brain;assigning at least one biotype to the patient based on the neurological model;generating a treatment database, where the treatment database associates biotypes with treatment protocols;retrieving a treatment protocol from the treatment database based on the at least one assigned biotype. 2. The method for predicting and assigning effective treatment protocols of claim 1, wherein treatment protocols in the treatment database comprise drug treatment. 3. The method for predicting and assigning effective treatment protocols of claim 1, wherein treatment protocols in the treatment database comprise behavioral therapies. 4. The method for predicting and assigning effective treatment protocols of claim 1, wherein treatment protocols in the treatment database comprise neuromodulation therapies. 5. The method for predicting and assigning effective treatment protocols of claim 1, wherein the drug database further comprises efficacy metrics describing the efficacy of a treatment protocol with respect to at least one biotype. 6. The method for predicting and assigning effective treatment protocols of claim 5, wherein the retrieved treatment protocol is associated with the efficacy metric signifying the highest likelihood of success of the treatment protocol for the individual based on the assigned at least one biotype. 7. The method for predicting and assigning effective treatment protocols of claim 1, further comprising utilizing a machine learning model to generate efficacy metrics. 8. The method for predicting and assigning effective treatment protocols of claim 1, further comprising providing a plurality of cognitive assessment scores representing the deviation of the individual's cognitive functions from an average cognitive function. 9. The method for predicting and assigning effective treatment protocols of claim 1, further comprising providing a plurality of emotional assessment scores representing the deviation of the individual's cognitive functions from an average cognitive function. 10. The method for predicting and assigning effective treatment protocols of claim 1, wherein the medical imaging system is a magnetic resonance imaging system. 11. A system for predicting and assigning effective treatment protocols, comprising: a medical imaging device;a processor; anda memory comprising an image processing application, where the image processing application directs the processor to: obtain a time-series sequence of image data of a patient's brain using the medical imaging device;generate at least one neurological model, where the at least one neurological model characterizes at least one neural circuit in the patient'... | The Board of Trustees of the Leland Stanford Junior University | A61B006/00 G06K009/00 G06T003/00 G06T007/00 A61B005/00 A61B005/16 A61B005/055 G06K009/34 G06K009/44
US10719866 | B2 | 2020-07-21 | Complementary activity based on availability of functionality | 2017-02-14 | In various example embodiments, a system and method for causing the performance of a complementary activity on a secondary device that is complementary to a device activity performed on a primary user device are presented. In an example embodiment, a device activity, being performed by a primary user device is detected. A secondary user device capable of performing a complementary activity corresponding to the device activity is identified. Instructions are generated for the secondary user device to perform the complementary activity based on the complementary activity including an activity component that utilizes a functionality of the secondary user device not available on the primary user device, wherein the functionality includes capturing data of a particular data type from a sensor. The instructions to perform the complementary activity are transmitted to the secondary user device. | 1. A method comprising: detecting a device activity being performed by a primary user device being used by a user;identifying, using a hardware processor of a machine, a secondary user device capable of performing a complementary activity corresponding to the device activity, the secondary user device comprising a mobile communication device, the identifying of the secondary user device including:determining the secondary user device is capable of capturing data of a particular data type from a sensor that is not available on the primary user device based on a set of functions retrieved from a third-party server based on a type of the secondary user device;determining whether the user is wearing the secondary user device in response to receiving, from the secondary user device, biometric information of the user captured by the secondary user device;generating, in response to determining that the user is wearing the secondary user device, instructions for the secondary user device to perform the complementary activity based on the complementary activity including an activity component that utilizes the particular data type from the sensor that is not available on the primary user device; andtransmitting, to the secondary user device, the instructions to perform the complementary activity. 2. The method of claim 1, further comprising: accessing attribute data associated with a user of the primary user device;inferring user characteristics from the attribute data associated with the user;identifying similar users, from among a plurality of other users, that are similar to the user based on the inferred user characteristics and respective user characteristics of the plurality of other users; andinferring a user preference of the user to perform the complementary activity corresponding to the device activity based on user preferences of the identified similar users indicated by the user characteristics of the identified similar users. 3. The method of claim 1, further comprising: accessing attribute data associated with a user of the primary user device; andinferring a user preference, from the attribute data, of the user to perform the complementary activity corresponding to the device activity. 4. The method of claim 1, wherein the identifying of the secondary user device is based on a device status of the secondary user device indicating a device capability to perform the complementary activity in real-time. 5. The method of claim 1, wherein the biometric information comprises biological information of the user, and wherein the secondary user device comprises a smart watch with a heart rate sensor, wherein the biological information comprises a heart rate of the user, and wherein determining whether the user is wearing the secondary user device comprises determining that the user is not wearing the smart watch when the heart rate sensor indicates no heart rate. 6. The method of claim 1, further comprising: determining, based on a comparison be... | eBay Inc. | H04L029/06 G06Q030/06 G06T011/20 H04L012/26 H04L029/12 H04L029/08 H04W076/14 H04W004/80 A61B005/04 H04W084/18
US10716478 | B2 | 2020-07-21 | Wearable device heart monitor systems | 2015-12-29 | In embodiments of wearable device heart monitor systems, a wearable device has electrical contacts integrated in a housing base of the wearable device as electrodes designed to contact skin of a user while wearing the wearable device. A housing bezel of the wearable device designed as an additional point of contact on the wearable device. The wearable device includes an electromyography (EMG) system to receive electrical signals from at least two of the electrodes and detect muscular movement of the user. Further, the wearable device includes an electrocardiogram (ECG) system to receive and combine the electrical signals from the electrodes when the user touch contacts the housing bezel while wearing the wearable device to complete an ECG loop between the electrodes and the housing bezel for a heart rate reading of the user. | 1. A wearable device, comprising: electrical contacts integrated in a housing base of the wearable device as electrodes configured to contact skin of a user while the user is wearing the wearable device, each of the electrodes being a point of contact on the wearable device;a housing bezel of the wearable device configured as an additional point of contact on the wearable device, the housing bezel grounded to a circuit board via a sequential ground path from the housing bezel to a ground clip and an internal puck housing, and via a ground connecting the internal puck housing to a gold pad on the circuit board;an electromyography (EMG) system configured to receive electrical signals from at least two of the electrodes and detect muscular movement of the user;an electrocardiogram (ECG) system configured to receive the electrical signals from the electrodes in response to the user contacting the housing bezel while wearing the wearable device to complete an ECG loop between the electrodes and the housing bezel for a heart rate reading of the user; anda memory and processing system that implements a heart monitor manager configured to control combining the electrical signals from the electrodes to complete the ECG loop of the ECG system and determine the heart rate reading of the user. 2. The wearable device as recited in claim 1, further comprising a photoplethysmogram (PPG) system that includes a light emitting diode configured to illuminate the skin of the user and an optical sensor configured to detect reflected light from which heart rate measurements of the user are determined. 3. The wearable device as recited in claim 2, wherein the heart monitor manager is configured to: dynamically switch between the EMG system and the ECG system based on detected user activity; andactivate the PPG system in conjunction with, or independently from, the EMG system or the ECG system. 4. The wearable device as recited in claim 1, wherein the heart monitor manager is configured to determine a gesture input to the wearable device based on the muscular movement of the user detected by the EMG system. 5. The wearable device as recited in claim 1, wherein the heart monitor manager is configured to determine the heart rate reading over time based on the combined electrical signals between a first extremity of the user in contact with the electrodes, through a heart of the user, and a second extremity of the user in contact with the housing bezel. 6. The wearable device as recited in claim 5, wherein the heart monitor manager is further configured to determine an electrocardiogram reading of the user. 7. The wearable device as recited in claim 1, wherein the electrodes are equally spaced and arranged in a pattern configured to facilitate any combination of the electrodes being utilized as the at least two electrodes by the EMG system. 8. The wearable device as recited in claim 1, further comprising: an electronic switch integrated in the circuit board and cont... | Motorola Mobility LLC | A61B005/0205 A61B005/024 G06F003/01 A61B005/0245 A61B005/026 A61B005/0488 A61B005/053 A61B005/00 A61B005/04 A61B005/0404 A61B005/0408 A61B005/0428 A61B005/0492
US10719945 | B2 | 2020-07-21 | System and method for face position tracking and alerting user | 2018-07-09 | Current technologies detect the alertness of a user using frontal face. However, there are no techniques to track the position of face from a side angle. A method and system for face position tracking of a user and alerting the user is disclosed. The method includes capturing images using a monocular camera and identifying a user's face from the images and storing the user's face as reference face. The user's face is divided into two regions and one or more corner points are identified. A centroid is created in the two regions and the corner points are joined using virtual lines to create a dual flexible spider model. One or more new corner points are created on fulfillment of pre-defined conditions. Angle and magnitude of flow vectors is determined using one or more new corner points and alerts are given to the user based on the angle and magnitude. | 1. A method for face position tracking and alerting a user, the method comprising: capturing one or more images through a camera;identifying the user's face from one or more faces based on space occupied in the one or more images to store the user's face as a reference face;dividing the user's face into (i) a first region and (ii) a second region and identifying one or more corner points from the first region and the second region of the user's face;creating a centroid in each of the first region and the second region and joining the one or more corner points with the centroid created in each of the first region and the second region using a plurality of virtual lines to create a dual flexible spider model;creating a dense plurality of virtual lines in the dual flexible spider model, based on movement of the user's face for a predefined time period;tracking variations in the dual flexible spider model due to the movement of the user's face by computing one or more flow vectors based on a magnitude and a phase value of the one or more corner points;creating one or more new corner points and creating one or more additional flow vectors by subjecting the one or more new corner points to optical flow frame processing, wherein the one or more new corner points are created based on one or more predefined conditions;determining one or more valid motion vectors, one or more invalid motion vectors and one or more static motion vectors based on an angle and magnitude of the one or more flow vectors; andalerting the user if the angle and magnitude of the one or more flow vectors of the one or more valid motion vectors is greater than a threshold. 2. The method of claim 1, wherein if the user's face is attentive for the predefined time period, the magnitude and the phase value of the one or more corner points is reset to avoid a false alert due to accumulation of noisy movements. 3. The method of claim 1, wherein the one or more valid motion vectors are the one or more flow vectors if (i) the magnitude of the one or more flow vectors is within a predefined magnitude or (ii) the plurality of virtual lines are moving in a same direction or exhibit homogeneous scaling or there is co-presence of vectors. 4. The method of claim 1, wherein the one or more invalid motion vectors are the one or more flow vectors if (i) the magnitude of the one or more flow vectors is greater than a predefined magnitude or (ii) the plurality of virtual lines are not moving in a same direction or do not exhibit homogeneous scaling or there is no co-presence of vectors. 5. The method of claim 1, wherein the one or more static motion vectors are the one or more flow vectors if the magnitude of the one or more flow vectors is constant. 6. The method of claim 1, wherein identifying the user's face based on the space occupied in the one or more images is performed using a haar cascade classifier algorithm. 7. The method of claim 1, wherein creating the dense plurality of virtual line... | Tata Consultancy Services Limited | G06K009/00 G06T007/246 G08B021/06 G06T007/149 B60K028/06 G06K009/46 G06K009/62 G06T007/73 G06T007/181 B60K028/02 G06T007/162 A61B005/18 G06F003/01
US10716482 | B2 | 2020-07-21 | Thermodilution catheter systems and methods for determining blood flow rates | 2014-04-24 | Catheter systems and methods for determining blood flow rates based on temperature measurements by thermodilution. The catheter may include a fluid lumen defined between inner and outer tubular members for delivering an indicator fluid, and a guidewire lumen defined by the inner tubular member. The catheter may include fluid infusion openings at the distal end region of the outer tubular member configured to permit the indicator fluid to exit the catheter from the fluid lumen, and a fluid hole located at the distal end region of the inner tubular member configured to permit the indicator fluid to pass from the fluid lumen into the guidewire lumen. A temperature sensor positioned on a guidewire may be positioned within the guidewire lumen to measure the temperature of the indicator fluid entering the guidewire lumen through the fluid hole. A blood flow rate may be calculated based on the measured temperature. | 1. A catheter system for determining blood flow in a body lumen, the system comprising: a catheter including: an outer tubular member;an inner tubular member disposed within the outer tubular member;wherein an inner wall surface of the outer tubular member is attached to an outer wall surface of the inner tubular member at a position adjacent to a distal end region of the catheter;a first, fluid lumen, at least a portion thereof defined between the inner tubular member and the outer tubular member;a second lumen defined by the inner tubular member, the second lumen having an open distal end such that the catheter is configured to be advanced over another medical device;one or more fluid infusion openings located at a distal end region of the catheter, the one or more fluid infusion openings configured to permit fluid to exit the catheter from the fluid lumen;one or more fluid holes located at the distal end region of the catheter, the one or more fluid holes configured to permit fluid to pass from the fluid lumen into the second lumen;an elongate member advanceable through the second lumen of the catheter, the elongate member having a temperature sensor disposed on a distal end portion thereof;wherein the temperature sensor is designed to shift between a first position where the temperature sensor is disposed within the second lumen of the catheter and is aligned with at least one of the one or more fluid holes in order to measure a temperature of fluid entering the second lumen of the catheter through the one or more fluid holes and a second position where the temperature sensor is disposed distal of the distal end of the catheter. 2. The catheter system of claim 1, wherein the one or more fluid infusion openings extend through a wall of the outer tubular member from an inner surface of the outer tubular member to an outer surface of the outer tubular member. 3. The catheter system of claim 2, wherein the one or more fluid infusion openings are configured to generate a jet of fluid exiting the catheter. 4. The catheter system of claim 1, wherein the one or more fluid holes extend through a wall of the inner tubular member from an outer surface of the inner tubular member to an inner surface of the inner tubular member. 5. The catheter system of claim 1, wherein the one or more fluid holes are one or more weeping holes configured to allow fluid to weep into the second lumen. 6. The catheter system of claim 1, wherein the one or more fluid infusion openings include four fluid infusion openings equidistantly spaced circumferentially around the outer tubular member. 7. The catheter system of claim 1, wherein the one or more fluid holes is a single weeping hole extending through the inner tubular member. 8. The catheter system of claim 1, wherein the catheter is a single-operator-exchange catheter in which the second lumen extends only through a distal portion of the catheter. 9. The catheter system of claim 1, wherein the second lumen extend... | HEXACATH | A61B005/026 A61B005/028
US10709344 | B2 | 2020-07-14 | Measurement apparatus | 2015-05-26 | A measurement apparatus includes a housing, a display disposed on the front face of the housing, a contact interface, to be contacted by a test site, disposed on a different surface than the front face of the housing, a laser light source that emits laser light from the contact interface, a light receiver that receives scattered laser light from the test site, a biological information generator that generates biological information based on output from the light receiver, and a controller that controls emission of laser light in the laser light source. | 1. A measurement apparatus comprising: a housing;a display disposed on a front face of the housing;an imaging device disposed on the front face of the housing;a contact interface, to be contacted by a test site, disposed on a different surface than the front face of the housing;a laser light source configured to emit laser light from the contact interface;a light receiver configured to receive scattered laser light from the test site; anda controller including a processor configured to: generate biological information based on output from the light receiver; andcontrol emission of laser light by the laser light source, whereinthe controller permits emission of the laser light from the laser light source when a face of a subject is included in an image captured by the imaging device and prohibits emission of the laser light from the laser light source when the face of the subject is not included in the image. 2. The measurement apparatus of claim 1, further comprising a guide indicating a position of the contact interface. 3. The measurement apparatus of claim 1, wherein when causing laser light to be emitted from the laser light source, the controller causes the display to display an image. 4. The measurement apparatus of claim 1, further comprising: a three-axis acceleration sensor configured to detect an inclination of the housing with reference to a direction of gravitational force, whereinthe controller judges whether an emission direction of the laser light is a predetermined angle or greater from a vertically upward direction based on the inclination detected by the three-axis acceleration sensor, permits emission of the laser light from the laser light source when the emission direction is the predetermined angle or greater, and prohibits emission of the laser light from the laser light source when the emission direction is less than the predetermined angle. 5. The measurement apparatus of claim 4, wherein the predetermined angle is 45°. 6. The measurement apparatus of claim 1, wherein the biological information includes information related to blood flow. | KYOCERA Corporation | A61B005/02 A61B005/026 A61B005/00 A61B005/0285 A61B005/145
US10722059 | B2 | 2020-07-28 | System for monitoring the liquid intake of a user and method of operating the system | 2019-02-19 | A system for monitoring liquid intake of a user includes a monitor configured to determine an amount of liquid removed from a container, an interface configured to receive a vital parameter of the user of the system from a vital parameter sensor, and a prompter configured to prompt the user to drink as a function of the amount of liquid removed and of the vital parameter. | 1. System for monitoring liquid intake of a user, comprising: a container,a vital parameter sensor,a monitor configured to determine an amount of liquid removed from the container;an interface configured to receive, from the vital parameter sensor, a vital parameter of the user of the system; anda prompter configured to prompt the user to drink as a function of the amount of liquid removed and of the vital parameter,wherein the monitor, the interface and the prompter are arranged within a socket configured to receive the container,wherein the container is formed of a transparent material, wherein the prompter is configured to couple light into the container for prompting, and wherein the container comprises scattering centers configured to scatter the light coupled in, andwherein the container comprises a projection formed in the container bottom, and wherein the prompter is configured to couple the light into the container at the projection. 2. System as claimed in claim 1, wherein the light is coupled into the projection perpendicularly to a viewing direction of a user looking into the container. 3. System as claimed in claim 1, wherein the monitor comprises a sensor configured to determine a current filling level of the container, the monitor being configured to determine the amount of liquid removed from the container by means of a time sequence of current filling levels of the container which have been determined. 4. System as claimed in claim 1, wherein the monitor is configured to determine the amount of liquid removed from the container by means of capacitive measurement, by means of optical measurement, by means of ultrasound, by means of radar, by means of a change in weight, by means of a run-time measurement, and/or by means of a conductivity measurement. 5. System as claimed in claim 1, wherein the prompter is configured to prompt the container, the vital parameter sensor or a mobile device to prompt the user to drink in a visual, auditive or tactile manner. 6. System as claimed in claim 1, wherein the prompter is arranged inside a mobile device. 7. System as claimed in claim 1, wherein the socket comprises a fastening mechanism configured to firmly connect the container to the socket in a mechanical or magnetic manner in an operating state. 8. System as claimed in claim 1, wherein the socket comprises a detection unit configured to identify the container in an operating state and to distinguish it from further containers each exhibiting a specific filling level, the distinction by the detection unit enabling determining of a container-specific amount of liquid removed and associating the container with the respective user. 9. System as claimed in claim 1, wherein the socket is configured to sense and set a temperature of the liquid comprised by the container. 10. System as claimed in claim 1, wherein the socket comprises a casing configured to protect the socket from external influences or to form an anti-slip bearing surfa... | BELENUS Verwaltungsgesellschaft mBH | A47G023/16 A61B005/00 A61B005/024 A47G019/22 A47G019/00 A47G023/00 A61B005/021
US10726295 | B2 | 2020-07-28 | Control method and recording medium | 2016-09-21 | A control method includes: receiving, from a case search system, a plurality of data including a plurality of images corresponding to a plurality of similar medical images having a certain similarity with a target medical image to be interpreted; displaying on a display a display screen including a first display area that displays thumbnails of the plurality of similar medical images; sensing one similar medical image selected from among the plurality of similar medical images displayed as thumbnails in the first display area; if the one selected similar medical image is a diffuse lesion, displaying the other plurality of medical images in a second display area included on the display screen; and if the one selected similar medical image is a localized lesion, successively displaying, in the second display area and in a first direction, the plurality of medical images including the localized lesion. | 1. A control method for an information terminal, the control method being executed by a computer of the information terminal, and comprising: receiving medical case data items for medical cases, the medical case data items including thumbnail images corresponding to medical images for the medical cases, the medical images being determined on the basis of a target medical image that is interpreted;displaying a display screen including the thumbnail images on a display;detecting a selected thumbnail image from among the thumbnail images; anddisplaying a first thumbnail image, a second thumbnail image, and a third thumbnail image in a first way or a second way, whereinthe medical case data items, the medical cases, the thumbnail images, and the medical images are in one-to-one relationship,the thumbnail images include the first thumbnail image, the second thumbnail image, and the third thumbnail image,each of the medical case data items includes first information indicating a corresponding medical image, second information indicating whether the corresponding medical case is diffuse or localized, and third information indicating first medical images for the medical case,the corresponding medical image and the first medical images are obtained through one medical examination performed on a subject,the corresponding medical image and the first medical images include a first image in a first tomographic plane of the subject, a second image in a second tomographic plane of the subject, and a third image in a third tomographic plane of the subject,the first tomographic plane is closest to a top of a head of the subject among the first tomographic plane, the second tomographic plane, and the third tomographic plane,the third tomographic plane is farthest away from the top among the first tomographic plane, the second tomographic plane, and the third tomographic plane,the first thumbnail image corresponds to the first image, the second thumbnail image corresponds to the second image, and the third thumbnail image corresponds to the third image,if the second information corresponding to the selected thumbnail image indicates that the medical case corresponding to the selected thumbnail image is diffuse, the first thumbnail image, the second thumbnail image, and the third thumbnail image are displayed in the first way,if the second information corresponding to the selected thumbnail image indicates that the medical case corresponding to the selected thumbnail image is localized, the first thumbnail image, the second thumbnail image, and the third thumbnail image are displayed in the second way,the first way is that the first thumbnail image, the second thumbnail image, and the third thumbnail image are displayed at a same time on the display, andthe second way is that (i) the second thumbnail image is displayed on the display after the first thumbnail image is displayed on the display and the third thumbnail image is displayed on the display after the second... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | G06F016/30 G06K009/62 G16H040/63 G16H050/70 G16H030/40 A61B005/055 A61B006/00 G06F019/00 G06F003/0482
US10702031 | B2 | 2020-07-07 | Panel bracelet with replaceable inserted panels | 2016-12-29 | A panel bracelet includes a plurality of panels configured to be selectively replaceable within the bracelet, and a pair of panel holding frames, wherein each open panel holding frame is configured to hold a panel therein. Each panel holding frame includes a retaining bead that is received within a perimeter groove of the panel which is held within the panel holding frame to receive and retain the panel therein. | 1. A reversible panel bracelet comprising: a reversible body;a plurality of two sided panels configured to be selectively replaceable within the body of the bracelet and with each side of the panel including an ornamental design;at least two open panel holding frames forming at least part of the reversible body of the bracelet, wherein each open panel holding frame includes a central opening through the reversible body of the bracelet and is designed to hold a two sided panel therein; andwherein each panel holding frame is one of i) generally “C” shaped in cross section to receive and retain one of the two sided panels therein and which are placed within the open panel holding frame by engaging with two opposing faces of the panel, or ii) includes a retaining bead that is received within a matching groove in a perimeter of a panel held within the open panel holding frame to receive and retain one of the two sided panels therein. 2. The reversible panel bracelet with inserted two sided panels according to claim 1 wherein exactly two substantially rectangular panel holding frames are provided in the reversible body of the panel bracelet. 3. The reversible panel bracelet with inserted two sided panels according to claim 1 wherein the open panel holding frames are formed of elastic material and wherein the reversible body of the bracelet is configured to be flipped in its entirety. 4. The reversible panel bracelet with inserted two sided panels according to claim 1 wherein at least one of the panels is formed of two separate facing halves coupled together. 5. The reversible panel bracelet with inserted two sided panels according to claim 1 wherein the panels include at least one of a functioning watch, a key-holding unit, an electronic device, and a sports trading card. 6. A panel bracelet comprising: a body having a width and a circumference;a plurality of panels configured to be selectively replaceable within the body of the bracelet, wherein each panel includes parallel ends configured to be perpendicular to the width of the body of the bracelet and longer parallel sides configured to extend along the circumference of the body of the bracelet, wherein each panel includes a perimeter groove around the perimeter of the panel;a pair of panel holding frames forming at least part of the body of the bracelet, wherein each open panel holding frame is configured to hold a panel therein, and wherein each panel holding frame includes a retaining bead that is received within the perimeter groove of a panel held within the panel holding frame to receive and retain the panel therein, and wherein each panel holding frame extends of over 130 degrees along the circumference of the body of the bracelet. 7. A panel bracelet comprising: a body having a width and a circumference;a plurality of panels configured to be selectively replaceable within the body of the bracelet, wherein each panel includes parallel ends configured to be perpendicular to the width of ... | - | A44C005/00 A44C017/02 A44C015/00 A61B005/00 F23Q002/00 G04G017/08 G04G021/02
US10709492 | B2 | 2020-07-14 | Effective parasitic capacitance minimization for micro ablation electrode | 2015-09-10 | A flexible catheter has an ablation electrode disposed in its distal segment. The ablation electrode a cavity formed in its external surface, a microelectrode configured to fit into the cavity, a conductive wire lead connecting the microelectrode to receiving circuitry, and an electrical shield surrounding the wire lead. A power generator is connected to the ablation electrode and the electrical shield in a generator circuit. A back patch electrode adapted to contact with the subject is connected in the generator circuit. The microelectrodes can be active while energizing the ablation electrode. | 1. An apparatus, comprising: a flexible catheter adapted for insertion into a heart of a living subject and having a distal segment;an ablation electrode disposed at the distal segment to be brought into contact with a target tissue in the heart; the ablation electrode having an external surface, and a cavity formed in the external surface;a microelectrode configured to fit into the cavity;a conductive wire lead connecting the microelectrode to receiving circuitry; andan electrical shield surrounding the conductive wire lead, the electric shield being electrically connected directly to an ablation electrode lead wire via a wire and configured to equalize potentials between the ablation electrode and the microelectrode to minimize capacitance leakage current flow. 2. The apparatus according to claim 1, further comprising: a generator circuit connecting a power generator to the ablation electrode and the electrical shield; anda back patch electrode adapted to contact with the living subject and connected to the generator circuit. 3. The apparatus according to claim 1, wherein the electrical shield comprises: a coaxial layer; anda dielectric layer disposed between the coaxial layer and the conductive wire lead. 4. The apparatus according to claim 3, wherein the electrical shield further comprises an insulating jacket that overlies the coaxial layer. 5. The apparatus according to claim 1, wherein the external surface of the ablation electrode has a curvature, and the microelectrode is contoured, located and oriented to conform to the curvature. 6. The apparatus according to claim 1, wherein the ablation electrode has a cylindrical portion, wherein the cavity comprises a plurality of cavities formed in the cylindrical portion. 7. The apparatus according to claim 1, wherein the ablation electrode has a distal annular portion, wherein the cavity comprises a plurality of cavities formed in the annular portion, and wherein the microelectrode comprises a plurality of microelectrodes, each one of the plurality of microelectrodes disposed in a respective one of the plurality of cavities. 8. The apparatus according to claim 1, wherein the microelectrode is linked to a thermocouple, wherein the thermocouple provides a signal representative of a temperature of the microelectrode. | Biosense Webster (Israel) Ltd. | A61B018/14 A61B018/12 A61B018/00 A61B005/0428 A61B018/16
US10724971 | B2 | 2020-07-28 | Display device and x-ray CT device | 2015-11-05 | A plurality of CT imaging conditions are created beforehand as presets, each preset is created beforehand as an icon, and each preset is associated with an icon and registered. Each icon is disposed in a CT imaging condition setting screen according to a spatial resolution and density resolution that serve as two-dimensional indexes relating to image quality. The selection of a given icon by a user results in the selection of the CT imaging conditions corresponding to the selected icon. The indexes relate to image quality and are two-dimensional, and this facilitates operation and makes it possible to intuitively select optimal CT imaging conditions with a small number of steps. | 1. A display device having a computer tomographic (CT) software function used in an X-ray CT device, comprising: control elements configured to register presets for a plurality of CT imaging conditions and arranged on a display screen using a two-dimensional index relating to image quality,wherein, as one of the control elements is selected, a CT imaging condition corresponding to the selected control element is set,wherein the control elements are sequentially arranged along a first axis according to a first index relating to image quality, and are sequentially arranged along a second axis according to a second index relating to image quality. 2. The display device according to claim 1, wherein the index includes a spatial resolution and a density resolution. 3. The display device according to claim 1, wherein, when the CT imaging condition is selected, a preview image is displayed as an image view serving as a reference of an output image corresponding to the selected condition. 4. The display device according to claim 1, wherein the control elements are icons arranged side by side across horizontal and vertical directions for each index on the display screen. 5. The display device according to claim 1, wherein the control elements are displayed independently for each index. 6. The display device according to claim 1, further comprising imaging condition display means configured to display the set CT imaging condition. 7. The display device according to claim 1, further comprising imaging condition input confirmation means configured to perform an input operation for confirming the set CT imaging condition, wherein CT imaging is performed on the basis of the confirmed CT imaging condition in response to the input operation to the imaging condition input confirmation means. 8. The display device according to claim 1, further comprising a control element that enables selection of a material on the display screen, wherein, when the control element is selected, a CT imaging condition is set for each material corresponding to the selected control element. 9. The display device according to claim 8, wherein the control element that enables selection of the material includes a dropdown list, a combobox, a radio button, a checkbox, or a tap. 10. The display device according to claim 1, wherein a preset of the registered CT imaging condition is configured editably. 11. An X-ray CT device comprising: an X-ray source that generates an X-ray for irradiating X-rays onto a target object;a stage for placing the target object;an X-ray detector that detects the X-ray irradiated from the X-ray source to the target object and transmitted through the target object;a display device having a CT software function including a function of setting a CT imaging condition; anda reconstruction processing unit that reconstructs a CT image of the target object from the data detected by the X-ray detector,wherein the display device comprises control elements configu... | Shimadzu Corporation | G01N023/04 A61B006/03 G06T011/60 G06T011/00 G01N023/046 A61B005/00 A61B006/00
US10718601 | B2 | 2020-07-21 | Optical coherence tomography device | 2018-02-20 | An optical coherence tomography device provided with: an OCT optical system configured to output OCT signal; and an analysis processing unit configured to process the OCT signal and generate motion contrast data of the specimen. The analysis processing unit is provided with: a first image data generation unit configured to process multiple OCT signals to generate first image data representing phase difference information of the multiple OCT signals, and a second image data generation unit configured to process the OCT signal to generate second image data representing amplitude information of the OCT signal. The analysis processing unit generates the motion contrast data based on the phase difference information in the first image data and the amplitude information in the second image data. | 1. An optical coherence tomography device comprising: an OCT optical system configured to detect measurement light irradiated onto a specimen and a reference light and output OCT signal based on the measurement light and the reference light; andan analysis processing unit configured to process the OCT signal and generate motion contrast data of the specimen,wherein the analysis processing unit comprises:a first image data generation unit configured to generate first motion contrast image data by applying first analysis processing on the OCT signal; anda second image data generation unit configured to generate second motion contrast image data by applying second analysis processing that is different from the first analysis processing on the OCT signal, andwherein the analysis processing unit applies the second motion contrast image data as a filter onto the first motion contrast image data to remove an artifact at a high reflection portion. 2. The optical coherence tomography device according to claim 1, wherein the analysis processing unit applies to the first motion contrast image data a weight based on the second motion contrast image data. 3. The optical coherence tomography device according to claim 1, wherein the high reflection portion is a nerve fiber layer (NFL). 4. The optical coherence tomography device according to claim 1, wherein the first image data generation unit generates a Doppler OCT image as the first motion contrast image data. 5. The optical coherence tomography device according to claim 1, wherein the second image data generation unit processes the multiple OCT signals which is a plurality of OCT signals output from the OCT optical system being obtained at a different timing for the same position on the specimen and generates, as the second motion contrast image data, image data representing difference between a first vector based on phase information and the amplitude information of a first OCT signal selected from the multiple OCT signals and a second vector based on phase information and the amplitude information of a second OCT signal selected from the multiple OCT signals. 6. The optical coherence tomography device according to claim 1, wherein the analysis processing unit applies, to one of the first motion contrast image data and the second motion contrast image data a filter based on another one of the first motion contrast image data and the second motion contrast image data. 7. The optical coherence tomography device according to claim 1, wherein the analysis processing unit applies, to one of the first motion contrast image data and the second motion contrast image data a mask based on another one of the first motion contrast image data and the second motion contrast image data. 8. The optical coherence tomography device according to claim 1, wherein the analysis processing unit obtains luminance with respect to each pixel of the motion contrast data by computing the phase difference information in the fir... | NIDEK CO., LTD. | G01B009/02 A61B005/00 A61B003/10 A61B003/00
US10713790 | B2 | 2020-07-14 | Methods of cardiac mapping and directional guidance | 2018-07-31 | Various embodiments provide a cardiac mapping method including: recording premature ventricular contraction (PVC) electrocardiogram (ECG) data during PVC of a heart; generating a PVC activation map of the heart, based on the PVC ECG data and a three-dimensional (3D) heart model generated based on two-dimensional (2D) images of the heart, the PVC activation map including an area of earliest activation; generating a reference image by modifying on one of the 2D images to identify the area of earliest activation; displaying both the PVC activation map and the reference image; and pacing the heart at a first pacing location disposed in the area of earliest activation. | 1. A cardiac mapping method comprising: recording premature ventricular contraction (PVC) electrocardiogram (ECG) data during PVC of a heart;generating a PVC activation map of the heart, based on the PVC ECG data and a three-dimensional (3D) heart model generated based on two-dimensional (2D) images of the heart, the PVC activation map including an area of earliest activation;generating a reference image by modifying on one of the 2D images to identify the area of earliest activation;displaying both the PVC activation map and the reference image;pacing the heart at a first pacing location disposed in the area of earliest activation;determining whether the first pacing location is an ablation location that may be ablated to prevent the PVC by comparing ECG data recorded during the pacing of the first pacing location to the PVC ECG data;adding guidance information related to a second pacing location to the reference image in response to determining that the first pacing location is not an ablation location;moving the pacing electrode to the second pacing location based on the guidance information; andpacing the heart at the second pacing location. 2. The method of claim 1, further comprising displaying a real-time image of the heart, wherein the generating a reference image comprises selecting one of the 2D heart images that most closely matches the real-time image. 3. The method of claim 1, wherein generating a reference image comprises adding a suggested pacing location to the reference image. 4. The method of claim 3, wherein the suggested pacing location is a suggested ablation location. 5. The method of claim 1, wherein the 2D images comprise magnetic resonance (MRI) images or computed tomography (CT) images of the heart. 6. The method of claim 1, wherein generating the reference image comprises identifying a location of a pacing catheter on the reference image. 7. The method of claim 1, wherein the area of earliest activation comprises an onset location of the PVC. 8. The method of claim 1, wherein the guidance information comprises a vector image pointing from the first pacing location toward the second pacing location. 9. The method of claim 1, wherein the guidance information comprises: a first point representing the first pacing location; anda vector image pointing from the first point toward the second pacing location. 10. The method of claim 1, wherein the guidance information comprises: a first point representing the first pacing location; anda second point representing the second pacing location. 11. The method of claim 1, further comprising determining whether the second pacing location is an ablation location based on a comparison of second pacing ECG data recorded during the pacing of the second pacing location and the PVC ECG data. 12. The method of claim 11, further comprising ablating the heart at the second pacing location if the second pacing location is determined to be a suitable ablation location. 13. A cardiac ... | CATHETER PRECISION, INC. | G06K009/00 G06T007/00 G06T007/73 G06T017/10 A61B005/0402 A61B005/00 G06T017/00
US10716489 | B2 | 2020-07-21 | Determining biological tissue type | 2015-10-12 | A method is described for determining biological tissue type based on a complex impedance spectra obtained from a probe with a conducting part adjacent a tissue region of interest, wherein the impedance spectra includes data from a number of frequencies. The method may include: obtaining, from the complex impedance spectra, a first data set representative of impedance modulus values, or equivalent admittance values, at one or more frequencies, obtaining, from the complex impedance spectra, a second data set representative of impedance phase angle values, or equivalent admittance values, at one or more different frequencies, applying a first discrimination criterion to the first data set, applying a second discrimination criterion to the second data set, and thereby determining if the tissue region of interest is a tissue type characterised by the discrimination criteria. | 1. A method of determining if a biological tissue region of interest includes intraneural tissue based on a complex impedance spectra that has been obtained from a probe with a conducting part adjacent the tissue region of interest, wherein the impedance spectra includes data from a plurality of frequencies, the method comprising: obtaining, from the complex impedance spectra, a first data set representative of impedance modulus values, or equivalent admittance values, at one or more frequencies;obtaining, from the complex impedance spectra, a second data set representative of impedance phase angle values, or equivalent admittance values, at one or more different frequencies;applying a first discrimination criterion to the first data set to determine whether or not the impedance modulus is below a value known to exclude subcutaneous fat tissue;applying a second discrimination criterion to the second data set to determine if the phase angle at a set frequency is above a threshold level that is known to exclude at least one of muscle tissue and perineural tissue;applying a third discrimination criterion to the second data set to determine if a gradient of the phase angle at a set frequency is larger than 3/126 degrees per kHz at a set frequency that is above 10 kHz and below 400 kHz; andthereby determining if the tissue region of interest includes intraneural tissue. 2. A method as claimed in claim 1, comprising obtaining and comparing impedance spectra for known tissue types in order to identify frequencies at which data parameters including two or more of impedance modulus, phase angle, and phase angle gradient have the greatest difference for different tissue types, and selecting discrimination criteria at these frequencies to thereby provide the impedance modulus value for the first discrimination criterion, the threshold phase angle for the second discrimination criterion, and the threshold gradient for the third discrimination criterion. 3. A method as claimed in claim 2, comprising using a compound variable based on the three discrimination criteria to determine if the tissue region of interest includes intraneural tissue. 4. A method as claimed in claim 1, wherein the first discrimination criterion requires an impedance of less than 2500 Ohms at a frequency in the range of 10 kHz to 50 kHz. 5. A method as claimed in claim 1, wherein the first discrimination criterion requires an impedance of less than 2200 Ohms at a frequency of about 39 kHz. 6. A method as claimed in claim 1, wherein the second discrimination criterion requires a phase angle of less than 25 degrees at a frequency selected from the range 10 kHz to 50 KHz. 7. A method as claimed in claim 6, wherein the second discrimination criterion requires a phase angle of less than 25 degrees at about 39 kHz. 8. A method as claimed in claim 1, wherein the third discrimination criterion requires an increase in phase angle of larger than 3 degrees for the phase angle at 251 kHz comp... | Oslo Universitetssykehus HF | A61B005/053 A61B005/00 A61B005/06
US10725550 | B2 | 2020-07-28 | Methods and apparatus for recognition of a plurality of gestures using roll pitch yaw data | 2015-03-03 | Described are apparatus and methods for reconstructing a plurality of gestures using roll pitch and yaw data, typically in combination with data for recognition of start and/or stop portions of the gesture using an auxiliary sensor, such as a capacitive touch sensor or a MEMS sensor. | 1. An apparatus capable of interacting with at least one controllable device based upon a pose of at least a portion of a human body, the apparatus comprising: one or more sensors that are sized for wearing on the human body, each of the one or more sensors emitting sensor data; anda detection unit that operates upon the sensor data to determine the pose of at least the portion of the human body and is capable of interacting with the at least one controllable device, the detection unit including: a memory that stores at least one or more characteristics of human anatomy that are associated with the human body using at least a partial skeletal rendering of a human; anda detection processor, automatically operating under software control, that inputs, aggregates and fuses the sensor data from each of the one or more sensors using the at least one or more characteristics of human anatomy stored in the memory to determine the pose of at least the portion of the human body based upon a locality of said one or more sensors, and wherein the pose is determined using at least a comparison of which one of an absolute value of a change in roll, a change in pitch and a change in yaw is greatest when compared to each other during a period of time. 2. The apparatus according to claim 1 further including an auxiliary sensor sized for wearing on the human body that receives a first specific input based on one of a tactile switch and capacitive touch input and generates a gesture start signal; and wherein the period of time is determined by when the detection processor begins to input, aggregate and fuse the sensor data upon receipt of the gesture start signal and ceases to input the sensor data upon receipt of a gesture stop signal; andwherein the auxiliary sensor receives a second specific input based on one of the tactile switch and the capacitive touch input and generates the gesture stop signal. 3. The apparatus according to claim 2 wherein the detection processor determines that the roll change is greatest, and that the roll change is greater than 0, thereby determining a clockwise gesture signal. 4. The apparatus according to claim 2 wherein the detection processor determines that the roll change is greatest, and that the roll change is less than 0, thereby determining a counter clockwise gesture signal. 5. The apparatus according to claim 2 wherein the detection processor determines that the yaw change is greatest, and that the yaw change is greater than 0, thereby determining a right gesture signal. 6. The apparatus according to claim 2 wherein the detection processor determines that the yaw change is greatest, and that the yaw change is less than 0, thereby determining a left gesture signal. 7. The apparatus according to claim 2 wherein the detection processor determines that the pitch change is greatest, and that the pitch change is greater than 0, thereby determining an up gesture signal. 8. The apparatus according to claim 2 wherein the detect... | NOD, INC. | G06F003/01 G06F003/044 G06K009/00 G06F001/3234 A61B005/11
US10702170 | B2 | 2020-07-07 | Apparatus and method for intravascular measurements | 2014-07-01 | Intravascular diagnosis apparatus and methods are disclosed. In one aspect of the disclosed technology, a intravascular diagnosis apparatus includes a monitoring guidewire and a display unit. The monitoring guidewire includes a core wire and a sensor disposed in a distal region of the core wire. The display unit includes a processor and a display screen, and is capable of receiving communication from the monitoring guidewire. The display unit is configured to perform computations using the processor based on communications received from the monitoring guidewire and is configured to display information on the display screen based on the computations. The display unit can be configured to be disposed after a predetermined number of uses or after a predetermined duration of use. | 1. A portable apparatus for intravascular diagnosis, the portable apparatus comprising: a monitoring guidewire comprising a core wire and a sensor disposed in a distal region of the core wire;a display unit configured to be inoperable after a duration of use, the display unit comprising a processor, a main housing, and a display screen, wherein the display unit is capable of receiving communication from the monitoring guidewire, is configured to perform computations using the processor based on communications received from the monitoring guidewire, and is configured to display information on the display screen based on the computations, the display unit further comprising one or more non-rechargeable batteries configured to power the display unit; anda connector connected to the display unit and configured to be attachable to and detachable from the monitoring guidewire, the connector configured to establish a communicative connection with the monitoring guidewire when the monitoring guidewire is attached to the connector,wherein the display screen is folded in the main housing in a folded position before use,wherein pivoting the display screen from the folded position to an open position acts as an ON switch that enables the one or more non-rechargeable batteries to power the display unit, wherein the display unit being configured to be inoperable after the duration of use includes, after the display screen is turned on, the display unit operating continuously until the one or more non-rechargeable batteries are depleted, andwherein the display unit has no capability of being turned off after the display screen is turned on. 2. The portable apparatus of claim 1, when the monitoring guidewire is not electrically or mechanically connected to the connector the display unit and the monitoring guidewire communicate wirelessly. 3. The portable apparatus of claim 1, further comprising a torquer configured to engage the monitoring guidewire to control the monitoring guidewire within a vasculature of a patient. 4. The portable apparatus of claim 1, wherein the monitoring guidewire further comprises a housing surrounding the sensor, the housing being laser etched to provide flexibility for the housing. 5. The portable apparatus of claim 1, wherein the monitoring guidewire further comprises a flexible coil surrounding the sensor, the coil having a portion over the sensor that creates a window where the sensor is located. 6. The portable apparatus of claim 1, wherein the sensor is a pressure sensor and communication from the monitoring guidewire includes measurements from the pressure sensor, and wherein the computations using the processor include computing fractional flow reserve based on pressure measurements from only the pressure sensor disposed in the distal region of the core wire. 7. The portable apparatus of claim 6, wherein the fraction flow reserve is computed based on: recorded pressure measurements proximal to a first stenosis, anddistal ... | ZURICH MEDICAL CORPORATION | A61B005/0215 A61B005/00 A61B005/02
US10726344 | B2 | 2020-07-28 | Diagnosis support apparatus and method of controlling the same | 2017-07-11 | A diagnosis support apparatus which supports diagnosis based on information associated with a diagnosis name in advance is provided. In the diagnosis support apparatus, an acquisition unit acquires the diagnosis name set by a user. A providing unit provides negative information for the diagnosis name set by the user based on the information. With the above arrangement, the diagnosis support apparatus selects and presents information influencing the diagnosis name expected by the user, thereby efficiently presenting information required by the user. | 1. A diagnosis support apparatus which supports diagnosis comprising: at least one processor and a memory operatively coupled to function as: an acquisition unit configured to acquire information of interpretation findings obtained by medical image interpretation, and acquire a diagnosis name set by a user; anda presentation unit configured to present, based on at least the interpretation findings, information which is positive for the diagnosis name set by the user and negative for an inference result different from the diagnosis name set by the user and information which is negative for the diagnosis name set by the user and positive for the inference result different from the diagnosis name set by the user. 2. The apparatus according to claim 1, wherein the presentation unit is configured to present from information, of the information configured to be input, which has been used for the inference. 3. The apparatus according to claim 1, wherein the presentation unit is configured to present from information, of the information configured to be input, which has not been used for the inference. 4. The apparatus according to claim 1, wherein the presentation unit is configured to display on a display unit the information positive for the diagnosis name presented by the presentation unit and the information negative for the diagnosis name on a display unit in different display forms. | Canon Kabushiki Kaisha | G06N005/04 G06Q050/22 A61B005/00 A61B006/00 G16H050/20 A61B006/03 A61B005/055
US10712919 | B2 | 2020-07-14 | Method for providing physiological state information and electronic device for supporting the same | 2016-08-10 | An apparatus and method for providing physiological state information are provided. The apparatus includes a wearable electronic device. The wearable electronic device includes a sensor module configured to measure motion of the wearable electronic device, a display configured to provide a user interface (UI) including movable particles displayed on the display, and a processor configured to reduce the number of the displayed movable particles based on the measured motion of the wearable device. | 1. A wearable electronic device comprising: a plurality of sensors;a display; anda processor configured to: obtain physiological state information related to a plurality of physiological states of a user of the wearable electronic device through the plurality of sensors,control the display to provide a user interface (UI) including a plurality of movable display objects corresponding to each of the plurality of physiological states, wherein, based upon properties of the physiological state information, the plurality of moveable display objects varies in at least one of a number, a location, a shape, a color, or a size,change at least one of the number, the location, the shape, the color, or the size of the plurality of movable display objects upon analyzing the obtained physiological information, andchange a state of the UI such that a position of the plurality of movable display objects corresponds to a direction of a sensed motion of the wearable electronic device. 2. The wearable electronic device of claim 1, wherein the at least one of a number, a location, a shape, a color, and a size of a display object corresponding to a first physiological state among the plurality of physiological states, is different from a number, a location, a shape, a color, and a size of a display object corresponding to a second physiological state among the plurality of physiological states. 3. The wearable electronic device of claim 1, wherein the plurality of sensors comprises an acceleration sensor to obtain the direction of the sensed motion of the wearable electronic device, andwherein the processor is further configured to: determine at least one of a walking state, a running state, a cycling state, a sitting state, or a sleeping state based on the sensed motion of the wearable electronic device, andchange at least one of the number of the plurality of movable display objects, a location, a shape, a color, or a size of each of the plurality of movable display objects in a different way based on the obtained physiological state information. 4. The wearable electronic device of claim 1, wherein the plurality of sensors comprises a heartbeat sensor to obtain heartbeat information of a user of the wearable electronic device, andwherein the processor is further configured to repeatedly change at least one of a shape, a color, or a size of at least one of the plurality of movable display objects in response to a heartbeat pattern of the user based on the heartbeat information. 5. The wearable electronic device of claim 1, wherein the plurality of sensors comprises a touch sensor to receive a user input associated with setting a fluid intake time, andwherein the processor is further configured to change at least one of the number of the plurality of movable display objects or a location, a shape, a color, or a size of each of the plurality of movable display objects based on the fluid intake time. 6. The wearable electronic device of claim 4, wherein the plura... | Samsung Electronics Co., Ltd. | G06F003/0484 G06F001/16 G06F003/0346 G06F003/01 G06F003/0488 G04G009/00 G06F003/0481 G06K009/00 G04G021/02 A61B005/11 A61B005/024 G16H050/20 A61B005/16 G16H020/60 A61B005/00 G16H040/63 G16H020/30
US10722292 | B2 | 2020-07-28 | Surgical device with an end-effector assembly and system for monitoring of tissue during a surgical procedure | 2014-06-02 | A medical instrument is provided and includes a housing and a shaft coupled to the housing. The shaft has a proximal end and a distal end. An end-effector assembly is disposed at the distal end of the shaft. The end-effector assembly includes first and second jaw members. At least one of the first and second jaw members is movable from a first position wherein the first and second jaw members are disposed in spaced relation relative to one another to at least a second position closer to one another wherein the first and second jaw members cooperate to grasp tissue therebetween. The medical instrument also includes one or more light-emitting elements and one or more light-detecting elements configured to generate one or more signals indicative of tissue florescence. The one or more light-emitting elements are adapted to deliver light energy to tissue grasped between the first and second jaw members. | 1. A medical instrument, comprising: a housing;a shaft coupled to the housing, the shaft having a proximal end and a distal end;an end-effector assembly disposed at the distal end of the shaft, the end-effector assembly including a first jaw member and a second jaw member pivotably coupled to each other defining an angle therebetween, at least one of the first jaw member or the second jaw member movable from a first position wherein the first and second jaw members are disposed in spaced relation relative to one another to at least a second position closer to one another wherein the first and second jaw members cooperate to grasp tissue therebetween, the end effector including an angle sensor disposed in at least one of the first jaw member or the second jaw member, the angle sensor configured to measure the angle between the first jaw member and the second jaw member;at least one light-emitting element coupled to at least one of the first jaw member or the second jaw member, the at least one light-emitting element adapted to deliver light energy at a plurality of wavelengths to the tissue grasped between the first and second jaw members;at least one light-detecting element configured to generate a plurality of fluorescence spectra indicative of tissue fluorescence at the plurality of wavelengths; anda controller configured to receive the plurality of fluorescence spectra and determine properties of a tissue seal based on a difference between the plurality of fluorescence spectra, the controller being coupled to the at least one light-emitting element and the angle sensor, the controller further configured to adjust the intensity of light energy emitted based on the angle. 2. The medical instrument of claim 1, wherein at least one of the first jaw member or the second jaw member includes a groove defined therein having a reflective surface. 3. The medical instrument of claim 2, wherein the at least one light-emitting element is disposed within the groove. 4. The medical instrument of claim 1, further comprising an optical assembly coupled to the at least one light-emitting element, the optical assembly configured to convey light energy emitted from the at least one light-emitting element to the tissue and to illuminate the tissue with a desired illumination pattern. 5. The medical instrument of claim 4, wherein the optical assembly includes at least one of an optical fiber, a refractive element, a reflective element, a diffracting element, and combinations thereof. 6. The medical instrument of claim 1, further comprising a first electrically-conductive tissue-contacting surface associated with the first jaw member and a second electrically-conductive tissue-contacting surface associated with the second jaw member, wherein one of the first electrically-conductive tissue-contacting surface or the second electrically-conductive tissue-contacting surface functions as an active electrode and the other one of the first or the second electrically-c... | Covidien LP | A61B018/00 A61B018/14 A61B005/00 A61B090/30 A61B005/01 A61B017/00 A61B017/29
US10719725 | B2 | 2020-07-21 | System and methods of monitoring driver behavior for vehicular fleet management in a fleet of vehicles using driver-facing imaging device | 2020-01-14 | Systems and methods monitor driver behavior for vehicular fleet management in a fleet of vehicles using driver-facing imaging device. The systems and methods herein relate generally to vehicular fleet management for enhancing safety of the fleet and improving the performance of the fleet drivers, and further relate to monitoring the operation of fleet vehicles using one or more driver-facing imaging devices disposed in the fleet vehicles for recording activities of the fleet drivers and their passengers, storing information relating to the monitored activities, selectively generating warnings related to the monitored activities, and reporting the monitored activities to a central fleet management system for use in enhancing the safety of the vehicles of the fleet and for helping to improve the performance of the fleet drivers. | 1. A system monitoring a driver attention condition of a driver of an associated vehicle during operation of the associated vehicle by the driver, the system comprising: an imaging device capturing an image of the driver disposed in the associated vehicle and generating image data representative of the captured image of the driver; anda control device comprising: a processor;an image data input operatively coupled with the processor, the image data input receiving the image data from the imaging device;a non-transient memory device operatively coupled with the processor, the non-transient memory device storing imaging device position data representative of a position of the imaging device relative to the one or more structures of the associated vehicle, and safe attention model data comprising a recommended value range of a driver road attention parameter of the monitored driver attention condition of the driver;driver head detection logic stored in the non-transient memory device, the driver head detection logic being executable by the processor to: process the image data to determine a head candidate area of the image captured by the imaging device likely above a predetermined threshold stored in the non-transient memory device to be representative of the head of the driver; andtag a portion of the image data corresponding to the head candidate area determined by the driver head detection logic as driver head image data;driver head direction logic stored in the non-transient memory device, the driver head direction logic being executable by the processor to: process the driver head image data to determine a head pose vector as a facing direction of the head of the driver; andgenerate driver head facing direction data, the driver head facing direction data being representative of the determined facing direction of the head of the driver;control logic stored in the non-transient memory device, the control logic being executable by the processor to: process the driver head facing direction data together with the imaging device position data to determine an operational value of the driver road attention parameter of the monitored driver attention condition of the driver; anddetermine a state of vehicle operation compliance as being a driver inattention state in accordance with the determined operational value of the driver road attention parameter of the monitored driver attention condition of the driver being outside of the recommended value range of the driver road attention parameter of the monitored driver attention condition of the driver, wherein the processor generates driver inattention data in accordance with the determined state of the vehicle operation compliance being the driver inattention state; andan output operatively coupled with the processor, the output selectively receiving the driver inattention data and generating a driver inattention signal representative of the determined operational value of the driver road attention parame... | Bendix Commercial Vehicle Systems LLC | G06K009/00 G07C005/08 G07C005/00 B60W040/08 B60W040/09 A61B005/18 A61B005/16
US10722122 | B2 | 2020-07-28 | Detection device for bedding for sleep monitoring | 2014-11-18 | A detection device, for placement under or on a mattress, configured for the movements of a person lying on the mattress, includes a sensing portion having an inflatable chamber intended to be positioned under the individual, an electronic unit arranged at a distance from the sensing portion and having a pressure sensor, intended to be positioned outside of the bedding, a transmitting portion interposed between the sensing portion and the electronic unit, including a channel establishing a fluid connection between the inflatable chamber and the sensor, the channel having a transverse dimension (T) that is much smaller than the width (L1) of the chamber. | 1. A detection device, intended to be placed on or under a mattress so as to detect by ballistography at least the movements, heart rates, and respiratory rate of an individual lying on the mattress, in order to monitor the sleep of the individual, comprising: a sensing portion comprising at least one inflatable chamber intended to contain a fluid consisting of air, intended to be positioned under the individual, under one of the head, chest, or lower torso of the individual,an electronic unit arranged at a distance from the sensing portion, comprising at least one pressure sensor, intended to be positioned externally to the mattress, anda pneumatic transmitting portion, distinct from the sensing portion, interposed between the sensing portion and the electronic unit, comprising at least one channel, establishing a fluid connection between the inflatable chamber and the sensor so that the electronic unit receives the pressure changes induced by movements of the individual;wherein the transmitting portion has upper and lower walls, wherein the channel of the transmitting portion comprises a solid flexible rod preventing the complete collapse of the upper wall of the transmitting portion against the lower wall of the transmitting portion, wherein the inflatable chamber has upper and lower walls,wherein a plurality of braces are provided in the inflatable chamber, joining at the locations where they are placed to the upper and lower walls of the inflatable chamber, andwherein the solid flexible rod extends along and defines a longitudinal axis, and a selected one of the plurality of braces is located at an end of the solid flexible rod and located opposite the solid flexible rod along the same longitudinal axis. 2. The device according to claim 1, further comprising an air pump and a discharge valve associated with the electronic unit, the air pump and the discharge valve able to be used by the electronic unit to inflate and/or deflate the sensing portion. 3. The device according to claim 1, wherein the chamber, even when under pressure, has a thickness that is smaller than its length (L2) and width (L1). 4. The device according to claim 1, wherein the electronic unit comprises a module for converting signals of pressure into electrical signals that can be filtered and analyzed and a wireless communication module configured to send data to a remote entity. 5. The device according to claim 1, further comprising a temperature sensor and/or a humidity sensor. 6. The device according to claim 1, wherein the device in its deflated state is in the form of a flexible sheet that is foldable, can be rolled up, and is washable. 7. The device according to claim 1, wherein the selected one of the plurality of braces which is arranged in alignment with the solid flexible rod comprises a V-shaped end, opened toward the solid flexible rod, and configured to act as a stop for the end of the rod. 8. The device according to claim 1, wherein the width of the tr... | WITHINGS | A61B005/0205 A61B005/11 A61B005/00 A61B005/08 A61G007/05
US10709848 | B2 | 2020-07-14 | Segmented safety cover for needle delivery | 2017-02-14 | A segmented safety cover may include a proximal segment combined with a needle hub; a needle cannula that extends from the proximal needle hub to a distal sharpened end; a distal end cap with a hollow bore through which the needle cannula passes; a centrally mounted toggle locking mechanism, separating the proximal needle hub from the distal end cap; the toggle locking mechanism is torsion spring assisted; a proximal cover segment is connected to the hub via a proximal hinge and a distal cover segment is connected to the end cap via a distal hinge; said proximal cover segment and said distal cover segment of said toggle locking mechanism being configured with abutting faces when the segmented safety cover is in a safety mode; and wherein the proximal hinge, abutting faces and distal hinge are in line in the safe mode. | 1. A segmented safety cover comprising: a) a proximal segment combined with a needle hub; b) a needle cannula that extends from the proximal needle hub to a distal sharpened end; c) a distal end cap with a hollow bore through which the needle cannula passes; d) a centrally mounted toggle locking mechanism, separating the proximal needle hub from the distal end cap, hingedly connected to both and maintaining the segmented safety cover in a position where it completely covers a needle tip on a needle; e) wherein the toggle locking mechanism is resistively bias-assisted and comprises a proximal cover segment and a distal cover segment which overlay the needle cannula and is linked together through a central hinge; f) said proximal cover segment and said distal cover segment of said toggle locking mechanism being configured with abutting faces that contact each other above said central hinge when the segmented safety cover is in a safety mode; g) wherein the proximal cover segment is connected to the hub via a proximal hinge and the distal cover segment is connected to the end cap via a distal hinge; and h) wherein the proximal hinge, abutting faces and distal hinge are in line and the central hinge is below the abutting faces in the safety mode so that a force exerted on the distal end cap is translated onto the abutting faces above the central hinge thereby maintaining the device in a secured state wherein the segmented safety cover is in a position where it completely covers the needle tip, while maintaining the distal end cap's position covering the needle tip extending through the needle cannula. 2. The segmented safety cover of claim 1 wherein the needle is present within the cover, and the segmented safety cover is configured so that the pointed injection end of the needle is completely covered by the distal end cap when the torsion spring is secured in the safety mode, and the pointed injection end of the needle is exposed from the distal end cap when the safety mode is removed and compressive force has been applied by the torsion spring, retracting the distal end cap by rotation of both the proximal cover segment and the distal cover segment about the central hinge. 3. The segmented safety cover of claim 1 wherein the proximal hinge, distal hinge and central hinge are configured to form a triangular configuration when the segmented safety cover is in a delivery mode for the needle within the needle cannula, and the distal cover segment and the distal end cap are configured to slide with respect to the needle cannula within the hollow bore, and wherein the abutting faces rotate away from each other, as the central hinge is elevated with top portions of the abutting faces having rotated with the central hinge farther away from each other than do lower portions of the abutting faces, when a force parallel to a plane formed between the abutting faces is applied upwardly towards the abutting faces. 4. The segmented safety cap of claim 3 wherei... | - | A61M005/32 A61B010/02 A61B005/15 A61B005/153
US10722125 | B2 | 2020-07-28 | Blood pressure signal acquisition using a pressure sensor array | 2017-01-09 | A system and method for blood pressure signal acquisition using a pressure sensor array. A solution is provided for a non-inflatable, non-invasive, continuous blood pressure waveform and blood pressure acquisition system. The system is operative to combine signals from various sensing elements where the less accurate sensor elements are calibrated utilizing the more accurate sensor elements. Blood pressure measurements are acquired using very sensitive pressure sensors that detect slight pressure changes through the skin, which are sampled and processed to yield a blood pressure signal, which is processed to yield actual systolic and diastolic continuous and or intermittent blood pressure readings. By sampling and detecting the signal in each of the sensors, the sensor that is best placed on the target artery is found and the signals from that sensor is used to weight the signals from other sensors based on signal quality. The sensor array is placed approximately on the target artery making it highly likely that more than one element will acquire signal from the artery. Combining a plurality of such correlated signals with uncorrelated noise yields a signal to noise enhancement and more accurate blood pressure readings. | 1. A computerized method of generating an output blood pressure (BP) signal, comprising: providing a plurality of N pressure sensors;acquiring a plurality of N blood pressure waveform signals, each waveform signal acquired from one of said pressure sensors; andcombining, via a processor, said N blood pressure waveform signals, by applying weighted coefficients to each of said N blood pressure waveforms, to generate said output blood pressure signal therefrom, said output blood pressure signal comprised of a weighted sum of each of said N blood pressure waveforms. 2. The method according to claim 1 further comprising deriving at least one of a systolic and diastolic blood pressure measurement from said output blood pressure signal. 3. The method according to claim 1, wherein said plurality of N pressure sensors are adapted to sense blood pressure from at least one of said radial, ulnar and brachial arteries. 4. The method according to claim 1, wherein said output blood pressure signal has a higher signal quality metrics than a signal quality metrics of each of the N blood pressure waveform signals acquired from said N pressure sensors signals. 5. A computerized method of obtaining an output blood pressure (BP) signal, comprising: providing a plurality of N pressure sensors;acquiring a plurality of N blood pressure waveform signals, each waveform signal acquired from one of said pressure sensors; andselecting, via a processor, one of said N blood pressure waveform signals from one of said plurality of N pressure sensors in accordance with one or more quality metrics as said output blood pressure signal;wherein non-selected pressure sensors and their related circuitry are powered down to reduce power consumption. 6. The method according to claim 5, wherein said one or more quality metrics is selected from a group comprising signal to noise ratio (SNR) and received signal strength indication (RSSI). 7. A computerized method of generating an output blood pressure (BP) signal, comprising: providing a plurality of N pressure sensors;acquiring a plurality of N blood pressure waveform signals, each waveform signal acquired from one of said pressure sensors;utilizing waveform signals from one or more relatively accurate pressure sensors from among said plurality of pressure sensors to calibrate one or more less accurate pressure sensors within said plurality of pressure sensors; andcombining, via a processor, said N blood pressure waveform signals including one or more calibrated blood pressure waveform signals to generate a composite output blood pressure signal therefrom. 8. The method according to claim 7, wherein said combining comprises calculating a higher signal to noise ratio (SNR) composite blood pressure signal. 9. The method according to claim 7, wherein said composite blood pressure signal comprises a weighted sum utilizing coefficients for each of said N blood pressure waveforms. 10. The method according to claim 9, wherein said coeff... | LIVEMETRIC (MEDICAL) S.A. | A61B005/0205 A61B005/021 A61B005/00
US10702185 | B2 | 2020-07-07 | Electronic device and body composition analyzing method | 2018-02-14 | An electronic device includes a memory storing profile information, at least one first electrode, at least one second electrode, and a processor connected to the at least one first electrode and the at least one second electrode. The processor is configured to determine a frequency of at least one signal for measuring body composition, based on the profile information, output a synthesis signal obtained by synthesizing the at least three signals to a body of the user through the at least one first electrode, wherein the synthesis signal includes a signal corresponding to the determined frequency and the at least three signals have different frequencies, receive the synthesis signal passing through the body of the user, through the at least one second electrode, measure body impedance of the user based on the received synthesis signal, and measure the body composition of the user based on the body impedance. | 1. An electronic device comprising: a memory configured to store profile information;at least one first electrode;at least one second electrode; anda processor electrically connected to the at least one first electrode and the at least one second electrode,wherein the processor is configured to:determine a frequency of at least one among at least three signals for measuring body composition, based on the profile information;output a synthesis signal to a body through the at least one first electrode, the synthesis signal being obtained by synthesizing the at least three signals, wherein the synthesis signal includes a signal corresponding to the determined frequency and the at least three signals each have different frequencies;receive the synthesis signal passing through the body, through the at least one second electrode;measure body impedance based on the received synthesis signal; anddetermine the body composition based on the body impedance. 2. The electronic device of claim 1, wherein the profile information includes at least one of: height, age, weight, gender, exercise history, food intake history, sleep history, body composition measuring history, body state, and stress index. 3. The electronic device of claim 1, wherein the processor is configured to: measure the body impedance for respective frequencies of the at least three signals. 4. The electronic device of claim 1, wherein the processor is configured to: measure a skin contact impedance based on the synthesis signal; andremove the skin contact impedance from the body impedance. 5. The electronic device of claim 4, wherein the synthesis signal includes at least three first signals, each of the at least three first signals having different frequencies in a first frequency range, and at least three second signals, each of the at least three second signals having different frequencies in a second frequency range, the second frequency range being smaller than the first frequency range, and wherein the processor is configured to:measure the skin contact impedance using the at least three second signals. 6. The electronic device of claim 1, wherein the processor is configured to: generate at least three signals corresponding to the determined frequency; andgenerate the synthesis signal by synthesizing the at least three signals. 7. The electronic device of claim 6, wherein the processor is configured to: adjust a phase of at least one among the at least three signals such that amplitude of the synthesis signal is minimized and/or reduced; andgenerate the synthesis signal using at least three signals, wherein the phase of each of the at least three signals is adjusted. 8. The electronic device of claim 1, wherein the memory is configured to store a plurality of synthesis signals obtained by synthesizing the at least three signals, each of the at least three signals having different frequencies, and wherein the processor is configured to:verify a synthesis signal, including a signal... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/00 A61B005/053 A61B005/024 A61B005/0428 A61B005/117
US10722151 | B2 | 2020-07-28 | Magnetic resonance device having a motion detection unit and a method for detecting a movement of a patient during a magnetic resonance examination | 2015-05-07 | A magnetic resonance device includes a radiofrequency unit that includes a radiofrequency antenna, at least one radiofrequency line and at least one radiofrequency injection point. Radiofrequency signals are transferred to the radiofrequency antenna by the at least one radiofrequency line and are coupled into the radiofrequency antenna at the at least one radiofrequency injection point. The magnetic resonance device also includes a patient receiving zone that is at least partially enclosed by the radiofrequency antenna, and a motion detection unit for detecting a movement of a patient that may be positioned within the patient receiving zone. At least one radiofrequency line includes at least one injection element by which at least one motion detection signal of the motion detection unit is coupled into the radiofrequency line. | 1. A magnetic resonance device comprising: a radiofrequency unit comprising: a radiofrequency antenna;at least one radiofrequency line; andat least one radiofrequency injection point, wherein radiofrequency signals generated by a radiofrequency antenna control unit are transferred to the radiofrequency antenna by the at least one radiofrequency line and injected into the radiofrequency antenna at the at least one radiofrequency injection point;a patient receiving zone that is at least partially enclosed by the radiofrequency antenna; anda motion detection unit comprising a radar unit and operable to detect a movement of a patient that is positionable within the patient receiving zone by at least one radar signal, the radar unit being configured to generate the at least one radar signal, wherein the at least one radar signal has a frequency of at least 3 GHz,wherein the at least one radiofrequency line comprises at least one injection element by which the at least one radar signal generated by the radar unit is coupled into the at least one radiofrequency line, andwherein the radiofrequency antenna is configured to: radiate the radiofrequency signals generated by the radiofrequency antenna control unit into the patient receiving zone via the at least one radiofrequency line and the at least one radiofrequency injection point;emit the at least one radar signal generated by the radar unit into the patient receiving zone, such that the movement of the patient is detectable, the at least one radar signal being transmittable from the radar unit to the radiofrequency antenna via the at least one radiofrequency line and the at least one radiofrequency injection point; andacquire at least one reflected radar signal, the at least one reflected radar signal being transmittable from the radiofrequency antenna to the radar unit via the radiofrequency line and the at least one injection element. 2. The magnetic resonance device of claim 1, wherein the at least one reflected radar signal is acquired by the radiofrequency antenna, and wherein the at least one reflected radar signal is configured to be reflected off the patient that is introducible into the patient receiving zone. 3. The magnetic resonance device of claim 1, wherein the at least one radar signal has a maximum frequency of 30 GHz. 4. The magnetic resonance device of claim 1, wherein the at least one radiofrequency line comprises two or more radiofrequency lines, and the at least one radiofrequency injection point comprises two or more radiofrequency injection points, and wherein a radiofrequency signal of the radiofrequency signals generated by the radiofrequency antenna control unit is transferred in each case to one radiofrequency injection point of the two or more radiofrequency injection points by the two or more radiofrequency lines, and each of the two or more radiofrequency lines has an injection element for injecting a radar signal of the at least one radar signal into the radiofrequenc... | Siemens Aktiengesellschaft | A61B005/11 G01R033/32 G01V003/00 G01S013/02 G01R033/36 G01R033/28 A61B005/08 A61B005/055 A61B005/00 A61B005/05 G01S013/88 G01S013/52 G01V003/175 G01R033/567 G01R033/34
US10716509 | B2 | 2020-07-21 | System and methods for determining nerve proximity, direction and pathology during surgery | 2018-04-03 | The present invention involves systems and methods for determining nerve proximity, nerve direction, and pathology relative to a surgical instrument based on an identified relationship between neuromuscular responses and the stimulation signal that caused the neuromuscular responses. | 1. A system for establishing an operative corridor for spinal surgery, comprising: a surgical access instrument deliverable through bodily tissue having neural structures in a selected path toward a targeted intervertebral disc of a spine, the surgical access instrument having a stimulation electrode positioned along a distal tip region to deliver a stimulation signal to the bodily tissue when the surgical access instrument is positioned in the selected path toward the targeted intervertebral disc of the spine; anda nerve monitoring system comprising a control unit having a display and a plurality of leg muscle sensor electrodes in communication with the control unit, the nerve monitoring system being configured to (a) deliver a plurality of stimulation signals to the stimulation electrode on the surgical access instrument positioned within tissue of a patient that contains a spinal nerve, (b) in response to delivery of each one of the plurality of stimulation current signals, receive a plurality of voltage response values detected by the plurality of leg muscle sensors coupled to different leg muscles, (c) determine the value of an unknown threshold nerve stimulation amplitude required to evoke a threshold response from any of the leg muscles, d) enter a monitoring mode in response to determining the value of the unknown threshold nerve stimulation amplitude, and e) detect whether stimulation current thresholds are changing as the surgical access instrument is advanced along the selected path toward a targeted intervertebral disc by incrementing or decrementing simulation current level delivered by the stimulation electrode on the surgical access instrument by an amount. 2. The system of claim 1, wherein the value of the unknown threshold nerve stimulation amplitude is determined using the steps of bracketing the unknown threshold nerve stimulation amplitude by determining a first stimulation amplitude that does not evoke the threshold response from any of the leg muscles and a second stimulation amplitude that does evoke the threshold response from at least one of the leg muscles, repeatedly bisecting the bracket that includes the threshold nerve stimulation amplitude until a final bracket of a predetermined width is reached, and selecting the threshold nerve stimulation amplitude from the final bracket. 3. The system of claim 1, wherein the nerve monitoring system is further configured to (d) determine which range from a predetermined group of ranges the determined threshold nerve stimulation amplitude is located in, and (e) communicate on the display an indicator corresponding to the determined range from the predetermined group of ranges as an indicator of nerve proximity to the surgical access instrument. 4. The system of claim 1, wherein whether the stimulation current level is incremented or decremented is determined as a function of evoked responses from one or more of the leg muscle sensor electrodes. 5. The system of claim 1, where... | NuVasive, Inc. | A61B005/00 A61B017/02 A61B005/0488 A61N001/08 A61B005/04 A61B017/00 A61N001/36 A61B005/11
US10702375 | B2 | 2020-07-07 | Electromyographic sensing and vision modification | 2016-09-17 | Adjustable vision modification using various systems, devices, and processes are provided. Adjustable vision modification may include sensing ocular related physiological activities of a user and making adjustments in an ocular device or ocular system to change the vision of the user. Various sensors and sensor locations may be employed in embodiments to sense or otherwise obtain these physiological activities. Likewise, various ocular devices, ocular processes, and ocular systems may be employed for improving vision. | 1. An electro-active intraocular lens comprising: an adjustable intraocular lens;a plurality of sensor arms extending away from the adjustable intraocular lens, each of the arms comprising a pair of electric sensors, the pair of electric sensors comprising a driving electrode and a sensing electrode; anda processor configured to receive EMG signals sensed by sensing electrodes of the plurality of sensor arms, and to control focus of the adjustable intraocular lens using the received EMG signals. 2. The electro-active intraocular lens of claim 1 wherein the sensor arms have a parabolic or conchoidal shape and are positioned to center the intraocular lens within the capsular bag when the adjustable intraocular lens is implanted in an eye of a user of the adjustable intraocular lens. 3. The electro-active intraocular lens of claim 1 wherein the adjustable intraocular lens is detachable and wherein the at least one pair of electric sensors are part of a wenner electrode array. 4. The electro-active intraocular lens of claim 1 wherein at least one of the electric sensors is positioned on a sensor arm such that when the adjustable intraocular lens is implanted in an eye of a user of the adjustable intraocular lens, the sensor is positioned in an intravitreal space. 5. The electro-active intra-ocular lens of claim 1 wherein the processor carries out steps comprising detecting the EMG signals, determining the need for accommodative correction based on the detected signals, providing a drive signal to affect a change in the adjustable intraocular lens, which at least partially restore lost accommodation of an eye of a user of the adjustable intraocular lens. 6. The electro-active intra-ocular lens of claim 1 wherein the processor determines impedance of a received physiological signal impedance cyclography and signals received from one or more of the electric sensors. 7. An ocular system for improving vision comprising: an adjustable intraocular lens;a pair of haptic sensor arms extending away from the adjustable intraocular lens, each of the arms comprising a mated pair of electric sensors, the mated pair of electric sensors comprising an anode sensor and a cathode sensor; anda processor configured to receive EMG signals sensed by the anode sensors and the cathode sensors and to control focus of the adjustable intraocular lens using the received EMG signals. 8. The system of claim 7 wherein at least one mated pair of electric sensors are positioned and configured to detect at least impedance changes occurring in a ciliary muscle during accommodation. 9. The system of claim 8 wherein the impedance changes are used as input for the adjustable intraocular lens and wherein the adjustable intraocular lens is a variable focus lens. 10. The system of claim 7 further comprising an electric sensor sized between 0.5 mm2 and 8.0 mm2, and configured to be implanted in at least one of the following ocular locations of a user of the adjustable intraocular len... | VISTA OCULAR, LLC | A61F002/16 G02C007/08 G02C007/04 A61B005/00 A61B005/0492 A61F002/48
US10716940 | B2 | 2020-07-21 | Implant unit for modulation of small diameter nerves | 2014-06-17 | An implant unit configured for implantation into a body of a subject is provided. The implant unit may include an antenna configured to receive a signal, and at least one pair of modulation electrodes configured to be implanted into the body of the subject in the vicinity of at least one nerve to be modulated. The at least one nerve may have a diameter less than about 2 millimeters at a modulation location, and the at least one pair of modulation electrodes may be configured to receive an applied electric signal in response to the signal received by the antenna and generate an electrical field to modulate the at least one nerve from a position spaced apart from the at least one nerve. | 1. An implant unit configured for implantation into a body of a subject, comprising: a carrier comprising a substantially planar central portion, a first elongated arm extending from a first side of the central portion, and a second elongated arm extending from a second side of the central portion, the carrier being configured to be implanted into the body of the subject with at least one of the first elongated arm and the second elongated arm extending in a vicinity of a terminal end of at least one nerve to be modulated, the carrier further being configured to be implanted into the body of the subject such that the at least one of the first elongated arm and the second elongated arm is not in physical contact with the at least one nerve and such that non-nerve tissue is interposed between the at least one nerve and the at least one of the first elongated arm and the second elongated arm, wherein the first elongated arm includes a suture hole;an antenna configured to receive a signal, the antenna disposed entirely on the substantially planar surface of the central portion of the carrier; anda first pair of modulation electrodes disposed on the first elongated arm and a second pair of modulation electrodes disposed on the second elongated arm, each of the first pair of modulation electrodes and the second pair of modulation electrodes being configured to receive an applied electric signal in response to the signal received by the antenna and generate an electric field, the electric field being configured for penetrating non-nerve tissue interposed between the at least one nerve and the modulation electrodes and for modulating the at least one nerve at the modulation location. 2. The implant unit of claim 1, wherein at the terminal end of the at least one nerve, the at least one nerve has a diameter of less than about one millimeter. 3. The implant unit of claim 1, wherein the first pair of modulation electrodes are spaced apart in a longitudinal direction of the at least one nerve to be modulated and configured to generate an electric field in the longitudinal direction. 4. The implant unit of claim 1, wherein at the terminal end of the at least one nerve, a diameter of the least one nerve is less than 1 mm. 5. The implant unit of claim 1, wherein at the terminal end of the at least one nerve, a diameter of the least one nerve is less than 500 microns. 6. The implant unit of claim 1, wherein at the terminal end of the at least one nerve, a diameter of the least one nerve is less than 200 microns. 7. The implant unit of claim 1, wherein at the terminal end of the at least one nerve, a diameter of the least one nerve is less than 100 microns. 8. The implant unit of claim 1, wherein at the terminal end of the at least one nerve, a diameter of the least one nerve is less than 50 microns. 9. The implant unit of claim 1, wherein at the terminal end of the at least one nerve, a diameter of the least one nerve is less than 25 microns. 10. The im... | Nyxoah SA | A61N001/36 A61N001/05 A61N001/375 A61N001/378 A61N002/00 A61B005/00 A61B005/11 A61B005/113 A61F005/56 A61N001/372 A61B005/08 A61N002/02 A61B005/0488
US10718728 | B2 | 2020-07-21 | In-vitro sensor using a tetrapolar impedance measurement | 2016-12-16 | Systems, methods, and devices are provided for determining an impedance measurement of a sample of fluid on a test strip having a base layer, at least two drive electrodes disposed on the base layer and in electrical communication with a current-source (Iac) for flowing an AC current provided by the current-source (Iac) through a fluid sample between the at least two drive electrodes, and at least two sense electrodes disposed on the base layer and positioned between the at least two drive electrodes, each of the at least two sense electrodes configured for measuring a difference in an AC potential therebetween to determine an impedance measurement of the fluid sample between the at least two sense electrodes. | 1. A test strip comprising: a base layer;a capillary chamber configured to receive a fluid sample;at least two drive electrodes disposed on the base layer to be placed in electrical communication with a current-source (Iac) for flowing an AC current provided by the current-source (Iac) through the fluid sample in the capillary chamber and between the at least two drive electrodes;at least two sense electrodes disposed on the base layer and positioned between the at least two drive electrodes, each of the at least two sense electrodes configured for measuring a difference in an AC potential therebetween to determine an impedance measurement of the fluid sample between the at least two sense electrodes; anda conductive trace extending from each of the at least two drive electrodes and running parallel and adjacent to at least one of the at least two sense electrodes to eliminate or reduce the stray capacitances by creating a capacitance bridge between the at least two drive electrodes,wherein when the at least two sense electrodes are in direct electrical communication with a high input-impedance voltage measurement circuit to measure the difference in AC potential between the at least two sense electrodes, an impedance measurement is indicative of a hematocrit level in the fluid sample. 2. The test strip of claim 1, wherein the current-source (Iac) is programmable in amplitude and frequency. 3. The test strip of claim 1, wherein the current-source (Iac) is produced from a power source. 4. The test strip of claim 1, wherein the AC current-source (Iac) is configured to apply the AC current through the fluid sample alternatively from a first one of the at least two drive electrodes to another one of the at least two drive electrodes as determined by a frequency of the AC current provided by the current-source (Iac). 5. A test strip comprising: a base layer;a capillary chamber configured to receive a fluid sample;at least two drive electrodes disposed on the base layer to be placed in electrical communication with a current-source (Iac) for flowing an AC current provided by the current-source (Iac) through the fluid sample in the capillary chamber and between the at least two drive electrodes;at least two sense electrodes disposed on the base layer and positioned between the at least two drive electrodes, each of the at least two sense electrodes configured for measuring a difference in an AC potential therebetween to determine an impedance measurement of the fluid sample between the at least two sense electrodes; andwherein when the at least two sense electrodes are configured to be connected to a high input-impedance voltage measurement circuit to measure the difference in AC potential between the at least two sense electrodes,wherein the voltage measurement circuit includes a voltmeter configured to reduce the AC current flowing between the at least two sense electrodes and the voltmeter when the at least two sense electrodes are in direct elec... | Trividia Health, Inc. | G01N027/327 G01N027/02 G01N033/49 A61B005/145
US10702206 | B2 | 2020-07-07 | Toothbrush for oral cavity position detection | 2016-06-28 | A toothbrush handle for providing oral cavity position detection during operation includes a handle body having a longitudinal axis, a receiver body elastically deformable relative to the longitudinal axis, and a three-axis force sensor that is operatively connected to the base portion and configured to obtain three-axis force data for oral cavity position detection during operation. | 1. An electric toothbrush handle for providing oral cavity position detection during operation, the handle comprising: a handle body having a longitudinal axis L;a receiver body extending from the handle body along said longitudinal axis L, wherein the receiver body has: an engaging portion distal the handle body and configured to removably engage a replaceable toothbrush head; anda base portion opposing the engaging portion, wherein the base portion adjoins the handle body;an electric motor contained in the handle body;a drive shaft operatively connected to the electric motor and extending along the longitudinal axis L from the motor and at least partially into the receiver body;wherein the receiver body is elastically deformable relative to said longitudinal axis L; anda three-axis force sensor operatively connected to the base portion and configured to obtain three-axis force data for oral cavity position detection during operation, wherein the three-axis force data comprises (a) lateral bending forces in at least one of an x-axis direction and a y-axis direction caused by elastic deformation of the receiver body and (b) traction or compression forces along the longitudinal axis, wherein the x-axis and y-axis directions are perpendicular to each other and to the longitudinal axis. 2. The handle of claim 1, wherein the three-axis force sensor comprises at least three strain gauges operatively and physically connected to a surface of the base portion. 3. The handle of claim 2, wherein the three-axis force sensor comprises at least four strain gauges operatively and physically connected to the surface of the base portion, wherein said surface of the base portion is an outside surface of the base portion but within a housing. 4. The handle of claim 2, wherein the at least three strain gauges are operatively and physically connected to the surface of the base portion. 5. The handle of claim 4, wherein the base portion has a cross section in a shape of a polygon; and wherein the cross section has a cross sectional area from about 20 mm2 to about 30 mm2. 6. The handle of claim 1, wherein the base portion has a cross sectional area from about 10 mm2 to about 40 mm2. 7. The handle of claim 1, further comprising an accelerometer and gyroscope contained within the handle body and configured to providing three-axis accelerometer and gyroscope data force for oral cavity position detection during operation. 8. The handle of claim 5, further comprising an accelerometer and gyroscope contained within the handle body and configured to providing three-axis accelerometer and gyroscope data force for oral cavity position detection during operation. 9. The electric toothbrush handle of claim 7, further comprising: a data transmitter, in communication with the 3-axis force sensor and accelerometer and gyroscope, configured for transmitting 3-axis force sensor data and the 3-axis accelerometer and gyroscope data. 10. The handle of claim 8, further comprisin... | BRAUN GMBH | A61B005/00 A61C017/16 A46B015/00 A46B009/04 A46B005/00 A61B005/053 A61C019/04 A61C017/22 A61B090/00
US10718618 | B2 | 2020-07-21 | Automatically creating a terrain mapping database | 2017-09-07 | A database for mapping off-road terrain of various characteristics in three-dimensional terms, comprising: a plurality of road segment entries, each containing data pertaining to the terrain characteristics of the segment; and a plurality of elbow entries, each containing (x, y, z) coordinates of the elbow and a record for each road segment having the elbow as one of its end points, the record comprising navigation directives for vehicles entering the segment from the elbow. | 1. A method of creating and continuously updating a database for mapping road terrain of various characteristics in three-dimensional terms, comprising: defining a plurality of road segments, each road segment defined by two bounding elbows, using at least one of maps and recorded trails, wherein the defining of the plurality of road segments further comprise: defining preliminary terrain characteristics for each road segment, wherein said preliminary terrain characteristics within each road segment are substantially homogenous and configured to be traversed using a single set of directives;defining preliminary navigation directives for vehicles entering each road segment from each of the two bounding elbows based on the preliminary terrain characteristics for each road segment; andcontinuously updating the database using data aggregated from sensors carried by travelers, wherein the sensors are position sensors including at least one of: Global Positioning System (GPS), European Geostationary Navigation Overlay Service (EGNOS), and Global Navigation Satellite System (GLONASS). 2. The method of claim 1, wherein currently updating the database comprises updating segment definitions according to recorded trails intersecting existing road segments. 3. The method of claim 1, wherein currently updating the database comprises updating segment definitions according to sensors data indicating changes in terrain characteristics within existing road segments. 4. The method of claim 1, wherein the data aggregated from the sensors is analyzed to determine terrain characteristics and levels of difficulty of trails. 5. The method of claim 4, wherein the data aggregated from the sensors is translated into navigation directives stored in the database in conjunction with the relevant road segments. 6. The method of claim 1, wherein the data aggregated from the sensors is used to update road segments and elbows definitions by identifying different characteristics of various parts of a segment and partitioning the road segment into two or more road segments according to different terrain characteristics of each portion which require different navigation directives. 7. The method of claim 1, wherein the navigation directives are translated into appropriate directives depending on a type vehicle used. 8. The method of claim 1, wherein the sensors are further selected from the group consisting of: motion sensors, environmental sensors, wearable physiological monitoring sensors, and cameras and microphones. 9. The method of claim 8, wherein the motion sensors are selected from the group consisting of: accelerometers, gravity sensors, gyroscopes and rotational vector sensors. 10. The method of claim 8, wherein said environmental sensors are selected from the group consisting of: barometers, photometers and thermometers. 11. The method of claim 8, wherein the wearable physiological monitoring sensors are configured to measure physiological parameters selected f... | Trailze, Ltd | G01C021/20 G01C021/10 G01S019/42 G01V099/00 A61B005/01 A61B005/021 A61B005/0404 A61B005/00 G06F016/29 G01C021/36 G01C021/32 G09B029/10
US10702205 | B2 | 2020-07-07 | Apparatus and method for monitoring rehabilitation from surgery | 2018-07-04 | Apparatus and methods are described for setting rehabilitation goals for a patient, measuring patient movements, storing the movement data for later transfer to a computer, displaying progress indicators and inspirational messages based on progress towards goals, reporting movement, skin temperature and swelling data to a caregiver so that they monitor compliance and be aware of potential infection. The invention consolidates data from motion, temperature and swelling sensors placed on the patient with data received directly from the patient in response to questions displayed on a tablet computer screen, to create one or more representative values indicating the patient's current status. These representative values may include the patient's degree of compliance with their prescribed exercise and icing regimen; their degree of wellness based on their self-reported pain scale and sensed temperature and limb swelling; and their current level of physical activity as detected by the motion sensors. | 1. Apparatus for monitoring a patient, comprising: a first motion sensor adapted to be attached to a patient's limb distally of a joint;a first temperature sensor adapted to be attached to the patient's limb in proximity to a surgical site;a second motion sensor adapted to be attached to the patient's limb proximately of the joint;a second temperature sensor adapted to be attached to the patient's limb proximate the joint and the spaced from the first temperature sensor;wherein the first and second motion sensors are configured for generating limb orientation data and transmitting the limb orientation data to a microprocessor that includes memory to store the limb orientation data and execute instructions thereupon, and to thereby generate a representation of the orientation of the patient's limb distal of the joint and the patient's limb proximate of the joint over time and to thereby count the number of flexions of the joint;a patient-assigned computer;wherein the microprocessor is configured to generate an activity score bya. using limb orientation data to determine the number of flexions and compare the number of flexions to a predetermined ideal number of flexions and to calculate a percentage of normal activity;b. determining the number of times the patient has interacted with the patient-assigned computer to calculate a percentage of normal interaction frequency;c. using limb orientation data to calculate the percentage of time the patient is moving; andd. weighting the percentage of normal activity, the percentage of normal interaction frequency, and the percentage of time the patient is moving to derive the activity score. 2. The apparatus according to claim 1 in which the limb orientation data further comprises the maximum and minimum flex angle of each flexion. 3. The apparatus according to claim 1 wherein the first and second temperature sensors are configured for generating temperature data and transmitting the temperature data to the microprocessor for comparison of the temperature data from the first temperature sensor with the temperature data from the second temperature sensor and thereby determine the difference in temperature data between the first and second temperature sensors. 4. The apparatus according to claim 3 in which the difference in temperature data between the first and second temperature sensors is indicative of infection at the surgical site. 5. The apparatus according to claim 3 in which the difference in temperature data between the first and second temperature sensors is indicative of icing at the surgical site. 6. The apparatus according to claim 1 further comprising calculating the percentage of normal interaction frequency by determining the number of times the patient has interacted with the patient-assigned computer and comparing that to a predetermined ideal number of interactions. 7. The apparatus according to claim 1 in which the activity score is calculated by weighting the percentage of normal a... | Claris Healthcare Inc. | A61B005/00 G06F019/00 G16H010/60 A61B005/11 A61B005/107 A61B005/01 G16H040/63 G16H020/30
US10726699 | B2 | 2020-07-28 | Method for defining access perimeters and handling perimeter breach events by residents of an assisted living facility | 2018-04-16 | One variation of a method for detecting breach events by residents of an assisted living facility includes: defining a first access perimeter within a facility; assigning the first access perimeter to a first resident, in a set of residents in the facility; tracking locations of resident wearable devices associated with residents, in the set of residents, within the facility; disabling a first subregion within the first access perimeter coincident a second location of a second resident, in the set of residents, flagged for separation from the first resident; and, in response to a first location of a first resident wearable device associated with the first resident falling outside of the first access perimeter, distributing a first breach event prompt to assist the first resident, proximal the first location, to care provider mobile devices associated with care providers affiliated with the facility. | 1. A method for detecting breach events by residents of an assisted living facility, the method comprising: defining a generic access perimeter within a facility;assigning the generic access perimeter to residents in a set of residents occupying the facility;tracking locations of a first care provider mobile device, associated with a first care provider, within the facility over a period of timedefining a first care provider response area containing and extending outwardly from the locations of the first care provider mobile device over the period of time; anddefining a first individual access perimeter excluding a first portion of the generic access perimeter and incorporating the first care provider response area;assigning the generic access perimeter to a first resident first resident in the set of during a daytime period;assigning the first individual access perimeter to the first resident, in the set of residents during a night time period distinct from the daytime period;tracking locations of resident wearable devices associated with residents, in the set of residents, within the facility;in response to a first location of a first resident wearable device associated with the first resident falling outside of the first individual access perimeter at a first time, distributing a first breach event prompt to assist the first resident, proximal the first location, to care provider mobile devices associated with care providers affiliated with the facility; andin response to a second location of a second resident wearable device associated with a second resident, in the set of residents, falling outside of the generic access perimeter, assigned to the second resident, at a second time, distributing a second breach event prompt to assist the second resident proximal the second location to the care provider mobile devices associated with care providers affiliated with the facility. 2. The method of claim 1: wherein assigning the generic access perimeter to residents occupying the facility comprises assigning the generic access perimeter to the first resident by default;further comprising: at an initial time, detecting an initial breach event in response to an initial location of the first resident wearable device falling outside of the generic access perimeter; andin response to detecting the initial breach event, distributing an initial breach event prompt to assist the first resident proximal the initial location to the care provider mobile devices associated with care providers affiliated with the facility; andwherein defining the first individual access perimeter comprises defining the first individual access perimeter excluding the first portion of the generic access perimeter containing the initial location responsive to the initial breach event. 3. The method of claim 1: wherein defining the first individual access perimeter comprises defining the first individual access perimeter spanning a first private room assigned to the first residen... | PHILIPS NORTH AMERICA LLC | A61B005/11 A61B005/00 G08B021/04 G08B021/02 G08B025/00 H04W004/02 G08B025/10 G08B025/08 G08B013/12 G06Q050/22 G06Q050/26 G08B021/00 A61B005/103 G16H040/63 G16H010/60
US10702181 | B2 | 2020-07-07 | Atrial fibrillation treatment systems and methods | 2018-05-31 | Methods for treating cardiac complex rhythm disorder in a patient can include receiving a plurality of electrical signals from a sensor system, wherein each electrical signal corresponds with a separate location on a cardiac wall of the heart of the patient, and wherein each electrical signal comprises an electrogram waveform; and ranking the electrical signals relative to each other based on at least a uniformity and a frequency of the electrogram waveform of each electrical signal. | 1. A method of treating cardiac complex rhythm disorder in a patient, the method comprising: receiving an electrical base signal of a heart comprising a base signal frequency and a base signal waveform;simultaneously positioning a sensor system inside the heart of a patient while receiving a plurality of electrical signals from the sensor system, wherein positioning comprises identifying a source of repetitive aberrant electrical activity in the plurality of electrical signals and placing the sensor system such that the source is within a sensing region, wherein the sensing region is defined by a plurality of sensor locations that are in contact with a cardiac wall of the heart, wherein each electrical signal is detected at each of the plurality of sensor locations for at least one heartbeat, and wherein each electrical signal comprises an electrogram waveform;determining that the source of repetitive aberrant electrical activity is propagating in an outward direction from within the sensing region based on at least a timing and a frequency of the electrogram waveform of each electrical signal that is distinguishable from the base signal frequency, wherein the sensor system is positioned at a location where the source of repetitive aberrant electrical activity is disposed between a plurality of sensor locations in at least two different directions; andelectrically isolating the source of repetitive aberrant electrical activity by relocating the sensor system and performing at least one focal ablation procedure inside the sensing region. 2. The method of claim 1, wherein a location of the source of repetitive aberrant electrical activity defines a target site within the sensing region and at least one focal ablation procedure is performed at the target site. 3. The method of claim 2, further comprising: repositioning the sensor system to another or the same sensing region that encompasses the target site after performing at least one focal ablation procedure and receiving a second plurality of electrical signals from the sensor system; and determining if the repetitive aberrant electrical activity has been eliminated from the electrical signals. 4. The method of claim 1, further comprising: generating a representative image of the heart of the patient; andidentifying a target position on the representative image that corresponds to a location of the source of aberrant electrical activity. 5. The method of claim 1, further comprising identifying a target site of the cardiac wall, wherein identifying the target site comprises interpolating a location of the target site from the sensor locations on the cardiac wall. 6. The method of claim 1, wherein the at least one focal ablation procedure comprises encircling the source with a plurality of focal ablations. 7. The method of claim 1, wherein relocating the sensor system comprises removing the sensor system from its position inside the heart. 8. A method of treating cardiac complex rhythm disor... | Intermountain Intellectual Asset Management, LLC | A61B005/046 A61B005/00 A61B005/04
US10702176 | B2 | 2020-07-07 | Multielectrode ECG sensor | 2015-08-21 | A multielectrode sensor for use in obtaining electrocardiograph measurements of a patient is disclosed. The sensor includes a substrate, at least a portion of the substrate being stretchable, and a plurality of electrodes coupled to the substrate. A distance between at least two of the electrodes is adjustable by stretching the substrate between the at least two electrodes. | 1. A multielectrode sensor for use in obtaining electrocardiograph measurements of a patient, the sensor comprising: a substrate, at least a portion of the substrate being stretchable; anda plurality of electrodes coupled to the substrate, the plurality of electrodes including V1, V2, V3, V4, V5, and V6 electrodes;wherein a distance between at least the V1 and V2 electrodes and between the V4 and V5 electrodes is adjustable by stretching the substrate between the V1 and V2 electrodes and between the V4 and V5 electrodes;wherein a first thickness of the substrate between the V1 and V2 electrodes is less than or equal to a second thickness of the substrate between the V4 and V5 electrodes such that a first stretchability of the substrate between the V1 and V2 electrodes is greater than or equal to a second stretchability of the substrate between the V4 and V5 electrodes; andwherein the first thickness is less than a third thickness of the substrate between the V2 and V3 electrodes such that the first stretchability is greater than a third stretchability of the substrate between the V2 and V3 electrodes. 2. The multielectrode sensor of claim 1, wherein the substrate comprises a silicone material. 3. The multielectrode sensor of claim 1, wherein the V1 electrode is adapted to be positioned on a first side of the patient's sternum and the V2 electrode is adapted to be positioned on a second side of the patient's sternum. 4. The multielectrode sensor of claim 1, wherein a distance between the V5 and V6 electrodes is adjustable by stretching the substrate between the V5 and V6 electrodes. 5. The multielectrode sensor of claim 4, wherein the first thickness is less than or equal to a fourth thickness of the substrate between the V5 and V6 electrodes such that the first stretchability is greater than or equal to a fourth stretchability of the substrate between the V5 and V6 electrodes. 6. The multielectrode sensor of claim 5, wherein the second thickness differs from the fourth thickness. 7. The multielectrode sensor of claim 5, wherein the second thickness is equal to the fourth thickness such that the second stretchability is substantially the same as a fourth stretchability. 8. The multielectrode sensor of claim 5, wherein the first thickness is less than at least one of the second thickness and the fourth thickness. 9. The multielectrode sensor of claim 1, further comprising a backing sheet removably attached to a patient-contacting side of the substrate. 10. The multielectrode sensor of claim 1, wherein each electrode comprises an adhesive on a patient-contacting side of the electrode. 11. The multielectrode sensor of claim 1, wherein each of the electrodes is connected to a wire, the wire being at least partially disposed in the substrate. 12. The multielectrode sensor of claim 11, wherein a distal end of each wire opposite the electrode is adapted to be connected to at least one of an electrocardiograph machine and a connector. 13. The ... | Children's Medical Center Corporation Theranova, LLC | A61B005/0408 A61B005/00
US10709325 | B2 | 2020-07-14 | Monitoring component and method for monitoring a visual capacity modification parameter of a user of at least one display device | 2016-06-30 | A method for monitoring a visual capacity modification parameter of a user of at least one display device, the visual capacity modification parameter being indicative of the modification of a parameter linked to the visual capacity of the user related to the use of the at least one display device, the method includes a display use data determining step during which display use data indicative of a use of the at least one display device by the user are determined, a visual capacity modification parameter determining step during which a modification parameter of the visual capacity of the user is determined based on said display use data, and an information generating step during which an information based on said modification parameter of the visual capacity of the user is generated. | 1. A method for monitoring a visual capacity modification parameter of a user of a plurality of display devices, the visual capacity modification parameter being indicative of modification of a parameter linked to a visual capacity of the user related to a use of the plurality of display devices, the method comprising: determining display use data indicative of a use of each display of the plurality of display devices by the user;determining a modification parameter of the visual capacity of the user based on said display use data; andgenerating information based on said modification parameter of the visual capacity of the user. 2. The method according to claim 1, wherein the visual capacity modification parameter is a visual capacity deterioration risk indicative of a risk of deterioration of the visual capacity of the user related to the use of the plurality of display devices. 3. The method according to claim 1, wherein the visual capacity modification parameter is a parameter linked to improved visual capacity of the user related to the use of the plurality of display devices. 4. The method according to claim 1, wherein the information comprises at least a control signal configured to be sent to and to control each of the plurality of display devices and is based on said visual capacity modification parameter of the user. 5. The method according to claim 1, wherein the generating further comprises generating recommendation data based at least on the determined visual capacity modification parameter. 6. The method according to claim 5, wherein the recommendation data comprises an alert indicative of the determined visual capacity modification parameter. 7. The method according to claim 1, further comprising obtaining display device data indicative of each display device of the plurality of the display devices used by the user,wherein:the visual capacity modification parameter is determined based on said display device data and on each determined display use data indicative of a use of each display device by the user. 8. The method according to claim 7, wherein the determining the display use data further comprises sending determined display use data, indicative of a use of each display device by the user, to an information generating unit, and the information generating unit is a distant entity from at least one of the display devices and is configured to receive the display use data and the display device data from each display device of the plurality of display devices. 9. The method according to claim 1, further comprising storing the determined display use data over time,wherein the visual capacity modification parameter is indicative of an evolution over time of the modification parameter of the visual capacity of the user related to the use of the plurality of display devices. 10. The method according to claim 1, wherein display use data indicative of a use of the plurality of display devices by the user includes at least time s... | Essilor International | A61B003/10 A61B003/00 G09G003/36 G02B027/00 A61B003/032 G02B027/01 A61B005/00
US10709406 | B2 | 2020-07-14 | Radiography apparatus, method for controlling radiography apparatus, and program | 2017-12-21 | A radiography apparatus includes a radiation emitting device that irradiates a subject with radiation, a camera that captures an image of the subject to acquire a captured image of the subject, and a radiation detector that detects the radiation transmitted through the subject and generates a radiographic image of the subject. The driving state of at least one of the radiation emitting device or the radiation detector is controlled on the basis of whether the radiation detector is included in the captured image. | 1. A radiography apparatus comprising: a radiation emitting device that irradiates a subject with radiation;a camera for capturing an image of the subject to acquire a captured image of the subject;a radiation detector that detects the radiation transmitted through the subject and generates a radiographic image of the subject; andat least one processor configured to operate as:a driving state control unit for controlling a driving state of at least one of the radiation emitting device or the radiation detector on a basis of whether the radiation detector is included in the captured image. 2. The radiography apparatus according to claim 1, wherein, in a case in which the radiation detector is included in the captured image, the driving state control unit controls the driving state of at least one of the radiation emitting device or the radiation detector such that power consumption is more than that in a case in which the radiation detector is not included in the captured image. 3. The radiography apparatus according to claim 1, wherein, in a case in which a state of the radiation detector changes from a state in which the radiation detector is not included in the captured image to a state in which the radiation detector is included in the captured image, the driving state control unit changes the driving state of the radiation detector from a power-off state to a power-on state. 4. The radiography apparatus according to claim 1, wherein, in a case in which a state of the radiation detector changes from a state in which the radiation detector is not included in the captured image to a state in which the radiation detector is included in the captured image, the driving state control unit changes the driving state of the radiation detector from a sleep state to a standby state. 5. The radiography apparatus according to claim 1, wherein, in the case in which a state of the radiation detector changes from a state in which the radiation detector is not included in the captured image to a state in which the radiation detector is included in the captured image, the driving state control unit changes the driving state of the radiation emitting device from a sleep state to a standby state. 6. The radiography apparatus according to claim 1, further comprising: a monitor for displaying the captured image. 7. The radiography apparatus according to claim 6, wherein the at least one processor is further configured to operate as: a display control unit for displaying at least one of identification information of the radiation detector, the driving state of the radiation detector, a vertical direction of the radiation detector, a remaining battery level of the radiation detector, a position of a center of the radiation detector in the case in which the radiation detector is included in the captured image, or a direction in which the radiation detector is present in the case in which the radiation detector is not included in the captured image so as to be su... | FUJIFILM Corporation | A61B006/00 A61B006/12 A61B006/06 A61B006/08 A61B005/00
US10725128 | B2 | 2020-07-28 | Adjustable local coil arrangement | 2018-04-25 | A local coil arrangement and a magnetic resonance apparatus are provided for examining an examination object by magnetic resonance. A local coil arrangement includes at least a first coil unit and a second coil unit and at least a first intermediary unit. In an assembled state, the first intermediary unit is arranged between the first coil unit and the second coil unit. The first coil unit, the second coil unit and the first intermediary unit enclose at least partially an object receiving region for accommodating the examination object. The object receiving region includes a central axis. The shape and/or size of the object receiving region may be changed by relative movements of the first coil unit, of the second coil unit and of the first intermediary unit perpendicular to the central axis. | 1. A local coil arrangement for examining an examination object by magnetic resonance, the local coil arrangement comprising: at least a first coil unit and a second coil unit, wherein the first coil unit is rigid; andat least a rigid first intermediary unit, wherein in an assembled state, the first intermediary unit is arranged between the first coil unit and the second coil unit,wherein the first coil unit, the second coil unit, and the first intermediary unit enclose at least partially an object receiving region for accommodating the examination object,wherein the object receiving region includes a central axis,wherein the first coil unit, the second coil unit, the first intermediary unit, or any combination thereof comprises at least a portion of a magnetic resonance antenna;wherein a shape and a size of the object receiving region is adjusted by relative movements between assembled states of the first coil unit, the second coil unit, the first intermediary unit, or any combination thereof;wherein relative movements of the first coil unit, relative movements of the second coil unit, and relative movements of the first intermediary unit are configured such that: the shape, the size, or the shape and the size of the object receiving region are adjustable by relative movements of the first coil unit in a direction perpendicular to the central axis, relative movements of the second coil unit in the direction perpendicular to the central axis, and relative movements of the first intermediary unit in the direction perpendicular to the central axis; wherein the relative movements of the first coil unit, the relative movements of the second coil unit, and the relative movements of the first intermediary unit are performed relative to one another. 2. The local coil arrangement of claim 1, wherein for the relative movements of the first coil unit in the direction perpendicular to the central axis, the relative movements of the second coil unit in the direction perpendicular to the central axis, and the relative movements of the first intermediary unit in the direction perpendicular to the central axis, distances between the central axis and the first coil unit, distances between the central axis and the second coil unit, and distances between the central axis and the first intermediary unit change in the same way. 3. The local coil arrangement of claim 1, wherein the first coil unit and the first intermediary unit have respective center planes that intersect in the central axis, and wherein the relative movements of the first coil unit in the direction perpendicular to the central axis and the relative movements of the first intermediary unit in the direction perpendicular to the central axis are performed parallel to the respective center planes. 4. The local coil arrangement of claim 1, further comprising a first part and a second part, wherein the first part comprises the first coil unit and the first intermediary unit,wherein the second part com... | Siemens Healthcare GmbH | G01R033/34 A61B005/055 G01R033/3415
US10709413 | B2 | 2020-07-14 | System and method for analysis of the upper airway and a respiratory pressure support system | 2016-07-04 | A system for analysis of the upper airway has a sensor arrangement with at least two sensor positions provided along a flow path leading to the mouth and/or nose of a user (4). A relation is derived between sensor signals at the two locations, and this is interpreted to detect at least the presence of upper airway obstructions, and preferably also the location and/or extent of such obstructions. The system is adapted to distinguish between inhalation and exhalation using the acoustic sensor arrangement signals at the first and second locations. | 1. A system for analysis of an upper airway of a user, comprising: a passageway for communicating a flow of gas to and from the upper airway;an acoustic sensor arrangement positioned at first and second locations within the passageway, wherein the second location is located along a length of the passageway different from the first location, wherein the acoustic sensor arrangement is configured to detect acoustic signals at the first and second locations, and further for use in distinguishing between signal components with different relative time shift between the first and second locations;a processor adapted to derive a parameter from a relation between the acoustic signals detected at the first and second locations, wherein the parameter varies in dependence on a presence of an upper airway obstruction and enables a detection of (i) the presence and (ii) a location of the upper airway obstruction within the upper airway,wherein the processor is adapted to process sound signals which comprise only ambient noise received from noise sources external to the analysis system, and noise generated by the user and is adapted to distinguish between inhalation and exhalation in response to a quantified change in the parameter during inhalation and exhalation based on the acoustic signals detected at the first and second locations. 2. The system as claimed in claim 1, wherein the acoustic sensor arrangement comprises a first acoustic sensor positioned at the first location and a second acoustic sensor positioned at the second location. 3. The system as claimed in claim 2, wherein the passageway comprises a tube and the first and second acoustic sensors are positioned at first and second locations along the tube, which has a first end which is for terminating at a mouth and/or nose of the user. 4. The system as claimed in claim 1, further comprising a detector that comprises a classifier based on supervised learning with an open airway impedance and with several known obstructed impedances, wherein the classifier is configured to interpret the parameter, via selecting, from a training set of spectrums, a corresponding spectrum that is a closest value to the parameter to detect the presence of the upper airway obstruction. 5. The system as claimed in claim 4, wherein the detector is further adapted to detect at least one of (i) the location and (ii) an extent of the upper airway obstruction. 6. The system as claimed in claim 1 wherein the processor is adapted to derive an acoustic transfer function from the acoustic signals detected at the first and second locations and derive the parameter in a form of an acoustic impedance from the acoustic transfer function. 7. The system as claimed in claim 6, comprising a Fourier transform processing arrangement for processing the acoustic signals detected at the first and second locations before processing to derive the acoustic transfer function, and further comprising an averaging arrangement for averaging the ... | KONINKLIJKE PHILIPS N.V. | A61B007/00 A61B005/00 A61B005/085 A61M016/00 A61M016/06
US10722171 | B2 | 2020-07-28 | Device for measuring a plurality of parameters in patient subject to a treatment with radiopharmaceuticals | 2015-02-16 | A device for measuring a plurality of parameters of a patient subject to a treatment with radiopharmaceuticals includes a main body constrained to the patient's body in such a way that the measurement device is of wearable type. The device includes a plurality of measurement modules, for example a first module and a second module, each of which is equipped with a respective sensor, configured to measure a predetermined parameter of the patient. The device is also equipped with a microprocessor configured to process a plurality of data measured by the plurality of sensors, obtaining a plurality of processed data. Furthermore, a transmission means is provided of wireless type arranged to send via wireless communication the data processed by the microprocessor to a remote control unit. | 1. A multiparametric device for measuring a plurality of parameters of a patient subject to a treatment with radiopharmaceuticals, said device comprising: a main body constrained to the patient's body, in such a way that said device for measuring, in use, is worn by the patient;a plurality of measurement modules, each measurement module of said plurality equipped with a sensor configured to measure a predetermined parameter of said patient;a microprocessor configured to process a plurality of data measured by each of said sensors of said plurality of measurement modules, obtaining a plurality of processed data;a wireless transmission means arranged to send via wireless communication said data processed by said microprocessor to a remote control unit;a supplier arranged to supply said microprocessor and each sensor of said plurality;engagement means arranged to engage with respective engagement means of each measurement module of said plurality, said engagement means arranged to operatively connect said microprocessor with each sensor of said plurality;wherein at least one measurement module of said plurality is a measurement module of radiations equipped with a radiation sensor configured to measure said radiations in a predetermined time range (Dt);wherein said measurement module is configured to be arranged at a distance from the patient's body in such a way that said radiation sensor is arranged to measure the radiations emitted by the source of radiations from a fixed distance; andwherein said control unit is arranged to carry out a processing procedure of said processed data for determining the time of discharge (td) of said patient subject to said treatment with radiopharmaceuticals;wherein said processing procedure comprises the steps of:constructing a characteristic curve (g) representative of the number of radiations detected by said sensors of radiations in function of the time;determining a time (t) at which said characteristic curve (g) assumes a value equal to, or lower than, a predetermined lower threshold value (t*), said time (t) corresponding to the time of discharge of the patient;wherein said constructing step of said characteristic curve (g) provides the steps of:counting the number of pulses emitted by said radiation sensor when said radiations are detected in a predetermined time range (Δt1) obtaining the number of pulses in said time range (n(Δt1));identifying a point representative of said number of pulses (n(Δt1)) on a graph (n(Δt1)) versus time (t);repeating said counting step and said identifying step for a predetermined number (k) of time range (Dt1) with a delay one from the subsequent one of a second predetermined time range (Dt2) for an overall time of detection (Dttot), at the end of said repeating step being determined the number (k) of radioactive particles in a predetermined time range (ni(Dt1)), with i=1 . . . k and identified a corresponding number (k) of points (Pi) on said graph;constructing said characteris... | WINMEDICAL S.R.L. | A61B005/00 G06F019/00 G16H040/63 G16H050/50 A61B005/0205 A61B006/00 A61N005/10 A61B005/021 A61B005/024 A61B005/0402 A61B005/08 A61B005/11 A61B005/145
US10705486 | B2 | 2020-07-07 | Magic gluco-wrist watch (MGW) | 2015-10-11 | The Magic Gluco Wrist Watch (MGW) is a time wrist watch with a capability to measure the blood glucose level. The fix watch clasp plays a role of a lancing device, and is attached to the end of a left strap of the MGW. The movement of a holding member from the left strap to a right strap causes new electric current test strip to be driven in to the test strip port. A test strip port is located in the lug. By opening the valve, the MGW turns from time watch into a blood glucose meter. | 1. A Magic Gluco Wrist Watch (MGW) apparatus for measuring and monitoring glucose level comprising: a wristwatch comprising a main body/lug, a left and right watch straps each having a left and right slot nestled inside, extending along a length of left and right strap bodies; a display screen located on top of said wristwatch; at least one watch clasp; a holding member, multiple new test strips, at least one lancing device, wherein said test strips are located and move inside said left and right strap bodies; an arced shaped push to slide open door covering said multiple new test strips, wherein when said door is in a closed setting said display screen has a full circular shape/design and wherein when said door is open display screen has a crescent design, said push to slide open door is slid open with a counter clockwise motion and is slid shut and closed with a clockwise motion; said left strap having an opening/window displaying the holding member and at least two holding member jags, wherein said holding member comprises multiple frames/cartridges. 2. The apparatus of claim 1, wherein used test strips are held inside said right strap body and is not discarded after each use, when all of said new multiple test strips are completely used the used test strips are removed from said right watch strap through said right slot. 3. The apparatus of claim 2, wherein said wristwatch is a smart digital watch where time, date, location, weather forecast, emails or/and user data transmitted via communication protocols from a smart phone or computer to a laptop to said wristwatch and received via a receiver inside said main body, the user data is displayed on said display screen of a watch interface; and the monitored glucose level, time or other health related data are displayed on said display screen of a glucose monitoring interface; wherein said watch interface and glucose monitoring interface are interchangeable, and wherein said watch interface is on display at all times and wherein glucose monitoring interface is active when said push to slide door is slid open. 4. The apparatus of claim 3, wherein said display is touchscreen and data is entered and sent through said transmitting protocol via said wristwatch including said monitored glucose level. 5. The apparatus of claim 4 wherein said glucose measurement data is saved in a memory located inside said main body and is transmitted to said smart phone and/or laptop and/or computer and/or physician's office or emergency room, wherein said transmission is initiated manually or automatically at certain time intervals and/or continuously and/or as needed; and wherein said glucose measurement data is calculated and displayed based on programs run by a processing unit located inside said wristwatch. 6. The apparatus of claim 5, wherein said MGW comprises at least one alarm system, comprising at least one flashing or pulsating light, vibration and/or sound and/or any combination thereof, flashing and p... | - | G04G021/02 G01N033/487 G01N033/49 A61B005/145
US10722154 | B2 | 2020-07-28 | Medical photometer and medical photometer control method | 2017-10-25 | A medical photometer includes a processor and a memory that stores an instruction readable by a computer. When the instruction is executed by the processor, the medical photometer acquires a first light attenuation of a first light beam that is transmitted through or reflected from a tissue of a living body, and that has a first wavelength, based on a first intensity signal corresponding to an intensity of the first light beam, acquires a second light attenuation of a second light beam that is transmitted through or reflected from the tissue, and that has a second wavelength, based on a second intensity signal corresponding to an intensity of the second light beam, calculates a tissue thickness based on a ratio of the first light attenuation and the second light attenuation, and detects an artifact of the living body based on a temporal variation of the tissue thickness. | 1. A medical photometer comprising: a first light emitter that emits a first light beam having a first wavelength;a second light emitter that emits a second light beam having a second wavelength;a detector that outputs a first intensity signal corresponding to an intensity of the first light beam that is transmitted through or reflected from a tissue of a living body, and a second intensity signal corresponding to an intensity of the second light beam that is transmitted through or reflected from the tissue;a processor; anda memory that stores an instruction readable by a computer,wherein, when the instruction is executed by the processor, the medical photometer acquires a first light attenuation of the first light beam based on the first intensity signal, acquires a second light attenuation of the second light beam based on the second intensity signal, calculates a tissue thickness of the tissue of the living body based on a ratio of the first light attenuation and the second light attenuation, detects a temporal variation of an intensity ratio of the first intensity signal and the second intensity signal, cancels a temporal variation of the intensity ratio corresponding to pulsation of blood in the living body from the temporal variation of the intensity ratio, detects an artifact in the temporal variation of the intensity ratio corresponding to a temporal variation of the tissue thickness of the tissue of the living body, and calculates at least one of arterial oxygen saturation, a fraction of carboxyhemoglobin, a fraction of methemoglobin, and a concentration of total hemoglobin based on the artifact, both the first light beam and the second light beam are infrared light beams or red light beams, a ratio of a first extinction coefficient of a first blood light absorber at the first wavelength and a second extinction coefficient of the first blood light absorber at the second wavelength is different from 1, a third extinction coefficient of a second blood light absorber at the first wavelength is higher than the first extinction coefficient, a fourth extinction coefficient of the second blood light absorber at the second wavelength is higher than the second extinction coefficient, and a ratio of the third extinction coefficient and the fourth extinction coefficient is approximable by 1. 2. The medical photometer according to claim 1 further comprising a notifying section that performs notification when the detected artifact exceeds a threshold value. 3. The medical photometer according to claim 1 further comprising a third light emitter that emits a third light beam having a third wavelength, wherein a ratio of an extinction coefficient of blood at the first wavelength and an extinction coefficient of the blood at the third wavelength is varied in accordance with a concentration of the second blood light absorber contained in the blood,a ratio of an extinction coefficient of blood at the second wavelength and the extinction coefficient of th... | NIHON KOHDEN CORPORATION | A61B005/1455 A61B005/145 A61B005/00 G01N021/59 G01N021/27 G01N021/31 G01J001/04 A61B005/107
US10702158 | B2 | 2020-07-07 | Light-based collagen measurement system and a skin treatment system | 2015-08-28 | A measurement system is disclosed for light-based measurement of collagen, using a first light intensity and a second light intensity. The measurement system comprises a light source for emitting a light beam having a source wavelength range and an optical system to polarize light within the source wavelength range, thereby generating a polarized light beam, and to direct and focus the polarized light beam to a target position inside the skin at a predetermined focus depth below an outer surface of the skin. The measurement system further comprises a first detector, and a second detector for detecting the first light intensity and the second light intensity within, respectively, a first detection wavelength range and a second detection wavelength range of light reflected from the target position. The measurement system further comprises a comparator, coupled to the first detector and the second detector and configured and arranged to determine a difference between the first light intensity and the second light intensity. When natural collagen is present at the target position, a reflected spectrum of light of the polarized light beam reflected by said natural collagen comprises a plurality of adjacent first and second wavelength ranges, wherein in said first wavelength ranges constructive interference prevails between the polarized light beam and the natural collagen, and wherein in said second wavelength ranges destructive interference prevails between the polarized light beam and the natural collagen. | 1. A measurement system for light-based measurement of collagen inside skin, using a first light intensity and a second light intensity, the measurement system comprising: an optical system having a light source configured and arranged to emit a light beam having a source wavelength range;the optical system being configured and arranged:to polarize light over the entire source wavelength range, thereby generating a polarized light beam; andto direct and focus the polarized light beam to a target position inside the skin at a predetermined focus depth below an outer surface of the skin such that a direction or orientation of the polarization is the same throughout the entire source wavelength range;the measurement system further comprising:a first detector configured and arranged to directly detect the first light intensity within a first predetermined detection wavelength range of light reflected from the target position;a second detector configured and arranged to directly detect the second light intensity within a second predetermined detection wavelength range of light reflected from the target position; anda comparator, coupled to the first detector and the second detector and configured and arranged to directly determine a difference between the first light intensity and the second light intensity based on the directly detected first and second light intensities;wherein, with reference to a reflected spectrum of light of the polarized light beam reflected by natural collagen when present at the target position, said reflected spectrum comprising a plurality of adjacent first and second wavelength ranges, wherein in said first wavelength ranges constructive interference prevails between the polarized light beam and the natural collagen, and wherein in said second wavelength ranges destructive interference prevails between the polarized light beam and the natural collagen:the first detection wavelength range is predetermined and/or controlled to be confined to one or more of said first wavelength ranges; andthe second detection wavelength range is predetermined and/or controlled to be confined to one or more of said second wavelengths ranges. 2. The measurement system according to claim 1, wherein the first detection wavelength range is predetermined and/or controlled to be confined to one or more of said plurality of first wavelength ranges wherein constructive interference of second order or higher order prevails between the polarized light beam and the natural collagen, and wherein the second detection wavelength range is predetermined and/or controlled to be confined to one or more of said plurality of second wavelength ranges wherein destructive interference of second order or higher order prevails between the polarized light beam and the natural collagen. 3. The measurement system according to claim 1, wherein the optical system is configured and arranged to focus the polarized light beam at a target position between 100 microns and 10... | KONINKLIJKE PHILIPS N.V. | A61B005/00
US10713822 | B2 | 2020-07-14 | Tomographic imaging apparatus, method of controlling the same, and computer program product | 2018-09-28 | A tomographic imaging apparatus includes an X-ray detector comprising a plurality of dual mode pixels and configured to detect radiation that has passed through an object, and at least one processor configured to obtain scan data from the X-ray detector, and control each pixel of the plurality of dual mode pixels to operate in one of a first mode and a second mode, wherein each pixel of the plurality of dual mode pixels includes a sensor configured to generate a scan signal by converting incident radiation into an electric signal, a first signal path circuit configured to transmit the scan signal in the first mode, a second signal path circuit configured to transmit the scan signal in the second mode, and a photon counter configured to count photons from the scan signal transmitted through one of the first and second signal path circuits. | 1. A tomographic imaging apparatus comprising: an X-ray detector comprising a plurality of dual mode pixels and configured to detect radiation that has passed through an object; andat least one processor configured to obtain scan data from the X-ray detector, and control each pixel of the plurality of dual mode pixels to operate in one of a first mode and a second mode,wherein each pixel of the plurality of dual mode pixels comprises: a sensor configured to generate a scan signal by converting incident radiation into an electric signal;a first signal path circuit configured to transmit the scan signal in the first mode;a second signal path circuit configured to transmit the scan signal in the second mode; anda photon counter configured to count photons from the scan signal transmitted through one of the first and second signal path circuits,wherein a minimum time interval of transmitting the scan signal of the second signal path circuit is smaller than that of the first signal path circuit. 2. The tomographic imaging apparatus of claim 1, wherein the first signal path circuit comprises: a first preamplifier configured to amplify the scan signal; anda pulse shaper configured to shape a waveform of the scan signal. 3. The tomographic imaging apparatus of claim 1, wherein the second signal path circuit comprises: a second preamplifier configured to amplify the scan signal; anda reset circuit configured to reset input and output levels of the second preamplifier when the output level of the second preamplifier exceeds a first level. 4. The tomographic imaging apparatus of claim 1, wherein the at least one processor is further configured to generate a model selection signal for selecting one of the first mode and the second mode and output the mode selection signal to the X-ray detector, wherein each of the plurality of dual mode pixels further comprises: a first switch configured to transmit, to one of the first signal path circuit and the second signal path circuit, the scan signal output from the sensor, according to the mode selection signal; anda second switch configured to connect, to one of the first signal path circuit and the second signal path circuit, an input terminal of the photon counter, according to the mode selection signal. 5. The tomographic imaging apparatus of claim 1, wherein the X-ray detector comprises a plurality of binning pixels comprising a plurality of sub-pixels, wherein the plurality of dual mode pixels are grouped as sub-pixels to be included in the plurality of binning pixels, and the at least one processor is further configured to control at least one first dual mode pixel in each of the plurality of binning pixels to operate in the first mode and control remaining at least one second dual mode pixel excluding the at least one first dual mode pixel to operate in the second mode. 6. The tomographic imaging apparatus of claim 1, wherein the X-ray detector further comprises an anti-scatter grid provided on the plura... | Samsung Electronics Co., Ltd. | A61B005/00 A61B006/00 G06T011/00 G01T001/164 G01T001/24 H04N005/3745 H04N005/32 G01T001/20 G01T001/17 A61B006/02 G01J001/44 G01T001/29 G01J001/46 G01N023/046 G01V005/00
US10716474 | B2 | 2020-07-21 | Patient support apparatuses with location/movement detection | 2018-11-16 | A patient support apparatus includes a first transceiver adapted to wirelessly communicate with a second transceiver of a headwall interface that is positioned off of the patient support apparatus. A communication link is automatically established between the first and second transceivers without requiring a user of the patient support apparatus to activate a designated control and without requiring the user to identify the headwall interface. The first transceiver includes a unique identifier assigned to the headwall interface in its messages to the headwall interface. The first transceiver may also automatically transmit a disconnect signal to the headwall interface indicating the termination of the communication link is not accidental. The disconnect signal is sent based on one or more of the following: (1) a brake being off, (2) an A/C power cord being unplugged; and/or (3) a signal strength between the transceivers decreasing. | 1. A patient support apparatus comprising: a support surface adapted to support a person;a first transceiver adapted to establish first wireless communication with a first headwall unit positioned at a fixed location in a room;a second transceiver adapted to establish second wireless communication with a second headwall unit positioned at a fixed location in the room; anda controller adapted to automatically determine from the first and second wireless communications when the patient support apparatus is leaving the room by monitoring an order in which the first and second transceivers discontinue communication with the first and second headwall units, respectively. 2. The patient support apparatus of claim 1 wherein the controller is further adapted to automatically determine when the patient support apparatus is entering the room, and the controller does so by monitoring an order in which the first and second transceivers establish communication with the first and second headwall units, respectively. 3. The patient support apparatus of claim 1 wherein the first transceiver is an infrared transceiver and the second transceiver is a radio frequency transceiver. 4. A patient support apparatus comprising: a support surface adapted to support a person;a first transceiver adapted to establish first wireless communication with a first headwall unit positioned at a fixed location in a room;a second transceiver adapted to establish second wireless communication with a second headwall unit positioned at a fixed location in the room;a controller adapted to automatically determine from the first and second wireless communications at least one of the following: (1) when the patient support apparatus is leaving the room, and (2) when the patient support apparatus is entering the room; andwherein the first transceiver is adapted to communicate first periodic heartbeat messages with the first headwall unit, the second transceiver is adapted to communicate second periodic heartbeat messages with the second headwall unit, and the controller is adapted to forward information to a remote device indicative of successful communication of the first and second heartbeat messages. 5. The patient support apparatus of claim 4 wherein the patient support apparatus is a bed, the first transceiver is an infrared transceiver, and the second transceiver is a radio frequency transceiver. 6. A patient support apparatus comprising: a support surface adapted to support a person;a first transceiver adapted to establish first wireless communication with a first headwall unit positioned at a fixed location in a room;a second transceiver adapted to establish second wireless communication with a second headwall unit positioned at a fixed location in the room;a controller adapted to automatically determine from the first and second wireless communications at least one of the following: (1) when the patient support apparatus is leaving the room, and (2) when the patient support app... | Stryker Corporation | A61B005/00 H04L012/26 A61B005/11 A61G007/018 A61G007/00 A61G007/015 A61G007/005 A61G007/012
US10716431 | B2 | 2020-07-21 | Beverage and food temperature regulation | 2015-10-28 | A device for food and/or beverage temperature regulation includes a cooler for providing a cooling mode in which the food and/or beverage is cooled, a heater for providing a heating mode in which the food and/or beverage is heated, and a pre-heating mode in which the food and/or beverage is pre-heated. A controller is configured to control the cooler and heater. The controller is further configured to activate the cooling or heating mode in response to signals indicative of a change in the behavior of a recipient of the food and/or beverage, and to activate the pre-heating mode in response to predicted behavior. The signals may represent sound and/or movement and may originate from sensors, such as microphones and/or radar sensors. | 1. A device for temperature regulation of food and/or beverage, comprising: a cooler configured to provide a cooling mode in which the food and/or beverage is cooled;a heater configured to provide a heating mode in which the food and/or beverage is heated and for providing a pre-heating mode in which the food and/or the beverage is pre-heated;a memory configured to store behavioral pattern data of a recipient of the food and/or beverage;a processor configured to predict, using the behavioral pattern data, a future state of a behavior of the recipient of the food and/or beverage; anda controller for controlling the cooler and the heater, wherein the controller is configured to:receive at least one signal indicative of a current state of the behavior of the recipient of the food and/or beverage,detect a change in the current state of the behavior of the recipient of the food and/or beverage,activate the cooling mode or the heating mode in response to the detecting of the change in the current state of the behavior of the recipient of the food and/or beverage, andactivate the pre-heating mode in response to the future state of the behavior predicted by the processor, wherein the pre-heating mode is a mode preceding the heating mode. 2. The device according to claim 1, wherein the at least one signal comprises a sound signal representing sound produced by the recipient. 3. The device according to claim 1, wherein the at least one signal comprises a movement and/or posture signal representing movement and/or posture of the recipient. 4. The device according to claim 1, wherein the at least one signal comprises a vital sign signal representing at least one vital sign of the recipient. 5. The device according to claim 1, wherein the at least one signal comprises a moisture signal representing moisture loss of the recipient. 6. The device according to claim 1, wherein the controller is configured to weigh the at least one signal using the behavior predicted by the processor with a weighting factor. 7. The device according to claim 1, wherein the controller is further arranged to provide an inactive mode in response to the predicted future behavior of the recipient. 8. The device according to claim 1, wherein the controller is configured to selectively respond to sound signals. 9. The device according to claim 8, wherein the controller comprises audio filters and/or a sound recognizer. 10. The device according to claim 1, wherein the controller is configured to detect the change being greater than a threshold, and wherein the threshold is different for different behaviors of the recipient. 11. The device according to claim 1, further comprising at least one sensor configured to produce the at least one signal indicative of the current state of the behavior of the recipient of the food and/or beverage. 12. A method for temperature regulation of food and/or beverage using a cooler for providing a cooling mode in which the food and/or beverage is ... | KONINKLIJKE PHILIPS N.V. | A47J036/24 G05D023/19 A61B005/00
US10722132 | B2 | 2020-07-28 | Body-worn system for continuous, noninvasive measurement of cardiac output, stroke volume, cardiac power, and blood pressure | 2011-12-27 | The invention provides a system for measuring stroke volume (SV), cardiac output (CO), and cardiac power (CP) from a patient that features: 1) impedance sensor connected to at least two body-worn electrodes and including an impedance circuit that processes analog signals from the electrodes to measure an impedance signal (e.g. a TBEV waveform); 2) an ECG sensor connected to at least two chest-worn electrodes and including an ECG circuit that processes analog signals from the electrodes to measure and ECG signal; 3) an optical sensor connected to a body-worn optical probe and including an optical circuit that processes signals from the probe to measure at least one optical signal (e.g. a PPG waveform) from the patient; 4) a processing system, typically worn on the patient's wrist and connected through a wired interface to the optical sensor, and through either a wired or wireless interface to the TBEV and ECG sensors. | 1. A system for measuring cardiac power from a patient, comprising: an electrocardiogram (ECG) module comprising: i) at least two electrodes configured to be positioned on the patient's torso; ii) an ECG circuit configured to process signals from the at least two electrodes to measure at least one ECG waveform from the patient; and iii) a first processing system configured to process the ECG waveform to determine a heart rate value;an optical module comprising: i) an optical probe configured to be positioned on a digit on a hand of the patient; and ii) an optical circuit configured to process signals from the optical probe and to measure at least one photoplesthymograph (PPG) waveform from the patient;a trans-brachial electrovelocimetry (TBEV) module comprising: i) at least two TBEV electrodes; ii) a TBEV circuit, configured to process TBEV signals from the at least two TBEV electrodes to measure a TBEV waveform from the patient; and iii) a second processing system configured to process the TBEV waveform to determine a time dependent derivative of the TBEV waveform (d(TBEV)dZ/dt)), a first zero value in d(TBEV)dZ/dt indicating onset of a pulse in d(TBEV)dZ/dt representing opening of the patient's aortic valve, a second zero value in d(TBEV)dZ/dt after onset of the pulse indicating closing of the patient's aortic valve, a systolic flow time (SFT) calculated using the first zero value in d(TBEV)dZ/dt and the second zero value in d(TBEV)dZ/dt, a d(TBEV)dZ/dtmax value, and a stroke volume value calculated using the SFT and the d(TBEV)dZ/dtmax value;a blood pressure module configured to: i) receive the ECG waveform from the ECG module; ii) receive the PPG waveform from the optical module; iii) receive the TBEV waveform from the impedance module; and iv) collectively process at least two of these waveforms to determine a blood pressure value;a third processing system configured to be worn on the patient's body, wherein the third processing system is configured to:i) receive the ECG waveform from the ECG module; ii) receive the TBEV waveform from the TBEV module; iii) receive the PPG waveform from the optical module; iii) collectively process at least two of these to determine a transit time value and a blood pressure value; iv) receive the stroke volume value from the TBEV module; v) receive the heart rate value from ECG module; vi) receive the blood pressure value from the blood pressure module; and vii) collectively process the stroke volume value, heart rate value, and blood pressure value to determine a cardiac power value; and,wherein the third processing system comprises a wireless system, configured to transmit the stroke volume value, the cardiac output value, and the cardiac power value to a remote location. 2. The system of claim 1, wherein the second processing system is further configured to determine a second stroke volume value based on estimating a pulse pressure value from the TBEV waveform. 3. The system of claim 1, wherein the secon... | SOTERA WIRELESS, INC. | A61B005/0295 A61B005/029 A61B005/053 A61B005/08 A61B005/11 A61B005/1455 A61B005/00 A61B005/01 A61B005/021 A61B005/02 A61B005/0205 A61B005/145 G06F019/00 A61B005/024 A61B005/0245
US10722124 | B2 | 2020-07-28 | Apparatus and method for physiological mechanical and electrical activity monitoring | 2016-04-11 | The invention describes a method for extracting several physiological parameters such as cardiac cycle parameters with one single measurement over a time period and for monitoring mechanical and electrical cardiac activity. The invention finds its use in providing continuous monitoring of heart mechanical and electrical activities over time. Such signal may be transmitted to clinics of family doctors to keep track of the subject's health. | 1. A method of using piezoelectric material to monitor a physiological activity of a subject and generate a physiological electrical signal corresponding to a physiological parameter the method comprising: coupling the piezoelectric material to a body part of the subject;obtaining a first piezoelectric electrical signal from the piezoelectric material, the first piezoelectric electrical signal is based on mechanical movement of the body part, the mechanical movement related to the physiological activity of the subject;mapping a segment of the first piezoelectric electrical signal to a conventional electrical signal obtained using conventional techniques for the physiological parameter of the subject, the conventional electrical signal is obtained only once and is obtained simultaneously while obtaining the segment of the first piezoelectric electrical signal, the segment of the first piezoelectric electrical signal and the conventional electrical signal are obtained for at least one full cycle of activity for the physiological parameter, the mapping is performed using signal processing techniques; the signal processing techniques comprise generating a transfer function representative of the physiological parameter; andgenerating solely from a subsequent segment of the first piezoelectric electrical signal a first physiological electrical signal corresponding to the physiological parameter based on the transfer function. 2. A method according to claim 1, where the mapping of the first piezoelectric electric signal to the conventional electrical signal is performed using a linear one-to-one mapping. 3. A method according to claim 1, the method further comprising storing the transfer function on a memory storage device. 4. A method according to claim 3, the method further comprising: storing in the memory storage device a plurality of conventional electrical signals obtained from the conventional techniques, wherein each of the plurality of conventional electrical signals corresponds to a different known physiological defect of the physiological parameter. 5. A method according to claim 4, wherein the method further comprises: comparing the first physiological electrical signal to each of the plurality of conventional electrical signals; andassessing whether the subject has a physiological defect of the physiological parameter based on the comparison. 6. A method according to claim 1, the method further comprising: obtaining a second piezoelectric electrical signal different from the first piezoelectric electrical signal, the second piezoelectric electrical signal based on a second mechanical movement of the body part at a time after obtaining the first piezoelectric electrical signal, the second mechanical movement related to the physiological activity of the subject;manipulating the second piezoelectric electrical signal by convolving the second piezoelectric electrical signal with the transfer function; andgenerating solely from the manipula... | United Arab Emirates University | A61B005/021 A61B005/029 A61B007/04 A61B005/00 A61B005/0205 A61B005/0245 A61B005/024 A61B005/11
US10722127 | B2 | 2020-07-28 | Photoplethysmography device | 2016-02-26 | The invention relates to a photoplethysmography device (20) of the reflectance type, comprising a light source (4), a light sensor (5), and an interface layer (21). The interface layer (21) has a recess between the source (4) and the sensor (5) in order to prevent reflection losses between the device (20) and the skin and to prevent that light from the source (4) can reach the sensor (5) directly via the interface layer (21). | 1. A photoplethysmography device of the reflectance type, comprising: at least one light source;at least one light sensor;at least one barrier wall; andan interface layer spaced apart from and covering the at least one light source and the at least one light sensor, the interface layer arranged for contacting the skin of a subject upon measurement, wherein the interface layer includes a recess between the at least one light source and the at least one light sensor, the recess preventing a light path between the at least one light source and the at least one light sensor within the interface layer. 2. The photoplethysmography device according to claim 1, wherein the interface layer is provided, on a face opposite to the at least one light source, and the at least one light sensor, with at least one projection adjacent to the recess. 3. The photoplethysmography device according to claim 1, wherein the recess is an elongated aperture or slit. 4. The photoplethysmography device according to claim 1, wherein the interface layer includes a plate of thermoplastic transparent to the light used for photoplethysmography. 5. The photoplethysmography device according to claim 1, wherein the interface layer includes or consists of silicone or rubber arranged for contacting the skin of the subject. 6. The photoplethysmography device according to claim 1, wherein the at least one light source and/or the at least one light sensor are provided with an encapsulation, the encapsulation of the at least one light source and/or the encapsulation of the at least one light sensor contacting the interface layer. 7. The photoplethysmography device according to claim 1, further comprising a barrier wall arranged between the at least one light source and the at least one light sensor preventing a light path between the at least one light source and the at least one light sensor between the at least one light source, the at least one light sensor and the interface layer. 8. A photoplethysmography device of the reflectance type, comprising: a light source and a light sensor arranged on a base;a barrier wall arranged between the light source and the light sensor;an interface layer spaced apart from the light source, the light sensor, and the base, the interface layer arranged for contacting the skin of a subject upon measurement; anda recess arranged in the interface layer in line with the barrier wall, the recess preventing a light path between the light source and the light sensor. 9. The photoplethysmography device according to claim 8, wherein the barrier wall extends between the base and the interface layer. 10. A photoplethysmography device of the reflectance type, comprising: a first side arranged to contact the skin of a subject;a second side, opposite the first side;at least one light source arranged on the second side;at least one light sensor arranged on the second side; andan interface layer including a recess arranged on the first side, wherein the interf... | KONINKLIJKE PHILIPS N.V. | A61B005/00 A61B005/024
US10722704 | B2 | 2020-07-28 | Implantable medical systems and methods including pulse generators and leads | 2018-05-07 | Methods for implanting a puke generator (PG) within a pectoral region of a chest of a patient and devices having the PG. The PG has a housing that includes a PG electrode. Methods also include implanting at least one lead having first and second electrode segments with the first electrode segment positioned along an anterior of the chest of the patient and the second electrode segment positioned along at least one of a posterior of the patient or a side of the patient. The first and second electrode segments are positioned subcutaneously at or below an apex of a heart of the patient, wherein the PG electrode and the first and second electrode segments are configured to provide electrical shocks for antiarrhythmic therapy. | 1. A method comprising: implanting a pulse generator (PG) within a pectoral region of a chest of a patient, the PG having a housing that includes a PG electrode; andimplanting at least one lead having first and second electrode segments with the first electrode segment positioned along an anterior of the chest of the patient and the second electrode segment positioned along at least one of a posterior of the patient or a side of the patient, the first and second electrode segments being positioned subcutaneously at or below an apex of a heart of the patient, wherein the PG electrode and the first and second electrode segments are configured to provide electrical shocks for antiarrhythmic therapy, wherein implanting the second electrode segment includes making a tunnel that extends along an intercostal gap between a first access incision proximate to a xiphoid process and a second access incision that is positioned at or beyond a midaxillary line. 2. The method of claim 1, wherein the first and second electrode segments are portions of a common lead coil that extends from a proximal end of the first electrode segment to a distal end of the second electrode segment, the common lead coil extending along the chest of the patient from the anterior of the chest toward the side of the patient, wherein the second electrode segment extends to or beyond a midaxillary line of the patient. 3. The method of claim 2, wherein an active length of the common lead coil is at least 35 cm from the proximal end of the first electrode segment to the distal end of the second electrode segment. 4. The method of claim 1, wherein the first and second electrode segments are spaced apart from one another, the second electrode segment being longer than the first electrode segment. 5. The method of claim 4, wherein the second electrode segment has an active length measured between proximal and distal ends of the second electrode segment, the active length being at least 12 centimeters (cm), wherein the second electrode segment extends beyond a posterior axillary line. 6. The method of claim 1, wherein implanting the at least one lead includes positioning one lead such that the one lead extends away from the pulse generator, extra-thoracically along a sternum, and along an intercostal gap so that a distal end of the one lead is positioned at or beyond a midaxillary line of the patient. 7. The method of claim 1, wherein implanting the first electrode segment includes tunneling from a pocket where the pulse generator is positioned along a sternum of the patient to an intercostal gap at or below the apex of the heart. 8. The method of claim 1, wherein implanting the pulse generator includes making a subcutaneous pocket and positioning the pulse generator within the subcutaneous pocket. 9. The method of claim 1, wherein the first and second electrode segments are portions of a continuous elongated body that wraps about the patient, wherein implanting the first electrode se... | Pacesetter, Inc. | A61N001/05 A61N001/39 A61B005/0245
US10709362 | B2 | 2020-07-14 | Analyte sensor | 2019-11-21 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A system for measuring a glucose concentration in a host, the system comprising: a transcutaneous glucose sensor configured to provide sensor data indicative of a glucose concentration in a host;sensor electronics configured to operatively connect to the transcutaneous analyte sensor; anda processor configured to calibrate the sensor data based on predetermined calibration information, wherein the predetermined calibration information is obtained prior to implantation, wherein the predetermined calibration information is derived at least in part from a batch of sensors meeting a predetermined limit of sensor sensitivity variation, wherein the predetermined calibration information comprises one or both of predetermined baseline information and predetermined sensitivity information associated with the glucose sensor, and wherein the processor is configured to calibrate the sensor data without requiring reference data from a reference glucose monitor. 2. The system of claim 1, wherein the calibration is based on a conversion function comprising a baseline value and a sensitivity value, and wherein the processor is configured to adjust the baseline value based on the predetermined baseline information. 3. The system of claim 1, wherein the predetermined sensitivity information is derived from in vitro testing and is provided as a calibration code. 4. The system of claim 1, further comprising: a transmitter unit operably coupled to the glucose sensor, the transmitter unit having a radio frequency transmitter configured to wirelessly transmit sensor data; anda receiver unit comprising a radio frequency receiver and a display, the radio frequency receiver configured to receive the transmitted sensor data and the display configured to display information based on the transmitted sensor data. 5. A method for processing data from a transcutaneous glucose sensor, the method comprising: receiving sensor data from the transcutaneous glucose sensor after the transcutaneous glucose sensor is implanted in a host, wherein the sensor data is indicative of a glucose concentration in the host;receiving predetermined calibration information, wherein the predetermined calibration information is derived at least in part from a batch of sensors meeting a predetermined limit of sensor sensitivity variation, wherein the predetermined calibration information comprises one or both of predetermined baseline information and predetermined sensitivity information associated with the transcutaneous glucose sensor; andcalibrating, using a processor, the sensor data received from the transcutaneous glucose sensor based at least in part on the predetermined calibration information, wherein the calibrating is performed without reliance on reference data from a reference glucose monitor. 6. The method of claim 5, wherein the calibrating is based on a conversion function comprising a sensitivity value and a baseline value, and wherein the method further comprises validating th... | DexCom, Inc. | A61B005/1473 A61B005/1495 A61B005/1486 A61B005/1468 A61B005/145 A61B005/00 A61B005/15 A61B005/05 A61B017/34 A61M005/172 A61M005/158 A61M005/142
US10716511 | B2 | 2020-07-21 | Subcutaneous device for monitoring and/or providing therapies | 2019-11-11 | A subcutaneously implantable device includes a housing, a clip attached to a top side of the housing that is configured to anchor the device to a muscle, a bone, and/or a first tissue, and a first prong with a proximal end attached to the housing and a distal end extending away from the housing that is configured to contact a first lung. A first electrode on the device is configured to contact the first lung, and a second electrode on the device is configured to contact the first lung or a second lung. Sensing circuitry in the housing in electrical communication with the first electrode and the second electrode is configured to measure an impedance in the first lung and/or the second lung, and/or a transthoracic impedance across the first lung and the second lung. | 1. A subcutaneously implantable device comprising: a housing;a clip attached to a top side of the housing that is configured to anchor the device to a muscle, a bone, and/or a first tissue;a prong and electrode structure comprising: a first prong with a proximal end attached to the housing and a distal end extending away from the housing that is configured to contact a first lung;a first electrode on the first prong that is configured to contact the first lung; anda second electrode on the first prong that is configured to contact the first lung;ora first prong with a proximal end attached to the housing and a distal end extending away from the housing that is configured to contact a first lung;a second prong with a proximal end attached to the housing and a distal end extending away from the housing that is configured to contact a second lung;a first electrode on the first prong that is configured to contact the first lung; anda second electrode on the second prong that is configured to contact the second lung; andsensing circuitry in the housing in electrical communication with the first electrode and the second electrode that is configured to measure an impedance in the first lung and/or the second lung, and/or a transthoracic impedance across the first lung and the second lung. 2. The device of claim 1, wherein the clip is configured to attach the device to a xiphoid process and/or a sternum of a patient. 3. The device of claim 1, wherein the clip further comprises: a top portion;a bottom portion; anda spring portion extending between and connecting the top portion to the bottom portion, wherein the spring portion is curved and is configured to act as a spring for the clip to push the top portion of the clip onto the bone, the muscle, and/or the first tissue to which it is anchored. 4. The device of claim 1, wherein the clip is configured to be positioned around the muscle, the bone, and/or the first tissue to anchor the device to the muscle, the bone, and/or the first tissue without piercing the muscle, the bone, and/or the first tissue. 5. The device of claim 1, wherein the first electrode and the second electrode are positioned on the first prong and are configured to contact the first lung. 6. The device of claim 1, wherein the prong and electrode structure comprises: the first prong with its proximal end attached to the housing and its distal end extending away from the housing that is configured to contact the first lung;the second prong with its proximal end attached to the housing and its distal end extending away from the housing that is configured to contact the second lung;the first electrode on the first prong that is configured to contact the first lung; andthe second electrode on the second prong that is configured to contact the second lung. 7. The device of claim 6, and further comprising: a third electrode on the first prong that is configured to contact the first lung; anda fourth electrode on the second prong that is ... | Manicka Institute LLC | A61B005/00 A61N001/05 A61B005/0205 A61M005/142
US10722162 | B2 | 2020-07-28 | Advanced analyte sensor calibration and error detection | 2019-09-27 | Systems and methods for processing sensor data and self-calibration are provided. In some embodiments, systems and methods are provided which are capable of calibrating a continuous analyte sensor based on an initial sensitivity, and then continuously performing self-calibration without using, or with reduced use of, reference measurements. In certain embodiments, a sensitivity of the analyte sensor is determined by applying an estimative algorithm that is a function of certain parameters. Also described herein are systems and methods for determining a property of an analyte sensor using a stimulus signal. The sensor property can be used to compensate sensor data for sensitivity drift, or determine another property associated with the sensor, such as temperature, sensor membrane damage, moisture ingress in sensor electronics, and scaling factors. | 1. An analyte sensor system comprising: a transcutaneous electrochemical analyte sensor comprising: a first portion configured to remain ex vivo during a sensor session; anda second portion configured to be disposed in vivo during the sensor session, wherein the second portion comprises an electrode configured to generate a signal associated with an analyte concentration; anda membrane disposed over at least a portion of the electrode; andsensor electronics operably connectable to the transcutaneous electrochemical analyte sensor, wherein the sensor electronics comprise a memory storing information associated with sensor sensitivity, wherein the information associated with sensor sensitivity is based at least in part on information based on a measurement of at least one sensor characteristic of the transcutaneous electrochemical analyte sensor during or after manufacturing of the transcutaneous electrochemical analyte sensor but prior to a sensor session;wherein the measured at least one sensor characteristic is associated with a thickness of the membrane. 2. The analyte sensor system of claim 1, wherein the sensor electronics are configured to process sensor data measured by the transcutaneous electrochemical analyte sensor to generate displayable sensor information. 3. The analyte sensor system of claim 1, further comprising a display device. 4. The analyte sensor system of claim 1, wherein the transcutaneous electrochemical analyte sensor is comprises a wire. 5. The analyte sensor system of claim 1, wherein the analyte is glucose. 6. The analyte sensor system of claim 1, further comprising a temperature sensor configured to measure an ex vivo temperature, wherein the analyte sensor system, during processing of the signal, is configured to perform temperature compensation by accounting for the measured ex vivo temperature. 7. The analyte sensor system of claim 1, wherein the sensor electronics are configured to process the signal. 8. The analyte sensor system of claim 7, wherein the sensor electronics are configured to process the signal based at least in part on the information associated with sensor sensitivity. 9. The analyte sensor system of claim 7, further comprising a receiver, wherein the receiver is configured to further process the processed signal based at least in part on the information associated with sensor sensitivity. 10. The analyte sensor system of claim 7, wherein the sensor electronics are configured to transmit the processed signal and the information associated with sensor sensitivity to a receiver for further processing to determine an estimated analyte value. 11. The analyte sensor system of claim 10, wherein the receiver is a handheld device, and wherein the handheld device is a smart phone. 12. The analyte sensor system of claim 1, wherein the membrane further comprises a polymer having a mediator. 13. The analyte sensor system of claim 1, wherein the analyte sensor system is configured to calibrate the tr... | DexCom, Inc. | G08B021/20 A61B005/1495 G01N027/327 A61B005/145 A61B005/1486 A61B005/00 G01D018/00 G01N027/02 G01N033/49 A61B005/1473
US10709873 | B1 | 2020-07-14 | Device for holding a catheter in place | 2017-02-04 | A device which is used in place of known catheters, such as the Foley catheter, in tests such as a Urodymics test. The device holds the catheter in place. | 1. A medical device configured for holding a catheter comprising: a body havinga conical portion having an apex, a base, a length extending between the apex and the base, a first end, a second end, and an inside surface,a first adhesive on the inside surface of the conical portion adjacent to the first end of the conical portion,a hollow cylindrical portion having an outer end, a second end which is unitary with the apex of the conical portion, a length which extends between the outer end and the second end of the cylindrical portion, an inside surface, and a hollow bore which extends from the outer end to the second end of the cylindrical portion, the inside surface of the cylindrical portion having an inner diameter,a pull tab having a first portion mounted on the inside surface of the cylindrical portion adjacent to the outer end of the cylindrical portion, the pull tab further having a pull portion located outside the hollow bore;a second adhesive releasably mounting the pull tab to the inside surface of the cylindrical portion, the second adhesive being spaced apart from the first adhesive on the conical portion so that the inside surface of the conical portion is free of adhesive between the adhesive adjacent to the first end of the conical portion and the cylindrical portion, the second adhesive including first and second portions, the second adhesive first portion being secured to the pull tab and the second adhesive second portion being located on the inside surface of the bore of the hollow cylindrical portion, the second adhesive first portion having adhesive properties selected so that the at least a portion of the second adhesive second portion remains in place on the inside surface of the cylindrical portion when the pull tab is separated from the cylindrical portion;a catheter having an outside surface having an outer diameter,the inside diameter of the hollow bore of the cylindrical portion being larger than the outer diameter of the catheter so that a gap is defined between the outside surface of the catheter and the inside surface of the cylindrical portion when the catheter is first inserted through the hollow bore;the second adhesive second portion remaining on the inside surface of the cylindrical portion of the body having adhesive properties to releasably secure the cylindrical portion to the outside surface of the catheter when the catheter is in abutting contact with the inside surface of the cylindrical portion of the body. 2. A medical device for holding a catheter comprising: a catheter having an outside surface and an outer diameter of the outside surface of the catheter; a device for holding a catheter, the device including a body havinga conical portion having an apex, a base, a length extending between the apex and the base, a first end, a second end and an inside surface,a first adhesive on the inside surface of the conical portion adjacent to the first end, anda hollow cylindrical portion having an outer end, a ... | - | A61B005/03 A61M025/01 A61M025/02 A61M039/10 A61M025/00 A61B005/00
US10702215 | B2 | 2020-07-07 | Systems and methods for dynamically and intelligently monitoring a host's glycemic condition after an alert is triggered | 2019-11-05 | Systems and methods for providing sensitive and specific alarms indicative of glycemic condition are provided herein. In an embodiment, a method of processing sensor data by a continuous analyte sensor includes: evaluating sensor data using a first function to determine whether a real time glucose value meets a first threshold; evaluating sensor data using a second function to determine whether a predicted glucose value meets a second threshold; activating a hypoglycemic indicator if either the first threshold is met or if the second threshold is predicted to be met; and providing an output based on the activated hypoglycemic indicator. | 1. A method of activating hypoglycemic indicators based on continuous glucose sensor data, the method comprising: evaluating sensor data using a first function to determine whether a real time glucose value meets one or more user settable first criteria;evaluating sensor data using a second function to determine whether a parameter indicative of glucose value meets one or more non user settable second criteria;providing an output including a first hypoglycemic indicator if the one or more user settable first criteria is met, and providing an output including a second hypoglycemic indicator if the one or more non user settable second criteria are met. 2. The method of claim 1, wherein the evaluating sensor data using a first function to determine whether a real time glucose value meets one or more user settable first criteria comprises determining whether the real-time glucose value passes a glucose threshold. 3. The method of claim 2, wherein the evaluating sensor data using a first function to determine whether a real time glucose value meets one or more user settable first criteria further comprises determining whether an amplitude of rate of change or direction of rate of change meets a rate of change criterion. 4. The method of claim 1, wherein the evaluating sensor data using a first function to determine whether the real time glucose value meets one or more user settable first criteria comprises evaluating a static risk of a substantially real time glucose value. 5. The method of claim 1, wherein the evaluating sensor data using a second function to determine whether the parameter indicative of glucose value meets one or more second criteria comprises evaluating a dynamic risk of the predicted glucose value. 6. The method of claim 1, wherein the second function comprises an artificial neural network model that uses at least one of exercise, stress, illness or surgery to determine the parameter indicative of glucose value, and wherein the parameter is a predicted glucose value. 7. The method of claim 1, wherein the second function utilizes a first order autoregressive model to determine the parameter indicative of glucose value, and wherein the parameter is a predicted glucose value. 8. The method of claim 7, wherein the first order autoregressive model comprises a parameter alpha, and wherein alpha is estimated recursively each time a sensor data point is received. 9. The method of claim 1, wherein evaluating the sensor data using the first function and the second function allows for increased warning time of hypoglycemic alerts without substantially increasing the number of alerts as compared to evaluating the sensor data using the first function without evaluating the sensor data using the second function. 10. The method of claim 1, wherein the one or more user settable first criteria comprises a first threshold that is configured to be user settable. 11. The method of claim 10, wherein the one or more non user settable second c... | DexCom, Inc. | A61B005/00 A61B005/145 G06F019/00
US10716495 | B1 | 2020-07-21 | Accelerometer-based gait analysis | 2017-02-16 | Embodiments of systems and methods determine one or more gait parameters using acceleration readings generated by an accelerometer coupled to a user. Gait parameters determined by various embodiments may include step length, step width, step time, single limb support time, or double limb support time. The accelerometer may be coupled to one of the user's arms. Exemplary systems may comprise an accelerometer, a processor, and a non-transitory computer readable medium (NTCRM). The processor may process acceleration readings from the accelerometer according to instructions stored on the NTCRM to determine the one or more gait parameters. The processor may communicate the one or more gait parameters to a receiving device. Some embodiments of the system may function as self-contained units and some embodiments may present information corresponding to gait parameters or recommendations regarding gait. | 1. An apparatus for determining a double limb support time of a user, the apparatus comprising an accelerometer, at least one processor, and at least one memory including instructions, the at least one memory and the instructions configured to, with the processor, cause the apparatus to at least: receive a plurality of acceleration readings, wherein the plurality of acceleration readings is generated by the accelerometer measuring acceleration of the user's arm;identify a first time at which a first reading of the plurality of acceleration readings was generated, the first reading being a lowest local minimum reading generated during a first window;identify a second time at which a second reading of the plurality of acceleration readings was generated, the second reading being a first occurring reading to exceed a threshold after the first reading;determine the double limb support time by determining a difference between the first time and the second time, wherein the double limb support time is indicative of at least a portion of a step in which the user's weight is at least partially supported on both limbs; andsend information corresponding to the double limb support time to a recipient device for display by the recipient device. 2. The apparatus of claim 1, wherein the at least one memory and the instructions are further configured to, with the processor, cause the apparatus to at least determine a fundamental frequency using an autocorrelation function applied to a plurality of readings of the plurality of acceleration readings and in which a duration of the first window is a fraction of the fundamental frequency. 3. The apparatus of claim 1, wherein the at least one memory and the instructions are further configured to, with the processor, cause the apparatus to at least compare double limb support times of one or more even numbered steps to double limb support times of one or more odd numbered steps to determine a measure of symmetry and send the measure of symmetry to the recipient device. 4. The apparatus of claim 1, wherein the at least one memory and the instructions are further configured to, with the processor, cause the apparatus to at least determine a variability measure of a plurality of double limb support times and send information corresponding to the variability measure to the recipient device. 5. The apparatus of claim 1, wherein the recipient device comprises at least one of: a non-transitory computer-readable medium;a device communicatively coupled to the processor using a network;an electromagnetic transmitter; ora user interface. 6. The apparatus of claim 5, wherein (a) the recipient device comprises the user interface, and (b) the user interface comprises one or more of: a visual display, an audio speaker, or a tactile output. 7. The apparatus of claim 1, wherein the processor is powered by one or more of: a battery, a photovoltaic cell, a body heat ambient energy collector, or a body motion ambient energy collect... | Fortify Technologies, LLC | A61B005/11 A61B005/00 G06F001/16
US10709903 | B2 | 2020-07-14 | Gantry image guided radiotherapy system and related treatment delivery methods | 2019-11-06 | A radiation treatment apparatus, comprising a gantry structure comprising a beam member extending between first and second ends of the gantry structure. The radiation treatment apparatus also includes a radiation treatment head movably mounted to the beam member in a manner that allows (i) translation of the radiation treatment head along the beam member between the first and second ends, and (ii) gimballing of the radiation treatment head relative to the beam member, the gimballing comprising pivotable movement in at least one independent pivot direction defined with respect to the beam member. The radiation treatment apparatus also includes a patient couch operative coupled with the radiation treatment head in manner to provide movement of the patient couch relative to the radiation treatment head. | 1. A radiation treatment apparatus, comprising: a gantry structure; and a radiation treatment head movably mounted to the gantry structure to allow gimballing of the radiation treatment head relative to the gantry structure, the gimballing comprising pivotable movement in at least one independent pivot direction defined with respect to the gantry structure; and a patient couch operatively coupled with the radiation treatment head to provide movement of the patient couch relative to the radiation treatment head; wherein the radiation treatment head has pivotable movement in a maximum angle range greater than +/−30 degrees. 2. The radiation treatment apparatus of claim 1, wherein the patient couch and the radiation treatment head are movable to provide non-coplanar radiation treatment of a tissue volume. 3. The radiation treatment apparatus of claim 2, further comprising a controller operatively coupled with the patient couch and the radiation treatment head, where to provide non-coplanar radiation treatment of a tissue volume the controller configured to control movement of the patient couch and gimbaling of the radiation treatment head to position a radiation beam generated by the radiation treatment head at the tissue volume. 4. The radiation treatment apparatus of claim 1, wherein the patient couch is configured to manipulate the patient couch with three or more degrees of freedom. 5. The radiation treatment apparatus of claim 4, where the three or more degrees of freedom comprise three orthogonal translations and one rotation. 6. The radiation treatment apparatus of claim 1, where a transverse isocentric plane being defined that passes through the isocenter in a direction orthogonal to the rotation axis, wherein the pivotable movement is in at least two independent pivot directions including a first pivot direction around a first pivot axis generally parallel to the transverse isocentric plane and a second pivot direction around a second pivot axis nonparallel to the first pivot axis. 7. The radiation treatment apparatus of claim 1, wherein the gantry structure is a rotatable gantry structure. 8. The radiation treatment apparatus of claim 1, wherein the gimballing comprises pivotable movement in at least two independent pivot directions defined with respect to a beam member. 9. A method of operating a radiation treatment apparatus, comprising: translating a radiation treatment head movably mounted to a gantry structure: gimballing, by a processing unit, the radiation treatment head relative to the gantry structure, the gimballing comprising pivotably moving the radiation treatment head in at least one independent pivot direction defined with respect to the gantry structure; and moving, by the processing unit, a patient couch relative to the radiation treatment head: wherein gimballing comprises pivotably moving the radiation treatment head with a maximum angle greater than +/−30 degrees. 10. The method of claim 9, wherein moving the p... | ACCURACY INCORPORATED | A61N005/10 A61B006/00 A61B006/06 A61B005/00 A61B005/055 A61B006/02 A61B006/03
US10708191 | B2 | 2020-07-07 | Systems and methods for performing electrophysiology (EP) signal processing | 2019-09-26 | Systems, methods, and computer program product embodiments are disclosed for performing electrophysiology (EP) signal processing. An embodiment includes an electrocardiogram (ECG) circuit board configured to process an ECG signal. The embodiment further includes a plurality of intracardiac (IC) circuit boards, each configured to process a corresponding IC signal. The embodiment further includes a communications interface communicatively coupled to a remote device, and a processor, coupled to the ECG circuit board, the plurality of IC circuit boards, and the communications interface. The processor is configured to receive, via the communications interface, feedback from the remote device. The processor is further configured to control, via the communication interface, the remote device based on the ECG signal, the IC signals, or the feedback from the remote device. | 1. A system for performing electrophysiology (EP) signal processing, comprising: an electrocardiogram (ECG) circuit board configured to receive an ECG signal, wherein the ECG circuit board comprises a large-signal detection circuit and a differential signal amplification stage, wherein the differential signal amplification stage is configured to amplify the ECG signal to output an output ECG signal and the large-signal detection circuit is coupled to the differential signal amplification stage and configured to remove signal saturation from the output ECG signal based on a time constant, wherein the time constant specifies a length of time that the output ECG signal is at maximum amplitude;a plurality of intracardiac (IC) circuit boards, each configured to receive a corresponding IC signal;a communications interface coupled to a remote device; anda processor, coupled to the ECG circuit board, the plurality of IC circuit boards, and the communications interface, configured to: receive, via the communications interface, feedback from the remote device; andcontrol, via the communication interface, the remote device based on the ECG signal, the IC signals, or the feedback from the remote device. 2. The system of claim 1, wherein the remote device is selected from the group consisting of an ultrasound machine, a radio frequency (RF) generator, a stimulator, a three-dimensional imaging device, an intra-cardiac echocardiography (ICE) machine, a fluoroscopy machine, and a defibrillator. 3. The system of claim 1, wherein the communications interface is coupled to the remote device using a communication protocol selected from the group consisting of Digital Imaging and Communications in Medicine (DICOM), Ethernet, Universal Serial Bus (USB), and Institute of Electrical and Electronics Engineers (IEEE) 802.11. 4. The system of claim 1, wherein the ECG circuit board further comprises a radio frequency filter circuit configured to attenuate an amplitude of the ECG signal between about 300 kHz and about 600 kHz. 5. The system of claim 4, wherein the ECG circuit board further comprises a low-frequency feedback circuit coupled to the radio frequency filter circuit, wherein the low-frequency feedback circuit is configured to drive a voltage of a reference node of the radio frequency filter circuit to increase input impedance of signal frequencies of the ECG signal such that the radio frequency filter circuit is configured as an open circuit at the signal frequencies of the ECG signal. 6. The system of claim 1, wherein the ECG circuit board comprises an input protection circuit configured to shunt a voltage of the ECG signal greater than or equal to 300 V. 7. The system of claim 6, wherein the ECG circuit board comprises a radio frequency filter circuit configured to attenuate an amplitude of the ECG signal between about 300 kHz and about 600 kHz. 8. The system of claim 7, wherein the ECG circuit board further comprises a low-frequency feedback circuit coup... | BioSig Technologies, Inc. Mayo Foundation for Medical Education and Research | A61K031/7105 H04L012/863 H04L012/26 A61B005/00 A61B005/0402 G16H040/63 A61B005/0428 H03F003/45 H03K005/125 A61B005/0215 A61B005/0245 A61B005/046 A61B005/0464 A61B018/14 A61B018/00 H02H009/04 H03F003/68 A61B005/024 A61B005/0456 A61B005/0472 A61B005/053
US10702132 | B2 | 2020-07-07 | System and method for using a capsule device | 2019-09-09 | The present invention discloses a method to move a magnetic capsule linearly both horizontally and linearly in a very controllable fashion, wherein the direction of a combined external magnetic field, the force experienced by the magnetic capsule, the movement direction of the capsule and magnetic direction of the magnetic capsule can be all aligned in parallel to each other. | 1. A method to navigate a magnetic capsule, comprises introducing the magnetic capsule into a target area, said magnetic capsule has a longitude direction and the magnetic dipole placed inside the magnetic capsule has a magnetization direction coincide with the longitude direction of the magnetic capsule;placing the magnetic capsule in a first orientation, including its magnetic direction in a z direction;providing an external magnetic control system comprising two magnetic balls, placed on opposing sides of the target area; androtating the magnetic capsule in the target area to adopt a second orientation while maintaining its position in xyz coordinates. 2. The method of claim 1, wherein the two magnetic balls are upper and lower magnetic balls. 3. The method of claim 2, wherein in the first orientation of the magnetic capsule, the two magnetic ball having a magnetic direction mirror to image to each other across a center dividing line between them, the magnetic directions of the two magnetic balls form an angle at 0-90 degrees between the center dividing line. 4. The method of claim 3, further comprising turning the upper magnetic ball counter clockwise while maintaining its position; andkeeping the position and orientation of the bottom magnetic ball unchanged. 5. The method of claim 3, further comprising turning both magnetic balls 0-360 degrees, simultaneously in a clockwise manner along xz plane. 6. The method of claim 2, wherein in the first orientation of the magnetic capsule, the two magnetic ball having a magnetic direction parallel to each other and point to the same direction. 7. The method of claim 6, wherein the second orientation differs in 0-90 degrees from the first orientation in a counter clock wise direction. 8. The method of claim 6, wherein the second orientation differs in 0-360 degrees from the first orientation in a clockwise direction. | Ankon Medical Technologies (Shanghai) Co., LTD. | A61B001/00 A61B034/00 A61B005/00 A61B005/06 A61B001/04
US10709364 | B2 | 2020-07-14 | System and methods for processing analyte sensor data for sensor calibration | 2019-11-21 | Systems and methods for processing sensor analyte data are disclosed, including initiating calibration, updating calibration, evaluating clinical acceptability of reference and sensor analyte data, and evaluating the quality of sensor calibration. The sensor can be calibrated using a calibration set of one or more matched sensor and reference analyte data pairs. Reference data resulting from benchtop testing an analyte sensor prior to its insertion can be used to provide initial calibration of the sensor data. Reference data from a short term continuous analyte sensor implanted in a user can be used to initially calibrate or update sensor data from a long term continuous analyte sensor. | 1. A glucose monitoring system comprising: a transcutaneous electrochemical glucose sensor comprising an in vivo portion and an ex vivo portion, wherein the transcutaneous electrochemical sensor is characterized by an in vivo sensitivity and an in vitro sensitivity, wherein the in vivo sensitivity is different from the in vitro sensitivity, and wherein the in vitro sensitivity is predictive of the in vivo sensitivity due to a relationship therebetween;sensor electronics configured to be operably connected to the transcutaneous electrochemical glucose sensor;a processor configured to be operably connected to the sensor electronics;a memory configured to be operably connected to the processor, wherein the memory has stored therein information associated with the in vivo sensitivity of the transcutaneous electrochemical glucose sensor, wherein the information associated with the in vivo sensitivity is based at least in part on corrections for the difference between the in vivo sensitivity and the in vitro sensitivity of the transcutaneous electrochemical glucose sensor. 2. The glucose monitoring system of claim 1, wherein the relationship between the in vivo sensitivity and the in vitro sensitivity is established empirically by performing a clinical study and comparing in vivo sensitivities with in vitro sensitivities of a plurality of transcutaneous glucose sensors. 3. The glucose monitoring system of claim 1, wherein the in vivo sensitivity is greater than the in vitro sensitivity. 4. The glucose monitoring system of claim 1, wherein the processor is configured to calibrate sensor data from the transcutaneous electrochemical glucose sensor using the information associated with in vivo sensitivity of the transcutaneous electrochemical glucose sensor without reliance on reference glucose concentration values obtained after insertion of at least a portion of the transcutaneous electrochemical glucose sensor in a host. 5. A method for calibrating data from a transcutaneous electrochemical glucose sensor, the method comprising: receiving sensor data from the transcutaneous electrochemical glucose sensor after at least a portion of the transcutaneous electrochemical glucose sensor is inserted in a host, wherein the sensor data is indicative of a glucose concentration in the host; andcalibrating, using a processor, the sensor data received from the transcutaneous electrochemical glucose sensor, wherein the calibrating is based at least in part on information associated with an in vivo sensitivity of the transcutaneous electrochemical glucose sensor, wherein the information associated with the in vivo sensitivity is based at least in part on corrections for a difference between the in vivo sensitivity and an in vitro sensitivity for the transcutaneous electrochemical glucose sensor. 6. The method for calibrating data from a transcutaneous electrochemical glucose sensor of claim 5, wherein the calibrating requires no user calibrations after the at lea... | DexCom, Inc. | A61B005/1473 A61B005/1486 A61B005/1468 A61B005/1495 A61B005/145 A61B005/00
US10722114 | B1 | 2020-07-28 | System and method for vision testing and/or training | 2019-11-03 | A system for vision testing and/or training of a user is disclosed herein. In one embodiment, the system includes a user input device, an eye movement tracking device, a visual display device having an output screen, and a data processing device operatively coupled to the user input device, the eye movement tracking device, and the visual display device. The data processing device is programmed to generate and display one or more visual objects on the output screen of the visual display device; determine, based upon a first signal received from the user input device, whether the user performed a correct action with respect to the one or more visual objects displayed on the output screen; and determine, based upon a second signal received from the eye movement tracking device, whether the user is generally looking in a direction that corresponds to a location of the one or more visual objects. | 1. A system for testing and/or training the vision of a user, said system comprising: a user input device, the user input device configured to output a first signal based upon an input response by a user;an eye movement tracking device, the eye movement tracking device configured to track eye movement and/or eye position of the user, and output a second signal based upon the tracked eye movement and/or eye position of the user;a visual display device having an output screen, the visual display device configured to display one or more visual objects on the output screen so that the one or more visual objects are visible to the user; anda data processing device, the data processing device operatively coupled to the user input device, the eye movement tracking device, and the visual display device, the data processing device being programmed to: generate and display one or more visual objects on the output screen of the visual display device;receive a first signal from the user input device based upon an input response by the user;determine, based upon the first signal received from the user input device, whether the user performed a correct action with respect to the one or more visual objects displayed on the output screen; anddetermine, based upon the second signal received from the eye movement tracking device, whether the user is generally looking in a direction that corresponds to a location of the one or more visual objects on the output screen of the visual display device. 2. The system according to claim 1, wherein the one or more visual objects generated by the data processing device and displayed on the output screen of the visual display device are displaced to different discrete locations on the output screen; and wherein the data processing device is further programmed to determine saccadic eye movements of the user from the second signal outputted by the eye movement tracking device as the user gazes at the one or more visual objects in the different discrete locations on the output screen. 3. The system according to claim 1, wherein the one or more visual objects generated by the data processing device and displayed on the output screen of the visual display device are displaced in a continuous manner on the output screen; and wherein the data processing device is further programmed to determine smooth pursuits eye movements of the user from the second signal outputted by the eye movement tracking device as the user attempts to follow the one or more visual objects being displaced on the output screen. 4. The system according to claim 1, wherein the one or more visual objects generated by the data processing device and displayed on the output screen of the visual display device are displaced to different discrete locations on the output screen; and wherein the data processing device is further programmed to determine anti-saccadic eye movements of the user from the second signal outputted by the eye movement tracking device as the... | Bertec Corporation | A61B003/00 A61B003/14 A61B003/113 A61H005/00 A61B005/16 A61B090/00
US10719941 | B1 | 2020-07-21 | Aerial device including translation or rotation measurement | 2019-12-23 | Disclosed is a system which uses one or more camera units with embedded processors to measure the relative translation and/or rotation between different members on an aerial device. Image data from each unit is processed and transmitted processing to the position control system of the aerial device, and used to determine the position of an aerial element of the device. | 1. A position detection system, the system comprising: a first member configured to be movable relative to a second member;a motion-sensing device on either the first or the second member, the motion-sensing device being oriented to receive image data relevant to motion of the first member relative to the second member; anda processing component configured to determine a position of one of the first and second members using the image data. 2. The system of claim 1 wherein the first member is translatable relative to the second member, and the processing component determines an extent of translation. 3. The system of claim 2 wherein the first member is in a telescopic relationship with the second member. 4. The system of claim 3 wherein the first member is translated in and out of the second member, the second member acting as a sleeve for the first member, the first member being configured to support a working platform. 5. The system of claim 4 wherein the motion-sensing device is located in a coaxial space existing between the first and second members and measures relative movement between the first and second members. 6. The system of claim 4 wherein the motion-sensing device is located at a sleeve end of the second member and aimed to obtain image data relating to the translation of the first member in and out of the second member. 7. The system of claim 2 wherein the first member rotates inside of the second member. 8. The system of claim 7 wherein the motion-sensing device is: (i) in a fixed relationship and rotates with the first member, and (ii) oriented to receive images from internal surfaces of the second member. 9. The system of claim 8 wherein the motion-sensing device includes a camera, the camera being suspended down from the first member and being oriented to receive images from the internal surfaces of the second member, the camera having a lens having a focal length corresponding to a distance from the camera to the internal surfaces of the second member. 10. The system of claim 7 wherein the motion-sensing device is mounted in fixed relation to the second member and aimed in a direction to receive images from a rotating surface of the first member. 11. The system of claim 10 wherein the motion-sensing device includes a camera, the camera having a lens, the lens having a focal length corresponding to a distance from the camera to a rotating surface of the second member. 12. The system of claim 7 wherein the motion-sensing device is: (i) in a fixed relationship and rotates with the first member, and (ii) oriented to receive images from a stationary surface a substantial distance from the first member. 13. The system of claim 7 wherein the motion-sensing device is a camera, and the camera: (i) is in a fixed relationship and rotates with the first member, (ii) is oriented to receive images from a stationary surface a substantial distance from the first member, and (iii) includes a lens having a focal length which is subs... | Altec Industries, Inc. | A61B005/00 G06T007/20 G05D003/20 B66F017/00 H04N005/225 G06T007/70 G06F003/00 B66F011/04
US10726588 | B2 | 2020-07-28 | Magnetic resonance imaging (MRI) image filtration according to different cardiac rhythms | 2018-10-31 | A method includes receiving a plurality of voxel values corresponding to respective locations in a heart, which are acquired using magnetic resonance imaging (MRI). Voxel values that, in spite of (i) corresponding to a same location in the heart and (ii) being gated to a same phase of an electrocardiogram (ECG) cycle of the heart, differ by more than a predefined difference, are identified. An image of at least a portion of the heart is reconstructed from the plurality of voxel values excluding at least the identified voxel values. | 1. A method, comprising: receiving a plurality of voxel values corresponding to respective locations in a heart, which are acquired using magnetic resonance imaging (MRI);identifying voxel values that, in spite of (i) corresponding to a same location in the heart and (ii) being gated to a same phase of an electrocardiogram (ECG) cycle of the heart, differ by more than a predefined difference; andreconstructing an image of at least a portion of the heart from the plurality of voxel values excluding at least the identified voxel values,wherein identifying the voxel values that differ by more than the predefined difference comprises:constructing a scatter plot comprising multiple data points in a plane whose first and second axes denote voxel values acquired in respective first and second MRI acquisitions, wherein each data point represents the voxel values at a given location in the heart in the first and second MRI acquisitions; andidentifying one or more data points that fall outside a predefined region in the plane. 2. The method according to claim 1, wherein reconstructing the image comprises skipping reconstruction of any image that comprises an acquisition taken during an identified arrhythmia. 3. The method according to claim 1, and comprising, in response to excluding the voxel values, adjusting a thickness of the reconstructed image. 4. The method according to claim 1, and comprising, in response to excluding the voxel values, adjusting a reconstruction filter used for reconstructing the image. 5. A system, comprising: a memory, which is configured to store a plurality of voxel values corresponding to respective locations in a heart, which are acquired using magnetic resonance imaging (MRI); anda processor, which is configured to:identify voxel values that, in spite of (i) corresponding to a same location in the heart and (ii) being gated to a same phase of an electrocardiogram (ECG) cycle of the heart, differ by more than a predefined difference; andreconstruct an image of at least a portion of the heart from the plurality of voxel values excluding at least the identified voxel values,wherein the processor is configured to identify the voxel values that differ by more than the predefined difference by:constructing a scatter plot comprising multiple data points in a plane whose first and second axes denote voxel values acquired in respective first and second MRI acquisitions, wherein each data point represents the voxel values at a given location in the heart in the first and second MRI acquisitions; andidentifying one or more data points that fall outside a predefined region in the plane. 6. The system according to claim 5, wherein the processor is configured to reconstruct the image by skipping reconstruction of any image that comprises an acquisition taken during an identified arrhythmia. 7. The system according to claim 5, wherein the processor is further configured to, in response to excluding the voxel values, adjust a thicknes... | Biosense Webster (Israel) Ltd. | A61B005/02 G06K009/00 G06T011/00 A61B005/055 G01R033/567
US10716486 | B1 | 2020-07-21 | System and method to estimate a signal hidden within a composite waveform | 2018-12-20 | A system is provided for isolating the value of a signal hidden within a composite electrical signal. The system comprises an input, a processor, and a memory configured to store instructions executable by the processor. The instructions cause the processor to estimate that portion of a received composite electrical signal that represents a hidden signal by subtracting a known first signal from the composite signal. | 1. A system for estimating a signal hidden within a composite electrical signal, the system comprising: an input for receiving an electrical signal and a first reading, the electrical signal comprising a search section SS comprising the sum of a first signal P1 and a hidden second signal P2 and the first reading describing an initial evaluation of the electrical signal based upon an assumption that SS equals P1 without regard to the hidden signal P2;a processor coupled to the input; anda memory configured to store instructions executable by the processor to cause the processor to: estimate the hidden signal P2; andcalculate a new signal P1 from which a second reading is provided, the second reading describing an updated evaluation of the electrical signal based upon subtracting the estimated hidden signal P2 from the search section SS;whereby, if the second reading is different from the first reading, the effect of the estimated hidden signal P2 is considered before taking action based on the first and second readings. 2. The system of claim 1, wherein: the composite electrical signal is an EKG signal;the first signal P1 comprises a portion of the EKG signal representing a ventricular depolarization wave depV;the hidden signal P2 comprises a portion of the EKG signal representing an atrial repolarization wave repA; andthe search section SS comprises a section of the EKG signal believed to include the ventricular depolarization wave depV and the atrial repolarization wave repA;whereby, the processor estimates depV and repA from the relations SS=depV+repA. 3. The system of claim 1, comprising further instructions causing the processor to determine whether the signals P1 and P2 are plausible. 4. The system of claim 3, comprising further instructions causing the processor to alternately change signal P1 and signal P2 until both are plausible based on SS=P1+P2. 5. The system of claim 3, comprising further instructions causing the processor to determine whether the search section SS is plausible. 6. A method for estimating a signal hidden within a composite electrical signal, the method comprising: receiving a composite electrical signal and a first reading, the composite electrical signal comprising a search section SS comprising the sum of a first signal P1 and a hidden second signal P2 and the first reading describing an initial evaluation of the composite electrical signal based upon an assumption that SS equals P1 without regard to the hidden signal P2; andestimating the hidden signal P2;calculating a new signal P1 from which a second reading is provided, the second reading describing an updated evaluation of the composite electrical signal based upon subtracting the estimated hidden signal P2 from the search section SS; andif the second reading is different from the first reading, considering the effect of the estimated hidden signal P2 before taking action based on the first and second readings. 7. The method of claim 6, wherein: the compo... | - | A61B005/04 A61B005/0472 A61B005/0468
US10716517 | B1 | 2020-07-21 | Biomechanics abnormality identification | 2015-11-25 | A system, method and article of manufacture are presented for assisting the fields of health care, kinesiology, and sports medicine. More specifically the method of the system measures the dynamics of the biomechanics of motion of a human patient or athlete and quantitatively determining the presence or absence of biomechanical abnormalities, classifying abnormalities that are present, developing or critiquing one or more diagnoses in terms of the biomechanics evidence supporting the classification, recommending an appropriate training or treatment regimen based on the diagnoses, and monitoring progress while the individual is under the training or treatment regimen. | 1. Computer-readable storage media having computer-executable instructions embodied thereon that when executed by a computer processor, facilitate a method of movement risk evaluation, the method comprising: receiving a subject array comprising a set of three-dimensional (3-D) subject motion data of motion variables for a test subject performing a first regimen of physical motion, the set of 3-D subject motion data captured using a 3-D motion-capture apparatus, the 3-D subject motion data associated with biomechanical kinetic and kinematic movements of at least one body member, the biomechanical kinetic and kinematic movements including at least one of body member angle, 3-D position, linear velocity, linear acceleration, jerk, impulse, work, force, power, momentum, angular velocity, angular jerk, angular impulse, angular acceleration, or torque;determining one or more biclusters from an expert database clustered using a biclustering technique, the expert database comprising at least a set of reference arrays, each reference array comprising 3-D reference motion data of the motion variables measured from a reference subject performing a second regimen of physical motion, wherein the biclustering technique produces biclusters from a two-dimensional array comprising a set of reference subjects associated with the set of reference arrays arranged in a first direction of the two-dimensional array and the motion variables arranged in a second direction of the two-dimensional array;determining one or more potential motion risks to a bodily area of the test subject by identifying a bicluster from the one or more biclusters based on a similarity between the set of 3-D subject motion data and the 3-D reference motion data of the identified bicluster; andissuing an electronic notification comprising information indicating the one or more potential motion risks relevant to the test subject. 2. The computer-readable media of claim 1, the method further comprising receiving measurements for a set of biological health variables associated with the test subject, wherein the subject array further comprises the measurements for the set of biological health variables. 3. The computer-readable media of claim 2, wherein the measurements for the set of biological health variables comprise biological health data derived from biological health sensors operative at a time of the receiving of the set of 3-D subject motion data of the test subject performing the first regimen of physical motion. 4. The computer-readable media of claim 3, wherein the biological health data comprises one or more of blood pressure, blood oxygen measure, heart rate, respiration rate, temperature, gender, age, weight, body mass index, or body size. 5. The computer-readable media of claim 1, wherein data associated with the identified bicluster comprises measurements for a subset of motion variables of the set of reference arrays, the set of reference arrays determined by eliminating one or... | Cerner Innovation, Inc. | A61B005/00 A61B005/0205
US10712924 | B2 | 2020-07-14 | Non-contacting monitor for bridges and civil structures | 2019-11-11 | A system for monitoring the movement of objects, structures, models of structures, cables and the like provides for the acquisition of images with an optical sensing device such as a video camera fixedly mounted at a selected distance from the item studied, in which the images are arranged into frames divided into pixels which are characterized by an intensity reflected or emitted over a selected time interval, and a data processing system to calculate a physical displacement as function of time of the item being studied or a portion of the item being studied based on an output from the video camera, and in some embodiments the system visually distinguishes one or more locations in the frame to indicate a difference in the phase of motion for multiple objects appearing in the frame. | 1. A system for monitoring the motion of a first object having at least one surface that reflects light from an environment surrounding the first object and having the capacity to store heat from said environment, comprising: a video camera which can be stably positioned at a selected distance from the first object with an unobstructed view of at least a portion of the first object, wherein the video camera acquires images of the first object which are able to be divided into a plurality of pixels, wherein the pixels are characterized by an intensity associated with one or both of visible light energy from light reflected from the at least one surface of the first object or infra-red energy emitted from thermal heating of the at least one surface of the first object over a selected time interval, and wherein the images are arranged in a plurality of frames; anda data processing system configured to obtain a first measurement of an intensity of a first location of at least one frame comprising at least one pixel and a second measurement of an intensity of a second location of the frame comprising at least one pixel apart from the first location, determine a difference in the first measurement and the second measurement, and display at least one image from the frame,wherein the data processing system is further configured to modify the displayed image to indicate a difference in a phase of motion between the first object and at least a second object based on the difference in the first measurement and the second measurement. 2. The system of claim 1, wherein the object in the displayed image is a first object, and a phase indication of the first object is binary such that a second object in the displayed image is represented to be either within a predetermined phase range with the first object or outside of the predetermined phase range. 3. The system of claim 2, wherein the data processing system is further configured to apply a color scheme to represent the second object as being within the predetermined phase range or outside the predetermined phase range. 4. The system of claim 3, wherein the phase indication is shown only for a location of pixels associated with an object within the predetermined phase range. 5. The system of claim 1, wherein the data processing system is further configured to apply a color scheme to represent a phase angle of the object in the displayed image. 6. A system for monitoring the condition of a structure, the structure having at least one surface that reflects light from an environment surrounding the structure and having the capacity to store heat from said environment, comprising: a video camera which can be stably positioned at a selected distance from the structure with an unobstructed view of a selected portion of the structure, wherein the video camera acquires images of the structure which are divisible into a plurality of pixels, wherein the images are arranged in a plurality of frames comprising at le... | RDI TECHNOLOGIES, INC. | G06K009/00 G06F003/0484 G06F016/732 A61B005/00 A61B005/01 A61B005/024 A61B005/08 A61B005/16 G06T007/00 G01N029/44 G06T007/11 G06T007/246 G06T007/262 H04N007/18
US10720038 | B1 | 2020-07-21 | Emergency response systems and methods of using the same | 2019-09-30 | The present disclosure relates to an emergency response system having an emergency response system server, a patient database, a communication system, and many patient response devices (PAD) for many patients. Each patient and each PAD are registered at emergency response system server and patient information is stored in patient database. When emergency occurs to a patient carrying PAD, patient presses an emergency button on PAD and PAD initiates an emergency call to emergency dispatch center to notify patient's GPS location and patient information. live emergency responder connects to one or more patient's helping hands and a medical facility to coordinate immediate medical assistance to the patient based on the patient information received and retrieved from the patient database. Patient specific medical assistance instructions are sent to patient, patient takes patient specific emergency medicines stored in PAD. The patient continues to communicate with the live emergency responder until ambulance arrives. | 1. An emergency response system, comprising: an emergency response system server, wherein each of a plurality of patients and a corresponding patient response device each patient carries are registered at the emergency response system server, and the emergency response system server provides immediate emergency assistance to each of the plurality of patients when emergencies occur;a patient database connected to and accessible by the emergency response system server, wherein the patient database stores patient information of the plurality of patients, wherein the patient information comprises personal information, medical history, patient contact information, and contact information of one or more patient's helping hands and local medical facilities to be notified of each patient;a communication system connected to the emergency response system server, wherein the communication system provides voice, text, and video communication over a communication network among the patient, one or more live emergency responders from a nearby emergency dispatch center, one or more patient's helping hands on record, and one or more nearby medical facilities to provide immediate emergency assistance to the patient when emergency occurs; wherein the patient's helping hands include all family members, relatives, friends, neighbors, guardians, bystanders, local emergency or non-emergency medical facilities, community rescue members and trained CPR volunteers, and anyone who is nearby and is able to offer assistance, anda plurality of patient response devices, wherein each patient carries a corresponding patient response device, and each patient response device comprises an emergency button and a microphone for the patient to activate the patient response device and initiate at least voice communication with the emergency response system server and the communication system directly, and an emergency medicine storage for storing one or more patient specific emergency medicines to be used when emergency occurs,wherein an emergency response system indicator of a patient response device is lit in green light indicating the patient response device is in normal operation state, when an emergency occurs to a patient, the patient activate the patient response device, the emergency response system indicator turns red indicating the patient is in an emergency, the patient response device initiates an emergency call to the nearby emergency dispatch center to notify a live emergency responder with the patient's GPS location information and patient information including contact information of one or more patient's helping hands on record, the live emergency responder connects to one or more patient's helping hands on record and a nearby medical facility to coordinate immediate medical assistance to the patient based on the patient information received and retrieved from the patient database, and the live emergency responder and the emergency response system server provide patient... | - | G08B023/00 G08B021/04 G08B025/01 G10L015/26 A61B005/00
US10709367 | B1 | 2020-07-14 | Multidepth tissue oximeter | 2015-01-20 | An oximeter measures oxygen saturation for two or more different tissue depths and shows these results on a screen. A probe of the oximeter has multiple different distances between source and detector sensors. One probe implementation has fixed sensor positions. Other implementations include sensors on a moveable platform or openings to accept sensors, which allow a user to vary a distance between sensors. | 1. A device comprising: a tissue oximetry probe including a probe surface, which will be placed against a tissue to be measured;a first structure of a first type formed on the probe surface;a second structure of the first type formed on the probe surface;a third structure of a second type formed on the probe surface;a fourth structure of the second type formed on the probe surface;a fifth structure of the second type formed on the probe surface; anda sixth structure of the second type formed on the probe surface, wherein a first distance extends between the first structure and the third structure without touching the second structure or the fifth structure, a second distance extends between the second structure and the fourth structure without touching the first structure or sixth structure, the first distance is different from the second distance,a third distance extends between the first structure and the fifth structure, a fourth distance extends between the second structure and the sixth structure, and the third distance is different from the fourth distance,a first line extends through centers of the third and fourth structures,a second line extends through centers of the fifth and sixth structures,the first line is not collinear with the second line,the third structure comprises an optical fiber and the fourth structure comprises a photodetector, andthe probe surface includes the first structure without an intervening structure of the second type between an edge of the probe surface and the first structure. 2. The device of claim 1 wherein the second distance is greater than the first distance, and the fourth distance is greater than the third distance. 3. The device of claim 1 wherein the first line passing through the third and fourth structures is parallel to the second line passing through the fifth and sixth structures. 4. The device of claim 3 wherein a third line passing through the first and second structures is not parallel to the first line. 5. The device of claim 1 wherein a third line passing through the first, third, and fifth structures is parallel with a fourth line passing through the second, fourth, and sixth structures. 6. The device of claim 1 wherein the first type of structure is a source and the second type of structure is a detector. 7. The device of claim 1 comprising: a first filter associated with the third and fourth structures allows a first spectrum of radiation to pass through to the third and fourth structures; anda second filter associated with the fifth and sixth structures allows a second spectrum of radiation to pass through to the fifth and sixth structures, wherein the first spectrum is different from the second spectrum. 8. The device of claim 7 wherein the first filter comprises a greater attenuation to input radiation than the second filter. 9. The device of claim 7 wherein the first filter comprises greater attenuation at a wavelength range to input radiation than the second filter. 10. The ... | ViOptix, Inc. | A61B005/1455 A61B005/00
US10720141 | B1 | 2020-07-21 | Tympanic membrane measurement | 2018-12-28 | An in-ear device includes a housing shaped to hold the in-ear device in an ear, and an audio package, disposed in the housing, to emit sound. A tympanic membrane measurement unit (TMMU) is structured to measure a movement of a tympanic membrane in the ear caused by external sound received by the tympanic membrane, and a controller is coupled to the audio package and the TMMU. The controller includes logic that when executed by the controller causes the in-ear device to perform operations. The operations include measuring a movement of the tympanic membrane, and in response to measuring the movement of the tympanic membrane, outputting sound from the audio package to destructively interfere with the external sound received by the tympanic membrane. | 1. A system, comprising: a magnetically active material deposited on a tympanic membrane of an ear, wherein the magnetically active material generates electromagnetic waves; andan in-ear device, comprising: a housing shaped to hold the in-ear device in the ear;an audio package, disposed in the housing, to emit sound;a tympanic membrane measurement unit (TMMU) structured to measure a movement of the tympanic membrane caused by external sound received by the tympanic membrane, wherein the TMMU includes: an electromagnetic detection system positioned in the housing to measure changes in the electromagnetic waves from the magnetically active material when the in-ear device is disposed in the ear, wherein the movement of the tympanic membrane is measured using the changes in the electromagnetic waves; anda controller coupled to the audio package and the TMMU, wherein the controller includes logic that when executed by the controller causes the in-ear device to perform operations, including: receiving the movement of the tympanic membrane measured by the TMMU;determining a waveform of cancellation sound based at least on the movement of the tympanic membrane; andoutputting the cancellation sound from the audio package to destructively interfere with the external sound received by the tympanic membrane. 2. The system of claim 1, wherein the TMMU and the controller are disposed in the housing. 3. The system of claim 1, further comprising a microphone coupled to the controller and mounted in or on the housing, wherein the controller further includes logic that when executed by the controller causes the in-ear device to perform operations, including: collecting external sound with the microphone;wherein determining the waveform of the cancellation sound based at least on the movement of the tympanic membrane includes determining the waveform of the cancellation sound based at least on the movement of the tympanic membrane and the external sound collected with the microphone. 4. The system of claim 3, wherein determining the waveform of the cancellation sound based at least on the movement of the tympanic membrane and the external sound collected with the microphone includes: analyzing the external sound collected with the microphone using predictive models. 5. The system of claim 1, wherein the TMMU includes: one or more optical emitters disposed in or on the housing and oriented to emit light towards the tympanic membrane when the in-ear device is disposed in the ear; andone or more optical receivers disposed in or on the housing and oriented to receive the light reflected from the tympanic membrane, wherein the movement of the tympanic membrane is measured using the light reflected from the tympanic membrane. 6. The system of claim 5, wherein the one or more optical receivers include an image sensor to output an image of the tympanic membrane, and wherein the movement of the tympanic membrane is measured using the image. 7. The system of claim 1, w... | X Development LLC | G10K011/16 G10K011/178 H04R001/10 A61F011/08 A61B001/227 A61B005/05 A61B005/00 A61B005/12 A61F011/14
US10702131 | B2 | 2020-07-07 | Imaging system and method | 2019-11-07 | An imaging connector includes a proximal side and a distal side. The proximal side includes a light input opening and an image output opening. The distal side includes a light output opening and an image input opening. The imaging connector is operable to (i) transmit light from the light input opening to the light output opening, and (ii) transmit an image from the image input opening to the image output opening. | 1. A system comprising: a processor; andmemory in communication with the processor, the memory storing instructions that when executed on the processor cause the processor to perform operations comprising: receiving an image from an image capturing device, the image corresponding to a target area of a patient;encrypting the received image of the target area of the patient;generating a modified image by editing the encrypted image; anddetermining a difference between the modified image of the target area of the patient and an ideal image, the ideal image corresponding to normality for the target area. 2. The system of claim 1, wherein the operations further comprise diagnosing the target area of the patient based on the difference between the modified image of the target area of the patient and the ideal image corresponding to normality for the target area. 3. The system of claim 1, wherein the operations further comprise: receiving the ideal image corresponding to the normality for the target area; andoverlaying the ideal image onto the modified image of the target area. 4. The system of claim 1, wherein editing the encrypted image comprises enhancing a portion of the encrypted image. 5. The system of claim 1, wherein editing the encrypted image comprises marking up a portion of the encrypted image. 6. The system of claim 1, further comprising a personal communication device comprising the processor and the memory. 7. The system of claim 1, wherein the operations further comprise storing the encrypted image in the memory. 8. The system of claim 7, wherein the operations further comprise retrieving the encrypted image from being stored in the memory based on a fingerprint. 9. The system of claim 1, wherein the operations further comprise communicating the modified image or the difference between the modified image and the ideal image to a graphical user interface. 10. The system of claim 9, wherein an external display comprises the graphical user interface, the external display in communication with a personal communication device comprising the processor and the memory. 11. A method comprising: receiving, at a processor, an image from an image capturing device, the image corresponding to a target area of a patient;encrypting, by the processor, the received image of the target area of the patient;generating, by the processor, a modified image by editing the encrypted image; anddetermining, by the processor, a difference between the modified image of the target area of the patient and an ideal image, the ideal image corresponding to normality for the target area. 12. The method of claim 11, further comprising generating, by the processor, a diagnosis for the target area of the patient based on the difference between the modified image of the target area of the patient and the ideal image corresponding to normality for the target area. 13. The method of claim 11, further comprising: receiving, at the processor, the ideal image correspond... | - | A61B005/00 A61B001/00 A61B001/227 A61B001/04 A61B001/233
US10719934 | B2 | 2020-07-21 | Devices, systems, and methods for medical imaging | 2018-10-17 | Devices, systems, and methods for generating a medical image receive scan data, wherein the scan data is defined in an acquisition space; perform a first reconstruction process to generate a lower-resolution image based on the scan data, wherein the first reconstruction process uses a subset of the scan data; and perform a second reconstruction process to generate a higher-resolution image based on the scan data, wherein the second reconstruction process uses more of the scan data than the subset of the scan data. | 1. A method for generating medical images, the method comprising: receiving scan data, wherein the scan data is defined in an acquisition space;performing a first reconstruction process to generate a lower-resolution image based on the scan data, wherein the first reconstruction process uses a subset of the scan data;performing a second reconstruction process to generate a higher-resolution image based on the scan data, wherein the second reconstruction process uses more of the scan data than the subset of the scan data, anddisplaying the lower-resolution image before the higher-resolution image is complete. 2. The method of claim 1, wherein the scan data is magnetic-resonance-imaging data. 3. The method of claim 1, further comprising: displaying the lower-resolution image as a thumbnail image. 4. The method of claim 3, further comprising: setting a scan range of a medical-imaging device based on the thumbnail image. 5. The method of claim 3, further comprising: sending the thumbnail image to a display device of a console of an operator of a medical-imaging device; andsending the higher-resolution image to a medical-image server. 6. The method of claim 1, wherein the second reconstruction process uses all of the scan data. 7. The method of claim 1, wherein the acquisition space is k-space. 8. The method of claim 7, wherein the subset of the scan data is a central part of the k-space. 9. The method of claim 1, wherein the lower-resolution image is defined in a two-dimensional image space. 10. The method of claim 1, wherein the lower-resolution image is defined in a three-dimensional image space. 11. The method of claim 1, wherein the first reconstruction process is different from the second reconstruction process. 12. The method of claim 11, wherein the first reconstruction process is faster than the second reconstruction process. 13. The method of claim 1, wherein the scan data includes data used by a compressed-sensing process. 14. A medical device comprising: one or more processors; andone or more computer-readable media that are in communication with the one or more processors and that include instructions for: obtaining scan data, wherein the scan data is defined in an acquisition space;generating a lower-resolution image based on a subset of the scan data;generating a higher-resolution image based on more of the scan data than the subset of the scan data; anddisplaying the lower-resolution image before the higher-resolution image is complete. 15. The device of claim 14, wherein the scan data was produced by magnetic resonance imaging, and wherein the acquisition space is k-space. 16. The device of claim 15, wherein the subset of the scan data is generated by undersampling the scan data. 17. The device of claim 14, wherein the higher-resolution image is generated based on all of the scan data. 18. The device of claim 14, wherein the generating of the lower-resolution image is started after the subset of the scan data has been... | CANON MEDICAL SYSTEMS CORPORATION | G06T007/00 A61B005/055 A61B006/03 A61B008/08 G06T011/00 G06T015/08 A61B005/00
US10716484 | B2 | 2020-07-21 | First-aid device for detecting myocardial infarction | 2019-07-12 | A first-aid device for detecting myocardial infarction, including a wearable textile, a first alarm detection unit, a second detection alarm unit, and a charging first-aid unit for actively detecting and alerting to myocardial infarction, instantly supplying crucial medicine, and providing a charging function. In particular, the wearable textile put on a patient carries the first alarm detection unit actively detecting myocardial infarction, the charging first-aid unit instantly provides the user medicine to alleviate painful symptom, and at the same time, the second detection alarm unit serves as a spare part while the first alarm detection unit is charged by the charging first-aid unit. As a result, the first alarm detection unit and the second detection alarm unit are alternatively used for actively and ceaselessly detecting myocardial infarction assuring first-aid for the patient. | 1. A first-aid device for detecting myocardial infarction, comprising: a wearable textile put on a user and provided with a plurality of sensing electrodes, a plurality of conductive connections, and a connection slot, the sensing electrodes attached onto a chest of the user for sensing and transmitting a physiological signal from the chest, the sensing electrodes connected to the connection slot through the conductive connections;a first detection alarm unit serving as an operating detection alarm unit and comprising a connection port for connecting the connection slot in a detachable manner, the first detection alarm unit receiving the physiological signal, generating and transferring an emergency alarm signal based on the physiological signal, and processing an alarm operation;a second detection alarm unit serving as a spare detection alarm unit, comprising a connection port, and remaining in a standby status for replacing the first detection alarm unit; anda charging first-aid unit comprising a charging space for accommodating and charging the second detection alarm unit, and a medicine box for accommodating at least one sublingual tablet, the charging space provided with a charging slot,wherein the alarm operation comprises generating a sound, vibration, or a light signal, the emergency alarm signal is received by a remote processing unit for an operator of the remote processing unit to perform an emergency treatment, each of the first detection alarm unit and the second detection alarm unit comprises a charging port, a controller, an alarm, and a rechargeable battery, the rechargeable battery supplies power to the controller and the alarm for operation, the charging port of the second detection alarm unit is electrically connected to the charging slot of the charging space of the charging first-aid unit for charging in an another detachable manner, the controller is electrically connected to the charging port and the alarm, receives the physiological signal, and determines if the user suffers from myocardial infarction based on the physiological signal, the controller generates and transfers the emergency alarm signal when the user suffers from myocardial infarction, the alarm receives the emergency alarm signal and performs an alarm operation based on the emergency alarm signal for informing and reminding the user of taking the sublingual tablet from the medicine box, the charging first-aid unit further comprises a charge controller and a rechargeable battery module, the charge controller is provided with a universal serial bus (USB) input port for connecting an external power source, and a power output port for connecting the charging slot, the charge controller is connected to the rechargeable battery module for charging the rechargeable battery module by an external power from the external power source through the USB input port, and the charge controller outputs an output power from the rechargeable battery module to the charging slot ... | Asiatic Fiber Corporation | A61B005/0408 G16H050/30 A61N001/04 A61B005/00 A61B005/0452 A61B005/024 A61B005/04 A61B005/0205
US10722218 | B2 | 2020-07-28 | Method and apparatus for three dimensional reconstruction of a joint using ultrasound | 2019-09-11 | A method of generating a 3-D patient-specific bone model, the method comprising: (a) acquiring a plurality of raw radiofrequency (“RF”) signals from an A-mode ultrasound scan of a patient's bone at a plurality of locations using an ultrasound probe that comprises a transducer array; (b) tracking the acquiring of the plurality of raw RF signals in 3-D space and generating corresponding tracking data; (c) transforming each of the plurality of raw RF signals into an envelope comprising a plurality of peaks by applying an envelope detection algorithm to each of the plurality of raw RF signals, each peak corresponding with a tissue interface echo; (d) identifying a bone echo from the tissue interface echoes of each of the plurality of raw RF signals to comprise a plurality of bone echoes by selecting the last peak having a normalized envelope amplitude above a preset threshold, wherein the envelope amplitude is normalized with respect to a maximum peak existing in the envelope; (e) determining a 2-D bone contour from the plurality of bone echoes corresponding to each location of the ultrasound probe to comprise 2-D bone contours; (f) transforming the 2-D bone contours into an integrated 3-D point cloud using the tracking data; and, (g) deforming a non-patient specific 3-D bone model corresponding to the patient's bone in correspondence with the integrated 3-D point cloud to generate a 3-D patient-specific bone model. | 1. A method of generating a virtual three dimensional subject-specific musculoskeletal model, the method comprising: generating an envelope for each of a plurality of raw signals obtained from an ultrasound scan of at least a portion of a subject, each envelope including one or more peaks;generating contour lines based on the peaks of the envelopes;transforming the contour lines into a three dimensional (3-D) point cloud using tracking information obtained by tracking a position of an ultrasound transducer during the ultrasound scan of at least the portion of the subject; andoptimizing a template tissue model with respect to the point cloud to generate a virtual three dimensional subject-specific musculoskeletal model. 2. The method according to claim 1, wherein the plurality of raw signals includes a first raw signal at a first frequency and a second raw signal at a second frequency, where the first frequency is different than the second frequency. 3. The method according to claim 1, further comprising: transmitting ultrasound signals during the ultrasound scan; andobtaining the plurality of raw signals responsive to transmitting the ultrasound signals. 4. The method according to claim 3, wherein obtaining the plurality of raw signals comprises: receiving reflected ultrasound energy from each of the transmitted ultrasound signals; and,converting the reflected ultrasound energy from each of the transmitted ultrasound signals into the plurality of raw signals. 5. The method according to claim 3, wherein transmitting the ultrasound signals comprises: transmitting a plurality of ultrasound pulses each having a frequency; and,changing the frequency of the ultrasound pulses as the raw signals are acquired. 6. The method according to claim 5, wherein the frequency of the ultrasound pulses is changed by sweeping the frequency. 7. The method of claim 1, wherein generating the contour lines based on the peaks of the envelope signals comprises: identifying a first peak of the one or more peaks in each envelope; andapplying a first filter to the first peaks in temporally adjacent envelopes. 8. The method according to claim 7, wherein the first peak is identified in each of the envelopes by using a second filter to estimate an optimal time delay. 9. The method according to claim 8, wherein the second filter is selected from a group consisting of a Kalman filter, a recursive Bayesian filter, and a particles filter. 10. The method according to claim 7, wherein the first filter is based on Bayesian model smoothing. 11. The method according to claim 1, wherein generating the contour lines based on the peaks of the envelopes comprises: calculating a first peak data matrix for the envelope of a first raw signal of the plurality of raw signals;estimating a second peak data matrix for the envelope of a second raw signal of the plurality of raw signals based on the first peak data matrix;calculating the second peak data matrix from the envelope of the secon... | JointVue, LLC | A61B008/08 A61B008/00 G06T017/00 A61B005/00 G06T019/20 A61B034/10 A61B008/14 G06T007/564 A61B090/00 A61B034/20 A61B005/055
US10716498 | B2 | 2020-07-21 | System and methods for processing analyte sensor data for sensor calibration | 2019-11-21 | Systems and methods for processing sensor analyte data are disclosed, including initiating calibration, updating calibration, evaluating clinical acceptability of reference and sensor analyte data, and evaluating the quality of sensor calibration. The sensor can be calibrated using a calibration set of one or more matched sensor and reference analyte data pairs. Reference data resulting from benchtop testing an analyte sensor prior to its insertion can be used to provide initial calibration of the sensor data. Reference data from a short term continuous analyte sensor implanted in a user can be used to initially calibrate or update sensor data from a long term continuous analyte sensor. | 1. A method of making one or more glucose monitoring systems, the method comprising: manufacturing one or more transcutaneous electrochemical glucose sensors, each transcutaneous electrochemical glucose sensor having an in vivo portion and an ex vivo portion;examining an in vitro signal response of at least one of the one or more transcutaneous electrochemical glucose sensors for each of a plurality of known glucose concentrations;calculating an in vitro sensitivity for at least one of the one or more transcutaneous electrochemical glucose sensors, based on the in vitro signal responses;predicting an in vivo sensitivity for at least one of the one or more transcutaneous electrochemical glucose sensors based at least in part on (1) the calculated in vitro sensitivity and (2) a relationship between in vivo responses and in vitro data empirically established in at least one previously performed clinical study;storing the predicted in vivo sensitivity in a memory; andconfiguring one or more processors to calibrate data generated by at least one of the one or more transcutaneous electrochemical glucose sensors using the predicted in vivo sensitivity without reliance on any manual calibrations by a user of at least one of the one or more glucose monitoring systems following insertion of at least a portion of the at least one of the one or more transcutaneous electrochemical glucose sensors. 2. One or more glucose monitoring systems made by the method of claim 1. 3. The method of making one or more glucose monitoring systems of claim 1, further comprising manufacturing one or more sensor electronics, wherein each of the sensor electronics is configured to be placed in electrical communication with one of the one or more transcutaneous electrochemical glucose sensors. 4. The method of making one or more glucose monitoring systems of claim 3, wherein each of the sensor electronics is configured to receive the data generated by the at least one of the one or more transcutaneous electrochemical glucose sensors. 5. A glucose monitoring system comprising: a transcutaneous electrochemical glucose sensor having an in vivo portion and an ex vivo portion, wherein the in vivo portion is configured for insertion into a host; andsensor electronics that are configured to provide calibrated glucose data based on sensor data provided by the transcutaneous electrochemical glucose sensor and a predicted in vivo sensitivity based at least in part on (1) a calculated in vitro sensitivity and (2) a relationship between in vivo responses and in vitro data empirically established in at least one previously performed clinical study associated with a plurality of transcutaneous electrochemical glucose sensors,wherein the sensor electronics are in electrical communication with the transcutaneous electrochemical glucose sensor. 6. The glucose monitoring system of claim 5, wherein the calibrated glucose data provided by the sensor electronics are not based on manual calibrati... | DexCom, Inc. | A61B005/1495 A61B005/1468 A61B005/1473 A61B005/1486 A61B005/145 A61B005/00
US10722172 | B1 | 2020-07-28 | Fitness systems and methods | 2019-08-29 | The invention is directed to systems and methods that can determine a user's fitness based on a single factor such as oxygen lung volume (pmVO2). A heart rate measurement is conducted on the user after the user performs a physical test specified by a fitness application of the system. The heart rate measurement is converted into a pmVO2 value and compared to pmVO2 values of other individuals. The system assigns the user a fitness grade based on the comparison result. The system further includes an artificial intelligent system that helps the user to move from a grade to a higher grade. | 1. A system for determining physical fitness comprising: a management system including a database and a central subsystem, wherein: the database is configured to store health information of a plurality of individuals and a fitness scale, wherein the health information includes predicted lung oxygen volume maxes (pmVO2s) of the individuals and health conditions associated with their pmVO2s and the fitness scale includes multiple grades for classifying pmVO2 or information representative of pmVO2 received from a fitness application; andthe central subsystem is configured to assign a threshold pmVO2 to each grade based on the health information, analyze pmVO2 or information representative of pmVO2 received from a fitness application in view of the pmVO2s and health conditions stored in the database, and determine a grade from the fitness scale to classify the received pmVO2 or information representative of pmVO2;wherein the management system is configured to: operate with a mobile device configured to receive a physical measurement from a wearable device, a geographical position calculating device configured to measure distance traveled, and a timer configured to set a period of time, the wearable device king configured to conduct a physical measurement and transmit the physical measurement to the mobile device;operate with a fitness application that provides options via a user interface allowing selection of one or more physical tests;verify whether persons attempting to use the fitness application are authorized users;receive a roster of individuals from an enterprise computer and the individuals on the roster are authorized users of the fitness application;wherein, in response to a selection of the one or more physical tests on the fitness application, the fitness application instructs the mobile device to: measure, via the geographical position calculating device, the distance traveled by an authorized user during performance of the selected test; or measure, via the timer, a time spent by an authorized user during performance of the selected test;instruct the wearable device to conduct a physical measurement; andtransmit a predicted lung oxygen volume max (pmVO2) or information representative of pmVO2 of the authorized user determined from the performance of the test to the management system;wherein the management system is further configured to: collect and aggregate pmVO2s or information representative of pmVO2s of one or more authorized users in the database and remove information identifying the one or more authorized users so that the pmVO2s or information representative of pmVO2s of the one or more authorized users are saved in the database anonymously;determine the grade based on the received pmVO2 or information representative of pmVO2 and the fitness scale, and classify each of the one or more authorized users to one of the multiple grades in the fitness scale based on the determined grade; andmonitor activities of the one or more auth... | COOPER HEALTH AND FITNESS APPLICATIONS, LLC | A61B005/00 A61B005/0205 A61B005/11 A61B005/091 A61B005/024 A61B005/083 A63B024/00
US10716634 | B2 | 2020-07-21 | 3D system and method for guiding objects | 2019-12-30 | An intraoral 3D scanner includes a probe light source configured to generate a probe light such that the probe light is transmitted towards the dental situation; a camera including an array of sensor elements, the camera being arranged such that the probe light from the dental situation is transmitted to the array of sensor elements, wherein the camera is configured to create images of the dental situation from which a point cloud is generated, and a guiding system configured to guide relative movement of the intraoral 3D scanner towards the dental situation, wherein the camera is part of the guiding system, the camera is configured to record images from which a relative position of the intraoral 3D scanner and the dental situation is determined, such that based on the relative position, the guiding system is configured to provide a positioning signal in the form of a positioning color code. | 1. A 3D guiding system for guiding a relative movement of a dental tool and a dental situation, the 3D guiding system comprises: a non-transitory computer-readable medium configured for at least temporary storing:a 3D-model of the dental situation, wherein the 3D-model is based on sub-surface scanning of the dental situation such that the 3D-model comprises data relating to an interior structure of the dental object;program code for determining a present relative arrangement of the dental tool and the dental situation based on the 3D-model of the dental situation; andprogram code for calculating information for guiding the relative movement of the dental tool and the dental situation towards a first preferred relative arrangement from said present relative arrangement; andan information displaying device configured for displaying the calculated information to an operator. 2. The 3D guiding system according to claim 1, wherein the sub-surface scanning is performed by Magnetic Resonance Imaging. 3. The 3D guiding system according to claim 1, wherein the sub-surface scanning is performed by X-ray scanning. 4. The 3D guiding system according to claim 1, wherein the sub-surface scanning is performed by CT scanning. 5. The 3D guiding system according to claim 1, wherein the non-transitory computer-readable medium further stores program code for deriving from said 3D-model the first preferred relative arrangement between the dental tool and the dental situation. 6. The 3D guiding system according claim 1, wherein the 3D guiding system comprises program code for deriving further preferred relative arrangements between the dental tool and the dental situation, and wherein the first preferred relative arrangement and the further preferred relative arrangements together form a preferred path of the relative movement of the dental tool and the dental situation, and said 3D guiding system is configured for guiding said relative movement according to said preferred path. 7. The 3D guiding system according to claim 6, wherein one or more of said further preferred relative arrangements corresponds to an arrangement where at least a portion of a distal end of the dental tool is located below the surface of the dental situation, wherein said 3D model comprises data relating to the interior structure of the dental situation. 8. The 3D guiding system according to claim 1, wherein the 3D-model of the dental situation comprises data relating to the dental situation and data relating to a planned modification of the dental situation. 9. The 3D guiding system according to claim 8, wherein the dental situation relates to the mandibular or maxillary of a patient and the planned modification corresponds to a hole which is planned to be defined in the mandibular or maxillary for accepting a dental implant. 10. The 3D guiding system according to claim 9, wherein said preferred path is configured to provide that the dental tool in the form of a drilling tool when gui... | 3SHAPE A/S | H04N007/18 A61B034/20 A61B005/00 A61C008/00 A61C003/02 A61C001/08 A61B017/17
US10711852 | B1 | 2020-07-14 | Locking clutch systems | 2019-01-18 | A clutch assembly includes an inner hub, an outer hub, a roller bearing, and a control collar. The inner hub includes a cam surface having a peak and a bottom. The roller bearing is interposed between the inner and outer hub and coupled to the inner hub at a predetermined position relative to the peak. In a locked position, the roller bearing is proximate the peak, wherein rotation of the inner hub relative to the outer hub is inhibited in a locked direction. In the unlocked position, the roller bearing is disposed a greater distance from the peak relative to the locked position, wherein rotation of the inner hub relative to the outer hub is not inhibited in the locked direction. | 1. A clutch assembly comprising: an inner hub rotatable about a hub axis, the inner hub comprising a cam surface comprising a peak and a bottom;an outer hub concentric with and disposed radially outward of the inner hub, the inner hub rotatable about the hub axis with respect to the outer hub;a roller bearing interposed between the inner and outer hub, the roller bearing coupled to the inner hub at a predetermined position relative to the peak; anda control collar configured to move the roller bearing between a locked position and an unlocked position, wherein: in the locked position, the roller bearing is proximate the peak, wherein rotation of the inner hub relative to the outer hub is inhibited in a locked direction but uninhibited in an unlocked direction,in the unlocked position, the roller bearing is disposed a greater distance from the peak relative to the locked position, wherein rotation of the inner hub relative to the outer hub is not inhibited in the locked direction or the unlocked direction. 2. The clutch assembly of claim 1, further comprising a control collar positioning member coupling the control collar to the inner hub and urging the control collar into the locked position, the control collar positioning member comprising a torsion spring disposed about the hub axis. 3. The clutch assembly of claim 1, comprising two roller bearings disposed 180 degrees apart from each other. 4. The clutch assembly of claim 1, further comprising a recess extending into an inner surface of the outer hub, the recess defining a free range in which the inner hub is free to rotate relative to the outer hub. 5. The clutch assembly of claim 1, further comprising a detent assembly coupling the inner hub and outer hub. 6. The clutch assembly of claim 1, further comprising a control collar actuation assembly coupled to the control collar and configured to urge the control collar to the unlocked position. 7. The clutch assembly of claim 1, further comprising a roller cage interposed between the inner hub and the outer hub, the roller bearing disposed within the roller cage, the control collar coupled to the roller cage and configured to rotate the roller cage relative to the hub axis. 8. A clutch assembly comprising: an inner hub rotatable about a hub axis, the inner hub comprising a cam surface comprising a peak and a bottom;an outer hub concentric with and disposed radially outward of the inner hub, the inner hub rotatable about the hub axis with respect to the outer hub, the outer hub having a recess extending into an inner surface of the outer hub, the recess defining a free range in which the inner hub is free to rotate relative to the outer hub;a roller bearing interposed between the inner and outer hub, the roller bearing coupled to the inner hub at a predetermined position relative to the peak; anda control collar configured to move the roller bearing between a locked position and an unlocked position, wherein: in the locked position, the ro... | General Electric Company | F16D041/08 A61B005/055 F16D041/067 F16D015/00
US10720239 | B2 | 2020-07-21 | Computer-assisted patient navigation and information systems and methods | 2019-10-24 | A computer-assisted patient navigational communication system for receiving electronic and oral communications from a patient, scanning data to determine the medical needs of the patient, and displaying relevant information to appropriate medical personnel who can immediately advise the patient of the most appropriate source of medical assistance relating to the patient's identified symptoms. | 1. A computer-assisted patient navigation and communication system that receives communications from a patient regarding the patient's symptoms, scans patient data to determine the medical needs of the patient, and displays information to navigation personnel who can immediately navigate the patient to the most appropriate medical provider and/or service responding to the patient's identified symptoms, said system comprising: a) a call center having said navigation personnel who navigate the patient to the most appropriate response by either (i) providing assessment, diagnosis, treatment, and/or education to said patient, or (ii) further navigating the patient to the most appropriate medical provider inside or outside the call center for assessing, diagnosing and/or treating the patient's symptoms; andb) said call center having facilities for communications between said patient and said navigation personnel, said facilities being arranged and configured for receiving, recording and displaying patient physiological data from one or more of the following sources: patient's oral or electronic communications, patient's video communications, physiological monitors possessed by the patient, patient physiological monitors located at remote facilities, patient physiological data in the patient's electronic health records, and/or data obtained from health care personnel associated with the patient;c) wherein said communications and patient physiological data enable said navigation personnel to assess, diagnose, treat, educate and/or refer said patient to one or more of the following medical providers: a laboratory, imaging facility, pharmacy, home health personnel, physical therapist, ambulance facility, hospital, emergency room, physician and/or medical specialist responding to the patient's identified symptoms. 2. The system according to claim 1, wherein personnel qualifications of said navigation personnel and said medical providers are input in the system and associated with patient symptoms, patient chronic diseases and/or patient communication barriers. 3. The system according to claim 2, wherein personnel qualifications of each navigation personnel and medical provider include one or more of the following: credentials, areas of special knowledge, and languages spoken. 4. The system according to claim 3, further comprising algorithms for first verifying the correctness of the incoming patient physiological data, and then weighting the patient communications and patient physiological data, and associating certain navigation personnel and medical providers with specific conditions and symptoms. 5. The system according to claim 1, wherein said navigation personnel consist of health care providers defined as physicians, physician assistants or nurses. 6. The system according to claim 1, wherein said call center is a virtual call center. 7. A virtual medical system for minimizing anxiety of a patient with health symptoms, said system designed to sp... | - | H04M003/51 G16H010/60 A61B005/00 G16H050/20 G16H040/63 G16H040/20 G06Q050/22 G06Q010/10 A61B005/021 H04M003/523 A61B005/145 A61B005/024
US10709391 | B1 | 2020-07-14 | Method of cell tagging for cancer detection | 2015-12-02 | Body-mountable devices are provided to detect the presence or status of a tumor in a body by detecting probes associated with circulating cells of the tumor that travel to subsurface vasculature of the body. The probe enters a tumor and associates with cells of the tumor before the cells metastasize. A wearable body-mountable device can be worn for a protracted period of time to detect the probe associated with circulating tumor cells in the vasculature at low concentrations and/or at low rates. A body-mounted device could detect the presence of such released, tumor-cell-associated probes to determine a presence or status of a tumor in the body. | 1. A system comprising: a probe, wherein the probe comprises at least one of (i) a tumor-binding agent that experiences a conformation change in response to exposure to a tumor environment such that the tumor-binding agent can bind to tumor cells when exposed to the tumor environment or (ii) an interfering element or protective layer that is capable of preventing the probe from binding to tumor cells and that is disabled in response to exposure to a tumor environment such that the probe can bind to tumor cells when exposed to the tumor environment, whereby the probe selectively associates with tumor cells before the tumor cells metastasize, and wherein the probe is configured to be able to remain associated with the tumor cells after the tumor cells have metastasized; anda body-mountable device, wherein the body-mountable device comprises: a sensor configured to be mounted to an external body surface of a body proximate a portion of subsurface vasculature, wherein the sensor is configured to detect the probe in the portion of subsurface vasculature; anda controller operably coupled to the sensor, wherein the controller comprises a computing device programmed to perform controller operations comprising: operating the sensor to detect the probe, wherein detecting the probe comprises measuring an amount of the probe in the portion of subsurface vasculature; anddetermining a tumor presence or status based on the measured amount of the probe in the portion of subsurface vasculature. 2. The system of claim 1, wherein determining a tumor presence or status comprises determining that a tumor is present in the body. 3. The system of claim 1, wherein determining a tumor presence or status comprises determining a degree of metastasis of a tumor in the body. 4. The system of claim 1, wherein the body-mountable device comprises a wearable device. 5. The system of claim 1, wherein the probe comprises a fluorophore, wherein measuring an amount of the probe in the portion of subsurface vasculature comprises detecting an amount of fluorescence light emitted from the portion of subsurface vasculature by the fluorophore, wherein the probe comprises a nanoparticle comprised of a protective material, and wherein the fluorophore is disposed within the protective material. 6. The system of claim 1, wherein the probe comprises a magnetic nanoparticle, wherein measuring an amount of the probe in the portion of subsurface vasculature comprises detecting a magnetic field produced by the magnetic nanoparticle. 7. The system of claim 1, wherein the probe comprises an interfering element that is capable of preventing the probe from binding to tumor cells and that is disabled in response to exposure to a tumor environment such that the probe can bind to tumor cells when exposed to the tumor environment, and wherein the interfering element is disabled by a protease that is characteristic of a tumor environment. 8. The system of claim 1, wherein the probe comprises a tumo... | Verily Life Sciences LLC | A61B005/00 A61B005/1455 A61B006/00 A61B005/055
US10722152 | B2 | 2020-07-28 | Analyte sensor | 2019-11-05 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A system for measuring a glucose concentration in a host, the system comprising: a mounting unit; an electronics unit; andan assembly comprising: a contact holder;an elastomeric sealing member at least partially disposed over the contact holder, wherein the elastomeric sealing member comprises an upper portion with a first hole and a second hole extending therethrough and a lower portion;a transcutaneous glucose sensor having a first sensor electrical contact and a second sensor electrical contact on an ex vivo portion of the transcutaneous glucose sensor, wherein at least a portion of the ex vivo portion of the transcutaneous glucose sensor is positioned between the upper portion of the sealing member and the lower portion of the sealing member;a first sealing member electrical contact positioned in the first hole in the upper portion of the elastomeric sealing member; anda second sealing member electrical contact positioned in the second hole in the upper portion of the elastomeric sealing member. 2. The system of claim 1, wherein the elastomeric sealing member has a durometer hardness of from about 5 Shore A to about 80 Shore A. 3. The system of claim 1, wherein the elastomeric sealing member comprises at least one gap maintained between a portion of the upper portion and a portion of the lower portion, when the first sealing member electrical contact is operably connected to the electronics unit. 4. The system of claim 3, wherein at least a portion of the ex vivo portion of the transcutaneous glucose sensor at least partially extends into the gap. 5. The system of claim 1, wherein the elastomeric sealing member comprises a material selected from the group consisting of silicone, silicone/polyurethane hybrid, polyurethane, polysulfide, and mixtures thereof. 6. The system of claim 1, wherein the elastomeric sealing member comprises a material selected from silicone, silicone/polyurethane hybrid, polyurethane, and polysulfide. 7. The system of claim 1, wherein the first sensor electrical contact is adjacent to the first hole extending through the upper portion of the elastomeric sealing member and the first sealing member electrical contact has a first end in contact with the first sensor electrical contact, wherein the second sensor electrical contact is adjacent to the second hole extending through the upper portion of the elastomeric sealing member and the second sealing member electrical contact has a first end in contact with the second sensor electrical contact. 8. The system of claim 7, wherein the first sealing member electrical contact has a second end exposed to an outside of the elastomeric sealing member to engage a first electrical contact of sensor electronics, wherein the second sealing member electrical contact has a second end exposed to the outside of the elastomeric sealing member to engage a second electrical contact of the sensor electronics. 9. The system of claim 8, wherein the second end of the first sealing me... | DexCom, Inc. | A61B005/1486 A61B005/1495 A61B005/1473 A61B005/1468 A61B005/145 A61B017/34 A61B005/00 A61B005/05 A61B005/15 A61M005/142 A61M005/158 A61M005/172
US10712837 | B1 | 2020-07-14 | Using geo-registered tools to manipulate three-dimensional medical images | 2019-07-29 | A virtual object that represents a tangible object is displayed in a three-dimensional image. Responsive to a first input which may be movement of the tangible object, the virtual object moves within the three-dimensional image. Alteration of how the three-dimensional image is displayed occurs responsive to a second input. For example, virtual medical tools can be controlled by manipulating tangible tools to perform procedures and tasks virtually. Eye tracking can be used to generate variable angular resolution in the field of view in a head display unit with which the three-dimensional image is displayed. Digital objects may be used to cause the viewer's eyes move from feature to feature in a saccades fashion or a smooth tracking fashion. | 1. An apparatus comprising: a processor, which processes a 3D image dataset of a patient;a head display unit (HDU) with both position tracking and orientation tracking on a geo-registration coordinate system;a geo-registered tool with both position tracking and orientation tracking on the geo-registration coordinate system;wherein the processor performs processing comprising: registering the HDU and the geo-registered tool on the geo-registration coordinate system;registering the 3D image dataset of the patient onto the geo-registered tool;processing HDU data comprising position tracking and orientation tracking of the HDU on the geo-registration coordinate system;processing geo-registered tool data comprising position tracking and orientation tracking of the geo-registered tool on the geo-registration coordinate system; anddisplaying the 3D image dataset of the patient on the HDU. 2. The apparatus of claim 1 further comprising moving the geo-registered tool and a geo-registration device tracks location and orientation of the geo-registered tool and the processor causes the 3D image dataset of the patient to move with the geo-registered tool. 3. The apparatus of claim 1 wherein the geo-registered tool comprises a focal point pen. 4. The apparatus of claim 3 further comprising wherein the focal point pen designates a focal point in the 3D image dataset of the patient. 5. The apparatus of claim 3 further comprising wherein the focal point pen adds annotations to the 3D image dataset of the patient. 6. The apparatus of claim 3 further comprising wherein the focal point pen draws in the 3D image dataset of the patient. 7. The apparatus of claim 1 wherein the geo-registered tool comprises a mobile platform. 8. The apparatus of claim 7 further comprising wherein the platform moves a selected volume of the 3D image dataset of the patient independently from a non-selected volume of the 3D image dataset of the patient. 9. The apparatus of claim 1 wherein the geo-registered tool comprises a mobile multi-function tool. 10. The apparatus of claim 9 further comprising wherein the multi-function tool virtually grasps and manipulates selected tissue in the 3D image dataset of the patient. 11. The apparatus of claim 9 further comprising wherein the multi-function tool virtually ablates voxels from a tissue type of interest in the 3D image dataset of the patient. 12. The apparatus of claim 1 wherein the geo-registered tool comprises a mobile cutting tool. 13. The apparatus of claim 12 further comprising wherein the mobile cutting tool virtually cuts tissue in the 3D image dataset of the patient. 14. The apparatus of claim 1 wherein the geo-registered tools is a master control platform wherein the master control platforms performs at least one of the group comprising: performing geo-registration of additional geo-registered tools;changing the orientation of the 3D image dataset of the patient;translating the 3D image dataset of the patient;enlarging t... | - | G06T015/00 G06F003/0346 G06F003/0354 G06F003/0481 G06F003/0484 A61B005/06 G06T019/20
US10716514 | B1 | 2020-07-21 | System and method for optimized independent component selection for automated signal artifact removal to generate a clean signal | 2018-02-07 | Described is a system for automated artifact removal to generate a clean signal. During operation, the system selects initial noise components from a multi-channel, pre-processed signal by performing independent component analysis decomposition on the pre-processed signal to separate and rank the independent components as noise components. A clean signal is then generated through optimized selection of the noise components based on a signal quality index in which the noise components are moved from the original pre-processed signal until a sufficient signal quality is received. | 1. A system for automated artifact removal to generate a clean signal, the system comprising: one or more processors and a memory, the memory being a non-transitory computer-readable medium having executable instructions encoded thereon, such that upon execution of the instructions, the one or more processors perform operations of: pre-processing a signal to generate a multi-channel, pre-processed signal;generating a plurality of independent components by performing independent component analysis decomposition on the multi-channel pre-processed signal;obtaining baseline signal quality measurements of the independent components based on a signal quality index (SQI);ranking the independent components based on known nuisance signals to generate a ranking of independent components from highest to lowest, otherwise ranking the independent components in random order;adding one or more independent components to a current noise components list according to the rankings derived from SOI changes upon removalgenerating a clean signal through optimized selection of the noise components based on a SQI and removal of independent components from the multi-channel, pre-processed signal, resulting in the clean signal; andrecording the clean signal on a computer program product. 2. The system as set forth in claim 1, wherein generating the clean signal further comprises operations of: removing the independent components that were added to the current noise components list from each channel in the multi-channel, pre-processed signal;evaluating a change in signal quality by comparing a current signal quality index with a baseline signal quality or a signal quality index from a previous iteration;if signal quality increases, adding a next highest ranking independent component in the ranking of independent components to the noise components list; andif signal quality decreases, removing a lowest ranking independent component in the ranking of independent components from the noise components list. 3. The system as set forth in claim 2, wherein the process of generating the clean signal is repeated until a minimum number of independent components remain unselected or a change in signal quality is below a predetermined threshold. 4. The system as set forth in claim 1, wherein the process of generating the clean signal is repeated until a minimum number of independent components remain unselected or a change in signal quality is below a predetermined threshold. 5. A computer program product for automated artifact removal to generate a clean signal, the computer program product comprising: a non-transitory computer-readable medium having executable instructions encoded thereon, such that upon execution of the instructions by one or more processors, the one or more processors perform operations of: pre-processing a signal to generate a multi-channel, pre-processed signal;generating a plurality of independent components by performing independent component analysis decomp... | HRL Laboratories, LLC | A61B005/00 A61B005/0476
US10715904 | B2 | 2020-07-14 | Wearable devices with increased adhesion | 2019-10-25 | Headphone ear tips are made with an outer surface that changes its coefficient of friction. The outer surface is coated with photochromic compound, hierarchical microstructures, fibers formed through electrostatic flocking, or a combination thereof. | 1. An apparatus comprising: a human-interface component of a wearable device comprising an ear tip for use with a headphone, the component having an outer surface shaped to contact a portion of the body of a user, wherein:(i) the outer surface is coated with hierarchical microstructures comprising a polymer, formed through chemical vapor deposition; wherein the microstructures are coated with a photochromic compound that is a zwitterion in its open form;(ii) the outer surface comprises a polymer and is coated with a photochromic compound that is a zwitterion in its open form;(iii) the component is made from a composition comprising a polymer and a photochromic compound that is a zwitterion in its open form; or(iv) the outer surface is coated with fibers, formed through electrostatic flocking, wherein the outer surface coated with fibers comprises a polymer and is further coated with a photochromic compound that is a zwitterion in its open form; or(v) any combination thereof,wherein the photochromic compound comprises a functional group suitable for attachment to the polymer. 2. The apparatus of claim 1, wherein the microstructures, photochromic compound, fibers, or a combination thereof, provide the outer surface of the component with increased adhesion. 3. The apparatus of claim 1, wherein the ear tip comprises an elastomer. 4. The apparatus of claim 3, wherein the elastomer is silicone, polynorbornene, fluoroelastomer, styrenic-based thermoplastic elastomer, polyacrylates, hydrogenated nitrile rubber, or a shape-memory polymer, or a mixture thereof. 5. The apparatus of claim 4, wherein the shape-memory polymer is polyurethane. 6. The apparatus of claim 3, wherein an outer surface of the ear tip is coated with further hierarchical microstructures, formed through chemical vapor deposition. 7. The apparatus of claim 3, wherein an outer surface of the ear tip is coated with further hierarchical microstructures, formed through chemical vapor deposition; wherein the further microstructures are coated with a photochromic compound that is a zwitterion in its open form. 8. The apparatus of claim 3, wherein an outer surface of the ear tip is coated with a photochromic compound that is a zwitterion in its open form. 9. The apparatus of claim 3, wherein the ear tip is made from a composition comprising a photochromic compound that is a zwitterion in its open form. 10. The apparatus of claim 9, wherein an outer surface of the ear tip is coated with a photochromic compound that is a zwitterion in its open form, wherein the photochromic compound on the outer surface of the ear tip is the same or different from the photochromic compound of the composition used to make the ear tip. 11. The apparatus of claim 3, wherein the photochromic compound provides the ear tip with improved grip. 12. The apparatus of claim 3, wherein the photochromic compound is: 1′-(2-Hydroxyethyl)-3′,3′-dimethyl-6-nitrospiro[1(2H)-benzopyran-2,2′-indoline];1′,3′,3′-Trimethylsp... | Bose Corporation | H04R001/10 A61B005/00 A61B005/024
US10713729 | B2 | 2020-07-14 | Sensor systems and methods for activity evaluation | 2019-09-09 | Systems and methods are discussed for providing sensor enhanced safety, recovery, and activity evaluation systems. Sensors that monitor user activity and behavior are worn by a user and/or placed in the user environment. Data from the sensors are processed to obtain a safety, recovery, and/or activity evaluation. Based on the evaluation, recommendations or adjustments to the terms of an insurance policy covering the user, the user's employer, or a facility providing health care to the user, aer generated, to accurately reflect the risks associated with the user, employer, and/or facility. In embodiments, an alert may be generated when a failure to conform with activity guidelines is detected. | 1. An activity evaluation system, comprising: a communications device configured to receive remote sensor data based on at least one sensor proximate to an individual associated with and distinct from a covered entity;a data storage device storing at least one activity guideline associated with the individual;a computer hardware server, in communication with the communications device and the data storage device, and operated by a risk management entity remote from the at least one sensor, configured to: analyze the received sensor data to determine activity characteristics of the individual;output an activity evaluation based on the activity characteristics and the at least one stored activity guideline;adjust data corresponding to a parameter of a risk management policy covering the covered entity based on the sensor data and the activity evaluation, wherein the computer hardware server is configured to adjust the data corresponding to the parameter so as to provide an adjustment favorable to the covered entity responsive to the activity evaluation being indicative of conformance by the individual with the at least one activity guideline; andresponsive to the activity evaluation being indicative of non-conformance by the individual with the at least one activity guideline, transmit an alert to a computing device of the covered entity. 2. The system of claim 1, wherein the individual associated with and distinct from the covered entity comprises an employee of the covered entity returning from an injury; wherein the at least one activity guideline comprises a recovery plan prescribed for the employee;wherein the sensor data comprises data indicative of one or more of body movement, body orientation, frequency of movement, duration of movement, movement distance, and calories expended, calculated based upon raw sensor data;wherein the computer hardware server is further configured to determine conformance or non-conformance by the individual with the at least one activity guideline, the determining comprising determining, based on the sensor data, whether the employee has conformed with the recovery plan prescribed for the employee; andwherein transmitting the alert to the covered entity comprises, responsive to a negative determination that the employee has conformed with the recovery plan, transmitting the alert to a local computer at the covered entity. 3. The system of claim 2, wherein the raw sensor data comprises data obtained from one or more three-axis accelerometers worn by the employee; and wherein the computer hardware server is further configured to: calculate a frequency of movement based on a series of the sensor data received from the one or more three-axis accelerometers; andcalculate a duration of movement based on a weighted average of movements over time identified in the series of the sensor data;wherein the computer hardware server is configured to adjust the data corresponding to the parameter of the risk management policy ... | Hartford Fire Insurance Company | G06Q040/08 A61B005/103 A61B005/11 A61B005/00 G16H050/30 G06Q010/10 G06F019/00
US10709347 | B1 | 2020-07-14 | Heart graphic display system | 2019-12-11 | A system is provided for displaying heart graphic information relating to sources and source locations of a heart disorder to assist in evaluation of the heart disorder. A heart graphic display system provides an intra-cardiogram similarity (“ICS”) graphic and a source location (“SL”) graphic. The ICS graphic includes a grid with the x-axis and y-axis representing patient cycles of a patient cardiogram with the intersections of the patient cycle identifiers indicating similarity between the patient cycles. The SL graphic provides a representation of a heart with source locations indicated. The source locations are identified based on similarity of a patient cycle to library cycles of a library cardiogram of a library of cardiograms. | 1. A method performed by one or more computing systems for presenting information to assist evaluation of an electromagnetic (“EM”) source of a body, the method comprising: accessing indications of a plurality of cycles within EM measurements of an EM field of the EM source;for pairs of cycles of the plurality of the cycles, calculating a similarity score for the pair of cycles, each pair of cycles including a first cycle of the pair and a second cycle of the pair, the similarity score for the pair of cycles indicating similarity between the first cycle of the pair and the second cycle of the pair;generating a graphic that provides a non-textual indication of the similarity scores for the pairs of cycles; andoutputting the graphic to an output devicewherein the EM measurements are generated based on a computational model that models EM activations of the EM source. 2. The method of claim 1 further comprising applying the computational model to perform a simulation that models the EM activations of the EM source and deriving the EM measurements from the modeled EM activations. 3. The method of claim 1 wherein the non-textual indications of similarity scores are based on varying a graphic characteristic. 4. The method of claim 3 wherein the graphic characteristic is based on color. 5. The method of claim 3 wherein the graphic characteristic is based on intensity level. 6. A method performed by one or more computing systems for presenting information to assist evaluation of an electromagnetic (“EM”) source of a body, the method comprising: accessing indications of a plurality of cycles within EM measurements of an EM field of the EM source;for pairs of cycles of the plurality of the cycles, calculating a similarity score for the pair of cycles, each pair of cycles including a first cycle of the pair and a second cycle of the pair, the similarity score for the pair of cycles indicating similarity between the first cycle of the pair and the second cycle of the pair;generating a graphic that provides a non-textual indication of the similarity scores for the pairs of cycles; andoutputting the graphic to an output devicewherein the graphic includes a map with a first axis representing each cycle and a second axis representing each cycle, and wherein an intersection of a pair of cycles on the map provides the non-textual indication of the similarity score for the pair of cycles. 7. The method of claim 6 wherein the EM measurements are collected from a person. 8. The method of claim 6 wherein the graphic further includes a third axis representing similarity score and wherein the non-textual indication of the similarity score for a pair of cycles is based on height along the third axis. 9. The method of claim 6 wherein the output device is a display device and further comprising receiving a selection of a pair of cycles and displaying information relating to the selected pair of cycles. 10. One or more computing systems for presenting information t... | VEKTOR MEDICAL, INC. | A61B005/00 A61B005/04 A61B005/0456 A61B005/0452 A61B005/044 A61B005/0472 G06N020/00 A61B005/046 G16H050/20 A61B005/0468
US10706602 | B2 | 2020-07-07 | Methods and apparatus to capture patient vitals in real time during an imaging procedure | 2018-11-21 | Apparatus, systems, and methods to capture and combine patient vitals and image data are disclosed. An example apparatus includes a video capturing device or an audio receiving device, a vitals data manager, and a vitals aggregator. The video capturing device captures visual vital information of a patient from a vital monitor during an imaging procedure. The audio receiving device captures audible vital information of the patient during the imaging procedure. The vitals data manager receives the captured visual vital information or the captured audible vital information, the captured visual vital information or the captured audible vital information to be tagged with an identifier of the patient to form tagged vitals information. The vitals aggregator receives the tagged vitals information and the image associated with the patient, the vitals aggregator to organize the tagged vitals information with the image associated with the patient to form a composite image. | 1. An apparatus to capture patient vitals in real time during an imaging procedure, the apparatus comprising: at least one of a video capturing device or an audio receiving device, the video capturing device configured to capture a visual vital information of a patient from a vital monitor during the imaging procedure and the audio receiving device configured to capture an audible vital information of the patient during the imaging procedure;a vitals data manager to: obtain at least one of the visual vital information or the audible vital information, the at least one of the visual vital information or the audible vital information tagged with an identifier of the patient;obtain a medical image of the patient, the medical image captured during the imaging procedure; anddetermine, utilizing the identifier of the patient, that the at least one of the visual vital information or the audible vital information correspond to the medical image of the patient; anda vitals aggregator to: when the at least one of the visual vital information or the audible vital information does not correspond to the medical image, trigger a recapture of the at least one of the visual vital information or the audible vital information; andembed the at least one of the visual vital information or the audible vital information with the medical image corresponding to the patient to form a composite image. 2. The apparatus as defined in claim 1, further including a composite image generator to output the composite image. 3. The apparatus as defined in claim 1, wherein the at least one of the video capturing device and the audio receiving device are to simultaneously capture the visual vital information and the audible vital information. 4. The apparatus as defined in claim 1, wherein the at least one of the video capturing device is to capture an image of the patient. 5. The apparatus as defined in claim 1, wherein the at least one of the video capturing device is a 360 degree rotatable camera that is to capture the visual vital information in one or more image frames and to capture an image of the patient. 6. The apparatus as defined in claim 1, wherein the video capturing device is to utilize optical character recognition to capture the visual vital information from the vital monitor. 7. The apparatus as defined in claim 1, wherein the vitals data manager is to: determine inconsistencies of the visual vital information and the audible vital information; andresolve the inconsistencies between the visual vital information and the audible vital information to generate consistent vital information. 8. The apparatus as defined in claim 1, wherein the visual vital information or the audible vital information is to be embedded in a header of a file, the file including the medical image of the patient. 9. The apparatus as defined in claim 1, further including a user interface to: provide instructions to select the patient from a plurality of patients;activate the video captu... | General Electric Company | G06T011/60 G06K009/03 H04R001/02 A61B005/00 A61B007/04 G16H040/63 G16H030/20 G06K009/00 G06K009/62 G06F003/0482 H04N005/232
US10702694 | B2 | 2020-07-07 | Systems, devices, and methods for treating vestibular conditions | 2019-08-30 | Apparatus and methods are described herein that provide a vibratory device that can apply a vibratory signal to a portion of a head of a user such that the vibratory signal can be conducted via bone to a vestibular system of the user and cause a portion of the vestibular system to move in a manner equivalent to that of a therapeutically effective vibratory signal applied to an area overlaying a mastoid bone of the user. The vibratory device can be associated with frequencies less than 200 Hz. The vibratory device can be effective at treating a physiological condition associated with the vestibular system. | 1. A method, comprising: applying, via a vibratory device positioned over an area of a head of a user, one or more vibratory signals to the area of the head of the user such that a portion of the one or more vibratory signals are conducted via a temporal bone to a vestibular system of the user, the one or more vibratory signals having a frequency between 30 and 80 Hz and a force level greater than 87 dB re 1 dyne and less than or equal to 101 dB re 1 dyne,the applying the one or more vibratory signals to the area of the head of the user causing movement of a portion of the vestibular system that treats a physiological condition associated with the vestibular system. 2. The method of claim 1, wherein the vibratory device has a resonant frequency, the method further comprising: supplying an electrical signal to the vibratory device to cause the vibratory device to vibrate to generate the one or more vibratory signal that are applied to the area of the head of the user, the electrical signal having a signal frequency;measuring, using a sensor, information including at least one of: a current of the electrical signal, a voltage change of the electrical signal across the vibratory device, a magnetic field generated near the vibratory device, and an acceleration of the vibratory device; andadjusting the signal frequency of the electrical signal based on the information such that the electrical signal causes the vibratory device to vibrate at the resonant frequency of the vibratory device. 3. The method of claim 1, wherein the physiological condition includes at least one of: vertigo, dizziness, motion sickness, virtual reality sickness, spatial discordance, sopite syndrome, nausea, headache, migraine, or tinnitus. 4. The method of claim 1, further comprising maintaining the vibratory device over the area of the head using at least one of: a rigid or elastic headband, a helmet, or an earcup. 5. The method of claim 1, wherein the vibratory device includes a magnet that is configured to vibrate to generate the one or more vibratory signal that are applied to the area of the head of the user. 6. The method of claim 1, further comprising: generating an electrical signal;amplifying the electrical signal to produce an amplified electrical signal; andsupplying the amplified electrical signal to the vibratory device to cause the vibratory device to generate the one or more vibratory signal that are applied to the area of the head of the user. 7. A method, comprising: generating, via a signal generator of a vibratory device, an electrical signal having a signal frequency;delivering the electrical signal to a set of coils of the vibratory device such that the set of coils generates an alternating magnetic field;causing, via the alternating magnetic field, a vibratory element of the vibratory device to vibrate and generate one or more mechanical vibratory signals; andtransferring, via a housing of the vibratory device in contact with an area of a head of a u... | Otolith Sound Inc. | A61H023/02 A61N001/36 A61B005/00 A61F011/04 A61F011/00 A61M021/00 A61N007/00 H04R005/033 H04R005/02
US10716534 | B1 | 2020-07-21 | Base station for a digital stethoscope, and applications thereof | 2019-10-21 | Embodiments disclosed herein improve digital stethoscopes and their application and operation. A first method detects of a respiratory abnormality using a convolution. A second method counts coughs for a patient. A third method predicts a respiratory event based on a detected trend. A fourth method forecasts characteristics of a future respiratory event. In a fifth embodiment, a base station is provided for a digital stethoscope. | 1. A base station for a digital stethoscope comprising: a housing configured to accept the digital stethoscope;a wireless charging unit, located within the housing, configured to charge the digital stethoscope when it rests on the exterior of the housing, and further configured to detect when the digital stethoscope becomes detached from the housing;a communication unit, coupled to the wireless charging unit, configured to receive a noise signal from the digital stethoscope when the wireless charging unit detects the digital stethoscope becoming detached from the housing; anda control unit, coupled to the communication unit, configured to analyze the noise signal for a respiratory event; andwhereinthe communication unit is further configured to communicate the respiratory event to a cloud-based service. 2. The base station of claim 1 further comprising: an air quality sensor, coupled to the control unit, configured to detect particulates in the air surrounding the housing; andwhereinthe communication unit is configured to communicate data from the air quality sensor to the cloud-based service. 3. The base station of claim 1 further comprising: an air quality sensor, coupled to the control unit, configured to detect particulates in the air surrounding the housing; andwhereinthe control unit is configured to predict a respiratory condition based on the detected particulates and the respiratory event. 4. The base station of claim 1 wherein the communication unit includes an Internet of Things (IoT) modem. 5. The base station of claim 1 wherein the communication unit includes a Bluetooth circuit. 6. The base station of claim 1 wherein the communication unit enables communication with a device associated with a user. 7. The base station of claim 1 wherein the control unit includes a field programmable gate array (FPGA) with logic trained to predict a respiratory condition from the respiratory event. 8. The base station of claim 7 further comprising a storage unit, coupled to the control unit; andwhereinthe control unit includes a microprocessor configured to use a convolutional neural network to determine whether the respiratory event includes the respiratory condition, and wherein the FPGA is configured to calculate convolutions for the convolutional neural network, and wherein the microprocessor receives the calculated convolutions and stores the calculated convolutions in the storage unit. 9. The base station of claim 8 wherein the control unit is configured to predict a future respiratory condition based on the respiratory event. 10. The base station of claim 9 wherein the control unit is configured to detect the respiratory condition over time, and wherein the control unit is configured to use a long short term memory (LSTM) network, including one or more weights, to predict the future respiratory condition based on the respiratory condition, and wherein the weights for the LSTM network are stored in the FPGA. 11. The base station of cl... | Sonavi Labs, Inc. | A61B005/08 A61B007/04 A61B007/00 A61B005/00 G16H040/63
US10709353 | B1 | 2020-07-14 | Detecting a respiratory abnormality using a convolution, and applications thereof | 2019-10-21 | Embodiments disclosed herein improve digital stethoscopes and their application and operation. A first method detects of a respiratory abnormality using a convolution. A second method counts coughs for a patient. A third method predicts a respiratory event based on a detected trend. A fourth method forecasts characteristics of a future respiratory event. In a fifth embodiment, a base station is provided for a digital stethoscope. | 1. A computer-implemented method for detection of a respiratory abnormality comprising: receiving, from a microphone, an auditory signal;generating an auditory spectrogram based on the auditory signal;performing a spatial convolution procedure on the auditory spectrogram to generate one or more convolution values, wherein the convolution values represent a compressed version of a matrix portion of the auditory spectrogram, and wherein the spatial convolution procedure is trained to generate one or more features for detecting the respiratory abnormality;applying one or more weights to the convolution values, wherein the weights are trained for detection of the respiratory abnormality;generating a classification value based on applying the weights to the convolution values, wherein the classification value indicates whether the auditory signal includes the respiratory abnormality; andtransmitting the classification value to a display device for outputting the classification value. 2. The method of claim 1 wherein performing the spatial convolution procedure enables the classification value to be output in real time after the auditory signal is received from the microphone. 3. The method of claim 1 further comprising: receiving a further auditory signal collected contemporaneously with the auditory signal, from a further microphone, wherein the further auditory signal represents a background noise; andadjusting, before generating the auditory spectrogram, the auditory signal based on the further auditory signal to reduce the background noise. 4. The method of claim 1 further comprising: receiving an accelerometer data, from an accelerometer coupled to the microphone, wherein the accelerometer data is collected contemporaneously with the auditory signal; andadjusting, before generating the auditory spectrogram, the auditory signal based on the accelerometer data to reduce noise from movement of the microphone from a patient's breathing. 5. The method of claim 1 further comprising: receiving an accelerometer data, from an accelerometer coupled to the microphone, wherein the accelerometer data is collected contemporaneously with the auditory signal, and wherein the accelerometer data represents a chest motion;analyzing the accelerometer data to extract a respiratory cycle based on the chest motion; andadjusting, before generating the auditory spectrogram, the auditory signal based on the respiratory cycle. 6. The method of claim 1 further comprising filtering, before generating the auditory spectrogram, the auditory signal according to a cochlear model. 7. The method of claim 6 further comprising: dividing the auditory signal into one or more frequency bands; andwhereinfiltering the auditory signal according to the cochlear model is based on filtering the frequency bands. 8. The method of claim 1 wherein performing the spatial convolution procedure is performed repeatedly to successively compress the auditory spectrogram into a convolutional mat... | Sonavi Labs, Inc. | A61B005/08 G06N003/08 G06K009/62 A61B007/00 A61B007/04 A61B005/00
US10722126 | B2 | 2020-07-28 | Heart rate detection method and device thereof | 2018-11-28 | A heart rate detection method includes a facial image data acquiring step, a feature points recognizing step, an effective displacement signal generating step and a heart rate determining step. The feature points recognizing step is for recognizing a plurality of feature points, wherein a number range of the feature points is from three to twenty, and the feature points include a center point between two medial canthi, a point of a pronasale and a point of a subnasale of the face. The effective displacement signal generating step is for calculating an original displacement signal, wherein the original displacement signal is converted to an effective displacement signal. The heart rate determining step is for transforming the effective displacement signals of each of the feature points to an effective spectrum, wherein a heart rate is determined from one of the effective spectrums corresponding to the feature points, respectively. | 1. A heart rate detection method, comprising: a facial image data acquiring step for acquiring a plurality of frames of facial image data of a face;a feature points recognizing step for recognizing a plurality of feature points, which are pre-determined, of the face from the plurality of frames of facial image data, wherein a number range of the feature points is from three to twenty, and the feature points comprise a center point between two medial canthi, a point of a pronasale and a point of a subnasale of the face;an effective displacement signal generating step for calculating an original displacement signal of each frame time of each of the feature points from the plurality of frames of facial image data, wherein the original displacement signal is established based on an original horizontal displacement and an original vertical displacement, and converted to an effective displacement signal; anda heart rate determining step for transforming the effective displacement signals of each of the feature points to an effective spectrum, wherein a heart rate is determined from one of the effective spectrums corresponding to the feature points, respectively. 2. The heart rate detection method of claim 1, wherein, in the feature points recognizing step, a plurality of landmark points of the face are recognized from the plurality of frames of facial image data by an algorithm of ensemble of regression tree, a number range of the landmark points is from forty to ninety, and the feature points being pre-determined of the face are recognized from the landmark points. 3. The heart rate detection method of claim 1, wherein the number range of the feature points is from seven to twenty, and the feature points further comprise a center point between two lateral canthi, a center point of a right eye, a point of a sellion and a point of a menton of the face. 4. The heart rate detection method of claim 1, wherein, in the effective displacement signal generating step, each of the original displacement signals is calculated and converted to the effective displacement signal by a weight of the original horizontal displacement and a weight of the original vertical displacement, and the weight of the original horizontal displacement is smaller than the weight of the original vertical displacement. 5. The heart rate detection method of claim 1, wherein, in the effective displacement signal generating step, each of the original displacement signals is calculated and converted to the effective displacement signal by a fixed-point algorithm of a fast independent component analysis and a bandpass filter. 6. The heart rate detection method of claim 1, wherein, in the heart rate determining step, the effective displacement signals of each of the feature points are transformed to the effective spectrum by a Fourier transform, there is a maximum relative power in each of the effective spectrums corresponding to the feature points, respectively, and the heart rate is a ... | NATIONAL YUNLIN UNIVERSITY OF SCIENCE AND TECHNOLOGY | A61B005/024 G06T007/73 G06K009/00 G06T007/00 A61B005/00 A61B005/11
US10722000 | B2 | 2020-07-28 | Dynamic fit footwear | 2019-11-15 | A foot presence sensor system for an active article of footwear can include a sensor housing configured to be disposed at or in an insole of the article, and a controller circuit, disposed within the sensor housing, configured to trigger one or more automated functions of the footwear based on a foot presence indication. In an example, the sensor system includes a capacitive sensor configured to sense changes in a capacitance signal in response to proximity of a body. A dielectric member can be provided between the capacitive sensor and the body to enhance an output signal from the sensor. | 1. A method for automatically adjusting an active article of footwear during a footwear use period that is between donning and doffing events, the method comprising: sensing, using a body sensor in the article of footwear, a change in a position of a foot inside the article of footwear; andin response to the change in the position of the foot as sensed by the body sensor, controlling a lacing engine to adjust a fit of the article of footwear relative to the foot. 2. The method of claim 1, further comprising generating, using a processor circuit, a lacing engine control signal for the lacing engine, wherein the lacing engine control signal indicates whether the change in the foot position as sensed by the body sensor meets a response criteria, and wherein controlling the lacing engine includes in response to the lacing engine control signal. 3. The method of claim 2, further comprising receiving a sensor signal from the body sensor, the sensor signal indicating a position of the foot over time; and wherein the processor circuit is configured to determine whether the change in the foot position meets the response criteria by comparing the sensor signal to a signal morphology limit. 4. The method of claim 2, wherein generating the lacing engine control signal includes using the processor circuit to characterize a location of the foot relative to the body sensor. 5. The method of claim 2, wherein generating the lacing engine control signal includes using the processor circuit to characterize an orientation of the foot relative to the body sensor. 6. The method of claim 2, wherein generating the lacing engine control signal includes using the processor circuit to characterize a rotation of the foot relative to a toothed of the article of footwear. 7. The method of claim 1, wherein controlling the lacing engine includes suppressing actuation of the lacing engine to maintain a tension characteristic of the article of footwear about the foot. 8. The method of claim 1, wherein controlling the lacing engine includes actuating the lacing engine to decrease a tension characteristic of the article of footwear about the foot. 9. The method of claim 1, further comprising detecting a gesture signal from the wearer, using a gesture sensor, and wherein controlling the lacing engine includes in response to the gesture signal. 10. The method of claim 1, wherein sensing the change in the foot position includes changing a baseline or reference characteristic of the body sensor based on an environment in which the article of footwear is used. 11. The method of claim 1, further comprising: generating a drive signal for the body sensor using a signal generator; andreceiving the drive signal at the body sensor and, in response, generating an electric field in the article of footwear and adjacent to the body sensor;wherein sensing the change in the foot position includes identifying a change in the electric field adjacent to the body sensor that is caused at leas... | NIKE, Inc. | A43C011/00 A43C011/16 A61B005/00 G01D005/12 G01L001/14 A43B001/00 A43B003/00 G01L005/12 G01L005/24 A43B013/14 A43B017/00 A43C001/00 G01D005/24 G05B015/02 A43C007/00 G01L001/12 G05B019/048 G01D005/14 G01D005/34
US10709363 | B2 | 2020-07-14 | Analyte sensor | 2019-11-21 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A device for measuring a glucose concentration in a host, the system comprising: a transcutaneous glucose sensor;sensor electronics configured to operably connect to the transcutaneous glucose sensor;at least one elastomeric electrical contact configured to operably connect the transcutaneous glucose sensor to the sensor electronics;an elastomeric sealing member at least partially surrounding a portion of the transcutaneous glucose sensor and at least partially surrounding the at least one elastomeric electrical contact;a contact holder, wherein each of the portion of the transcutaneous glucose sensor, the at least one elastomeric electrical contact, and the elastomeric sealing member is at least partially disposed over the contact holder; anda mounting unit configured to operably connect to the contact holder,wherein a durometer hardness of the at least one elastomeric electrical contact is higher than a durometer hardness of the elastomeric sealing member. 2. The device of claim 1, wherein the at least one electrical contact is formed from a material having a durometer hardness of from about 5 Shore A to about 80 Shore A. 3. The device of claim 2, wherein the sealing member is formed from a material having a durometer hardness of from about 5 Shore A to about 80 Shore A. 4. The device of claim 2, wherein the durometer hardness of the at least one electrical contact is about 50 Shore A. 5. The device of claim 1, wherein the durometer hardness of the sealing member is about 20 Shore A to about 50 shore A. 6. The device of claim 1, wherein the durometer hardness of the sealing member is about 50 Shore A. 7. The device of claim 1, further comprising a sealant adjacent to the sealing member. 8. The device of claim 1, wherein the at least one electrical contact extends through a portion of the sealing member. 9. The device of claim 1, wherein the mounting unit and the contact holder are connected to the sensor electronics. 10. A device for transcutaneous measurement of a glucose concentration in a host, the device comprising: a transcutaneous glucose sensor;sensor electronics configured to operably connect to the transcutaneous glucose sensor;one or more electrical contacts configured to operably connect the transcutaneous glucose sensor to the sensor electronics;a sealing member at least partially surrounding a portion of the transcutaneous glucose sensor and at least partially surrounding the one or more electrical contacts;a contact holder over which the portion of the transcutaneous glucose sensor, the one or more electrical contacts, and the sealing member are at least partially disposed; anda mounting unit configured to operably connect to the contact holder,wherein a durometer hardness of at least one of the one or more electrical contacts is higher than a durometer hardness of the sealing member, wherein the sealing member comprises an upper portion and a lower portion, wherein the portion of the transcutaneous glucose sensor is p... | DexCom, Inc. | A61B005/1468 A61B005/1473 A61B005/1486 A61B005/1495 A61B005/145 A61B005/00 A61B017/34 A61B005/05 A61B005/15 A61M005/158 A61M005/172 A61M005/142
US10722165 | B1 | 2020-07-28 | Systems and methods for reaction measurement | 2019-09-30 | This disclosure relates to systems, media, and methods for quantifying and monitoring visual acuity, motor skill, and cognitive ability of a user through measuring user activities while the user responds to prompts on an electronic device. Disclosed embodiments may receive image data from a user-facing camera, motion data from a sensor device attached to the user, location data from a touchscreen, and time data measuring a user's response times. Disclosed embodiments may provide graphical user interfaces to which the user responds and calculate health metrics based on characteristics of the user's response. Disclosed embodiments may calculate visual acuity, motor skill, and cognitive ability metrics based on data measured while the user responds to the graphical user interfaces. Disclosed embodiments may include comparing the metrics to respective baseline values and providing an alert when the metrics deviate from threshold values indicating a change in user health. | 1. A system for measuring reaction times of a user, comprising: an electronic device comprising a touchscreen, a user-facing camera, at least one memory storing operation instructions, and at least one processor;a first sensor device attached to the user, the first sensor device comprising an inertial measurement unit;the at least one processor executing the operation instructions to perform operations comprising: providing, on the touchscreen, a graphical user interface with at least one selection region;updating, at an interval, the graphical user interface to provide at least one object related to the at least one selection region;receiving a user selection on the graphical user interface of the touchscreen;recording, between the updating and the receiving, data comprising: image data, from the user-facing camera, including at least a portion of a face of the user;first motion data, from the first sensor device, representing first motions of the user;location data, from the touchscreen, representing a location of the user selection within the graphical user interface; andtime data, indicating a time period between the updating and the receiving;determining, based on the location data, if the user performed reaction test instructions comprising selecting one of the at least one selection region on the touchscreen when at least one corresponding object appears;calculating, based on the image data, a visual acuity metric as an average percent of a field of view of the camera that includes the user's face;calculating, based on the first motion data, a motor skill metric as an average amount of time the user moves between the updating and the receiving;calculating, based on the time data, a cognitive ability metric as an average time period between the updating and the receiving,comparing the visual acuity metric, the motor skill metric, and the cognitive ability metric to a visual acuity threshold, a motor skill threshold, and a cognitive ability threshold, respectively; andproviding, based on the comparing, an alert based on at least one of the following occurring:the visual acuity metric exceeding the visual acuity threshold,the motor skill metric exceeding the motor skill threshold, orthe cognitive ability metric exceeding the cognitive ability threshold. 2. The system of claim 1, wherein the operations further comprise: storing the visual acuity metric, the motor skill metric, and the cognitive ability metric;setting a baseline value of visual acuity, a baseline value of motor skill, and a baseline value of cognitive ability as an average of the stored visual acuity metrics, stored motor skill metrics, and stored cognitive ability metrics, respectively; andsetting the visual acuity threshold, the motor skill threshold, and the cognitive ability thresholds based on respective baseline values. 3. The system of claim 1, wherein the operations further comprise waiting for a predefined period before updating the graphical user interface. 4. The ... | BioMech Sensor LLC | G09B019/00 A61B005/16 G06F003/01 G06F003/0484 A61B005/11 A61B005/00 G06K009/00
US10709359 | B2 | 2020-07-14 | Systems and methods for dynamically identifying a patient support surface and patient monitoring | 2019-07-12 | Various patient monitoring systems can include a sensor configured to collect three dimensional information. The systems can identify a location of a patient support surface based on the three dimensional information. The systems can set a two dimensional planar threshold based on the patient support surface. The systems can identify a patient location above the patient support surface based on the three dimensional information and compare the patient location to the two dimensional planar threshold. Exceeding the threshold can be indicative of a high risk of a patient fall. An alert can be generated based on the threshold being exceeded. The systems can repeat the identification of the patient support surface location and the setting of the threshold to account for changes in the patient area. | 1. A patient monitoring system for monitoring a patient area, the system comprising: a sensor configured to output a plurality of frames, the plurality of frames containing three dimensional pixel information of the patient area;a user interface; andcontrol circuitry configured to: receive the three dimensional pixel information of the plurality of frames;identify a location and dimensions of a patient support surface based on the three dimensional pixel information;set at least one threshold that is extended above, below, or vertically beyond the location and dimensions of the patient support surface;determine a repositioning of the patient support surface;adjust the at least one threshold based on the repositioning of the patient support surface;identify one or more patient locations based on the three dimensional pixel information;compare the one or more patient locations to the at least one threshold; andgenerate an output with the user interface based on the one or more patient locations traversing the at least one threshold. 2. The patient monitoring system of claim 1, wherein the control circuitry is configured to: detect a change in illumination of the patient area; andidentify the one or more patient locations and set the at least one threshold in response to the change in illumination. 3. The patient monitoring system of claim 1, wherein the control circuitry is configured to: set a first of the at least one threshold in a first area of the patient support surface corresponding to a foot of the patient support surface; andset a second of the at least one threshold in a second area of the patient support surface corresponding to a head of the patient support surface. 4. The patient monitoring system of claim 3, wherein the control circuitry is configured to assign a greater warning for traversing the first of the at least one threshold than traversing the second of the at least one threshold. 5. The patient monitoring system of claim 1, wherein the control circuitry is configured to: set the at least one threshold at a plurality of heights corresponding to a plurality of levels of risk to a patient; andissue notifications and alarms corresponding to the plurality of levels of risk to the patient based on traversal of the at least one threshold set at the plurality of heights. 6. The patient monitoring system of claim 1, wherein the control circuitry is configured to: identify a plurality of planar surfaces of the patient support surface; andset the at least one threshold parallel with the plurality of planar surfaces. 7. The patient monitoring system of claim 1, wherein the control circuitry is configured to set the at least one threshold beyond the length or width of the patient support surface. 8. The patient monitoring system of claim 1, wherein the control circuitry is configured to set the at least one threshold at a height that is below the patient support surface and above a floor of the patient area. 9. The patient monito... | CareView Communications, Inc. | H04N007/18 H04N013/02 A61B005/11 A61B005/00 G06T007/136 G06T007/00 G06T007/70 G06F019/00
US10702162 | B2 | 2020-07-07 | Guidewire with internal pressure sensor | 2019-11-21 | The present document describes a pressure guidewire. It comprises a shaft tube having a proximal section and a sensor housing having a distal end, the proximal section and the sensor housing being the continuity of the same shaft tube. A fiber optic pressure sensor is embedded in the sensor housing and comprises an optical fiber extending within the sensor housing, the sensor adapted to measure a pressure of a fluid which is substantially applied in an axis collinear with the longitudinal axis of the guidewire. A marker band, to help in localizing the pressure guidewire in a patient's vessels, is embedded inside the sensor housing and fixed to the optical fiber for holding the pressure sensor inside the sensor housing. | 1. A pressure guidewire having a longitudinal axis comprising: a shaft tube having a proximal section and a sensor housing having a distal end;a fiber optic pressure sensor assembly embedded in the sensor housing and comprising a pressure sensor and an optical fiber extending within the sensor housing, the pressure sensor adapted to measure a pressure of a fluid which is substantially applied in an axis collinear with the longitudinal axis of the guidewire; anda band to support the fiber optic pressure sensor assembly, the band being embedded inside the sensor housing and fixed to the fiber optic pressure sensor assembly for holding the pressure sensor inside the sensor housing. 2. The pressure guidewire of claim 1, further comprising a tip portion extending distally relative to the sensor housing, the tip portion having a proximal end attached to the distal end of the sensor housing, the tip portion being the most distal portion of the pressure guidewire. 3. The pressure guidewire of claim 1, wherein the shaft tube comprises a proximal section, the proximal section and the sensor housing being the continuity of the same shaft tube. 4. The pressure guidewire of claim 1, further comprising a middle section extending further distally relative to the proximal section, the middle section being cut according to a cut pattern thereby resulting in greater flexibility in the middle section than in the proximal section. 5. The pressure guidewire of claim 4, further comprising an inner hypotube and a window giving access to the inner hypotube for welding, soldering or bonding the inner hypotube within the middle section of the pressure guidewire. 6. The pressure guidewire of claim 5, further comprising a small piece of material in contact with a wall of the inner hypotube and protruding within the window, the small piece of material being welded to the inner hypotube. 7. The pressure guidewire of claim 1, wherein the sensor housing defines a cavity, inside the sensor housing, which is distal relative to the pressure sensor and delimited at a proximal end thereof by the pressure sensor. 8. The pressure guidewire of claim 7, wherein the pressure sensor comprises a diaphragm sensitive to a pressure of a fluid. 9. The pressure guidewire of claim 1, wherein the band is a radio opaque marker band. 10. The pressure guidewire of claim 1, wherein the band is spaced apart from the pressure sensor. 11. The pressure guidewire of claim 1, wherein the band is disposed immediately proximal to the pressure sensor. 12. The pressure guidewire of claim 1, wherein the band is bonded to the sensor housing. 13. The pressure guidewire of claim 12, wherein the band is bonded by an adhesive. 14. The pressure guidewire of claim 13, wherein the band is bonded by an adhesive to the fiber optic pressure sensor assembly. 15. The pressure guidewire of claim 13, wherein the band is bonded to the optical fiber of the fiber optic pressure sensor assembly. 16. The pressure gu... | Opsens Inc. | A61B005/00 A61B005/0215 A61M025/09 A61M025/00
US10722118 | B2 | 2020-07-28 | Electronic ecosystem for medical examination room | 2019-07-26 | A system comprises a storage device, a piece of medical equipment, a portable computing device, and a hub portal. The piece of medical equipment has at least one electrically powered feature and comprises either a medical examination table or a storage cabinet. The storage device is operable to store data and is remotely located relative to the first piece of medical equipment and relative to the portable computing device. The hub portal is operable to provide communication of one or both of data or commands between the storage device, the piece of medical equipment, and the portable computing device. A method includes enablement or activation of an electrically powered feature of medical equipment upon entry of a portable computing device into a medical examination room. The method also includes disablement or further activation of the electrically powered feature upon exit of the portable computing device from the medical examination room. | 1. A method comprising: (a) emanating a wireless signal from a communication hub assembly located in an examination room, wherein the communication hub assembly is in communication with a remote server;(b) sensing, by a proximity sensor located in the examination room, the presence or absence of a portable computing device in the examination room;(c) forming a wireless connection between the portable computing device and the communication hub assembly via the wireless signal in response to the proximity sensor sensing the presence of the portable computing device in the examination room; and(d) in response to forming the wireless connection between the portable computing device and the communication hub, enabling access to a medical device located in the examination room,wherein the communication hub assembly maintains communication with the remote server in the absence of the portable computing device in the examination room. 2. The method of claim 1, wherein the enabling access to the medical device comprises activating an electrically powered feature of the medical device. 3. The method of claim 1, further comprising: (a) after enabling access to the medical device, collecting patient data from a patient with the medical device; and(b) transmitting the patient data from the medical device to the communication hub assembly. 4. The method of claim 3, further comprising transmitting the patient data from the communication hub assembly to the portable computing device via the wireless connection. 5. The method of claim 1, further comprising: (a) after enabling access to the medical device, collecting data from the medical device; and(b) transmitting the data from the medical device to the communication hub assembly. 6. The method of claim 5, further comprising transmitting the data from the communication hub assembly to the portable computing device via the wireless connection. 7. The method of claim 1, wherein the medical device is an electrocardiogram (ECG) device. 8. The method of claim 7, further comprising: (a) after enabling access to the medical device, collecting electrocardiogram data from a patient with the medical device; and(b) transmitting the echocardiogram data from the medical device to the communication hub assembly. 9. The method of claim 1, wherein the medical device is a storage cabinet. 10. The method of claim 9, further comprising after enabling access to the medical device, unlocking the storage cabinet. 11. The method of claim 1, wherein the medical device is an examination table. 12. The method of claim 11, further comprising: (a) after enabling access to the medical device, collecting weight data from a patient with the medical device; and(b) transmitting the weight data from the medical device to the communication hub assembly. 13. The method of claim 11, further comprising after enabling access to the medical device, adjusting an electrically powered feature of the medical device. 14. The method of claim 11... | Midmark Corporation | A61B005/00 H04L029/08 G01G019/44 G16H010/60 H04W004/80 A61G013/00 G16H040/63 G16H040/67
US10709345 | B1 | 2020-07-14 | Method and apparatus for noninvasive absolute (mean) intracranial pressure (A-ICP) measurement and/or monitoring | 2019-11-26 | Non-invasive measurement of intracranial pressure (ICP), for example mean ICP. A probe adjacent the head emits energy, such as ultrasound, and receives reflected signals. A processing unit derives ICP waveform from the signals. A pressure mechanism applies external pressure intermittently to outer surface of the head and incrementally increases the external pressure. The processing unit is configured to detect a decrease in amplitude of the ICP waveform (occurring in some embodiments only after an intermediate period of ICRS compensation), the processing unit configured to determine the ICP of the person from a sum of applied external pressures from a time of the initial value A1 until a final value at which the amplitude remains stable with additional increase in applied external pressure. In some cases, the final value is earlier than that but the processing unit extrapolates the sum to when the amplitude remains stable with additional increases in pressure. | 1. A non-invasive method for measuring a mean intracranial pressure (ICP) of a person, comprising: using at least one probe positioned adjacent an exterior tissue layer of the person's head to emit energy at pulsating one or more of brain tissue, brain vessels and brain ventricles in the head and to receive reflected signals;deriving from the signals or from images associated with the signals, using a processor, an ICP waveform having a discernable amplitude whose variations correspond to the pulsating one or more of brain tissue, brain vessels and brain ventricles;applying external pressure intermittently to an outer surface of the head using a pressure mechanism in communication with the processor, the mechanism including a surface, by applying an initial external pressure when the surface of the mechanism contacts the outer surface and then incrementally increasing the external pressure to the outer surface using the pressure mechanism;after a decrease in amplitude of the ICP waveform, the decrease beginning at an initial value A1 of the amplitude, determining, by the processor, the ICP of the person from a sum, or an extrapolated sum, of applied external pressures applied beginning on or after a time of the initial value A1 and ending by a time of the final value A2,wherein one of the following is true:(i) the final value is a value at which the amplitude reaches a predefined absolute or relative amount with additional increase in applied external pressure and the processor is configured to determine the ICP from the sum of applied external pressures,(ii) the final value A2 is a value to which the amplitude has decreased from the initial value A1 such that at the final value the amplitude is moving towards but has not yet reached the amplitude that is the predefined absolute or relative amount with additional increase in applied external pressure, and the processor is configured to determine the ICP or an estimated ICP by extrapolating the sum to a point at which the amplitude reaches the predefined absolute or relative amount with additional increase in applied external pressure. 2. The method of claim 1, wherein the energy is applied using a steerable one-dimensional ultrasound probe placed on a forehead of the person and focused on an area in the brain where the pulsations occur. 3. The method of claim 2, wherein the one dimensional ultrasound probe is operated at a frequency of 0.5-1.5 MHz. 4. The method of claim 2, wherein the area in the brain is at least one of the following: the third ventricle, lateral ventricles, thalamic striate veins and central cerebral veins of the 3rd ventricular cavity, edges of the 3rd and lateral ventricles, 4th ventricle, choroid plexus veins of ventricles, and cortical pulsations. 5. The method of claim 2, wherein the signals are ultrasound signals that are registered on the sagittal plane 20-50 degrees to the horizontal (axial) plane. 6. The method of claim 1, wherein the energy is applied using a st... | - | A61B005/03 A61B005/00 A61B008/08 A61B008/00
US10712172 | B1 | 2020-07-14 | Accelerometer-based systems and methods for quantifying steps | 2019-12-11 | Systems and methods detect steps from one or more sensed accelerometer signals. Systems comprise an accelerometer and a non-transitory computer readable medium, each of which communicates with a processor. The accelerometer is coupled to an individual and generates outputs received by the processor. The non-transitory computer readable medium stores instructions controlling the processor to perform steps of a method. The processor determines a minimum and a maximum reading for respective time periods. Counted peak heights are maxima that exceed a rest threshold. The rest threshold may be a function of the rest maximum. The processor determines a peak threshold for each time period having a counted peak height. A counted amplitude comprises the difference between the maximum reading and minimum readings of a time period in which the counted peak height exceeds the peak threshold. The processor increments a step counter when a counted amplitude exceeds an amplitude threshold. | 1. A system for counting steps, the system comprising: an accelerometer coupled to an individual, the accelerometer generating accelerometer outputs;a processor in communication with the accelerometer; anda non-transitory computer readable medium in communication with the processor, the non-transitory computer-readable medium storing instructions controlling the processor to: determine readings using the accelerometer outputs;identify counted amplitudes, each of the counted amplitudes comprising a difference between a maximum reading and a minimum reading during a time period in which the maximum reading exceeds a rest threshold, the rest threshold determined by: receiving readings from the accelerometer during a period of rest,identifying a rest maximum, the rest maximum comprising a greatest reading generated by the accelerometer during the period of rest, anddetermining the rest threshold using the rest maximum;determine an amplitude threshold for each time period having a counted amplitude, the amplitude threshold comprising a first weighted average of two or more counted amplitudes;identify counted peak heights, each of the counted peak heights comprising the maximum reading during a time period in which the amplitude is greater than the amplitude threshold;determine a peak threshold for each time period having a counted peak height, the peak threshold comprising a first weighted average of two or more counted peak heights; andincrement a counter if the counted peak height exceeds the peak threshold associated with the time period in which the counted peak height occurred, a number stored by the counter representing the number of steps taken by the individual. 2. The system of claim 1, wherein the accelerometer measures linear acceleration along three axes and the accelerometer outputs from the three axes are combined. 3. A method of counting steps from accelerometer outputs of an accelerometer, the accelerometer coupled to an individual and in communication with a processor, the processor in communication with a non-transitory computer readable medium, the non-transitory computer readable medium storing instructions controlling the processor to execute the steps of the method, the method comprising: determining readings using the accelerometer outputs of the accelerometer;identifying counted amplitudes, each of the counted amplitudes comprising a difference between a maximum reading and a minimum reading during a time period in which the maximum reading exceeds a rest threshold, the rest threshold determined by: receiving readings from the accelerometer during a period of rest,identifying a rest maximum, the rest maximum comprising a greatest reading generated by the accelerometer during the period of rest, anddetermining the rest threshold using the rest maximum;determining an amplitude threshold for each time period having a counted amplitude, the amplitude threshold comprising a first weighted average of two or more counted amplitudes;... | SAVVYSHERPA, LLC | G01C022/00 G01P013/00 A61B005/11 A61B005/00
US10702239 | B1 | 2020-07-07 | Predicting characteristics of a future respiratory event, and applications thereof | 2019-10-21 | Embodiments disclosed herein improve digital stethoscopes and their application and operation. A first method detects of a respiratory abnormality using a convolution. A second method counts coughs for a patient. A third method predicts a respiratory event based on a detected trend. A fourth method forecasts characteristics of a future respiratory event. In a fifth embodiment, a base station is provided for a digital stethoscope. | 1. A computer-implemented method for forecasting characteristics of a future respiratory event comprising: receiving, from a microphone of a digital stethoscope, a first noise signal capturing a first respiratory event at a first time;generating, by one or more computing devices, a first feature space representing the first noise signal;encoding, by the one or more computing devices, the first feature space into a first convolution vector using a spatial convolution procedure;receiving, from the microphone of the digital stethoscope, a second noise signal capturing a second respiratory event at a second time;generating, by the one or more computing devices, a second feature space representing the second noise signal;encoding, by the one or more computing devices, the second feature space into a second convolution vector using the spatial convolution procedure;determining, by the one or more computing devices, a predicted convolution vector based on the first and second feature spaces;decoding, by the one or more computing devices, the predicted convolution vector into a predicted feature space representing a sound made by the future respiratory event; andtransmitting, by the one or more computing devices, the predicted feature space to a display unit for output, wherein the output indicates the likelihood of the future respiratory event. 2. The method of claim 1 wherein the first, second, and predicted feature spaces are three-dimensional spaces with values each signifying a frequency, rate, and severity of the first noise signal, the second noise signal, and the sound made by the future respiratory event. 3. The method of claim 1 wherein the first noise signal, second noise signal, and the sound made by the future respiratory event are coughs. 4. The method of claim 1 wherein the first noise signal, second noise signal, and the sound made by the future respiratory event are wheezes. 5. The method of claim 1 wherein encoding the first and second feature spaces comprises repeatedly performing the spatial convolution procedure on the first and second feature spaces to generate one or more convolution values, wherein the convolution values represent a compressed version of a portion of the first and second feature spaces. 6. The method of claim 5 wherein decoding the predicted convolution vector comprises repeatedly performing an inverse of the spatial convolution procedure to generate the predicted feature space. 7. The method of claim 1 wherein determining the predicted convolution vector based on the first and second feature spaces comprises: applying a first weight to the first convolutional vector to generate a first intermediate vector; andcalculating a second intermediate vector based on applying a second weight to the second convolutional vector, wherein the first and second weights are values trained based on a history of respiratory events. 8. A non-transitory computer readable medium including instructions for forecasting character... | Sonavi Labs, Inc. | A61B005/08 A61B007/00 G16H050/20 A61B007/04 A61B005/00
US10709414 | B1 | 2020-07-14 | Predicting a respiratory event based on trend information, and applications thereof | 2019-10-21 | Embodiments disclosed herein improve digital stethoscopes and their application and operation. A first method detects of a respiratory abnormality using a convolution. A second method counts coughs for a patient. A third method predicts a respiratory event based on a detected trend. A fourth method forecasts characteristics of a future respiratory event. In a fifth embodiment, a base station is provided for a digital stethoscope. | 1. A computer-implemented method for predicting a respiratory event comprising: receiving, from a microphone of a digital stethoscope, a first noise signal at a first time;analyzing, by one or more computing devices, the first noise signal to determine a first severity score indicating how severe a first respiratory event of a patient captured in the first noise signal is;receiving, from the microphone of the digital stethoscope, a second noise signal captured at a second time after the first time;analyzing, by the one or more computing devices, the second noise signal to determine a second severity score indicating how severe a second respiratory event of the patient captured in the second noise signal is;generating, by the one or more computing devices, a prediction indicating a likelihood of a third respiratory event occurring for the patient at a future time based on applying a first weight to the first severity score and a second weight to the second severity score, wherein the first and second weights are trained based on a history of respiratory events for the patient, wherein the first and second weights are set such that the second severity score impacts the generation of the prediction more heavily than the first severity score; andtransmitting, by the one or more computing devices, the prediction to a display device for outputting the prediction. 2. The method of claim 1 further comprising: generating, by the one or more computing devices, a first feature map, including the first severity score, describing one or more conditions at the first time;generating, by the one or more computing devices, a second feature map, including the second severity score, describing the conditions at the second time; andgenerating the prediction based on the first feature map and the second feature map. 3. The method of claim 2 wherein the first feature map includes a first environmental data describing a first environmental factor at the first time, and wherein the second feature map includes a second environmental data describing a second environmental factor at the second time. 4. The method of claim 3 wherein the first and second environment data include a weather data, a pollen data, or a pollution level data. 5. The method of claim 2 wherein the first feature map includes a first count of coughs detected from the microphone during the first time, and the second feature map includes a second count of coughs detected from the microphone during the second time. 6. The method of claim 1 wherein generating the prediction is based on a long short term memory (LSTM) procedure trained based on the history of respiratory events for the patient. 7. The method of claim 1 wherein the first severity score and the second severity score are each one of (1) not present, (2) intermittent, (3) mild, (4) moderate, (5) severe, (6) acute, and wherein the first, second, and third respiratory events are wheezes. 8. The method of claim 1 wherein analyzing the first... | Sonavi Labs, Inc. | A61B005/08 A61B007/00 G16H050/20 A61B007/04 A61B005/00
US10709355 | B2 | 2020-07-14 | Systems, devices, and methods for performing active auscultation and detecting sonic energy measurements | 2019-08-15 | Active auscultation may be used to determine organ (e.g., lung or heart) characteristics of users. An acoustic or piezo-electric signal (e.g., a pulse, a tone, and/or a broadband pulse) may be projected into an animal (typically human) body or thorax. The signal interacts with the body, or lungs, and in some cases may induce resonance within the body/lungs. A resultant signal may be emitted from the body which may be analyzed to determine, for example, a lung's resonant frequency or frequencies and/or how the sound is otherwise absorbed, reflected, or modified by the body. This information may be indicative of lung characteristics such as lung capacity, a volume of air trapped in the lungs, and/or the presence of COPD. | 1. A method of performing active auscultation comprising: directing, by a processor, a broadband acoustic signal into user's body toward a user's lung, wherein the broadband acoustic signal is a pulse that comprises a plurality of frequencies;receiving, by the processor, an acoustic signal from the user's body, the received acoustic signal corresponding to the broadband acoustic signal;determining, by the processor, a characteristic of the received acoustic signal, wherein the characteristic is at least one of an intensity and a frequency of the received acoustic signal; anddetermining, by the processor, a volume of trapped air present in the user's lung using the determined characteristic, the trapped air being air that is trapped within discrete pockets of lung tissue of the user's lung following the user's exhalation of air from the lung. 2. The method of claim 1, wherein the broadband acoustic signal is a first broadband acoustic signal and the received acoustic signal is a first received acoustic signal, the method further comprising: receiving, by the processor, a second acoustic signal from the user's body, the received acoustic signal corresponding to a second broadband acoustic signal projected into the user's body;determining, by the processor, a characteristic of the second received acoustic signal, wherein the characteristic of the second received acoustic signal is at least one of an intensity and a frequency of the second received acoustic signal;comparing, by the processor, a characteristic of the first and second received acoustic signals; andproviding, by the processor, an indication of the comparing of the characteristic of the first and second received acoustic signals to an operator. 3. The method of claim 1, further comprising: comparing, by the processor, the characteristic of the received acoustic signal to a predetermined value. 4. The method of claim 1, furthering comprising: adjusting, by the processor, at least one of a duration of the broadband signal, an intensity of the broadband acoustic signal, and frequencies included in the broadband acoustic signal responsively to the determined characteristic of the broadband acoustic signal. 5. The method of claim 1, wherein the broadband acoustic signal comprises a plurality of frequencies between 2,000 Hz and 30,000 Hz. 6. The method of claim 1, wherein a duration of the pulse within a range of 0.1 seconds and 2 seconds. 7. A method comprising: receiving, by a processor, a detected acoustic signal from a receiver, the detected acoustic signal corresponding to an acoustic signal directed toward a lung within a body of a user in a pulse and responsively emanating from the user, the acoustic signal comprising a plurality of frequencies;analyzing, by the processor, the received detected acoustic signal to determine a resonant frequency for the lung; anddetermining, by the processor, a volume of trapped air present in the lung following the user's exhalation of air from the... | Respira Labs, Inc. | A61B005/091 A61B007/00 G10L025/51
US10722128 | B2 | 2020-07-28 | Heart rate detection system and method | 2019-07-31 | Computerized eyewear and corresponding methods measure a wearer's heart rate using a first optical sensor and a second optical sensor at least partially embedded in an eyewear temple. The first optical sensor transmits a first signal to a temple of the wearer and the second optical sensor transmits a second signal to the temple of the wearer. Reflections of the first signal are used to measure a raw heart rate delta and reflections of the second signal are used to measure a noise delta. The raw heart rate delta and the noise delta are used to determine a measured heart rate of the wearer of the computerized eyewear. | 1. Computerized eyewear, comprising: a frame, wherein the frame comprises a frame first end, anda frame second end;a first temple, wherein the first temple comprises a first temple first end pivotally coupled to the frame first end, anda first temple second end configured to rest on a first ear of a wearer of the computerized eyewear;a second temple, wherein the second temple comprises a second temple first end pivotally coupled to the frame second end,a second temple second end configured to rest a second ear of the wearer of the computerized eyewear,a first optical transmitter at least partially embedded in the second temple and operative to transmit a first signal to a temple of the wearer of the computerized eyewear to measure a raw heart rate delta,a second optical transmitter at least partially embedded in the second temple proximate to the first optical transmitter and operative to transmit a second signal to the temple of the wearer of the computerized eyewear to measure a noise delta, anda detector at least partially embedded in the second temple proximate to the first optical transmitter and to the second optical transmitter, the detector operative to receive a reflected first signal originating from the first optical transmitter or a reflected second signal originating from the second optical transmitter and reflected back to the detector from a surface of the temple of the wearer;at least one processor coupled to one of the first temple, the second temple, or the frame and communicatively coupled to the first optical transmitter, the second optical transmitter, and the detector, the at least one processor operative to use the reflected first signal over time representative of a movement of a heart rate detection area of the surface of the temple of the wearer to generate the raw heart rate delta, to use the reflected second signal over time representative of a movement of a wide field of detection on the surface of the temple of the wearer that is larger than the heart rate detection area to generate the noise delta, and to use the raw heart rate delta and the noise delta to determine a measured heart rate of the wearer of the computerized eyewear. 2. The computerized eyewear of claim 1, wherein the first optical transmitter comprises a narrow field emitter operative to focus the first signal on the heart rate detection area of the temple of the wearer, the heart rate detection area comprising a superficial temporal artery of the wearer. 3. The computerized eyewear of claim 2, wherein the second optical transmitter comprises a wide field emitter operative to focus the second signal on the wide field of detection that is larger than the heart rate detection area of the first signal. 4. The computerized eyewear of claim 3, wherein the first optical transmitter comprises a narrow field emitting Light Emitting Diode (LED) and the second optical transmitter comprises a wide field emitting LED. 5. The computerized eyewear of claim 4, wh... | Vision Service Plan | A61B005/024 A61B005/00 A61B005/22 G02C011/00 A61B005/0205 G02C005/14 G02C005/00
US10720029 | B1 | 2020-07-21 | Medical device alert, optimization, personalization, and escalation | 2019-02-05 | An alert is used to inform a user that their blood glucose level has dropped below a threshold (e.g., the user is hypoglycemic) or has increased above a threshold (e.g., the user is hyperglycemic). There is a hierarchy of alerts from lowest priority to highest priority. The alert is communicated by a user device (e.g., a mobile device), which is, for example, a smartphone, smart watch, home automation device, or the like. The alert is modified in order to increase the likelihood that the user receives or acknowledges the alert within a minimal amount of time. An intensity level (e.g., a volume) of an alert is modified based on, for example, whether the user has acknowledged a previous alert. A modality or a sensory channel of an alert is changed if an initial alert does fails to elicit a response from the user. | 1. A method comprising: detecting a triggering event for communicating a diabetes-related alert to a user;detecting a keyword from at least one of a calendar event title or a calendar event location;determining a first predefined alert profile based on the keyword from the at least one of the calendar event title or the calendar event location;determining a second predefined alert profile based on a keyword from a location of a mobile device associated with the user;generating the diabetes-related alert using a first alert type associated with the first predefined alert profile in response to the first alert type matching a second alert type associated with the second predefined alert profile; andgenerating the diabetes-related alert using one of the first alert type or the second alert type that has a higher priority in response to the first alert type failing to match the second alert type. 2. The method of claim 1, wherein the first alert type and the second alert type comprise at least one different vibration setting or volume setting. 3. The method of claim 1, further comprising: using a higher-priority alert profile in response to a failure of the mobile device to receive a user acknowledgement of the diabetes-related alert after a predefined period of time. 4. The method of claim 3, further comprising: receiving an indication of a user selection of a modality for at least one of the first predefined alert profile or the second predefined alert profile; andsetting, based on the modality, a number of repetitions that the diabetes-related alert is communicated using at least one of the first alert type or the second alert type before progressing to a higher-priority modality. 5. The method of claim 1, wherein the first predefined alert profile is determined in response to the calendar event location comprising a street address that matches the street address of a location type. 6. The method of claim 1, the method further comprising: comparing the keyword in the calendar event title to keywords associated with a plurality of predefined alert profiles, wherein the first predefined alert profile is determined in response to the keyword in the calendar event title matching a keyword associated with the first predefined alert profile. 7. The method of claim 1, further comprising: comparing the keyword in the calendar event location to keywords associated with a plurality of predefined alert profiles, and wherein the first predefined alert profile is determined in response to the keyword in the calendar event location matching a keyword associated with the first predefined alert profile. 8. The method of claim 1, wherein the keyword in the calendar event location comprises a street address, the method further comprising: comparing the street address in the calendar event location to street addresses of a plurality of location types having predefined alert profiles, and wherein the first predefined alert profile is determined in response to ... | ROCHE DIABETES CARE, INC. | A61B005/00 G08B003/10 H04L029/08 A61B005/1171 A61B005/12 G16H010/60 G16H050/30 G16H040/63 H03G003/32 A61B005/145
US10722120 | B1 | 2020-07-28 | Sensor platform array that moves sensor platforms to conform to the shape of the subject to be monitored | 2019-02-11 | The sensor platform array includes a plurality of sensors and one or more sensor platforms. Each sensor may be mounted on a sensor platform. The sensor platform may be mounted on a support panel. The support panel may be mounted on a toilet lid. The one or more sensor platforms may be extendable towards a user seated on the toilet. Each sensor may be independently extended so that it is positioned flush against a user's body. The sensor platform array may include bendable arms which reach around a user and place sensors on the ventral side of the user. The sensors may collect measurements which are relevant to the user's health and well-being. In some embodiments, the sensors may be removable and replaceable so that different sensors may be added according to a user's individual needs. | 1. A sensor platform array comprising: a support panel;a plurality of sensor-inflatable cell assemblies, each assembly comprising: an inflatable cell; anda sensor, wherein the sensor is mounted on the inflatable cell, wherein each of the plurality of sensor-inflatable cell assemblies is connected to the support panel, and wherein the sensor collects physiological measurements;a gas pump, wherein in the gas pump is in fluid connection with the air inflatable cell in each of the plurality of sensor-inflatable cell assemblies and is configured to pump gas independently to each of the plurality of inflatable cells; anda toilet lid, wherein the support panel is mounted on the toilet lid. 2. The sensor platform array of claim 1, further comprising a conformation control system, the conformation control system comprising: a plurality of pressure sensors, wherein each of the plurality of pressure sensors is in communication with the inflatable cell within one of the plurality of sensor-inflatable cell assemblies;a controller, wherein the controller is in electronic communication with each of the plurality of pressure sensors; anda non-transitory computer readable medium comprising instructions for actuating the pump and directing the pump to provide a supply of gas to the inflatable cell in each of the plurality of assemblies in the absence of a signal from the pressure sensor. 3. The sensor platform array of claim 2, wherein the non-transitory computer readable medium further comprises instructions for independently halting the supply of gas to each of the at least one inflatable cell in response to a signal from the pressure sensor on the sensor platform. 4. The sensor platform array of claim 1, wherein the sensor is independently selected from the following list: a stethoscope, an ultrasound probe, an echocardiogram probe, a temperature sensor, a durometer, an electrocardiogram lead, and a ballistocardiography sensor. 5. The sensor platform array of claim 1, wherein the sensors are removable and exchangeable with other sensors. 6. The sensor platform array of claim 1, further comprising at least one bendable arm, wherein each of the at least one bendable arm comprises: at least one sensor; andat least one hinge;wherein the at least one bendable arm extends from the support panel. 7. The sensor platform array of claim 1, wherein the sensor within each of plurality of sensor-inflatable cell assemblies is in electrical communication with the controller. 8. The sensor platform array of claim 7 wherein the controller comprises a memory, and wherein the controller stores a signal produced by the sensor within each of the plurality of sensor-inflatable cell assemblies on the memory. 9. The sensor platform array of claim 8, further comprising a non-transitory computer readable medium comprising instructions for creating an analysis of the signal produced by the sensor within each of the plurality of sensor-inflatable cell assemblies. 10. A sensor pla... | Hall Labs LLC | A61B005/00 A47K013/24 G01L001/26