## Sheet1
公告(開)號 | KindCode | 公告(開)日 | 名稱 | 申請日 | 摘要 | 申請範圍1 | 申請範圍 | 專利權人 | IPC
US06925326 | B1 | 2005-08-02 | Implantable cardiac stimulation system and method for measuring atrioventricular conduction and adjusting atrioventricular hystersis | 2002-01-04 | An implantable cardiac stimulation device and method measure atrioventricular conduction times and automatically adjust an atrioventricular delay time based on the measured conduction time values. | 1. A method for setting an atrioventricular delay in a cardiac stimulation device, the method comprising: monitoring for atrial events;adjusting an atrioventricular delay until an intrinsic ventricular event is detected;maintaining the adjusted atrioventricular delay for a predetermined number of cardiac cycles; andstoring the adjusted atrioventricular delay if at least a predetermined percentage of ventricular events during the predetermined number of cardiac cycles comprise intrinsic ventricular events. | 1. A method for setting an atrioventricular delay in a cardiac stimulation device, the method comprising: monitoring for atrial events;adjusting an atrioventricular delay until an intrinsic ventricular event is detected;maintaining the adjusted atrioventricular delay for a predetermined number of cardiac cycles; andstoring the adjusted atrioventricular delay if at least a predetermined percentage of ventricular events during the predetermined number of cardiac cycles comprise intrinsic ventricular events. 2. The method of claim 1, wherein monitoring comprises monitoring for intrinsic atrial events. 3. The method according to claim 2, further comprising decreasing a base stimulation rate to inhibit delivery of atrial stimulation pulses; measuring an average atrioventricular conduction time following sensing of intrinsic atrial events; andcalculating an atrial-ventricular hysteresis based on the measured average atrioventricular conduction time. 4. The method of claim 1, wherein monitoring comprises monitoring for stimulated atrial events. 5. The method according to claim 4, further comprising increasing a base stimulation rate to induce delivery of atrial stimulation pulses; measuring an average atrioventricular conduction time following delivery of atrial stimulation pulses; andcalculating an atrial-ventricular hysteresis based on the measured average atrioventricular conduction time. 6. The method of claim 1, wherein adjusting an atrioventricular delay comprises adjusting a hysteresis value. 7. The method according to claim 6, wherein adjusting the atrial-ventricular hysteresis comprises: calculating an average atrioventricular conduction time from a plurality of atrioventricular conduction time measurements;calculating a measure of variability of the atrioventricular conduction time measurements; andcalculating the hysteresis value based on the average atrioventricular conduction time and the measure of variability of the atrioventricular conduction time measurements. 8. The method according to claim 1, further comprising determining an atrioventricular delay on a periodic basis. 9. A cardiac stimulation device for automatically measuring an atrioventricular conduction time, comprising: means for monitoring for atrial events;means for monitoring for intrinsic ventricular events;means for determining atrioventricular conduction times for a plurality of cardiac cycles;means for determining, based on the atrioventricular conduction times, a conduction time value for which at least a predetermined percentage of ventricular events that occur during the plurality of cardiac cycles comprise intrinsic ventricular events; andmeans for setting an atrioventricular delay to a value based on the conduction time value. 10. The stimulation device of claim 9, wherein the means for monitoring monitors for intrinsic atrial events. 11. The stimulation device of claim 9, wherein the means for monitoring monitors for stimulated atrial events. 12. The sti... | Pacesetter, Inc. | A61B005/04
US06935889 | B2 | 2005-08-30 | Electrode-pad package that is removable from an electrode-pad lead and method for opening the package | 2001-02-28 | A package stores at least one electrode pad having a lead. The package includes an interior for storing the pad, and includes a portion that is openable to allow removal of the package from the leads. For example, the openable portion of the package may include a sealed seam through which the lead extends. The operator can peel apart the seam to both open the package and free the package from the lead. Where the pad is a defibrillator electrode pad, the ability to quickly and easily free the package from the lead may save precious seconds during the rescue of a patient in cardiac arrest. For example, freeing the package from the lead typically allows the operator a fuller view of and easier access to the defibrillator and the patient, and eliminates a potential distraction. | 1. An electrode-pad storage system, comprising: an electrode-pad package comprising, an interior, and a seam that can be opened; and an electrode pad disposed in the interior, the electrode pad having a lead that extends through the seam of the package. | 1. An electrode-pad storage system, comprising: an electrode-pad package comprising, an interior, and a seam that can be opened; and an electrode pad disposed in the interior, the electrode pad having a lead that extends through the seam of the package. 2. The electrode-pad storage system of claim 1 wherein the seam can be peeled apart to allow removal of the electrode pad from the interior of the package and the package from the lead. 3. The electrode-pad storage system of claim 1 wherein the electrode pad comprises a defibrillator electrode pad. 4. The electrode-pad storage system of claim 1 wherein the electrode pad comprises a monitor electrode pad. 5. The electrode-pad storage system of claim 1, wherein the seam is sealed to inhibit moisture loss from the interior. 6. The electrode-pad storage system of claims 1, wherein the seam is tapered toward the location where the lead extends through the seam of the package. 7. An electrode-pad storage system, comprising: an electrode-pad package comprising, a sealed seam which is designed to be opened to open the package, an opening disposed along the seam, first and second side portions attached to and separable from one another along the seam as the package is opened, and a package interior defined by the first and second side portions; and an electrode pad disposed in the package interior and having a lead that extends through the opening, which is sealed around the lead. 8. The electrode-pad storage system of claim 7 wherein the first and second sides respectively include first and second tabs that facilitate peeling apart the first and second sides along the seam. 9. The electrode-pad storage system of claim 7 wherein the first and second sides respectively include first and second tabs that are aligned with the opening and that facilitate peeling apart the first and second sides along the seam. 10. The electrode-pad storage system of claim 7, wherein the seam is sealed to inhibit moisture loss from the package interior. 11. An electrode-pad storage system, comprising: an electrode-pad package comprising, an opening, first and second seams that join the opening, first and second sides attached to and separable from one another along the first and second seams, and an interior disposed between the first and second sides; and an electrode pad disposed in the interior and having a lead that extends through the opening, which is sealed around the lead. 12. The electrode-pad storage system of claim 11 wherein the first and second sides respectively include first and second tabs that are aligned with the opening and that facilitate peeling apart the first and second sides along the first and second seams. 13. The electrode-pad storage system of claim 11 wherein the first and second seams taper toward the opening. 14. The electrode-pad storage system of claim 11 wherein the electrode-pad package further comprises: an edge that is contiguous with the first and second seams; wherein th... | Koninklijke Philips Electronics N.V. | H01R013/52 B65D081/124 A61B005/04
US06944496 | B2 | 2005-09-13 | Apparatus for positioning and marking a location of an EMG electrode | 2002-03-13 | An apparatus for positioning and marking a location of an electromyography (EMG) electrode includes a housing having a back surface, an EMG sensor for receiving the EMG signal, an amplifying circuit for amplifying the EMG signal to generate an amplified EMG signal and an audio output terminal for outputting the amplified EMG signal as the audio EMG signal. The examinee may directly position and mark the location of the EMG electrode and may easily detect the fatigue rate of the muscle or the abnormal condition of the muscle by using the sound of the audio equipment. | 1. An apparatus for positioning and marking a location of an electromyography (EMG) electrode to convert an EMG signal to an audio EMG signal, comprising: a housing having a back surface; an EMG sensor, installed on the back surface of the housing, for receiving the EMG signal, wherein the EMG sensor has a sensing surface capable of being contacted with a human body; an amplifying circuit for amplifying the EMG signal to generate an amplified EMG signal; an audio output terminal, installed on a side of the housing, for outputting the amplified EMG signal as the audio EMG signal, wherein the audio output terminal has a structure capable of being connected with an audio input terminal of a audio device; a marking lever installed to be capable of moving along a lever guide conduit installed through the housing; and a marking tap, integratedly installed on an end of the marking lever, capable of being projected outside of the housing due to the movement of the marking lever. | 1. An apparatus for positioning and marking a location of an electromyography (EMG) electrode to convert an EMG signal to an audio EMG signal, comprising: a housing having a back surface; an EMG sensor, installed on the back surface of the housing, for receiving the EMG signal, wherein the EMG sensor has a sensing surface capable of being contacted with a human body; an amplifying circuit for amplifying the EMG signal to generate an amplified EMG signal; an audio output terminal, installed on a side of the housing, for outputting the amplified EMG signal as the audio EMG signal, wherein the audio output terminal has a structure capable of being connected with an audio input terminal of a audio device; a marking lever installed to be capable of moving along a lever guide conduit installed through the housing; and a marking tap, integratedly installed on an end of the marking lever, capable of being projected outside of the housing due to the movement of the marking lever. 2. The apparatus of claim 1, further comprising a gain control unit, installed on a side of the housing, for controlling a gain of the amplifying circuit. 3. The apparatus of claim 2, wherein the EMG sensor has one or more EMG sub-sensors for receiving one or more EMG sub-signals, respectively, and the amplifying circuit includes: one or more amplifying units for amplifying said one or more EMG sub-signals to generate one or more amplified EMG sub-signals, respectively; a differential amplifying unit for receiving said one or more amplified EMG sub-signals to output a differential amplified EMG signal which is in proportion to a difference between said one or more amplified EMG sub-signals; a variable amplifying unit for amplifying the differential amplified EMG signal based on a variable gain which corresponds to a variable resister of the gain control unit to generate a variable amplified EMG signal; and a signal output unit for outputting the variable amplified EMG signal to the audio output terminal. 4. The apparatus of claim 1, further comprising an elastic material for allowing the marking lever to elastically move and wherein the housing has an elastic material housing space for receiving the elastic material. | Electronics and Telecommunications Research Institute | A61B005/0488
US06957097 | B2 | 2005-10-18 | Rapid measurement of time-averaged blood flow using ungated spiral phase-contrast MRI | 2002-04-17 | Pulsatile flow is measured using magnetic resonance imaging without cardiac gating using a phase-contrast excitation method to rapidly quantify blood flow and using a spiral k-space trajectory for image data read-out to mitigate deleterious effects of pulsatility. Post-processing of the read-out data provides a cumulative-average velocity plot from which a period of a cardiac cycle is obtained. Time-averaged blood flow rates can be rapidly and robustly measured and is more repeatable than conventional gated techniques. | 1. A method of imaging pulsatile blood flow in a patient using magnetic resonance imaging without cardiac gating comprising the steps of: a) placing the patient in a static magnetic field, b) exciting nuclear spins in the presence of a gradient (Gz) for a slice defined by the gradient, c) applying a phase contrast flow encoding magnetic pulse following step b), d) applying magnetic read-out gradients (Gx, Gy) after step c) whereby spiral k-space type sampling is used during read-out, and e) acquiring read-out data, f) repeating steps b)-e) with rotated magnetic readout gradients and without cardiac gating over at least one cardiac cycle, and g) processing the acquired data to obtain time averaged flow rate information. | 1. A method of imaging pulsatile blood flow in a patient using magnetic resonance imaging without cardiac gating comprising the steps of: a) placing the patient in a static magnetic field, b) exciting nuclear spins in the presence of a gradient (Gz) for a slice defined by the gradient, c) applying a phase contrast flow encoding magnetic pulse following step b), d) applying magnetic read-out gradients (Gx, Gy) after step c) whereby spiral k-space type sampling is used during read-out, and e) acquiring read-out data, f) repeating steps b)-e) with rotated magnetic readout gradients and without cardiac gating over at least one cardiac cycle, and g) processing the acquired data to obtain time averaged flow rate information. 2. The method as defined by claim 1 wherein the gradient magnetic field (Gz) in step b) includes a gradient spoiler. 3. The method as defined by claim 2 wherein the readout gradients in step d) rewind the k-space trajectory to k-space origin. 4. The method as defined by claim 3 wherein step g) includes use of a cumulative-average velocity plot from which a period of a cardiac cycle is obtained. 5. The method as defined by claim 4 wherein step c) applies a flow encoding magnetic pulse having a non-zero first moment. 6. The method as defined by claim 5 wherein in step d) k-space is sampled along spiral k-space trajectories. 7. The method as defined by claim 6 wherein the k-space trajectories are interleaved. 8. The method as defined by claim 7 wherein the k-space trajectories have variable density. 9. The method as defined by claim 5 wherein in step d) k-space sampling is along radial lines. 10. The method as defined by claim 5 wherein in step d) k-space sampling is along twisted radial lines. 11. The method as defined by claim 1 wherein step g) includes use of a cumulative-average velocity plot from which a period of cardiac cycle is obtained. 12. The method as defined by claim 1 wherein step c) applies a flow encoding magnetic pulse having a non-zero first moment. 13. The method as defined by claim 1 wherein in step d) k-space sampling is along spiral k-space trajectories. 14. The method as defined by claim 13 wherein the k-space trajectories are interleaved. 15. The method as defined by claim 14 wherein the k-space trajectories have variable density. 16. The method as defined by claim 1 wherein in step d) k-space sampling is along radial lines. 17. The method as defined by claim 1 wherein in step d) k-space sampling is along twisted radial lines. 18. The method as defined by claim 1 wherein step d) applies three readout gradients (Gx, Gy, Gz) and step g) obtains three dimensional time averaged flow information. | The Board of Trustees of the Leland Stanford Junior University | A61B005/05
US07123954 | B2 | 2006-10-17 | Method for classifying and localizing heart arrhythmias | 2002-12-18 | Analyzes surface electrocardiographic and intracardiac signals to identify and separate electrical activity corresponding to distinct but superimposed events in the heart. Assesses the spatial phase, temporal phase, rate, spectrum and reproducibility of each event to determine uniformity of activation in all spatial dimensions. Uses numerical indices derived from these analyses to diagnose arrhythmias. Uses these indices to determine the location of an arrhythmia circuit, and to direct the movement of an electrode catheter to this location for ablation or permanent catheter positioning. Subsequently, uses these indices to determine whether ablation has successfully eliminated the circuit. Uses variability in these indices from the surface electrocardiogram to indicate subtle beat-to-beat fluctuations which reflect the tendency towards atrial and ventricular arrhythmias. | 1. A method of analyzing a cardiac signal, comprising: providing a digital template and a digitized cardiac signal with sufficient fidelity and resolution to facilitate identification of atrial fibrillatory components of activity from distinct regions of the heart and further having sufficient fidelity and resolution to facilitate identification of the atrial fibrillatory components as originating from a left side or a right side of the heart, the identification being sufficient to facilitate a separation of atrial tachycardia, atrial fibrillation, atrial flutter and ventricular tachycardia, the digitized cardiac signal comprising a plurality of amplitudes corresponding to consecutive time samples, wherein the time samples span a plurality of cardiac cycles;generating a plurality of correlation values correlating the digital template to successive time samples of the cardiac signal, wherein a plurality of correlation values are generated for each cardiac cycle that contain information necessary to identify and differentiate distinct components of activity and to discern whether those components have origins in the left or right sides of the heart;mapping the correlation values; andidentifying a treatable disease, based upon the mapping of the correlation values. | 1. A method of analyzing a cardiac signal, comprising: providing a digital template and a digitized cardiac signal with sufficient fidelity and resolution to facilitate identification of atrial fibrillatory components of activity from distinct regions of the heart and further having sufficient fidelity and resolution to facilitate identification of the atrial fibrillatory components as originating from a left side or a right side of the heart, the identification being sufficient to facilitate a separation of atrial tachycardia, atrial fibrillation, atrial flutter and ventricular tachycardia, the digitized cardiac signal comprising a plurality of amplitudes corresponding to consecutive time samples, wherein the time samples span a plurality of cardiac cycles;generating a plurality of correlation values correlating the digital template to successive time samples of the cardiac signal, wherein a plurality of correlation values are generated for each cardiac cycle that contain information necessary to identify and differentiate distinct components of activity and to discern whether those components have origins in the left or right sides of the heart;mapping the correlation values; andidentifying a treatable disease, based upon the mapping of the correlation values. 2. The method as set forth in claim 1, wherein the correlation values are mapped against time. 3. The metbod as set forth in claim 1, wherein the correlation values are plotted. 4. The method as set forth in claim 1, wherein successive time samples are consecutive time samples. 5. The method as set forth in claim 1, wherein the cardiac signal is an ECG. 6. The method as set forth in claim 1, wherein the cardiac signal is an intracardiac electrogram. 7. The method as set forth in claim 1, and thither comprising computing at least one Eigen vector. 8. The method as set forth in claim 1, and further comprising: comparing correlation values with a threshold value; andidentifying features as correlation values that exceed the threshold value. 9. The method as set forth in claim 8, and further comprising defining features from the correlation values. 10. The method as set forth in claim 9, wherein the features comprise activity from distinct regions of cardiac chambers. 11. The method as set forth in claim 9, wherein the features comprise activity from an atrium. 12. The method as set forth in claim 9, wherein the features comprise activity from a ventricle. 13. The method as set forth in claim 9, wherein the features comprise repolarization. 14. The method as set forth in claim 1, and further comprising generating frequency representation of the digitized cardiac signal and of the correlation values. 15. The method as set forth in claim 14, wherein the frequency representation of the digitized cardiac signal comprises a plot of frequency versus power. 16. The method as set forth in claim 14, and further comprising identifying an non-harmonic frequencies having relatively low amp... | - | A61B005/046
US07254435 | B2 | 2007-08-07 | Method and magnetic resonance apparatus for calibrating coil sensitivities | 2003-01-31 | In a method and magnetic resonance imaging apparatus wherein magnetic resonance signals are simultaneously received from an examination subject by multiple reception coils, a single, uninterrupted pulse sequence is executed which includes reference scans of the subject with a first sequence kernel that is optimized for coil sensitivity calibration, immediately followed by a series of accelerated image scans with a second sequence kernel, different from the first sequence kernel, that is optimized for imaging. Coil sensitivity maps for the respective coils are calculated from the data acquired in the reference scans, and an image of the subject is reconstructed by operating on the image data with a parallel reconstruction algorithm employing the calculated coil sensitivity maps. | 1. A method for operating a magnetic resonance imaging apparatus having an antenna array comprised of a plurality of reception coils for simultaneously receiving magnetic resonance signals from a subject, each of said coils having a coil sensitivity, comprising the steps of: acquiring a plurality of sets of reference data of a subject by executing a series of reference scans with a first sequence kernel optimized for coil sensitivity calibration;obtaining a plurality of sets of image data of said subject by executing a series of accelerated image scans with a second sequence kernel, different from said first sequence kernel, optimized for imaging, by obtaining magnetic resonance signals from said subject simultaneously with multiple reception coils in said plurality of reception coils;for each coil, calculating a coil sensitivity map from said sets of reference data; andreconstructing an image of said subject by operating on said sets of image data with a parallel reconstruction algorithm employing said coil sensitivity maps. | 1. A method for operating a magnetic resonance imaging apparatus having an antenna array comprised of a plurality of reception coils for simultaneously receiving magnetic resonance signals from a subject, each of said coils having a coil sensitivity, comprising the steps of: acquiring a plurality of sets of reference data of a subject by executing a series of reference scans with a first sequence kernel optimized for coil sensitivity calibration;obtaining a plurality of sets of image data of said subject by executing a series of accelerated image scans with a second sequence kernel, different from said first sequence kernel, optimized for imaging, by obtaining magnetic resonance signals from said subject simultaneously with multiple reception coils in said plurality of reception coils;for each coil, calculating a coil sensitivity map from said sets of reference data; andreconstructing an image of said subject by operating on said sets of image data with a parallel reconstruction algorithm employing said coil sensitivity maps. 2. A method as claimed in claim 1 comprising employing a FLASH pulse sequence as said first sequence kernel. 3. A method as claimed in claim 1 comprising employing a TrueFISP pulse sequence as said second sequence kernel. 4. A method as claimed in claim 1 comprising employing a FLASH pulse sequence as said first sequence kernel and a TrueFISP pulse sequence as said second sequence kernel. 5. A method as claimed in claim 1 comprising executing said series of reference scans including generating respective RF pulses each having a first flip angle, and executing said series of accelerated image scans including generating respective RF pulses each having a second flip angle with a magnitude of said second flip angle being different from a magnitude of said first flip angle. 6. A method as claimed in claim 1 comprising executing said series of accelerated image scans with a different contrast than said series of reference scans. 7. A method as claimed in claim 1 comprising executing said series of reference scans with a larger field of view than said accelerated image scans. 8. A method as claimed in claim 1 comprising executing said reference scans to obtain said reference data from a slice of said subject and executing said accelerated image scans to obtain said image data from said slice of said subject, and employing identical information designating a position and an orientation of said slice in each of said series of reference scans and said series of accelerated image scans. 9. A magnetic resonance imaging apparatus comprising: a magnetic resonance scanner having an antenna array including a plurality of reception coils, each of said reception coils having a coil sensitivity;a sequence controller connected to said magnetic resonance scanner for operating said magnetic resonance scanner to execute a continuous, uninterrupted pulse sequence comprising a plurality of reference scans of said subject with a first sequen... | Siemens Aktiengesellschaft | A61B005/055
US07020507 | B2 | 2006-03-28 | Separating motion from cardiac signals using second order derivative of the photo-plethysmogram and fast fourier transforms | 2003-01-30 | The present invention is directed toward a pulse oximetry system for the determination of a physiological parameter capable of removing motion artifacts from physiological signals comprises a hardware subsystem and a software subsystem. The software subsystem is used in conjunction with the hardware subsystem to perform a method for removing a plurality of motion artifacts from the photo-plethysmographic data and for obtaining a measure of at least one physiological parameter from the data. The method comprises acquiring the raw photo-plethysmographic data, transforming the data into the frequency domain, analyzing the transformed data to locate a series of candidate cardiac spectral peaks (primary plus harmonics), reconstructing a photo-plethysmographic signal in the time domain with only the candidate cardiac spectral peaks (primary plus harmonics), computing the second order derivative of the reconstructed photo-plethysmographic signal, analyzing the candidate second order derivative photo-plethysmographic signal to determine the absence or presence of cardiac physiologic signal characteristics, and finally selecting the best physiologic candidate from the series of potential cardiac spectral peaks (primary plus harmonics) based upon a second derivative scoring system. This scoring system is preferentially based upon second derivative processing analysis, but can be equally applied using the first, third, fourth or other similar derivative processing analysis. | 1. A method for use in a system comprising a sensor for acquiring photo-plethysmographic data of a patient and a signal processing system for removing a plurality of motion artifacts from said photo-plethysmographic data and for obtaining a measure of at least one physiological parameter from said data, the method comprising: (a) acquiring said photo-plethysmographic data;(b) transforming said data;(c) analyzing said transformed data to locate candidate spectral peaks;(d) reconstructing a photo-plethysmographic signal;(e) computing the second order derivative of said reconstructed photo-plethysmographic signal;(f) analyzing said second order derivative photo-plethysmographic signal to determine a plurality of parameters;(g) identifying signal representations of a plurality of motion artifacts; and(h) selecting a candidate spectral peak from a plurality of potential cardiac spectral peaks, wherein analyzing said transformed data to locate candidate spectral peaks comprises the steps of assigning a largest power amplitude from said data as a primary candidate spectral peak and assigning a next largest power amplitude that is not a harmonic of said primary candidate spectral peak as a secondary candidate spectral peak, wherein said analyzing transformed data to locate candidate spectral peaks is performed using a first derivative criteria, and wherein said candidate spectral peak is a frequency bin having a slope value equal to approximately zero. | 1. A method for use in a system comprising a sensor for acquiring photo-plethysmographic data of a patient and a signal processing system for removing a plurality of motion artifacts from said photo-plethysmographic data and for obtaining a measure of at least one physiological parameter from said data, the method comprising: (a) acquiring said photo-plethysmographic data;(b) transforming said data;(c) analyzing said transformed data to locate candidate spectral peaks;(d) reconstructing a photo-plethysmographic signal;(e) computing the second order derivative of said reconstructed photo-plethysmographic signal;(f) analyzing said second order derivative photo-plethysmographic signal to determine a plurality of parameters;(g) identifying signal representations of a plurality of motion artifacts; and(h) selecting a candidate spectral peak from a plurality of potential cardiac spectral peaks, wherein analyzing said transformed data to locate candidate spectral peaks comprises the steps of assigning a largest power amplitude from said data as a primary candidate spectral peak and assigning a next largest power amplitude that is not a harmonic of said primary candidate spectral peak as a secondary candidate spectral peak, wherein said analyzing transformed data to locate candidate spectral peaks is performed using a first derivative criteria, and wherein said candidate spectral peak is a frequency bin having a slope value equal to approximately zero. 2. The method of claim 1, wherein said first derivative criteria comprises a first criterion for detecting extrema, having a plurality of slopes, and a second criterion for verifying maxima or minima, wherein said extrema are considered maxima or minima based upon whether a second derivative of the slopes of extrema is lesser than or greater than zero. 3. A method for use in a system comprising a sensor for acquiring photo-plethysmographic data of a patient and a signal processing system for removing a plurality of motion artifacts from said photo-plethysmographic data and for obtaining a measure of at least one physiological parameter from said data, the method comprising: (a) acquiring said photo-plethysmographic data;(b) transforming said data;(c) analyzing said transformed data to locate candidate spectral peaks;(d) reconstructing a photo-plethysmographic signal;(e) computing the second order derivative of said reconstructed photo-plethysmographic signal;(f) analyzing said second order derivative photo-plethysmographic signal to determine a plurality of parameters;(g) identifying signal representations of a plurality of motion artifacts; and(h) selecting a candidate spectral peak from a plurality of potential cardiac spectral peaks, wherein the second order derivative photo-plethysmographic signal is analyzed to determine the absence or presence of cardiac physiologic signal characteristics by performing a morphological analysis of the second order derivative photo-plethysmographic signal to determin... | Dolphin Medical, Inc. | A61B005/00
US07177675 | B2 | 2007-02-13 | Electroencephalography based systems and methods for selecting therapies and predicting outcomes | 2002-07-11 | A method and system for utilizing neurophysiologic information obtained by techniques such as quantitative electroencephalography (QEEG), electrode recordings, MRI in appropriately matching patients with therapeutic entities is disclosed. The present invention enables utilization of neurophysiologic information, notwithstanding its weak correlation with extant diagnostic schemes for mental disorders, for safer and expeditious treatment for mental disorders, discovering new applications for therapeutic entities, improved testing of candidate therapeutic entities, inferring the presence or absence of a desirable response to a treatment, and deducing the mode of action of one or more therapeutic entities. In particular, methods for effectively comparing neurophysiologic information relative to a reference set are disclosed along with database-based tools for deducing therapeutic entity actions on particular patients such that these tools are readily accessible to remote users. | 1. A method for identifying an outcome of at least a first treatment comprising: a) providing a first electroencephalogram;b) scaling said first electroencephalogram to enable comparison with a treatment-response database, wherein said database comprises stored electroencephalographic information obtained from a first plurality of subjects;c) computing at least one indicative variable from said first electroencephalogram; andd) identifying a cluster derived from said database under conditions that predict at least a first outcome of a first treatment prior to actually administering said first treatment, wherein said cluster comprises a second plurality of subjects and at least eighty percent of said second plurality of subjects have a common response to said first treatment. | 1. A method for identifying an outcome of at least a first treatment comprising: a) providing a first electroencephalogram;b) scaling said first electroencephalogram to enable comparison with a treatment-response database, wherein said database comprises stored electroencephalographic information obtained from a first plurality of subjects;c) computing at least one indicative variable from said first electroencephalogram; andd) identifying a cluster derived from said database under conditions that predict at least a first outcome of a first treatment prior to actually administering said first treatment, wherein said cluster comprises a second plurality of subjects and at least eighty percent of said second plurality of subjects have a common response to said first treatment. 2. The method of claim 1, further comprising identifying said at least one indicative variable by screening said treatment-response database. 3. The method of claim 1, further comprising identifying said at least one indicative variable within said cluster. 4. The method of claim 1, wherein said treatment-response database comprises a clinical global improvement score, said score comprising an integer selected from the group consisting of negative one (−1) indicating an adverse therapeutic effect, zero (0) indicating no improvement, positive one (+1) indicating a minimal improvement, positive two (+2) indicating a moderate improvement, and positive three (+3) indicating a complete absence of symptoms. 5. The method of claim 1, wherein said indicative variable is calculated by the difference between said at least one treatment-response database univariate value within said treatment-response database and at least one reference database univariate value, said reference database comprising age-matched reference subjects. 6. The method of claim 1, wherein said first outcome is capable of improving at least one symptom of a traditionally diagnosed mental disease. 7. The method of claim 6, further comprising identifying a second treatment by comparing said first outcome to a second outcome, said second outcome resulting from said second treatment. 8. The method of claim 7, wherein said traditionally diagnosed mental disease is a major depressive disorder and said second treatment is selected from the group consisting of glutamine, phenylalanine, and tyrosine. 9. The method of claim 7, wherein said traditionally diagnosed mental disease comprises psychological factors affecting atypical asthma and said second treatment is selected from the group consisting of glutamine, phenylalanine, tyrosine, bupropion, pamate, moclobemide, phenelzine, selegiline, venlafaxine, carbamazapine, gabapentin, lamotrigine, ginko biloba, dexedrine, methamphetamine, methylphenidate, and pemoline. 10. The method of claim 7, wherein said traditionally diagnosed mental disease comprisies anxiety disorders and said second treatment is selected from the group consisting of gaba, glutamine, phenylalanin... | CNS Response, Inc | A61B005/02
US07291117 | B2 | 2007-11-06 | Method and apparatus for penetrating tissue | 2002-12-31 | These and other objects of the present invention are achieved in a body fluid sampling system for use on a tissue site that includes a voice coil based penetrating member driver. A plurality of penetrating members are included. A transport mechanism is configured to engage the cartridges. Each of the cartridges is operatively engaged with the penetrating member driver when moved into position by the transport mechanism. | 1. A body fluid sampling system for use on a tissue site, the system comprising: a voice coil based penetrating member driver;a plurality of penetrating members housed in or on a cartridge;a transport mechanism configured to move the cartridge to a position for an alignment with a penetrating member with the voice-coil based penetrating member driver to drive the penetrating member along a path into the tissue site, wherein each of said penetrating member is operatively engaged with said penetrating member driver when moved into position by said transport mechanism. | 1. A body fluid sampling system for use on a tissue site, the system comprising: a voice coil based penetrating member driver;a plurality of penetrating members housed in or on a cartridge;a transport mechanism configured to move the cartridge to a position for an alignment with a penetrating member with the voice-coil based penetrating member driver to drive the penetrating member along a path into the tissue site, wherein each of said penetrating member is operatively engaged with said penetrating member driver when moved into position by said transport mechanism. 2. A system as in claim 1 further comprising a penetrating member sensor positioned to monitor a penetrating member coupled to said force generator, the penetrating member sensor configured to provide information relative to a depth of penetration of a penetrating member through a skin surface. 3. The system of claim 2, wherein the depth of penetration is about 100 to 2500 microns. 4. The system of claim 2, wherein the depth of penetration is 500 to 750 microns. 5. The system of claim 2, wherein the depth of penetration is no more than about 1000 microns beyond a stratum corneum thickness of a skin surface. 6. The system of claim 2, wherein the depth of penetration is no more than about 500 microns beyond a stratum corneum thickness of a skin surface. 7. The system of claim 2, wherein the depth of penetration is no more than about 300 microns beyond a stratum corneum thickness of a skin surface. 8. The system of claim 2, wherein the depth of penetration is less than a sum of a stratum corneum thickness of a skin surface and 400 microns. 9. The system of claim 2, wherein the penetrating member sensor is further configured to provide an indication of velocity of a penetrating member. 10. The system of claim 2, wherein the penetrating member sensor is coupled to a processor with control instructions for the penetrating member driver. 11. The system of claim 10, wherein the processor includes a memory for storage and retrieval of a set of penetrating member profiles utilized with the penetrating member driver. 12. The system of claim 10, wherein the processor is utilized to monitor position and speed of a penetrating member as the penetrating member moves in a first direction. 13. The system of claim 10, wherein the processor is utilized to adjust an application of force to a penetrating member to achieve a desired speed of the penetrating member. 14. The system of claim 10, wherein the processor is utilized to adjust an application of force to a penetrating member when the penetrating member contacts a target tissue so that the penetrating member penetrates the target tissue so that the penetrating member penetrates the target tissue within a desired range of speed. 15. The system of claim 10, wherein the processor is utilized to monitor position and speed of a penetrating member as the penetrating member moves in the first direction toward a target tissue, wherein the appl... | Pelikan Technologies, Inc. | A61B005/00
US07184819 | B2 | 2007-02-27 | User interface system for use in ECG signal derivation and management | 2003-01-30 | The system provides a Graphical User Interface which automatically updates labels associated with waveform information in multiple different windows within an individual image and in different images for use in a medical information telemetry system in synthesis and non-synthesis modes. A user interface display system for use in ECG signal management includes a display generator for generating at least one display image enabling assignment of two measured ECG chest lead signals to any two signals of the six ECG chest lead signals of a conventional 12 lead ECG signal set. A command processor automatically updates labels associated with the six ECG chest lead signals displayed in a display image to be compatible with the assignment of the two measured ECG chest lead signals in response to an assignment command made via the at least one display image. | 1. A user interface display system for use in ECG signal management, comprising: a display generator for generating at least one display image enabling assignment of two measured EGG chest lead signals to any two signals of the six EGG chest lead signals of a conventional 12 lead EGG signal set; anda command processor for automatically updating labels for identifying, (a) individual non-synthesized measured EGG chest lead signals and (b) individual synthesized non-measured EGG chest lead signals, associated with said six EGG chest lead signals displayed in a display image to be compatible with said assignment of said two measured EGG chest lead signals in response to an assignment command made via said at least one display image. | 1. A user interface display system for use in ECG signal management, comprising: a display generator for generating at least one display image enabling assignment of two measured EGG chest lead signals to any two signals of the six EGG chest lead signals of a conventional 12 lead EGG signal set; anda command processor for automatically updating labels for identifying, (a) individual non-synthesized measured EGG chest lead signals and (b) individual synthesized non-measured EGG chest lead signals, associated with said six EGG chest lead signals displayed in a display image to be compatible with said assignment of said two measured EGG chest lead signals in response to an assignment command made via said at least one display image. 2. The system of claim 1 wherein said command processor automatically updates labels associated with said six EGG chest lead signals to identify two measured chest lead signals and four derived chest lead signals, in response to an assignment command made via said at least one display image. 3. The system of claim 1 wherein said two measured chest lead signals are the V and V+ lead signals of a conventional 12 lead EGG signal set. 4. The system of claim 1 wherein said command processor automatically updates labels associated with said six EGG chest lead signals to identify two measured chest lead signals and four derived chest lead signals in multiple display images including at least one of, (a) an EGG waveform display image, (b) a configuration screen for use in configuring a parameter display image for a particular patient and (c) a telemetry system configuration image, in response to an assignment command made via said at least one display image. 5. The system of claim 1 wherein said command processor automatically updates labels associated with said six ECG chest lead signals in a menu window used for assigning a particular chest lead EGG signal to a window used for displaying a corresponding particular chest lead EGG signal waveform. 6. The system of claim 1 wherein said command processor automatically updates labels associated with said six ECG chest lead signals in a menu window used for assigning a particular chest lead EGG signal to a display channel. 7. The system of claim 1 including a data processor for, in a synthesis mode, identifying and selecting coefficients from a plurality of stored coefficients and applying a transformation, using said selected coefficients, to data comprising said two measured EGG chest lead signals to derive data representing a plurality of non-measured EGG chest lead signals and whereinsaid command processor, in said synthesis mode automatically updates labels to identify individual synthesized non-measured EGG chest lead signals and individual non-synthesized measured EGG chest lead signals in response to said assignment command made via said at least one display image. 8. The system of claim 7 wherein said command processor, in a non-synthesis mode, automatically updates ... | Draeger Medical Systems, Inc. | A61B005/044
US07263399 | B2 | 2007-08-28 | Apparatus and methods for analysis of cardiac device stored episodes containing sensed signals and waveforms | 2002-04-29 | Apparatus and methods for analyzing implantable medical device (IMD) data and to distinguish specific morphologies for review. EGM waveforms of a cardiac cycle can be grouped according to waveform shape similarities, preferably accomplished using clustering algorithms. The EGM waveform can be divided into segments represented by a one dimensional vector for each segment, and nearest neighbor vector clustering used to group the waveforms according to morphological similarity. The waveforms can be further grouped at a second level into chronologically contiguous similar EGM waveforms. Interval data for a cardiac cycle may also be grouped according to interval similarity and chronology. The chronologically contiguous waveform and interval groupings may be further grouped. Large quantities of EGM data are automatically analyzed and annotated to focus the attention of a cardiologist on likely areas of concern, reducing the need to visually inspect endless streams of EGM waveforms. | 1. A method for analyzing stored cardiac EGM waveform data for a plurality of a first heart chamber cycles, the data including a plurality of waveform cycle segments, the data initially originating from an implanted device, the method comprising: grouping the waveform cycle segments into a first set of cycle groups as a function of waveform shape similarity between the waveform cycle segments, wherein each of the first set group members have similar waveform shapes to the other members of the same first set group;wherein the grouping includes clustering the waveform cycle segments based on waveform shape similarity;wherein the waveform cycle segments are treated as vectors; andwherein the clustering clusters includes using the Euclidean distance between the vectors. | 1. A method for analyzing stored cardiac EGM waveform data for a plurality of a first heart chamber cycles, the data including a plurality of waveform cycle segments, the data initially originating from an implanted device, the method comprising: grouping the waveform cycle segments into a first set of cycle groups as a function of waveform shape similarity between the waveform cycle segments, wherein each of the first set group members have similar waveform shapes to the other members of the same first set group;wherein the grouping includes clustering the waveform cycle segments based on waveform shape similarity;wherein the waveform cycle segments are treated as vectors; andwherein the clustering clusters includes using the Euclidean distance between the vectors. 2. A method as in claim 1, wherein the clustering includes nearest neighbor clustering. 3. A method as in claim 1, further comprising: grouping the waveform cycle segments into a second set of cycle groups, wherein the second set group members are substantially chronologically contiguous to the other second set group members and are members of the same first set group. 4. A method as in claim 3, further comprising processing the first set groups by executing one of a step of marking, storing, displaying, and outputting the group memberships, and combinations thereof. 5. A method for analyzing stored cardiac EGM data from an implantable device, the method comprising: retrieving the stored data, wherein the data includes EGM waveform data;dividing the EGM waveform data into a plurality of cardiac cycle waveform segments, wherein each heartbeat waveform segment is associated with a heartbeat; andgrouping the EGM cardiac cycle waveform segments into waveform identity groups as a function of waveform shape similarities between the waveform segments with members of the same waveform identity group having similar waveform shapes to the other waveform identity group members;wherein the grouping includes calculating a waveform shape distance between the cardiac cycle waveform segments and clustering the cardiac cycle waveform segments as a function of the shape distance between the heartbeat waveform segments;wherein each cardiac cycle waveform segment is represented as a vector and the shape distance is a distance between the vectors; andwherein the vector distance is a Euclidean distance. 6. A method as in claim 5, further comprising marking the EGM heartbeat waveform segments with an indicia of identity group membership and storing the marked EGM identity waveform segments in nonvolatile storage. 7. A method as in claim 5, further comprising the step of visually displaying the EGM heartbeat waveform segments together with a visual indicia of the waveform identity group membership for each EGM heartbeat waveform segment. 8. A method as in claim 5, further comprising collecting temporally adjacent heartbeat waveform segments having the same heartbeat waveform identity group membership ... | Medtronic, Inc. | A61B005/04
US07076282 | B2 | 2006-07-11 | Bioadhesive compositions and biomedical electrodes containing them | 2002-10-29 | Bioadhesive Compositions which comprise a hydrophobic polymer wherein the concentration of the polymer at the surface of the adhesive is greater than its concentration in the bulk of the adhesive are described; and biomedical electrodes, fixation products and wound dressings containing them. | 1. A bioadhesive composition comprising: (a) an aqueous plasticiser, and(b) a copolymer of: a first hydrophilic unsaturated water-soluble monomer, which has the formula wherein R1 is an optionally substituted hydrocarbon moiety, R2 is hydrogen or an optionally substituted methyl or ethyl group, and M is hydrogen or a cation, anda second hydrophilic unsaturated water-soluble monomer, which has the formula wherein R5 is hydrogen or an optionally substituted methyl or ethyl group, R6 is hydrogen or a cation and R7 is an optionally substituted alkylene moiety of 1 to 4 carbon atoms, anda cross-linking agent;the composition having: (i) a water activity in the range of 0.4 to 0.9;(ii) an elastic modulus at 1 rad/s in the range of 700 to 15,000 Pa;(iii) an elastic modulus at 100 rad/s in the range of 2000 to 40,000 Pa;(iv) a viscous modulus at 1 rad/s in the range of 400 to 14,000 Pa;(v) a viscous modulus at 100 rad/s in the range of 1000 to 35,000 Pa;wherein the viscous modulus is less than the elastic modulus in a frequency range of 1 to 100 rad/s. | 1. A bioadhesive composition comprising: (a) an aqueous plasticiser, and(b) a copolymer of: a first hydrophilic unsaturated water-soluble monomer, which has the formula wherein R1 is an optionally substituted hydrocarbon moiety, R2 is hydrogen or an optionally substituted methyl or ethyl group, and M is hydrogen or a cation, anda second hydrophilic unsaturated water-soluble monomer, which has the formula wherein R5 is hydrogen or an optionally substituted methyl or ethyl group, R6 is hydrogen or a cation and R7 is an optionally substituted alkylene moiety of 1 to 4 carbon atoms, anda cross-linking agent;the composition having: (i) a water activity in the range of 0.4 to 0.9;(ii) an elastic modulus at 1 rad/s in the range of 700 to 15,000 Pa;(iii) an elastic modulus at 100 rad/s in the range of 2000 to 40,000 Pa;(iv) a viscous modulus at 1 rad/s in the range of 400 to 14,000 Pa;(v) a viscous modulus at 100 rad/s in the range of 1000 to 35,000 Pa;wherein the viscous modulus is less than the elastic modulus in a frequency range of 1 to 100 rad/s. 2. A bioadhesive composition according to claim 1, wherein R1 is an optionally substituted alkyl, cycloalkyl or aromatic moiety containing from 3 to 12 carbon atoms. 3. A bioadhesive composition according to claim 1, wherein R1 iswherein R3 is hydrogen or an optionally substituted straight or branched chain alkyl group possessing from 1 to 6 carbon atoms and R4 is an optionally substituted straight or branched chain alkyl group possessing from 1 to 6 carbon atoms. 4. A bioadhesive composition according to claim 1, wherein said first monomer is 2-acrylamido-2-methyl-propanesulfonic acid or a salt thereof. 5. A bioadhesive composition according to claim 4, wherein said salt is a sodium, potassium or lithium salt. 6. A bioadhesive composition according to claim 1, wherein substituents of R1, R2, R5 and R7, when present, are selected from the group consisting of hydroxyl, amino and ammonium groups and halogen atoms. 7. A bioadhesive composition according to claim 1, wherein said second monomer is a salt of acrylic acid (3-sulfopropyl) ester. 8. A bioadhesive composition according to claim 7, wherein said salt is a sodium, potassium or lithium salt. 9. A bioadhesive composition according to claim 1, wherein said aqueous plasticiser includes a polymeric or non-polymeric polyhydric alcohol, an ester derived therefrom, a polymeric alcohol, or any mixture thereof. 10. A bioadhesive composition according to claim 1, wherein said aqueous plasticiser includes glycerol or an ester derived from boric acid and a polyhydric alcohol. 11. A bioadhesive composition according to claim 1, further comprising a surfactant. 12. A bioadhesive composition according to claim 1, wherein said copolymer further comprises an additional monomer. 13. A bioadhesive composition according to claim 12, wherein said additional monomer is selected from the group consisting of an N-disubstituted acrylamide, (3-acrylamido-propyl)-trime... | First Water Limited | A61B005/04
US07276254 | B2 | 2007-10-02 | Emulsion/aggregation polymeric microspheres for biomedical applications and methods of making same | 2002-05-07 | A method of forming polymeric microspheres for biomedical applications includes forming polymeric microspheres by an emulsion/aggregation process from a precursor monomer species, and treating the polymeric microspheres to attach a biomedical functional material to the polymeric microspheres, where the polymeric microspheres have an average particle diameter of from about 1 to about 15 microns with a narrow particle geometric size distribution. The biomedical functional material may be, for example, a radioactive material, a radioactive precursor material, a bioactive agent, or a ligand. | 1. A method of forming polymeric microspheres for biomedical applications, comprising: forming polymeric microspheres by an emulsion/aggregation process from a precursor monomer species; andattaching a biomedical functional material selected from the group consisting of a radioactive material, radioactive precursor material, a bioactive agent, and a ligand to said polymeric microspheres,wherein said polymeric microspheres have an average particle diameter of from about 1 to about 15 microns with a narrow particle geometric size distribution of less than about 1.25, andwherein said emulsion/aggregation process comprises forming a polymeric resin from said precursor monomer species and aggregating and coalescing said polymeric resin into polymeric microspheres. | 1. A method of forming polymeric microspheres for biomedical applications, comprising: forming polymeric microspheres by an emulsion/aggregation process from a precursor monomer species; andattaching a biomedical functional material selected from the group consisting of a radioactive material, radioactive precursor material, a bioactive agent, and a ligand to said polymeric microspheres,wherein said polymeric microspheres have an average particle diameter of from about 1 to about 15 microns with a narrow particle geometric size distribution of less than about 1.25, andwherein said emulsion/aggregation process comprises forming a polymeric resin from said precursor monomer species and aggregating and coalescing said polymeric resin into polymeric microspheres. 2. The method of claim 1, wherein said biomedical functional material is a ligand. 3. The method of claim 2, wherein said ligand is a pharmacologically active compound selected from the group consisting of peptides, enzymes, analytes, antigens, and antibodies. 4. The method of claim 2, wherein said ligand is an antigen selected from the group consisting of protein, peptide, polysaccharide, glycoprotein, glycolipid, nucleic acid, and mixtures thereof. 5. The method of claim 1, wherein said biomedical functional material is a radioactive material or a radioactive precursor material. 6. The method of claim 5, wherein said biomedical functional material is selected from the group consisting of yttrium-89, yttrium-90, phosphorus-31, and phosphorus-32. 7. The method of claim 1, wherein said biomedical functional material is a bioactive agent. 8. The method of claim 7, wherein said bioactive agent is a medicament. 9. The method of claim 7, wherein said bioactive agent is selected from the group consisting of antibiotics, antimicrobials, antiseptics, bacteriocins, bacteriostats, disinfectants, steroids, anesthetics, fungicides, anti-inflammatory agents, antibacterial agents, antiviral agents, antitumor agents, growth promoters, and mixtures thereof. 10. The method of claim 1, wherein the biomedical functional material is attached to said polymeric microspheres by at least one of covalent bonding, complexation, physical adsorption and physical absorption. 11. The method of claim 1, wherein the polymeric microspheres are formed from said precursor monomer species and one or more additives. 12. The method of claim 11, wherein the one or more additives are selected from the group consisting of colorants, fluorescent materials, magnetic materials, superparamagnetic materials, and bioactive agents. 13. The method of claim 12, wherein the one or more additives comprises a colorant. 14. The method of claim 12, wherein the one or more additives comprises a magnetic material. 15. The method of claim 14, wherein the magnetic material is a magnetite. 16. The method of claim 12, wherein the one or more additives comprises a superparamagnetic material. 17. The method of claim 12, wherein the one o... | Xerox Corporation | A61K009/14 A61K051/00 A61K009/16 A61M036/14 A61B005/055
US07118534 | B2 | 2006-10-10 | Methods for monitoring and optimizing central venous pressure and intravascular volume | 2001-09-21 | Central Venous Pressure (CVP) is non-invasively determined with accuracy comparable to invasive measurement techniques. To do so, curves are plotted based on non-invasively determined patient information obtained by applying a controllable variable (pressure) to a vein of interest at a non-distal point and taking certain measurements (such as pressure and volume measurements) from the patient. An example of a controllable variable is voltage applied in incremental inflation/deflation of a vascular cuff (1). A curve is plotted based on datapoints (such as a volume increase curve or a volume decline curve). Pertinent, accurate CVP and/or blood volume information is obtained from the slope of the non-invasive-based curve. Accurate CVP information is provided without the risks and disadvantages of invasive measurements. | 1. A method for non-invasively determining and optimizing Central Venous Pressure (CVP) comprising the steps of: applying a controllable pressure to a vein of interest at a non-distal point of a patient;taking pressure and volume measurements from the patient to generate data points;plotting intravascular measurements based on the data points as pressure and volume change curves;determining from the pressure and volume and change curves an overall slope of the volume change curve and a maximum slope of the volume change curve; anddetermining the CVP based on pressure at the volume change curve maximum slope. | 1. A method for non-invasively determining and optimizing Central Venous Pressure (CVP) comprising the steps of: applying a controllable pressure to a vein of interest at a non-distal point of a patient;taking pressure and volume measurements from the patient to generate data points;plotting intravascular measurements based on the data points as pressure and volume change curves;determining from the pressure and volume and change curves an overall slope of the volume change curve and a maximum slope of the volume change curve; anddetermining the CVP based on pressure at the volume change curve maximum slope. 2. The method of claim 1, including the step of absolutely determining central vascular pressure. 3. The method of claim 1, wherein the step of applying a controllable pressure includes inflating and/or deflating a vascular cuff surrounding an extremity upper part. 4. The method of claim 1, wherein the step of applying a controllable pressure includes inflating and/or deflating a vascular cuff surrounding an extremity upper part and further including the step of determining the rate of venous volume change upon deflation of said cuff. 5. The method of claim 1, wherein the step of applying a controllable pressure includes inflating and/or deflating a vascular cuff surrounding an extremity upper part and further including the step of determining the rate of venous volume change upon inflation of the vascular cuff. 6. The method of claim 1, wherein the step of applying a controllable pressure includes inflating and/or deflating a vascular cuff surrounding an extremity upper part and further including the step of determining the threshold of venous volume change upon deflation of the vascular cuff. 7. The method of claim 1, including the step of measuring pressure for a large vein. 8. The method of claim 1, wherein the step of including the step of wherein the step of applying a controllable pressure includes surrounding an upper extremity with an inflatable cuff, and where the method further includes making the volume and pressure measurements while inflating the cuff. 9. The method of claim 8, wherein the inflatable cuff is electrically and pneumatically connected to a programmed cuff inflator device. 10. The method of claim 9, wherein a computer processor receives pressure measurements from the inflatable cuff. 11. The method of claim 10, wherein a volume sensitive detector sends information to the computer processor via a signal amplifier. 12. The method of claim 1, further including the step of non-invasively detecting volume sensitive data for a distal part of an extremity by operating a volume sensitive detector selected from the group consisting of a mercury-in-silastic strain gauge, a bioimpedance device, a photo device, a plethysmography device, a laser Doppler device and a spectroscopic volume detection device. 13. The method of claim 1, wherein the steps of plotting intravascular measurements, determining an overall slope ... | Virginia Commonwealth University | A61B005/02
US07135000 | B2 | 2006-11-14 | Heart state monitor method | 2003-01-17 | A heart state monitor method is proposed. The physiologic parameter of a life body is measured and then preformed by Fourier transform analysis. Next, the frequency domain data of the physiologic parameter are converted into a power spectral density (PSD). A heart beat noise index of the life body is then obtained according to the PSD. Finally, the heart state of the life body is diagnosed according to the heart beat noise index. The heart beat noise index may be quickly discriminated to let the life body and medical staffs grasp the current heart state of the life body at any time. | 1. A heart state monitor method comprising the steps of: (a) measuring a physiologic parameter of a life body by using an infrared physiology measurement device; (b) converting said physiologic parameter from the time domain to the frequency domain by means of Fourier transform and then converting the frequency domain data into a power spectral density; (c) obtaining a heart beat noise index of said life body according to said power spectral density; and (d) diagnosing the heart state of said life body according to said heart beat noise index. | 1. A heart state monitor method comprising the steps of: (a) measuring a physiologic parameter of a life body by using an infrared physiology measurement device; (b) converting said physiologic parameter from the time domain to the frequency domain by means of Fourier transform and then converting the frequency domain data into a power spectral density; (c) obtaining a heart beat noise index of said life body according to said power spectral density; and (d) diagnosing the heart state of said life body according to said heart beat noise index. 2. The heart state monitor method as claimed in claim 1, wherein said physiologic parameter is a heart beat signal. 3. The heart state monitor method as claimed in claim 1, wherein said physiologic parameter can be obtained by measuring an arm, wrist, finger or ankle of said life body. 4. The heart state monitor method as claimed in claim 1, wherein said physiologic parameter of said life body before analysis and said frequency domain data and said power spectral density after analysis can further be stored into a database. 5. The heart state monitor method as claimed in claim 1 further comprising a step of transmitting said physiologic parameter to a medical center via a transmission circuit for performing said Step (b) after said step (a). 6. The heart state monitor method as claimed in claim 5, wherein said transmission circuit is a telephone circuit, an ISDN, an ADSL, a HDSL, a TCP/IP, a cable circuit, an infrared transmission circuit or an radio-frequency transmission circuit. 7. The heart state monitor method as claimed in claim 1, wherein many noise waveforms between each harmonic in said power spectral density can further be used to calculate out said heart beat noise index of said life body. 8. A heart state monitor method comprising the steps of: (a) measuring a physiologic parameter of a life body; (b) analyzing said physiologic parameter to obtain envelope functions thereof; (c) calculating a roughness integral and a roughness differential according to said envelope functions; and (d) diagnosing the heart state of said life body according to said roughness integral and said roughness differential. 9. The heart state monitor method as claimed in claim 8, wherein said physiologic parameter can be obtained by measuring an arm, wrist, finger or ankle of said life body. | Kuo-Yuan Chang Kuo-Ku Chang | A61B005/02
US06975897 | B2 | 2005-12-13 | Short/long axis cardiac display protocol | 2002-03-15 | A method for generating views of a heart along anatomically useful planes includes generating a prescription image of a heart, receiving data regarding a position of at least one landmark on the prescription image to define a series of slices through the heart along at least one anatomical plane lane, generating a plurality of images under software control, and saving the generated images. | 1. A method for generating views of a heart along anatomically useful planes, comprising: generating a prescription image of a heart;receiving data regarding a position of at least one landmark on the prescription image to define a series of slices through the heart along at least one anatomical plane;generating a plurality of images;saving the generated images; andselecting a short/long axis protocol. | 1. A method for generating views of a heart along anatomically useful planes, comprising: generating a prescription image of a heart;receiving data regarding a position of at least one landmark on the prescription image to define a series of slices through the heart along at least one anatomical plane;generating a plurality of images;saving the generated images; andselecting a short/long axis protocol. 2. A method in accordance with claim 1 wherein said generating a prescription image of a heart further comprises loading at least one a volume data set that represents a heart at a desired phase and a plurality of volume data sets of the heart at a plurality of desired phases in the computer. 3. A method in accordance with claim 1 wherein said receiving data regarding a position of at least one landmark on the prescription image further comprises assisting a user with a plurality of visual aids to position the landmarks on the prescription image. 4. A method in accordance with claim 3 wherein said assisting a user with visual aids to position the landmarks further comprises assisting a user with at least one of an example image and a detailed text to position the landmarks. 5. A method in accordance with claim 1 wherein said generating a plurality of images under software control further comprises generating a plurality of 4D images under software control. | GE Medical Systems Global Technology Company, LLC | A61B005/05
US06934574 | B1 | 2005-08-23 | MRI scanner and method for modular patient handling | 2002-06-20 | An MRI scanner that includes a magnet with poles having substantially vertical pole surfaces and a removable positioner for supporting a patient within a gap in the pole surfaces. A method of positioning a patient within a gap formed by vertical pole surfaces in an MRI scanner. The method includes selecting a first patient positioner and inserting the first patient positioner into the gap between vertical pole surfaces in the MRI scanner. | 1. An MRI scanner comprising: a magnet including poles having substantially vertical pole surfaces such that there is a gap between the pole surfaces; anda removable positioner for supporting a patient within the gap; andan elevator positioned below the gap to raise and lower the positioner relative to the poles of the magnet, wherein the elevator includes a support platform and a scissor mechanism that raises and lowers the support platform. | 1. An MRI scanner comprising: a magnet including poles having substantially vertical pole surfaces such that there is a gap between the pole surfaces; anda removable positioner for supporting a patient within the gap; andan elevator positioned below the gap to raise and lower the positioner relative to the poles of the magnet, wherein the elevator includes a support platform and a scissor mechanism that raises and lowers the support platform. 2. An MRI scanner comprising: a magnet including poles having substantially vertical pole surfaces such that there is a gap between the pole surfaces; anda removable positioner for supporting a patient within the gap; andan elevator positioned below the gap to raise and lower the removable positioner relative to the poles of the magnet, wherein the elevator includes a sensor and a control coupled to the sensor to position the height of the elevator. 3. An MRI scanner comprising: a magnet including poles having substantially vertical pole surfaces such that there is a gap between the pole surfaces; anda removable positioner for supporting a patient within the gap, wherein the removable positioner includes a control and radio frequency (RF) connection ports that exchange signals with the control to orient the removable positioner. 4. An MRI scanner comprising: a magnet including poles having substantially vertical pole surfaces such that there is a gap between the pole surfaces; anda removable positioner for supporting a patient within the gap wherein the magnet includes a front side and a back side, the gap communicating with the front and back sides of the magnet such that the removable positioner can enter the gap from the front side and the back side of the magnet. 5. An MRI scanner comprising: a magnet including poles having substantially vertical pole surfaces such that there is a gap between the pole surfaces; anda removable positioner for supporting a patient within the gap, wherein the removable positioner includes a pair of gates, each gate including a support section such that a patient can grasp the support section of one gate and lean against the support section of the other gate. 6. The MRI scanner of claim 5 wherein the removable positioner includes an elevator such that the pair of gates are rotatably mounted on the elevator. 7. A method of positioning a patient within a gap formed by vertical pole surfaces in an MRI scanner, the method comprising: selecting a first patient positioner;inserting the first patient positioner into the gap between the vertical pole surfaces; andremoving a second patient positioner from between the gap before inserting the first patient positioner into the gap, wherein inserting the first patient positioner into the gap between the vertical pole surfaces includes inserting the first patient positioner from a front side of the MRI scanner and removing a second patient positioner from between the gap includes removing the second patient positioner from a back si... | Fonar Corporation | A61B005/05
US07223388 | B2 | 2007-05-29 | Modified reoviral therapy | 2002-08-02 | The present invention relates to hyperproliferative diseases. Specifically, the present invention encompasses pharmaceutical compositions comprising a modified Reoviridae virus, wherein the Reoviridae virus is conjugated to a hydroxylated hydrocarbon or a polycationic polymer to reduce the clearance of the composition and reduce the immunogenicity of the composition. Yet further, the invention relates to methods of treating a hyperproliferative disease by administering to a patient an effective amount of the modified Reoviridae virus. | 1. A pharmaceutical composition comprising a Reoviridae virus conjugated to one or more carbohydrates dispersed in a pharmaceutically acceptable carrier. | 1. A pharmaceutical composition comprising a Reoviridae virus conjugated to one or more carbohydrates dispersed in a pharmaceutically acceptable carrier. 2. The pharmaceutical composition of claim 1, wherein the Reoviridae virus is a reovirus. 3. The pharmaceutical composition of claim 2, wherein the reovirus is human serotype 3 reovirus. 4. The pharmaceutical composition of claim 1 further comprising a cell-targeting moiety conjugated with said virus within the pharmaceutical composition. 5. The pharmaceutical composition of claim 4, wherein the targeting moiety is an antibody. 6. The pharmaceutical composition of claim 5, wherein the antibody is C225, Trastuzuab, 1D5, PB3, TA-1, 4d5, Edrecolomab, CC49, Mc5, N901, CD56, MDX447, HMFG1, or IDEC-C2B8. 7. The pharmaceutical composition of claim 1, wherein the carbohydrate is mannose, galactose, dextran, cellulose, starch, glycogen, agarose, guar gum, pullulan, inulin, xanthan gum, carrageenin, pectin, or alginic acid hydrolysates. 8. The pharmaceutical composition of claim 1, wherein the carbohydrate is dextran. | Board of Regents, The Univeristy of Texas System | A61K039/15 A61K048/00 A61K047/00 A61K047/02 A61K039/00 A61B005/055
US07248916 | B2 | 2007-07-24 | Automated system and method for establishing a patient status reference baseline | 2001-02-20 | A system and method for determining a reference baseline of individual patient status for use in an automated collection and analysis patient care system is described. A set of collected measures is retrieved from a medical device adapted to be implanted in a patient. The collected device measures set includes individual measures which each relate to patient information recorded by the medical device adapted to be implanted during an initial time period. The collected device measures set is received from the medical device adapted to be implanted over a communications link which is interfaced to a network server. The collected device measures set is stored into a patient care record for the individual patient within a database server organized to store one or more patient care records. The collected device measures set is processed into a set of reference measures. Each reference measure is representative of at least one of measured or derived patient information. The reference measures set is stored into the patient care record as data in a reference baseline indicating an initial patient status. | 1. A system for determining a patient wellness status reference baseline, comprising: a medical device autonomously monitoring at least one of telemetered and derived physiological measures providing quantitative health care data indicators continuously and directly during an initial observation period;a database periodically storing and processing the monitored physiological measures;a reference baseline generated from the stored physiological measures upon completion of the initial observation period indicating an initial patient wellness status; andthe medical device obtaining the physiological measures provided at least one of in situ, exogenous, and proximate to a patient monitoring point. | 1. A system for determining a patient wellness status reference baseline, comprising: a medical device autonomously monitoring at least one of telemetered and derived physiological measures providing quantitative health care data indicators continuously and directly during an initial observation period;a database periodically storing and processing the monitored physiological measures;a reference baseline generated from the stored physiological measures upon completion of the initial observation period indicating an initial patient wellness status; andthe medical device obtaining the physiological measures provided at least one of in situ, exogenous, and proximate to a patient monitoring point. 2. A system according to claim 1, further comprising: a terminal monitoring quality of life measures providing qualitative health care data indicators; andthe database consolidating the monitored quality of life measures into the reference baseline. 3. A system according to claim 1, further comprising: the medical device providing at least one of electrical or drug interventions during the initial observation period. 4. A system according to claim 1, further comprising: the medical device being dynamically reprogrammed to provide modified at least one of electrical or drug interventions. 5. A system according to claim 1, further comprising: the database defining acceptance parameters comprising normalized physiological measures substantially conforming to a wellness status acceptance criteria and identifying substantial non-conformity by comparing each stored physiological measure to a corresponding acceptance parameter. 6. A method for determining a patient wellness status reference baseline, comprising: autonomously monitoring at least one of telemetered and derived physiological measures providing quantitative health care data indicators continuously and directly during an initial observation period;periodically storing and processing the monitored physiological measures into a database;generating a reference baseline from the stored physiological measures upon completion of the initial observation period indicating an initial patient wellness status; andobtaining the physiological measures through a medical device provided at least one of in situ, exogenous, and proximate to a patient monitoring point. 7. A method according to claim 6, further comprising: monitoring quality of life measures providing qualitative health care data indicators; andconsolidating the monitored quality of life measures into the reference baseline. 8. A method according to claim 6, further comprising: providing at least one of electrical or drug interventions through the medical device during the initial observation period. 9. A method according to claim 6, further comprising: dynamically reprogramming the medical device to provide modified at least one of electrical or drug interventions. 10. A method according to claim 6, further comprising: defining acceptance para... | Cardiac Intelligence Corporation | A61N001/00 A61B005/00
US07128718 | B2 | 2006-10-31 | Guidewire with deflectable tip | 2002-08-20 | A bi-directional steerable guidewire having a deflectable distal tip which comprises a longitudinal hypotube and a spring coil attached to the distal end of the hypotube and also includes a longitudinally movable deflection member which is attached to the distal end of the spring coil and a tip retaining member which extends from the distal end of the hypotube to the distal end of the spring coil for providing very precise deflection of the distal tip. | 1. A bi-directional steerable guidewire having a deflectable tip which comprises: an elongated flexible tubing having proximal and distal portions;a flexible helical coil having proximal and distal ends, the proximal end of said helical coil is attached to the distal portion of the flexible tubing;an elongated deflection member having proximal and distal portions and being slidably disposed within said tubing and within said helical coil, the distal portion of said deflection member being flattened to form a deflection ribbon which extends in a plane;a retaining ribbon having proximal and distal ends, the proximal end of the retaining ribbon is attached to the distal portion of the flexible tubing and the retaining ribbon is oriented to extend in a plane which is generally parallel to the plane of the deflection ribbon; and,an attachment member engaging the distal end of the helical coil, the distal portion of the deflection member and the distal end of the retaining ribbon so that longitudinal movement of the deflection member in a distal direction causes the distal end of the helical coil to be deflected in one direction and longitudinal movement of the deflection member in a proximal direction causes the distal end of the helical coil to deflect in another opposite direction, said attachment member takes the form of a rounded bead which contacts the distal end of the helical coil to define a circular surface at the distal end of the coil and the deflection ribbon engages the rounded bead at a location offset from the center of the circular surface of the rounded bead, said distal end of the retaining ribbon engages the rounded bead at a location offset from the center of the circular surface in an opposite direction from the offset location of the deflection ribbon. | 1. A bi-directional steerable guidewire having a deflectable tip which comprises: an elongated flexible tubing having proximal and distal portions;a flexible helical coil having proximal and distal ends, the proximal end of said helical coil is attached to the distal portion of the flexible tubing;an elongated deflection member having proximal and distal portions and being slidably disposed within said tubing and within said helical coil, the distal portion of said deflection member being flattened to form a deflection ribbon which extends in a plane;a retaining ribbon having proximal and distal ends, the proximal end of the retaining ribbon is attached to the distal portion of the flexible tubing and the retaining ribbon is oriented to extend in a plane which is generally parallel to the plane of the deflection ribbon; and,an attachment member engaging the distal end of the helical coil, the distal portion of the deflection member and the distal end of the retaining ribbon so that longitudinal movement of the deflection member in a distal direction causes the distal end of the helical coil to be deflected in one direction and longitudinal movement of the deflection member in a proximal direction causes the distal end of the helical coil to deflect in another opposite direction, said attachment member takes the form of a rounded bead which contacts the distal end of the helical coil to define a circular surface at the distal end of the coil and the deflection ribbon engages the rounded bead at a location offset from the center of the circular surface of the rounded bead, said distal end of the retaining ribbon engages the rounded bead at a location offset from the center of the circular surface in an opposite direction from the offset location of the deflection ribbon. 2. A steerable guidewire as defined in claim 1, wherein the deflection member and the retaining ribbon are joined to each other within the rounded bead. 3. A steerable guidewire as defined in claim 2, wherein the deflection ribbon and the retaining ribbon are formed as a single unitary element. 4. A steerable guidewire as defined in claim 3, wherein the deflection ribbon and the retaining ribbon are joined to form a generally U-shaped configuration to thereby provide a predetermined spacing between the deflection ribbon and the retaining ribbon and to maintain the deflection ribbon and the retaining ribbon in planes which are parallel to each other. 5. A steerable guidewire as defined in claim 4, wherein the deflection ribbon is formed by flattening an intermediate portion of the deflection member and the retaining ribbon is formed by flattening a distal portion of the deflection member. 6. A steerable guidewire as defined in claim 5, wherein the retaining ribbon is of a thickness which is less than the thickness of the deflection ribbon. 7. A steerable guidewire as defined in claim 6, wherein the deflection ribbon is of a thickness equal to about 0.002 inches and the retaini... | Cordis Corporation | A61B005/00 A61M025/00
US07294356 | B2 | 2007-11-13 | Performance enhancing coating on intraluminal devices | 2002-12-18 | This invention comprises guidewires, stents and other intraluminal devices having a performance enhancing coating deposited using physical vapor deposition (PVD) or chemical vapor deposition (CVD). Preferably, a radiopaque coating comprising platinum, tungsten, iridium, tantalum or the like is deposited on a desired portion of the device. Alternatively, the performance enhancing coating is a wear resistant coating of carbides such as tungsten carbide, titanium carbide, or nitrides such as titanium nitride. | 1. A method of providing an intraluminal device with a radiopaque coating, comprising: forming an intraluminal device;recognizing that the intraluminal device is to include a radiopaque coating for improved radiopacity; anddepositing a radiopaque coating having a thickness of not more than about 15 μm to at least a portion of the intraluminal device by applying the coating in a plurality of layers, each layer having a thickness of not more than 5 μm;wherein the coating is applied in the plurality of layers to prevent the intraluminal device from rising significantly in temperature during coating;wherein the plurality of layers are applied to prevent the intraluminal device from rising in temperature above 400° C. during coating;wherein the depositing of the radiopaque coating involves employing a voltage of 610±5 volts DC, a power of approximately 5 kW, a sputtering rate of approximately 2600 Å/min and a current of 8.2±0.1 Amp DC. | 1. A method of providing an intraluminal device with a radiopaque coating, comprising: forming an intraluminal device;recognizing that the intraluminal device is to include a radiopaque coating for improved radiopacity; anddepositing a radiopaque coating having a thickness of not more than about 15 μm to at least a portion of the intraluminal device by applying the coating in a plurality of layers, each layer having a thickness of not more than 5 μm;wherein the coating is applied in the plurality of layers to prevent the intraluminal device from rising significantly in temperature during coating;wherein the plurality of layers are applied to prevent the intraluminal device from rising in temperature above 400° C. during coating;wherein the depositing of the radiopaque coating involves employing a voltage of 610±5 volts DC, a power of approximately 5 kW, a sputtering rate of approximately 2600 Å/min and a current of 8.2±0.1 Amp DC. 2. The method of claim 1 wherein the portion of the intraluminal device comprises a relatively thin adhesion layer of the coating and a subsequent thicker layer of the coating. 3. The method of claim 1, wherein the intraluminal device is a stent. 4. The method of claim 1, wherein the intraluminal device is a guidewire. 5. The method of claim 1, wherein the radiopaque coating is deposited by physical vapor deposition. 6. The method of claim 1, wherein the coating is selected from the group consisting of platinum, tungsten, iridium, and tantalum. 7. A method for providing radiopacity to an intraluminal device, comprising: providing a guidewire;providing a flexible helical coil;disposing the flexible helical coil on the guidewire; andafter providing the flexible helical coil, applying a relatively high radiopaque coating along at least a portion of the coil, wherein the relatively high radiopaque coating includes a thickness of not more than about 15 μn, wherein the coating is applied in a plurality of layers, each layer having a thickness of not more than 5 μm;wherein the guide wire includes a core member and further comprising coating the core member with a radiopaque material;wherein the coating is applied in the plurality of layers to prevent the coil from rising significantly in temperature during coating;wherein the depositing of the radiopague coating involves employing a voltage of 610±5 volts DC, a power of approximately 5 kW, a sputtering rate of approximately 2600 Å/mm and a current of 8.2±0.1 Amp DC. 8. The method of claim 7, wherein the method includes coating the helical coil. 9. The method of claim 7, wherein the method includes coating the guidewire with a relatively uniform thickness and a smooth surface. 10. The method of claim 7, wherein the coating is selected from the group consisting of platinum, tungsten, iridium, and tantalum. 11. The method of claim 7, wherein less than an entirety of the coil is coated with radiopaque material. 12. The method of claim 7, wherein the plurality of layers ... | Advanced Cardiovascular Systems, Inc. | A61L033/00 B05D003/00 A61B005/00 A61M025/00
US07198606 | B2 | 2007-04-03 | Method and apparatus for a multi-use body fluid sampling device with analyte sensing | 2002-12-18 | A device for use with a penetrating member driver to penetrate tissue is provided. A plurality of penetrating members are coupled to a single cartridge and are operatively couplable to the penetrating member driver. The penetrating members are movable to extend radially outward from the cartridge to penetrate tissue. A plurality of analyte sensors are coupled to the single cartridge and are positioned on the cartridge to receive body fluid from a wound in the tissue created by the penetrating member. | 1. A device for use with a penetrating member driver to penetrate tissue, the device comprising: a single cartridge;a plurality of penetrating members coupled to said single cartridge and operatively couplable to a penetrating member driver, said penetrating members movable to extend radially outward from the cartridge to penetrate tissue;a plurality of analyte sensors each in a sample chamber with the sample chambers being coupled to said single cartridge, each of a sample chamber being positioned on the cartridge to receive body fluid from a wound in the tissue created by each of said penetrating members and having a volume of no greater than 1 microliter; anda position sensor coupled to the plurality of penetrating members, the position sensor configured to provide information relative to a depth of penetration of a penetrating member through a skin surface, the position sensor further configured to provide an indication of velocity of a penetrating member. | 1. A device for use with a penetrating member driver to penetrate tissue, the device comprising: a single cartridge;a plurality of penetrating members coupled to said single cartridge and operatively couplable to a penetrating member driver, said penetrating members movable to extend radially outward from the cartridge to penetrate tissue;a plurality of analyte sensors each in a sample chamber with the sample chambers being coupled to said single cartridge, each of a sample chamber being positioned on the cartridge to receive body fluid from a wound in the tissue created by each of said penetrating members and having a volume of no greater than 1 microliter; anda position sensor coupled to the plurality of penetrating members, the position sensor configured to provide information relative to a depth of penetration of a penetrating member through a skin surface, the position sensor further configured to provide an indication of velocity of a penetrating member. 2. The device of claim 1 wherein said penetrating members are slidably coupled to said cartridge. 3. The device of claim 1 wherein at least one analyte sensor is associated with at least one of said penetrating members. 4. The device of claim 1 wherein said cartridge is a flat radial disc. 5. The device of claim 1 wherein said cartridge has a diameter of less than 6 cm. 6. The device of claim 1 wherein said cartridge is a unitary body. 7. The device of claim 1 wherein said penetrating members are not attached by a resilient member to the cartridge. 8. The device of claim 1 wherein said analyte sensors are electrochemical sensors. 9. The device of claim 1 wherein said analyte sensors are potentiometric sensors. 10. The device of claim 1 wherein said sample chambers are configured to determine analyte levels using a body fluid sample of less than about 300 nanoliters. 11. The device of claim 1 wherein said analyte sensors are mounted on said cartridge. 12. The device of claim 1 wherein each of said analyte sensors comprises an array of sensors. 13. The device of claim 1 wherein each of said analyte sensors comprises an array of sensors wherein a plurality of said array sensors have different ranges of analyte sensitivity. 14. The device of claim 1 wherein each of said analyte sensors comprises an array of sensors formed from nanowires. 15. The device of claim 1 wherein each of said penetrating members are elongate members without molded attachments. 16. The device of claim 1 wherein each of said penetrating members are elongate wires of substantially constant diameter. 17. The device of claim 1 wherein each of said penetrating members is made of only one material selected from the following: a metal or a metal alloy. 18. A device for use with a penetrating member driver to penetrate tissue, the device comprising: a single cartridge having a plurality of openings;a plurality of penetrating members having sharpened tips movable to penetrate tissue;a plurality of analyte sensors... | Pelikan Technologies, Inc. | A61B005/00
US07274958 | B2 | 2007-09-25 | Registration pointer with interchangeable tip and method | 2002-10-04 | A registration pointer for registering points of surfaces of bodily elements in computer-assisted surgery, comprising a handle for being manipulated by a pointer operator. A stem is connected to the handle and has a first and a second tip configurations each having a respective use relating to a registration of points on specified bodily elements. A detectable device is secured to the stem, and is adapted to be tracked in space for position and orientation by a sensing apparatus. A position of the tip configurations with respect to the detectable device is determinable such that a position of the tip configurations is calculable as a function of the position and orientation of the detectable device. The tip configurations are chosen to register points on given bodily elements as a function of the respective use, by recording a position of points of contact between the tip configurations and the given bodily element. | 1. A registration pointer for registering points of surfaces of bodily elements in computer-assisted surgery, comprising: a handle for being manipulated by a pointer operator;a stem connected to the handle;a first tip configuration at a first tip of the stem;a second tip configuration being a removable hollow stem portion that selectively matingly covers the first tip and a portion of the stem, with each tip configuration having a respective use relating to a registration of points on specified bodily elements; anda detectable device secured to any one of the handle, the stem and the tip configurations, the detectable device being adapted to be tracked in space for position and orientation by a sensing apparatus;whereby a position of any one of the tip configurations with respect to the detectable device is determinable such that a position of any one of the tip configurations is calculable as a function of the position and orientation of the detectable device, any one of the tip configurations being chosen to register points on given bodily elements as a function of the respective use, by recording a position of points of contact between any one of the tip configurations and the given bodily element. | 1. A registration pointer for registering points of surfaces of bodily elements in computer-assisted surgery, comprising: a handle for being manipulated by a pointer operator;a stem connected to the handle;a first tip configuration at a first tip of the stem;a second tip configuration being a removable hollow stem portion that selectively matingly covers the first tip and a portion of the stem, with each tip configuration having a respective use relating to a registration of points on specified bodily elements; anda detectable device secured to any one of the handle, the stem and the tip configurations, the detectable device being adapted to be tracked in space for position and orientation by a sensing apparatus;whereby a position of any one of the tip configurations with respect to the detectable device is determinable such that a position of any one of the tip configurations is calculable as a function of the position and orientation of the detectable device, any one of the tip configurations being chosen to register points on given bodily elements as a function of the respective use, by recording a position of points of contact between any one of the tip configurations and the given bodily element. 2. The registration pointer according to claim 1, wherein the hollow stem portion has a tapping portion, and the stem has a corresponding threading portion for detachable threading engagement with the tapping portion. 3. The registration pointer according to claim 1, wherein the stem has an intermediate portion of circular cross-section with a given diameter connected to the handle, and an end portion having the first tip, the end portion having a circular cross-section with a smaller diameter than said given diameter. 4. The registration pointer according to claim 3, wherein the hollow stem portion has an annular cross-section with an outer diameter equal to the given diameter such that stem and the hollow stem portion form a continuous smooth surface when secured to one another. 5. The registration pointer according to claim 1, wherein the first tip configuration has a blunt tip and the second tip configuration has a pointed tip, the respective use of the first tip configuration being to register points on a bone element following displacement of the blunt tip through soft tissue to reach the bone element, so as to substantially limit damage to the soft tissue, and the respective use of the second tip configuration being to register points on thin soft tissue covering bone element. 6. The registration pointer according to claim 1, wherein the detectable device is passive. 7. The registration pointer according to claim 6, wherein the detectable device has three retro-reflective spheres in a determined pattern. 8. The registration pointer according to claim 1, the number of the at least one stem is two, the first tip configuration being at an end of a first one of the stems, and the second tip configuration being at an end of a second one of ... | Orthosoft Inc. | A61B005/05
US07228159 | B2 | 2007-06-05 | Optical sensor containing particles for in situ measurement of analytes | 2002-06-27 | The invention relates to a sensor for the in vivo measurement of an analyte, comprising a plurality of particles of suitable size such that when implanted in the body of a mammal the particles can be ingested by macrophages and transported away from the site of implantation, each particle containing the components of an assay having a readout which is an optical signal detectable transdermally by external optical means, and either each particles being contained within a biodegradable material preventing ingestion by the macrophages, or each particle being non-biodegradable. The invention relates to a process for the detection of an analyte using such a sensor, comprising implantation of the sensor into the skin of a mammal, transdermal detection of analyte using external optical means, degradation of the biodegradable material, ingestion of the particles by macrophages, and removal of the particles from the site of implantation by macrophages. | 1. A sensor for in vivo measurement of an analyte, comprising a plurality of particles having a size such that when implanted in a body of a mammal at an implantation site the particles can be ingested by macrophages and transported away from the implantation site, wherein each particle of the sensor contains the components of an assay having a readout which is an optical signal detectable transdermally by external optical means, and wherein the particles are each contained within a biodegradable material preventing ingestion by the macrophages. | 1. A sensor for in vivo measurement of an analyte, comprising a plurality of particles having a size such that when implanted in a body of a mammal at an implantation site the particles can be ingested by macrophages and transported away from the implantation site, wherein each particle of the sensor contains the components of an assay having a readout which is an optical signal detectable transdermally by external optical means, and wherein the particles are each contained within a biodegradable material preventing ingestion by the macrophages. 2. A sensor as claimed in claim 1, wherein the particles are less than 5 μm in their largest dimension. 3. A sensor as claimed in claim 1, wherein the assay is a binding assay. 4. A sensor as claimed in claim 3, wherein the binding assay is a competitive binding assay, the components of which include an analyte binding agent and an analyte analogue. 5. A sensor as claimed in claim 4, wherein the analyte analogue is labelled with a first chromophore and the analyte binding agent is labelled with a second chromophore, an emission spectrum of the first chromophore or the second chromophore overlapping with an absorption spectrum of the second chromophore or the first chromophore respectively. 6. A sensor as claimed in claim 4, wherein the binding agent is an antibody, a Fab fragment, a lectin, a hormone receptor, a drug receptor, an aptamer or a molecularly-imprinted polymer. 7. A sensor as claimed in claim 1, wherein the optical signal is generated by fluorescence resonance energy transfer, fluorescence polarisation, fluorescence quenching, phosphorescence, luminescence enhancement, luminescence quenching, diffraction or plasmon resonance. 8. A sensor as claimed in claim 1, wherein the analyte is glucose. 9. A sensor as claimed in claim 1, wherein the optical signal is quantitatively measurable. 10. A process for the detection of an analyte using a sensor as claimed in claim 1, comprising the steps of: implanting the sensor into skin of a mammal at an implantation site, andtransdermally detecting analyte using external optical means,where implantation is followed by degradation of the biodegradable material, ingestion of the particles by macrophages, and removal of the particles from the implantation site. 11. A sensor for in vivo measurement of an analyte, comprising a plurality of particles having a size such that when implanted in a body of a mammal at an implantation site the particles can be ingested by macrophages and transported away from the implantation site, wherein each particle of the sensor contains the components of an assay having a readout which is an optical signal detectable transdermally by external optical means, and wherein the particles are each contained within a biodegradable material preventing ingestion by the macrophages, the biodegradable material being a gel. | Precisense A/S | A61B005/00
US06898453 | B2 | 2005-05-24 | Method and apparatus for calculating blood flow parameters | 2001-10-25 | A method for determining tissue type includes quantitatively determining a tissue blood flow (TBF) by deconvoluting Q(t) and Ca(t), where Q(t) represents a curve of specific mass of contrast, and Ca(t) represents an arterial curve of contrast concentration, and quantitatively determining a tissue blood volume (TBV) by deconvoluting Q(t) and Ca(t). The method also includes quantitatively determining a tissue mean transit time (TMTT) by deconvoluting Q(t) and Ca(t), and quantitatively determining a tissue capillary permeability surface area product (TPS) by deconvoluting Q(t) and Ca(t). The method also includes determining a tissue type based on the TBF, the TBV, the TMTT, and the TPS. | 1. A method for determining tissue type, said method comprising: quantitatively determining a tissue blood flow (TBF) by deconvoluting Q(t) and Ca(t), where Q(t) is a tissue residue function and represents a curve of specific mass of contrast in tissue, and Ca(t) represents an arterial curve of contrast concentration for a tissue having a blood stream containing a contrast without leaking the contrast into an interstitial space of the tissue by solving a matrix equation of Q=Ah for a vector h, and determining a least squares solution for the vector h under an equality constraint, wherein vector h includes a plurality of elements comprising an impulse residue function at different times Q comprises a vector including elements comprising values of a tissue residue function at different times, A comprises a matrix formed by values of the arterial curve of contrast concentration at different times; quantitatively determining a tissue blood volume (TBV) by deconvoluting Q(t) and Ca(t), quantitatively determining a TBV comprises quantitatively determining a TBV for the tissue by solving the matrix equation of Q=Ah for the vector h, and determining a least squares solution for the vector h under an equality constraint; quantitatively determining a tissue mean transit time (TMTT) by deconvoluting Q(t) and Ca(t), quantitatively determining a TMTT comprises quantitatively determining a TMTT for the tissue by solving the matrix equation of Q=Ah for the vector h, and determining a least squares solution for the vector h under an equality constraint; and determining a tissue type based on the TBF, the TBV, and the TMTT. | 1. A method for determining tissue type, said method comprising: quantitatively determining a tissue blood flow (TBF) by deconvoluting Q(t) and Ca(t), where Q(t) is a tissue residue function and represents a curve of specific mass of contrast in tissue, and Ca(t) represents an arterial curve of contrast concentration for a tissue having a blood stream containing a contrast without leaking the contrast into an interstitial space of the tissue by solving a matrix equation of Q=Ah for a vector h, and determining a least squares solution for the vector h under an equality constraint, wherein vector h includes a plurality of elements comprising an impulse residue function at different times Q comprises a vector including elements comprising values of a tissue residue function at different times, A comprises a matrix formed by values of the arterial curve of contrast concentration at different times; quantitatively determining a tissue blood volume (TBV) by deconvoluting Q(t) and Ca(t), quantitatively determining a TBV comprises quantitatively determining a TBV for the tissue by solving the matrix equation of Q=Ah for the vector h, and determining a least squares solution for the vector h under an equality constraint; quantitatively determining a tissue mean transit time (TMTT) by deconvoluting Q(t) and Ca(t), quantitatively determining a TMTT comprises quantitatively determining a TMTT for the tissue by solving the matrix equation of Q=Ah for the vector h, and determining a least squares solution for the vector h under an equality constraint; and determining a tissue type based on the TBF, the TBV, and the TMTT. 2. A method in accordance with claim 1 wherein determining a least squares solution comprises determining a least squares solution for the vector h under a time causality constraint and a minimum transit time constraint. 3. A method in accordance with claim 2 wherein determining a least squares solution further comprises determining a least squares solution for the vector h under a smoothness constraint, a nonnegativity constraint, and a monotonicity constraint, wherein the smoothness constraint forces h to be smoothly varying, and the monotonicity and nonnegativity constraints force h to start at a maximum and then monotonically decreases towards a zero baseline. 4. A method in accordance with claim 1 further comprising quantitatively determining a partial volume averaging scaling factor for the arterial curve of contrast concentration by: deconvoluting the measured arterial curve of contrast concentration with a venous curve of contrast concentration to determine a transit time spectrum through a tissue of interest; extrapolating the arterial curve; and convolving the extrapolated arterial curve with the transit time spectrum to generate an extrapolated venous curve, wherein the partial volume averaging scaling factor is the ratio of an area underneath the extrapolated arterial curve to an area underneath the extrapolated venous curve. ... | The John P. Robarts Research Institute | A61B005/00
US06847836 | B1 | 2005-01-25 | Emergency ECG electrode chest pad | 2002-02-08 | An ECG electrode chest pad particularly adapted for use in emergency room situations having upper fit portions with upper limb electrodes, and elongated central or medial base fit portion with a plurality of precordal unipolar electrodes and lower fit portions with lower limb electrodes, said electrodes being attached to leads which are internal to the base chest pad and terminate into at least one lead branch adapted to plug into an ECG monitor and having a perforation in the base pad material such that one group of electrodes may be separated from a second group of electrodes to facilitate ease of patient monitoring and complimentary medical procedures. | 1. An electrocardiogram diagnostic chest pad comprising: a base pad comprised of a lightweight material and having a central body portion, a central fit portion extending from the central body portion, a plurality of upper fit portions, and a plurality of lower fit portions; a first set of a plurality of electrodes attached to said base pad, each of said first set of plurality of electrodes being connected to a lead capable of conducting electric signals and each said lead embedded with the base pad material and terminating at a first lead bundle having a plug adapted input to an ECG monitor; a second set of a plurality of electrodes attached to said base pad, each of said second set of plurality of electrodes being connected to a lead capable of conducting electric signals and each said lead passing through the base pad material and terminating at a second lead bundle having a plug adapted input to an ECG monitor; and a perforation in the material of the base pad between the first set of plurality of electrodes and the second set of plurality of electrodes, whereby the chest pad can be used as a whole or separated into two distinct sections for varying ECG measurement and monitoring functions. | 1. An electrocardiogram diagnostic chest pad comprising: a base pad comprised of a lightweight material and having a central body portion, a central fit portion extending from the central body portion, a plurality of upper fit portions, and a plurality of lower fit portions; a first set of a plurality of electrodes attached to said base pad, each of said first set of plurality of electrodes being connected to a lead capable of conducting electric signals and each said lead embedded with the base pad material and terminating at a first lead bundle having a plug adapted input to an ECG monitor; a second set of a plurality of electrodes attached to said base pad, each of said second set of plurality of electrodes being connected to a lead capable of conducting electric signals and each said lead passing through the base pad material and terminating at a second lead bundle having a plug adapted input to an ECG monitor; and a perforation in the material of the base pad between the first set of plurality of electrodes and the second set of plurality of electrodes, whereby the chest pad can be used as a whole or separated into two distinct sections for varying ECG measurement and monitoring functions. 2. The electrocardiogram diagnostic chest pad of claim 1 further comprising a first and second upper fit portion and a first and second lower fit portion. 3. The electrocardiogram diagnostic chest pad of claim 2 wherein said first set of plurality of electrodes comprises a first electrode on said first upper fit portion, a second electrode on said second upper fit portion, a third electrode on said central body portion, a fourth electrode on said first lower fit portion and a fifth electrode on said second lower fit portion. 4. The electrocardiogram diagnostic chest pad of claim 2 wherein said second set of plurality of electrodes comprises first and second electrodes on said central body portion, and third, fourth and fifth electrodes on said central fit portion. 5. The electrocardiogram diagnostic chest pad of claim 1 wherein each of the first and second sets of electrodes further comprise a solid metal stud embedded within a silver/silver chloride conductor gel. 6. The electrocardiogram diagnostic chest pad of claim 1 further comprising a plurality of pacer pads, each of said plurality of pacer pads being connected to a lead embedded within the material of the base pad and terminating in a plug bundle adapted for connection to a pacer control instrument. 7. The electrocardiogram diagnostic chest pad of claim 1 wherein the base pad further comprises an external side and a patient contact side and said patient contact side further comprises a plurality of adhesive portions, each containing a removable backing, whereby a health care professional can remove the removable backing, expose each adhesive portion and adhere the chest pad to a patient. 8. The electrocardiogram diagnostic chest pad of claim 2 wherein the base pad is radiologically transpare... | - | A61B005/0408
US06970738 | B1 | 2005-11-29 | Complex impedance spectrometer using parallel demodulation and digital conversion | 2003-01-30 | A spectrometer for measuring the complex impedance spectrum of tissue comprises: a multi-frequency excitation current generator; a demodulation signal generator; two identical amplification/demodulation circuits; an A/D converter; and a microprocessor for signal processing. In use, the current generator excites the tissue sample and a series-connected reference impedance. The voltages generated in the tissue and reference are measured, demodulated, and digitized in parallel using the demodulation signal generator, the two amplification/demodulation circuits, and the A/D converter. Demodulation is done using the same demodulation signal generated at a frequency with a preset difference from the excitation signal, which allows measurements to be made at a low frequency independent of the excitation frequency. The microprocessor then calculates the complex impedance spectrum in relation to the reference signal. Because the measurements are relative, they are independent of both the excitation current amplitude of the phase of the excitation and demodulation signals. | 1. A system for measuring the complex impedance spectrum of a sample, said system comprising: a. an excitation circuit comprising: a sample connection mechanism; a reference impedance electrically connected in series to the sample connection mechanism; and a multi-frequency excitation device electrically connected in series to the sample connection mechanism and reference impedance;b. first and second amplification/demodulation circuits respectively connected in parallel to the sample connection mechanism and reference impedance;c. a demodulation signal generator electrically connected to the first and second amplification/demodulation circuits; andd. a signal processor electrically connected to the first and second amplification/demodulation circuits. | 1. A system for measuring the complex impedance spectrum of a sample, said system comprising: a. an excitation circuit comprising: a sample connection mechanism; a reference impedance electrically connected in series to the sample connection mechanism; and a multi-frequency excitation device electrically connected in series to the sample connection mechanism and reference impedance;b. first and second amplification/demodulation circuits respectively connected in parallel to the sample connection mechanism and reference impedance;c. a demodulation signal generator electrically connected to the first and second amplification/demodulation circuits; andd. a signal processor electrically connected to the first and second amplification/demodulation circuits. 2. The system of claim 1 wherein: a. the first amplification/demodulation circuit comprises: a first amplifier having an input electrically connected in parallel to the sample connection mechanism; a first multiplier electrically connected to an output of the first amplifier and to the demodulation signal generator; and a first bandpass filter electrically connected to the first multiplier and to the signal processor; andb. the second amplification/demodulation circuit comprises: a second amplifier having an input electrically connected in parallel to the reference impedance; a second multiplier electrically connected to an output of the second amplifier and to the demodulation signal generator; and a second bandpass filter electrically connected to the second multiplier and to the signal processor. 3. The system of claim 1 wherein: a. the excitation device is configured to generate an excitation signal at a first frequency; andb. the demodulation signal generator is configured to generate a demodulation signal at a second frequency different from the first frequency, wherein the difference between the first frequency and the second frequency is substantially constant. 4. The system of claim 3 wherein each amplification/demodulation circuit has a bandpass filter set at a third frequency about the same as the difference between the first frequency and the second frequency. 5. The system of claim 3 wherein: a. the first amplification/demodulation circuit comprises: a first amplifier having an input electrically connected in parallel to the sample connection mechanism; a first multiplier electrically connected to an output of the first amplifier and to the demodulation signal generator; and a first bandpass filter electrically connected to the first multiplier and to the signal processor and set at a third frequency about the same as the difference between the first frequency and the second frequency; andb. the second amplification/demodulation circuit comprises: a second amplifier having an input electrically connected in parallel to the reference impedance; a second multiplier electrically connected to an output of the second amplifier and to the demodulation signal generator; and a second bandp... | Innovamedica S.A. de C.V. | A61B005/05
US07107091 | B2 | 2006-09-12 | Multiple bone tracking | 2002-07-25 | A surgical device adapted for use with an image guided surgical system to facilitate monitoring inter-dependently mobile bone elements. A rigidly connected multiple-point anchor assembly is releasably engageable to at least two inter-dependently mobile bone elements and immobilizes the bone elements such that no relative movement therebetween is possible, and a complex rigid body comprising the immobilized bone elements and the multiple-point anchor assembly is formed. A trackable member is fastened to the multiple-point anchor assembly and comprises a detectable element adapted to be located and tracked in three dimensional space by the image guided surgical system, thereby defining position and movement of the complex rigid body. | 1. A surgical device adapted for use with computer assisted surgical system to facilitate monitoring inter-dependently mobile bone elements, comprising: a rigidly connected multiple-point anchor assembly releasably engageable to at least two inter-dependently mobile bone elements, immobilizing the bone elements such that no relative movement therebetween is possible, and thereby forming a complex rigid body comprising the immobilized bone elements and the multiple-point anchor assembly; and a trackable member, fastened to the multiple-point anchor assembly and comprising a detectable element adapted to be located and tracked in three dimensional space by the computer assisted surgical system, thereby defining position and movement of the complex rigid body. | 1. A surgical device adapted for use with computer assisted surgical system to facilitate monitoring inter-dependently mobile bone elements, comprising: a rigidly connected multiple-point anchor assembly releasably engageable to at least two inter-dependently mobile bone elements, immobilizing the bone elements such that no relative movement therebetween is possible, and thereby forming a complex rigid body comprising the immobilized bone elements and the multiple-point anchor assembly; and a trackable member, fastened to the multiple-point anchor assembly and comprising a detectable element adapted to be located and tracked in three dimensional space by the computer assisted surgical system, thereby defining position and movement of the complex rigid body. 2. The device as defined in claim 1, wherein the multiple-point anchor assembly provides one of a compressive and a tensile force between the immobilized bone elements engaged therewith. 3. The device as defined in claim 1, wherein the multiple-point anchor assembly comprises at least two discrete anchoring members. 4. The device as defined in claim 3, wherein the anchoring members are inter-fastenable and adjustable substantially longitudinally relative to the bone elements. 5. The device as defined in claim 3, wherein at least one of the anchoring members comprises a clamp adapted for fixation to one of the bone elements. 6. The device as defined in claim 5, wherein the clamp comprises opposing jaws operable to close together for releasable attachment with the bone element. 7. The device as defined in claim 4, wherein the anchoring members, having bores therethrough and being arranged such that the bores substantially correspond, are inter-fastened by a linking bar extending through the bores to which each anchoring member is fastenable. 8. The device as defined in claim 7, wherein the linking bar has a pre-determined shape adapted to conform to the shape of a longitudinal axis of the bone elements. 9. The device as defined in claim 1, wherein the bone elements comprise vertebrae. 10. The device as defined in claim 9, wherein the vertebrae include spinous processes, and the discrete anchoring members are engageable thereto such that pedicles of the vertebrae remain accessible. 11. The device as defined in claim 3, wherein at least one of the anchoring members comprises a sacrum press member. 12. The device as defined in claim 11, wherein the sacrum press member is longitudinally extendable relative to the spinal column. 13. The device as defined in claim 12, wherein the extending sacrum press member comprises a pinion gear driven rack providing substantially rectilinear displacement thereof. 14. The device as defined in claim 9, wherein the anchoring members comprise three vertebral clamps, each adapted for clamping to a vertebra and being inter-fastened such that the three vertebrae are rigidly clamped together, and thereby trackable as a single unit by the computer assisted s... | Orthosoft Inc. | A61B005/05 A61B019/00
US07004928 | B2 | 2006-02-28 | Autonomous, ambulatory analyte monitor or drug delivery device | 2002-04-23 | The invention relates to analyte monitoring/drug (pharmaceutical agent) delivery device. The invention is suited for monitoring various blood constituents such as glucose. The device has a housing that at least partially encloses a plurality of microneedles disposed on a carrier and an electronics portion. Each microneedle is in fluid communication with a corresponding microchannel. Each microneedle is individually addressable. That is, each microneedle can be extended and retracted individually via an actuator. The electronics portion includes a processor and associated circuitry (e.g., memory, supporting electronics and the like), a motor or the like, a sensor, a power supply (e.g., battery) and optionally an interface. In general, the processor controls the operation of the device and is data communication with the actuator, motor, sensor and interface. The invention provides for autonomous operation, that is, without intervention of the user. The invention can optionally provide for calibration without intervention of the user. The invention can also provide for semi-continuous monitoring for day and night time. The invention can provide for up to four, or more, weeks of operation. The invention can provide for a device that is relative small in size, and therefore unobtrusive. The invention can also provide for device with remote control and interactive electronics. The invention may be also used for the delivery of various pharmaceutical agents including high potency drugs to minimize patient intervention and minimize discomfort. | 1. An analyte monitoring device operable to draw a fluid sample from a subject, the device comprising: a first plurality of microneedles,a plurality of monitoring microchannels, wherein each of the first plurality of microneedles is at least intermittently in fluid communication with a corresponding monitoring microchannel, andwherein each monitoring microchannel is associated with a reagent,at least one detector associated with the monitoring microchannels, the detector comprising a CMOS imager;at least one actuator operable to extend each microneedle to draw the fluid sample from the subject, anda controller operable to initiate analyte testing of the fluid sample. | 1. An analyte monitoring device operable to draw a fluid sample from a subject, the device comprising: a first plurality of microneedles,a plurality of monitoring microchannels, wherein each of the first plurality of microneedles is at least intermittently in fluid communication with a corresponding monitoring microchannel, andwherein each monitoring microchannel is associated with a reagent,at least one detector associated with the monitoring microchannels, the detector comprising a CMOS imager;at least one actuator operable to extend each microneedle to draw the fluid sample from the subject, anda controller operable to initiate analyte testing of the fluid sample. 2. The device of claim 1 wherein the actuator is operable to extend at least one of the first plurality of microneedles omnidirectionally. 3. The device of claim 1 wherein each of the first plurality of microneedles has an internal diameter in the range of about 25 to 200 micrometers. 4. The device of claim 1 wherein each of the first plurality of microneedles is fabricated from at least one of metal, plastic, glass and crystal. 5. The device of claim 1 wherein each of the first plurality of microneedles has a distal end that is operable to penetrate into a skin surface to a maximum of about 2.5 mm. 6. The device of claim 1 wherein the fluid sample is substantially blood. 7. The device of claim 1 wherein each of the first plurality of microneedles is at least intermittently in fluid communication with a monitoring microchannel via a conduit. 8. The device of claim 1 wherein each monitoring microchannel is operable to store the fluid sample and accumulation of the fluid sample in the monitoring microchannel is entirely dependent on capillary forces. 9. The device of claim 1 wherein each of the monitoring microchannels has at least one internal surface that is at least partially coated with at least one insoluble material to enhance the capillary forces and minimize coagulation. 10. The device of claim 1 wherein the plurality of monitoring microchannels are fabricated in an array. 11. The device of claim 1 wherein the monitoring microchannels are fabricated in an array having approximately 50–150 microchannels formed in a maximum diameter of approximately 5 cm. 12. The device of claim 1 wherein each of the first plurality of microneedles is dimensioned for volumes of fluid in the range of about 50–500 nanoliters. 13. The device of claim 1 comprising a detector operable to determine when the fluid sample completely fills at least one monitoring microchannel such that the accumulation of the fluid sample with the associated microneedle may be terminated. 14. The device of claim 1 wherein at least one monitoring microchannel is in fluid communication with at least one reagent operable to assay for analytes selected from the group of (a) glucose, (b) cholesterol, (c) ethanol, (d) digoxin, (e) HDL cholesterol, (f) lithium, (g) sodium, (h) phenytoin, (i) therophylline, (j) cyclo... | Rosedale Medical, Inc. | A61M005/00 A61M005/315 A61B005/00 B65D081/00
US07226426 | B2 | 2007-06-05 | Apparatus and method for the detection and quantification of joint and tissue inflammation | 2002-07-25 | The apparatus and method for the detection and quantification of joint and tissue inflammation (100) of the present invention comprises a sensing component (102) for obtaining and collecting data indicative of the surface or dimensions of the joint or tissue. The sensing component (102) comprises various components for obtaining measurements of the cardinal signs of inflammation. Preferably the sensing component (102) includes a device for detecting swelling (104), a color analyzing device (108), a temperature measuring device (110), a pain or tenderness detection device (112), an archival and retrieval device (114), and a display device (116). In operation a patient is placed in position for examination of the joint or tissue area measurements are taken of one or more of the cardinal signs of inflammation at the joint or tissue area, the measurements are automatically stored in a archival storage and retrieval device (114), and from the measurements an inflammation score is derived which can be compared against a baseline. | 1. An apparatus for detecting and quantifying inflammation of a joint or tissue area comprising a sensing component; wherein said sensing component includes means for directly obtaining measurements of at least two of the cardinal signs of inflammation. | 1. An apparatus for detecting and quantifying inflammation of a joint or tissue area comprising a sensing component; wherein said sensing component includes means for directly obtaining measurements of at least two of the cardinal signs of inflammation. 2. The Apparatus of claim 1 further comprising a device for storing the measurements and a display device for displaying the measurements. 3. The Apparatus of claim 1 further comprising means for comparing the measurements with previous measurements of the joint or tissue area. 4. The Apparatus of claim 1 wherein said sensing component comprises means for detecting swelling of the joint or tissue area. 5. The Apparatus of claim 4 wherein said means for detecting swelling is selected from the group consisting of mechanical contact devices, video-based digitizing scanning systems, optical scanning systems, laser-based systems, position tracking based system devices, glossmeter-based devices, ultrasound devices, and magnetic resonance imaging systems. 6. The Apparatus of claim 1 wherein said sensing component comprises a range-of-motion device determining the range of motion of the joint or tissue. 7. The Apparatus of claim 6 wherein said range-of-motion device is selected from the group consisting of geometric chart devices, geometric gage devices, optical grid devices, video-based digitizing scanning system devices, optical scanning system devices, laser-based system devices, position tracking based system devices, ultrasound-based devices, and magnetic resonance imaging and measuring system devices. 8. The Apparatus of claim 1 where said sensing component comprises a color analyzing device for analyzing the color of the area of the joint or tissue area. 9. The Apparatus of claim 8 wherein said color analyzing device is a light-sensitive sensor device. 10. The Apparatus of claim 8 wherein said color analyzing device is selected from the group consisting of colorimeters, spectrophotometers, and glossmeters. 11. The Apparatus of claim 1 wherein said sensing component comprises a temperature measuring device measuring the temperature of the joint or tissue area. 12. The Apparatus of claim 11 wherein said temperature measuring device is selected from the group consisting of mechanical contact devices, optical, laser-based, thermistor type, thermocouple-type, thermometer devices, thermography-type devices, infrared-based devices, light sensitive devices, and surface electrical conductance-resistance based devices. 13. The Apparatus of claim 1 wherein said sensing component comprises a pain or tenderness detection device determining the threshold and tolerance of pain or tenderness at the joint or tissue area. 14. The Apparatus of claim 13 wherein said pain or tenderness detection device Is selected from the group consisting of dolorimeter-type pressure devices, heating devices, and electrical stimuli devices. 15. The Apparatus of claim 1 further comprising means for generating an image of ... | - | A61B005/117 A61B005/103
US07115095 | B2 | 2006-10-03 | Systems and methods for detecting, diagnosing and treating congestive heart failure | 2002-04-19 | The invention provides improved apparatus and methods for treating congestive heart failure in a medical patient. The apparatus includes a pressure transducer permanently implantable within the left atrium of the patient's heart and operable to generate electrical signals indicative of fluid pressures within the patient's left atrium. The pressure transducer is connected to a flexible electrical lead, which is connected in turn to electrical circuitry, which in the preferred embodiment includes digital circuitry for processing electrical signals. The electrical circuitry processes the electrical signals from the pressure transducer and, based at least in part on those signals, generates a signal that indicates a desired therapeutic treatment for treating the patient's condition. That signal is then communicated to the patient via a patient signaling device, following which the patient administers to him or herself the prescribed therapeutic treatment indicated by the signal. | 1. A method for treating congestive heart failure in a medical patient, the method comprising: implanting a pressure transducer within the medical patient;operating the pressure transducer to generate a signal indicative of fluid pressure within the left atrium;communicating the signal indicative of the fluid pressure within the left atrium to signal processing apparatus;operating the signal processing apparatus to generate an instructive signal indicative of an appropriate therapeutic treatment, whereinoperating the signal processing apparatus comprises generating one of at least two distinct instructive signals, each said signal indicative of a different therapeutic treatment; andcommunicating the instructive signal indicative of the appropriate therapeutic treatment to the patient. | 1. A method for treating congestive heart failure in a medical patient, the method comprising: implanting a pressure transducer within the medical patient;operating the pressure transducer to generate a signal indicative of fluid pressure within the left atrium;communicating the signal indicative of the fluid pressure within the left atrium to signal processing apparatus;operating the signal processing apparatus to generate an instructive signal indicative of an appropriate therapeutic treatment, whereinoperating the signal processing apparatus comprises generating one of at least two distinct instructive signals, each said signal indicative of a different therapeutic treatment; andcommunicating the instructive signal indicative of the appropriate therapeutic treatment to the patient. 2. The method of claim 1, further comprising implanting at least a portion of the signal processing apparatus substantially permanently within the patient's body. 3. The method of claim 1, further comprising anchoring the pressure transducer to the patient's atrial septum. 4. The method of claim 2, wherein implanting the signal processing apparatus within the patient's body comprises implanting the signal processing apparatus in a subcutaneous pocket in the region of the patient's shoulder. 5. The method of claim 3, wherein anchoring the pressure transducer to the patient's atrial septum comprises advancing a helical screw into the patient's atrial septum. 6. The method of claim 3, wherein anchoring the pressure transducer to the patient's atrial septum comprises deploying at least one anchor by spring force to bear against the patient's atrial septum. 7. The method of claim 1, further comprising implanting the signal processing apparatus within the patient's body. 8. The method of claim 1, wherein communicating the instructive signal indicative of the appropriate therapeutic treatment comprises generating an electrical shock perceptible by the patient. 9. The method of claim 1, wherein generating one of at least two distinct signals, each said signal indicative of a different therapeutic treatment, camprises generating one of at least four distinct signals. 10. The method of claim 1, wherein the transducer is situated within the left atrium. 11. The method of claim 1, wherein the transducer is situated within a pulmonary vein. 12. The method of claim 1, wherein the transducer is situated within the left atrial appendage. 13. The method of claim 1, wherein the transducer is located in the septal wall. 14. The method of claim 1, wherein the transducer is situated at a location within the heart having a pressure that is at least nearly the same as the pressure in the left atrium. 15. The method of claim 1, wherein the transducer is situated within the heart. 16. The method of claim 1, wherein the transducer is situated outside of the heart. 17. The method of claim 1, wherein the transducer is associated with the heart in a manner that is sufficient to ... | Cedars-Sinai Medical Center | A61B005/02
US07297122 | B2 | 2007-11-20 | Method and apparatus for penetrating tissue | 2002-12-31 | These and other objects of the present invention are achieved in a body fluid sampling system for use on a tissue site that includes a single drive force generator. A plurality of penetrating members are operatively coupled to the force generator. The force generator moves each of the members along a path out of a housing with a penetrating member exit, into the tissue site, stops in the tissue site, and withdraws out of the tissue site. A flexible support member couples the penetrating members to define a linear array. The support member is movable and configured to move each of the penetrating members to a launch position associated with the force generator. A user interface is configured to relay at least one of, penetrating member performance or a penetrating member setting. | 1. A body fluid sampling system for use on a tissue site, the system comprising: a single drive force generator;a plurality of penetrating members operatively coupled to said force generator, said force generator moving each of said penetrating members along a path out of a housing through a penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing each penetrating member out of said tissue site back into said housing;a plurality of analyte sensors, each of an analyte sensor of the plurality of analyte sensors being associated with a penetrating member;a flexible support member coupling said penetrating members to define a linear array, said support member being movable and configured to move each of said penetrating members to a launch position associated with said force generator; anda user interface coupled to a processor, wherein in response to an input at the user interface by a user the processor provides an input to the drive force generator of a lancing penetration parameter selected from at least one of, penetrating member depth of penetration, velocity, braking, or retraction. | 1. A body fluid sampling system for use on a tissue site, the system comprising: a single drive force generator;a plurality of penetrating members operatively coupled to said force generator, said force generator moving each of said penetrating members along a path out of a housing through a penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing each penetrating member out of said tissue site back into said housing;a plurality of analyte sensors, each of an analyte sensor of the plurality of analyte sensors being associated with a penetrating member;a flexible support member coupling said penetrating members to define a linear array, said support member being movable and configured to move each of said penetrating members to a launch position associated with said force generator; anda user interface coupled to a processor, wherein in response to an input at the user interface by a user the processor provides an input to the drive force generator of a lancing penetration parameter selected from at least one of, penetrating member depth of penetration, velocity, braking, or retraction. 2. The system of claim 1, wherein the user interface is configured to provide a user with at least one input selected from, depth of a penetrating member penetration, velocity of a penetrating member, a desired velocity profile, a velocity of a penetrating member into the target tissue, velocity of the penetrating member out of the target tissue, dwell time of the penetrating member in the target tissue, and a target tissue relaxation parameter. 3. The system of claim 1, wherein the user interface provides t least one output to the user selected from, number of penetrating members available, number of penetrating members used, actual depth of penetrating member penetration on a target tissue, stratum corneum thickness, force delivered on a target tissue, energy used by a penetrating member driver to drive a penetrating member into the target tissue, dwell time of the penetrating member, battery status, system status, consumed energy, speed profile of a penetrating member, information relative to contact of a penetrating member with target tissue before penetration by the penetrating member, and information relative to a change of speed of a penetrating member as in travels in the target tissue. 4. The system of claim 1, further comprising: a data interface configured to couple the tissue penetrating system to at least one of, support equipment with a data interface and the internet. 5. The system of claim 4, wherein the data interface is selected from at least one of, Serial RS-232, modem interface, USB, HPNA, Ethernet, optical interface, IRDA, RF interface, BLUETOOTH interface, cellular telephone interface, 2 way pager interface, parallel port interface standard, near field magnetic coupling, RF transceiver and a telephone systems. 6. The system of claim 1, wherein the support equipment is selected from at least one of, a ... | PeliKan Technologies, Inc. | A61B005/00
US06980842 | B2 | 2005-12-27 | Ocular analyte sensor | 2002-10-11 | An ophthalmic lens comprising a receptor moiety can be used to determine the amount of an analyte in an ocular fluid. The receptor moiety can bind either a specific analyte or a detectably labeled competitor moiety. The amount of detectably labeled competitor moiety which is displaced from the receptor moiety by the analyte is measured and provides a means of determining analyte concentration in an ocular fluid, such as tears, aqueous humor, or interstitial fluid. The concentration of the analyte in the ocular fluid, in turn, indicates the concentration of the analyte in a fluid or tissue sample of the body, such as blood or intracellular fluid. | 1. An ophthalmic lens for detecting glucose in an ocular fluid, comprising: competitor moieties; glucose oxidase molecules comprising glucose/competitor moiety binding sites to which glucose and competitor moieties can reversibly bind; and detectable labels which are associated with the glucose oxidase molecules and/or the competitor moieties and which are capable of generating an optical signal that changes in a concentration-dependent manner when glucose reversibly binds to the glucose/competitor moiety binding sites. | 1. An ophthalmic lens for detecting glucose in an ocular fluid, comprising: competitor moieties; glucose oxidase molecules comprising glucose/competitor moiety binding sites to which glucose and competitor moieties can reversibly bind; and detectable labels which are associated with the glucose oxidase molecules and/or the competitor moieties and which are capable of generating an optical signal that changes in a concentration-dependent manner when glucose reversibly binds to the glucose/competitor moiety binding sites. 2. An ophthalmic lens of claim 1, wherein the competitor moieties are selected from the group consisting of 2-deoxy-D-glucose, D-mannose and D-galactose. 3. An ophthalmic lens of claim 2, wherein the competitor moieties are tethered to a polymer meshwork, wherein pores of the polymer meshwork (a) permit glucose to bind reversibly to the glucose/competitor moiety binding sites and (b) prevent the glucose oxidase molecules from diffusing out of the ophthalmic lens. 4. An ophthalmic lens of claim 1, wherein at least one of the detectable labels is a fluorescent label. 5. A glucose sensor system, comprising: an ophthalmic lens for detecting glucose in an ocular fluid, wherein the ophthalmic lens includes competitor moieties, glucose oxidase molecules having glucose/competitor moiety binding sites to which glucose and the competitor moieties can reversibly bind, and detectable labels which are associated with the glucose oxidase molecules and/or the competitor moieties and which are capable of generating an optical signal that changes in a concentration-dependent manner when glucose reversibly binds to the glucose/competitor moiety binding sites; and a detector configured to detect the optical signal. 6. A glucose sensor system of claim 5, wherein the competitor moieties are selected from the group consisting of 2-deoxy-D-glucose, D-mannose and D-galactose. 7. A glucose sensor system of claim 6, wherein the competitor moieties are tethered to a polymer meshwork, wherein pores of the polymer meshwork (a) permit glucose to bind reversibly to the glucose/competitor moiety binding sites and (b) prevent the glucose oxidase molecules from diffusing out of the ophthalmic lens. 8. A glucose sensor system of claim 5, wherein the detector is a fluorophotometer. 9. A glucose sensor system of claim 5, further comprising an insulin pump and a transmitter coupled to the detector and configured to transmit to the insulin pump a signal indicating a glucose concentration in the ocular fluid, wherein the pump is configured to vary the level of glucose in a body fluid or tissue according to the glucose concentration in the ocular fluid. 10. A glucose sensor system of claim 9, wherein the transmitter is contained in a personal accessory. 11. A glucose sensor system of claim 10, wherein the transmitter is further configured to transmit the signal to the insulin pump wirelessly. 12. A glucose sensor system of claim 5, wherein the ophthalmic len... | Novartis | A61B005/00
US07001345 | B2 | 2006-02-21 | Wire guide | 2002-08-23 | A wire guide includes a mandrel that has a proximal portion and a distal portion. A first coating with a low coefficient of friction is disposed on the proximal portion of the mandrel. A second coating is disposed on the distal portion of the mandrel, where the second coating provides a sub-structure. A third coating is disposed on the second coating, where the third coating comprises a surface that allows for easy maneuverability of the wire guide. The first coating on the proximal portion provides enough lubricity to keep the wire guide from becoming bound or stuck in a catheter or medical device while still allowing a user to have a good grip of the wire guide. The second coating on the distal portion provides lubricity for the wire guide, which allows the user to easily maneuver the wire guide through a vascular anatomy. | 1. A wire guide apparatus comprising: a mandrel having a proximal portion and a distal portion;a first coating having a first coefficient of friction greater than 0.1 and less than or equal to 0.7 and being disposed on the proximal portion of the mandrel;a second coating disposed on the distal portion of the mandrel and being configured to provide a sub-structure thereto; anda third coating disposed on the second coating and having a second coefficient of friction of between about 0.01 and 0.1, the second coefficient of friction being lower than the first coefficient of friction, allowing for easy maneuverability of the wire guide. | 1. A wire guide apparatus comprising: a mandrel having a proximal portion and a distal portion;a first coating having a first coefficient of friction greater than 0.1 and less than or equal to 0.7 and being disposed on the proximal portion of the mandrel;a second coating disposed on the distal portion of the mandrel and being configured to provide a sub-structure thereto; anda third coating disposed on the second coating and having a second coefficient of friction of between about 0.01 and 0.1, the second coefficient of friction being lower than the first coefficient of friction, allowing for easy maneuverability of the wire guide. 2. The wire guide of claim 1, wherein the first coating allows a user to have a good grip of the wire guide. 3. The wire guide of claim 1, wherein the first coating comprises a polymer. 4. The wire guide of claim 3, wherein the polymer comprises polytetrafluoroethylene. 5. The wire guide of claim 3, wherein the polymer comprises polyethylene. 6. The wire guide of claim 1, wherein the third coating comprises a hydrophilic material. 7. The wire guide of claim 6, wherein the second coating comprises a polymeric material. 8. The wire guide of claim 7, wherein the polymeric material is polyurethane. 9. The wire guide of claim 1, wherein the mandrel is made of Nitinol. 10. The wire guide of claim 7, wherein the mandrel is super-elastic at or below 0 degrees Celsius. 11. The wire guide of claim 1, wherein a thickness of the first coating is about 0.0002 inches to about 0.20 inches. 12. The wire guide of claim 1, wherein the distal portion of the mandrel comprises a tapered portion and an un-tapered portion. 13. The wire guide of claim 12, further comprising a safety wire connected to the un-tapered portion. 14. The wire guide of claim 13, further comprising a tip coil connected to the tapered portion. 15. The wire guide of claim 14, wherein the tip coil comprises a radiopaque material. 16. The wire guide of claim 15, wherein the tip coil comprises a to least on one of the following material platinum, platinum alloy or stainless steel or a mixture thereof. 17. The wire guide apparatus of claim 1 wherein the first coefficient of friction is between about 0.04 and 0.7. 18. The wire guide apparatus of claim 8 wherein the hydrophilic material comprises a solvent in an uncured state and is attracted to the polyurethane for bonding between the second coating and third coating. 19. The wire guide apparatus of claim 1 wherein the first coefficient of friction is about 0.6. | Cook Incorporated | A61B005/00 A61M025/00
US06993379 | B1 | 2006-01-31 | System and method for emulating a surface EKG using an implantable cardiac stimulation device | 2002-12-30 | A surface electrocardiogram (EKG) is emulated using signals detected by the internal leads of an implanted device. The emulation is performed using a matrix-based technique that separately emulates each of the individual signals of a multiple-lead EKG, rather than merely emulating a single combined EKG. In one example, each of the twelve signals of a standard 12-lead EKG are individually emulated, allowing for separate processing and display. The emulation technique takes into account factors affecting the relative locations of the internal leads, such as respiration and posture, to thereby provide a more accurate emulation. A calibration technique is provided for calibrating the EKG emulation for use with a particular patient and a verification technique is provided for automatically verifying the reliability of the emulation. Any significant loss in emulation reliability is likely caused by lead dislodgment and so automatic detection of possible lead dislodgment is also achieved. | 1. A method for emulating a multiple-lead surface electrocardiogram (EKG) of a patient in which an implantable cardiac stimulation device is implanted, the method comprising: sensing electrical cardiac signals using a combination of pairs of electrodes implanted within the patient; andemulating each of a plurality of separate signals associated with the multiple-lead surface EKG using the electrical cardiac signals. | 1. A method for emulating a multiple-lead surface electrocardiogram (EKG) of a patient in which an implantable cardiac stimulation device is implanted, the method comprising: sensing electrical cardiac signals using a combination of pairs of electrodes implanted within the patient; andemulating each of a plurality of separate signals associated with the multiple-lead surface EKG using the electrical cardiac signals. 2. The method of claim 1 wherein emulating a plurality of signals associated with the multiple-lead surface EKG is performed to emulate each of twelve signals associated with a 12-lead surface EKG. 3. The method of claim 1 further comprising converting the plurality of signals generated into a single surface EKG signal. 4. The method of claim 1 wherein emulating the plurality of signals associated with the multiple-lead surface EKG comprises: converting input electrical signals into a time-varying vector F(t) having individual elements corresponding to the pairs of electrodes;inputting a conversion matrix M of weighting factors representative of the relative extent to which the signals derived from the pairs of electrodes influence surface voltages at locations corresponding to the multiple-lead surface EKG being emulated; andgenerating a time-varying vector E(t), representative of the surface voltages at the locations corresponding to the multiple-lead surface EKG being emulated, by calculating: E(t)=M*F(t). 5. The method of claim 4wherein the combination of electrodes includes N pairs of electrodes and the surface EKG to be emulated includes M surface leads;wherein F(t) has N elements, n=1, 2, . . . , N; E(t) has M elements, m=1, 2, . . . , M;and M is an N by M matrix containing weighting factors; andwherein inputting conversion matrix M comprises inputting a set of individual weighting factors representative of the extent to which the internal voltage derived from electrode pair n influences the surface voltage at surface lead m. 6. The method of claim 4 further comprising: inputting a signal representative of factors affecting the positions of one or more electrodes included in the pairs of electrodes; andwherein emulating is performed to take into account the factors affecting the positions of the electrodes. 7. The method of claim 6 wherein the factors affecting the positions of the electrodes include one or more of: phase of a respiratory cycle of the patient, phase of a cardiac cycle of the patient, and posture of the patient. 8. The method of claim 6 wherein emulating comprises adjusting the weighting factors of conversion matrix M to take into account the factor that affects the positions of the electrodes. 9. The method of claim 6 wherein emulating comprises selecting a different conversion matrix M based on a current status of the factor that affects the positions of the electrodes. 10. The method of claim 9 wherein two or more factors affecting the positions of the electrodes are detected and wherein emulating co... | Pacesetter, Inc. | A61B005/0402
US07066883 | B2 | 2006-06-27 | Method and system for monitoring the course of therapy of a patient being therapeutically treated | 2002-05-07 | In a method and system for monitoring the course of therapy administered to a patient, at least one information value that requires a therapeutic and/or diagnostic measure and that is a direct or indirect measure of the course of the therapy is registered continuously or in time intervals, and at least one comparison value, to be compared to the information value or to a calculated value determined based on the information value, is allocated to the information value. An alarm signal is emitted given an impermissible deviation of the information value or of the calculated value from the comparison value and is detected in the course of the comparison. The comparison value is varied dependent on the course of the therapy and is adapted to the course of the therapy, the adaptation of ensuing automatically or by means of an authorized person or entity. | 1. A method for monitoring therapy being administered to a patient, comprising the steps of: (a) acquiring an information value by obtaining a measurement from a patient representing a course of a therapy being administered to the patient;(b) designating a designated value selected from the group consisting of said information value and a calculated value calculated from said information value;(c) allocating a comparison value to said designated value and comparing said designated value to said comparison value in a comparison;(d) generating an alarm if said comparison indicates an impermissible deviation of said designated value from said comparison value;(e) at intervals during said course of said therapy selected from the group consisting of predetermined time intervals and variable time intervals, automatically determining, dependent on said designated value, whether a need exists to adapt said comparison value and, if so, automatically taking action selected from the group consisting of automatically changing said comparison value, and informing an agent of said need to adapt said comparison value; and(f) if said comparison value is to be adapted by an agent, automatically generating an indication to said agent of said need to adapt said comparison value. | 1. A method for monitoring therapy being administered to a patient, comprising the steps of: (a) acquiring an information value by obtaining a measurement from a patient representing a course of a therapy being administered to the patient;(b) designating a designated value selected from the group consisting of said information value and a calculated value calculated from said information value;(c) allocating a comparison value to said designated value and comparing said designated value to said comparison value in a comparison;(d) generating an alarm if said comparison indicates an impermissible deviation of said designated value from said comparison value;(e) at intervals during said course of said therapy selected from the group consisting of predetermined time intervals and variable time intervals, automatically determining, dependent on said designated value, whether a need exists to adapt said comparison value and, if so, automatically taking action selected from the group consisting of automatically changing said comparison value, and informing an agent of said need to adapt said comparison value; and(f) if said comparison value is to be adapted by an agent, automatically generating an indication to said agent of said need to adapt said comparison value. 2. A method as claimed in claim 1 wherein step (e) comprises adapting said comparison value by an authorized person as said agent. 3. A method as claimed in claim 1 wherein step (c) of adapting said comparison value comprises adapting said comparison value by an authorized entity as said agent. 4. A method as claimed in claim 1 wherein, in step (c), allocating a comparison value to said designated value comprises allocating an upper comparison value and a lower comparison value to said designated value, and wherein comparing said designated value to said comparison value comprises comparing said designated value to each of said upper comparison value and said lower comparison value. 5. A method as claimed in claim 4 wherein step (d) comprises generating an alarm if said designated value is above said upper comparison value or below said lower comparison value. 6. A method as claimed in claim 1 wherein said designated value is a current designated value, and further comprising determining said impermissible deviation for said comparison by analyzing a trend of said course of said therapy from a plurality of designated values preceding said current designated value. 7. A method as claimed in claim 1 wherein step (a) comprises acquiring a plurality of information values respectively from a plurality of measurements from said patient, and wherein the step of designating a designated value comprises designating a plurality of designated values respectively selected from the group consisting of said plurality of information values and a plurality of calculated values respectively calculated from said plurality of information values, and wherein step (c) comprises allocating a comparison val... | Siemens Aktiengesellschaft | A61B005/00
US07041068 | B2 | 2006-05-09 | Sampling module device and method | 2002-04-19 | A tissue penetration device and method of using same that may include a lancet module or sampling module. The sampling module may optionally be in a cartridge configuration and include sampling and analyzing functions, which may be integrated. | 1. A sampling module, comprising a module body portion comprising a lancet channel with a plurality of lancet bearing guides disposed within the lancet channel; a lancet comprising a sharpened distal tip and shaft portion which is slidably disposed within the lancet channel and lancet bearing guides with the lancet bearing guides configured to confine a lancet to substantially linear axial movement; and a sample reservoir in fluid communication with a sample input port of the module body portion. | 1. A sampling module, comprising a module body portion comprising a lancet channel with a plurality of lancet bearing guides disposed within the lancet channel; a lancet comprising a sharpened distal tip and shaft portion which is slidably disposed within the lancet channel and lancet bearing guides with the lancet bearing guides configured to confine a lancet to substantially linear axial movement; and a sample reservoir in fluid communication with a sample input port of the module body portion. 2. The sampling module of claim 1 wherein the module body portion comprises three lancet bearing guides separated spatially along the lancet channel with a lancet bearing guide disposed at either end of the lancet channel. 3. The sampling module of claim 1 wherein a longitudinal axis of the lancet channel is substantially parallel to a longitudinal axis of the module body portion. 4. The sampling module of claim 1 wherein the lancet channel further comprises at least one flow stop chamber disposed within a distal end portion of the lancet channel that has a transverse dimension that is significantly larger than a transverse dimension of a distally adjacent portion of the lancet channel and a rapid transition in transverse dimension that will interrupt capillary action from a distal end of the lancet channel. 5. The sampling module of claim 1 wherein the module body portion is configured to be mechanically registered and secured adjacent a lancet driver. 6. A sampling module, comprising: a module body portion configured to be mechanically registered and secured adjacent a lancet driver; a lancet comprising a sharpened distal tip and shaft portion which is slidably disposed within the module body portion; and a sample reservoir in fluid communication with a sample input port of the module body portion that has a fill detector for determining when the sample reservoir is full of a sample. 7. The sampling module of claim 6 wherein the fill detector is electrical. 8. The sampling module of claim 6 wherein the fill detector is optical. 9. The sampling module of claim 6 wherein the fill detector is visual. 10. A sampling module, comprising: a module body portion comprising a lancet channel; a lancet comprising a sharpened distal tip and shaft portion which is slidably disposed within the lancer channel of the module body portion; a first sample reservoir having an input end and a terminal end, the input end of the sample reservoir being in fluid communication with a sample input port of the module body portion; and a second sample reservoir having an input end in fluid communication with the terminal end of the first sample reservoir and having a cross sectional area that is substantially smaller than a cross sectional area of the first sample reservoir such that a sample flowing from the simple input port into the first sample reservoir will fill the first sample reservoir and then rapidly fill the second sample reservoir by capillary action. 11. The... | Pelikan Technologies, Inc. | A61B005/00
US07269456 | B2 | 2007-09-11 | Repetitive visual stimulation to EEG neurofeedback protocols | 2002-05-30 | An EEG neurofeedback and total evoked brain activity measurement methods utilize minimum ambient EEG activity as stimulant frequencies. A method of using repetitive stimulation in conjunction with EEG neurofeedback protocols is described. Electrodes, attached to a subject's scalp, transmit electroencephalographic (EEG) signals from the subject. These signals are in response to the visual and/or auditory stimuli being displayed to the subject. The resultant EEG signals are then filtered at pre-defined frequencies or frequency bands. The output from the filtered EEG signals is then analyzed and monitored for short-term state changes. The invention also uses flicker stimulation, real-time signal filtration and feedback, feedback during audio and visual stimulation derived from filtered outputs, and fundamental and integral harmonics in combination with total evoked response. | 1. A method of measuring sensory and perceptual mechanisms of the brain, and training these mechanisms with feedback, comprising the steps of: securing electrodes to a subject's scalp;displaying stimuli to the subject;receiving an electroencephalographic (EEG) signal from the subject while the stimuli is being displayed;transmitting the resultant EEG signal to a filter;narrow band filtering the EEG signal to a predefined frequency;reconstructing the total evoked brain activity by combining the filter outputs;feeding back in real-time of the total evoked response comprising at least one steady-state visual evoked potential (SSVEP) component amplitude; andmonitoring, in real-time, the short-term changes in each of the SSVEP component amplitudes, for the interpretation of waxing and waning behavior to reflect variations in sensory and perceptual vigilance for use as criteria for providing feedback for at least one of training, monitoring, and assessment purposes. | 1. A method of measuring sensory and perceptual mechanisms of the brain, and training these mechanisms with feedback, comprising the steps of: securing electrodes to a subject's scalp;displaying stimuli to the subject;receiving an electroencephalographic (EEG) signal from the subject while the stimuli is being displayed;transmitting the resultant EEG signal to a filter;narrow band filtering the EEG signal to a predefined frequency;reconstructing the total evoked brain activity by combining the filter outputs;feeding back in real-time of the total evoked response comprising at least one steady-state visual evoked potential (SSVEP) component amplitude; andmonitoring, in real-time, the short-term changes in each of the SSVEP component amplitudes, for the interpretation of waxing and waning behavior to reflect variations in sensory and perceptual vigilance for use as criteria for providing feedback for at least one of training, monitoring, and assessment purposes. 2. The method of claim 1 wherein the EEG signal is used to control a stimulator which produces a visual stimulus, wherein the control trains an increase or decrease in the evoked response. 3. The method of claim 1 wherein auditory feedback consisting of heard sounds is employed to reflect to the subject the subject's brain response to the applied visual, wherein the feedback trains the sensory/perceptual pathways of the brain. 4. The method of claim 1 wherein the subject receives, in real-time feedback specific to particular sensory and perceptual pathways of the brain, and receives feedback, of evoked-potential components that differentially reflect waxing and waning of independent generators in the brain, with real-time information facilitating operant or learned (volitional) conditioning of the brain. 5. The method of claim 1 wherein particular components of the SSVEP are extracted in real-time and used to make a control decision that can result in either a change in feedback, an automated decision, a switch closure, the triggering of an actuator, or another event. 6. The method of claim 1 wherein the composition of harmonic activity in the SSVEP is used in real-time to provide variations in an auditory feedback signal, so that the trainee hear, separately, the brain's visual primary and secondary sensory and perceptual response mechanisms in real time via an audio amplifier for training purposes. 7. A method of measuring sensory and perceptual mechanisms of the brain, either separately or together, and training these mechanisms with feedback, comprising the steps of: securing electrodes to a subject's scalp;displaying stimuli to the subject;receiving an electroencephalographic (EEG) signal from the subject while the stimuli is being displayed;transmitting the resultant EEG signal to a filter;narrow band filtering the EEG signal to a predefined frequency;reconstructing the total evoked brain activity by combining the filtered waveform outputs;feeding back in real-time of the total ... | - | A61B005/04
US07151958 | B2 | 2006-12-19 | System and method of classifying tachyarrhythmia episodes as associated or disassociated | 2002-08-02 | A method and system for classifying tachyarrhythmia episodes as associated or disassociated. Primary and auxiliary depolarizations are detected from different cardiac regions during a tachycardia episode. A first time interval is positioned to surround each of one or more primary depolarizations, including auxiliary depolarizations that occur in the first time interval. The auxiliary depolarizations occurring in the first time interval are then counted. Based on the number and the location of the auxiliary depolarizations counted in the first time intervals, the one or more primary depolarizations and the auxiliary depolarizations of the tachycardia episode are classified as disassociated or associated. | 1. A system, comprising: control circuitry where the control circuitry receives a first cardiac signal and a second cardiac signal sensed during a cardiac episode, where the first cardiac signal includes indications of depolarizations in a first cardiac area and the second cardiac signal includes indications of depolarizations in a second cardiac area, the control circuitry including: a first depolarization detector circuit to detect occurrence of primary depolarizations in the first cardiac signal;a second depolarization detector circuit to detect occurrence of auxiliary depolarizations in the second cardiac signal;a cardiac depolarization counting circuit coupled to the first depolarization detector circuit and the second depolarization detector circuit, where the cardiac depolarization counting circuit counts a first number (K) of auxiliary depolarizations occurring during a first portion of a first time interval surrounding each of one or more primary depolarizations of the primary depolarizations and counts a total number (N) of auxiliary depolarizations during the first time interval; anda microprocessor coupled to the cardiac depolarization counting circuit, where the microprocessor calculates K/N and determines whether the primary depolarizations and the auxiliary depolarizations in the cardiac episode are associated or disassociated based on K/N. | 1. A system, comprising: control circuitry where the control circuitry receives a first cardiac signal and a second cardiac signal sensed during a cardiac episode, where the first cardiac signal includes indications of depolarizations in a first cardiac area and the second cardiac signal includes indications of depolarizations in a second cardiac area, the control circuitry including: a first depolarization detector circuit to detect occurrence of primary depolarizations in the first cardiac signal;a second depolarization detector circuit to detect occurrence of auxiliary depolarizations in the second cardiac signal;a cardiac depolarization counting circuit coupled to the first depolarization detector circuit and the second depolarization detector circuit, where the cardiac depolarization counting circuit counts a first number (K) of auxiliary depolarizations occurring during a first portion of a first time interval surrounding each of one or more primary depolarizations of the primary depolarizations and counts a total number (N) of auxiliary depolarizations during the first time interval; anda microprocessor coupled to the cardiac depolarization counting circuit, where the microprocessor calculates K/N and determines whether the primary depolarizations and the auxiliary depolarizations in the cardiac episode are associated or disassociated based on K/N. 2. The system of claim 1, where the first time interval has an approximate midpoint location positioned at each of the one or more primary depolarizations, where the first portion of the first time interval is substantially a first-half of the first time interval, and where the microprocessor determines that the primary depolarizations and the auxiliary depolarizations in the cardiac episode are disassociated when K/N is equal to approximately one-half (0.5). 3. The system of claim 1, where the first time interval has an approximate midpoint location positioned at each of the one or more primary depolarizations, where the first portion of the first time interval is substantially a first-half of the first time interval, and where the microprocessor determines that the primary depolarizations and the auxiliary depolarizations in the cardiac episode are associated when K/N is not equal to approximately one-half (0.5). 4. The system of claim 1, where the first time interval has an approximate midpoint location positioned at each of the one or more primary depolarizations, where the first portion of the first time interval is substantially a first-half of the first time interval, and where the microprocessor determines that the primary depolarizations and the auxiliary depolarizations in the cardiac episode are associated when K/N is statistically different from one-half (0.5) at a predetermined confidence level and where the microprocessor determines that the primary depolarizations and the auxiliary depolarizations in the cardiac episode are disassociated when K/N is not statistically different ... | Cardiac Pacemakers, Inc. | A61B005/046
US07082326 | B2 | 2006-07-25 | Method of magnetic resonance imaging | 2002-09-27 | A method of magnetic resonance (MR) imaging of blood perfusion in vascularized organs of interest, e.g. the myocardium, the kidneys, the brain, the liver wherein a contrast agent composition comprising a particulate superparamagnetic metal oxide is administered. | 1. A method of first pass, contrast enhanced organ perfusion magnetic resonance imaging of a human or vascularized animal subject, comprising administering a MR contrast agent composition comprising a particulate superparamagnetic metal oxide into the vasculature of said subject and generating magnetic resonance images of slices through a vascularized organ of said subject during the passage of said agent through the vasculature of said organ, wherein pre-saturation slabs in the phase encoding direction and a T2-weighted single shot Turbo (fast) spin-echo sequence are used to generate said images. | 1. A method of first pass, contrast enhanced organ perfusion magnetic resonance imaging of a human or vascularized animal subject, comprising administering a MR contrast agent composition comprising a particulate superparamagnetic metal oxide into the vasculature of said subject and generating magnetic resonance images of slices through a vascularized organ of said subject during the passage of said agent through the vasculature of said organ, wherein pre-saturation slabs in the phase encoding direction and a T2-weighted single shot Turbo (fast) spin-echo sequence are used to generate said images. 2. The method as claimed in claim 1 wherein said vascularized organ is the myocardium, the kidneys, the brain or the liver. 3. The method as claimed in claim 1 wherein said vascularized organ is the myocardium. 4. The method according to claim 1 wherein said superparamagnetic metal oxide is a superparamagnetic iron oxide. 5. The method as claimed in claim 4 wherein said superparamagnetic iron oxide is an ultra small superparamagnetic iron oxide (USPIO). 6. The method as claimed in claim 5 wherein said ultra small superparamagnetic iron oxide (USIPO) having on its surface an optionally modified carbohydrate or polysaccharide or derivative thereof. 7. The method as claimed in claim 1 wherein said MR contrast agent composition is administered by bolus injection over a period of less than 3 minutes. | Amersham Health AS | A61B005/055
US06968223 | B2 | 2005-11-22 | System and method for wireless voice control of an interventional or diagnostic medical device | 2002-02-01 | A system and method of controlling a medical device through voice commands is provided that includes an input unit, such as a microphone, for receiving a voice command identifying a function associated with one of a diagnostic and interventional procedure. The system also includes a control module for directing a medical device to perform the function based on the voice command. The system includes a voice decoder that decodes the voice command into a basic signal, and a protocol translator that converts the basic signal into a signal code associated with the voice command. The control module also includes a processing unit. The control module receives the signal code and the processing unit directs the medical device to perform the function designated by the signal code. | 1. A system for controlling a medical device through voice commands, comprising: a medical device for performing at least one of interventional and diagnostic procedures; an input unit for receiving a voice command identifying a function associated with one of a diagnostic and interventional procedure, wherein the function includes movement of said medical device; and a control module for directing said medical device to perform the function based on said voice command. | 1. A system for controlling a medical device through voice commands, comprising: a medical device for performing at least one of interventional and diagnostic procedures; an input unit for receiving a voice command identifying a function associated with one of a diagnostic and interventional procedure, wherein the function includes movement of said medical device; and a control module for directing said medical device to perform the function based on said voice command. 2. The system of claim 1 wherein said input unit is a microphone. 3. The system of claim 1 further including a voice decoder for decoding said voice command into a basic signal. 4. The system of claim 3 further including a protocol translator for converting said basic signal into a signal code representing the function identified by said voice command. 5. The system of claim 4 wherein said control module includes a processing unit for directing said medical device to perform said function designated by said signal code. 6. The system of claim 4 further including a transmitter for transmitting said signal code to said control module. 7. The system of claim 1 wherein said medical device is an interventional medical device. 8. The system of claim 1, further comprising additional functions including positioning and imaging characteristics of said medical device, wherein said control module is configured to direct said medical device to perform the function and the additional functions based on said voice command and additional voice commands, respectively. 9. A system for controlling a medical device through voice commands, comprising: a medical device for performing one of a medical diagnostic and interventional procedure; a microphone for receiving a voice command identifying a function associated with one of said medical diagnostic and interventional procedure, wherein the function includes movement of said medical device; a voice decoder for decoding said voice command into a basic signal; and a protocol translator for converting said basic signal into a signal code representing said function identified by said voice command; and a processing unit for directing said medical device to perform said function designated by said signal code. 10. The system of claim 9 wherein said medical device is an interventional medical device. 11. The system of claim 9 wherein said signal code is an infrared (IR) signal code. 12. The system of claim 9 wherein said signal code is a radio frequency (RF) signal code. 13. The system of claim 9 wherein said medical device is a magnetic resonance imaging (MRI) device. 14. The system of claim 9 wherein said medical device is a computerized tomography imaging device. 15. The system of claim 9 wherein said medical device is a fluoroscopic imaging device. 16. The system of claim 9 further including: a transmitter associated with said microphone, wherein said transmitter transmits said signal code, and a receiver provided at said medical devi... | GE Medical Systems Global Technology Company, LLC | A61B005/05
US06878262 | B2 | 2005-04-12 | Analytical element and measuring device and substrate quantification method using the same | 2002-08-13 | The present invention provides an analytical element which is free from evaporation of a sample during measurement and therefore capable of quantifying a substrate using a very small amount of sample with high accuracy and which is free from scattering of the sample during and after the measurement and therefore hygienically excellent; and a measuring device and a substrate quantification method using the same. The analytical element comprises a cavity for accommodating a sample, a working electrode and a counter electrode exposed to an inside of the cavity, a reagent layer which comprises at least an oxidoreductase and is formed inside or in the vicinity of the cavity, an opening communicating with the cavity and a member covering the opening. | 1. An analytical element comprising a cavity for accommodating a sample, a working electrode and a counter electrode exposed to an inside of said cavity, a reagent layer which comprises at least an oxidoreductase and is formed inside or in the vicinity of said cavity, an opening communicating with said cavity and a member for covering said opening, and a driving unit coupled to the member, wherein the driving unit drives the member to cover and/or uncover the opening. | 1. An analytical element comprising a cavity for accommodating a sample, a working electrode and a counter electrode exposed to an inside of said cavity, a reagent layer which comprises at least an oxidoreductase and is formed inside or in the vicinity of said cavity, an opening communicating with said cavity and a member for covering said opening, and a driving unit coupled to the member, wherein the driving unit drives the member to cover and/or uncover the opening. 2. The analytical element in accordance with claim 1, wherein the member covering the opening is a puncture member. 3. The analytical element in accordance with claim 1, wherein the member covering the opening is transparent. 4. The analytical element in accordance with claim 1, wherein the member covering the opening also serves as the working electrode or the counter electrode. 5. The analytical element in accordance with one of claims 1 though 4, wherein the member covering the opening is provided with the reagent layer. 6. The analytical element in accordance with claim 1, further comprising a position controlling unit coupled to the driving unit, said position controlling unit controlling the position of said member. 7. An analytical element comprising: a rod-like element body comprising an insulating member which has a depression for accommodating a sample at its side face in the vicinity of its tip end and a working electrode and a counter electrode arranged so as to be exposed to an inside of said depression; a cover which covers the face of the element body having said depression, said cover having a groove along a longitudinal direction of said element body at a portion facing said depression; a reagent layer which comprises at least an oxidoreductase and is formed inside or in the vicinity of said depression; and a puncture member which moves in the longitudinal direction of the element body in said groove and sticks out of the tip end of the element body to puncture a subject, wherein a sample supply opening communicating with said depression is formed between the tip end of the element body and the cover, an air vent is formed at or above a position of the cover corresponding to said depression, and the puncture member is capable of substantially covering said depression. 8. An analytical element comprising: an element body comprising an insulating base plate on which an electrode system comprising a working electrode and a counter electrode is formed and a cover member which covers said base plate and forms a sample supply pathway communicating with said electrode system between the cover and the base plate; a reagent layer which comprises at least an oxidoreductase and is formed inside said sample supply pathway; and a top cover, wherein said sample supply pathway extends from a sample supply port at one end of the element body to an air vent formed in the cover member, and the top cover is configured so as to be combined to the element body by sliding it in a ... | Matsushita Electric Industrial Co., Ltd. | G01N027/327 A61B005/05
US06931270 | B2 | 2005-08-16 | Method and system for conducting medical imaging transactions | 2002-02-11 | Method and system for remotely conducting medical imaging transactions conducive to meeting the individualized imaging needs of a plurality of users are provided. The method provides a user interface configured to present a menu of prompts to a user of a respective medical imager and/or imaging equipment associated with the imager. The menu of prompts is configured to elicit information indicative of imaging performance of the medical imager and/or equipment relative to imaging preferences of the user. The method further allows transmitting the information indicative of the imaging performance over a communications network to an imaging service center. A database is configured to store the information indicative of the imaging performance of the medical imager and/or imaging equipment. The database is further configured to store the imaging preferences of the user. The information stored in the database is processed relative to a set of business rules to select imaging upgrades configured to improve the operational performance of the medical imager. The imaging upgrades are transmitted to reconfigure the imager and/or associated equipment so that the individualized imaging needs of each user are achieved based on the imaging upgrades. | 1. A method for remotely conducting medical imaging transactions conducive to meeting the individualized imaging needs of a plurality of users, the method comprising: providing a user interface configured to present a menu of prompts to a user of a respective medical imager and/or imaging equipment associated with the imager, the menu of prompts configured to elicit information indicative of imaging performance of the medical images and/or equipment relative to imaging preferences of the user; transmitting the information indicative of the imaging performance over a communications network to an imaging service center; providing a database accessible to the service center, the database configured to store the information indicative of the imaging performance of the medical imager and/or imaging equipment wherein the information indicative of the imaging performance transmitted to the imaging service center comprises an image generated subject to the imaging preferences of the user, the database further configured to store the imaging preferences of the user; processing the information stored in the database relative to a set of business rules to select imaging upgrades configured to improve the operational performance of the medical imager; transmitting the imaging upgrades to reconfigure the imager and/or associated equipment so that the individualized imaging needs of each user are achieved based on the imaging upgrades; and, comparing the transmitted image relative to reference images corresponding to the imaging preferences of the user, and, based on the results of the comparison, issuing a recommendation to the user regarding the medical imager and/or imagine equipment. | 1. A method for remotely conducting medical imaging transactions conducive to meeting the individualized imaging needs of a plurality of users, the method comprising: providing a user interface configured to present a menu of prompts to a user of a respective medical imager and/or imaging equipment associated with the imager, the menu of prompts configured to elicit information indicative of imaging performance of the medical images and/or equipment relative to imaging preferences of the user; transmitting the information indicative of the imaging performance over a communications network to an imaging service center; providing a database accessible to the service center, the database configured to store the information indicative of the imaging performance of the medical imager and/or imaging equipment wherein the information indicative of the imaging performance transmitted to the imaging service center comprises an image generated subject to the imaging preferences of the user, the database further configured to store the imaging preferences of the user; processing the information stored in the database relative to a set of business rules to select imaging upgrades configured to improve the operational performance of the medical imager; transmitting the imaging upgrades to reconfigure the imager and/or associated equipment so that the individualized imaging needs of each user are achieved based on the imaging upgrades; and, comparing the transmitted image relative to reference images corresponding to the imaging preferences of the user, and, based on the results of the comparison, issuing a recommendation to the user regarding the medical imager and/or imagine equipment. 2. The method of claim 1 wherein the imaging upgrades comprise an upgraded set of imaging parameters for controlling the imager and/or upgraded imaging processing algorithms. 3. The method of claim 1 wherein the imaging equipment comprises an imaging probe and the imaging upgrade comprises selecting a probe from a database of available probes for the imager. 4. The method at claim 1 wherein the imaging equipment comprises an imaging probe and the imaging upgrade comprises selecting a new probe design configured to meet a desired imaging preference of the user. 5. The method of claim 1 wherein the medical imager is selected from the group consisting of ultrasound, magnetic resonance, X-ray, computer tomography, and nuclear imagers. 6. The method of claim 1 further comprising accumulatin, historical information indicative of the imaging performance of the medical imager and/or equipment relative to imaging preferences of each user for a plurality of corresponding imagers and/or equipment. 7. A system for remotely conducting medical imaging transactions conducive to meeting the individualized imaging needs of a plurality of users, the system comprising: a user interface configured to present a menu of prompts to a user of a respective medical imager and/or imaging equipmen... | General Electric Company | A61B005/00 G06F017/00
US06944492 | B1 | 2005-09-13 | Patient bed support for an open MRI system | 2002-06-20 | A patient bed support for maneuvering a patient in an open MRI system that includes an upper pole assembly and a lower pole assembly. The patient bed support includes a frame for attaching the patient bed support to the lower pole assembly in the MRI system and a member movably attached to the frame such that a patient placed on the member may be moved relative to the lower and upper pole assemblies of the MRI system. | 1. A patient bed support for maneuvering a patient in an open MRI system that includes an upper pole assembly and a lower pole assembly, the patient bed support comprising: a frame attaching the patient bed support to the lower pole assembly in the MRI system;a member movably attached to the frame to maneuver patients that are placed on the member relative to the lower and upper pole assembly of the MRI system, wherein the member rotates relative to the frame and the frame is a first ring and the member is a second ring, the first and second rings being sized in order to position the first and second rings around an upper portion of the lower pole assembly. | 1. A patient bed support for maneuvering a patient in an open MRI system that includes an upper pole assembly and a lower pole assembly, the patient bed support comprising: a frame attaching the patient bed support to the lower pole assembly in the MRI system;a member movably attached to the frame to maneuver patients that are placed on the member relative to the lower and upper pole assembly of the MRI system, wherein the member rotates relative to the frame and the frame is a first ring and the member is a second ring, the first and second rings being sized in order to position the first and second rings around an upper portion of the lower pole assembly. 2. The patient bed support of claim 1, wherein the first ring and the second ring have a common longitudinal axis. 3. The patient bed support of claim 2, wherein the second ring rotates about the longitudinal axis of the first and second rings. 4. The patient bed support of claim 2, wherein the first ring includes a first semi-circular section and a second semi-circular section that are coupled together to form the first ring. 5. The patient bed support of claim 4, wherein the second ring includes a third semi-circular section and a fourth semi-circular section that are coupled together to form the second ring. 6. The patient bed support of claim 1, further comprising bearings between the first and second rings to facilitate rotating the second ring relative to the first ring. 7. The patient bed support of claim 1, wherein the first and second rings are substantially horizontal when attached to the lower pole assembly of the MRI system. 8. The patient bed support of claim 1, further comprising one or more supports extending from opposing sides of the second ring, and a platform extending between the supports to receive a patient to be analyzed using the MRI system. 9. The patient bed support of claim 8, wherein one set of two supports extends from one side of the second ring, and another set of two supports extends from an opposing side of the second ring. 10. The patient bed support of claim 8, wherein at least one of the supports includes lifting means to maneuver the platform relative to the second ring. 11. The patient bed support of claim 8, wherein the platform includes a recessed section that receives the upper portion of the lower pole assembly when the patient bed support is mounted to the lower pole assembly. 12. A patient bed support for maneuvering a patient in an open MRI system that includes an upper pole assembly and a lower pole assembly, the patient bed support comprising: a first ring attaching the patient bed support to the lower pole assembly in the MRI system;a second ring rotatably attached to the first ring;a platform for receiving patients that are to be analyzed by the MRI system, the platform engaged with the second ring such that the patient is maneuvered within the MRI system by rotating the second ring relative to the first ring. 13. The patient bed sup... | Fonar Corporation | A61B005/055
US06936016 | B2 | 2005-08-30 | Method for analysis of abnormal body tremors | 2002-05-17 | A method useful for evaluating abnormal body tremors wherein a subject is positioned onto a support surface and one or more components of the load exerted by the subject upon the support surface are measured. Simultaneously, the cardiac activity of the subject is also recorded and the time interval between each consecutive heartbeat is computed. The measured load output is modified to normalize the weight of the subject and then plotted against a modified frequency wherein each time interval between heartbeats represent a unit of time and a selected equal number of force measurement samples are interpolated between each heartbeat. The frequency spectrum analysis of such a modified power spectrum graph yields information about body tremor which can be compared to a prior analysis of the same subject or to the same analysis of a standard. The standard is created using the identical method upon, preferably, a sufficient number of subjects known to be free of abnormal body tremor to yield an average measure useful for comparison and diagnostic purposes. | 1. A method for evaluating body tremor comprising the steps of: a) measuring one or more selected components of the load exerted by a human subject upon a support surface and simultaneously determining the time interval between each of a selected number of consecutive heartbeat events corresponding to each heartbeat of the subject; b) computing a modified frequency distribution of the load components measured in step (a) by interpolation of selected samples of said measurements wherein the samples of the measurements of the selected components made step (a) are equally spaced within each time interval determined in step (a) and of equal number from one interval to the next regardless of the heart rate of the subject being evaluated to minimize the effect of heart beat modulation of the subject upon the frequency distribution computed. | 1. A method for evaluating body tremor comprising the steps of: a) measuring one or more selected components of the load exerted by a human subject upon a support surface and simultaneously determining the time interval between each of a selected number of consecutive heartbeat events corresponding to each heartbeat of the subject; b) computing a modified frequency distribution of the load components measured in step (a) by interpolation of selected samples of said measurements wherein the samples of the measurements of the selected components made step (a) are equally spaced within each time interval determined in step (a) and of equal number from one interval to the next regardless of the heart rate of the subject being evaluated to minimize the effect of heart beat modulation of the subject upon the frequency distribution computed. 2. The method defined in claim 1 further comprising performing the identical steps on a selected number of human subjects known to be free of any significant abnormal body tremor and computing a modified frequency distribution representing an average of the one or more components of the load generated by each of the subjects to determine a modified frequency distribution standard and comparing said standard to a similar modified frequency distribution obtained performing the method of claim 1 upon a selected subject for body tremor diagnostic purposes. 3. The method defined in claim 1 wherein heartbeat events recorded represent the QRS peaks obtained using an electrocardiograph device. 4. The method defined in claim 1 wherein the heartbeat events recorded represent the blood pressure peak during each heartbeat. 5. The method defined in claim 1 wherein the components of the load measured are those forces exerted in a vertical direction relative to said support surface. | Bertec Corporation | A61B005/103
US06920349 | B2 | 2005-07-19 | System and method for predicting the onset of cardiac pathology using fractal analysis | 2002-07-01 | The onset of cardiac pathology can be predicted by first acquiring a plurality of lead values as a function of time for set of electrocardiogram leads and defining a spatial curve from the lead values for at least three leads. A fractal index for the spatial curve is calculated as a function of time and the time rate of change of the fractal index is monitored. A negative time rate of change is indicative of normal cardiac activity, while a positive time rate of change is indicative of pathological cardiac activity. | 1. A method for predicting cardiac pathology comprising the steps of: acquiring a plurality of lead values as a function of time for a set of electrocardiogram leads; defining a spatial curve from the lead values for at least three leads; calculating a fractal index as a function of time for the spatial curve; and monitoring the time rate of change of the fractal index. | 1. A method for predicting cardiac pathology comprising the steps of: acquiring a plurality of lead values as a function of time for a set of electrocardiogram leads; defining a spatial curve from the lead values for at least three leads; calculating a fractal index as a function of time for the spatial curve; and monitoring the time rate of change of the fractal index. 2. The method of claim 1, further comprising the step of calculating the lead values from electrode readings taken from electrodes attached to a patient. 3. The method of claim 1, wherein a negative time rate of change of the fractal index is indicative of normal cardiac activity, and a positive time rate of change of the fractal index is indicative of pathological cardiac activity. 4. The method of claim 1, wherein the set of electrocardiogram leads comprises from 3 to about 80 leads, and wherein the full set of electrocardiogram leads can be calculated from a subset of at least 3 electrocardiogram leads by using a universal transformation matrix. 5. The method of claim 1, wherein the fractal index is defined by the equation FD(k)=log(k)/(log(k)+log(DD(tk)/LSUM(tk))), wherein: tk=kδt is the time after k intervals of time unit δt; LSUM⁡(tk)=LSUM⁡(tk-1)+(∑i=1N⁢(xi⁡(tk)-xi⁡(tk-1))2)12;DD⁡(tk)=max⁡((∑i=1N⁢xi2⁡(tk))12,DD⁡(tk-1));N represents the number of electrocardiogram leads; and xi(tk) represents the value of the ith lead in the N-lead set at time tk. 6. The method of claim 1, wherein the fractal index is defined by the equation K1=(LSUM(tk)^FD(k))/SumArea(tk) wherein tk=kδt is the time after k intervals of time unit δt; LSUM⁡(tk)=LSUM⁡(tk-1)+(∑i=1N⁢(xi⁡(tk)-xi⁡(tk-1))2)12; FD(k)=log(k)/(log(k)+log(DD(tk)/LSUM(tk))); DD⁡(tk)=max⁡((∑i=1N⁢xi2⁡(tk))12,DD⁡(tk-1));SumArea⁡(tk)=⁢SumArea⁢(tk-1)+⁢14⁢((D⁡(tk)+D⁡(tk-1)+L⁡(tk))⁢(D⁡(tk)-D⁡(tk-1)+⁢L⁢(tk))⁢(-D⁡(tk)+D⁡(tk-1)+L⁡(tk))⁢(D⁡(tk)+⁢D⁡(tk-1)-L⁡(tk)))12,whereD⁡(tk)=(∑i=1N⁢xi2⁡(tk))12⁢ ⁢and L⁡(tk)=(∑i=1N⁢(xi⁡(tk)-xi⁡(tk-1))2)12;N represents the number of electrocardiogram leads; and xi(tk) represents the value of the ith lead in the N-lead set at time tk. 7. The method of claim 1, wherein the fractal index is defined by the equation K1=(PerimTot(tk)^FD(k))/SumArea(tk) wherein tk=kδt is the time after k intervals of time unit δt; PerimTot⁡(tk)=(∑i=1N⁢xi2⁡(tk))12+LSUM⁡(tk)LSUM⁡(tk)=LSUM⁡(tk-1)+(∑i=1N⁢(xi⁡(tk)-xi⁡(tk-1))2)12; FD(k)=log(k)/(log(k)+log(DD(tk)/LSUM(tk))); DD⁡(tk)=max⁡((∑i=1N⁢xi2⁡(tk))12,DD⁡(tk-1));SumArea⁡(tk)=⁢SumArea⁢(tk-1)+⁢14⁢((D⁡(tk)+D⁡(tk-1)+L⁡(tk))⁢(D⁡(tk)-D⁡(tk-1)+⁢L⁢(tk))⁢(-D⁡(tk)+D⁡(tk-1)+L⁡(tk))⁢(D⁡(tk)+⁢D⁡(tk-1)-L⁡(tk)))12;whereD⁡(tk)=(∑i=1N⁢xi2⁡(tk))12⁢ ⁢and L⁡(tk)=(∑i=1N⁢(xi⁡(tk)-xi⁡(tk-1))2)12;N represents the number of electrocardiogram leads; and xi(tk) represents the value of the ith lead in the N-lead set at time tk. 8. The method of claim 1, further comprising the step of displaying the spatial loop in a heart monitor. 9. The method of claim 1, further comprising the steps of displaying t... | - | A61B005/0402
US06850206 | B2 | 2005-02-01 | Local radio-frequency antenna for a magnetic resonance apparatus | 2002-11-26 | A local radio-frequency antenna for a magnetic resonance apparatus is fashioned for a region to be imaged in an edge region of and/or outside an imaging volume of the magnetic resonance apparatus, and the radio-frequency antenna has a passive shim device. | 1. A local radio-frequency antenna for a magnetic resonance apparatus, said magnetic resonance apparatus having an imaging volume, defined by a spatial extent of a homogeneous basic magnetic field, said local radio-frequency antenna comprising: a radio-frequency antenna structure disposed and shaped for obtaining signals from a volume encompassing space outside of said imaging volume; and a passive shim device physically supported by said antenna structure for shimming said volume encompassing space outside of said imaging volume. | 1. A local radio-frequency antenna for a magnetic resonance apparatus, said magnetic resonance apparatus having an imaging volume, defined by a spatial extent of a homogeneous basic magnetic field, said local radio-frequency antenna comprising: a radio-frequency antenna structure disposed and shaped for obtaining signals from a volume encompassing space outside of said imaging volume; and a passive shim device physically supported by said antenna structure for shimming said volume encompassing space outside of said imaging volume. 2. The local radio-frequency antenna as claimed in claim 1 wherein said antenna structure forms a structural unit with said shim device. 3. The local radio-frequency antenna as claimed in claim 1 wherein said shim device comprises a carrier device and shim elements, composed of magnetic material, carried by said carrier device. 4. The local radio-frequency antenna as claimed in claim 3 wherein said antenna structure has a receptacle space, and wherein said carrier device is introducible into and withdrawable from said receptacle space. 5. The local radio-frequency antenna as claimed in claim 3 wherein said shim elements are composed of ferromagnetic material. 6. The local radio-frequency antenna as claimed in claim 1 wherein said antenna structure is adapted for use at a prescribed position relative to said imaging volume. 7. The A local radio-frequency antenna as claimed in claim 1 wherein said antenna structure is rigid. | Siemens Atiengesellschaft | H01Q001/00 A61B005/055
US07283869 | B2 | 2007-10-16 | Apparatus for measuring body fat | 2002-02-15 | An apparatus has a first electrode group (1) including at least one electrode and arranged on the abdominal surface of a subject (human) with the navel of the subject serving as a reference position, a second electrode group (2) including at least one electrode and arranged on the back surface of the subject, a third electrode group (3) including at least two electrodes and arranged on the surface of the subject at an intermediate position between the first and second electrode groups (1, 2), and a controller (8) to supply a current between an electrode selected from the first electrode group (1) and an electrode selected from the second electrode group (2), measure a voltage generated between two electrodes of the third electrode group (3), and compute an abdominal fat quantity of the subject according to the measured voltage. | 1. An apparatus for measuring body fat, comprising: a first electrode group including at least an electrode and configured to be arranged on the abdominal surface of a subject with the navel of the subject serving as a reference position;a second electrode group including at least an electrode and configured to be arranged on the back surface of the subject;a third electrode group including at least two electrodes and configured to be arranged on one side of the surface of the subject at an intermediate position between the first and second electrode groups;a measuring unit configured to measure a voltage generated between predetermined two electrodes of the third electrode group, the voltage being generated in response to a current supplied between an electrode selected from the first electrode group and an electrode selected from the second electrode group; anda computing unit configured to compute an abdominal fat quantity of the subject according to the measured voltage. | 1. An apparatus for measuring body fat, comprising: a first electrode group including at least an electrode and configured to be arranged on the abdominal surface of a subject with the navel of the subject serving as a reference position;a second electrode group including at least an electrode and configured to be arranged on the back surface of the subject;a third electrode group including at least two electrodes and configured to be arranged on one side of the surface of the subject at an intermediate position between the first and second electrode groups;a measuring unit configured to measure a voltage generated between predetermined two electrodes of the third electrode group, the voltage being generated in response to a current supplied between an electrode selected from the first electrode group and an electrode selected from the second electrode group; anda computing unit configured to compute an abdominal fat quantity of the subject according to the measured voltage. 2. The apparatus of claim 1, further comprising an input unit to enter a circumferential length of the abdominal part of the subject,the computing unit computing an abdominal fat quantity of the subject according to the measured voltage and entered circumferential length. 3. The apparatus of claim 2, wherein: the first, second, and third electrode groups are installed on a belt that is substantially inelastic; anda distance between the first and third electrode groups is substantially equal to a distance between the third and second electrode groups. 4. The apparatus of claim 2, wherein: the first, second, and third electrode groups are installed on a belt that is elastic; anddistances between the electrode groups are adjustable due to the elasticity of the belt. 5. The apparatus of claim 2, wherein: the first, second, and third electrode groups are installed on a belt that is substantially inelastic; andthe belt has adjusters to adjust distances between the electrode groups. 6. The apparatus of claim 2, further comprising joints connecting the first and third electrode groups to each other and the second and third electrode groups to each other. 7. The apparatus of claim 2, further comprising attachments configured to attach the first, second, and third electrode groups to the surface of the subject. 8. The apparatus of claim 1, wherein the second electrode group is positioned on the back surface of the subject relative to a spine of the subject. 9. The apparatus of claim 1, wherein: the first, second, and third electrode groups are installed on a belt that is substantially inelastic; anda distance between the first and third electrode groups is substantially equal to a distance between the third and second electrode groups. 10. The apparatus of claim 1, wherein: the first, second, and third electrode groups are installed on a belt that is elastic; anddistances between the electrode groups are adjustable due to the elasticity of the belt. 11. The apparatus of claim... | Kao Corporation | A61B005/05
US07087034 | B2 | 2006-08-08 | Vascular shunt with audio flow indication | 2002-09-19 | A vascular shunt apparatus (10) includes a tubular member having first and second end portions (16, 20) and an aperture extending through the first and second end portions (16, 20). A transducer (64) can be associated with the tubular member to provide a signal in response to the flow of fluid through the tubular member. One or both of the end portions (16, 20) also can be adapted to form respective sealing connections with different parts of a patient's vascular system. | 1. A vascular shunt apparatus comprising: an elongate tubular member having first and second end portions spaced apart by a length of an intermediate portion, an aperture extending axially through the tubular member between the first and second portions;the first end portion being adapted to form a first sealing connection with a fluid carrying system at a first location, the second end portion being adapted to form a second sealing connection with the fluid carrying system at a second location spaced apart from the first location to facilitate fluid flow through the aperture of the tubular member between the first and second locations;a single piezo-electric crystal connected with the tubular member for monitoring the fluid flow through the tubular member, the single piezo-electric crystal being configured provide an electrical signal in response to flow of fluid through the tubular member; andan electrical circuit electrically connected with the single piezoelectric crystal, the electrical circuit being configured to convert the electrical signal provided by the single piezo-electric crystal to an audible sound corresponding to the flow of fluid through the tubular member. | 1. A vascular shunt apparatus comprising: an elongate tubular member having first and second end portions spaced apart by a length of an intermediate portion, an aperture extending axially through the tubular member between the first and second portions;the first end portion being adapted to form a first sealing connection with a fluid carrying system at a first location, the second end portion being adapted to form a second sealing connection with the fluid carrying system at a second location spaced apart from the first location to facilitate fluid flow through the aperture of the tubular member between the first and second locations;a single piezo-electric crystal connected with the tubular member for monitoring the fluid flow through the tubular member, the single piezo-electric crystal being configured provide an electrical signal in response to flow of fluid through the tubular member; andan electrical circuit electrically connected with the single piezoelectric crystal, the electrical circuit being configured to convert the electrical signal provided by the single piezo-electric crystal to an audible sound corresponding to the flow of fluid through the tubular member. 2. The apparatus of claim 1, further comprising a length of an electrical wire electrically coupled with the single piezo-electric crystal, the wire extending from the single piezoelectric crystal and terminating in a connector configured to couple to electrical circuit, the electrical signal provided by the single piezo-electric crystal being communicated through the wire to the connector. 3. The apparatus of claim 2, wherein the electrical circuit further comprises: an amplifier coupled to the connector to receive the electrical signal from the single piezo-electric crystal, the amplifier amplifying the electrical signal to provide an amplified electrical signal; andan audio device that converts the amplified electrical signal to the audible sound corresponding to the flow of fluid through the tubular member. 4. The apparatus of claim 3, wherein the connector is a first connector part, the apparatus further comprising a second connector part configured to electrically connect the first connector part with the amplifier. 5. The apparatus of claim 1, the first end portion further comprising an genetally fixed, enlarged portion having an enlarged cross-section relative to portions of the tubular member on axially opposed sides of the enlarged portion, the enlarged portion being dimensioned and configured to form the first sealing connection. 6. The apparatus of claim 5, the enlarged portion comprising a generally soft, flexible material that is more compliant than the elongate tubular member. 7. The apparatus of claim 5, the second end portion further comprising an inflatable collar that can, upon being inflated, form the second sealing connection. 8. The apparatus of claim 5, further comprising an opening extending through a sidewall of the first end portion of the tub... | - | A61M005/00 A61B005/02 A61F002/06
US06990372 | B2 | 2006-01-24 | Programmable signal analysis device for detecting neurological signals in an implantable device | 2002-04-11 | A programmable signal analysis device for detecting and counting neurological and muscular events in living tissue that incorporates one or more signal threshold levels to identify an event, a counter to count events, and a timer to determine event frequency as events per unit time. The threshold levels and counter/timer operating modes/parameters are digitally programmable to identify and count events, i.e., electrical signals having defined parameters. Such a circuit is of particular use in a microstimulator/sensor which is capable of being injected into living tissue at the site of interest. The microstimulator/sensor is used to stimulate a neural pathway or muscle and/or to block a neural pathway to alleviate pain or block stimulation of a muscle. | 1. A signal analysis device for use within an implantable device, wherein said implantable device is configured at least in art for sensing a biological signal within a patient's body, said signal analysis device comprising: a sealed elongate housing having an axial dimension of less than 60 mm and a lateral dimension of less than 6 mm;sensing circuitry contained within said sealed elongate housing for receiving a biological signal within said implantable device and generating a sensed voltage output in response thereto;event detection circuit contained within said sealed elongate housing for detecting an attribute of said sensed output according to one or more designated criteria and designating said detected attribute as an event, wherein said event detection circuit is alternatively configurable for determining whether an event occurred according to one or according to two or more programmable criteria;an event counter contained within said sealed elongate housing configured for accumulating detected events;a clock counter contained within said sealed elongate housing for accumulating clock pulses;event analysis circuitry contained within said sealed elongate housing for determining a processed value corresponding to said detected events and said accumulated clock pulses; and whereinsaid programmable criteria include at least first and second threshold levels and is further alternatively configurable to determine an event when:said sensed voltage output increases above said first threshold level and then decreases below said first threshold level; orsaid sensed voltage output increases above said first threshold level and then decreases below said first threshold level without exceeding said second threshold level. | 1. A signal analysis device for use within an implantable device, wherein said implantable device is configured at least in art for sensing a biological signal within a patient's body, said signal analysis device comprising: a sealed elongate housing having an axial dimension of less than 60 mm and a lateral dimension of less than 6 mm;sensing circuitry contained within said sealed elongate housing for receiving a biological signal within said implantable device and generating a sensed voltage output in response thereto;event detection circuit contained within said sealed elongate housing for detecting an attribute of said sensed output according to one or more designated criteria and designating said detected attribute as an event, wherein said event detection circuit is alternatively configurable for determining whether an event occurred according to one or according to two or more programmable criteria;an event counter contained within said sealed elongate housing configured for accumulating detected events;a clock counter contained within said sealed elongate housing for accumulating clock pulses;event analysis circuitry contained within said sealed elongate housing for determining a processed value corresponding to said detected events and said accumulated clock pulses; and whereinsaid programmable criteria include at least first and second threshold levels and is further alternatively configurable to determine an event when:said sensed voltage output increases above said first threshold level and then decreases below said first threshold level; orsaid sensed voltage output increases above said first threshold level and then decreases below said first threshold level without exceeding said second threshold level. 2. The signal analysis device of claim 1 wherein said signal analysis device additionally comprises: an average peak detector contained within said sealed elongate housing for determining an average peak value of voltage peaks of said sensed voltage output; and whereinsaid first and second threshold levels are effectively adjusted according to changes in said average peak value. 3. The signal analysis device of claim 1 additionally comprising: a peak detector contained within said sealed elongate housing for determining peaks of said sensed voltage output; and whereinsaid event detection circuitry designates said peaks as events. 4. The signal analysis device of claim 1 wherein said programmable criteria includes at least first and second threshold levels and wherein said signal analysis device additionally comprises: a peak detector contained within said sealed elongate housing for determining peaks of said sensed voltage output; and whereinsaid event detection circuitry designates said peaks as events when said peaks correspond to voltages that are greater than said first threshold level and less than said second threshold level. 5. The signal analysis device of claim 1 wherein said event detection circuitry comprises: first a... | Alfred E. Mann Foundation for Scientific Research | A61B005/04
US07113623 | B2 | 2006-09-26 | Methods and systems for display and analysis of moving arterial tree structures | 2002-10-08 | Methods and systems for reconstruction of a three-dimensional representation of a moving arterial tree structure from a pair of sequences of time varying two-dimensional images thereof and for analysis of the reconstructed representation. In one aspect of the invention, a pair of time varying arteriographic image sequences are used to reconstruct a three-dimensional representation of the vascular tree structure as it moves through a cardiac cycle. The arteriographic image sequences maybe obtained from a biplane imaging system or from two sequences of images using a single plane imaging system. Another aspect of the invention then applies analysis methods and systems utilizing the three-dimensional representation to analyze various kinematic and deformation measures of the moving vascular structure. Analysis results may be presented to the user using color coded indicia to identify various kinematic and deformation measures of the vascular tree. | 1. A method for analysis of dynamic arterial tree structures comprising the steps of: providing a time varying 3D representation of a moving arterial tree structure; anddisplaying analysis of said time varying 3D representation in response to interactive user input,wherein the step of providing a 3D representation includes the steps of:providing a first sequence of time varying 2D arteriograms derived from a first angular view;providing a second sequence of time varying 2D arteriograms derived from a second angular view;reconstructing said time varying 3D representation of said moving arterial tree structure using said first sequence and using said second sequence; andestablishing temporal correspondence of points in said time varying 3D representation the temporal correspondence being computed in accordance with Hamilton's Principle. | 1. A method for analysis of dynamic arterial tree structures comprising the steps of: providing a time varying 3D representation of a moving arterial tree structure; anddisplaying analysis of said time varying 3D representation in response to interactive user input,wherein the step of providing a 3D representation includes the steps of:providing a first sequence of time varying 2D arteriograms derived from a first angular view;providing a second sequence of time varying 2D arteriograms derived from a second angular view;reconstructing said time varying 3D representation of said moving arterial tree structure using said first sequence and using said second sequence; andestablishing temporal correspondence of points in said time varying 3D representation the temporal correspondence being computed in accordance with Hamilton's Principle. 2. The method of claim 1 wherein the step of establishing comprises the step of: providing smoothness constraints,wherein the step of establishing temporal correspondence of points in said time varying 3D representation computes the temporal correspondence in accordance with Hamilton's Principle and using said smoothness constraints. 3. The method of claim 1 wherein the step of displaying analysis comprises the steps of: determining a magnitude of displacement for a selected segment of said time varying 3D representation using the temporal correspondence; anddisplaying said magnitude of displacement on a user display. 4. The method of claim 3 wherein the step of displaying said magnitude of displacement comprises the step of: displaying said magnitude of displacement as color coded magnitudes of displacement at each of a plurality of points along said selected segment. 5. The method of claim 1 wherein the step of displaying analysis comprises the steps of: determining a velocity for a selected segment of said time varying 3D representation using the temporal correspondence; anddisplaying said velocity on a user display. 6. The method of claim 5 wherein the step of displaying said velocity comprises the step of: displaying said velocity as color coded velocity at each of a plurality of points along said selected segment. 7. The method of claim 1 wherein the step of displaying analysis comprises the steps of: determining an acceleration for a selected segment of said time varying 3D representation using the temporal correspondence; anddisplaying said acceleration on a user display. 8. The method of claim 7 wherein the step of displaying said acceleration comprises the step of: displaying said acceleration as color coded acceleration at each of a plurality of points along said selected segment. 9. The method of claim 1 wherein the step of displaying analysis comprises the steps of: determining a curvature for a selected segment of said time varying 3D representation using the temporal correspondence; anddisplaying said curvature on a user display. 10. The method of claim 9 wherein the step of displaying said c... | The Regents of the University of Colorado | G06K009/00 G06T015/00 G09G005/02 A61B005/02
US07113819 | B2 | 2006-09-26 | Method and apparatus for monitoring the condition of a fetus | 2002-04-01 | A method and apparatus for monitoring the condition of a fetus to assess a degree of risk of developing a permanent neurological condition is provided. A signal indicative of a fetal heart is processed to derive a degree of risk of developing a permanent neurological condition. The data indicative of the degree of risk of developing a permanent neurological condition indicates a likelihood that the condition of the fetus belongs to a class in a group of classes, where each class in the group of classes is associated with a pre-defined fetal condition. Data indicative of the degree of risk of developing a permanent neurological condition is then released. In one example, a neural network is used to obtain data indicating the likelihood that the condition of the fetus belongs to a class in the group of classes. | 1. An apparatus suitable for monitoring the condition of a fetus, said apparatus comprising: a) an input for receiving a signal indicative of a fetal heart rate;b) a processing unit coupled to said input, said processing unit being operative for processing said signal indicative of the fetal heart rate to derive data indicative of a degree of risk of developing a permanent neurological condition, the data indicative of the degree of risk of developing a permanent neurological condition indicating a likelihood that the condition of the fetus belongs to a class in a group of classes, the classes in the group of classes being associated with pre-defined fetal conditions characterized at least in part by: i. a measure of base deficit; andii. a presence or absence of new-born encephalopathy;c) an output for releasing the data indicative of the degree of risk of developing a permanent neurological condition. | 1. An apparatus suitable for monitoring the condition of a fetus, said apparatus comprising: a) an input for receiving a signal indicative of a fetal heart rate;b) a processing unit coupled to said input, said processing unit being operative for processing said signal indicative of the fetal heart rate to derive data indicative of a degree of risk of developing a permanent neurological condition, the data indicative of the degree of risk of developing a permanent neurological condition indicating a likelihood that the condition of the fetus belongs to a class in a group of classes, the classes in the group of classes being associated with pre-defined fetal conditions characterized at least in part by: i. a measure of base deficit; andii. a presence or absence of new-born encephalopathy;c) an output for releasing the data indicative of the degree of risk of developing a permanent neurological condition. 2. An apparatus as defined in claim 1, wherein said input is a first input, said apparatus comprising a second input for receiving a signal indicative of uterine activity, said processing unit being coupled to said first and second inputs and being operative for processing said signal indicative of the fetal heart rate and said signal indicative of uterine activity to derive the data indicative of a degree of risk of developing a permanent neurological condition. 3. An apparatus as defined in claim 1, wherein for each class in the group of classes, the data indicates a likelihood value that the condition of the fetus belongs to the corresponding class. 4. An apparatus as defined in claim 1, wherein said group of classes includes at least four classes comprising: a) a first class and a second class associated to respective pre-defined fetal conditions indicative of a presence of new-born encephalopathy; andb) a third class and a fourth class associated to respective pre-defined fetal conditions indicative of an absence of new-born encephalopathy. 5. An apparatus as defined in claim 4, wherein said first class is associated to a pre-determined fetal condition indicative of the presence of new-born encephalopathy and a measure of base deficit above a pre-determined level. 6. An apparatus as defined in claim 4, wherein said second class is associated to a pre-determined fetal condition indicative of the presence of new-born encephalopathy and a measure of base deficit below a pre-determined level. 7. An apparatus as defined in claim 4, wherein said third class is associated to a pre-determined fetal condition indicative of the absence of new-born encephalopathy and a base deficit measurement above a pre-determined level. 8. An apparatus as defined in claim 4, wherein said fourth class is associated to a pre-determined fetal condition indicative of the absence of new-born encephalopathy and a base deficit measurement below a pre-determined level. 9. An apparatus as defined in claim 1, wherein said group of classes includes at least four classes ... | LMS Medical Systems Ltd. | A61B005/04
US07308295 | B2 | 2007-12-11 | Compilation of image information and mammography apparatus for performing biopsy | 2002-04-16 | A mammography apparatus has a bracket having one of its arms fitted with an X-ray tube head. The X-ray tube head is formed of an X-ray tube for producing X-radiation and a collimator for collimating the X-radiation. The bracket arm opposite the X-ray tube head is fitted with an image forming element. The bracket is further provided with a compression means for compressing a breast during an imaging procedure. The image forming element has a digital full field detector. The mammography apparatus has integrable therewith a biopsy device with a biopsy needle. The detector required in a biopsy imaging process for imaging the volume space of biopsy is formed of a desired element of the full field detector. A needle guiding unit is provided for a guided movement of the biopsy needle in the volume space of biopsy on the basis of three-dimensional information compiled in the biopsy imaging process. | 1. A mammography apparatus comprising: a bracket (1, 5, 6) having first and second spaced bracket arms;an X-ray tube head (12) fitted to said first bracket arm (1) and comprising an X-ray tube (3) for producing X-radiation and a collimator (4) for collimating the X-radiation produced by the X-ray tube;compression means (7) for positioning a breast in the path of the collimated X-radiation of the X-ray tube head and for compressing the breast for imaging;a biopsy device (20), integrable with said mammography apparatus proximate to said compression means, said biopsy device having a biopsy needle (21) for carrying out a biopsy in a localized volume space of biopsy within the breast;an image forming component (23) fitted to said second bracket arm (6) to receive the X-radiation of said X-ray tube head passing through the breast, said image forming component comprising a digital full field detector (11) of a size that is sufficiently large to present a single image of the entire breast;means for employing a small section of said full field detector as a detector element when the localized volume space of biopsy is imaged with collimated radiation from said X-ray tube head, said image forming component of said apparatus thus serving as an X-radiation detector for both full scale breast imaging and localized biopsy volume space imaging; anda needle guiding unit (22) for guiding movement of the biopsy needle (21) in said localized volume space of biopsy on the basis of three-dimensional information compiled in the imaging of the localized volume space of biopsy. | 1. A mammography apparatus comprising: a bracket (1, 5, 6) having first and second spaced bracket arms;an X-ray tube head (12) fitted to said first bracket arm (1) and comprising an X-ray tube (3) for producing X-radiation and a collimator (4) for collimating the X-radiation produced by the X-ray tube;compression means (7) for positioning a breast in the path of the collimated X-radiation of the X-ray tube head and for compressing the breast for imaging;a biopsy device (20), integrable with said mammography apparatus proximate to said compression means, said biopsy device having a biopsy needle (21) for carrying out a biopsy in a localized volume space of biopsy within the breast;an image forming component (23) fitted to said second bracket arm (6) to receive the X-radiation of said X-ray tube head passing through the breast, said image forming component comprising a digital full field detector (11) of a size that is sufficiently large to present a single image of the entire breast;means for employing a small section of said full field detector as a detector element when the localized volume space of biopsy is imaged with collimated radiation from said X-ray tube head, said image forming component of said apparatus thus serving as an X-radiation detector for both full scale breast imaging and localized biopsy volume space imaging; anda needle guiding unit (22) for guiding movement of the biopsy needle (21) in said localized volume space of biopsy on the basis of three-dimensional information compiled in the imaging of the localized volume space of biopsy. 2. A mammography apparatus as set forth in claim 1, wherein said X-ray tube head is movable to image the localized volume space of biopsy from at least two different angles for compiling three-dimensional information on the basis of imaging procedures performed from said at least two different angles. 3. A mammography apparatus as set forth in claim 1, wherein the X-ray tube head (12) is adapted to position itself at least at two different angles relative to the full field detector (11). 4. A mammography apparatus as set forth in claim 3, wherein the mammography apparatus comprises means to select said detector element of the full field detector (11) within the field of the full field detector as a function of an angle by which the volume space is imaged relative to the full field detector. 5. A mammography apparatus as set forth in claim 3, wherein the mammography apparatus comprises means to maneuver collimated X-radiation of the X-ray tube head within the field of the full field detector (11) as a function of an angle by which the volume space is imaged relative to the full field detector. 6. A method for performing biopsy of a breast, said method comprising the steps of: compressing the breast for retaining it for the duration of a biopsy procedure in an essentially static position in a mammography apparatus;placing a localized volume space of biopsy in the breast within an imaging fiel... | Instrumentarium Corporation | A61B005/00
US07272433 | B2 | 2007-09-18 | Transcutaneous monitor and method of use, using therapeutic output from an implanted medical device | 2002-04-30 | An external physiologic or implanted device monitor/controller (“monitor”) designed to be attached to the body, as for example, to the skin of the patient. The monitor is designed to detect the therapeutic outputs actually produced by the implanted medical device. Having knowledge of the operation of the implanted medical device, the monitor may then deduce, or decode, the physiologic conditions and/or devices conditions sensed by the implanted medical device. The monitor is then able to perform an action appropriate to the sensed condition and the specific implementation. | 1. A monitoring system for monitoring a patient, comprising: an implantable medical device configurable to be implanted into said patient, said implantable medical device being capable, when implanted, of generating a therapeutic effect to said patient according to known parameters in response to a physiologic condition of said patient, said implantable medical device having a therapeutic output; anda monitor, externally attachable to said patient, containing a sensor for monitoring said therapeutic output, deducing an underlying condition from said therapeutic output and acting in response thereto. | 1. A monitoring system for monitoring a patient, comprising: an implantable medical device configurable to be implanted into said patient, said implantable medical device being capable, when implanted, of generating a therapeutic effect to said patient according to known parameters in response to a physiologic condition of said patient, said implantable medical device having a therapeutic output; anda monitor, externally attachable to said patient, containing a sensor for monitoring said therapeutic output, deducing an underlying condition from said therapeutic output and acting in response thereto. 2. A monitoring system as in claim 1 wherein said therapeutic output is a therapeutic stimulus. 3. A monitoring system as in claim 2 wherein said therapeutic stimulus is an electrical signal. 4. A monitoring system as in claim 3 wherein said implanted medical device is configured to generate said therapeutic stimulus to said patient according to known predetermined parameters in response to said physiologic condition. 5. A monitoring system as in claim 4 wherein said monitor is attachable to said patient with a skin patch electrode. 6. A monitoring system as in claim 5 wherein said skin parch electrode is attached to said skin of said patient through the use of an adhesive. 7. A monitoring system as in claim 1 wherein said monitor acts in response thereto by communicating said underlying condition externally from said monitor using a signaling device. 8. A monitoring system as in claim 7 wherein said signaling device utilizes radiotelemetry communication to an external receiver. 9. A monitoring system as in claim 7 wherein said monitor communicates via a visual indication. 10. A monitoring system as in claim 1 wherein said monitor acts in response thereto by controlling a secondary medical device. 11. A monitoring system as in claim 1 wherein said underlying condition is indicative of said physiologic condition. 12. A monitoring system as in claim 1 wherein said underlying condition is of a condition of said implantable medical device. 13. A monitoring device for monitoring a patient having an implanted medical device configured to generate a therapeutic effect to said patient according to known parameters in response to said physiologic condition of said patient, said implanted medical device having a therapeutic output, comprising a monitor, externally attachable to said patient, a sensor for monitoring said output, deducing an underlying condition from said therapeutic output and acting in response thereto. 14. A monitoring device as in claim 13 wherein said therapeutic output is a therapeutic stimulant. 15. A monitoring device as in claim 14 wherein said patient has skin and wherein said monitor is externally attachable to said skin of said patient. 16. A monitoring device as in claim 15 wherein said monitor is attachable to said patient with a skin patch electrode. 17. A monitoring device as in claim 16 wherein said skin patch ele... | Medtronic, Inc. | A61B005/02 A61N001/18
US07139604 | B1 | 2006-11-21 | Cardiac stimulation system and method for discriminating sinus from non-sinus events | 2002-03-28 | A cardiac stimulation device and method discriminates sinus events from non-sinus events and provide a uniquely prescribed response upon detection of a specific event. | 1. A method of discriminating sinus events from non-sinus events using an implantable cardiac stimulation device, comprising: acquiring a reference signal value for a sinus event;acquiring a reference signal value for a non-sinus event;sampling a sensed signal resulting from an event in a cardiac chamber during a predetermined time period following an event in another cardiac chamber to obtain a characteristic signal value;identifying the sensed signal as a sinus event when the characteristic signal value substantially approximates the sinus-event reference signal value;identifying the sensed signal as a non-sinus event when the characteristic signal value substantially approximates the non-sinus-event reference signal value; andeffecting operation of the implantable cardiac stimulation device in response to the identified sensed signal. | 1. A method of discriminating sinus events from non-sinus events using an implantable cardiac stimulation device, comprising: acquiring a reference signal value for a sinus event;acquiring a reference signal value for a non-sinus event;sampling a sensed signal resulting from an event in a cardiac chamber during a predetermined time period following an event in another cardiac chamber to obtain a characteristic signal value;identifying the sensed signal as a sinus event when the characteristic signal value substantially approximates the sinus-event reference signal value;identifying the sensed signal as a non-sinus event when the characteristic signal value substantially approximates the non-sinus-event reference signal value; andeffecting operation of the implantable cardiac stimulation device in response to the identified sensed signal. 2. The method according to claim 1, further comprising setting the predetermined time period as an atrial refractory period and an atrial alert period following delivery of a ventricular stimulation pulse. 3. The method according to claim 2, wherein identifying the sensed signal as a sinus event comprises comparing a characteristic value of a signal sensed during any of the atrial refractory period or the atrial alert period to a reference signal value for a sinus P-wave. 4. The method according to claim 3, further comprising identifying the sensed signal as a sinus P-wave if the characteristic value of the sensed signal substantially matches a reference signal value for a sinus P-wave. 5. The method according to claim 4, wherein effecting operation of the implantable cardiac stimulation device in response to a detection of a sinus P-wave during the atrial refractory window comprises any of: inhibiting a scheduled atrial stimulation pulse;ignoring the P-wave for ventricular tracking purposes;automatic mode-switching;atrial rate determination for tachycardia detection; oratrial rate determination for dynamic atrial overdrive pacing. 6. The method according to claim 5, wherein effecting operation of the implantable cardiac stimulation device in response to a detection of a sinus P-wave during the atrial alert period comprises any of: inhibiting a scheduled atrial stimulation pulse;ventricular tracking of the P-wave;determining atrial rate for tachycardia detection; ordetermining atrial rate for dynamic atrial overdrive pacing. 7. The method according to claim 6, wherein, if the sensed signal does not substantially match the reference signal value for a sinus P-wave, identifying the sensed signal as a non-sinus event if the characteristic value of the sensed signal approximately matches the reference signal value for a non-sinus event; and providing a non-sinus event response. 8. The method according to claim 7, wherein, if the sensed signal is identified as an ectopic P-wave, the step of providing a non-sinus event response comprises any of: ventricular tracking of the ectopic P-wave;adjusting a P-V interval... | Pacesetter, Inc. | A61B005/04
US06882876 | B2 | 2005-04-19 | Diagnosis of normal pressure hydrocephalus by automated processing of MR images | 2002-03-22 | A method useful in diagnosing normal pressure hydrocephalus and monitoring treatment is presented. A database of mean brain diffusion constants and distribution width is created. Three-dimensional diffusion images of the subject are obtained and algorithms are applied to the images, which include reducing the three-dimensional images to one-dimensional maps of the average diffusion content for the entire brain. Distribution maps (diffusion histograms) are obtained for the brain, and the histograms are fit to a model. From the fit are obtained a mean brain diffusion constant, which is characteristic of the diffusion of the brain tissue in the three compartment model, and a distribution width for the tissue. These numbers are compared to the database for normals, and the significant deviation from the normals (higher in the case of mean brain diffusion constant and wider in the case of distribution width) is an indicator of the presence of normal pressure hydrocephalus. | 1. A method to diagnose normal pressure hydrocephalus comprising the steps of: processing three-dimensional diffusion magnetic resonance images of the entire brain of an individual to produce a diffusion distribution map of the entire brain for the individual; fitting the diffusion distribution map to a model and determining from the model a mean brain diffusion constant for the entire brain of the individual; comparing the mean brain diffusion constant with at least one normal brain diffusion constant for normal individuals to determine if the individual is likely to have normal pressure hydrocephalus or is likely to be a normal individual. | 1. A method to diagnose normal pressure hydrocephalus comprising the steps of: processing three-dimensional diffusion magnetic resonance images of the entire brain of an individual to produce a diffusion distribution map of the entire brain for the individual; fitting the diffusion distribution map to a model and determining from the model a mean brain diffusion constant for the entire brain of the individual; comparing the mean brain diffusion constant with at least one normal brain diffusion constant for normal individuals to determine if the individual is likely to have normal pressure hydrocephalus or is likely to be a normal individual. 2. The method of claim 1 further comprising the step of determining from the model a distribution width and wherein the step of comparing the mean brain diffusion constant with at least one normal brain diffusion constant comprises comparing the mean brain diffusion constant and the distribution width with the at least one normal brain diffusion constant and at least one normal distribution width for normal individuals. 3. The method of claim 2 wherein the at least one normal distribution width comprises a plurality of normal distribution widths stored in a database, the method further comprising the steps of: adding the distribution width to the database if the individual is likely to be a normal individual. 4. The method of claim 1 wherein the at least one normal brain diffusion constant comprises a plurality of normal brain diffusion constants stored in a database, the method further comprising the step of adding the mean brain diffusion constant to the database if the individual is likely to be diagnosed as a normal individual. 5. The method of claim 4 further comprising the steps of: comparing the mean brain diffusion constant and the distribution width with a plurality of normal pressure hydrocephalus brain diffusion constants and a plurality of normal pressure hydrocephalus distribution widths for individuals diagnosed as having normal pressure hydrocephalus that are stored in the database; if the mean brain diffusion constant fits within the plurality of normal pressure hydrocephalus brain diffusion constants and the distribution width fits within the plurality of normal pressure hydrocephalus distribution widths: diagnosing the individual as likely having normal pressure hydrocephalus; adding the mean brain diffusion constant to the plurality of normal pressure hydrocephalus brain diffusion constants; and adding the distribution width to the plurality of normal pressure hydrocephalus distribution widths. 6. The method of claim 1 further comprising the steps of: if the individual is likely to have normal pressure hydrocephalus: obtaining new three-dimensional diffusion magnetic resonance images of the entire brain of the individual after the individual has undergone a spinal test procedure; processing the new three-dimensional diffusion magnetic resonance images to produce a new diffusion di... | Cornell Research Foundation, Inc. | A61B005/05
US06860858 | B2 | 2005-03-01 | Ventilator patient synchronization | 2002-05-22 | A method and apparatus that provides an expert system for determining respiratory phase during ventilatory support of a subject. Discrete phase states are partitioned and prior probability functions and observed probability functions for each state are defined. The probability functions are based upon relative duration of each state as well as the flow characteristics of each state. These functions are combined to determine phase probabilities for each state using Bayes' theorem. The calculated probabilities for the states may then be compared to determine which state the subject is experiencing. A ventilator may then conform respiratory support in accordance with the most probable phase. To provide a learning feature, the probability functions may be adjusted during use to provide a more subject specific response that accounts for changing respiratory characteristics. | 1. A method for determining a subject's respiratory phase during ventilatory support comprising the steps of: identifying a plurality of phase states from a respiration cycle; measuring airflow from a subject during the subject's respiratory cycle; calculating phase-state probabilities for said phase states with a measure of said airflow; and determining a phase in the subject's respiratory cycle from the calculated phase-state probabilities of said phase states. | 1. A method for determining a subject's respiratory phase during ventilatory support comprising the steps of: identifying a plurality of phase states from a respiration cycle; measuring airflow from a subject during the subject's respiratory cycle; calculating phase-state probabilities for said phase states with a measure of said airflow; and determining a phase in the subject's respiratory cycle from the calculated phase-state probabilities of said phase states. 2. The method of claim 1 wherein said calculating step includes the sub-steps of: calculating prior probabilities for said phase states; and calculating observed probabilities for said phase states given said measure of airflow. 3. The method of claim 2 wherein said prior probabilities are functions of a phase state duration relative to respiratory cycle duration. 4. The method of claim 3 wherein said prior probabilities are a further function of previously calculated phase-state probabilities. 5. The method of claim 3 wherein said observed probabilities are defined by a comparing function that relates said phase states to sub-ranges selected from a range of measures of an airflow cycle. 6. The method of claim 5 wherein said comparing function changes as a result of changing airflow cycles of the subject. 7. The method of claim 3 wherein said determining a phase step involves comparing said calculated phase-state probabilities to identify a calculated phase-state probability that is more probable than other calculated phase-state probabilities. 8. The method of claim 3 wherein said identifying step involves partitioning a respiration cycle in relation to a peak flow. 9. The method of claim 3 wherein said identifying step is repeated to adjust said phase states to account for the subject's changing respiratory characteristics as assessed by a measure of said airflow over time. 10. The method of claim 3 wherein said step of calculating phase-state probabilities involves an evaluation of the prior probabilities and the observed probabilities as a function of Bayes' theorem. 11. The method of claim 3 further comprising the step of adjusting ventilatory support based upon results from said determining a phase step. 12. The method of claim 3 wherein said step of calculating phase-state probabilities includes adjusting a phase-state probability based upon state succession. 13. The method of claim 3 wherein said step of determining a phase involves the sub-step of calculating a continuous phase variable from said phase-state probabilities. 14. The method of claim 1 wherein said step of measuring airflow includes the sub-step of smoothing said airflow by performing a low pass filtering operation on said airflow. 15. An apparatus for determining a subject's respiratory phase during ventilatory support comprising: a ventilator to provide a controlled supply of breathable gas to a subject's airway; a pressure transducer for generating an air flow signal representative of respiratory ... | ResMed Limited | A61B005/00
US06932894 | B2 | 2005-08-23 | Biosensor membranes composed of polymers containing heterocyclic nitrogens | 2002-05-14 | Novel membranes comprising various polymers containing heterocyclic nitrogen groups are described. These membranes are usefully employed in electrochemical sensors, such as amperometric biosensors. More particularly, these membranes effectively regulate a flux of analyte to a measurement electrode in an electrochemical sensor, thereby improving the functioning of the electrochemical sensor over a significant range of analyte concentrations. Electrochemical sensors equipped with such membranes are also described. | 1. A membrane for use in an electrochemical sensor, comprising: a crosslinker; and a polymer having the formula: wherein the solid horizontal line represents a polymer backbone; A is an alkyl substituted with a water-soluble constituent; D is selected from a group consisting of phenylalkyl, alkoxystyrene, hydroxyalkyl, alkoxyalkyl, alkoxycarbonylalkyl, a poly(ethylene glycol)-containing constituent, and a polyhydroxyl-containing constituent; and each of n, l, p and q is independently a positive number. | 1. A membrane for use in an electrochemical sensor, comprising: a crosslinker; and a polymer having the formula: wherein the solid horizontal line represents a polymer backbone; A is an alkyl substituted with a water-soluble constituent; D is selected from a group consisting of phenylalkyl, alkoxystyrene, hydroxyalkyl, alkoxyalkyl, alkoxycarbonylalkyl, a poly(ethylene glycol)-containing constituent, and a polyhydroxyl-containing constituent; and each of n, l, p and q is independently a positive number. 2. The membrane of claim 1, wherein A is negatively charged. 3. The membrane of claim 1, wherein A is selected from a group consisting of a sulfonate, a carboxylate, and a phosphate. 4. The membrane of claim 1, wherein A is selected from a group consisting of sulfopropyl, sulfobutyl, carboxypropyl, and carboxypentyl. 5. The membrane of claim 1, wherein A is of the formula L-G, where L is a C2-C12 linear or branched alkyl linker and G is a negatively charged carboxy or sulfonate. 6. The membrane of claim 5, wherein L is substituted with an aryl, alkoxy, alkenyl, alkynyl, —F, —Cl, —OH, aldehyde, ketone, ester, or amide. 7. The membrane of claim 1, wherein D is styrene or C1-C18 alkyl methacrylate. 8. The membrane of claim 1, wherein the polymer has a formula selected from a group consisting of: 9. The membrane of claim 1, wherein an average molecular weight of the polymer is above about 50,000. 10. The membrane of claim 1, wherein an average molecular weight of the polymer is above about 200,000. 11. The membrane of claim 1, wherein an average molecular weight of the polymer is above about 1,000,000. 12. The membrane of claim 1, wherein the polymer further comprises a B-containing copolymer such that the polymer has the formula: wherein B is a modifier and m is a positive number. 13. The membrane of claim 12, wherein B is selected from a group consisting of a chelator, a negatively charged constituent, a hydrophobic hydrocarbon constituent, a hydrophilic hydroxyl or polyhydroxy constituent, a silicon polymer, and a poly(ethylene glycol). 14. The membrane of claim 12, wherein B is a poly(ethylene glycol) having a molecular weight of from about 100 to about 20,000. 15. The membrane of claim 12, wherein the polymer has a formula selected from a group consisting of: 16. The membrane of claim 12, wherein [n/(n+m+l+p)]×100% is from about 1% to about 25%. 17. The membrane of claim 12, wherein [m/(n+m+l+p)]×100% is from about 1% to about 15%. 18. The membrane of claim 12, wherein [l/(n+m+l+p)]×100% is from about 20% to about 90%. 19. The membrane of claim 12, wherein [p/(n+m+l+p)]×100% is from about 1% to about 50%. 20. The membrane of claim 1 or 12, wherein at least one heterocyclic nitrogen constituent of the polymer is independently selected from a group consisting of pyridine, imidazole, oxazole, thiazole, pyrazole, and any derivative thereof. 21. The membrane of claim 1 or 12, wherein at least one heterocyclic nitrogen consti... | TheraSense, Inc. | G01N027/327 A61B005/05 B01D039/00
US06997884 | B2 | 2006-02-14 | System and method for assessing urinary function | 2002-06-27 | A test module is provided for coupling with a control device to form a portable system for assessing urinary function. The test module includes a test module housing, at least one coupling element for removably coupling the test module to the control device, a tubing assembly located at least partially within the test module housing, and a pressure interface in fluid communication with at least a portion of the tubing assembly. When the test module is coupled to the control device, the pressure interface is positioned relative to a pressure sensor in the control device so as to transmit pressure data thereto. The tubing assembly may also couple with a pump device in the control device so as to enable the pump device to pump fluid through the tubing assembly. | 1. A test module for coupling with a control device to form a portable system for assessing urinary function, comprising: a test module housing;at least one coupling element for removably coupling the test module housing to the control device;a tubing assembly including a first fluid conduit extending from a fluid inlet located exterior of the test module housing, to a first portion located entirely within the test module housing, to a second portion that is exposed to an exterior of the test module housing at an upper side thereof, back through the test module housing to a fluid outlet located exterior of the test module housing;a pressure interface in fluid communication with at least a portion of the tubing assembly;wherein when the test module is coupled to the control device, the pressure interface is positioned relative to a pressure sensor in the control device so as to transmit pressure data thereto. | 1. A test module for coupling with a control device to form a portable system for assessing urinary function, comprising: a test module housing;at least one coupling element for removably coupling the test module housing to the control device;a tubing assembly including a first fluid conduit extending from a fluid inlet located exterior of the test module housing, to a first portion located entirely within the test module housing, to a second portion that is exposed to an exterior of the test module housing at an upper side thereof, back through the test module housing to a fluid outlet located exterior of the test module housing;a pressure interface in fluid communication with at least a portion of the tubing assembly;wherein when the test module is coupled to the control device, the pressure interface is positioned relative to a pressure sensor in the control device so as to transmit pressure data thereto. 2. The test module according to claim 1, further comprising an insert member dimensioned for at least partial insertion into a patient's body, the insert member being coupled to the first fluid outlet so that fluid infused through the first fluid conduit passes through the insert member. 3. The test module according to claim 2, wherein the insert member is a meatus plug dimensioned for at least partial insertion into the patient's urethral canal. 4. The test module according to claim 3, wherein the meatus plug has a channel therethrough coupled with the first fluid outlet, and is dimensioned so that, when inserted into the patient's urethral canal distal of the patient's urethral sphincter, it substantially blocks fluid flow into or out of the urethral canal other than through the meatus plug channel. 5. The test module according to claim 2, wherein the insert member is a catheter. 6. The test module according to claim 1, wherein the pressure interface is in fluid communication with the first fluid conduit. 7. The test module according to claim 1, wherein the second portion of the first fluid conduit is configured to engage a pump device in the control device so as to enable the pump device to pump fluid therethrough. 8. The test module according to claim 1, wherein the test module housing has at least one port therein configured to receive therein the pressure sensor. 9. The test module according to claim 8, wherein the pressure information includes Urethral Resistance Pressure information. 10. The test module according to claim 8, wherein the pressure information includes bladder pressure information. 11. The test module according to claim 8 further comprising a second pressure interface, wherein the test module housing further comprises a second port therein configured to receive a second control device pressure sensor. 12. The test module according to claim 11, wherein the second pressure interface is positioned relative to the second port so that when the second control device pressure sensor is received therein the second pre... | Ethicon, Inc. | A61B005/00
US07048692 | B2 | 2006-05-23 | Method and apparatus for estimating auditory filter shape | 2003-01-15 | A method for estimating the shape of an auditory filter by finding a coefficient p of a roex (p, r) filter which is used as a model for the shape of the auditory filter is provided, which comprises the steps of: generating a tone S′; generating a masker of a notch width g and level Nx′ including in a notch the frequency f in which frequency characteristics of the external and middle ears of the subject are considered; transmitting to the subject a sound which superposes the masker on the tone S′; measuring the minimum notch width gx−a of the subject while varying the notch width g; calculating the coefficient p from an upper limit value gmax of the notch width g suitable for the subject and estimating the shape of the auditory filter from the coefficient p calculated above and the value x corresponding to the coefficient r. | 1. A method for estimating the shape of an auditory filter by finding a coefficient p of a roex (p, r) filter which is used as a model for the shape of the auditory filter, comprising the steps of: determining a threshold masking level Nx based on a tone S which adds a given value x to a threshold of hearing T of a subject in frequency f;generating a tone S′ which subtracts a given value a from the tone S;generating a masker of notch width g and a level Nx including the frequency f in a notch;transmitting to the subject an inspection sound which superposes the masker on the tone S′;measuring a minimum notch width gx−a of the subject while varying the notch width g;estimating an upper limit value gmax of the notch width g suitable for the subject from the minimum notch width gx−a;calculating the coefficient p from the upper limit value gmax, the minimum notch width gx−a, the value x corresponding to a coefficient r, and a value a; andestimating the shape of the auditory filter from the coefficient p calculated above and the value x corresponding to the coefficient r. | 1. A method for estimating the shape of an auditory filter by finding a coefficient p of a roex (p, r) filter which is used as a model for the shape of the auditory filter, comprising the steps of: determining a threshold masking level Nx based on a tone S which adds a given value x to a threshold of hearing T of a subject in frequency f;generating a tone S′ which subtracts a given value a from the tone S;generating a masker of notch width g and a level Nx including the frequency f in a notch;transmitting to the subject an inspection sound which superposes the masker on the tone S′;measuring a minimum notch width gx−a of the subject while varying the notch width g;estimating an upper limit value gmax of the notch width g suitable for the subject from the minimum notch width gx−a;calculating the coefficient p from the upper limit value gmax, the minimum notch width gx−a, the value x corresponding to a coefficient r, and a value a; andestimating the shape of the auditory filter from the coefficient p calculated above and the value x corresponding to the coefficient r. 2. The method for estimating the shape of an auditory filter according to claim 1, wherein the step of measuring the minimum notch width gx−a of the subject includes using the value x as a parameter. 3. The method for estimating the shape of an auditory filter according to claim 1 or claim 2, wherein when at least one of the threshold masking level Nx and the minimum notch width gx−a is measured, transmission of the tone is started at predetermined time intervals after starting the transmission of the masker to the subject. 4. An apparatus for estimating the shape of an auditory filter by finding a coefficient p of a roex (p, r) filter which is used as a model for the shape of the auditory filter comprising: a tone generation element for generating a tone of a predetermined frequency;a tone level setting element for amplifying or attenuating the tone generated at the tone generation element to a predetermined level;a noise generation element for generating noise which is not provided with a notch;a noise level setting element for amplifying or attenuating the noise generated at the noise generation element to a predetermined level;a notch width setting element for providing the noise with a notch including the frequency of the tone;a notched noise superposition element for superposing the notched noise output from the notch width setting element on the tone output from the tone level setting element;an inspection sound transmission element for transmitting to a subject an inspection sound output from the notched noise superposition element;an upper limit value calculation element for calculating an upper limit value of the notch width suitable for the subject based on the notch width in the case where the subject can perceive the inspection sound;an auditory filter calculation element for calculating the coefficient p of the roex (p, r) filter based on the notch width and the upper... | Rion Co., Ltd. | A61B005/00
US07223245 | B2 | 2007-05-29 | Method and apparatus for automatic non-cooperative frequency specific assessment of hearing impairment and fitting of hearing aids | 2003-01-29 | A method and device for automatically assessing loss of hearing sensitivity and compression (recruitment) with user defined frequency resolution by means of extrapolated DPOAE I/O functions and ABRs as well as for automatically fitting hearing aids without any cooperation of the subject tested using a device having a display screen attached to a handheld device generating and collecting otoacoustic emission signals and brain stem response signals into a programmed with a clinical audiogram with fitting parameters for hearing aids calculated on the basis of assessed hearing threshold and compression and identifying the type of hearing required for the individual. | 1. An apparatus for automatically assessing loss of hearing sensitivity and loss of compression (recruitment) with user defined frequency resolution by means of extrapolated OAE I/O-functions and ABRs as well as for automatically fitting hearing aids comprising: means for generating one or more stimuli with acoustic transmitters in each ear canal of a subject,signal collecting means for collecting any transient evoked and distortion product otoacoustic emissions generated by the cochlea in response to the stimulus with microphone means placed in the ear canal for generating a frequency mixed product electric signal, and/or brain stem responses via scalp attached electrodes,computer means with computer display for analyzing the response signal distortion product otoacoustic emission input/output functions (DPOAE I/O-functions), which mirror the sensitivity and non-linear compression of cochlear outer hair cells, and/or auditory brainstem responses, especially, high rate click evoked ABRs and amplitude modulated following responses (AMFRs),means for representing neural sound processing in afferent nerve fibers using binomial statistics, different artifact categories by a digital signal processor associated with the signal collecting means,means for displaying the signal responses of the subject,means for employing a hardware interface for hearing instruments (HIPRO box) associated with the signal collecting means allowing the test subject to interact with a computer display of test responses, and calculating hearing aid fitting parameters at selected f2 from DPOAE I/O-functions.means for fitting a hearing aid to supply frequency specific and input level dependent amplification to provide for a subject's needs, andmeans for displaying the results. | 1. An apparatus for automatically assessing loss of hearing sensitivity and loss of compression (recruitment) with user defined frequency resolution by means of extrapolated OAE I/O-functions and ABRs as well as for automatically fitting hearing aids comprising: means for generating one or more stimuli with acoustic transmitters in each ear canal of a subject,signal collecting means for collecting any transient evoked and distortion product otoacoustic emissions generated by the cochlea in response to the stimulus with microphone means placed in the ear canal for generating a frequency mixed product electric signal, and/or brain stem responses via scalp attached electrodes,computer means with computer display for analyzing the response signal distortion product otoacoustic emission input/output functions (DPOAE I/O-functions), which mirror the sensitivity and non-linear compression of cochlear outer hair cells, and/or auditory brainstem responses, especially, high rate click evoked ABRs and amplitude modulated following responses (AMFRs),means for representing neural sound processing in afferent nerve fibers using binomial statistics, different artifact categories by a digital signal processor associated with the signal collecting means,means for displaying the signal responses of the subject,means for employing a hardware interface for hearing instruments (HIPRO box) associated with the signal collecting means allowing the test subject to interact with a computer display of test responses, and calculating hearing aid fitting parameters at selected f2 from DPOAE I/O-functions.means for fitting a hearing aid to supply frequency specific and input level dependent amplification to provide for a subject's needs, andmeans for displaying the results. 2. An apparatus for automatically assessing loss of hearing sensitivity and loss of compression (recruitment) with user defined frequency resolution by means of extrapolated OAE I/O-functions and ABRs as well as for automatically fitting hearing aids without any cooperation of the subject tested according to claim 1, wherein the otoacoustic emissions cochlea scan testing data and brain stem audiologic testing data is inputted into a computer for further analysis for entry into a HIPRO-box to prescribe the required hearing aid for a patient. 3. An apparatus for automatically assessing loss of hearing sensitivity and loss of compression (recruitment) with user defined frequency resolution by means of extrapolated OAE I/O-functions and ABRs as well as for automatically fitting hearing aids without any cooperation of the subject tested according to claim 1, including switching means selectively activating measuring equipment means which perform: DPOAE Cochlea-Scans,ABR Cochlea-Scans,Automatic Non Cooperative (ANC) Hearing Aid Fitting,Screening,ABR-Inter Peak Latency assessment, (IPL) andTypanometry differentiation of middle ear, cochlear, and neural disorders. 4. An apparatus for automatically assessing los... | NATUS Medical, Inc. | A61B005/00 A61B005/04 A61B005/12 C21B007/24 G01H001/00 G01N029/00 G01N029/04 G01N003/32 G01N009/24 G01S003/80
US07003340 | B2 | 2006-02-21 | Electrochemical analyte sensor | 2002-11-11 | An electrochemical analyte sensor formed using conductive traces on a substrate can be used for determining and/or monitoring a level of analyte in in vitro or in vivo analyte-containing fluids. For example, an implantable sensor may be used for the continuous or automatic monitoring of a level of an analyte, such as glucose, lactate, or oxygen, in a patient. The electrochemical analyte sensor includes a substrate and conductive material disposed on the substrate, the conductive material forming a working electrode. In some sensors, the conductive material is disposed in recessed channels formed in a surface of the sensor. An electron transfer agent and/or catalyst may be provided to facilitate the electrolysis of the analyte or of a second compound whose level depends on the level of the analyte. A potential is formed between the working electrode and a reference electrode or counter/reference electrode and the resulting current is a function of the concentration of the analyte in the body fluid. | 1. A transcutaneous electrochemical sensor comprising: a substrate having a proximal end portion having a width defined by a first side and a second side, and a distal end portion having a width defined by a third side and a fourth side, the width being less than the width of the proximal end portion, the distal end portion of the distal end portion being adjacent to the proximal end portion and extending therefrom;at least one contact pad located on the proximal end portion; andat least one conductive trace located on the distal end portion and the proximal end portion such that the at least one conductive trace electrically communicates with the at least one contact pad,wherein the first side and third side together form a continuous straight edge, and wherein the second side and the fourth side do not form a continuous straight edge. | 1. A transcutaneous electrochemical sensor comprising: a substrate having a proximal end portion having a width defined by a first side and a second side, and a distal end portion having a width defined by a third side and a fourth side, the width being less than the width of the proximal end portion, the distal end portion of the distal end portion being adjacent to the proximal end portion and extending therefrom;at least one contact pad located on the proximal end portion; andat least one conductive trace located on the distal end portion and the proximal end portion such that the at least one conductive trace electrically communicates with the at least one contact pad,wherein the first side and third side together form a continuous straight edge, and wherein the second side and the fourth side do not form a continuous straight edge. 2. The sensor of claim 1 having a length between the proximal end portion and the distal end portion of 0.3 to 5 cm. 3. The sensor of claim 1, wherein the distal end portion has a length of 0.25 to 2 cm. 4. The sensor of claim 1, wherein the proximal end portion has a length of 0.05 to 3 cm. 5. The sensor of claim 1, wherein the width of the proximal end portion between the first side and the second side is 0.5 mm to 15 mm. 6. The sensor of claim 1, wherein the width of the distal end portion between the third side and the fourth side is 2 mm or less. 7. The sensor of claim 1 wherein the width of the proximal end portion between the first side and second side is at least about three times greater than the width of the distal end portion between the third side and fourth side. 8. The sensor of claim 1 wherein the width of the proximal end portion between the first side and second side is at least about six times greater than the width of the distal end portion between the third side and fourth side. 9. The sensor of claim 1 wherein the proximal end portion has a length and the width of the proximal end portion is at least about twenty percent greater than the length. 10. The sensor of claim 1 wherein the proximal end portion and distal end portion each have a length, the length of the proximal end portion being at least about ninety percent of the length of the distal end portion. 11. The sensor of claim 1 wherein at least one of the third side and fourth side of the distal end portion has a barb. 12. A transcutaneous electrochemical sensor comprising: a thin flexible substrate having a generally rectangular proximal end portion having a first side and a second side and a generally rectangular distal end portion having a third side and a fourth side, the distal end portion being adjacent to the proximal end portion and extending therefrom;a plurality of contact pads located within the proximal end portion;a plurality of conductive traces located on the distal end portion and the proximal end portion such that the conductive traces electrically communicate with the contact pads; anda sensing layer adjacen... | Abbott Diabetes Care Inc. | A61B005/05 A61B005/04
US06988992 | B2 | 2006-01-24 | Blood pressure cuffs with resilient support sleeves | 2002-11-12 | Inflatable blood pressure cuff assemblies include an inflatable elongate cuff member having opposing long edges and opposing short edge portions with a fluid chamber therein and a resilient support sleeve attached to a respective one of the opposing short edge portions of the inflatable elongate cuff member. The sleeve may be configured with a sensor chamber and a cable routing channel. Related systems and methods of fabricating cuff assemblies and resilient support sleeves are also described. | 1. An inflatable blood pressure cuff assembly comprising: an inflatable elongate cuff member having opposing long edges and opposing short edge portions with an inflatable fluid chamber therein; anda resilient sleeve attached to a minor portion of the cuff member proximate a single one of the opposing short edge portions of the inflatable elongate cuff member, wherein the sleeve has a body with at least a major portion of the body sized and configured to elastically expand to snugly and generally conformably fit on a limb of a patient when in position on the patient prior to inflation of the cuff member, and wherein, in position, the sleeve is configured to define a first closed member before the cuff member is wrapped over the sleeve and the cuff member defines a second closed member after the cuff member is wrapped over the sleeve. | 1. An inflatable blood pressure cuff assembly comprising: an inflatable elongate cuff member having opposing long edges and opposing short edge portions with an inflatable fluid chamber therein; anda resilient sleeve attached to a minor portion of the cuff member proximate a single one of the opposing short edge portions of the inflatable elongate cuff member, wherein the sleeve has a body with at least a major portion of the body sized and configured to elastically expand to snugly and generally conformably fit on a limb of a patient when in position on the patient prior to inflation of the cuff member, and wherein, in position, the sleeve is configured to define a first closed member before the cuff member is wrapped over the sleeve and the cuff member defines a second closed member after the cuff member is wrapped over the sleeve. 2. A blood pressure cuff assembly according to claim 1, wherein the sleeve is air permeable and comprises a fabric that includes stretch fibers, and wherein the sleeve is configured to generally radially elastically compress against the limb of the patient under the cuff member when the cuff member is inflated. 3. A blood pressure cuff assembly according to claim 1, wherein the sleeve has a closed perimeter configuration defining an aperture extending in the axial direction, wherein the sleeve aperture is sized and configured to stretch to receive a limb of a patient therein during use. 4. A blood pressure cuff assembly according to claim 3, wherein the sleeve aperture has a first configuration with a first width during periods of non-use and a second configuration with an expanded second width when in position on a patient, wherein, in the second configuration, the sleeve is substantially conformable to and resides securely against a desired portion of the limb of the patient with sufficient compressive force so that it is able to maintain its desired longitudinal position to thereby inhibit slippage during use. 5. A blood pressure cuff assembly according to claim 2, wherein the sleeve comprises nylon fibers. 6. A blood pressure cuff assembly according to claim 2, wherein the sleeve comprises spandex fibers. 7. A blood pressure cuff assembly according to claim 4, wherein the sleeve has opposing upper and lower edge portions, and wherein, in position on a patient, the sleeve is configured to have an elastic lateral stretch of at least about 15% at the lower edge portion to provide the second configuration width. 8. A blood pressure cuff assembly according to claim 7, wherein the sleeve is configured to accommodate patients having limbs that vary in width by up to at least about 150%. 9. A blood pressure cuff assembly according to claim 8, wherein the sleeve is sized and configured to accommodate patients having limbs that vary in width between about 100–205%. 10. A blood pressure cuff assembly according to claim 1, wherein the sleeve has a frustoconical shape. 11. A blood pressure cuff assembly according to... | SunTech Medical, Inc. | A61B005/02
US07052467 | B2 | 2006-05-30 | Stethoscopic systems and methods | 2002-10-04 | The invention provides a stethoscope device comprising a chest piece, an amplifier and a mixer compatible with a conventional voice quality telephone system. Electrical signals are generated that are representative of physiological sounds such as breath sounds, heart sounds and bowel sounds. The electrical signals can be combined with voice communications at the remote site and transmitted over the telephone system. The invention also provides a system and method for diagnosis by a trained health care professional with the patient being located at a remote site. In another embodiment, the invention also provides a system and method for training students to interpret stethoscope sounds. In addition, the system allows for simultaneous listening by the health care professional and another individual at the remote location of the stethoscope. | 1. A stethoscopic system comprising: a stethoscopic device having a chest piece that collects a first signal representative of physiological sounds from a subject;an amplifier that amplifies the signal collected by the stethoscopic chest piece to result in an amplified signal, wherein the amplified signal has the same frequency spectrum as the first signal;a mixer that combines the amplified signal and an output of a telephone compatible ear piece and mouthpiece, wherein the mixer output signal has the same frequency spectrum as the first signal; andan operative link to a telephone system, wherein the mixer output signal indicative of physiological sounds is accessed over the telephone system. | 1. A stethoscopic system comprising: a stethoscopic device having a chest piece that collects a first signal representative of physiological sounds from a subject;an amplifier that amplifies the signal collected by the stethoscopic chest piece to result in an amplified signal, wherein the amplified signal has the same frequency spectrum as the first signal;a mixer that combines the amplified signal and an output of a telephone compatible ear piece and mouthpiece, wherein the mixer output signal has the same frequency spectrum as the first signal; andan operative link to a telephone system, wherein the mixer output signal indicative of physiological sounds is accessed over the telephone system. 2. The stethoscopic system of claim 1 wherein the chest piece comprises an acoustic transducer operatively coupled to a microphone by an acoustic conduit. 3. The stethoscopic system of claim 2 wherein the microphone comprises an electret microphone. 4. The stethoscopic system of claim 1 wherein the chest piece comprises a piezoelectric transducer. 5. The stethoscopic system of claim 1 wherein the chest piece comprises an electromagnetic transducer. 6. The stethoscopic system of claim 1 further comprising a telephone at a site remote from the subject being accessed. 7. The stethoscopic system of claim 1 wherein the operative link further comprises at least one wireless link. 8. The stethoscopic system of claim 7 wherein the at least one wireless link comprises a radio frequency link. 9. The stethoscopic system of claim 1 wherein the operative link comprises a wired connection to the telephone system. 10. The stethoscopic system of claim 1 wherein the signals indicative of physiological sounds comprises signals associated with respiration and cardiac activity collected from lungs, trachea and heart regions of the subject. 11. A method of listening to acoustic signals indicative of a physiological condition produced by a subject at a first site comprising the steps of: placing a stethoscopic chest piece in contact with the body surface of the subject producing sounds indicative of the physiological condition at the first site;converting the sounds into a first electrical signal representative of the physiological sounds;amplifying the first electrical signal, wherein the amplified signal has the same frequency spectrum as the first electrical signal;combining the amplified electrical signal with electrical signals received and transmitted by a telephone handset, wherein the combined signal has the same frequency spectrum as the first electrical signal;transmitting the combined signal via an operative link to a telephone system;accessing the telephone system at a second site using conventional voice telephone equipment; andlistening to the sounds indicative of the physiological condition produced by the subject at the first site to diagnose a physiological condition. 12. The method of claim 11 wherein the acoustic signals comprise signals associated ... | Shawn D. Johnson David S. Geller | A61B005/02
US07094209 | B2 | 2006-08-22 | Method for introducer replacement | 2003-01-14 | The invention provides a method for replacing a first introducer, whose distal end is inside a vessel and whose proximal end is outside the skin of a patient, with a second introducer. The method comprises the steps of inserting a guide rod into the first introducer, removing the first introducer, passing the second introducer over the guide rod, and removing the guide rod. The guide rod to be used according to the method has preferably a flexible distal portion with a blunt distal end and a stiff proximal portion with a tapered proximal end. | 1. A method for replacing a first introducer, whose distal end is inside a vessel and whose proximal end is outside the skin of a patient, with a second introducer, said method comprising the steps of: inserting a guide rod into the first introducer until the distal end of the guide rod is inside the vessel,removing the first introducer, thereby leaving only the guide rod in place such that the guide rod plugs a hole in the vessel such that there is substantially no flow of blood out from the vessel,passing the second introducer over the guide rod until the distal end of the second introducer is inside the vessel, andremoving the guide rod, thereby leaving only the second introducer in place. | 1. A method for replacing a first introducer, whose distal end is inside a vessel and whose proximal end is outside the skin of a patient, with a second introducer, said method comprising the steps of: inserting a guide rod into the first introducer until the distal end of the guide rod is inside the vessel,removing the first introducer, thereby leaving only the guide rod in place such that the guide rod plugs a hole in the vessel such that there is substantially no flow of blood out from the vessel,passing the second introducer over the guide rod until the distal end of the second introducer is inside the vessel, andremoving the guide rod, thereby leaving only the second introducer in place. 2. The method according to claim 1, wherein the guide rod is solid. 3. The method according to claim 2, wherein a distal portion of the guide rod has a blunt end. 4. The method according to claim 2, wherein a proximal portion of the guide rod has a tapered end. 5. The method according to claim 2, wherein a distal portion of the guide rod can be pre-bent in such a way that it resumes its bent shape after having passed through the first introducer. 6. The method according to claim 2, wherein a diameter of the guide rod is approximately the same size as the hole in a wall of the vessel. 7. The method according to claim 2, wherein the guide rod further comprises a measurement scale. 8. The method according to claim 2, wherein the guide rod comprises a flexible distal portion and a stiff proximal portion. 9. A method for replacing a first introducer, whose distal end is inside a vessel and whose proximal end is outside the skin of a patient, with a second introducer, said method comprising the steps of: inserting a guide rod into the first introducer until the distal end of the guide rod is inside the vessel,removing the first introducer, thereby leaving the guide rod in place such that the guide rod plugs a hole in the vessel such that there is substantially no flow of blood out from the vessel,passing the second introducer over the guide rod until the distal end of the second introducer is inside the vessel, andremoving the guide rod, thereby leaving the second introducer in place. | Radi Medical Systems AB | A61B005/00 A61M025/00
US07252640 | B2 | 2007-08-07 | Detection of disordered breathing | 2002-12-04 | Devices and methods for detecting disordered breathing involve determining that the patient is asleep and sensing one or more signals associated with disordered breathing indicative of sleep-disordered breathing while the patient is asleep. Sleep-disordered breathing is detected using the sensed signals associated with disordered breathing. The sensed signals associated with disordered breathing may also be used to acquire a respiration pattern of one or more respiration cycles. Characteristics of the respiration pattern are determined. The respiration pattern is classified as a disordered breathing episode based on the characteristics of the respiration pattern. One or more processes involved in the detection of disordered breathing are performed using an implantable device. | 1. A method for detecting sleep-disordered breathing of a patient, comprising: detecting a respiration pattern of one or more respiration cycles;determining characteristics of the respiration pattern, comprising: determining one or more breath intervals and one or more tidal volume measurements of the one or more respiration cycles;identifying an ongoing disordered breathing event as one or both of an apnea event and a hypopnea event based on comparison of the breath intervals to a breath interval threshold or the tidal volume measurements to a tidal volume threshold; anddetermining a duration of the ongoing disordered breathing event; andif the ongoing disordered breathing duration exceeds an event duration threshold, discriminating between apnea, hypopnea and mixed apnea and hypopnea based on the one or more characteristics of the respiration pattern, wherein at least one of detecting, determining and discriminating is performed at least in part implantably. | 1. A method for detecting sleep-disordered breathing of a patient, comprising: detecting a respiration pattern of one or more respiration cycles;determining characteristics of the respiration pattern, comprising: determining one or more breath intervals and one or more tidal volume measurements of the one or more respiration cycles;identifying an ongoing disordered breathing event as one or both of an apnea event and a hypopnea event based on comparison of the breath intervals to a breath interval threshold or the tidal volume measurements to a tidal volume threshold; anddetermining a duration of the ongoing disordered breathing event; andif the ongoing disordered breathing duration exceeds an event duration threshold, discriminating between apnea, hypopnea and mixed apnea and hypopnea based on the one or more characteristics of the respiration pattern, wherein at least one of detecting, determining and discriminating is performed at least in part implantably. 2. The method of claim 1, further comprising determining that the patient is asleep. 3. The method of claim 1, wherein at least two of detecting, determining, and classifying is performed at least in part implantably. 4. The method of claim 1, wherein at each of detecting, determining, and classifying is performed at least in part implantably. 5. The method of claim 1, wherein detecting a respiration pattern comprises measuring transthoracic impedance. 6. The method of claim 1, further comprising establishing one or both of the breath interval threshold or the tidal volume threshold for a particular patient. 7. The method of claim 6, wherein establishing the breath interval threshold or the tidal volume threshold comprises modifying the breath interval threshold or the tidal volume threshold based on a condition of the patient. 8. The method of claim 7, wherein the condition of the patient comprises a cardiopulmonary disease condition. 9. The method of claim 6, further comprising adapting a sensitivity of the breath interval threshold or the tidal volume threshold based on one or more previously detected episodes of sleep-disordered breathing. 10. The method of claim 9, wherein adapting the sensitivity of the breath interval threshold or the tidal volume threshold comprises adapting the sensitivity based on a frequency of the one or more previously detected episodes of sleep-disordered breathing. 11. The method of claim 9, wherein adapting the sensitivity of the breath interval threshold or the tidal volume threshold comprises adapting the sensitivity based on a severity of the one or more previously detected episodes of sleep-disordered breathing. 12. The method of claim 9, wherein adapting the sensitivity of the breath interval threshold or the tidal volume threshold comprises adapting the sensitivity based on an interval between two or more previously detected episodes of sleep-disordered breathing. 13. The method of claim 1, wherein: determining the one or more characteristi... | Cardiac Pacemakers, Inc. | A61B005/02
US07160258 | B2 | 2007-01-09 | Capsule and method for treating or diagnosing the intestinal tract | 2001-06-26 | A device and method for mapping, diagnosing and treating the intestinal tract is provided using a capsule passing through the intestinal tract. Further, a capsule tracking system is provided for tracking a capsule's location along the length of an intestinal tract as various treatment and/or sensing modalities are employed. In one variation, an acoustic signal is used to determine the location of the capsule. A map of sensed information may be derived from the pass of a capsule. Capsules may be subsequently passed through to treat the intestinal tract at a determined location along its length. One variation uses an electrical stimulation capsule to treat and/or diagnose a condition in the intestinal tract. | 1. A system for treating or diagnosing the intestinal tract comprising: an autonomous capsule comprising an acoustic transducer, the transducer being arranged to emit an acoustic signal detectable externally of a patient's body as the capsule passes through at least a portion of the intestinal tract;at least one external acoustic receiver configured to sense the acoustic signal transmitted by the capsule; anda reference generator arranged to generate a time reference from which the time of the acoustic signal generation is determined. | 1. A system for treating or diagnosing the intestinal tract comprising: an autonomous capsule comprising an acoustic transducer, the transducer being arranged to emit an acoustic signal detectable externally of a patient's body as the capsule passes through at least a portion of the intestinal tract;at least one external acoustic receiver configured to sense the acoustic signal transmitted by the capsule; anda reference generator arranged to generate a time reference from which the time of the acoustic signal generation is determined. 2. The system of claim 1 wherein the system comprises a plurality of external acoustic receivers configured to sense the acoustic signal transmitted by the capsule. 3. The system of claim 1 wherein the reference is a trigger signal arranged to trigger generation of the acoustic signal. 4. The system of claim 1 wherein the capsule comprises a plurality of acoustic transducers, each of the plurality of transducer being arranged to emit an acoustic signal detectable by the at least one acoustic receiver as the capsule passes through at least a portion of the intestinal tract, to provide information from which the orientation of the capsule may be derived. 5. The system of claim 1 wherein the at least one external acoustic receiver comprises a transducer for converting the acoustic signal to a representative electrical signal and an output for communicating the electrical signal. 6. The system of claim 5 further comprising: a processor configured to communicate with the output of the at least one acoustic receiver, wherein the processor is arranged to determine location of the capsule on a coordinate system based at least in part on the electrical signal of the at least one acoustic receiver. 7. The system of claim 6 wherein the reference generator arranged to generate a reference signal at a predetermined time interval from transmission of the acoustic signal and at least one external reference signal receiver, wherein the at least one external reference receiver is arranged to receive the reference signal and communicate the reference signal to the processor,wherein the processor is arranged to use the reference signal to determine transmission time of the acoustic signal from the capsule to the at least one acoustic receiver and to determine the location of the capsule on the coordinate system. 8. The system of claim 7 wherein the reference signal generator is an RF telemetry device. 9. The system of claim 6 wherein the processor comprises a tracking system arranged to determine a location of the capsule along a length of the portion of the intestinal tract. 10. The system of claim 9 wherein the location along the length is based on the determination of a plurality of locations of the capsule as the capsule passes through the portion of the intestinal tract. 11. The system of claim 1 wherein the capsule further comprises a sensor for sensing a characteristic of the intestinal tract adjacent the location of ... | Entrack, Inc. | A61B005/117 A61B005/103 A61B005/05
US06879853 | B2 | 2005-04-12 | Method and apparatus for magnetic resonance arteriography using contrast agents | 2003-01-24 | The present invention is a technique of, and system for, imaging vascular anatomy over distance considerably greater than the maximum practical field of view of a magnetic resonance imaging system while using substantially one contrast agent injection. The technique and system of the present invention acquires image data of a plurality of image volumes which are representative of different portions of the patient's body. The image data of each image volume includes image data which is representative of the center of k-space. The acquisition of image data which is representative of the center of k-space is correlated with a concentration of contrast agent in the artery(ies) residing in the image volume. This provides preferential enhancement of arteries relative to adjacent veins and background tissue for each acquisition, wherein each acquisition is representative of a different portion of the arterial system (e.g., abdominal aorta, femoral, popliteal, and tibial arteries). | 1. A method of imaging arteries in a human patient using a magnetic resonance imaging system and a magnetic resonance contrast agent, wherein the arteries include a first artery which is located in a first image volume, a second artery which is located in a second image volume and a third artery which is located in a third image volume, the method comprising: injecting the human patient with the magnetic resonance contrast agent via substantially one injection; collecting magnetic resonance image data of the first image volume, wherein at least a portion of the image data of the first image volume is collected during an arterial phase of contrast enhancement of the first artery and wherein the first artery is in the abdomen or the pelvis of the human patient; collecting magnetic resonance image data of the second image volume, wherein at least a portion of the image data of the second image volume is collected during an arterial phase of contrast enhancement of the second artery and wherein the second artery is in a leg of the human patient; and collecting magnetic resonance image data of the third image volume, wherein at least a portion of the image data of the third image volume is collected during an arterial phase of contrast enhancement of the third artery and wherein the third artery is in a foot of the human patient. | 1. A method of imaging arteries in a human patient using a magnetic resonance imaging system and a magnetic resonance contrast agent, wherein the arteries include a first artery which is located in a first image volume, a second artery which is located in a second image volume and a third artery which is located in a third image volume, the method comprising: injecting the human patient with the magnetic resonance contrast agent via substantially one injection; collecting magnetic resonance image data of the first image volume, wherein at least a portion of the image data of the first image volume is collected during an arterial phase of contrast enhancement of the first artery and wherein the first artery is in the abdomen or the pelvis of the human patient; collecting magnetic resonance image data of the second image volume, wherein at least a portion of the image data of the second image volume is collected during an arterial phase of contrast enhancement of the second artery and wherein the second artery is in a leg of the human patient; and collecting magnetic resonance image data of the third image volume, wherein at least a portion of the image data of the third image volume is collected during an arterial phase of contrast enhancement of the third artery and wherein the third artery is in a foot of the human patient. 2. The method of claim 1 further including constructing an image of the first artery using the image data of the first image volume and an image of the second artery using the image data of the second image volume. 3. The method of claim 2 further including combining the image of the first artery with the image of the second artery in a mosaic manner. 4. The method of claim 1 further including applying a tourniquet to the patient to reduce blood flow in a calf or a thigh of the human patient. 5. The method of claim 1 wherein collecting the magnetic resonance image data of the second image volume includes collecting image data which is representative of the center of k-space substantially at a beginning of a second imaging sequence. 6. The method of claim 1 further including administering the magnetic resonance contrast agent to the patient over a substantial portion of the period of collection of the image data of the first image volume. 7. A method of imaging arteries in a patient using a magnetic resonance imaging system, wherein the arteries include a first artery which is located in a first image volume and a second artery which is located in a second image volume, the method comprising: injecting a magnetic resonance contrast agent of substantially one injection into the patient, wherein the magnetic resonance contrast agent is injected using, a coiled conduit having a first end, a second end and an inner diameter; a multiport adapter connected to the first end of the coiled conduit, the multiport adapter having at least two input ports wherein each input port includes a one way valve; and a locking mechanism cou... | - | A61B005/055
US06880701 | B2 | 2005-04-19 | Storage container for at least one hypodermic needle | 2002-11-07 | The connecting end of this storage container for at least one hypodermic needle has connectors complementary to the connectors on the tip of an injection instrument, one of the connectors having radial elasticity and devices for converting an axial force exerted between the needle and the injection instrument into at least one radial component which can deform the elastic connectors. The container has a housing fore the needle designed to hold this needle in a given position and an opening giving access to the housing. The shape and dimensions of the exterior wall bounding the housing are selected to create with the needle an ergonomic positioning and/or handling interface so as to facilitate the connection of the complementary connectors of the needle and the tip, respectively. | 1. A storage container for a hypodermic needle, the hypodermic needle having a connecting end which has connectors complementary with connectors on the tip of an injection instrument, one of the connectors having radial elasticity and devices for converting an axial force exerted between the needle and the injection instrument into at least one radial component capable of deforming the elastic connectors, the container having a housing designed for holding the needle in a determined position, and an opening giving access to the housing; wherein an exterior wall bounding the housing has a cylindrical guiding surface coaxial with the longitudinal axis of the housing and designed to engage with a complementary guiding surface of the instrument so that the engagement of these guiding surfaces with each other allows the complementary connectors to be put in the assembling position, following the exertion on the devices for converting an axial force into at least one radial component of a pressure coaxial with the longitudinal axis of the housing, wherein the housing holding the hypodermic needle in a determined position has two coaxial openings at a distance from each other, the first of these openings having a fixed diameter approximately equal to the diameter of the connecting end, while the second opening is delimited by elastically radial devices so as to allow its diameter to vary between the minimum diameter and the maximum diameter, and in that a ring having an internal diameter approximately equal to that of the first opening grips the connecting end while projecting from its rear face, thus keeping back the locking elements, this ring having, at a distance from its projecting end equal to the distance between the two said openings, passages whose dimensions allow the locking elements to pass through them and to project within the ring, so that, once the complementary connectors are connected together displacing the injection instrument outwards along its axis causes the hypodermic needle and the ring to shift until the ring rests against the rim of the first opening, in which position the locking elements pass through the passages and the ring separates from the connecting end, while the reinsertion of the connecting end into the housing causes this connecting end to be held in position by the locking elements. | 1. A storage container for a hypodermic needle, the hypodermic needle having a connecting end which has connectors complementary with connectors on the tip of an injection instrument, one of the connectors having radial elasticity and devices for converting an axial force exerted between the needle and the injection instrument into at least one radial component capable of deforming the elastic connectors, the container having a housing designed for holding the needle in a determined position, and an opening giving access to the housing; wherein an exterior wall bounding the housing has a cylindrical guiding surface coaxial with the longitudinal axis of the housing and designed to engage with a complementary guiding surface of the instrument so that the engagement of these guiding surfaces with each other allows the complementary connectors to be put in the assembling position, following the exertion on the devices for converting an axial force into at least one radial component of a pressure coaxial with the longitudinal axis of the housing, wherein the housing holding the hypodermic needle in a determined position has two coaxial openings at a distance from each other, the first of these openings having a fixed diameter approximately equal to the diameter of the connecting end, while the second opening is delimited by elastically radial devices so as to allow its diameter to vary between the minimum diameter and the maximum diameter, and in that a ring having an internal diameter approximately equal to that of the first opening grips the connecting end while projecting from its rear face, thus keeping back the locking elements, this ring having, at a distance from its projecting end equal to the distance between the two said openings, passages whose dimensions allow the locking elements to pass through them and to project within the ring, so that, once the complementary connectors are connected together displacing the injection instrument outwards along its axis causes the hypodermic needle and the ring to shift until the ring rests against the rim of the first opening, in which position the locking elements pass through the passages and the ring separates from the connecting end, while the reinsertion of the connecting end into the housing causes this connecting end to be held in position by the locking elements. 2. A container according to claim 1, wherein the shape and dimensions of the exterior wall bounding the housing are selected to create, with the needle, an ergonomic positioning or handling interface, in order to facilitate the connection of the complementary connectors on the needle and on the tip, respectively. 3. A container according to claim 2, wherein a number of housings are arranged side by side on a shared stand. 4. A container according to claim 2, wherein the wall of the housing is of a tubular form and that in it a number of hypodermic needles are placed successively in equivalent positions so as to be accessible to the... | Ares Trading S.A. | B65D083/10 A61B005/00 A61M005/32
US07252643 | B2 | 2007-08-07 | Guide wire | 2003-01-28 | A guide wire includes a wire body, a first coil, a second coil, and an elongate member. The first coil has a helical shape, and is disposed on a distal end side of the guide wire. The second coil has a helical shape, and is disposed on a distal end side of the first coil. The elongate member has a section that extends a longitudinal direction of the wire body. At least a portion of the section is located inside the second coil, and the elongate member is formed integrally with the first coil or the second coil. | 1. A guide wire comprising: a wire body;a first coil having a helical shape, and being disposed on a distal end side of said wire body;a second coil having a helical shape, and being disposed on a distal end side of said first coil; andan elongate member having a section that extends in a longitudinal direction of said wire body, and made of a reshapeable material, said material being the same as a material of one of said first coil and said second coil,wherein at least a portion of said section is located inside said second coil, and said elongate member and said one of said first coil or said second coil are a single unitary part,wherein a distal end of said wire body is inserted into an inside of said first and second coils substantially at an internal center thereof, andwherein a distal end of said elongate member is fixed on said second coil. | 1. A guide wire comprising: a wire body;a first coil having a helical shape, and being disposed on a distal end side of said wire body;a second coil having a helical shape, and being disposed on a distal end side of said first coil; andan elongate member having a section that extends in a longitudinal direction of said wire body, and made of a reshapeable material, said material being the same as a material of one of said first coil and said second coil,wherein at least a portion of said section is located inside said second coil, and said elongate member and said one of said first coil or said second coil are a single unitary part,wherein a distal end of said wire body is inserted into an inside of said first and second coils substantially at an internal center thereof, andwherein a distal end of said elongate member is fixed on said second coil. 2. A guide wire comprising: a wire body;a first coil having a helical shape, and being disposed on a distal end side of said wire body;a second coil having a helical shape, and being disposed on a distal end side of said first coil; andan elongate member having a section that extends in a longitudinal direction of said wire body, and made of a reshapeable material, said material being the same as a material of one of said first coil and said second coil,wherein at least a portion of said section is located inside said second coil, and said elongate member and said one of said first coil or said second coil are a single unitary part, andwherein a portion of said elongate member in a longitudinal direction branches off. 3. A guide wire comprising: a wire body;a first coil having a helical shape, and being disposed on a distal end side of said wire body;a second coil having a helical shape, and being disposed on a distal end side of said first coil; anda plurality of elongate members each having a section that extends in a longitudinal direction of said wire body, and made of a reshapeable material, said material being the same as a material of one of said first coil and said second coil,wherein at least a portion of said section is located inside said second coil, and said plurality of elongate members and said one of said first coil or said second coil are a single unitary part. 4. A guide wire comprising: a wire body;a helical coil having a distal end and a proximal end and being disposed on a distal end side of said wire body; andan elongate member that extends from the proximal end of said coil, is disposed inside said coil, and is made of a material that can be reshaped, said material being the same as a material of the coil, said elongate member and said coil being a single unitary part,in which a distal end of said elongate member is fixed internally at a central portion of said coil. 5. A guide wire comprising: a wire body;a first coil having a helical shape, and being disposed on a distal end side of said wire body;a second coil having a helical shape, and being disposed on a distal end side ... | Terumo Kabushiki Kaisha | A61B005/00 A61M025/00
US07251610 | B2 | 2007-07-31 | Clinical documentation system for use by multiple caregivers | 2001-09-10 | A computer-based system for recording, storing, and accessing clinical documentation in an acute care setting is provided. In an embodiment of the invention, a single electronic database or repository for storing clinical patient notes, provides multiple points of read/write access via a user interface operating on one or more client computers that are in real-time communication with the repository. | 1. A system for documenting patient information from multiple caregivers comprising: a health record system including a plurality of data elements, the data elements including patient note data acquired from multiple caregivers, the patient note data including at least a first patient note and a second patient note related to the care of a given patient; anda first interface and a second interface coupled to the health record system each including a graphic display, the first interface and second interface operable to retrieve, based upon information received from a first and second user, patient note data associated with a specified patient, and to graphically depict the patient note data in list form and text form within the graphic displays, each of the first and second interfaces further including an editor operable with the interface to facilitate editing of the patient notes within the graphic display;wherein the first interface is adapted to lock the first patient note when the first user is editing the first patient note such that the second interface maintains view access to the patient notes but is denied write access to the first patient note while the first user is editing the first patient note, and wherein the second interface is adapted to simultaneously lock the second patient note when the second user is editing the second patient note such that the first interface maintains simultaneous view access to the patient notes but is denied write access to the second patient note while the second user is editing the second patient note; andwherein each interface further includes a role-based security scheme that limits at least one of the first and second user's viewing and editing access to the patient note data associated with the specified patient to only limited types of information based on a role of the user. | 1. A system for documenting patient information from multiple caregivers comprising: a health record system including a plurality of data elements, the data elements including patient note data acquired from multiple caregivers, the patient note data including at least a first patient note and a second patient note related to the care of a given patient; anda first interface and a second interface coupled to the health record system each including a graphic display, the first interface and second interface operable to retrieve, based upon information received from a first and second user, patient note data associated with a specified patient, and to graphically depict the patient note data in list form and text form within the graphic displays, each of the first and second interfaces further including an editor operable with the interface to facilitate editing of the patient notes within the graphic display;wherein the first interface is adapted to lock the first patient note when the first user is editing the first patient note such that the second interface maintains view access to the patient notes but is denied write access to the first patient note while the first user is editing the first patient note, and wherein the second interface is adapted to simultaneously lock the second patient note when the second user is editing the second patient note such that the first interface maintains simultaneous view access to the patient notes but is denied write access to the second patient note while the second user is editing the second patient note; andwherein each interface further includes a role-based security scheme that limits at least one of the first and second user's viewing and editing access to the patient note data associated with the specified patient to only limited types of information based on a role of the user. 2. The system of claim 1, wherein the editor is operable with the interface to facilitate creation of patient note data. 3. The system of claim 1, wherein the editor is operable independent of the depiction of the patient note data within the graphic display. 4. The system of claim 1, wherein the plurality of caregivers are associated with a plurality of medical services. 5. The system of claim 1, wherein the plurality of caregivers are associated with a plurality of related medical services. 6. The system of claim 1, wherein the interface further comprises a filter for selecting particular patient note data for depiction within the graphic display. 7. The system of claim 6, further comprising filter criteria associated with the filter, the filter criteria comprising at least one of: author, author type and time. 8. The system of claim 1, wherein the patient note data is linked by time stamp data. 9. The system of claim 8, wherein the patient note data is graphically depicted within the graphic display according to the time stamp data. 10. The system of claim 1, wherein the interface further comprises a searcher oper... | Epic Systems Corporation | A61B005/00
US07112176 | B2 | 2006-09-26 | Compensation method for thermodilution catheter having an injectate induced thermal effect in a blood flow measurement | 2002-02-20 | A catheter for retrograde orientation in a blood flow is used to determine the blood flow rate by thermodilution measurements. The determination of the blood flow rate accommodates injectate induced thermal influences on a dilution thermal sensor, wherein the thermal influences can occur prior to introduction of the injectate into the blood flow. | 1. A method of determining blood flow in a conduit, comprising compensating for an injectate induced thermal offset of a thermal dilution sensor connected to a catheter, the compensating including one of (i) pre-calibrating a thermal property of the catheter or the thermal dilution sensor; (ii) determining a calibration coefficient for the catheter and (iii) determining a thermal transfer coefficient for the catheter. | 1. A method of determining blood flow in a conduit, comprising compensating for an injectate induced thermal offset of a thermal dilution sensor connected to a catheter, the compensating including one of (i) pre-calibrating a thermal property of the catheter or the thermal dilution sensor; (ii) determining a calibration coefficient for the catheter and (iii) determining a thermal transfer coefficient for the catheter. 2. The method of claim 1, wherein compensating for an injectate induced thermal offset of a thermal dilution sensor includes thermally insulating the thermal dilution sensor from the injectate prior to introduction of the injectate into the blood flow in the conduit. 3. The method of claim 1, further comprising determining a calibration coefficient for the catheter. 4. The method of claim 3, further comprising adjusting the calibration coefficient in response to a blood flow rate in the conduit or an injection rate of the induced injectate. 5. The method of claim 3, further comprising increasing the calibration coefficient in response to a reduced blood flow rate. 6. The method of claim 3, further comprising decreasing the calibration coefficient in response to an increased blood flow. 7. The method of claim 1, further comprising disposing the catheter in a retrograde orientation in the conduit. 8. A method of determining blood flow in a conduit, comprising compensating for an injectate induced thermal offset of a thermal dilution sensor connected to a catheter by pre-calibrating a thermal property of the catheter. 9. The method of claim 8, wherein pro-calibrating a thermal property of the catheter includes determining a thermal transfer coefficient Ki, such thatKi=Δ⁢⁢Ti(Tb-Ti);where Tb corresponds to the temperature of the blood, Ti corresponds to the temperature of the injectate and ΔTi is the change in the thermal dilution sensor temperature from the injectate induced cooling. 10. The method of claim 8, further comprising disposing the catheter in a retrograde orientation in the conduit. 11. A method of determining a blood flow in a conduit, the method comprising: (a) passing an injectate through a lumen in a catheter, the passing injectate inducing a measurement offset in a blood parameter sensor; and(b) compensating for the measurement offset of the blood parameter sensor by one of (i) pre-calibrating a thermal property of the catheter or the blood parameter sensor; (ii) determining a calibration coefficient for the catheter and (iii) determining a thermal transfer coefficient for the catheter. 12. The method of claim 11, wherein compensating for measurement offset includes thermally isolating the blood parameter sensor from the injectate passing through the lumen in the catheter. 13. The method of claim 11, further comprising disposing the catheter in a retrograde orientation in the conduit. 14. A method of determining a blood flow in a conduit, the method comprising: (a) passing an injectate through a lumen in a ... | Transonic Systems, Inc. | A61B005/00
US07184818 | B2 | 2007-02-27 | Method and system for characterizing a representative cardiac beat using multiple templates | 2002-03-25 | The present invention provides a method and system for characterizing one beat of a patient's supraventricular rhythm. A plurality of templates is provided and updated using a plurality of qualified beats. Updating occurs by temporally aligning the shock channel waveforms of the template beats using rate channel fiducial points. The template beats are combined by point-by-point addition of the shock channel waveforms. The resultant updated template characterizes one of the patient's supraventricular conducted cardiac beats. | 1. A method of characterizing a patient's supraventricular rhythm, comprising: providing a plurality of templates;detecting, using an implantable lead system, rate channel signals and shock channel signals of a plurality of beats;determining if the beat signals meet predetermined criteria:in a body implantable system, selectively updating the plurality of templates, comprising: selecting a template of the plurality of templates for updating based on a similarity between the template and a beat of the plurality of beats; andupdating the selected template using the beat; andcharacterizing, in the body implantable system, the patient's supraventricular rhythm by identifying a particular template of the plurality of selectively updated templates as representative of the patient's supraventricular rhythm. | 1. A method of characterizing a patient's supraventricular rhythm, comprising: providing a plurality of templates;detecting, using an implantable lead system, rate channel signals and shock channel signals of a plurality of beats;determining if the beat signals meet predetermined criteria:in a body implantable system, selectively updating the plurality of templates, comprising: selecting a template of the plurality of templates for updating based on a similarity between the template and a beat of the plurality of beats; andupdating the selected template using the beat; andcharacterizing, in the body implantable system, the patient's supraventricular rhythm by identifying a particular template of the plurality of selectively updated templates as representative of the patient's supraventricular rhythm. 2. The method of claim 1, wherein providing the plurality of templates further comprises retrieving at least a first template from memory and forming one or more other templates using the beats. 3. The method of claim 1, wherein selecting the template of the plurality of templates for updating, comprises: determining if the beat is correlated to the template; and updating the selected template with the correlated beat. 4. The method of claim 3, wherein updating the selected template further comprises combining the beat and the selected template by any method selected from point-by-point addition of the beat to the selected template, performing a running average of the beat and the selected template, performing a weighted running average of the beat and the selected template, and using a median filtering technique to combine the beat and the selected template. 5. The method of claim 3, further comprising incrementing a counter associated with the selected template each time the selected template is updated. 6. The method of claim 3, wherein determining if the beat is correlated to the template further comprises: determining the fiducial points of the beat and the template, respectively;aligning shock channel signals of the beat and the template;identifying a plurality of features of shock channel signals of the beat and the template; andusing the plurality of identified features to determine if the beat is correlated to the template. 7. The method of claim 1, wherein characterizing the patient's supraventricular rhythm using the particular template of the plurality of updated templates further comprises storing the particular template. 8. The method of claim 1, further comprising: providing a plurality of counters associated with the plurality of templates;incrementing one of the plurality of counters when the template associated with the counter is updated; andusing the template which has an associated counter with the highest count as a characterization of the patient's supraventricular rhythm. 9. The method of claim 1, wherein updating the plurality of templates further comprises replacing one of the plurality of templates with a beat when ... | Cardiac Pacemakers, Inc. | A61B005/0402
US06937888 | B2 | 2005-08-30 | Signal evaluation method for detecting QRS complexes in electrocardiogram signals | 2002-02-07 | A signal evaluation method for detecting QRS complexes in electrocardiogram (ECG) signals comprises the following steps: sampling of the ECG signal (4) and conversion into discrete signal values (x(n)) in chronological order;comparing the signal values (xf(n), xfq(n)) to a threshold function (K(n)) adaptively determined therefrom;determining a frequency number (D(n)) within a defined segment of the consecutive signal values, by which signal values (xf(n), xfq(n)) preferably fall short of the threshold function (K(n));comparing the determined frequency number (D(n)) to a defined threshold (Θ), wherein an undershoot of the threshold (Θ) is significant for apresence of a QRS complex (5, 6, 7) in the defined segment of the ECG signal (4). | 1. A signal evaluation method for detecting QRS complexes in an electrocardiogram (ECG) signal, comprising the following steps: sampling the ECG signal to produce consecutive sampled signal values; converting the ECG sampled signal values into consecutive discrete signal values of chronological order; comparing the discrete signal values to a threshold function adaptively determined from the discrete signal values; determining a frequency number within a defined segment of the consecutive discrete signal values, the frequency number being representative of the number of discrete signal values that are below the threshold function; and comparing the determined frequency number to a defined frequency number threshold, wherein the presence of a QRS complex in the defined segment of the ECG signal is indicated when the determined frequency number is less than the frequency number threshold. | 1. A signal evaluation method for detecting QRS complexes in an electrocardiogram (ECG) signal, comprising the following steps: sampling the ECG signal to produce consecutive sampled signal values; converting the ECG sampled signal values into consecutive discrete signal values of chronological order; comparing the discrete signal values to a threshold function adaptively determined from the discrete signal values; determining a frequency number within a defined segment of the consecutive discrete signal values, the frequency number being representative of the number of discrete signal values that are below the threshold function; and comparing the determined frequency number to a defined frequency number threshold, wherein the presence of a QRS complex in the defined segment of the ECG signal is indicated when the determined frequency number is less than the frequency number threshold. 2. The signal evaluation method according to claim 1, wherein said step of converting comprises subjecting the sampled ECG signal values to a high-pass filtering. 3. The signal evaluation method according to claim 1, wherein said step of converting comprises subjecting the sampled EGG signal values to a band-pass filtering. 4. The signal evaluation method according to claim 3, wherein upper and lower limiting pass frequencies of the band-pass filter are approximately 18 Hz and approximately 27 Hz. 5. The signal evaluation method according to claim 3, wherein said step of converting further comprises generating absolute values of the filtered signal values. 6. The signal evaluation method according to claim 5, wherein said step of generating absolute values is carried out by mathematically squaring the filtered signal values. 7. The signal evaluation method according to claim 6, wherein the value of the threshold function is determined adaptively from a flowing averaging of the squared signal values for an averaging period determined by a memory factor. 8. The signal evaluation method according to claim 1, wherein the value of the threshold function is determined adaptively from a flowing averaging of the discrete signal values, for an averaging period determined by a memory factor. 9. The signal evaluation method according to claim 1, wherein the frequency number threshold is variably set as an adaptive threshold from the frequency number itself. | Biotronik GmbH & Co. KG | A61B005/04
US06954665 | B2 | 2005-10-11 | Catheter system | 2003-02-04 | A catheter system is suitable for the simultaneous, continuous, mutually unaffected measurement of the centrovenous oxygen saturation and the local concentration of injected indocyanine green. It has a centrovenous catheter with a fiber-optic lumen and a fiber-optic probe that can be inserted into the fiber-optic lumen for taking reflecto-oximetric measurements at a first wavelength of 660 nm, a second measuring wavelength of 805 nm and a reference wavelength of 880 nm. To prevent a longitudinal shift of the fiber-optic probe relative to the fiber-optic lumen, a connecting piece, securely connected to the fiber-optic probe, and a counterpart, securely connected to the catheter, are provided which can be joined together. The fiber-optic lumen which continues in the interior of the connecting piece can be cleaned when the connecting piece is attached to the counterpart. For this purpose, the connecting piece has a cleansing attachment for connecting a cleansing device. The cleaning takes place through the cleansing channel. | 1. A catheter system for continuously measuring the centrovenous oxygen saturation and/or measuring the local concentration of injected indocyanine green, having a flexible, elongated, centrovenously applicable basic body, a fiber-optic probe (8) for taking reflecto-oximetric measurements, a fiber-optic lumen for housing the fiber-optic probe and fastening means for preventing a longitudinal shift of the fiber-optic probe relative to the fiber-optic lumen, the fastening means being detachable to enable a longitudinal shift of the fiber-optic probe relative to the fiber-optic lumen for removal of the fiber-optic probe, wherein the fastening means has a connecting piece securely connected with the fiber-optic probe and a counterpart securely connected with the centrovenously applicable basic body which can be joined together, and that, when connected, the distance between the distal end of the centrovenously applicable basic body and the distal end of the fiber-optic probe is set at a predetermined value by the secure connections between probe and connecting piece as well as basic body and counterpart. | 1. A catheter system for continuously measuring the centrovenous oxygen saturation and/or measuring the local concentration of injected indocyanine green, having a flexible, elongated, centrovenously applicable basic body, a fiber-optic probe (8) for taking reflecto-oximetric measurements, a fiber-optic lumen for housing the fiber-optic probe and fastening means for preventing a longitudinal shift of the fiber-optic probe relative to the fiber-optic lumen, the fastening means being detachable to enable a longitudinal shift of the fiber-optic probe relative to the fiber-optic lumen for removal of the fiber-optic probe, wherein the fastening means has a connecting piece securely connected with the fiber-optic probe and a counterpart securely connected with the centrovenously applicable basic body which can be joined together, and that, when connected, the distance between the distal end of the centrovenously applicable basic body and the distal end of the fiber-optic probe is set at a predetermined value by the secure connections between probe and connecting piece as well as basic body and counterpart. 2. Catheter system according to claim 1, wherein the fiber-optic lumen can be cleaned when the connecting piece is attached to the counterpart. 3. Catheter system according to claim 2, wherein the connecting piece has a cleansing attachment for connection to a cleansing device. 4. Catheter system according to claim 1, wherein the predetermined value of the distance between distal end of the centrovenously applicable basic body and the distal end of the fiber-optic probe is between 10 and 40 mm, the distal end of the fiber-optic probe protruding beyond the distal end of the centrovenously applicable basic body. 5. Catheter system according to claim 1, wherein the predetermined value of the distance between the distal end of the centrovenously applicable basic body and the distal end of the fiber-optic probe is selected in such a way that the distal end of the fiber-optic probe extends flush with the basic body. 6. Catheter system according to claim 1, wherein the fiber-optic probe has a thin, flexible shaft consisting of a bio-compatible and hemo-compatible material as well as at least one afferent and at least one efferent optical fiber. 7. Catheter system according to claim 1, wherein the fiber-optic probe has an antithrombogenic cover in the area of a possible contact with blood. 8. Catheter system according to claim 1, further comprising insertion aids for the centrovenous application of the basic body. 9. Catheter system according to claim 8, wherein the insertion aids have a guide wire. 10. Catheter system according to claim 9, wherein the insertion aids also have a puncture cannula suitable for inserting the guide wire. 11. Catheter system according to claim 10, wherein the insertion aids also have a tapered dilator adapted to the diameter of the guide wire, whose shaft diameter is at least as large as the outside diameter of the cen... | Pulsion Medical Systems AG | A61B005/00
US06934577 | B2 | 2005-08-23 | Risk monitoring | 2001-09-24 | A risk assessment device, in particular for the risk of a severe heart functional disorder such as sudden death due to heart disease, comprising an input unit which is adapted to record or read in an electrocardiographic input signal and to output a corresponding output signal comprising a plurality of cardiac cycles and corresponding intervals, and a detection unit which is connected to the input unit and which is adapted to detect events of predetermined nature in the output signal, wherein there are provided a summing unit which is connected to the detection unit and the input unit and is adapted to add up the length of predetermined intervals connected to a detected event within the input signal and to output a corresponding sum parameter, and an evaluation unit which is adapted to evaluate the sum parameter under predetermined conditions and to output a risk parameter which is dependent on the evaluation. | 1. A risk assessment device, for assessing the risk of a severe heart functional disorder, comprising: an input unit to record, or read in, an electrocardiographic input signal and to output a corresponding processed output signal; a detection unit, connected to the input unit, to detect events of a predetermined nature in the output signal; a summing unit connected to the detection unit and the input unit, the summing unit adding up a length of predetermined intervals connected to a detected event within the input signal and outputting a corresponding sum parameter; and an evaluation unit to evaluate the sum parameter under predetermined conditions and to output a risk parameter which is dependent on the evaluation. | 1. A risk assessment device, for assessing the risk of a severe heart functional disorder, comprising: an input unit to record, or read in, an electrocardiographic input signal and to output a corresponding processed output signal; a detection unit, connected to the input unit, to detect events of a predetermined nature in the output signal; a summing unit connected to the detection unit and the input unit, the summing unit adding up a length of predetermined intervals connected to a detected event within the input signal and outputting a corresponding sum parameter; and an evaluation unit to evaluate the sum parameter under predetermined conditions and to output a risk parameter which is dependent on the evaluation. 2. The risk assessment device of claim 1 wherein the predetermined events are from a group consisting of a QRS-complex, an extrasystole, and a combination thereof. 3. The risk assessment device of claim 2, further comprising a selection unit, connected to the detection unit, to select among the detected events such events which have one or more predetermined properties. 4. The risk assessment device of claim 3, wherein the selection unit selects ventricular extrasystoles which involve a change in a morphology of a QRS-complex associated with an extrasystole in relation to QRS-complexes not linked to an extrasystole, in the output signal of the input unit. 5. The risk assessment device of claim 4, wherein the selection unit selects ventricular VES-2 extrasystoles that are followed by an interval which is increased in length by a predetermined limit value in comparison with preceding intervals not containing any extrasystoles. 6. The risk assessment device claim 5, wherein the selection unit selects, to form a reference interval, a predetermined number of intervals not containing any extrasystoles prior to and after a VES-2 extrasystole. 7. The risk assessment device of claim 6, further comprising an interval selection unit, connected to the selection unit, to select intervals that are associated with at least two VES-2 extrasystoles which occur within a predetermined period of time, to form the sum parameter. 8. The risk assessment device of claim 7, wherein the predetermined period of time is of the order of magnitude of one hour. 9. The risk assessment device of claim 8, wherein the evaluation unit forms a trend pattern by evaluating the sum parameter, in such a way that the risk parameter corresponds to a trend of the sum parameter. 10. The risk assessment device of claim 9, further comprising a second evaluation unit, connected to the evaluation unit, to take over the risk parameter as a first parameter, the second evaluation unit having at least one input for at least one further parameter, the second evaluation unit calculating a risk signal corresponding to a probability of a severe heart functional disorder based on a linear combination of at least two of the parameters, as is commonly determined by logistical regres... | Biotronik Mess- und Therapiegeraete GmbH & Co. Ingeieurbuero Berlin | A61B005/04
US07153265 | B2 | 2006-12-26 | Anti-inflammatory biosensor for reduced biofouling and enhanced sensor performance | 2002-04-22 | A biosensor including an external surface, and an accessory material in close proximity to the external surface. The accessory material includes a coating containing a hydrophilic material and/or a fiber modified to deliver a therapeutic agent. The biosensor modifies a biological response to the biosensor upon contact with a tissue, such as upon implantation into the skin of a subject, thereby reducing biofouling, inflammation and other undesirable tissue responses that interfere with biosensor performance. The biosensor can be any biocompatible sensor, suitable for short- or long-term use. Preferably, the biosensor is an enzymatic or electrochemical sensor, such as a glucose sensor. Also provided are a method of producing a biosensor and a method of delivering a biologically active substance to a subject. | 1. A biosensor comprising an analyte sensor having an external surface, and an accessory material coated with a hydrophilic polymer provided in proximity to the external surface, wherein the accessory material modifies the biological response of a tissue that is in contact with the biosensor, wherein the accessory material is a fiber modified to deliver dexamethasone or a salt thereof. | 1. A biosensor comprising an analyte sensor having an external surface, and an accessory material coated with a hydrophilic polymer provided in proximity to the external surface, wherein the accessory material modifies the biological response of a tissue that is in contact with the biosensor, wherein the accessory material is a fiber modified to deliver dexamethasone or a salt thereof. 2. The biosensor of claim 1, wherein the biological response comprises protein deposition, inflammation or proliferation of macrophages or foreign body giant cells. 3. The biosunsor of claim 1, wherein the hydrophilic polymer comprises polyhydroxyethylmethacrylate (PHEMA), polyurethane, polysaccharide, polyacrylamide, or polyurea. 4. The biosensor of claim 1, wherein the hydrophilic polymer comprises polyethylene oxide (PEO). 5. The biosensor of claim 4, wherein the PEO comprises a polyurethane, polyurea or a cross-linked PEO. 6. The biosensor of claim 1, wherein the salt is a sodium phosphate or acetate salt. 7. The biosensor of claim 1, which comprises an enzymatic, molecular recognition, oprochemical or electrochemical sensor. 8. The biosensor of claim 1, which comprises a glucose sensor. 9. The biosensor of claim 1, wherein the fiber comprises cotton, polypropylene, polyurethane, polyester, polylacric acid (PLA), polyglycolic acid (PGA), or a co-polymer of lactic acid and glycolic acid (PLGA). 10. The biosensor of claim 1, further comprising an additional therapeutic agent, which additional agent comprises an anti-inflammatory, anti-bacterial, anti-viral, anti-coagulant, anti-proliferative or disinfecting agent, or a growth factor. 11. The biosensor of claim 1, wherein the fiber is affixed to a distal end of the biosensor. 12. The biosensor of claim 11, wherein the fiber is affixed to the distal end of the biosensor by passage through a hole in the biosensor. 13. The biosensor of claim 1, wherein the fiber is affixed to the biosensor by insertion into a groove in the biosensor. 14. A method of delivering a biologically active substance to a subject comprising implanting a biosensor of claim 1 into a tissue of the subject, wherein the accessory niaterial comprises the biologically active substance. 15. The method of claim 14, wherein the biologically active substance additionally comprises a cytokine, growth factor or therapeutic agent. 16. A biosensor comprising a sensor having an external surface, and an accessory material provided in proximity to the external surface, wherein the accessory material modifies the biological response of a tissue that is in contact with the biosensor, wherein the accessory material is a fiber modified to deliver a therapeutic agent, and wherein the therapeutic agent comprises cefazolin or benzalkonium chloride, wherein the fiber comprises cotton, polypropylene polyurethane, polyesrer, polylactic acid (PLA), polyglycolic acid (PGA), or a co-polymer of lactic acid and glycolic acid (PLGA). 17. The biosensor of claim... | Medtronic MiniMed, Inc. | A61B005/05 A61B005/00
US07289839 | B2 | 2007-10-30 | Implantable marker with a leadless signal transmitter compatible for use in magnetic resonance devices | 2002-12-30 | A leadless marker for localizing the position of a target within a patient. In one embodiment, the marker includes a casing, a resonating circuit, and a ferromagnetic element. The casing is configured to be positioned at a selected location relative to a target site in the patient; the casing, for example, can be configured to be permanently or semi-permanently implanted into the patient. The resonating circuit has an inductor within the casing comprising a plurality of windings of a conductor, but it does not have external electrical lead lines extending through the casing. The ferromagnetic element is at least partially within the inductor. The ferromagnetic element has a volume such that when the marker is in an imaging magnetic field having a field strength of 1.5 T and a gradient of 3 T/m, then the force exerted on the marker by the imaging magnetic field is not greater than gravitational force exerted on the marker. | 1. A lead less marker for localizing the position of a target within a patient, comprising: a casing configured to be positioned at a selected location relative to a target site in the patient;a resonating circuit without external electrical lead lines extending through the casing, the resonating circuit having an inductor within the casing comprising a plurality of windings of a conductor; anda ferromagnetic element at least partially within the inductor, the ferromagnetic element having a volume such that when the marker is in an imaging magnetic field having a field strength of 1.5 T and a gradient of 3 T/m, then force exerted on the marker by the imaging magnetic field is not greater than gravitational force exerted on the marker. | 1. A lead less marker for localizing the position of a target within a patient, comprising: a casing configured to be positioned at a selected location relative to a target site in the patient;a resonating circuit without external electrical lead lines extending through the casing, the resonating circuit having an inductor within the casing comprising a plurality of windings of a conductor; anda ferromagnetic element at least partially within the inductor, the ferromagnetic element having a volume such that when the marker is in an imaging magnetic field having a field strength of 1.5 T and a gradient of 3 T/m, then force exerted on the marker by the imaging magnetic field is not greater than gravitational force exerted on the marker. 2. The marker of claim 1 wherein (a) the inductor further comprises a capacitor electrically coupled to the conductor, (b) the ferromagnetic element comprises a ferrite rod having a diameter not greater than approximately 0.5 mm, and (c) the casing has an outer diameter not greater than approximately 2 mm. 3. The marker of claim 1 wherein the ferromagnetic element comprises a ferrite rod having a diameter of approximately 0.2 mm-0.7 mm and a length of approximately 2 mm-12 mm. 4. The marker of claim 1 wherein the ferromagnetic element comprises a ferrite rod having a volume of approximately 0.5 mm3-19.0 mm3. 5. The marker of claim 1 wherein the ferromagnetic element comprises a ferrite rod and the volume of the rod is less than a volume of ferrite that produces an artifact of 1500 mm2 in an image produced by a resonating magnetic field of approximately 1.5 T. 6. The marker of claim 1 wherein the ferromagnetic element comprises a ferrite rod and the volume of the rod is less than a volume of ferrite that produces an artifact of 400 mm2-1200 mm2 in an image produced by a resonating magnetic field of approximately 1.5 T. 7. The marker of claim 1 wherein (a) the ferromagnetic element comprises a ferrite rod having a diameter of approximately 0.2 mm-0.7 mm, (b) the windings comprise a coil having an inner diameter of approximately 0.2 mm-0.8 mm and an outer diameter of approximately 0.8 mm-1.9 mm, and (c) the casing has an outer diameter of approximately 1.0 mm-2.5 mm. 8. The marker of claim 1 wherein (a) the ferromagnetic element comprises a ferrite rod having a diameter of approximately 0.3 mm-0.5 mm, (b) the windings comprise a coil having an inner diameter of approximately 0.3 mm-0.6 mm and an outer diameter of approximately 1.2 mm-1.9 mm, and (c) the casing has an outer diameter of approximately 2 mm. 9. The marker of claim 1 wherein (a) the inductor further comprises a capacitor at one end of the ferromagnetic element which is electrically coupled to the conductor, and (b) the marker further comprises a module at the other end of the capacitor which is symmetrical with respect to the capacitor. 10. A leadless implantable marker for localizing the position of a target within a patient, comprising: a casing ... | Calypso Medical Technologies, Inc. | A61B005/00
US07085599 | B2 | 2006-08-01 | Characterization of supraventricular rhythm using collected cardiac beats | 2002-10-23 | The supraventricular rhythm of a patient's heart is characterized using cardiac beats collected during a plurality of time periods, some of which are noncontiguous. Pacing parameters may be modified to allow the development of a spontaneous rhythm. Waveforms of the collected cardiac beats are stored for later processing, or are combined with previously collected beats and the combination is stored. When a sufficient number of cardiac beats are collected, the patient's supraventricular rhythm is characterized using the collected cardiac beats. | 1. A method for characterizing a patient's supraventricular rhythm, comprising: collecting cardiac beats representative of the patient's supraventricular rhythm during a plurality of periods, of which some are noncontiguous; andgenerating a morphology template characterizing the patient's supraventricular rhythm using the collected cardiac beats. | 1. A method for characterizing a patient's supraventricular rhythm, comprising: collecting cardiac beats representative of the patient's supraventricular rhythm during a plurality of periods, of which some are noncontiguous; andgenerating a morphology template characterizing the patient's supraventricular rhythm using the collected cardiac beats. 2. The method of claim 1, wherein collecting the cardiac beats comprises storing the cardiac beats. 3. The method of claim 1, wherein collecting the cardiac beats comprises combining each cardiac beat with previously collected cardiac beats and storing the result of the combination. 4. The method of claim 3, wherein combining each cardiac beat with previously collected cardiac beats comprises performing point-by-point addition of a cardiac beat waveform to a template waveform. 5. The method of claim 1, wherein collecting the cardiac beats comprises: modifying a pacing parameter; andcollecting the cardiac beats during a time in which the pacing parameter is modified. 6. The method of claim 5, wherein collecting the cardiac beats further comprises returning the pacing parameter to an original value. 7. The method of claim 5, wherein modifying the pacing parameter comprises modifying a pacemaker mode. 8. The method of claim 5, wherein modifying the pacing parameter comprises modifying a pacing rate. 9. The method of claim 5, wherein modifying the pacing parameter comprises modifying a pacing interval. 10. The method of claim 5, wherein modifying the pacing parameter comprises suspending one or more pacemaker functions. 11. The method of claim 5, wherein modifying the pacing parameter comprises increasing a pacing escape interval. 12. The method of claim 5, wherein modifying the pacing parameter comprises increasing an AV interval. 13. The method of claim 5, wherein modifying the pacemaker parameter comprises changing a pacemaker mode from a tracking, rate responsive mode to a non-tracking, non-rate responsive mode. 14. The method of claim 5, wherein collecting the cardiac beats during a time in which the pacing parameter is modified comprises collecting the cardiac beats during periodic sessions in which the pacing parameter is modified. 15. The method of claim 14, wherein the periodic sessions in which the pacing parameter is modified occur every 32 to 1024 beats. 16. The method of claim 5, wherein modifying the pacing parameter comprises modifying the pacing parameter to a pacing parameter limit. 17. The method of claim 5, wherein modifying the pacing parameter comprises: increasing an AV interval;modifying a pacemaker mode if no intrinsic beats are detected during a period having the increased AV interval; andslowing a pacing rate to allow intrinsic beats to develop. 18. The method of claim 5, wherein modifying the pacing parameter comprises switching a pacemaker mode from a dual chamber mode to a single chamber mode. 19. The method of claim 5, wherein modifying the pacing parameter com... | Cardiac Pacemakers, Inc. | A61B005/0452
US07065396 | B2 | 2006-06-20 | System and method for non-invasive monitoring of physiological parameters | 2002-07-30 | A method and apparatus for enabling the non-invasive acquisition of physiological parameters of a mammal, while allowing the mammal to move around a predetermined area is provided. In accordance with one example embodiment, a monitoring device for non-invasively acquiring physiological measurements of a mammal is provided. The device includes a first electrode configured to attach to an ear of the mammal in a manner suitable for obtaining a first electrical signal. A second electrode is configured to attach to a first alternate location of the mammal in a manner suitable for obtaining a second electrical signal. A third electrode is configured to attach to a second alternate location of the mammal in a manner suitable for obtaining a third electrical signal. The first electrode, second electrode, and third electrode are in communication with a signal receiver. Physiological parameters, such as those associated with electrocardiograms, can be obtained using the monitoring device. | 1. A method of non-invasively acquiring a physiological measurement of a mammal, comprising: establishing an electrical connection between a first electrode and a first ear of the mammal in a manner suitable for obtaining a first electrical signal;establishing an electrical connection between a second electrode and a first alternate location of the mammal in a manner suitable for obtaining a second electrical signal;establishing an electrical connection between a means for electrical signal detection and the mammal in a manner suitable for obtaining a third electrical signal;receiving the first electrical signal, the second electrical signal, and the third electrical signal at a receiver; andrectifying the first electrical signal, the second electrical signal, and the third electrical signal into the physiological measurement;wherein the physiological measurement comprises an electrocardiogram. | 1. A method of non-invasively acquiring a physiological measurement of a mammal, comprising: establishing an electrical connection between a first electrode and a first ear of the mammal in a manner suitable for obtaining a first electrical signal;establishing an electrical connection between a second electrode and a first alternate location of the mammal in a manner suitable for obtaining a second electrical signal;establishing an electrical connection between a means for electrical signal detection and the mammal in a manner suitable for obtaining a third electrical signal;receiving the first electrical signal, the second electrical signal, and the third electrical signal at a receiver; andrectifying the first electrical signal, the second electrical signal, and the third electrical signal into the physiological measurement;wherein the physiological measurement comprises an electrocardiogram. 2. The method of claim 1, wherein the mammal comprises a rodent. 3. The method of claim 1, wherein establishing an electrical connection comprises positioning the first electrode and the second electrode with at least one of a removable fastener and/or structure supporting at least one of the first electrode and/or the second electrode for passive contact with the mammal. 4. A method of non-invasively acquiring a physiological measurement of a mammal, comprising: receiving a first electrical signal originating from a first ear of the mammal;receiving a second electrical signal originating from a first alternate location of the mammal;receiving a third electrical signal originating from a second alternate location of the mammal; andrectifying the first electrical signal, the second electrical signal, and the third electrical signal into the physiological measurement;wherein the physiological measurement comprises an electrocardiogram. 5. The method of claim 4, wherein the mammal comprises a rodent. 6. The method of claim 1, wherein the mammal is conscious while the physiological measurements are obtained. | The Curavita Corporation | A61B005/04
US06890306 | B2 | 2005-05-10 | Active medical device for the diagnosis of the sleep apnea syndrome | 2002-12-16 | An active medical device have an improved diagnosis of a sleep apnea syndrome. This device measures the respiratory activity of the patient, determines a state of activity, this state being likely to take, according to satisfaction of predetermined criteria, a value representative of a state of sleep of the patient, and analyzes a detected signal corresponding to the respiratory activity to detect, when the aforementioned state is a state of sleep, the presence of respiratory pauses, and thereby to produce an indicating signal of sleep apnea in the event of the occurrence of a respiratory pause of duration longer than a first predetermined duration. The analysis also includes inhibiting the production of the aforesaid indicating signal, or a treatment to resolve an apnea, when the duration of the detected respiratory pause is longer than a second predetermined duration, typically of at least one minute. | 1. An active medical device, comprising: means for detecting a respiratory activity of a patient and producing a signal corresponding to said detected respiratory activity; means for determining a state of activity of the patient comprising means for acquiring an activity parameter, means for comparing said activity parameter to a predetermined criteria, and means for determining that said state is a sleep state on satisfaction of said predetermined criteria, means for analyzing the detecting means signal comprising means for identifying a respiratory pause and for determining a duration of said detected respiratory pause, and means for producing an indicating signal of sleep apnea in response to an identified respiratory pause having a duration longer than a first predetermined duration during a determined sleep state; wherein the analyzing means further comprises means for inhibiting the production of the indicating signal in response to said respiratory pause duration being longer than a second predetermined duration. | 1. An active medical device, comprising: means for detecting a respiratory activity of a patient and producing a signal corresponding to said detected respiratory activity; means for determining a state of activity of the patient comprising means for acquiring an activity parameter, means for comparing said activity parameter to a predetermined criteria, and means for determining that said state is a sleep state on satisfaction of said predetermined criteria, means for analyzing the detecting means signal comprising means for identifying a respiratory pause and for determining a duration of said detected respiratory pause, and means for producing an indicating signal of sleep apnea in response to an identified respiratory pause having a duration longer than a first predetermined duration during a determined sleep state; wherein the analyzing means further comprises means for inhibiting the production of the indicating signal in response to said respiratory pause duration being longer than a second predetermined duration. 2. The device of claim 1, wherein said second predetermined duration is at least one minute. 3. The device of claim 1, wherein said device is an implantable medical device and further comprises means for treating a sleep apnea syndrome by delivering an electro-cardiac stimulation in response to said indicating signal, and means for inhibiting said electro-cardiac stimulation treating means from delivering a stimulation in response to the detected respiratory pause being longer than said second predetermined duration. 4. An active medical device, comprising: a respiratory activity monitor having an output signal corresponding to a patient's respiratory activity; a patient state detector including an activity sensor having an output corresponding to a patient activity state, said detector including a predetermined criteria corresponding to a sleep state, and said output signal having a sleep state output in response to said patient activity signal satisfying said predetermined criteria; a controller able to analyze the respiratory output signal and said state detector wherein the controller analyzes the respiratory output signal and identifies a respiratory pause therein and a duration of said detected respiratory pause, and produces an indicating signal of sleep apnea in response to an identified respiratory pause having a duration longer than a first predetermined duration during a determined sleep state; wherein the controller operates to inhibit the production of the indicating signal in response to said respiratory pause duration being longer than a second predetermined duration. 5. The device of claim 4, wherein said second predetermined duration is at least one minute. 6. The device of claim 4, wherein said device is an implantable medical device and further comprises means for treating a sleep apnea syndrome by delivering an electro-cardiac stimulation in response to said indicating signal, and means for inhibiting s... | ELA Medical S.A. | A61B005/08 A61B005/05
US06979299 | B2 | 2005-12-27 | Measuring guide for use in orthopedic procedure | 2002-02-05 | A measuring guide for use in various surgical procedures. The measuring guide is a noninvasive tool that may be repeatedly applied to a desired bone member to make appropriate measurements. The measuring guide allows for easy adjustment of the distance between a stop plate designed to engage an end of a given bone member and a marking guide used to facilitate marking of the bone. Additionally, the stop plate may be pivotably mounted to facilitate ease and accuracy of measurement. | 1. A measuring guide for assisting in locating a prosthetic device during an orthopedic procedure, comprising: a marking guide adapted to facilitate marking of a bone tissue at a desired location;a stop plate;a ruler, including a top, a bottom and two sides, the ruler being coupled to the marking guide and to the stop plate, and the ruler indicating the distance between the marking guide and the stop plate, wherein the stop plate includes an abutment surface, configured to abut a bone, and further wherein the abutment surface is positioned on both sides of the ruler; anda locking mechanism that cooperates with the ruler to permit selective adjustment of the distance between the marking guide and the stop plate by moving the stop plate with respect to the ruler, wherein the marking guide and the stop plate are utilized in a manner noninvasive to the bone tissue, and further wherein the stop plate is pivotably coupled to the locking mechanism to permit positioning of the stop plate at desired angles with respect to the ruler. | 1. A measuring guide for assisting in locating a prosthetic device during an orthopedic procedure, comprising: a marking guide adapted to facilitate marking of a bone tissue at a desired location;a stop plate;a ruler, including a top, a bottom and two sides, the ruler being coupled to the marking guide and to the stop plate, and the ruler indicating the distance between the marking guide and the stop plate, wherein the stop plate includes an abutment surface, configured to abut a bone, and further wherein the abutment surface is positioned on both sides of the ruler; anda locking mechanism that cooperates with the ruler to permit selective adjustment of the distance between the marking guide and the stop plate by moving the stop plate with respect to the ruler, wherein the marking guide and the stop plate are utilized in a manner noninvasive to the bone tissue, and further wherein the stop plate is pivotably coupled to the locking mechanism to permit positioning of the stop plate at desired angles with respect to the ruler. 2. The measuring guide as recited in claim 1, wherein the abutment surface of the stop plate is configured to abut the distal femoral condyles. 3. The measuring guide as recited in claim 1, wherein the abutment surface of the stop plate is configured to abut the proximal end of a tibia. 4. The measuring guide as recited in claim 1, wherein the stop plate is adapted to be locked at desired angles with respect to the ruler. 5. The measuring guide as recited in claim 4, wherein the desired angles are approximately 84°, 90°, and 96°. 6. The measuring guide as recited in claim 1, wherein the locking mechanism comprises an opening through which the ruler is received and a release mechanism to selectively release the ruler for sliding movement through the opening. 7. The measuring guide as recited in claim 6, wherein the ruler comprises a flat side and a series of periodic grooves, and the release mechanism comprises at least one corresponding protrusion to engage selected grooves of the series of periodic grooves. 8. The measuring guide as recited in claim 7, further comprising a handle coupled to the locking mechanism and a lever coupled to the release mechanism to permit selective engagement and disengagement of the release mechanism and the ruler. 9. The measuring guide as recited in claim 1, wherein said locking mechanism includes an opening, and said ruler is configured to slide within said opening. 10. The measuring guide as recited in claim 1, wherein said ruler is a rigid member that extends the full length between the marking guide and the stop plate and said ruler includes a series of markings to indicate the distance between the marking guide and the stop plate, and further wherein when the stop plate is moved along the ruler towards the marking guide, the distance between the stop plate and the marking guide decreases. 11. The measuring guide as recited in claim 1, wherein said ruler is a rigid member that exte... | Zimmer Austin, Inc. | A61B005/103
US06892087 | B2 | 2005-05-10 | Vascular introducer with mapping capabilities | 2001-11-01 | A percutaneous vascular introducer is disclosed which includes a dilator having an elongated tubular body defining opposed proximal and distal end portions, the distal end portion of the dilator having, among other things, radiopaque characteristics to facilitate the safe placement of the distal tip of the dilator, and an elongated sheath having an interior lumen dimensioned and configured to accommodate the dilator. | 1. A percutaneous vascular introducer having mapping capabilities comprising: a) a dilator having an elongated tubular body defining opposed proximal and distal end portions, the distal end portion of the dilator including a tapered distal tip having radiopaque characteristics to facilitate safe positioning of the distal tip of the dilator within a patient's body, and wherein the dilator includes electrically conductive means to facilitate cardiac mapping; and b) an elongated sheath having an interior lumen dimensioned and configured to accommodate the dilator and provide a pathway for the introduction of an endocardial device in the absence of the dilator. | 1. A percutaneous vascular introducer having mapping capabilities comprising: a) a dilator having an elongated tubular body defining opposed proximal and distal end portions, the distal end portion of the dilator including a tapered distal tip having radiopaque characteristics to facilitate safe positioning of the distal tip of the dilator within a patient's body, and wherein the dilator includes electrically conductive means to facilitate cardiac mapping; and b) an elongated sheath having an interior lumen dimensioned and configured to accommodate the dilator and provide a pathway for the introduction of an endocardial device in the absence of the dilator. 2. A percutaneous vascular introducer as recited in claim 1, wherein the dilator has an axial lumen extending therethrough to accommodate a guidewire. 3. A percutaneous vascular introducer as recited in claim 1, wherein the tapered distal portion of the dilator is formed from a plastic material having a radiopaque material combined therewith. 4. A percutaneous vascular introducer as recited in claim 1, wherein the tapered distal portion of the dilator has at least one radiopaque band of material extending thereabout. 5. A percutaneous vascular introducer as recited in claim 1, wherein a distal portion of the sheath has radiopaque characteristics to facilitate placement thereof. 6. A percutaneous vascular introducer as recited in claim 5, wherein the distal portion of the sheath has at least one radiopaque band of material extending thereabout. 7. A percutaneous vascular introducer as recited in claim 1, wherein a distal portion of the sheath is electrically conductive to facilitate cardiac mapping. 8. A percutaneous vascular introducer as recited in claim 7, wherein the proximal portion of the sheath includes an electrical connector operatively associated with the electrically conductive distal portion of the sheath. 9. A percutaneous vascular introducer as recited in claim 1, wherein the proximal portion of the dilator and the proximal portion the sheath include locking means to fasten the dilator to the sheath. 10. A percutaneous vascular introducer as recited in claim 1, wherein the sheath includes means to facilitate separation thereof from the dilator. 11. A percutaneous vascular introducer as recited in claim 1, wherein the proximal portion of the dilator includes an electrical connector operatively associated with the electrically conductive distal tip of the dilator. 12. A percutaneous vascular introducer as recited in claim 1, wherein the proximal portion of the sheath includes a handle to facilitate manipulation thereof. 13. A percutaneous vascular introducer as recited in claim 1, wherein the tapered distal portion of the dilator includes means for measuring arterial blood flow parameters to facilitate placement of the dilator. 14. A percutaneous vascular introducer as recited in claim 1, wherein the tapered distal portion of the dilator is formed form a material that ... | Oscor Inc. | A61B005/04
US06920346 | B2 | 2005-07-19 | Pre-operative device for localizing marked tissues and process using such a device | 2002-04-10 | A pre-operative device includes at least one elongated probe (1) integrating at least one element for detecting radioactive radiation (3) emitted by tissues (4) marked by radioactive isotopes or radioactive colloids, generating a first representative electrical signal (5), at least one element (6) for emitting luminous radiation toward the tissues (4) and at least one element (7) for receiving and transmitting reflected luminous radiation (8) reflected by the tissues (4) marked by vital colorants or magnetic markers, toward at least one element (9) for converting the reflected luminous radiation (8) into a second representative electrical signal (10), electrical connection elements for conveying the first and second electrical signals (5, 10) and, at least one electronic means (11) for processing electronic signals. | 1. Device for pre-operative location of marked tissues, which comprises at least one probe (1) having an elongated shape and integrating at least two of three types of detection means selected from the group comprised by: at least one means (1) for the detection of radioactive radiation (3) emitted by tissues (4) marked with isotopes or radioactive colloids, generating a corresponding representative electrical signal (5); at least one optical detection means comprising at least one means (6) for emitting luminous radiation toward the tissues (4) and at least one means (7) for receiving and transmitting luminous radiation (8) reflected or emitted by fluorescence or luminescence by said tissues (4) marked by vital colorants toward at least one means for conversion (9) of said recovered luminous radiation (8) into a corresponding electrical signal (10); a means (15) for detecting magnetization or measuring magnetic susceptibility of the tissues (4) marked by superparamagnetic substances, generating a corresponding representative electrical signal (16); and in addition to said at least two types of detection means, means for transmitting or conveying said electrical signals (5, 10, 16) and at least one electronic means (11) for processing electrical signals (5, 10, 16) conveyed by the last-named means, so as to display, superimpose, inlay and/or edit a reconstituted image (12), at least partial, of the tissues (4) marked by said colorants and simultaneously supplying a physically perceptible expression (12′) of the representative electrical signal (5) delivered by the detection means (2) of the radioactive radiation (3) emitted by the radioactively marked tissues (4) and/or of the representative electrical signal (16) delivered by the means (15) for detecting the magnetization or measuring the magnetic susceptibility of the tissues (4) marked by superparamagnetic substances. | 1. Device for pre-operative location of marked tissues, which comprises at least one probe (1) having an elongated shape and integrating at least two of three types of detection means selected from the group comprised by: at least one means (1) for the detection of radioactive radiation (3) emitted by tissues (4) marked with isotopes or radioactive colloids, generating a corresponding representative electrical signal (5); at least one optical detection means comprising at least one means (6) for emitting luminous radiation toward the tissues (4) and at least one means (7) for receiving and transmitting luminous radiation (8) reflected or emitted by fluorescence or luminescence by said tissues (4) marked by vital colorants toward at least one means for conversion (9) of said recovered luminous radiation (8) into a corresponding electrical signal (10); a means (15) for detecting magnetization or measuring magnetic susceptibility of the tissues (4) marked by superparamagnetic substances, generating a corresponding representative electrical signal (16); and in addition to said at least two types of detection means, means for transmitting or conveying said electrical signals (5, 10, 16) and at least one electronic means (11) for processing electrical signals (5, 10, 16) conveyed by the last-named means, so as to display, superimpose, inlay and/or edit a reconstituted image (12), at least partial, of the tissues (4) marked by said colorants and simultaneously supplying a physically perceptible expression (12′) of the representative electrical signal (5) delivered by the detection means (2) of the radioactive radiation (3) emitted by the radioactively marked tissues (4) and/or of the representative electrical signal (16) delivered by the means (15) for detecting the magnetization or measuring the magnetic susceptibility of the tissues (4) marked by superparamagnetic substances. 2. Device according to claim 1, characterized in that the probe (1) comprises means (2) for detecting radioactive radiation (3) and at least one optical detection means (6, 7). 3. Device according to claim 1, characterized in that the probe (1) comprises means (15) for detecting magnetization or for measuring magnetic susceptibility and at least one optical detection means (6, 7). 4. Device according to claim 1, characterized in that the probe (1) comprises means (2) for detecting radioactive radiation (3), means (15) for measuring magnetic susceptibility and at least one optical detection means (6, 7). 5. Device according to claim 1, characterized in that it comprises moreover means (14) for storing in digital form data relative to one or several reconstituted images (12) and to said representative electrical signals (5, 16) generated simultaneously, on an electronic storage support, such as a computer memory, as well as if desired data relating to one or more anatomical images (12″) of the operating field enclosing the marked tissues (4), taken previously or subsequently. ... | Eurorad 2-6 (Societe Anonyme) | A61B005/05
US06952606 | B2 | 2005-10-04 | Combined electrical impedance and ultrasound scanner | 2002-07-26 | An electrical impedance scanner having a probe in its contact surface [lacuna] an electrode matrix for detection of the surface distribution of the body current and/or body potentials originating from an opposing electrode which is fitted at some other point on the body of the patient, in which case additional ultrasound transducers (12) for the transmission and reception of ultrasound waves are installed in the contact surface (9) of the probe (3) and are connected to an ultrasound investigation appliance which has a monitor. | 1. An electrical impedance scanner having a probe in its contact surface [lacuna] an electrode matrix for detection of the surface distribution of the body current and/or body potentials originating from an opposing electrode which is fitted at some other point on the body of the patient, characterized in that additional ultrasound transducers (12) for the transmission and reception of ultrasound waves are installed in the contact surface (9) of the probe (3) and are connected to an ultrasound investigation appliance which has a monitor. | 1. An electrical impedance scanner having a probe in its contact surface [lacuna] an electrode matrix for detection of the surface distribution of the body current and/or body potentials originating from an opposing electrode which is fitted at some other point on the body of the patient, characterized in that additional ultrasound transducers (12) for the transmission and reception of ultrasound waves are installed in the contact surface (9) of the probe (3) and are connected to an ultrasound investigation appliance which has a monitor. 2. The electrical impedance scanner as claimed in claim 1, characterized in that the ultrasound transducers (12) are arranged at the side, alongside the electrode matrix (11). 3. The electrical impedance scanner as claimed in claim 1, characterized in that the electrode matrix (11) has cutouts in which the ultrasound transducers (12) are installed. 4. The electrical impedance scanner as claimed in claim 1, wherein the electrical impedance scanning function and the ultrasound scanning function (EI and US) can selectively be operated either individually or combined. 5. The electrical impedance scanner as claimed in claim 1, characterized by a picture-in-picture display with both the electrical impedance scanning function and the ultrasound scanning function being carried out at the same time. 6. The electrical impedance scanner as claimed in claim 1, characterized in that the results of the electrical impedance scanning function and the ultrasound scanning function can be displayed at the same time on one monitor (6) by means of image fusion. 7. The electrical impedance scanner as claimed in claim 1, characterized in that the electronics for the electrical impedance scanning function and the ultrasound scanning function are combined in one appliance (5). 8. The scanning device of claim 1, wherein both the electrode matrix and the ultrasonic transducers are arranged on one said surface of the probe such that the scanning device can be used with the probe surface positioned on skin of a patient. 9. A scanning device comprising: an electronics unit comprising an ultrasound analyzer and an electrical impedance analyzer; a probe comprising ultrasound transducers and an electrode matrix, both the ultrasound transducers and the electrode matrix being disposed on a single contact surface of the probe; means for transmitting ultrasound data and electrical impedance data from the probe to the ultrasound analyzer and electrical impedance analyzer, respectively, in the electronics unit; and an opposing electrode; wherein the electrode matrix is structured and arranged to detect current from the opposing electrode and generate the electrical impedance data, the ultrasound transducers being structured and arranged to transmit and receive ultrasound waves and generate the ultrasound data. 10. The scanning device of claim 9, wherein the ultrasound transducers and the electrode matrix are arranged on the probe surface so ... | Siemens Aktiengesellschaft | A61B005/05 A61B010/00
US06993378 | B2 | 2006-01-31 | Identification by analysis of physiometric variation | 2002-06-25 | Methods involving extraction of information from the inherent variability of physiometrics, including data on cardiovascular and pulmonary functions such as heart rate variability, characteristics of ECG traces, pulse, oxygenation of subcutaneous blood, respiration rate, temperature or CO2 content of exhaled air, heart sounds, and body resonance, can be used to identify individual subjects, particularly humans. Biometric data for use in the methods can be obtained either from contact sensors or at a distance. The methods can be performed alone or can be fused with previous identification algorithms. | 1. A method of identifying a subject, comprising steps of: (a) comparing a first and a second set of biometric parameters extracted from a heart rate variability measurement, wherein (1) the first set of biometric parameters is associated with a first subject, and(2) the second set of biometric parameters is associated with a second subject; and(b) identifying the first subject as the second subject if the first set and second set correlate, and identifying the first subject as different from the second subject if the first and second sets do not correlate. | 1. A method of identifying a subject, comprising steps of: (a) comparing a first and a second set of biometric parameters extracted from a heart rate variability measurement, wherein (1) the first set of biometric parameters is associated with a first subject, and(2) the second set of biometric parameters is associated with a second subject; and(b) identifying the first subject as the second subject if the first set and second set correlate, and identifying the first subject as different from the second subject if the first and second sets do not correlate. 2. The method of claim 1 further comprising the step of analyzing the heart rate variability measurement of the first subject to extract the first set of biometric parameters. 3. The method of claim 2 further comprising the step of acquiring the heart rate variability measurement from the first subject. 4. The method of claim 1 wherein the first set of biometric parameters is compared to a plurality of second sets of biometric parameters, wherein individual members of the plurality are each associated with a different, third subject. 5. The method of claim 1 further comprising applying a low pass filter to the heart rate variability measurement. 6. The method of claim 5 wherein the low pass filter is a 40 Hz low pass filter. 7. The method of claim 1 wherein the heart rate variability is measured using a standoff measuring technique. 8. A method of compiling a database comprising sets of biometric parameters each associated with an individual subject, comprising steps of: (a) acquiring heart rate variability measurements from a plurality of individual subjects;(b) analyzing the heart rate variability measurements to extract sets of biometric parameters, wherein each set of biometric parameters is associated with an individual subject; and(c) compiling the sets of biometric parameters into a database, wherein individual sets of biometric parameters are each correlated with the individual subjects from which the heart rate variability measurements were obtained. 9. The method of claim 8 further comprising applying a low pass filter to the heart rate variability measurements. 10. The method of claim 9 wherein the low pass filter is a 40 Hz low pass filter. 11. A method for identifying a subject according to a periodic physiological process of the subject, the method comprising steps of: measuring the periodic physiological process for a period of time to generate biometric data, the period of time being sufficient such that the biometric data contains data corresponding to a plurality of cycles of the physiological process;extracting a plurality of the cycles of the biometric data; andestimating a first plurality of coefficients based on the plurality of cycles. 12. The method of claim 11, further including the step of determining a periodicity of the biometric data. 13. The method of claim 11, further including the step of normalizing the plurality of cycles, the step of estimating in... | Science Applications International Corporation | A61B005/0402
US07043293 | B1 | 2006-05-09 | Method and apparatus for waveform assessment | 2002-12-24 | An improved method and apparatus for assessing time-variant waveforms and identifying artifacts of interest therein. In one exemplary embodiment, an iterative interval search technique is applied to ECG waveform data in order to identify one or more artifacts (e.g., right atrial and right ventricular “spikes”) within the waveform in conjunction with a fuzzy logic noise threshold analysis. This technique allows for robust artifact identification in waveforms where significant variations in noise and artifact periodicity may exist. The identified artifact(s) may then be used as an input to another process, such as being substituted as an ECG “Q” point for subsequent fiducial point detection within a cardiographic impedance waveform. Apparatus including computer programs for implementing the aforementioned techniques are also disclosed. | 1. A method of identifying an artifact in a cardiographic waveform, comprising: providing at least one cardiographic waveform;determining the shape of at least one of said artifacts within said waveform using an iterative interval search; andidentifying said artifact based at least in part on said shape;wherein said act of identifying further comprises analyzing said shape with respect to a variable noise threshold. | 1. A method of identifying an artifact in a cardiographic waveform, comprising: providing at least one cardiographic waveform;determining the shape of at least one of said artifacts within said waveform using an iterative interval search; andidentifying said artifact based at least in part on said shape;wherein said act of identifying further comprises analyzing said shape with respect to a variable noise threshold. 2. A method of identifying an artifact in a cardiographic waveform, comprising: providing at least one cardiographic waveform;determining the shape said artifacts within said waveform; andidentifying said artifact based at least in part on said shape;wherein said act of identifying further comprises analyzing the noise threshold associated with said at least one waveform using fuzzy logic. 3. A method of identifying at least one artifact within a waveform, comprising: sampling said waveform at a first frequency to produce sampled data;filtering said sampled data according to a first pass band to produce first filtered data;filtering said first filtered data according to a second pass band to produce second filtered data;determining at least one noise threshold based at least in part on said second filtered data; andidentifying at least one artifact from said waveform based at least in part on said at least one noise threshold. 4. The method of claim 3, wherein said act of identifying at least one artifact comprises performing at least one golden section search. 5. The method of claim 3, wherein said act of determining at least one noise threshold comprises analyzing at least a portion of said second filtered using fuzzy logic. 6. A method of processing waveform data, comprising: sampling said waveform at a first frequency to produce sampled data;filtering said sampled data according to at least one filter scheme to produce filtered data; anddetermining at least one noise threshold associated with said waveform based at least in part on fuzzy logic analysis of said filtered data. 7. The method of claim 6, wherein said act of filtering comprises: high pass filtering said sampled data according to a first pass band; andlow pass filtering said high pass filtered data according to a second pass band. 8. The method of claim 6, wherein said act of determining at least one noise threshold comprises: fuzzification of at least a portion of said filtered data to produce a plurality of fuzzy inputs;mapping at least a portion of said fuzzy inputs to a plurality of fuzzy outputs; anddetermining said at least one noise threshold based at least on part on said plurality of fuzzy outputs. 9. The method of claim 8, wherein said act of fuzzification comprises generating said plurality of fuzzy inputs based on at least one membership function and said at least portion of filtered data. 10. The method of claim 8, wherein said act of mapping comprises mapping using a rule based inference (RBI) process. 11. The method of 8, wherein said act of det... | CardioDynamics International Corporation | A61B005/0402
US07297108 | B2 | 2007-11-20 | Disease management system and method including analysis of disease specific changes | 2002-10-01 | A system and method for allowing a patient to access an automated process for managing a specified health problem called a disease. The system performs disease management in a fully automated manner, using periodic interactive dialogs with the patient to obtain health state measurements from the patient, to evaluate and assess the progress of the patient's disease, to review and adjust therapy to optimal levels, and to give the patient medical advice for administering treatment and handling symptom flare-ups and acute episodes of the disease. The medical records are updated, the progression of the disease is stored and tracked, and the patient's preferences for treatment are stored and then used to offer medical advice based on the current state of the disease. A prestored general disease trend curve is compared against a patient specific disease trend curve, and the system makes an automated response such as adjusting therapy. | 1. A computerized critical curve assessment method, comprising:providing a critical curve for a particular disease;providing a plurality of health parameters in an electronic medical record corresponding to a particular patient having been diagnosed with the particular disease; andcomparing at least one of the health parameters to the critical curve to obtain health assessment information. | 1. A computerized critical curve assessment method, comprising:providing a critical curve for a particular disease;providing a plurality of health parameters in an electronic medical record corresponding to a particular patient having been diagnosed with the particular disease; andcomparing at least one of the health parameters to the critical curve to obtain health assessment information. 2. The method of claim 1, wherein a recommendation is made if, based on the comparison, a threshold is reached. 3. The method of claim 2, wherein a disease therapy is adjusted if, based on the comparison, a threshold is reached. 4. The method of claim 1, wherein the health parameters include objective parameters. 5. The method of claim 1, wherein the health parameters include subjective parameters. 6. The method of claim 5, wherein the subjective parameters comprise health data perceived or felt by the particular patient. 7. The method of claim 1, wherein at least one of the health parameters is measured by use of a medical instrument. 8. The method of claim 1, additionally comprising generating a patient specific critical curve, comprising: determining disease specific health parameters from the patient's electronic medical record;adding current health parameters to the patient's electronic medical record; andplotting a patient specific curve consisting of health parameters measured against time. 9. The method of claim 8, additionally comprising performing statistical analysis on the patient specific critical curve. 10. The method of claim 9, additionally comprising: comparing values from the patient specific critical curve to the critical curve to determine if a threshold has been met; andindicating medical treatment if the value threshold has been met. 11. The method of claim 10, additionally comprising comparing the slope of the patient specific critical curve to the slope of the critical curve if the value threshold has not been met, wherein if the critical curve slope exceeds the patient specific critical curve slope, medical treatment is indicated. 12. The method of claim 1, wherein a critical curve is a plurality of critical curves and providing a critical curve for a particular disease comprises providing critical curves for a plurality of health parameters of interest for the particular disease, each critical curve corresponding to a selected one of the health parameters of interest. 13. The method of claim 12, wherein comparing at least one of the health parameters to the critical curve comprises comparing at least one of the health parameters of the electronic medical record to the critical curve of the corresponding health parameter of interest. 14. A system for performing a computerized critical curve assessment, the system comprising: a computer;a storage, in data communication with the computer, configured to store an electronic medical record corresponding to each particular patient of a plurality of patients; anda disease manageme... | Clinical Decision Support, LLC | A61B005/00
US07001346 | B2 | 2006-02-21 | Apparatus and methods for making intraoperative orthopedic measurements | 2002-11-14 | Apparatus and methods for making intraoperative orthopedic measurements are provided. Using telemetry devices attached to a patient, relative measurements of the positions and orientations of the patient's bones may be determined. Using the relative measurements of the positions and orientations of the patient's bones, a differential measurement may be determined in connection with the orthopedic medical procedure. | 1. A method for making orthopedic measurements during an orthopedic medical procedure, the method comprising: attaching a plurality of telemetry devices to a patient's body;prior to a first portion of the orthopedic medical procedure, while moving the patient's body through a first range of motion, tracking movement of the plurality of telemetry devices relative to each other, and calculating the position of at least one anatomical feature of the patient's body relative to the plurality of telemetry devices based on the movement of the plurality of telemetry devices;after the first portion of the orthopedic medical procedure, while moving the patient's body through a second range of motion, tracking movement of the plurality of telemetry devices relative to each other, and calculating the position of the at least one anatomical feature relative to the plurality of telemetry devices based on the movement of the plurality of telemetry devices; anddetermining a differential measurement in connection with the first portion of the orthopedic medical procedure, wherein the differential measurement is based at least in part on the calculated position of the at least one anatomical feature before the first portion of the orthopedic medical procedure and on the calculated position of the at least one anatomical feature after the first portion of the orthopedic medical procedure. | 1. A method for making orthopedic measurements during an orthopedic medical procedure, the method comprising: attaching a plurality of telemetry devices to a patient's body;prior to a first portion of the orthopedic medical procedure, while moving the patient's body through a first range of motion, tracking movement of the plurality of telemetry devices relative to each other, and calculating the position of at least one anatomical feature of the patient's body relative to the plurality of telemetry devices based on the movement of the plurality of telemetry devices;after the first portion of the orthopedic medical procedure, while moving the patient's body through a second range of motion, tracking movement of the plurality of telemetry devices relative to each other, and calculating the position of the at least one anatomical feature relative to the plurality of telemetry devices based on the movement of the plurality of telemetry devices; anddetermining a differential measurement in connection with the first portion of the orthopedic medical procedure, wherein the differential measurement is based at least in part on the calculated position of the at least one anatomical feature before the first portion of the orthopedic medical procedure and on the calculated position of the at least one anatomical feature after the first portion of the orthopedic medical procedure. 2. The method of claim 1, further comprising providing on a display an output that corresponds to the differential measurement. 3. The method of claim 1, wherein a first one of the plurality of telemetry devices is a first electromagnetic receiver, wherein a second one of the plurality of telemetry devices is a second electromagnetic receiver, and wherein a third one of the plurality of telemetry devices is a magnetic field generator. 4. The method of claim 1, wherein the plurality of telemetry devices consists of a first telemetry device and a second telemetry device, wherein the first telemetry device is an electromagnetic receiver, and wherein the second telemetry device is a magnetic field generator. 5. The method of claim 1, wherein the plurality of telemetry devices are a plurality of radio frequency telemetry devices. 6. The method of claim 1, wherein the orthopedic medical procedure is a hip joint replacement procedure, and wherein the differential measurement is a leg length difference. 7. The method of claim 1, wherein the orthopedic medical procedure is a hip joint replacement procedure, and wherein the differential measurement is an offset. 8. The method of claim 1, wherein the step of tracking movement of the plurality of telemetry devices prior to the first portion of the orthopedic medical procedure comprises: receiving and recording a first set of multiple relative positions and orientations of the plurality of telemetry devices. 9. The method of claim 8, wherein the step of calculating the position of the at least one anatomical feature prior to the first ... | Michael R. White | A61B005/103 A61B005/117
US07236936 | B2 | 2007-06-26 | Security infusion pump with bar code reader | 2002-12-30 | A medical infusion pump with a bar code reader having simplified and enhanced security features. The infusion pump with bar code reader is comprised of an infusion pump with a housing, a control display screen and control software for operating the infusion pump according to input infusion data. A bar code reader is contained in the infusion pump housing and operatively attached to communicate scanned bar code infusion data to the pump display screen and to the control software. At least one initialization button for starting execution of a bar code scanning program is operatively connected through circuitry in the pump, the program activating the bar code reader to scan for an authorized user code, for an authorized user ID code, for a patient ID code, and for infusion data compatible for operating the infusion pump. A confirmation program and confirmation buttons manually operable by the authorized user. Acceptable scanning of authorization code, authorized user ID code, patient ID code and infusion data must be manually confirmed by the authorized user prior to activating the pump to infuse fluids according to the infusion data. | 1. A medical infusion pump having simplified and enhanced security features comprising: a pump housing;a pumping mechanism in the pump housing including electronic controls;a display screen on the pump housing;a manual entry control panel on the pump housing; control software operatively connected to the electronic controls of the pumping mechanism to operate the pumping mechanism according to input infusion data;a bar code reader contained in the pump housing and operatively attached to communicate scanned bar code infusion data to said display screen and to said control software, said bar code reader including means for data capture including capture ofan authorized user code, an authorized user ID code; a patient ID code; and infusion operating control data; and the bar code reader further including means for activating the bar code reader to scan for the information to be captured including for an authorized user code, for an authorized user ID code, for a patient ID code, and for medicinal infusion operating data including intended patient ID data; andmeans for automatically confirming a match of the scanned authorization code to an acceptable authorization code and for confirming a match of the scanned patient ID code to the intended patient ID in the scanned medicinal infusion data before operating the infusion pump according to said medicinal infusion data; anda display and control program connected to the bar code reader and the display panel to display the scanned infusion data and requiring manual entry by the authorized user to validate the displayed medicinal infusion data prior to activating said pump mechanism to infuse medicines according to said scanned infusion data. | 1. A medical infusion pump having simplified and enhanced security features comprising: a pump housing;a pumping mechanism in the pump housing including electronic controls;a display screen on the pump housing;a manual entry control panel on the pump housing; control software operatively connected to the electronic controls of the pumping mechanism to operate the pumping mechanism according to input infusion data;a bar code reader contained in the pump housing and operatively attached to communicate scanned bar code infusion data to said display screen and to said control software, said bar code reader including means for data capture including capture ofan authorized user code, an authorized user ID code; a patient ID code; and infusion operating control data; and the bar code reader further including means for activating the bar code reader to scan for the information to be captured including for an authorized user code, for an authorized user ID code, for a patient ID code, and for medicinal infusion operating data including intended patient ID data; andmeans for automatically confirming a match of the scanned authorization code to an acceptable authorization code and for confirming a match of the scanned patient ID code to the intended patient ID in the scanned medicinal infusion data before operating the infusion pump according to said medicinal infusion data; anda display and control program connected to the bar code reader and the display panel to display the scanned infusion data and requiring manual entry by the authorized user to validate the displayed medicinal infusion data prior to activating said pump mechanism to infuse medicines according to said scanned infusion data. 2. The medical infusion pump of claim 1 further comprising: means for manually controlling the infusion pumping operation from the manual control panel; andmeans for selecting between the barcode reader for entry of infusion data to control infusion by the pump and the manual control panel for entry of the infusion data to control infusion by the pump. 3. The medical infusion pump of claim 1 further comprising means for automatically requiring a scanned authorized user ID code before operating the infusion pump according to said medicinal infusion data. 4. A medical infusion pump bar code reading system having simplified and enhanced security features comprising: an infusion pump with a housing, a control display screen and control software for operating the infusion pump according to input infusion data;a bar code reader contained in said infusion pump housing and operatively attached to communicate scanned bar code infusion data to said control display screen and to said control software;means for manually controlling the infusion pumping operation from the manual control panel; andmeans for selecting between the barcode reader for entry of infusion data to control infusion by the pump and the manual control panel for entry of the infusion data to control infus... | B. Braun Medical, Inc. | A61B005/00
US07065397 | B2 | 2006-06-20 | Multi-parameter acquisition of ECG and related physiologic data employing multi-parameter sensor and conventional ECG lead conductors, and enabled for remote operational management communication | 2002-06-18 | ECG-related adaptor/sensor structure and methodology which enables various selectable operational communication to occur both unidirectionally and bidirectionally between the adaptor/sensor structure, and remote structure which is intended to receive ECG-related physiologic data from the adaptor/sensor structure. | 1. Adaptor/sensor apparatus constructed for operative connection to a set of ECG lead conductors having distal ends for collecting, from a subject, plural-parameter physiologic data including EGG data, said apparatus comprising an adaptor constructed for direct connection operatively to the distal end of such EGG lead conductors,unitary plural-parameter sensor structure organized internally with plural specific parameter sensor devices in the forms of data-collection transducer structures that are designed each to collect, respectively, different physiologic parameter data from a subject's common anatomical site, which sensor structure is discardably connectable operatively to said adaptor to communicate information therewith and therebetween, andremotely accessible electrical response structure operatively connected to at least one of said unitary sensor structure and/or said adaptor, capable of being interrogated electrically over conductors in such EGG lead conductors, and to provide, with respect to such an interrogation, an interrogation-response report regarding selected informational characteristics relating to said adaptor and/or sensor structure. | 1. Adaptor/sensor apparatus constructed for operative connection to a set of ECG lead conductors having distal ends for collecting, from a subject, plural-parameter physiologic data including EGG data, said apparatus comprising an adaptor constructed for direct connection operatively to the distal end of such EGG lead conductors,unitary plural-parameter sensor structure organized internally with plural specific parameter sensor devices in the forms of data-collection transducer structures that are designed each to collect, respectively, different physiologic parameter data from a subject's common anatomical site, which sensor structure is discardably connectable operatively to said adaptor to communicate information therewith and therebetween, andremotely accessible electrical response structure operatively connected to at least one of said unitary sensor structure and/or said adaptor, capable of being interrogated electrically over conductors in such EGG lead conductors, and to provide, with respect to such an interrogation, an interrogation-response report regarding selected informational characteristics relating to said adaptor and/or sensor structure. 2. Adaptor/sensor apparatus constructed for operative connection to a conventional set of EGG lead conductors having distal ends for collecting, from a subject, plural-parameter physiologic data including EGG data, said apparatus comprising an adaptor constructed for direct connection operatively to the distal ends of such EGG lead conductors,unitary plural-parameter sensor structure organized internally with plural-specific-parameter sensor devices in the forms of data-collection transducer structures that are designed each to collect, respectively, different physiologic parameter data from a subject's common anatomical site, which sensor structure is discardably connectable operatively to said adaptor to communicate information therewith and therebetween, andremotely accessible communication-enabling structure operatively disposed within said unitary sensor structure for the purpose of enabling at least unidirectional operative communication between the sensor structure and an outside communication source for the purpose of addressing selected operational qualities and characteristics of components within the sensor structure. 3. The apparatus of claim 2 wherein said communication-enabling structure permits bidirectional communication. 4. Adaptor/sensor apparatus constructed for operative connection to a set of conventional ECG lead conductors having distal ends for collecting, from a person, plural-parameter physiologic data including ECG data, said apparatus comprising: an adaptor which is adapted for direct connection operatively to the distal ends of such ECG lead conductors,a unitary plural-parameter sensor organized internally with plural specific parameter sensor devices in the forms of data-collection transducer structures that are designed each to collect, respectively, different p... | Inovise Medical, Inc. | A61B005/02
US07027855 | B2 | 2006-04-11 | R-peak detection corrections for improved cardiac gated reconstruction | 2002-08-27 | A method and apparatus for use with a gated CT imaging system for correcting for Rpeak errors in a heart beat signal, the method including identifying heart rate corresponding to each specific heart beating cycle and comparing the heart rate for each specific cycle to moving averages to determine when a likely Rpeak error has occurred and, when a likely error has occurred, modifying the heart beat signal to generate a more accurate signal and thereby render more accurate gated images. | 1. A method to be used with a CT imaging system including a heart beat monitor that senses R-peaks and generates a heart beat signal including R-peak pulses where each R-peak pulse indicates a sensed peak heart cycle magnitude, the method for correcting for heart beat signal errors corresponding to data collected over the course of several heart cycles, the method comprising the steps of: for each specific heart cycle where the specific heart cycle corresponds to a period between first and second R-peak pulses occurring at times t(1) and t(2), respectively:identifying a heart rate;using the heart rate to determine when an R-peak error has likely occurred; andwhen an R-peak error has likely occurred, modifying the heart beat signal Rpeak pulse times by at least one of storing and eliminating at least one R-peak pulses;wherein the step of using includes the step of determining when at least one undetected R-peak has likely occurred between times t(1) and t(2) and where the step of modifying includes, when at least one undetected R-peak has likely occurred, storing at least one estimated R-peak pulse time t(e) corresponding to the heart beat signal between times t(1) and t(2). | 1. A method to be used with a CT imaging system including a heart beat monitor that senses R-peaks and generates a heart beat signal including R-peak pulses where each R-peak pulse indicates a sensed peak heart cycle magnitude, the method for correcting for heart beat signal errors corresponding to data collected over the course of several heart cycles, the method comprising the steps of: for each specific heart cycle where the specific heart cycle corresponds to a period between first and second R-peak pulses occurring at times t(1) and t(2), respectively:identifying a heart rate;using the heart rate to determine when an R-peak error has likely occurred; andwhen an R-peak error has likely occurred, modifying the heart beat signal Rpeak pulse times by at least one of storing and eliminating at least one R-peak pulses;wherein the step of using includes the step of determining when at least one undetected R-peak has likely occurred between times t(1) and t(2) and where the step of modifying includes, when at least one undetected R-peak has likely occurred, storing at least one estimated R-peak pulse time t(e) corresponding to the heart beat signal between times t(1) and t(2). 2. The method of claim 1 wherein preceding cycles are heart cycles that precede the specific heart cycle and the step of using further includes the steps of identifying a moving average rate corresponding to the moving average heart rate of a sub-set of the preceding cycles and determining that at least one undetected R-peak pulse has likely occurred between times t(1) and t(2) when the heart rate is less than Y % of the moving average rate. 3. The method of claim 2 further including the step of identifying a moving average period corresponding to the moving average heart cycle period of a sub-set of the preceding cycles and wherein the step of determining further includes the steps of identifying a preceding heart rate corresponding to at least one preceding heart cycle, identifying the estimated R-peak time t(e) as the time that precedes time t(2) by the moving average period, identifying an estimated heart rate corresponding to times t(e) and t(1) and determining that at least one undetected R-peak pulse has likely occurred between times t(1) and t(2) when the estimated heart rate is less than Z times the preceding heart rate. 4. The method of claim 3 further including the step of: (a) determining when at least another undetected R-peak has likely occurred between times t(1) and t(2); and(b) when at least another undetected R-peak has likely occurred, storing another estimated R-peak pulse time corresponding to the heart beat signal between times t(e) and t(1). 5. The method of claim 4 further including the step of repeating steps (a) and (b) until it is unlikely that another undetected R-peak pulse has occurred between times t(1) and t(2). 6. The method of claim 4 wherein the step of determining when at least another undetected R-peak has likely occurred further inclu... | General Electric Company | A61B005/04
US06931271 | B2 | 2005-08-16 | System for adaptively deriving ECG chest lead signal data | 2002-11-01 | A system provides synthesized chest lead signals of a conventional 12 lead electrocardiogram (ECG) signal set. The system adaptively derives ECG chest lead signals for a plurality of the six ECG chest lead signals of a conventional 12 lead ECG signal set. The system includes an input processor for receiving data comprising two measured ECG chest lead signals comprising a particular two of the six ECG chest lead signals of a conventional 12 lead ECG signal set and receiving associated information identifying the particular two signals of the six ECG chest lead signals. A data processor in the system uses the information identifying the particular two signals of the six ECG chest lead signals for identifying and selecting coefficients from a plurality of stored coefficients and applying a transformation using the selected coefficients to the data comprising the two measured ECG chest lead signals to derive data representing a plurality of non-measured ECG chest lead signals. An output processor processes the data representing the plurality of non-measured ECG chest lead signals for output. The selected coefficients are patient non-specific. | 1. A system for adaptively deriving ECG chest lead signals for a plurality of the six ECG chest lead signals of a conventional 12 lead ECG signal set, comprising: an input processor for receiving data comprising two measured ECG chest lead signals comprising a particular two of the six ECG chest lead signals of a conventional 12 lead ECG signal set and receiving associated information identifying said particular two signals of said six ECG chest lead signals; a data processor for using said information identifying said particular two signals of said six ECG chest lead signals for identifying and selecting coefficients from a plurality of stored coefficients, and applying a transformation using said selected coefficients to said data comprising said two measured ECG chest lead signals to derive data representing a plurality of non-measured ECG chest lead signals; and an output processor for processing said data representing said plurality of non-measured ECG chest lead signals for output. | 1. A system for adaptively deriving ECG chest lead signals for a plurality of the six ECG chest lead signals of a conventional 12 lead ECG signal set, comprising: an input processor for receiving data comprising two measured ECG chest lead signals comprising a particular two of the six ECG chest lead signals of a conventional 12 lead ECG signal set and receiving associated information identifying said particular two signals of said six ECG chest lead signals; a data processor for using said information identifying said particular two signals of said six ECG chest lead signals for identifying and selecting coefficients from a plurality of stored coefficients, and applying a transformation using said selected coefficients to said data comprising said two measured ECG chest lead signals to derive data representing a plurality of non-measured ECG chest lead signals; and an output processor for processing said data representing said plurality of non-measured ECG chest lead signals for output. 2. A system according to claim 1 wherein said input processor also receives data representing conventional 12 lead ECG signal set signals I and II, and said data processor applies said transformation using said selected coefficients to said signals I and II in deriving said data representing a plurality of non-measured remaining ECG chest lead signals. 3. A system according to claim 1 wherein said plurality of non-measured ECG chest lead signals comprises the four remaining non-measured ECG chest lead signals of said six conventional ECG chest lead signals. 4. A system according to claim 1 wherein said plurality of stored coefficients comprises a plurality of sets of coefficients and an individual set of coefficients is associated with a corresponding particular selection of two of the six conventional ECG chest lead signals. 5. A system according to claim 4 wherein said plurality of sets of coefficients comprises thirty sets of coefficients associated with thirty available ways of selecting two of the six conventional ECG chest lead signals. 6. A system according to claim 5 wherein said plurality of sets of coefficients comprises thirty coefficient matrices. 7. A system according to claim 1 wherein said data processor identifies and selects coefficients using said information identifying the particular two signals of the six conventional ECG chest lead signals for any patient. 8. A system according to claim 1 wherein said transformation is a linear transformation. 9. A system according to claim 8 wherein said selected coefficients comprise a coefficient matrix and said data processor performs a matrix multiplication operation in performing said linear transformation. 10. A system for adaptively deriving ECG chest lead signals for a plurality of the six ECG chest lead signals of a conventional 12 lead ECG signal set, comprising: an input processor for receiving data comprising two measured ECG chest lead signals comprising a particular two of th... | Draeger Medical Systems, Inc | A61B005/04
US06871089 | B2 | 2005-03-22 | Portable ECG monitor and method for atrial fibrillation detection | 2002-06-05 | An apparatus and method for detecting an atrial fibrillation through monitoring the R—R intervals of a patient's QRS complexes. Ratios of the current R—R interval are made to previous R—R intervals. Multiple moving averages are taken of the ratio results. The ratios are compared to a validating threshold and if the particular ratio is within a selected range, a moving average is calculated with inclusion of the present R—R interval. If the ratio is outside the range, a moving average is calculated without inclusion of the current R—R interval. The moving averages are combined and this difference average is compared to a trigger threshold. If the difference average exceeds the threshold, an atrial fibrillation is determined to exist, a memory trigger is provided, and the QRS complex data within a selected period of time about the atrial fibrillation event are recorded and removed from overwrite status in the memory. A reset threshold, differing from the trigger threshold, is used and if the difference average falls below the reset threshold, the method and apparatus are re-enabled to record a new AFIB event. | 1. A method of analyzing the cardiac rhythms of a patient to detect the presence of atrial fibrillation, the method comprising: measuring R—R intervals between a plurality of QRS complexes of the patient, including present and preceding QRS complexes; comparing the present R—R interval to the preceding R—R interval to result in a first ratio; comparing the present R—R interval to the second preceding R—R interval to result in a second ratio; comparing the first ratio and the second ratio to a validating threshold; calculating a first average of R—R intervals of the plurality of QRS complexes without including the present R—R interval in the first average if the first ratio exceeds the validating threshold; calculating the first average of R—R intervals with inclusion of the present R—R interval in the first average if the first ratio is less than or equal to the validating threshold; calculating a second average of R—R intervals of the plurality of QRS complexes without inclusion of the present R—R interval in the second average if the second ratio exceeds the validating threshold; calculating the second average of R—R intervals with inclusion of the present R—R interval in the second average if the second ratio is less than or equal to the validating threshold; and determining the presence of an atrial fibrillation by analyzing the first and second averages. | 1. A method of analyzing the cardiac rhythms of a patient to detect the presence of atrial fibrillation, the method comprising: measuring R—R intervals between a plurality of QRS complexes of the patient, including present and preceding QRS complexes; comparing the present R—R interval to the preceding R—R interval to result in a first ratio; comparing the present R—R interval to the second preceding R—R interval to result in a second ratio; comparing the first ratio and the second ratio to a validating threshold; calculating a first average of R—R intervals of the plurality of QRS complexes without including the present R—R interval in the first average if the first ratio exceeds the validating threshold; calculating the first average of R—R intervals with inclusion of the present R—R interval in the first average if the first ratio is less than or equal to the validating threshold; calculating a second average of R—R intervals of the plurality of QRS complexes without inclusion of the present R—R interval in the second average if the second ratio exceeds the validating threshold; calculating the second average of R—R intervals with inclusion of the present R—R interval in the second average if the second ratio is less than or equal to the validating threshold; and determining the presence of an atrial fibrillation by analyzing the first and second averages. 2. The method of analyzing cardiac rhythms of claim 1 wherein the steps of calculating the first average comprise calculating exponential moving averages for the first average. 3. The method of analyzing cardiac rhythms of claim 1 wherein the steps of calculating the second average comprises calculating exponential moving averages for the second average. 4. The method of analyzing cardiac rhythms of claim 1 wherein the step of comparing the first ratio and the second ratio to a validating threshold comprises selecting the validating threshold so that R—R intervals caused by premature ventricular contractions and other irregular rhythms, other than AFIB, are excluded from the first and second averages. 5. The method of analyzing cardiac rhythms of claim 4 wherein the first and second ratios are formed in accordance with the following: dRR1⁡(n)=RR⁡(n)RR⁡(n-1)dRR2⁡(n)=RR⁡(n)RR⁡(n-2)where: RR(n)=the R—R interval between the current and preceding QRS complexes RR(n−1)=the R—R interval between the preceding and second preceding QRS complexes RR(n−2)=the R—R interval between the second preceding and third preceding QRS complexes dRRx(n)=the current ratio of the respective R—R intervals. 6. The method of analyzing cardiac rhythms of claim 1 further comprising the step of determining a difference average in accordance with: Dev′=Dev1−α∥Dev1−Dev2∥where: Dev′=the difference average Dev1=the first average α=a compensating factor preventing false triggering on irregular rhythms other than AFIB Dev2=the second average. 7. The method of analyzing cardiac rhythms of claim 1 further comprising: com... | Card Guard Technologies, Inc. | A61B005/0404 A61B005/0456 A61B005/046
US06945943 | B2 | 2005-09-20 | Analyte concentration determination devices and methods of using the same | 2002-05-01 | Devices and methods for analyte concentration determination are provided. The subject devices have a housing having an aperture, a lancing element having a lancet held therein disposed within the housing, a manner for activating the lancing element to displace the lancet to provide an incision in an area of skin to provide physiological sample, and a manner for determining whether a sufficient amount of the physiological sample is present at the surface of the skin. The subject methods include lancing an area of skin to provide physiological sample at the surface of the area of skin, illuminating the physiological sample present at the surface of skin, detecting light reflected from the physiological sample, and determining whether the physiological sample is present in an amount sufficient for analyte concentration determination based upon the detected light. The subject invention also includes kits for use in practicing the subject methods. | 1. A method for determining the concentration of an analyte in a physiological sample, said method comprising: (a) lancing an area of skin to provide a physiological sample at the surface of said area of skin; (b) illuminating said physiological sample present at said surface of skin; (c) detecting light reflected from said physiological sample present at said surface of skin; and (d) determining whether said physiological sample is present at said surface of said skin in an amount sufficient for analyte concentration determination based upon said detected reflected light. | 1. A method for determining the concentration of an analyte in a physiological sample, said method comprising: (a) lancing an area of skin to provide a physiological sample at the surface of said area of skin; (b) illuminating said physiological sample present at said surface of skin; (c) detecting light reflected from said physiological sample present at said surface of skin; and (d) determining whether said physiological sample is present at said surface of said skin in an amount sufficient for analyte concentration determination based upon said detected reflected light. 2. The method according to claim 1, further comprising determining the concentration of an analyte in said physiological sample determined to be present in a sufficient amount, wherein analyte concentration is not determined if an insufficient amount of sample is not present. 3. The method according to claim 2, wherein said analyte concentration determination comprises moving a tester in contact with said physiological sample determined to be present in a sufficient amount, wherein a tester is not moved into contact with said physiological sample determined to be present in an insufficient amount. 4. The method according to claim 1, further comprising increasing the amount of physiological fluid expressed from said incised area of skin by applying a force to the perimeter of incised said area, thereby increasing the amount of said physiological fluid at said incised site. | LifeScan, Inc. | A61B005/00 B65D081/00
US06950687 | B2 | 2005-09-27 | Isolation and communication element for a resposable pulse oximetry sensor | 2003-01-24 | A sensor has a connector adapted to electrically communicate with a physiological measurement instrument. A breakable conductor incorporated by the sensor transitions from a continuity state to a discontinuity state as the result of sensor wear. An isolation and communications element (ICE) has an instrument port and an electrically isolated loop port. The instrument port is in communications with the connector and the loop port is in communications with the breakable conductor. The ICE generates a control output responsive to the discontinuity state to render the sensor inoperable. | 1. A sensor comprising: a connector adapted to electrically communicate with a physiological measurement instrument; a breakable conductor that transitions from a continuity state to a discontinuity state as the result of sensor wear; and an isolation and communications element (ICE) having an instrument port and an electrically isolated loop port, said instrument port in communications with said connector and said loop port in communications with said breakable conductor, wherein: said ICE further comprises a control output responsive to said discontinuity state, said sensor further comprises an electrical component, said ICE being connected in parallel with said connector and said electrical component, and said component being operable from said instrument via said connector according to said control output. | 1. A sensor comprising: a connector adapted to electrically communicate with a physiological measurement instrument; a breakable conductor that transitions from a continuity state to a discontinuity state as the result of sensor wear; and an isolation and communications element (ICE) having an instrument port and an electrically isolated loop port, said instrument port in communications with said connector and said loop port in communications with said breakable conductor, wherein: said ICE further comprises a control output responsive to said discontinuity state, said sensor further comprises an electrical component, said ICE being connected in parallel with said connector and said electrical component, and said component being operable from said instrument via said connector according to said control output. 2. The sensor according to claim 1 further comprising a switch in series between said component and said connector, said switch responsive to said control output so as to disconnect said component from said connector. 3. The sensor according to claim 2 wherein said ICE further comprises a converter which supplies DC power to said ICE, said converter connected to said instrument port so as to draw modulated power from said instrument. 4. The sensor according to claim 2 wherein said ICE further comprises: a processor having an input indicating said discontinuity state and providing said control output in response; a transceiver in communications with said processor and said instrument port and adapted to transfer a plurality of data items between said instrument and said processor; and a memory readable and writeable by said processor and adapted to store and retrieve said data items. 5. The sensor according to claim 2 wherein said ICE further comprises: an opto-isolator driver connected to said loop port and adapted to supply a current to said breakable conductor; and an opto-isolator receiver connected to said loop port and adapted to receive said current so as to detect said discontinuity state. 6. A sensor comprising: a connector adapted to electrically communicate with a physiological measurement instrument; a breakable conductor configured as a current loop that becomes an open circuit as the result of sensor wear, said conductor electrically isolated from said connector; an electrical component operable from said instrument via said connector said electrical component rendered inoperable in response to said open circuit; a driver configured to provide a current to said breakable conductor; and a receiver responsive to said current so as to provide a control output indicating said open circuit, wherein said driver and said receiver are adapted to provide electrical isolation between said conductor and said connector. 7. The sensor according to claim 6 further comprising a switch between said component and said connector, said switch actuated by said control output so as to disconnect said component from said connector. 8. The... | Masimo Corporation | A61B005/00
US06941166 | B2 | 2005-09-06 | Software controlled electrophysiology data management | 2001-08-30 | Methods for conducting an electrophysiology procedure using a programmed machine are described, including methods for managing discrete data capture requests received throughout the course of an electrophysiology procedure and for discretizing the display of independent data capture events that occur during the course of the electrophysiology procedure. These methods enable an operator to review location maps of sites visited during the course of a procedure and to review data from discrete moments in the procedure and to play back data over a sequence of times throughout the procedure, one moment at a time. The invention further includes software constructs that enable data management. | 1. In a programmed system having a memory, a method for conducting an electrophysiology procedure using one or more catheters at least one of which is a roving catheter which includes: one or more first means supported on the catheter for obtaining physiologic data; and one or more second means for obtaining location data concerning a location of the first means when the catheter is positioned indwelling, the method comprising the steps of: a) displaying the location data obtained by the second means in a first window; b) displaying the physiologic data obtained by the first means in a second window; c) freezing the first and second windows to a first moment in time during the electrophysiology procedure in response to operator action; d) disassociating in the memory of the programmed system any physiologic data from location data for any moment in time other than the first moment in time; and e) creating a set of data points for display in the first window, each data point associating the physiologic data obtained by each first means at the first moment in time with the location data obtained by the second means at the first moment in time. | 1. In a programmed system having a memory, a method for conducting an electrophysiology procedure using one or more catheters at least one of which is a roving catheter which includes: one or more first means supported on the catheter for obtaining physiologic data; and one or more second means for obtaining location data concerning a location of the first means when the catheter is positioned indwelling, the method comprising the steps of: a) displaying the location data obtained by the second means in a first window; b) displaying the physiologic data obtained by the first means in a second window; c) freezing the first and second windows to a first moment in time during the electrophysiology procedure in response to operator action; d) disassociating in the memory of the programmed system any physiologic data from location data for any moment in time other than the first moment in time; and e) creating a set of data points for display in the first window, each data point associating the physiologic data obtained by each first means at the first moment in time with the location data obtained by the second means at the first moment in time. 2. The method of claim 1, wherein the first window and the second window are active panes presented on a single display. 3. In a programmed system, a method for managing discrete data capture requests received throughout the course of an electrophysiology procedure, comprising the steps of: a) obtaining physiologic data from sites adjacent one or more locations of a roving catheter distal tip portion; b) obtaining location data concerning a location of each of said sites; c) providing an index value in response to a request, at any given point in time, to capture the data obtained from steps (a) and (b); d) associating in a first memory the obtained physiologic data with the index value; e) associating in a second memory the obtained location data with the index value; and f) updating the index value prior to repeating steps (c) through (e). 4. The method of claim 3, including the further step of creating a set of data points, the data points associating the physiologic data in the first memory with the location data in the second memory for any one of said any given point in time. 5. The method of claim 4, including the further step of displaying the set of data points in a first window. 6. The method of claim 5, including the further step of displaying in the first window a map of the locations visited during the course of the electrophysiology procedure using the obtained location data for a plurality of index values. 7. In a programmed system, a method for managing discrete data capture requests received throughout the course of an electrophysiology procedure, comprising the steps of: a) obtaining physiologic data from sites adjacent one or more locations of a roving catheter distal tip portion; b) obtaining location data concerning a location of each of said sites; c) providing an index value in ... | C.R. Bard, Inc. | A61B005/04
US07016732 | B2 | 2006-03-21 | Cardiac rhythm management system with remotely activated capture verification for CHF and other patients | 2003-01-07 | A cardiac rhythm management device capable of delivering multiple uni-chamber stimulation pulses to a patient's heart and suitable for verifying capture independently for each uni-chamber stimulation pulse. The uni-chamber capture verification mode of the cardiac rhythm management device may be activated via telemetry or by applying a magnetic field proximate the device. During the capture verification mode, bi-chamber pacing, for example, may precede or follow uni-chamber pacing to allow for pacing support. Also, the energy levels of the pacing stimulus over several beats may be varied, thereby verifying the programmed safety margins. | 1. A method for verifying capture of stimulation pulses delivered to pre-selected chambers of a patient's heart, by a cardiac rhythm management device of the type having means for sensing cardiac electrogram signals, a programmable controller coupled to receive sensed electrogram signals, and stimulation means for applying cardiac stimulation pulses to a patient's heart in response to control signals from the controller, said method comprising the steps of: (a) setting predetermined parameters of the cardiac rhythm management device;(b) delivering a predetermined number of first stimulation pulses to an electrically continuous area of the heart in accordance with the predetermined parameters prior to delivering a second stimulation pulse to said area;(c) identifying depolarization waveforms corresponding to each stimulation pulse; and(d) determining whether each stimulation results in capture. | 1. A method for verifying capture of stimulation pulses delivered to pre-selected chambers of a patient's heart, by a cardiac rhythm management device of the type having means for sensing cardiac electrogram signals, a programmable controller coupled to receive sensed electrogram signals, and stimulation means for applying cardiac stimulation pulses to a patient's heart in response to control signals from the controller, said method comprising the steps of: (a) setting predetermined parameters of the cardiac rhythm management device;(b) delivering a predetermined number of first stimulation pulses to an electrically continuous area of the heart in accordance with the predetermined parameters prior to delivering a second stimulation pulse to said area;(c) identifying depolarization waveforms corresponding to each stimulation pulse; and(d) determining whether each stimulation results in capture. 2. The method according to claim 1, wherein determination of whether each stimulation results in capture is initiated by an externally applied magnetic field which activates a switch coupled to a controller contained within said cardiac rhythm management device. 3. The method according to claim 2, wherein said first stimulation pulses are directed to a right ventricle of the patient and the second stimulation pulse is directed to a left ventricle of the patient. 4. The method according to claim 1, wherein said first stimulation pulses are directed to a right ventricle of the patient and the second stimulation pulse is directed to a left ventricle of the patient. 5. The method according to claim 1, wherein said first stimulation pulses are directed to a ventricle of the patient and the second stimulation pulse is directed to a ventricle of the patient. 6. The method according to claim 1, wherein said first stimulation pulses are directed to a right atrium of the patient and the second stimulation pulse is directed to a left atrium of the patient. 7. The method according to claim 1, wherein said first stimulation pulses are directed to an atrium of the patient and the second stimulation pulse is directed to an atrium of the patient. 8. The method according to claim 1, wherein a stimulation energy output of the first stimulation pulses is varied by a predetermined amount. 9. The method according to claim 1, wherein a stimulation energy of the second stimulation pulse is varied by a predetermined amount. | Cardiac Pacemakers, Inc. | A61N001/365 A61B005/04
US07134994 | B2 | 2006-11-14 | Multipurpose host system for invasive cardiovascular diagnostic measurement acquisition and display | 2002-05-20 | A multifunctional invasive cardiovascular diagnostic measurement host is disclosed that interfaces a variety of sensor devices, such as guide wire-mounted pressure sensors, flow sensors, temperature sensors, etc, and provides a multi-mode graphical user interface providing a plurality of displays in accordance with the various types of sensors and measurements rendered by the sensors. | 1. A multipurpose host system for invasive cardiovascular diagnostic measurement acquisition and display incorporating a component based arrangement, the system comprising: an external input signal bus interface for receiving data arising from cardiovascular diagnostic measurement sensors;a plurality of measurement processing components, that operate at a user mode level in the multipurpose host system, for facilitating receiving darn of particular sensor types and rendering diagnostic measurement parameter values according to the received data;a multi-mode graphical user interface host comprising diagnostic measurement user interfaces including display components corresponding to data output rendered by specified ones of the plurality of measurement processing components; andone or more kernel mode drivers that extract processed sensor data from as, peripheral interface card providing a hardware interface for one or more invasive diagnostic measurement devices. | 1. A multipurpose host system for invasive cardiovascular diagnostic measurement acquisition and display incorporating a component based arrangement, the system comprising: an external input signal bus interface for receiving data arising from cardiovascular diagnostic measurement sensors;a plurality of measurement processing components, that operate at a user mode level in the multipurpose host system, for facilitating receiving darn of particular sensor types and rendering diagnostic measurement parameter values according to the received data;a multi-mode graphical user interface host comprising diagnostic measurement user interfaces including display components corresponding to data output rendered by specified ones of the plurality of measurement processing components; andone or more kernel mode drivers that extract processed sensor data from as, peripheral interface card providing a hardware interface for one or more invasive diagnostic measurement devices. 2. The multipurpose host system of claim 1 wherein the plurality of measurement processing components includes a blood pressure processing component and a blood velocity processing component. 3. The multipurpose host system of claim 2 wherein the plurality of measurement processing components includes an invasive device position sensor processing component. 4. The multipurpose host system of claim 2 wherein the diagnostic measurement user interfaces include a pressure measurement display interface wherein a ratio of a first measured pressure and a second measured pressure is rendered. 5. The multipurpose host system of claim 2 wherein the diagnostic measurement user interfaces include a pressure measurement display interface wherein a pressure gradient between a first measured pressure and a second measured pressure is rendered. 6. The multipurpose host system of claim 2 wherein the diagnostic measurement user interfaces include a flow measurement display interface wherein a ratio of flow velocity under high flow conditions and low velocity conditions is rendered. 7. The multipurpose host system of claim 1 wherein the plurality of measurement processing components includes a temperature processing component and an invasive device position sensor processing component. 8. The multipurpose host system of claim 1 wherein the multi-mode graphical user interface host includes a first display interface made wherein blood pressure measurements are displayed and a second display interface mode wherein blood flew measurements are displayed. 9. The multipurpose host system of claim 8 wherein the multi-mode graphical user interface host further includes a third display interface mode graphically depicting a combination of display elements including one or more pressure measurement display elements from the first display interface mode and one or more flow measurement display elements from the second display interface. 10. The multipurpose host system of claim 1 wherein the plurality of measu... | Volcano Corporation | A61B005/00 A61B005/02
US07020515 | B2 | 2006-03-28 | Presentation of data stored in an active implantable medical device to assist a practitioner's diagnosis | 2002-03-27 | A process for the presentation of data stored in an active implantable medical device to assist a practitioner in making a diagnosis. This process includes the steps of: a) reading the raw implant data in the stored memory of the device, b) processing the raw data to elaborate a synthesis of the data, preferably through the use of an expert system, c) indexing the elaborated data synthesis according to a plurality of headings, each heading having associated therewith a set of specific implant data having contributed to the development of the heading, d) displaying on a screen (40), at the destination of a user, the aforementioned headings, in the form of a plurality of distinct display fields (41, 42, 43, 44) grouped on the same summary page (40), e) putting at the disposal of the user a tool for the individual selection of each one of the plurality of fields, and f) when a field of the summary page is selected at the choice of the user, displaying on the summary page (40) the aforementioned specific implant data (60) having contributed to the development of the heading corresponding to this selected field (41). | 1. A process for processing raw implant data stored in the memory of an active implantable medical device, in particular of a pacemaker, a defibrillator, a cardiovertor and/or a multisite device, comprising: a) reading the raw implant data stored in the memory of the device;b) processing the raw data to elaborate a data synthesis by means of an expert system;c) generating and indexing from the elaborated data synthesis a plurality of headings, each heading having associated therewith a set of specific implant data having contributed to the development of said heading;d) displaying on a screen a plurality of headings in a corresponding plurality of distinct display fields grouped on a summary page;e) selecting one of said plurality of fields; andf) displaying, in response to a selected field, said specific implant data having contributed to the development of the heading corresponding to said selected field. | 1. A process for processing raw implant data stored in the memory of an active implantable medical device, in particular of a pacemaker, a defibrillator, a cardiovertor and/or a multisite device, comprising: a) reading the raw implant data stored in the memory of the device;b) processing the raw data to elaborate a data synthesis by means of an expert system;c) generating and indexing from the elaborated data synthesis a plurality of headings, each heading having associated therewith a set of specific implant data having contributed to the development of said heading;d) displaying on a screen a plurality of headings in a corresponding plurality of distinct display fields grouped on a summary page;e) selecting one of said plurality of fields; andf) displaying, in response to a selected field, said specific implant data having contributed to the development of the heading corresponding to said selected field. 2. The process of claim 1, wherein displaying the specific implant data further comprises displaying at least one of a chain of markers, a cardiac frequency histogram, a cardiac depolarization/stimulation amplitude histogram, a curve of stored cardiac events, a time-stamped list of cardiac events, and a time-stamped list of sensor data events. 3. The process of claim 1, wherein displaying the specific implant data further comprises displaying said implant data in superposition on the fields of the non-selected headings. 4. The process of claim 1, wherein step e) further comprises providing each said field with a zone inside the field to be activated to effect the selection. | ELA Medical S.A. | A61B005/044
US06923764 | B2 | 2005-08-02 | Analyte monitor | 2003-01-22 | Provided is an analyte monitoring device having a housing, the device comprising: a plurality of needles, each having a tip, a retracted position, a position wherein the tip is extended from the housing a distance adapted to pierce skin; an electrically or spring powered needle pushing apparatus movable to separately engage each of the needles to move each from the retracted position to the extended position; an energy source located within the housing; a plurality of analysis sites comprising an analysis preparation, each adapted to receive liquid from the needles to wet the analysis preparation; one or more light sources adapted to direct light at the analysis sites; one or more light detectors adapted to receive light from the analysis sites; and a processor. | 1. A wearable device constructed for automated detection of the concentration of an analyte present in a sampled body fluid, the device comprising: a housing; at least one needle adapted to pierce the skin; at least one analysis site associated with the at least one needle; a processor; detection elements wherein the processor accepts data from the detection elements and derives normalized values therefrom; and a controller constructed to monitor and control the device in a manner such that repeated detection of the concentration of the analyte present in the sampled body fluid can occur without the necessity of intervention by a wearer of the device. | 1. A wearable device constructed for automated detection of the concentration of an analyte present in a sampled body fluid, the device comprising: a housing; at least one needle adapted to pierce the skin; at least one analysis site associated with the at least one needle; a processor; detection elements wherein the processor accepts data from the detection elements and derives normalized values therefrom; and a controller constructed to monitor and control the device in a manner such that repeated detection of the concentration of the analyte present in the sampled body fluid can occur without the necessity of intervention by a wearer of the device. 2. The device of claim 1, wherein the wearer is a human, and device is wearable on one or more of the flowing areas of the body: wrists, forearms, upper arms, torso, thighs, calves, and ankles. 3. The device of claim 1, further comprising a strap for securing the device to the body. 4. The device of claim 1, further comprising an adhesive for securing the device to the body. 5. The device of claim 1, wherein the analyte comprises glucose. 6. The device of claim 5, wherein the body fluid comprises at least one of: interstitial fluid and blood. 7. The device of claim 1, wherein the housing is constructed such that it is not reversibly openable by the wearer, whereby the entire device is disposable. 8. The device of claim 1, further comprising: an energy source; and detection elements. 9. The device of claim 8, wherein the energy source is removable from the housing. 10. The device of claim 1, further comprising a cassette which is removable from the housing. 11. The device of claim 10, wherein the at least one needle and the at least one analysis site are mounted in the cassette. 12. The device of claim 11, wherein the power source is also mounted in the cassette. 13. The device of claim 1, wherein the at least one needle comprises a microneedle. 14. The device of claim 13, wherein the microneedle has a diameter of 200 μm or less. 15. The device of claim 13, wherein the microneedle has a diameter of 150 μm or less. 16. The device of claim 13, wherein the microneedle has a diameter of 100 μm or less. 17. The device of claim 13, wherein the microneedle has a diameter of 50 μm or less. 18. The device of claim 13, wherein the microneedle has a diameter of 25 μm or less. 19. The device of claim 13, wherein the microneedle is constructed from plastic, glass, silicon wafer, sapphire, quartz, or metal. 20. The device of claim 1, further comprising at least one actuator associated with the at least one needle. 21. The device of claim 20, wherein the actuator comprises at least one of a spring or electrically-driven device. 22. The device of claim 1, wherein the at least one analysis site comprises a support carrying a chemical reagent. 23. The device of claim 22, wherein the support comprises a porous membrane. 24. The device of claim 23, wherein the porous membrane has a pore density ... | Rosedale Medical, Inc. | A61B005/00
US07003344 | B2 | 2006-02-21 | Installation for scanning a leg or other limb of an animal with a magnetic resonance imaging apparatus | 2002-12-18 | An installation for enabling the scanning of a leg (5) of a horse (8) with a magnetic resonance imaging (MRI) device comprises a floor section (90) on which the horse (8) stands with the leg to be scanned. The floor section (90) is mounted above a support surface (10) for the MRI device and includes at least one floor member which is moveable to expose an opening for accommodating a magnetic scanning element (1) adjacent the leg. The floor section (90) may have three moveable floor members and either pair of adjacent floor members may be moveable to expose a pair of slots or openings for accommodating the magnetic scanning element (1) on opposite sides of the leg supported on the floor section (90) between the pair of openings. | 1. An installation for enabling the scanning of a leg or other limb (4, 5) of an animal (8) with a magnetic resonance imaging (MRI) device, comprising a floor section (90) on which the animal (8) stands with the limb (4, 5) to be scanned, characterized in that said floor section (90) is mountable above a support surface (10) for the MRI device and includes at least one floor member (41) moveable in order to expose an opening (50) for accommodating a magnetic scanning element (1) of the MRI device adjacent the limb to be scanned. | 1. An installation for enabling the scanning of a leg or other limb (4, 5) of an animal (8) with a magnetic resonance imaging (MRI) device, comprising a floor section (90) on which the animal (8) stands with the limb (4, 5) to be scanned, characterized in that said floor section (90) is mountable above a support surface (10) for the MRI device and includes at least one floor member (41) moveable in order to expose an opening (50) for accommodating a magnetic scanning element (1) of the MRI device adjacent the limb to be scanned. 2. An installation as claimed in claim 1, wherein the floor section (90) has a free edge (91) for positioning adjacent the MRI device and the floor member (41) is moveable towards and away from the free edge (91) of the floor section. 3. An installation as claimed in claim 1, wherein the floor member (41) is slideable. 4. An installation as claimed in claim 1, wherein the floor member (41) is moveable in response to movement of the MRI device towards and away from the floor section (90). 5. An installation as claimed in claim 1, wherein the floor section (90) includes two floor members (41) moveable to expose two spaced openings or slots (50) for accommodating the magnetic scanning element (1) on opposite sides of the a limb (4, 5) which is supported on the floor section between the two openings or slots (50). 6. An installation as claimed in claim 1, wherein the floor section (90) includes three floor members (41) moveable to expose three spaced openings or slots (50), and wherein either pair of adjacent floor members is moveable to expose a pair of openings or slots (50) for accommodating the magnetic scanning element (1) on opposite sides of a limb supported on the floor section between the respective pair of openings or slots (50). 7. An installation as claimed in claim 1, including a biasing means for restoring the moveable floor member (41) to the position in which the slot or opening (50) is closed in response to movement of the MRI device away from the floor section. 8. An installation as claimed in claim 1, including a trolley or carriage (100) for manoeuvring the magnetic scanning element (1) of the MRI device into and out of the slot(s) or opening(s) (50) and towards and away from the floor section (90). 9. An installation as claimed in claim 8, wherein the trolley or carriage (100) includes means for moving the magnetic scanning element (1) of the MRI device in the horizontal and vertical plane. 10. An installation as claimed in claim 9, wherein the magnetic scanning element (1) is moveable along three substantially orthogonal axes. | Hallmarq Veterinary Imaging Ltd. | A61B005/05
US07052480 | B2 | 2006-05-30 | Access disconnection systems and methods | 2002-04-10 | Apparatuses, devices, systems and methods for detecting access disconnection during medical therapy are provided. The present invention includes an electric circuit coupled to, for example, a blood circuit used during dialysis therapy. The electric circuit can inject an electrical signal into the blood circuit such that changes in an electrical value in response to access disconnection, such as needle drop-out or catheter drop-out, can be effectively detected. | 1. A method of detecting access disconnection during an extracorporeal blood treatment, the method comprising the steps of: coupling an extracorporeal blood system including a plurality of components to a patient via an extracorporeal blood circuit including a first blood line and a second blood line allowing blood to flow into, through and out of the patient along the extracorporeal blood circuit during treatment;injecting an electrical signal into the blood circuit through a direct conductive connection between a pair of electrodes in direct contact with blood as it flows through the extracorporeal blood circuit to define a loop; andmeasuring a change in an electrical value at a third contact point attached to the extracorporeal blood circuit in response to access disconnection. | 1. A method of detecting access disconnection during an extracorporeal blood treatment, the method comprising the steps of: coupling an extracorporeal blood system including a plurality of components to a patient via an extracorporeal blood circuit including a first blood line and a second blood line allowing blood to flow into, through and out of the patient along the extracorporeal blood circuit during treatment;injecting an electrical signal into the blood circuit through a direct conductive connection between a pair of electrodes in direct contact with blood as it flows through the extracorporeal blood circuit to define a loop; andmeasuring a change in an electrical value at a third contact point attached to the extracorporeal blood circuit in response to access disconnection. 2. The method of claim 1 wherein an induction coil is coupled to the extracorporeal blood circuit at the third contact point allowing the change in the electrical value to be measured. 3. The method of claim 1 wherein the third contact point includes an electrical contact attached to the blood circuit in fluid communication with the blood. 4. The method of claim 1 wherein the third contact point includes a capacitive coupler. 5. The method of claim 1 wherein the electrical value is selected from the group consisting of voltage, resistance, impedance, current, rate of changes thereof and combinations thereof. 6. The method of claim 1 wherein the change in electrical value is measured in response to dislodgment of a needle in the patient that is coupled to the extracorporeal blood circuit. 7. An apparatus for detecting dislodgment of an access device from a patient as blood flows through an extracorporeal blood circuit including a first blood line and a second blood line allowing connection of the patient to an extracorporeal blood system including a plurality of components, the apparatus comprising: a direct conductive connection between a pair of electrodes in direct contact with blood as it flows through the extracorporeal blood circuit and through which an electrical signal can be injected into the blood circuit to define a loop; anda measuring device coupled to the extracorporeal blood circuit at a third contact point wherein the measuring device can be used to measure a change in an electrical value in response to dislodgment detection. 8. The apparatus of claim 7 wherein an induction coil is coupled to the extracorporeal blood circuit at the third contact point allowing the change in the electrical value to be measured. 9. The apparatus of claim 7 wherein the third contact point includes an electrical contact attached to the blood circuit in fluid communication with the blood. 10. The apparatus of claim 7 wherein the third contact point includes a capacitive coupler attached to the blood circuit. 11. The apparatus of claim 7 wherein the electrical value is selected from the group consisting of voltage, resistance, impedance, current, rate of changes there... | Baxter International Inc. Baxter Healthcare S.A. | A61M037/00 B01D011/00 C02F001/00 A61B005/00
US07027858 | B2 | 2006-04-11 | Methods and apparatus for cardiac R-wave sensing in a subcutaneous ECG waveform | 2002-09-11 | The present invention uses a R-wave sensing algorithm that uniquely combines an automatic threshold adjustment method with a new noise rejection technique. This algorithm has significant advantages in avoiding the sensing of T-waves, P-waves, and noise/artifacts. Detecting the presence of noise bursts uses features that determine if an R-R interval adjacent to or within the noise signal is valid. Circuitry that discriminates noise signals from R-waves can use any one of several features including, but not limited to, the following: detection events occurring so close together that they are outside normal physiologic heart rates; frequency content that is wider than that of QRS complexes; amplitudes that are different than the adjacent or encompassing R-waves; and amplitudes that display greater than normal variability. The present invention employs multiple discrete thresholds optionally with different decay constants, alone or in combination with one or more substantially constant magnitude sensing threshold. | 1. A method of detecting R-waves produced by a depolarization event of a heart in an ECG waveform, comprising the steps of: programming a minimum threshold;applying a first threshold for a first period of time;applying a second threshold for a second period of time beginning at the end of the first period of time;applying a third threshold for a third period of time beginning at the end of the second period of time;sensing ECG signals emanating from near a heart with at least one electrode; andrecording each event as an R-wave, noise, an arrhythmia, or an asystole based on whether features of the sensed ECG waveform exceed or fail to exceed one or more of the minimum, first, second or third thresholdssensing a peak amplitude of an R-wave and,applying the peak amplitude in combination with the programmed sensitivity (PS) setting;calculating a threshold curve; anddisplaying the threshold curve in relation to a set of detected cardiac activity. | 1. A method of detecting R-waves produced by a depolarization event of a heart in an ECG waveform, comprising the steps of: programming a minimum threshold;applying a first threshold for a first period of time;applying a second threshold for a second period of time beginning at the end of the first period of time;applying a third threshold for a third period of time beginning at the end of the second period of time;sensing ECG signals emanating from near a heart with at least one electrode; andrecording each event as an R-wave, noise, an arrhythmia, or an asystole based on whether features of the sensed ECG waveform exceed or fail to exceed one or more of the minimum, first, second or third thresholdssensing a peak amplitude of an R-wave and,applying the peak amplitude in combination with the programmed sensitivity (PS) setting;calculating a threshold curve; anddisplaying the threshold curve in relation to a set of detected cardiac activity. 2. A method according to claim 1, wherein the first threshold operates according to a relatively slow decay time constant. 3. A method according to claim 2, wherein the relatively slow decay time constant has a value between about 0.5 seconds and about 10.0 seconds. 4. A method according to claim 3, wherein the second threshold has a decay time constant is shorter than the slow decay time constant of the first threshold. 5. A method according to claim 4, wherein the shorter decay time constant is between about 0.5 seconds and 5.0 seconds. 6. A method according to claim 5, wherein the third threshold has a decay time constant is shorter than the slow decay time constant of the first threshold. 7. A method according to claim 1, wherein the method is implemented in an analog electronic circuit, a digital electronic circuit, a hybrid electronic circuit, in a firmware memory location, or in an executable software routine. 8. A method according to claim 1, wherein the set of cardiac activity of the heart is essentially “real-time” cardiac activity. 9. A method according to claim 8, wherein the set of cardiac activity is filtered or rectified prior to displaying same. 10. A method according to claim 1, wherein the method is substantially completely performed by a device and said device is a one of the following: an AED, a ILR, a pacemaker, an ICD, an IPG. 11. A method according to claim 10, further comprising the steps of automatically storing a portion of the waveform after detection of an arrhythmia or an asystole. 12. A method according to claim 11, wherein the storing step occurs based on a priority sequence for each type of arrhythmia or asystole. 13. A method according to claim 12, wherein in the event that a particular higher priority bin is full, a record scheduled for storage in the particular higher priority bin instead displaces a record in a lower priority bin. 14. A method according to claim 1, wherein the at least one electrode comprises at least one of: a subcutaneous electrode, a transcu... | Medtronic, Inc. | A61B005/04
US07022072 | B2 | 2006-04-04 | System for monitoring physiological characteristics | 2001-12-27 | Apparatuses and methods for medical monitoring physiological characteristics values such as blood glucose levels for the treatment of diabetes, are presented. The apparatuses and methods provide dynamic glucose monitoring functions that perform predictive analysis to anticipate harmful conditions, such as glucose crash and hyperglycemic incidents for a patient. The dynamic functions can also be used to maximize athletic performance and warn of inadequate nocturnal basal rate. Other aspects include advanced alarm and reminder functions, as well as advanced data presentation tools to further facilitate convenient and efficient management of various physiological conditions. | 1. A physiological characteristic monitor, comprising: an input device capable of receiving a signal from a sensor, the signal being based on a sensed physiological characteristic value of a user; anda processor for analyzing the received signal; wherein the processor determines a dynamic behavior of the physiological characteristic value;wherein the processor provides an observable indicator based upon the dynamic behavior of the physiological characteristic value so determined;wherein analyzing the received signal and determining a dynamic behavior includes repeatedly measuring the physiological characteristic value to obtain a series of physiological characteristic values and analyzing a rate of change of the physiological characteristic value over time from the series of physiological characteristic values;wherein analyzing the received signal comprises determining whether a most recent of the series of physiological characteristic values is within a qualifying range; andwherein a slope of a line fit to the series of physiological characteristic values is calculated if the most recent of the series of physiological characteristic values is within the qualifying range. | 1. A physiological characteristic monitor, comprising: an input device capable of receiving a signal from a sensor, the signal being based on a sensed physiological characteristic value of a user; anda processor for analyzing the received signal; wherein the processor determines a dynamic behavior of the physiological characteristic value;wherein the processor provides an observable indicator based upon the dynamic behavior of the physiological characteristic value so determined;wherein analyzing the received signal and determining a dynamic behavior includes repeatedly measuring the physiological characteristic value to obtain a series of physiological characteristic values and analyzing a rate of change of the physiological characteristic value over time from the series of physiological characteristic values;wherein analyzing the received signal comprises determining whether a most recent of the series of physiological characteristic values is within a qualifying range; andwherein a slope of a line fit to the series of physiological characteristic values is calculated if the most recent of the series of physiological characteristic values is within the qualifying range. 2. The physiological characteristic monitor of claim 1, wherein the physiological characteristic value is a measure of the concentration of blood glucose in the user. 3. The physiological characteristic monitor of claim 1, wherein each of the series of physiological characteristic values includes a smoothing filtered group of repeated physiological characteristic value readings. 4. The physiological characteristic monitor of claim 1, wherein the slope is negative. 5. The physiological characteristic monitor of claim 1, wherein the indicator includes a warning alarm; and wherein the warning alarm is responsive to the dynamic behavior profile of the physiological characteristic value. 6. The physiological characteristic monitor of claim 5, wherein the physiological characteristic value is a measure of the concentration of blood glucose in the user. 7. The physiological characteristic monitor of claim 6, wherein the warning alarm announces an anticipated glucose crash. 8. The physiological characteristic monitor of claim 1, wherein the series of physiological characteristic values spans a time period of approximately ten minutes. 9. The physiological characteristic monitor of claim 1, wherein the indicator includes an alarm warning of anticipated low glucose levels. 10. The physiological characteristic monitor of claim 1, wherein the series of physiological characteristic values spans a time period of approximately thirty minutes. 11. The physiological characteristic monitor of claim 1, wherein the indicator is provided if the slope is steeper than a threshold rate. 12. The physiological characteristic monitor of claim 11, wherein the threshold rate is approximately 3% per minute and the qualifying range is approximately 100 to 150 mg/dL. 13. The physiological character... | Medtronic MiniMed, Inc. | A61B005/05 A61B005/00
US07050845 | B2 | 2006-05-23 | Projecting patient image data from radioscopic imaging methods and/or tomographic imaging methods onto video images | 2002-04-29 | The invention relates to a device for projecting patient image data from radioscopic imaging methods and/or tomographic imaging methods onto video images, comprising an image display device (10), at least one video camera (14) and a computer-assisted navigation system which can detect the spatial positions of the image display device (10) and/or the video camera (14) and the spatial positions of a part of a patient's body (1) via tracking means (12, 3) attached to it, wherein the image display device (10) and the video camera (14) are assigned to each other and are realized as a portable unit. | 1. A device for projecting patient image data from radioscopic imaging methods and/or tomographic imaging methods onto video images, comprising an image display device, at least one video camera, and a computer-assisted navigation system which can detect the spatial positions of said image display device and/or said video camera and the spatial positions of a part of a patient's body via tracking means attached to it, wherein said image display device and said video camera are assigned to each other and are realized as a portable unit. | 1. A device for projecting patient image data from radioscopic imaging methods and/or tomographic imaging methods onto video images, comprising an image display device, at least one video camera, and a computer-assisted navigation system which can detect the spatial positions of said image display device and/or said video camera and the spatial positions of a part of a patient's body via tracking means attached to it, wherein said image display device and said video camera are assigned to each other and are realized as a portable unit. 2. The device as set forth in claim 1, wherein said video camera is arranged on the rear side of said image display device. 3. The device as set forth in claim 1, wherein said video camera is spatially assigned to said image display device in a way which is predetermined and/or known to the system. 4. The device as set forth in claim 1, wherein said video camera is spatially assigned to said image display device in a way which is predetermined and/or known to the system. 5. The device as set forth in claim 1, wherein said image display device is a portable LCD flat screen. 6. The device as set forth in claim 1, wherein the image data from said radioscopic imaging method and/or tomographic imaging method to be projected onto said video images are communicated to said image display device via radio interfaces of said navigation system. 7. The device as set forth in claim 1, wherein the image data from said radioscopic imaging method and/or tomographic imaging method to be projected onto said video images are communicated to said image display device via a data cable of said navigation system. 8. The device as set forth in claim 1, wherein said video camera exhibits a small aperture and a low depth of field. 9. The device as set forth in claim 1, wherein a data transmission means is arranged on said video camera or on said image display device, by means of which information on the image shown on said image display device is transmitted to said navigation system, wherein the image data to be projected, from said radioscopic imaging methods and/or tomographic imaging methods, communicated from said navigation system to said image display device, are adapted in size and position on said image display device such that they are in registration with said video images. 10. The device as set forth in claim 9, wherein said navigation system comprises a contour-matching unit which superimposes said video images and said projected images, in particular via outer contour matching of said part of the patient's body. 11. The device as set forth in claim 1, wherein an illuminating device for said part of the patient's body is provided on said camera or in its vicinity on said image display device. 12. The device as set forth in claim 1, wherein control apparatus and command input apparatus for altering the combined image shown are attached to the portable unit. 13. The device as set forth in claim 12, wherein the control... | BrainLAB AG | A61B005/05
US06923772 | B2 | 2005-08-02 | Apparatus and method for determining responders to cardiac resynchronization therapy using implantable accelerometers | 2002-09-06 | A device for measuring a synchronicity of contraction of a heart to determine if an individual would be a responder to cardiac resynchronization therapy. The device may include a first electrode positioned at a first ventricular wall location to measure movement of the first ventricular wall location and generate a first signal, as well as a second electrode positioned at a second ventricular wall location to measure movement of the second ventricular wall location and generate a second signal. A processing module may process the first and second signals, for example, integrate the signals multiple times, and generate an output based on the processed signals. Based on this output, an individual can be labeled a responder or non-responder to cardiac resynchronization therapy. The electrodes may each include an accelerometer to measure acceleration of the heart wall. | 1. A device for measuring a synchronicity of contraction of a heart to determine if an individual would be a responder to cardiac resynchronization therapy, the device comprising: a first electrode adapted to be positioned at a first ventricular wall location and configured to mechanically measure movement of the first ventricular wall location; a second electrode adapted to be positioned at a second ventricular wall location and configured to mechanically measure movement of the second ventricular wall location; a processing module coupled to the first and second electrodes, the processing module adapted to process a first signal produced by the first electrode at the first ventricular wall location and a second signal produced by the second electrode at the second ventricular wall location and generate an output based on the processed signals; and a comparator module coupled to the processing module to indicate that the individual would be a responder if the output of the processing module exceeds a threshold and to indicate that the individual would be a non-responder if the output of the processing module is below the threshold. | 1. A device for measuring a synchronicity of contraction of a heart to determine if an individual would be a responder to cardiac resynchronization therapy, the device comprising: a first electrode adapted to be positioned at a first ventricular wall location and configured to mechanically measure movement of the first ventricular wall location; a second electrode adapted to be positioned at a second ventricular wall location and configured to mechanically measure movement of the second ventricular wall location; a processing module coupled to the first and second electrodes, the processing module adapted to process a first signal produced by the first electrode at the first ventricular wall location and a second signal produced by the second electrode at the second ventricular wall location and generate an output based on the processed signals; and a comparator module coupled to the processing module to indicate that the individual would be a responder if the output of the processing module exceeds a threshold and to indicate that the individual would be a non-responder if the output of the processing module is below the threshold. 2. The device of claim 1, wherein the first electrode includes a first accelerometer and the second electrode includes a second accelerometer. 3. The device of claim 1, wherein the processing module further comprises: an integrator module that integrates the first signal and the second signal and plots the integrated first and second signals to form a displacement loop; and a synchronicity index module that measures a phase difference of contraction based on the displacement loop. 4. A device for determining whether an individual would be a responder to cardiac resynchronization therapy, the device comprising: a first accelerometer adapted to be positioned at a first ventricular wall location; a second accelerometer adapted to be positioned at a second ventricular wall location; a processing module coupled to the first and second accelerometers, the processing module adapted to process a first signal produced by acceleration of the first accelerometer at the first ventricular wall location and a second signal produced by acceleration of the second accelerometer at the second ventricular wall location and generate an output; and a comparator module coupled to the processing module to indicate that the individual would be a responder if the output of the processing module exceeds a threshold and to indicate that the individual would be a non-responder if the output of the processing module is below the threshold. 5. The device of claim 4, wherein the processing module and the comparator module of the device are in vitro. 6. The device of claim 4, wherein the entire device is implanted into the individual. 7. The device of claim 4, wherein the processing module further comprises: an integrator module that integrates the first signal and the second signal and plots the integrated first and second signals to form ... | Cardiac Pacemakers, Inc. | A61B005/02
US06901282 | B2 | 2005-05-31 | Scoutless whole-body imaging with fast positioning | 2003-02-04 | A method and apparatus are provided for forming a magnetic resonance angiographic image of a human body. A plurality of spatially non-selective radio-frequency pulses and a plurality of different combinations of phase-encoding gradients are applied to the human body, that are temporally non-coincident with the radio-frequency pulses and where each combination includes a pulse value in a slice selective direction and a pulse value in an in-plane direction and magnetic resonance imaging data are detected. A slice processor and/or a thickness processor identify the presence, location and/or thickness of a body portion of the human body. Identification of a body portion in a first imaging volume becomes the basis of application of a second plurality of spatially non-selective radio-frequency pulses to a second imaging volume of the human body. | 1. An apparatus for forming a magnetic resonance angiographic image of a human body comprising: means for applying a first plurality of spatially non-selective radio-frequency pulses to a first imaging volume of the human body; means for applying a first plurality of combinations of magnitude of phase-encoding gradients in slice-selective and in-plane directions through the first imaging volume of the human body in conjunction with application of the first plurality of spatially non-selective radio-frequency pulses; means for detecting first magnetic resonance imaging data from the first imaging volume of the human body based upon the first plurality of spatially non-seleqtive radio-frequency pulses and the applied first plurality of combinations of magnitude of phase-encoding gradients; means for identifying a location of a body portion of the human body within the first imaging volume based upon the detected first magnetic resonance imaging data; means for applying a second plurality of spatially non-selective radio-frequency pulses to a second imaging volume of the human body based upon the identified location of the body portion; means for applying a second plurality of combinations of magnitude of phase-encoding gradients in slice-selective and in-plane directions through the second imaging volume of the human body in conjunction with application of the second plurality of spatially non-selective radio-frequency pulses; means for detecting magnetic resonance imaging data from the second imaging volume of the human body based upon the second plurality of spatially non-selective radio-frequency pulses and second plurality of combinations of magnitude of phase-encoding gradients; and means for forming the magnetic resonance angiographic image of the second imaging volume. | 1. An apparatus for forming a magnetic resonance angiographic image of a human body comprising: means for applying a first plurality of spatially non-selective radio-frequency pulses to a first imaging volume of the human body; means for applying a first plurality of combinations of magnitude of phase-encoding gradients in slice-selective and in-plane directions through the first imaging volume of the human body in conjunction with application of the first plurality of spatially non-selective radio-frequency pulses; means for detecting first magnetic resonance imaging data from the first imaging volume of the human body based upon the first plurality of spatially non-seleqtive radio-frequency pulses and the applied first plurality of combinations of magnitude of phase-encoding gradients; means for identifying a location of a body portion of the human body within the first imaging volume based upon the detected first magnetic resonance imaging data; means for applying a second plurality of spatially non-selective radio-frequency pulses to a second imaging volume of the human body based upon the identified location of the body portion; means for applying a second plurality of combinations of magnitude of phase-encoding gradients in slice-selective and in-plane directions through the second imaging volume of the human body in conjunction with application of the second plurality of spatially non-selective radio-frequency pulses; means for detecting magnetic resonance imaging data from the second imaging volume of the human body based upon the second plurality of spatially non-selective radio-frequency pulses and second plurality of combinations of magnitude of phase-encoding gradients; and means for forming the magnetic resonance angiographic image of the second imaging volume. 2. The apparatus for forming a magnetic resonance angiographic image as in claim 1 further comprising means for injecting a contrast agent into the human body. 3. The apparatus for forming a magnetic resonance angiographic image as in claim 1 further comprising means for encoding in-plane spatial information using one of the group consisting of vastly undersampled projection reconstruction, Hadamard encoding and wavelet encoding. 4. The apparatus for forming a magnetic resonance angiographic image as in claim 1 further comprising means for deleting any imaging slices within the first imaging volume without the identified body portion of the human body. 5. A method of forming a magnetic resonance angiographic image of a human body comprising the steps of: applying a first plurality of spatially non-selective radio-frequency pulses to a first imaging volume of the human body; applying a first plurality of combinations of magnitude of phase-encoding gradients in slice-selective and in-plane directions through the first imaging volume of the human body in conjunction with application of the first plurality of spatially pon-selective radio-frequency pulses; detecting first mag... | Evanston Northwestern Healthcare, Inc. | A61B005/055
US07020513 | B2 | 2006-03-28 | Testing the electrical continuity of the connecting cables for a recorder of physiological signals, in particular a holter recorder of ECG signals | 2003-01-23 | Detection of a discontinuity, a disconnection or loss of conductivity, of cables connected to a physiological signal recorder, in particular a Holter recorder of ECG signals by measuring an impedance of the line(s). To measure the impedance of the line comprising the connecting cable connecting an external electrode placed on a patient to a signal terminal (12) of the recorder (10), one generates a current impulse, applies this impulse to the line, measures the variation of voltage resulting on the terminal from the signal during the application duration of the current impulse, and determines the impedance of the line based upon the voltage variations thus measured. The current impulse is a biphasic impulse including two successive cycles of opposite polarities, the durations and the amplitudes of these two cycles being selected so as to define approximately equal respective loads and of contrary signs. A switch (36) connects the base of the pull-down resistance (32) of the signal terminal (12) to the output of a voltage or a current generator (38) during the impedance measurement, and connects it to the potential ground (34) during the remainder of time. | 1. A process for testing the electrical continuity of a line by measuring an impedance of a line, the line comprising a connecting cable connecting an external electrode placed on a patient to a signal terminal of a recorder of physiological signals comprising the steps of: generating a biphasic current impulse having a first cycle of a first polarity and a second cycle of a second polarity, each cycle having a duration and an amplitude, selecting said durations and the amplitudes of said first and second cycles to define approximately equal respective loads of contrary signs on said line;applying said impulse to said line.measuring a voltage variation on the signal terminal in response to said biphasic current impulse; anddetermining an impedance of the line based upon said measured voltage variations. | 1. A process for testing the electrical continuity of a line by measuring an impedance of a line, the line comprising a connecting cable connecting an external electrode placed on a patient to a signal terminal of a recorder of physiological signals comprising the steps of: generating a biphasic current impulse having a first cycle of a first polarity and a second cycle of a second polarity, each cycle having a duration and an amplitude, selecting said durations and the amplitudes of said first and second cycles to define approximately equal respective loads of contrary signs on said line;applying said impulse to said line.measuring a voltage variation on the signal terminal in response to said biphasic current impulse; anddetermining an impedance of the line based upon said measured voltage variations. 2. The process of claim 1, wherein measuring the voltage variation comprises measuring a first voltage variation at the end of said first cycle and measuring a second voltage variation at the end of said second cycle. 3. The process of claim 2, wherein measuring the voltage variation further comprises obtaining an absolute value of the voltage variations measured at the end of the first cycle and at the end of second cycle and adding them together. 4. The process of claim 2, further comprising using said line to collect an ECG signal, sampling the collected ECG signal at a first frequency and obtaining a plurality of samples in a first measurement cycle, measuring the ECG signal using at least one of said samples, wherein measuring said voltage variation further comprises using at least two of said samples for the measurement of the voltage variations at the end of the first cycle and the end of the second cycle. 5. The process of claim 4, wherein measuring the ECG signal further comprises averaging selected ones of said plurality of samples. 6. The process of claim 4, further comprising selecting said first frequency of sampling the ECG signal to produce at least two samples for use in the measurement of the voltage variation at the end of each of the first cycle and the second cycle, and at least one sample for use in the measurement of the ECG signal. 7. The process of claim 6, wherein selecting said first frequency further produces an intervening sample following said at least two samples used in the voltage variation measurement, wherein said intervening sample is not used in measurement of said voltage variations. 8. The process of claim 1, wherein the patient is equipped with an implanted pacemaker able to deliver stimulation pulses, further comprising eliminating from the measured voltage variations signals corresponding to said stimulation pulses. 9. The process of claim 8, wherein eliminating signals corresponding to said stimulation pulses comprises detecting a stimulation pulse in the measured voltage variation, defining a temporal post-stimulation window, and measuring the impedance of said line during said window. 10. The pr... | Ela Medical S.A. | A61B005/04
US07111980 | B2 | 2006-09-26 | System and method using thermal image analysis and slope threshold classification for polygraph testing | 2001-11-13 | Thermal image data of at least a region of a face of a person is provided. The thermal image data is transformed to blood flow rate data which includes a slope representative of the change of blood flow rate over time. The slope is compared to a slope threshold to determine a physiological state of the person. For example, it may be determined whether the person is deceptive or non-deceptive, e.g., deceptive with respect to an elicited response from the person. | 1. A method for use in detecting deception of a person, the method comprising: providing thermal image data of at least a region of a face of a person; transforming the thermal image data to blood flow rate data, wherein the blood flow rate data comprises a slope representative of change of blood flow rate over time; providing a slope threshold representative of a demarcation between change of blood flow rate over time for a deceptive person and change of blood flow rate over time for a non-deceptive person; and comparing the slope representative of change of blood flow rate over time to the slope threshold to determine whether the person is deceptive or non-deceptive. | 1. A method for use in detecting deception of a person, the method comprising: providing thermal image data of at least a region of a face of a person; transforming the thermal image data to blood flow rate data, wherein the blood flow rate data comprises a slope representative of change of blood flow rate over time; providing a slope threshold representative of a demarcation between change of blood flow rate over time for a deceptive person and change of blood flow rate over time for a non-deceptive person; and comparing the slope representative of change of blood flow rate over time to the slope threshold to determine whether the person is deceptive or non-deceptive. 2. The method of claim 1, wherein providing a slope threshold comprises generating a slope threshold based on at least answer slopes representative of the change of blood flow rate over time for a plurality of subjects during at least a portion of an answer period of time in which the plurality of subjects answer questions posed thereto. 3. The method of claim 1, wherein providing a slope threshold comprises generating a slope threshold based on answer slopes representative of the change of blood flow rate over time for a plurality of subjects during at least a portion of an answer period of time in which the plurality of subjects answer questions posed thereto during a question period of time and also based on question slopes representative of the change of blood flow rate over time for a plurality of subjects during at least a portion of the question period of time. 4. The method of claim 3, wherein providing a slope threshold comprises generating a slope threshold based on a product of corresponding question and answer slopes. 5. The method of claim 1, wherein providing the thermal image data comprises providing thermal image data of at least a region proximate an eye of the person. 6. The method of claim 1, wherein providing thermal image data comprises: asking the person a question to elicit a response therefrom; focusing a thermal infrared image device operable to provide thermal image data on at least the region of the face of the person; and capturing thermal image data during at least a period of time during the response from the person. 7. The method of claim 6, wherein determining whether the person is deceptive or non-deceptive comprises determining whether the person is being deceptive or non-deceptive with respect to the response to the question. 8. The method of claim 1, wherein providing thermal image data comprises: focusing a thermal infrared image device operable to provide thermal image data on at least the region of the face of the person; capturing frames of thermal image data during at least a period of time; and tracking movement of at least the region of the face of the person. 9. A system for use in detecting deception of a person, the system comprising: a thermal infrared image device operable to provide thermal image data of at least a region ... | Honeywell International Inc. | G01K013/00 A61B005/02
US07225006 | B2 | 2007-05-29 | Attachment and optical probe | 2003-01-23 | A medical monitoring system, such as an oximetry system, applies an attachment for securing an optical probe to a measurement site. The attachment has an elongated support with a first end and a second end, and a dedicated area in proximity of the first end. The dedicated area receives an optical probe and includes a material that is transparent for light emitted and received by the optical probe. The dedicated area mountably receives the optical probe on the material so that in use, the material is positioned between the optical probe and a surface of a measurement site. The optical probe may be factory-mounted to the dedicated area of the attachment as a ready-to-use sensor. In the alternative, the attachment may be available as an individual component, or as a part of a kit including the attachment and the optical probe. | 1. A probe for non-adhesively attaching to a measurement site, comprising: an elongated support having a first end and a second end, the support comprising:at least one attachment element in proximity of one of said ends, wherein the at least one attachment element is configured to attach with a second portion of said support to non-adhesively secure an optical probe to a measurement site; anda dedicated area in proximity of the first end and configured to receive the optical probe, the dedicated area comprising a material different from said support that is transparent for light emitted and received by the optical probe; anda disposable optical probe comprising a standalone adhesive-type optical probe which is configured to adhesively attach to a measurement site independent of said support;wherein the disposable optical probe is mounted on the dedicated area so that during measurement, the material is positioned between the optical probe and a surface of a measurement site. | 1. A probe for non-adhesively attaching to a measurement site, comprising: an elongated support having a first end and a second end, the support comprising:at least one attachment element in proximity of one of said ends, wherein the at least one attachment element is configured to attach with a second portion of said support to non-adhesively secure an optical probe to a measurement site; anda dedicated area in proximity of the first end and configured to receive the optical probe, the dedicated area comprising a material different from said support that is transparent for light emitted and received by the optical probe; anda disposable optical probe comprising a standalone adhesive-type optical probe which is configured to adhesively attach to a measurement site independent of said support;wherein the disposable optical probe is mounted on the dedicated area so that during measurement, the material is positioned between the optical probe and a surface of a measurement site. 2. The probe of claim 1, wherein the support is configured to allow wrapping the support around the measurement site. 3. The probe of claim 2, said attachment element comprising a tongue extending from the second end of the support. 4. The probe of claim 1, wherein the support includes a soft foam padding. 5. The probe of claim 4, said attachment element comprising a tongue extending from the second end of the support, the tongue comprising hooks to cooperate with a surface of the soft foam padding in a hook-loop securing mechanism. 6. The probe of claim 1, wherein the support includes a tape. 7. The probe of claim 6, said attachment element in proximity of the second end and comprising an adhesive to cooperate with a surface of the tape. 8. The probe of claim 1, wherein the dedicated area is sized to allow passage of light emitted and received by an optical probe. 9. The probe of claim 1, wherein the support has at least one opening, and wherein the material is coupled to the at least one opening. 10. The probe of claim 9, wherein the material is mounted within the at least one opening by at least one of sewing, gluing and welding. 11. The probe of claim 10, wherein the material is a two-part element, and wherein the parts are spaced apart at a predetermined distance. 12. The probe of claim 9, wherein the material includes an opaque portion that prevents light from traveling from one side of the material to another side of the material. 13. The probe of claim 12, wherein the material and the dedicated area have an elongate shape, and wherein the opaque portion comprises a stripe that extends between longitudinal sides of the material. 14. The probe of claim 1, wherein the material is a two-part element, and wherein the support has two openings, each opening sized to receive one of the parts. 15. The probe of claim 1, wherein the support comprises a bottom layer having a back surface and a bottom surface facing of the measurement site, and wherein at least a po... | Masimo Corporation | A61B005/00
US07149570 | B2 | 2006-12-12 | Alarm activated acoustic measuring signals for patient monitoring | 2002-10-17 | A system and a method for monitoring a physiological parameter of a patient providing a measurement-modulated acoustic signal (7) after/when a measurement value of the monitored physiological parameter exceeds an alarm limit. | 1. A patient monitoring system for monitoring at least two physiological parameters of a patient, the system comprising: an electrocardiograph ECG which monitors the patient's cardiac heart rate;a second physiological parameter monitor which monitors values of a second physiological parameter of the patient;an acoustic signal generator which generates acoustic signals;a beep generator connected with the ECG monitor and the second physiological parameter monitor which generates a beep tone at a characteristic point in each cycle of said cyclic heart rate at a tone which varies in accordance with the monitored values of the second physiological parameter, the beep generator being connected with the acoustic signal generator to generate a series of beeps whose rate is indicative of the cyclic heart rate and whose tone is indicative of the monitored values of the second physiological parameter; and,an alarm unit which causes a volume of the acoustic signal to change from switched off or a lower volume to a predefined higher acoustic volume in response to the monitored values of one or both of the second monitored physiological parameters or the monitored heart rate exceeding an alarm limit. | 1. A patient monitoring system for monitoring at least two physiological parameters of a patient, the system comprising: an electrocardiograph ECG which monitors the patient's cardiac heart rate;a second physiological parameter monitor which monitors values of a second physiological parameter of the patient;an acoustic signal generator which generates acoustic signals;a beep generator connected with the ECG monitor and the second physiological parameter monitor which generates a beep tone at a characteristic point in each cycle of said cyclic heart rate at a tone which varies in accordance with the monitored values of the second physiological parameter, the beep generator being connected with the acoustic signal generator to generate a series of beeps whose rate is indicative of the cyclic heart rate and whose tone is indicative of the monitored values of the second physiological parameter; and,an alarm unit which causes a volume of the acoustic signal to change from switched off or a lower volume to a predefined higher acoustic volume in response to the monitored values of one or both of the second monitored physiological parameters or the monitored heart rate exceeding an alarm limit. 2. The system of claim 1, wherein the pre-defined acoustic volume is independent of previous settings for the volume before the alarm limit has been exceeded. 3. The system according to claim 1, wherein the acoustic signal is only generated as long as one or more alarm limits are exceeded, and is switched off otherwise. 4. The system according to claim 1, wherein the acoustic signal is generated at the pre-defined higher acoustic volume as long as and as soon as the second monitored values or the heart rate are at an alarm condition and the acoustic signal is switched off or is generated at the lower acoustic volume as long as and as soon as the second measurement values and the heart rate do not exceed one or more given alarm limits. 5. A method for monitoring at least two physiological parameters of a patient and providing a measurement-modulated acoustic signal after/when a measurement value of the monitored physiological parameters exceeds an alarm limit, the method comprising: measuring a value of a first cyclic monitored physiological parameter;measuring a value of a second monitored physiological parameter;in response to a measured value of any physiological parameter exceeding an alarm limit, providing the measurement-modulated acoustic signal at a predefined high acoustic volume, the acoustic signal including a series of beeps whose rate varies in accordance with the measurement values of the first cyclic monitored physiological parameter and whose tone varies in accordance with the measurement values of the second monitored physiological parameter; and,in response to no alarm limits being exceeded, reducing the acoustic volume of the measurement modulated acoustic signal to a low volume or switching the measurement-modulated acoustic signal off. 6. ... | Koninklijke Philips Electronics N.V. | A61B005/04 A61B005/00
US07162292 | B2 | 2007-01-09 | Beam scanning probe system for surgery | 2002-05-16 | With the beam scanning probe system for surgery, a pointer indicating the observation point of a beam scanning probe is superimposed upon an image of a lesion to be operated on in the head which is obtained via a TV camera or a surgery microscope, and the superimposed image information is registered as such in an image recording device. With the beam scanning probe system for surgery, information of a cytological picture is also registered in the image recording device. With the beam scanning probe system for surgery, the image recording device registers the two image information in a paired fashion. Through this arrangement, the beam scanning probe system for surgery can smoothly locate, for a given cytological picture, a site of a tumor to be treated from which the picture was obtained, which will ease the operation. | 1. A beam scanning probe system for surgery comprising: image acquiring means for acquiring an image of a lesion to be operated on in the head;a laser beam scanning probe for acquiring optical image information, including a cytological picture, of a specific site of the lesion by confocally scanning light reflected by the specific site;a detection means for detecting the position of an observation point of the laser beam scanning probe relative to the site of the lesion;a recording means for recording, in a paired fashion, two image information: one relating to the cytological picture obtained via the beam scanning probe, and the other to an image of the lesion upon which the information regarding the position of the observation point of the beam scanning probe is superimposed;a transmitting means for transmitting image information and observation point information of the beam scanning probe, determination information of the cytological picture, and instruction information introduced based on the determination information of the cytological picture;a terminal for transmitting/receiving information transmitted by the transmitting means;a determination result feeding means for feeding a determination result of the cytological picture;a determination result displaying means for displaying at least two kinds of determination based on determination results provided by the determination result feeding means; anda superimposing means for superimposing the determination results provided by the determination result displaying means on the image of the lesion, based on the observation point position of the beam scanning probe relative to the site of the lesion. | 1. A beam scanning probe system for surgery comprising: image acquiring means for acquiring an image of a lesion to be operated on in the head;a laser beam scanning probe for acquiring optical image information, including a cytological picture, of a specific site of the lesion by confocally scanning light reflected by the specific site;a detection means for detecting the position of an observation point of the laser beam scanning probe relative to the site of the lesion;a recording means for recording, in a paired fashion, two image information: one relating to the cytological picture obtained via the beam scanning probe, and the other to an image of the lesion upon which the information regarding the position of the observation point of the beam scanning probe is superimposed;a transmitting means for transmitting image information and observation point information of the beam scanning probe, determination information of the cytological picture, and instruction information introduced based on the determination information of the cytological picture;a terminal for transmitting/receiving information transmitted by the transmitting means;a determination result feeding means for feeding a determination result of the cytological picture;a determination result displaying means for displaying at least two kinds of determination based on determination results provided by the determination result feeding means; anda superimposing means for superimposing the determination results provided by the determination result displaying means on the image of the lesion, based on the observation point position of the beam scanning probe relative to the site of the lesion. 2. A beam scanning probe system for surgery according to claim 1, further comprising: a computation means for determining a specific area based on determination results provided by the determination result displaying means; anda specific area displaying means for displaying the specific area determined by the computation means, wherein: the superimposing means superimposes at least one chosen from the specific area to be displayed and the determination results to be displayed, upon an image of a lesion to be operated on. 3. A beam scanning probe system for surgery according to claim 2, wherein: the computation means takes two marker points having different determination results, determines a point between the two according to a predetermined ratio, connect the mid-points thus determined to obtain a closed outline, and takes the area surrounded by the outline as the specific area. 4. A beam scanning probe system for surgery according to claim 1, wherein: the recording means stores the determination results to be displayed provided by the determination result displaying means in connection with cytological pictures, comprising: an input selection means for selectively feeding input from among the determination results to be displayed; anda read means for reading an appropriate cytological picture f... | Olympus Corporation | A61B005/05
US07083578 | B2 | 2006-08-01 | Device and method for examining a body lumen | 2002-07-11 | An in vivo examining device and method are described. The in vivo examining device has two operational phases; an initial phase in which the device is of initial dimensions and a final phase in which the device is of final dimensions. In the initial phase the device can pass freely through a normally configured body lumen whereas it may not be able to pass freely through an abnormally configured lumen. In the final phase the device can pass freely through a body lumen even if it is abnormally configured. | 1. A method for testing a body lumen, the method comprising the steps of: introducing into the body lumen a device, said device having initial dimensions for a predetermined period and reduced dimensions after the predetermined period, said device including at least a partially dissolvable plug covered by an impermeable coating defining an opening; and monitoring the device. | 1. A method for testing a body lumen, the method comprising the steps of: introducing into the body lumen a device, said device having initial dimensions for a predetermined period and reduced dimensions after the predetermined period, said device including at least a partially dissolvable plug covered by an impermeable coating defining an opening; and monitoring the device. 2. The method according to claim 1 wherein the body lumen is the GI tract. 3. The method according to claim 2 wherein the step of introducing the device comprises ingesting the device. 4. The method according to claim 1 wherein the predetermined period is a period of approximately one hundred hours or more. 5. The method according to claim 1 wherein the device comprises a filling disposed within said coating, the filling being capable of absorbing fluid from the body lumen. 6. The method according to claim 5 wherein, after the predetermined period, the filling swells enough to burst the coating. 7. The method according to claim 1 wherein, after the predetermined period, the coating is burst. 8. The method according to claim 1 wherein the coating comprises at least one layer. 9. The method according to claim 1 wherein the coating includes a material or combination of materials selected from the group consisting of: parylene C, hard gelatin, soft gelatin, vegetable gelatin and a hydrogel. 10. The method according to claim 9 wherein the hydrogel is ethyl cellulose acetate. 11. The method according to claim 5 wherein the filling comprises at least one layer. 12. The method according to claim 5 wherein the filling includes a material or combination of materials selected from the group consisting of: a filler, a biodegradable polymer, microparticles of a hydrophilic substance and a hydrogel. 13. The method according to claim 12 wherein the biodegradable polymer is a polymer of lactide and glycollide. 14. The method according to claim 12 wherein the filling further comprises adhesives. 15. The method according to claim 5 wherein the coating is thinner than the filling. 16. The method according to claim 1 wherein the initial dimensions comprise a diameter of approximately 1-12 mm and wherein the reduced dimensions comprise a diameter of approximately 2-10 μm. 17. The method according to claim 1 wherein the step of monitoring the device comprises detecting a signal emanated from a monitoring device, said device including at least the monitoring device. 18. The method according to claim 17 wherein the monitoring device is of smaller dimensions than the device initial dimensions. 19. The method according to claim 17 wherein the monitoring device is a passive device. 20. The method according to claim 19 wherein the monitoring device is selected from the group consisting of: an electronic ID tag, a magnetized device or an acoustic device. 21. The method according to claim 19 wherein the step of detecting a signal emanated from a monitoring device comprises the steps ... | Given Imaging Ltd. | A61B005/103 A61B005/117
US07239906 | B1 | 2007-07-03 | Magnetic resonance imaging system including a transpolar fixture | 2002-04-18 | A magnetic resonance imaging system is provided with a fixture which can be mounted to the static field magnet within the patient-receiving space of the magnet. The fixture can support and stabilize the patient. | 1. A magnetic resonance imaging system; comprising: a magnetic resonance imaging apparatus having a first pole face and a second pole face disposed opposite the said first pole face; anda fixture comprising: (a) a structure;(b) first and second pads connected to said structure, said first and second pads being movable between a collapsed condition in which said pads are disposed at a first distance from one another and an expanded condition in which said pads are at a second distance from one another greater than said first distance, said structure being disposed between said first and second pole faces such that said first and second pads engage said first and second pole faces of the imaging apparatus and secure said structure to said first and second pole faces when said pads are in said expanded condition. | 1. A magnetic resonance imaging system; comprising: a magnetic resonance imaging apparatus having a first pole face and a second pole face disposed opposite the said first pole face; anda fixture comprising: (a) a structure;(b) first and second pads connected to said structure, said first and second pads being movable between a collapsed condition in which said pads are disposed at a first distance from one another and an expanded condition in which said pads are at a second distance from one another greater than said first distance, said structure being disposed between said first and second pole faces such that said first and second pads engage said first and second pole faces of the imaging apparatus and secure said structure to said first and second pole faces when said pads are in said expanded condition. 2. A system as claimed in claim 1 wherein said structure includes a structural element which is substantially in the form of an elongated rod extending between said pads at least when said pads are in said expanded condition. 3. A system as claimed in claim 2, wherein at least one of said pads is pivotally connected to said structural element for pivoting movement about at least one pad pivot axis transverse to the direction of elongation of said structural element. 4. A system as claimed in claim 3, wherein said first pad and said second pad include planar surfaces adapted to bear on said pole faces. 5. A system as claimed in claim 2 wherein said structural element includes a first elongated section coupled to said first pad, a second elongated section coupled to said second pad, and a connector connecting said sections so that said sections are movable relative to one another to move the pads between said collapsed and expanded conditions. 6. A system as claimed in claim 5, wherein said connector includes a hinge allowing pivoting movement of said first and said second sections about a hinge pivot axis in a first direction from a folded position in which said elongated sections are not collinear and an unfolded position in which said elongated sections are substantially collinear, and a stop limiting pivoting movement in said first direction beyond said unfolded position. 7. A system as claimed in claim 6, wherein said first pad is adjustably connected to said first section so that said first pad can be adjusted relative to said first section in the direction of elongation of said first section. 8. A system as claimed in claim 6, further comprising a locking mechanism for locking said first and second sections in said unfolded position. 9. A system as claimed in claim 5, further comprising a screwjack, said first pad being connected to said first section through said screwjack, said screwjack providing an adjustable connection. 10. A system as claimed in claim 2, wherein said structural element includes a first elongated section and a second section slidably movable relative said first section between a retracted position and an ... | Fonar Corporation | A61B005/055
US06892090 | B2 | 2005-05-10 | Method and apparatus for virtual endoscopy | 2002-08-19 | A surgical instrument navigation system is provided that visually simulates a virtual volumetric scene of a body cavity of a patient from a point of view of a surgical instrument residing in the cavity of the patient. The surgical instrument navigation system includes: a surgical instrument; an imaging device which is operable to capture scan data representative of an internal region of interest within a given patient; a tracking subsystem that employs electro-magnetic sensing to capture in real-time position data indicative of the position of the surgical instrument; a data processor which is operable to render a volumetric, perspective image of the internal region of interest from a point of view of the surgical instrument; and a display which is operable to display the volumetric perspective image of the patient. | 1. A surgical instrument navigation system, comprising: a surgical instrument; a tracking subsystem operable to capture in real-time position data indicative of the position of the surgical instrument; a data processor adapted to receive scan data representative of a region of interest of a given patient and the position data from the tracking subsystem, the data processor being operable to render an image of the region of interest from a point of view which relates to position of the surgical instrument, the image being derived from the scan data; and a display in data communication with the data processor, the display being operable to display the image of the patient. | 1. A surgical instrument navigation system, comprising: a surgical instrument; a tracking subsystem operable to capture in real-time position data indicative of the position of the surgical instrument; a data processor adapted to receive scan data representative of a region of interest of a given patient and the position data from the tracking subsystem, the data processor being operable to render an image of the region of interest from a point of view which relates to position of the surgical instrument, the image being derived from the scan data; and a display in data communication with the data processor, the display being operable to display the image of the patient. 2. The surgical navigation system of claim 1 wherein a volumetric perspective image is rendered from a point of view of the surgical instrument. 3. The surgical navigation system of claim 1 wherein the surgical instrument is further defined as at least one of a catheter, a guide wire, a pointer probe, a stent, a seed, an implant, or an endoscope. 4. The surgical navigation system of claim 1 further comprising an imaging device operable to capture and provide the scan data to the data processor. 5. The surgical navigation system of claim 4 wherein the imaging device is operable to capture volumetric scan data or surface data representative of the region of interest. 6. The surgical navigation system of claim 4 wherein the imaging device is further defined as at least one of a magnetic resonance imaging device, a computed tomography imaging device, a positron emission tomography imaging device, a fluoroscopic imaging device, or an ultrasound imaging device. 7. The surgical navigation system of claim 1 wherein the tracking subsystem is further defined as an electro-magnetic localizing device having one or more electro-magnetic sensors attached to the surgical instrument. 8. The surgical navigation system of claim 7 wherein a volumetric perspective image is rendered from a point of view which correlates to one of the electro-magnetic sensors attached to the surgical instrument. 9. The surgical navigation system of claim 1 wherein the data processor is operable to render a second image of the region of interest based on the scan data, and to superimpose an indicia of the surgical instrument onto the second image of the patient. 10. The surgical navigation system of claim 9 wherein the data processor is further operable to track in real-time the position of the surgical instrument as it is moved within the region of interest and update the corresponding position of the indicia of the surgical instrument on the second image of the patient. 11. The surgical navigation system of claim 9 wherein the data processor is further operable to track in real-time the location and orientation of the surgical instrument as it is moved within the region of interest and the display is further operable to display the location and orientation of the surgical instrument. 12. A surgical instru... | Surgical Navigation Technologies, Inc. | A61B005/05
US06929611 | B2 | 2005-08-16 | Device for measuring the body temperature | 2002-06-13 | A device method and system are provided for measuring the body temperature in the area of the head. Reliable measured values are furnished even during movement and physical exercise. The device includes a head strap (1), whose strap elements (2, 3) are in contact at least partially as strap sections (7, 10) with the scalp of a user (4) of the headband and in which at least one of the strap sections (7) is designed for receiving a first temperature sensor (12) measuring the skin temperature. | 1. A device for measuring a body temperature of a subject, the device comprising: a head positioning element with a section positioned adjacent to the scalp of a subject or adjacent to the fontanelle of the subject; and a double temperature sensor with a first temperature sensor and a second temperature sensor arranged as an assembly unit opposite each other, wherein said double temperature sensor is positioned by said head positioning element such that said first temperature sensor measures a temperature th1 near the skin on the forehead of the said subject substantially in the middle between the eyes or on the head in the area of the fontanelle and said second temperature sensor measures a temperature th2 close in space to the head or the forehead; and a body temperature evaluating unit determining the body temperature of the subject from the calculation formula tc=th1+(th1−th2) * Ks/Kg in which Ks=coefficient of thermal conduction for the sensor and Kg=coefficient of thermal conduction for the tissue. | 1. A device for measuring a body temperature of a subject, the device comprising: a head positioning element with a section positioned adjacent to the scalp of a subject or adjacent to the fontanelle of the subject; and a double temperature sensor with a first temperature sensor and a second temperature sensor arranged as an assembly unit opposite each other, wherein said double temperature sensor is positioned by said head positioning element such that said first temperature sensor measures a temperature th1 near the skin on the forehead of the said subject substantially in the middle between the eyes or on the head in the area of the fontanelle and said second temperature sensor measures a temperature th2 close in space to the head or the forehead; and a body temperature evaluating unit determining the body temperature of the subject from the calculation formula tc=th1+(th1−th2) * Ks/Kg in which Ks=coefficient of thermal conduction for the sensor and Kg=coefficient of thermal conduction for the tissue. 2. A device in accordance with claim 1, wherein said head positioning element includes a strap holding said first temperature sensor fastened in an area of the forehead of the subject or in the area of the top side of the head of the subject, said strap accommodating said double temperature sensor and said double temperature sensor including an element that is at least partially self-sticking. 3. A device in accordance with claim 2, wherein said bead positioning element strap as a head strap. 4. A device in accordance with claim 3, wherein said head strap has strap elements at least partially in contact as strap sections with the scalp of the subject. 5. A device in accordance with claim 1, wherein said head positioning element includes strap sections as elements of an inner part of a protective helmet with a hard outer shell. 6. A device in accordance with claim 2, wherein said head positioning element includes strap sections as elements of a protective helmet wherein said strap sections are elements of an inner part of a protective helmet with a fix outer shell. 7. A device in accordance with claim 6, wherein said inner art is a helmet inner shell. 8. A device in accordance with claim 4, wherein said strap elements are elements of an inner part of a protective helmet with a fixed outer shell wherein said second temperature sensor is fastened in an area between the inner part an inner wall of the outer shell. 9. A device in accordance with claim 4, wherein said first temperature sensor and said second temperature sensor are arranged thermally insulated from one another. 10. A device in accordance with claim 4, further comprising a receiving station, wherein said first temperature sensor and said second temperature sensor are connected to said evaluating unit for transmitting measured temperature values to said receiving station. 11. A device in accordance with claim 10, further comprising a wireless transmission section provided betw... | Dräger Medical AG & Co. KGaA | A61B005/00
US07225011 | B2 | 2007-05-29 | Heart modeling using a template | 2002-03-28 | A device, a method and a computer program to model and report an anatomic structure, for example for coronary angiography is provided. The graphical editor models and reports the morphology/topology of the coronary anatomy with its clinical relevant items, i.e. grafts, stents, stenosis/occlusion and turtuosity. From the model different presentations (views) with adjustable display attribute settings, like scale and shape of the arteries can be constructed. | 1. A device for describing and evaluating a three-dimensional anatomical structure, the device comprising: an input means for inputting examination data of said anatomical structure;an output means for outputting a simplified two-dimensional image representation of said anatomical structure, wherein the simplified two-dimensional image representation is based on a pre-stored image template the pre-stored image template including a graphical model of clinically relevant structures of the anatomical structure and a corresponding nomenclature, the template further utilizing a hierarchical data structure for modeling the clinically relevant structures;means for graphically supplementing said simplified two-dimensional image representation with individualized details, wherein graphically supplementing with the individualized details includes annotating with relevant clinical information, including nomenclature of the anatomical structure; andmeans for carrying out an automated assignment of the graphically supplemented individualized details to the simplified two-dimensional image representation and with respect to the examination data of the anatomical structure. | 1. A device for describing and evaluating a three-dimensional anatomical structure, the device comprising: an input means for inputting examination data of said anatomical structure;an output means for outputting a simplified two-dimensional image representation of said anatomical structure, wherein the simplified two-dimensional image representation is based on a pre-stored image template the pre-stored image template including a graphical model of clinically relevant structures of the anatomical structure and a corresponding nomenclature, the template further utilizing a hierarchical data structure for modeling the clinically relevant structures;means for graphically supplementing said simplified two-dimensional image representation with individualized details, wherein graphically supplementing with the individualized details includes annotating with relevant clinical information, including nomenclature of the anatomical structure; andmeans for carrying out an automated assignment of the graphically supplemented individualized details to the simplified two-dimensional image representation and with respect to the examination data of the anatomical structure. 2. The device as claimed in claim 1, wherein said device comprises evaluation means arranged to evaluate shape, size, nomenclature and/or type of the details in an automated fashion, said device further comprising means arranged to produce a text report comprising evaluation results. 3. The device as claimed in claim 1, wherein the device is arranged to undertake an automated amendment of an alpha-numerical information to the details or their parts. 4. The device as claimed in claim 1, wherein the device is arranged to undertake an automated assignment of the details to a front or a back surface of the anatomical structure. 5. The device as claimed in claim 1, wherein the device is arranged to undertake an automated calculation of a functionality of the anatomical structure with reference to abnormal regions. 6. The device as claimed in claim 1, wherein the anatomical structure comprises a heart and the details comprise a cardiovascular system comprising arteries. 7. The device as claimed in claim 6, arranged to divide images of the arteries into individual segments, said device being further arranged to undertake an automated naming of the segments of the arteries. 8. A method for describing a three-dimensional anatomical structure, comprising: performing an inputting of individual examination data of the anatomical structure into a device for processing the data;performing an outputting of a simplified two-dimensional image representation of the anatomical structure, wherein the simplified two-dimensional image representation is based on the individual examination data and a template, the template including a graphical model of clinically relevant structures of the anatomical structure and a corresponding nomenclature, the template further utilizing a hierarchical data structure for... | Koninklijke Philips Electronics, N.V. | A61B005/05
US07144495 | B2 | 2006-12-05 | Electrochemical test strip with an integrated micro-needle and associated methods | 2002-08-23 | Methods and devices for electrochemically detecting a change in the viscosity of a fluid are provided. In the subject methods, a fluid sample (e.g., a whole blood sample) is introduced into an electrochemical test strip's electrochemical cell using a micro-needle integrated into one of oppositely spaced apart working and reference electrodes. An electric potential is applied to the electrochemical cell to first achieve a steady state cell current. A decrease in the steady state cell current is then detected and related to a change in viscosity of the sample. In many embodiments, the sample is blood and the change in viscosity is related to the onset of coagulation in the blood sample, and often the PT of the blood sample. An electrochemical test strip includes an electrochemical cell and an integrated micro-needle. The electrochemical cell includes oppositely spaced apart working and reference electrodes and a reagent mixture (e.g., a redox couple and a coagulation catalyzing agent). The micro-needle is integrated into one of the working and reference electrodes of the electrochemical cell. | 1. A method for detecting a change in the viscosity of a fluid sample, said method comprising: (a) introducing a fluid sample into an electrochemical cell of an electrochemical test strip comprising oppositely spaced apart working and reference electrodes, wherein one of the working and reference electrodes includes an integrated micro-needle for obtaining the fluid sample from a user and a capillary channel for introducing the fluid sample from the micro-needle to the electrochemical cell;(b) applying an electric potential to said electrochemical cell to produce a steady state current between said oppositely spaced apart working and reference electrodes;(c) detecting a change in said steady state current; and(d) relating said change in steady state current to a change in viscosity of said fluid sample. | 1. A method for detecting a change in the viscosity of a fluid sample, said method comprising: (a) introducing a fluid sample into an electrochemical cell of an electrochemical test strip comprising oppositely spaced apart working and reference electrodes, wherein one of the working and reference electrodes includes an integrated micro-needle for obtaining the fluid sample from a user and a capillary channel for introducing the fluid sample from the micro-needle to the electrochemical cell;(b) applying an electric potential to said electrochemical cell to produce a steady state current between said oppositely spaced apart working and reference electrodes;(c) detecting a change in said steady state current; and(d) relating said change in steady state current to a change in viscosity of said fluid sample. 2. The method of claim 1, wherein the introducing step introduces a fluid sample with a volume in the range of 100 nano-liters to 150 nano-liters. 3. The method of claim 1, wherein said fluid sample is a physiological sample. 4. The method of claim 3, wherein said physiological sample is blood. 5. The method of claim 4, wherein said method further comprises relating said change in viscosity to the prothrombin time (PT) of said blood. 6. The method of claim 1, wherein said electrochemical cell comprises a redox couple. 7. A method for detecting the onset of coagulation of a blood sample, said method comprising: (a) introducing said blood sample into an electrochemical cell of an electrochemical test strip comprising: (i) oppositely spaced apart working and reference electrodes; and(ii) a reagent mixture comprising a redox couple;wherein one of the working and reference electrodes includes an integrated micro-needle for obtaining the blood sample from a user and a capillary channel for introducing the fluid sample from the micro-needle to the electrochemical cell; (b) applying an electric potential to said electrochemical cell to produce a steady state current between said oppositely spaced apart working and reference electrodes;(c) detecting a change in said steady state current; and(d) relating said change in steady state current to the onset of coagulation in said blood sample. 8. The method of claim 7, wherein said change is a decrease. 9. The method of claim 7, wherein said reagent comprises a coagulation catalyzing agent. 10. The method of claim 9, wherein said coagulation catalyzing agent comprises thromboplastin. 11. The method of claim 10, wherein said method further comprises relating said onset of coagulation to the prothrombin time of said blood sample. | Lifescan, Inc. | G01N027/327 A61B005/00
US06986749 | B2 | 2006-01-17 | Device for inserting a guide wire and/or for handling a catheter shaft | 2002-04-05 | The invention relates to a device for handling at least one guide wire (51) in order to guide an invasive medical instrument or in order to handle a catheter shaft using invasive medical techniques. A fixing arm (13) which can be fixed to an introduction valve (1) and a clamping device (16) attached to the fixing arm (13) are provided, wherewith the relevant guide wire or the catheter shaft can be secured against movements in the longitudinal direction. This enables the introduction valves (1) to be operated with one hand, while the other hand can be used for the manipulation of the other invasive medical instruments as opposed to having to fix elements such as a guide wire or catheter shaft which cannot be moved in a longitudinal direction. | 1. A device for handling at least one guide wire in order to guide an invasive medical instrument or for handling a catheter shaft in invasive medical techniques, with a fixing arm which can be secured on an introduction valve, and with a clamping device which is arranged on the fixing arm and with which the particular guide wire or catheter shaft can be secured against longitudinal displacement, wherein the fixing arm protrudes beyond an introduction end of the introduction valve. | 1. A device for handling at least one guide wire in order to guide an invasive medical instrument or for handling a catheter shaft in invasive medical techniques, with a fixing arm which can be secured on an introduction valve, and with a clamping device which is arranged on the fixing arm and with which the particular guide wire or catheter shaft can be secured against longitudinal displacement, wherein the fixing arm protrudes beyond an introduction end of the introduction valve. 2. A device according to claim 1, characterized in that the fixing arm can be secured on the introduction valve in a releasable manner via a groove-and-tongue connection. 3. A device according to claim 1, characterized in that the fixing arm can be secured on the introduction valve in a releasable manner via a plug connection. 4. A device according to claim 2, characterized in that an adapter piece is provided with which the fixing arm can be secured in a releasable manner on the introduction valve. 5. A device according to claim 1, characterized in that the fixing arm is connected rigidly to the introduction valve. 6. A device according to claim 1, characterized in that the fixing arm has a securing portion and a holding portion which is angled off in relation to the securing portion. 7. A device according to claim 1, characterized in that the fixing arm has a securing portion and a holding portion designed as a holding plate. 8. A device according to claim 1, characterized in that the clamping device has a clamping unit for fixing a guide wire. 9. A device according to claim 1, characterized in that the clamping device has at least two clamping units which are arranged laterally spaced apart from one another and are used for fixing two guide wires or one guide wire and a catheter shaft. 10. A device according to claim 1, characterized in that the clamping device has three or four clamping units which are arranged laterally spaced apart from one another and are used for fixing guide wires and/or catheter shafts. 11. A device for handling at least one guide wire in order to guide an invasive medical instrument or for handling a catheter shaft in invasive medical techniques, with a fixing arm which can be secured on an introduction valve, and with a clamping device which is arranged on the fixing arm and with which the particular guide wire or catheter shaft can be secured against longitudinal displacement, wherein the clamping device has, as clamping unit, at least one elongated clamping rail which has a holder part with a base plate and two side cheeks and has a clamping body inserted into the holder part and designed with a longitudinally extending clamping slit. 12. A device according to claim 11, characterized in that the base plate and the side cheeks are connected rigidly to one another and at right angles to one another. 13. A device according to claim 11, characterized in that the side cheeks are each arranged pivotably on the base plate via a respe... | - | A61B005/04
US07182935 | B2 | 2007-02-27 | Bacterial plaque evidencing composition based on natural colorants | 2002-06-19 | The present invention deals with a composition for evidencing bacterial plaque based on natural colorants comprising at least one concentrated solution of natural colorant selected from the group consisting of colorants extracted from the aai (Euterpe oleracea) and colorants extracted from urucum (Bixa orellana), in a pharmaceutically acceptable vehicle and, optionally, acceptable pharmaceutical additives. Preferentially, in the case of the aai solution, the concentration varies from 60 to 95% of concentrated extract, whilst for the urucum solution, the concentration varies between 56 and 90% of concentrated extract.;Also preferentially, in solutions containing a mixture of the extracts, these will be present in a proportion that varies from 10:1 to 1:1. | 1. A bacterial plaque evidencing composition comprising: (i) at least one concentrated solution of natural colorants, said concentrated solution of natural colorants being obtained from aai (Euterpe oleracea) by an extraction process comprising adding alcohol-containing agueous solution to ripe aai fruit, cold maceration in the absence of light, and concentration of the resulting extract; and(ii) a pharmaceutically acceptable vehicle, wherein the composition is effective to color bacterial plaque formations. | 1. A bacterial plaque evidencing composition comprising: (i) at least one concentrated solution of natural colorants, said concentrated solution of natural colorants being obtained from aai (Euterpe oleracea) by an extraction process comprising adding alcohol-containing agueous solution to ripe aai fruit, cold maceration in the absence of light, and concentration of the resulting extract; and(ii) a pharmaceutically acceptable vehicle, wherein the composition is effective to color bacterial plaque formations. 2. The composition according to claim 1 further comprising pharmaceutically acceptable additives. 3. The composition according to claim 1 wherein the pharmaceutically acceptable vehicle is selected from the group consisting of water, alcohol or their mixtures. 4. The composition according to claim 2 having at least one pharmaceutically acceptable additive selected from the group consisting of thickeners, flavorings, sweeteners, dispersants and preservatives. 5. The composition according to claim 4 wherein the thickener is selected from the group consisting of glycerine, glycol propylene, gelatine and their mixtures. 6. The composition according to claim 4 wherein the flavoring is selected from the group consisting of natural flavoring oils, synthetic flavorings and their mixtures. 7. The composition according to claim 2 comprising from 60 to 95% of the at least one concentrated solution of natural colorants, from 5 to 25% of pharmaceutically acceptable vehicles, and from 0 to 13% of pharmaceutically acceptable additives. 8. A method of disclosing oral bacterial plaque comprising contacting teeth with the evidencing composition according to claim 1. 9. A process of using the evidencing composition of claim 1 comprising: (i) conventional cleaning of the oral cavity;(ii) direct application of said composition to the surface of the teeth; and (iii) mechanical removal of the bacterial plaques evidenced by the said evidencing composition. | Empresa Brasileira De Pesquisa Agropecuaria Embrapa Universidade Federal Do Para | A61B005/00 A61K036/00
US07116810 | B2 | 2006-10-03 | Method and system for airway measurement | 2002-11-27 | A method for measuring tubular anatomical structures, for example lung airways, comprises isolating by at least one segmentation process a given tubular anatomical structure of interest and measuring at least one attribute of the structure of interest. A system for measuring lung airways using acquired image data is provided. The system comprises an imaging device for acquiring the image data and an image processing device coupled to the imaging device. The imaging processing device is configured for isolating by at least one segmentation process a given airway of interest, fitting an inner ellipse to an inner boundary of the given airway and an outer ellipse to an outer boundary of the airway structure using statistical techniques at a given point in the airway, and further configured for generating measurements of the given airway using the inner and outer ellipses. | 1. A method for measuring lung airways from acquired image data comprising: isolating by at least one segmentation process a given airway of interest;fitting an inner ellipse to an inner boundary of the given airway and an outer ellipse to an outer boundary of the airway structure using statistical techniques at a given point in the airway; and,generating measurements of the given airway using the inner and outer ellipses. | 1. A method for measuring lung airways from acquired image data comprising: isolating by at least one segmentation process a given airway of interest;fitting an inner ellipse to an inner boundary of the given airway and an outer ellipse to an outer boundary of the airway structure using statistical techniques at a given point in the airway; and,generating measurements of the given airway using the inner and outer ellipses. 2. The method of claim 1 wherein the at least one segmentation process comprises segmenting, detecting and grouping of homogeneous regions within the acquired image data. 3. The method of claim 1 wherein the measurements comprises at least one of average wall thickness, wall thickness variations, volume measurements, two-dimensional (2D)) area measurements and volume area distribution. 4. The method of claim 1 wherein the measurements are used for at least one of disease diagnosis and tracking of disease progression, and wherein the disease is chronic obstructive pulmonary disease. 5. The method of claim 1 wherein the steps of isolating, fitting and generating measurements are repeated for image data corresponding to a length of the given airway to generate measurements along the length of the given airway. 6. The method of claim 1 further comprising generating an output, wherein the output is used for at least one of staging the given disease in a patient, measuring response to therapy, phenotyping for patient selection to participate in drug trials, measuring stability of an anatomical structure and prediction of rate of change of the given disease. 7. A system for measuring lung airways using acquired image data comprising: an imaging device for acquiring the image data; and,an image processing device coupled to the imaging device and configured for isolating by at least one segmentation process a given airway of interest, fitting an inner ellipse to an inner boundary of the given airway and an outer ellipse to an outer boundary of the airway structure using statistical techniques at a given point in the airway, and further configured for generating measurements of the given airway using the inner and outer ellipses. 8. The system of claim 7 wherein the measurements comprises at least one of average wall thickness, wall thickness variations, volume measurements, two-dimensional (2D) area measurements and volume area distribution. 9. The system of claim 7 wherein the measurements are used for at least one of disease diagnosis and tracking of disease progression, and wherein the disease is at least one of chronic obstructive pulmonary disease and asthma. 10. The system of claim 7 further comprising a display device coupled to the image processing device for reporting measurements to a system user. 11. The system of claim 7 wherein the image processor is further adapted to generate an output and wherein the output is used for at least one of staging the given disease in a patient, measuring response to therapy, phenot... | General Electric Company | G06K009/00 A61B005/08
US06892088 | B2 | 2005-05-10 | Computer-assisted bone densitometer | 2002-09-18 | An x-ray densitometry system provides computer assistance to the operator in identifying possible sources of scanning or analysis error through computer review of the acquired data, operator input, and the ultimate diagnostic outputs. | 1. A computer aided bone densitometry system comprising: an x-ray source and detector opposable about a patient to produce signals indicating x-ray attenuation by bone of the patient; a computer receiving the signals and executing a stored program to: (a) control the x-ray source and detector to signals for a plurality of points over a scan area; (b) calculate, for the plurality of points, a bone mineral data set indicating x-ray attenuation caused by bone; (c) measure the bone mineral data set to evaluate the likelihood that acquisition of the signals was faulty; and (d) output an indication of faulty data acquisition to the operator when the evaluation of the bone mineral data set indicates that the acquired data is likely faulty. | 1. A computer aided bone densitometry system comprising: an x-ray source and detector opposable about a patient to produce signals indicating x-ray attenuation by bone of the patient; a computer receiving the signals and executing a stored program to: (a) control the x-ray source and detector to signals for a plurality of points over a scan area; (b) calculate, for the plurality of points, a bone mineral data set indicating x-ray attenuation caused by bone; (c) measure the bone mineral data set to evaluate the likelihood that acquisition of the signals was faulty; and (d) output an indication of faulty data acquisition to the operator when the evaluation of the bone mineral data set indicates that the acquired data is likely faulty. 2. The computer aided densitometer of claim 1 wherein the measurement of the bone mineral data set compares peak attenuations in the bone mineral data set with an expected peak attenuation to detect faulty acquisition caused by foreign objects with high density. 3. The computer aided densitometer of claim 1 wherein the measurement of the bone mineral data set detects patient motion causing faulty acquisition. 4. The computer aided densitometer of claim 3 wherein the measurement of the bone mineral data set calculates discontinuities in bone edges to detect patient motion. 5. The computer aided densitometer of claim 3 wherein the measurement of the bone mineral data set calculates jumps in density within a bone to detect patient motion. 6. The computer aided densitometer of claim 1 wherein the measurement of the bone mineral data set detects mispositioning of the scan area on the patient causing faulty acquisition. 7. The computer aided densitometer of claim 6 wherein the measurement of the bone mineral data set calculates a difference between the location of bones represented by the bone mineral data set and an expected location of the bones to detect mispositioning of the scan area on the patient. 8. The computer aided densitometer of claim 6 wherein the measurement of the bone mineral data set detects mispositioning of the scan area with respect to the vertebrae of the spine. 9. The computer aided densitometer of claim 6 wherein the measurement of the bone mineral data set detects mispositioning of the scan area with respect to the proximal femur. 10. The computer aided densitometer of claim 1 wherein the patient is supported horizontally on a table and wherein the scan area encompasses the spine wherein the measurement of the bone mineral data set detects excessive displacement of the spine from a center of the table causing faulty acquisition. 11. The computer aided densitometer of claim 1 wherein the patient is supported horizontally on a table and wherein the scan area encompasses the spine wherein the measurement of the bone mineral data set detects angulation of the spine from parallel with a long axis of the table causing faulty acquisition. 12. The computer aided densitometer of claim 1 wherein t... | General Electric Company | A61B005/05 G01N023/06
US06878116 | B2 | 2005-04-12 | Blood pressure analyzer | 2002-10-23 | A blood pressure analyzer having an interconnected pressure sensor, a signal processor, a microprocessor, and a display. The microprocessor is also connected to memory and a computer transmission interface. The signal processor including a first signal amplifier, a low pass filter, a second signal amplifier, a high pass filter, a voltage clamp, a comparator, a third signal amplifier, an accumulator circuit, and a crystal oscillator. The analyzed and calculated systolic and diasystolic blood pressure values are shown on the display and, furthermore, accurately represents pulse wave form fluctuations to afford the user a complete understanding of blood pressure conditions and thereby provide physicians a reference for the diagnosis of cardiovascular changes. | 1. A blood pressure analyzer comprising: an interconnected pressure sensor, a signal processor, and a microprocessor, with said microprocessor also connected to memory and said signal processor comprising a first signal amplifier, a low pass filter, a second signal amplifier, a high pass filter, a voltage clamp, a comparator, a third signal amplifier, an accumulator circuit, and a crystal oscillator such that said pressure sensor registers changes in pressure that are converted by said signal processor into high resolution, digital pulse width modulation (PWM) signal values and said microprocessor controls the saving into said memory of the digital signal values aggregated during the measurement process as well as the calculation of said digital signal values to obtain the average value, high pressure value, and low pressure value of blood pressure, wherein said microprocessor executes the equation below to solve for the respective values Pa, Ph, and P1 of the blood pressure average curve, the high pressure base curve, and the low pressure base curve: Pa=(An+A (n+1))/2, wherein An is the nth point before the average blood pressure value, and A(n+1) is the (n+1) point of the blood pressure value, with the calculation of Pa consisting of first sequentially locating all points before the average blood pressure value and then ascertaining the final point of the average blood pressure value, said value being the value of Pa; taking Pa as a base value, each An point is sequentially integrated once to obtain a respective Kn value and each Kn is compared to Pa, and when a Kn is smaller than Pa, the value of Ph is such that Ph=An; taking Pa as a base value, each An point is sequentially differentiated twice to obtain a respective Km value and each Km is compared to Pa, and when a Km is larger than Pa, the value of P1 is such that P1=An. | 1. A blood pressure analyzer comprising: an interconnected pressure sensor, a signal processor, and a microprocessor, with said microprocessor also connected to memory and said signal processor comprising a first signal amplifier, a low pass filter, a second signal amplifier, a high pass filter, a voltage clamp, a comparator, a third signal amplifier, an accumulator circuit, and a crystal oscillator such that said pressure sensor registers changes in pressure that are converted by said signal processor into high resolution, digital pulse width modulation (PWM) signal values and said microprocessor controls the saving into said memory of the digital signal values aggregated during the measurement process as well as the calculation of said digital signal values to obtain the average value, high pressure value, and low pressure value of blood pressure, wherein said microprocessor executes the equation below to solve for the respective values Pa, Ph, and P1 of the blood pressure average curve, the high pressure base curve, and the low pressure base curve: Pa=(An+A (n+1))/2, wherein An is the nth point before the average blood pressure value, and A(n+1) is the (n+1) point of the blood pressure value, with the calculation of Pa consisting of first sequentially locating all points before the average blood pressure value and then ascertaining the final point of the average blood pressure value, said value being the value of Pa; taking Pa as a base value, each An point is sequentially integrated once to obtain a respective Kn value and each Kn is compared to Pa, and when a Kn is smaller than Pa, the value of Ph is such that Ph=An; taking Pa as a base value, each An point is sequentially differentiated twice to obtain a respective Km value and each Km is compared to Pa, and when a Km is larger than Pa, the value of P1 is such that P1=An. 2. A blood pressure analyzer as claimed in claim 1 in which said microprocessor is also connected to a display storage device and said display storage device is connected to a display. 3. A blood pressure analyzer as claimed in claim 1 in which said microprocessor is also connected to a computer transmission interface. 4. A blood pressure analyzer comprising: an interconnected pressure sensor, a signal processor, and a microprocessor, with said microprocessor also connected to memory and said signal processor comprising a first signal amplifier, a low pass filter, a second signal amplifier, a high pass filter, a voltage clamp, a comparator, a third signal amplifier, an accumulator circuit, and a crystal oscillator such that said pressure sensor registers changes in pressure that are converted by said signal processor into high resolution, digital pulse width modulation (PWM) signal values and said microprocessor controls the saving into said memory of the digital signal values aggregated during the measurement process as well as the calculation of said digital signal values to obtain the average value, high pressure va... | - | A61B005/02
US06909912 | B2 | 2005-06-21 | Non-invasive perfusion monitor and system, specially configured oximeter probes, methods of using same, and covers for probes | 2002-06-20 | The present invention relates to a novel non-invasive perfusion/resistance status monitor system and methods of using the same, and more specifically, a vascular perfusion status monitor system receiving and processing signals from at least two pulse oximeter probes, where each of the at least two pulse oximeter probes are situated at advantageously different locations in a patient. Novel pulse oximeter probes are configured to be placed, respectively, across the lip or cheek, across the septum or nares of the nose, and on the tongue. These probes are fabricated to provide signals to estimate arterial oxygen saturation. Conventional oximeter probes also can be configured to function according to the novel methods of determining differences in peripheral blood flow and/or resistance described herein. | 1. A method for analyzing pulse oximeter probe signals from at least two sites on a living vertebrate to determine the presence or absence of, or to monitor changes in, impaired peripheral perfusion, comprising the steps of: a. removably affixing a first pulse oximeter probe to a central source site located in the head of a patient in need of said method; b. removably affixing a second pulse oximeter probe to a first peripheral site of interest on said patient; c. measuring signals from said central source site pulse oximeter and said first peripheral site pulse oximeter, d. averaging sufficient sequential signals from each site to obtain a statistically reliable average, and from said average calculating a time-set estimate of the arterial blood oxygen saturation, and repeating this averaging to obtain sequential time-set estimates from each site; e. comparing a time-set estimate of the arterial blood oxygen saturation from the central source site with an estimate, taken at a similar time, of the arterial blood oxygen saturation from the first peripheral site; and f. determining the presence or absence of, or the time-based changes in, impaired peripheral perfusion proximal to said first peripheral site. | 1. A method for analyzing pulse oximeter probe signals from at least two sites on a living vertebrate to determine the presence or absence of, or to monitor changes in, impaired peripheral perfusion, comprising the steps of: a. removably affixing a first pulse oximeter probe to a central source site located in the head of a patient in need of said method; b. removably affixing a second pulse oximeter probe to a first peripheral site of interest on said patient; c. measuring signals from said central source site pulse oximeter and said first peripheral site pulse oximeter, d. averaging sufficient sequential signals from each site to obtain a statistically reliable average, and from said average calculating a time-set estimate of the arterial blood oxygen saturation, and repeating this averaging to obtain sequential time-set estimates from each site; e. comparing a time-set estimate of the arterial blood oxygen saturation from the central source site with an estimate, taken at a similar time, of the arterial blood oxygen saturation from the first peripheral site; and f. determining the presence or absence of, or the time-based changes in, impaired peripheral perfusion proximal to said first peripheral site. 2. The method of claim 1 wherein said central source site is selected from the group consisting of cheek, nasal septum, alar nares, and tongue. 3. The method of claim 1 wherein said first peripheral site is selected from the group consisting of a finger, a toe, and a section of tissue distal from an area of potential damage or disease adversely affecting, suspected or known to have compromised the peripheral vascular resistance and/or peripheral blood perfusion, and/or peripheral vascular disease. 4. The method of claim 1 wherein said monitoring assesses changes over time in oxygenation of tissues proximal to said second pulse oximeter during drug dosing/effectiveness and/or reoxygenation procedures. 5. The method of claim 1, additionally comprising estimating the blood pressure of said living vertebrate through analysis of pulse time delay of said signals between said at least two sites. 6. The method of claim 5, additionally comprising comparing the analysis of said pulse time delay with blood pressure estimates obtained by other means, and applying a factor to enhance the accuracy of said estimating the blood pressure. 7. The method of claim 1, wherein said presence of impaired peripheral perfusion is related to decreased peripheral circulation caused by a condition selected from the group consisting of atherosclerosis, Raynaud's disease, Buerger's disease, chronic obstructive pulmonary disease, embolic occlusive disease, and progressive diabetes. 8. The method of claim 1, wherein said presence of impaired peripheral perfusion is related to lower circulation to the limbs of said living vertebrate, caused by an acute condition selected from the group consisting of shock, trauma, and mechanical injury. 9. Apparatus for analyzing pulse ... | University of Florida Beta Biomed, Inc. | A61B005/00
US06978167 | B2 | 2005-12-20 | Video pose tracking system and method | 2002-07-01 | A measurement system and method for detecting and tracking the pose of an object displaceable in a coordinate reference frame. In the system and method a visible target pattern is provided on the object. The visible target pattern includes a series of contrast regions for providing at least two feature points. Each of these feature points is located at a juncture of an optically detectable plurality of edges. Each edge in the optically detectable plurality of edges separates different contrast regions. The method and system determine the location of the feature points by first locating the edges using the change in contrast between contrast regions, and then determining junctures of multiple edges. The method and system also involve a stereoscopic digital camera for generating a pair of digital images of the target pattern and a marker template comprising a set of reference characteristics including a relationship between the feature points. | 1. A measurement system for tracking a pose of an object displaceable in a coordinate reference frame, the system comprising: a) an optically visible target pattern secured to the object, the target pattern containing an arrangement of contrast regions for defining at least two feature points, each feature point being located at a juncture of an optically detectable plurality of edges, wherein each edge in the optically detectable plurality of edges separates different contrast regions;b) an optical sensor assembly for generating a pair of digital images of the target pattern containing the feature points, the pair of images being generated from dissimilar viewing angles;c) a marker template comprising a unique set of reference characteristics including a relationship between the feature points; and,d) a processor for testing the pixels of the digital images to determine the spatial location in the reference frame for each of the feature points of the target pattern, for comparing the determined spatial locations of the feature points to the set of reference characteristics to confirm the recognition of the target pattern, and for calculating the object pose from the spatial location of the feature points. | 1. A measurement system for tracking a pose of an object displaceable in a coordinate reference frame, the system comprising: a) an optically visible target pattern secured to the object, the target pattern containing an arrangement of contrast regions for defining at least two feature points, each feature point being located at a juncture of an optically detectable plurality of edges, wherein each edge in the optically detectable plurality of edges separates different contrast regions;b) an optical sensor assembly for generating a pair of digital images of the target pattern containing the feature points, the pair of images being generated from dissimilar viewing angles;c) a marker template comprising a unique set of reference characteristics including a relationship between the feature points; and,d) a processor for testing the pixels of the digital images to determine the spatial location in the reference frame for each of the feature points of the target pattern, for comparing the determined spatial locations of the feature points to the set of reference characteristics to confirm the recognition of the target pattern, and for calculating the object pose from the spatial location of the feature points. 2. The system of claim 1 further comprising at least three feature points to define a pair of non-collinear target vectors. 3. The system of claim 2, wherein at least four feature points define the pair of non-collinear target vectors. 4. The system of claim 2, wherein a plurality of the feature points defines at least three non-collinear target vectors. 5. The system of claim 2 further comprising a plurality of target patterns with at least one of the feature points present on each one of the plurality of target patterns. 6. The system of claim 2, wherein the reference characteristics are selected from the group comprising a distance between the adjacent feature points, the number of feature points, a distance between adjacent target vectors, and an angle between adjacent target vectors. 7. The system of claim 1 wherein the contrast regions include a plurality of alternating dark and bright regions. 8. The system of claim 7 further comprising an intensity gradient between adjacent pixels in the digital images, one of the adjacent pixels being in one of the dark regions and the other of the adjacent pixels being in one of the bright regions. 9. The system of claim 8, wherein a distribution of intensity gradients corresponding to the alternating dark and bright regions defines an arrangement of alternating intensity peaks and valleys. 10. The system of claim 9, wherein each of the feature points is an intensity saddle located in the center of the arrangement of alternating peaks and valleys. 11. The system of claim 10, wherein the type of the feature points is determined by the number of peaks and valleys associated with each of the feature points. 12. The system of claim 11, wherein the feature point type is included in the reference... | Claron Technology Inc. | A61B005/05
US07155377 | B2 | 2006-12-26 | Method and system for calculating the spatial-temporal effects of climate and other environmental conditions on animals | 2001-08-23 | The present invention is a system and method for accurately calculating the spatial-temporal effects of a variety of environmental conditions on animal individual, population and community dynamics, given the animal's temperature-dependent behaviors, morphology and physiology, by running integrated microclimate and animal models to calculate the discretionary energy and water available to the animal and its activity time. The methodology requires relatively few, easily measured data to perform the calculations. The microclimate model translates a set of climate and other environmental conditions into a set of microclimate conditions experienced by an animal, given a set of the animal's characteristics. The animal model uses the microclimate conditions data and the set of animal characteristics data to solve for several animal conditions. Further calculations and several display options are available to a user, including spatial-temporal analyses. The method is implemented by use of a client-server system employing a graphical user interface, and a programming structure that allows for use of the methodology with personal computers, despite the large sizes of the databases that must be accessed. The method may also be implemented on individual computers. | 1. A method of using a computer for calculating an at least one effect of a plurality of environmental conditions on an animal, using said computer to perform the following steps comprising: providing a data set of the plurality of environmental conditions for an at least one locality during an at least one time interval;providing a data set of characteristics of the animal;translating the data set of environmental conditions into a set of microclimate conditions experienced by the animal, by running a microclimate model to generate values for the set of microclimate conditions using the data set of environmental conditions and the data set of animal characteristics, the microclimate model comprising a solar radiation model to enable the calculation of an input of solar radiation when the sun is above the horizon and when the sun is below the horizon; and,solving for a set of animal conditions by running an animal model using the set of microclimate conditions and the data set of animal characteristics, for each of the at least one time intervals at each of the at least one localities, the set of animal conditions comprising an amount of discretionary energy and an amount of discretionary water available to the animal, and a total activity time of the animal. | 1. A method of using a computer for calculating an at least one effect of a plurality of environmental conditions on an animal, using said computer to perform the following steps comprising: providing a data set of the plurality of environmental conditions for an at least one locality during an at least one time interval;providing a data set of characteristics of the animal;translating the data set of environmental conditions into a set of microclimate conditions experienced by the animal, by running a microclimate model to generate values for the set of microclimate conditions using the data set of environmental conditions and the data set of animal characteristics, the microclimate model comprising a solar radiation model to enable the calculation of an input of solar radiation when the sun is above the horizon and when the sun is below the horizon; and,solving for a set of animal conditions by running an animal model using the set of microclimate conditions and the data set of animal characteristics, for each of the at least one time intervals at each of the at least one localities, the set of animal conditions comprising an amount of discretionary energy and an amount of discretionary water available to the animal, and a total activity time of the animal. 2. The method of claim 1, further comprising storing the set of animal conditions data in an animal conditions database. 3. The method of claim 2, further comprising calculating an at least one user-specifiable variable using the data stored in the animal conditions database. 4. The method of claim 3, wherein the at least one user-specifiable variable comprises a reproductive potential of the animal calculated by using the amounts of discretionary energy and water available to the animal. 5. The method of claim 3, wherein the at least one user-specifiable variable comprises a growth potential of the animal calculated by using the amounts of discretionary energy and water available to the animal. 6. The method of claim 3, wherein the at least one user-specifiable variable comprises a mortality rate of the animal calculated by using the total activity time of the animal. 7. The method of claim 3, wherein the at least one user-specifiable variable comprises an optimal body size of the animal for the at least one locality and for the at least one time interval. 8. The method of claim 1, further comprising displaying the set of animal conditions data. 9. The method of claim 8, wherein displaying the set of animal conditions comprises plotting the set of animal conditions by locality to analyze a spatial variation in the animal conditions. 10. The method of claim 1, wherein running the animal model comprises: calculating an amount of metabolic heat generation required to maintain a core temperature of the animal and noting an activity state of the animal, by performing a heat balance analysis, the activity state being chosen from the group consisting of active and non-active;calculating t... | Wisconsin Alumni Research Foundation | G06G007/48 A61B005/00 C05F001/00
US07191004 | B2 | 2007-03-13 | Capture verification using an evoked response reference | 2002-12-31 | A method and system for verifying capture in the heart involves the use of pacing artifact templates. One or more pacing artifact templates characterizing a post pace artifact signal associated with a particular pace voltage or range of voltages are provided. A pacing artifact template is canceled from a cardiac signal sensed following a pacing pulse. Capture is detected by comparing the pacing artifact canceled cardiac signal to an evoked response reference. Fusion/pseudofusion detection involves determining a correlation between a captured response template and a sensed cardiac signal. | 1. A method for detecting capture of a patient's heart, comprising: providing one or more pacing artifact templates, each pacing artifact template formed using a cardiac signal of at least one non-captured cardiac cycle;providing at least one template characterizing an evoked response;sensing a cardiac signal following a pacing pulse of a cardiac cycle, the cardiac cycle subsequent to the non-captured cardiac cycle;canceling a particular pacing artifact template of the one or more pacing artifact templates from the cardiac signal sensed following the pacing pulse of the subsequent cardiac cycle; anddetecting capture of the heart for the subsequent cardiac if the pacing artifact canceled cardiac signal is correlated to the evoked response template. | 1. A method for detecting capture of a patient's heart, comprising: providing one or more pacing artifact templates, each pacing artifact template formed using a cardiac signal of at least one non-captured cardiac cycle;providing at least one template characterizing an evoked response;sensing a cardiac signal following a pacing pulse of a cardiac cycle, the cardiac cycle subsequent to the non-captured cardiac cycle;canceling a particular pacing artifact template of the one or more pacing artifact templates from the cardiac signal sensed following the pacing pulse of the subsequent cardiac cycle; anddetecting capture of the heart for the subsequent cardiac if the pacing artifact canceled cardiac signal is correlated to the evoked response template. 2. The method of claim 1, wherein: providing the one or more pacing artifact templates comprises providing a number of pacing artifact templates respectively associated with a number of pacing voltages; andcanceling the particular pacing artifact template comprises canceling a pacing artifact template associated with a pacing voltage used for the pacing pulse. 3. The method of claim 1, wherein canceling the particular pacing artifact template further comprises normalizing the particular pacing artifact template with respect to the sensed cardiac signal. 4. The method of claim 1, wherein providing the one or more pacing artifact templates further comprises updating the one or more pacing artifact templates. 5. The method of claim 1, wherein providing each of the one or more pacing artifact templates comprises: delivering one or more pace pulses to the heart, wherein the one or more pace pulses do not produce capture of the heart;sensing one or more pacing artifact waveforms produced by the one or more pulses; andforming the pacing artifact template using the sensed pacing artifact waveforms. 6. The method of claim 5, wherein delivering the one or more pace pulses to the heart comprises delivering the one or more pace pulses to the heart at a particular energy level. 7. The method of claim 5, wherein delivering the one or more pace pulses to the heart comprises delivering the one or more pace pulses to the heart at a particular range of energy levels. 8. The method of claim 5, wherein forming the pacing artifact template further comprises combining pacing artifact waveforms. 9. The method of claim 5, wherein sensing the one or more pacing artifact waveforms comprises sensing the pacing artifact waveforms during a capture verification window. 10. The method of claim 9, wherein the capture verification window is a time interval beginning at a first predetermined time following a pacing pulse and continuing for a second predetermined time. 11. The method of claim 10, wherein the first predetermined time is about 25 ms and the second predetermined time is about 50 ms. 12. The method of claim 1, wherein providing each of the one or more pacing artifact templates further comprises: delivering one or ... | Cardiac Pacemakers, Inc. | A61N001/37 A61N001/372 A61B005/04
US07221974 | B1 | 2007-05-22 | Uneven-counter-rotational coil based MRI RF coil array | 2002-02-27 | A MRI RF coil array is formed from a first coil having a null B1 point and a quasi-one-peak sensitivity profile, and a second coil oriented with respect to the first coil to reduce coupling. | 1. A MRI RF coil array, said array comprising: a first coil having a null B1 point and a quasi-one-peak sensitivity profile with only one peak; anda second coil oriented with respect to said first coil to reduce coupling. | 1. A MRI RF coil array, said array comprising: a first coil having a null B1 point and a quasi-one-peak sensitivity profile with only one peak; anda second coil oriented with respect to said first coil to reduce coupling. 2. A MRI RF coil array according to claim 1, wherein said second coil is within said first coil. 3. A MRI coil array according to claim 1, wherein said second coil overlaps said first coil. 4. A MRI coil array according to claim 1, wherein said second coil is cascaded with said first coil. 5. A MRI coil array according to claim 1, wherein said second coil is solenoidal. 6. A MRI RF coil array according to claim 1, wherein said second coil is placed near the null B1 point. 7. A MRI RF coil array according to claim 1, wherein said first coil comprises a plurality of winding sections with the B1 field produced by one winding section stronger than the B1 field produced by another winding section. 8. A MRI RF coil array according to claim 1, wherein said first coil comprises a plurality of winding sections, a number of the winding sections based on sensitivity parameters. 9. A MRI RF coil array according to claim 1, wherein said first coil comprises a plurality of winding sections and at least one of (i) a separation between winding sections and (ii) a diameter of the winding sections is based on sensitivity parameters. 10. A MRI RF coil array according to claim 1, wherein said second coil is positioned relative to said first coil based on B1 field strength. 11. A MRI coil array, said array comprising: a first solenoidal coil having a first section and a second section, said first section having more turns than said second section and said second section having a counter-rotational orientation with respect to said first section, said first coil having a quasi-one-peak sensitivity profile with only one peak; anda second solenoidal coil, said second coil being oriented with respect to said first coil to reduce coupling. 12. A MRI coil array according to claim 11, wherein said second coil is oriented between said first and second sections. 13. A MRI coil array according to claim 11, wherein said second coil is oriented about said second section. 14. A MRI coil array according to claim 11, wherein said second coil is cascaded with said first coil. 15. A MRI coil array according to claim 11, further comprising an orthogonal coil forming a quadrature pair with each of said solenoidal coils. 16. A method for providing a MRI coil array, said method comprising: configuring a first coil having a null B1 point and a quasi-one-peak sensitivity profile with only one peak; andconfiguring a second coil oriented with respect to said first coil to reduce coupling. 17. A method according to claim 16, wherein said second coil overlaps said first coil. 18. A method according to claim 16, wherein said second coil is cascaded with said first coil. | General Electric Company | A61B005/055
US07215993 | B2 | 2007-05-08 | Cardiac rhythm management systems and methods for detecting or validating cardiac beats in the presence of noise | 2002-08-06 | This document describes systems, devices, and methods for detecting or validating cardiac beats, such as in the presence of myopotential or other noise. In one example, an amplitude peak, which is a candidate for a detected cardiac beat, is used in a weighted average, along with a preceding and subsequent sample. The weighted average is compared to a noise threshold. Based on the result of comparison, the amplitude peak is either deemed an actual cardiac beat, or otherwise is deemed noise. The described systems, devices, and methods improve the accuracy of detecting an actual cardiac beat in the presence of noise, during normal sinus rhythm or during an arrhythmia such as ventricular fibrillation. This, in turn, improves the accuracy with which therapy is delivered or withheld by an implantable cardiac rhythm management device. In one example, such as where the system includes a cardiac signal detector with automatic gain control (AGC) circuitry, the weighted average is normalized. | 1. A system comprising: a cardiac signal detector, comprising a detector input configured to be coupled to a first electrode associated with a heart, and comprising a detector output providing a sampled cardiac signal; anda signal processor circuit, coupled to the detector output, the signal processor circuit configured to distinguish, using a peak sample, a preceding sample to the peak sample, and a subsequent sample to the peak sample, a cardiac depolarization from noise, and in which the signal processor circuit is configured to form an average statistic computed using as data points an amplitude of the peak sample from a baseline, an amplitude of the preceding sample from the same baseline, and an amplitude of the subsequent sample from the same baseline, and to distinguish between a cardiac depolarization and noise by comparing the statistic to a predetermined threshold value. | 1. A system comprising: a cardiac signal detector, comprising a detector input configured to be coupled to a first electrode associated with a heart, and comprising a detector output providing a sampled cardiac signal; anda signal processor circuit, coupled to the detector output, the signal processor circuit configured to distinguish, using a peak sample, a preceding sample to the peak sample, and a subsequent sample to the peak sample, a cardiac depolarization from noise, and in which the signal processor circuit is configured to form an average statistic computed using as data points an amplitude of the peak sample from a baseline, an amplitude of the preceding sample from the same baseline, and an amplitude of the subsequent sample from the same baseline, and to distinguish between a cardiac depolarization and noise by comparing the statistic to a predetermined threshold value. 2. The system of claim 1, in which the signal processor circuit is configured to compute a weighted average using the peak sample, the preceding sample, and the subsequent sample, and to distinguish between a cardiac depolarization and noise by comparing the weighted average to a predetermined threshold value. 3. The system of claim 2, in which the signal processor circuit is configured to compute a normalized weighted average using the peak sample, the preceding sample, and the subsequent sample, and to distinguish between a cardiac depolarization and noise by comparing the normalized weighted average to a predetermined threshold value. 4. The system of claim 3, in which the normalized weighted average is computed as NWA=[P(n−1)+2*P(n)+P(n+1)]/[4*P(n)]. 5. The system of claim 4, in which the predetermined threshold value is about 0.62. 6. The system of claim 1, further comprising an electrode coupled to the detector input. 7. The system of claim 6, further comprising an intracardiac leadwire carrying the electrode. 8. The system of claim 1, further comprising: an implantable device including the cardiac signal detector and the signal processing circuit; anda remote user interface, configured to be communicatively coupled to the implantable device. 9. The system of claim 1, in which the signal processor circuit is configured to distinguish a normal sinus rhythm cardiac depolarization from noise. 10. The system of claim 1, in which the signal processor circuit is configured to distinguish an arrhythmic cardiac depolarization from noise. 11. The system of claim 10, in which the signal processor circuit is configured to distinguish a ventricular fibrillation cardiac depolarization from noise. 12. An implantable cardiac rhythm management device comprising: a cardiac signal detector, comprising a cardiac signal detector input configured to be coupled to a first electrode associated with a heart to receive an intrinsic cardiac signal therefrom, and comprising a cardiac signal detector output providing a sampled cardiac signal; anda depolarization detector circuit,... | Cardiac Pacemakers, Inc. | A61B005/0456
US06990368 | B2 | 2006-01-24 | Method and apparatus for virtual digital subtraction angiography | 2002-04-04 | A medical imaging system is provided that detects any change in relative position between the patient and the imaging device and corrects the image data for any motion that occurs. The medical imaging system includes: an imaging device for capturing two or more image data sets representative of a patient; a tracking subsystem for capturing patient position data that is indicative of the position of the patient and device position data that is indicative of the position of the imaging device; an image subtraction subsystem for performing a digital subtraction operation between at least two image data sets; and a motion correction subsystem configured to detect a change in the relative position between the patient and the imaging device and, upon detecting a change in the relative position, compensate in at least one of the first image data set and the second image data set for the change in relative position prior to performing the digital subtraction operation. | 1. A medical imaging system, comprising: an imaging device operable to capture two or more image data sets representative of a patient; a tracking subsystem operable to detect patient position data corresponding to each image data set and device position data corresponding to each image data set, where the patient position data is indicative of the position of the patient and the device position data is indicative of the position of the imaging device; a motion correction subsystem adapted to receive patient position data corresponding to the first and second image data set and device position data corresponding to the first and second image data set, and the motion correction subsystem is operable to detect a change in relative position between the patient and the imaging device; and an image subtraction subsystem adapted to receive at least a first image data set and a second image data set from the imaging device, the image subtraction subsystem is operable to perform a digital subtraction operation between the first image data set and the second image data set when no substantial change occurs in the relative position between the patient and the imaging device. | 1. A medical imaging system, comprising: an imaging device operable to capture two or more image data sets representative of a patient; a tracking subsystem operable to detect patient position data corresponding to each image data set and device position data corresponding to each image data set, where the patient position data is indicative of the position of the patient and the device position data is indicative of the position of the imaging device; a motion correction subsystem adapted to receive patient position data corresponding to the first and second image data set and device position data corresponding to the first and second image data set, and the motion correction subsystem is operable to detect a change in relative position between the patient and the imaging device; and an image subtraction subsystem adapted to receive at least a first image data set and a second image data set from the imaging device, the image subtraction subsystem is operable to perform a digital subtraction operation between the first image data set and the second image data set when no substantial change occurs in the relative position between the patient and the imaging device. 2. The medical imaging system of claim 1 wherein the imaging device includes an image source operable to emanate radiation towards the patient and an image receiver positioned to receive radiation from the image source. 3. The medical imaging system of claim 1 wherein the imaging device is selected from the group consisting of an x-ray imaging device, a computed tomography imaging device and a magnetic resonance imaging device. 4. The medical imaging system of claim 1 wherein the tracking subsystem is further defined as an infrared tracking sensor. 5. The medical imaging system of claim 1 wherein the imaging device further includes a plurality of tracking targets that are detectable by the tracking subsystem. 6. The medical imaging system of claim 1 further comprises a reference frame marker positioned in a fixed location relative to the patient, where the reference frame marker includes tracking targets that are detectable by the tracking subsystem. 7. The medical imaging system of claim 6 wherein the reference frame marker is physically attached to a portion of the patient. 8. The medical imaging system of claim 1 wherein the motion correction subsystem is further operable to generate an operator alarm upon detection of a change in the relative position between the patient and the imaging device. 9. The medical imaging system of claim 1 further comprising a surgical instrument having tracking targets that are detectable by the tracking subsystem, such that the tracking subsystem detects in real-time instrument position data indicative of the position of the surgical instrument. 10. The medical imaging system of claim 9 wherein the motion correction subsystem is adapted to receive instrument position data from the tracking subsystem and operable to compensate for a change in... | Surgical Navigation Technologies, Inc. | A61B005/05 A61B006/00
US07113825 | B2 | 2006-09-26 | Method and apparatus for detecting acoustic oscillations in cardiac rhythm | 2002-05-03 | A cardiac rhythm management device is configured to detect oscillations in cardiac rhythm by measuring the amplitudes of heart sounds during successive heart beats. Upon detection of acoustic alternans, the device may adjust its operating behavior to compensate for the deleterious effects of the condition. | 1. An implantable cardiac rhythm management device, comprising: an accelerometer for generating a phonocardiogram signal;a controller for processing the phonocardiogram signal to detect heart sounds within the phonocardiogram signal and to measure oscillations in the amplitudes of the heart sounds during successive heart beats; and,wherein the controller is programmed to detect acoustic alternans when the measured oscillations exceed a specified magnitude and to adjust the operating behavior of the device when acoustic alternans is detected. | 1. An implantable cardiac rhythm management device, comprising: an accelerometer for generating a phonocardiogram signal;a controller for processing the phonocardiogram signal to detect heart sounds within the phonocardiogram signal and to measure oscillations in the amplitudes of the heart sounds during successive heart beats; and,wherein the controller is programmed to detect acoustic alternans when the measured oscillations exceed a specified magnitude and to adjust the operating behavior of the device when acoustic alternans is detected. 2. The device of claim 1 wherein the controller is programmed to detect oscillations in the amplitude of the first heart sound S1 in order to detect acoustic alternans. 3. The device of claim 1 wherein the controller is programmed to detect oscillations in the amplitude of the second heart sound S2 in order to detect acoustic alternans. 4. The device of claim 1 further comprising: a sensing/pacing channel; and,wherein the controller is programmed to deliver pacing therapy in accordance with a bradycardia mode, and further wherein the controller is programmed to decrease a programmed pacing rate when acoustic alternans is detected. 5. The device of claim 1 further comprising: a sensing/pacing channel; and,wherein the controller is programmed to deliver anti-tachycardia pacing therapy when a measured heart rate exceeds a specified tachycardia threshold rate, and further wherein the controller is programmed to decrease the tachycardia threshold rate when acoustic alternans is detected. 6. The device of claim 1 further comprising: a first ventricular pacing channel;a second ventricular pacing channel; and,wherein the controller is programmed to deliver pacing therapy in accordance with a bradycardia mode, and further wherein the controller is programmed to initiate ventricular resynchronization pacing by delivering paces through both ventricular pacing channels when acoustic alternans is detected. 7. The device of claim 1 further comprising: a first ventricular pacing channel;a second ventricular pacing channel; and,wherein the controller is programmed to deliver ventricular resynchronization pacing by delivering paces through both ventricular channels, and further programmed to adjust a resynchronization operating parameter when acoustic alternans is detected. 8. The device of claim 7 wherein the resynchronization operating parameter adjusted by the controller is an atrio-ventricular interval. 9. The device of claim 7 wherein the resynchronization operating parameter adjusted by the controller is a biventricular delay interval. 10. The device of claim 1 further comprising: a first atrial pacing channel;a second atrial pacing channel; and,wherein the controller is programmed to deliver pacing therapy in accordance with a bradycardia mode, and further wherein the controller is programmed to initiate atrial resynchronization pacing by delivering paces through both atrial channels when acoustic alternans is... | Cardiac Pacemakers, Inc. | A61N001/39 A61B005/02
US07226422 | B2 | 2007-06-05 | Detection of congestion from monitoring patient response to a recumbent position | 2002-10-09 | Systems and methods obtain information that allows detection of whether a patient suffers from congestion by monitoring a response of the patient to a recumbent position of the patient's body. The patient may be monitored to determine a respiration pattern for a non-recumbent position such as standing and a respiration pattern for a recumbent position such as lying down. The two patterns may be compared, either by a processing device or a physician, to determine a difference in the two respiration patterns. Furthermore, the congestion may be inferred from detecting an amount of time that the patient spends in a recumbent position of from detecting the recumbent angle that the patient obtains, either of which is presumed to indicate whether a recumbent position presents discomfort to the patient due to the presence of congestion. | 1. A method of determining whether a patient is suffering from congestion, comprising: detecting whether the patient is in a recumbent position;when the patient is not in the recumbent position, monitoring the respiration of the patient to detect a first respiration pattern;when the patient is in the recumbent position, monitoring the respiration of the patient to detect a second respiration pattern;comparing the first respiration pattern to the second respiration pattern to detect whether a difference is present between the first respiration pattern and the second respiration pattern; andin response to the difference between the first respiration pattern and the second respiration pattern determining whether the patient is suffering from congestion. | 1. A method of determining whether a patient is suffering from congestion, comprising: detecting whether the patient is in a recumbent position;when the patient is not in the recumbent position, monitoring the respiration of the patient to detect a first respiration pattern;when the patient is in the recumbent position, monitoring the respiration of the patient to detect a second respiration pattern;comparing the first respiration pattern to the second respiration pattern to detect whether a difference is present between the first respiration pattern and the second respiration pattern; andin response to the difference between the first respiration pattern and the second respiration pattern determining whether the patient is suffering from congestion. 2. The method of claim 1, wherein detecting whether the patient is in the recumbent position comprises comparing a reclination of a torso of the patient to a threshold. 3. The method of claim 2, further comprising detecting the reclination of the torso with a tilt switch. 4. The method of claim 2, further comprising detecting the reclination of the torso with an accelerometer. 5. The method of claim 1, wherein monitoring the respiration comprises monitoring trans-thoracic impedance. 6. The method of claim 1, wherein comparing the first respiration pattern to the second respiration pattern to detect whether a difference is present comprises comparing a rate of respiration of the first pattern to a rate of respiration of the second pattern. 7. The method of claim 1, wherein comparing the first respiration pattern to the second respiration pattern to detect whether a difference is present comprises comparing a tidal volume of respiration of the first respiration pattern to a tidal volume of respiration of the second respiration pattern. 8. The method of claim 1, wherein comparing the first respiration pattern to the second respiration pattern to detect whether a difference is present comprises comparing a combination of tidal volume of respiration and respiration rate for the first respiration pattern to a combination of tidal volume of respiration and respiration rate for the second respiration pattern. 9. The method of claim 1, wherein comparing the first respiration pattern to the second respiration pattern to detect whether a difference is present comprises comparing an inspiratory time the first respiration pattern to an inspiratory time for the second respiration pattern and comparing an expiratory time the first respiration pattern to an expiratory time for the second respiration pattern. 10. The method of claim 1, further comprising: detecting over multiple periods whether the patient is in a recumbent position;when the patient is not in the recumbent position during each period, monitoring the respiration of the patient to detect a first respiration pattern for a current period; andwhen the patient is in the recumbent position during each period, monitoring the respiration of the patie... | Cardiac Pacemakers, Inc. | A61B005/05 A61B005/08
US07010337 | B2 | 2006-03-07 | Method and apparatus for monitoring blood condition and cardiopulmonary function | 2002-10-24 | Sensors located on a sensor carrier are placed adjacent one or more of a surgical patient's major thoracic blood-containing structures such as the aorta or pulmonary artery, and characteristics of blood in the blood-containing structures are determined noninvasively by measuring transmission or reflection of light or other types of energy by the blood. Emitters and receptors included in the sensors are connected electrically with suitable electronic signal generating and processing components in a package remote from the sensor carrier. | 1. A method of at least partially evaluating a selected aspect of a patient's metabolic function, wherein said selected aspect is an abnormal shunt of blood, indicating either or both of an anatomic or a physiological deficiency of at least one of said patient's heart and lungs, the method comprising: measuring separately at least one selected characteristic of blood present in each of at least two selected major thoracic blood-containing structures by using at least one electronic sensor located within said patient's thoracic cavity, but outside said selected major thoracic blood-containing structures, to observe said blood through a respective wall of each of said selected major thoracic blood-containing structures without removing any of said blood from said selected major thoracic blood-containing structures, and comparing respective resulting values of said at least one selected characteristic as measured in each of said at least two selected major thoracic blood-containing structures to determine whether an abnormal shunt is present. | 1. A method of at least partially evaluating a selected aspect of a patient's metabolic function, wherein said selected aspect is an abnormal shunt of blood, indicating either or both of an anatomic or a physiological deficiency of at least one of said patient's heart and lungs, the method comprising: measuring separately at least one selected characteristic of blood present in each of at least two selected major thoracic blood-containing structures by using at least one electronic sensor located within said patient's thoracic cavity, but outside said selected major thoracic blood-containing structures, to observe said blood through a respective wall of each of said selected major thoracic blood-containing structures without removing any of said blood from said selected major thoracic blood-containing structures, and comparing respective resulting values of said at least one selected characteristic as measured in each of said at least two selected major thoracic blood-containing structures to determine whether an abnormal shunt is present. 2. A method of at least partially evaluating a selected aspect of a patient's metabolic function, comprising: introducing an identifying agent into said patient's bloodstream to enable an electronic sensor to measure at least one selected characteristic of said patient's blood, and thereafter measuring separately said at least one selected characteristic of blood present in each of at least two selected major thoracic blood-containing structures by using said electronic sensor, located within said patient's thoracic cavity, but outside said selected major thoracic blood-containing structures, to observe said blood through a respective wall of each of said selected major thoracic blood-containing structures without removing any of said blood from said selected major thoracic blood-containing structures, and comparing respective resulting values of said at least one selected characteristic as measured in each of said at least two selected major thoracic blood-containing structures. 3. The method of either claim 1 or claim 2, wherein one of said at least two selected major thoracic blood-containing structures is said patient's aorta. 4. The method of claim 3 wherein another one of said at least two selected major thoracic blood-containing structures is said patient's main pulmonary artery. 5. The method of either claim 1 or claim 2, wherein one of said at least two selected major thoracic blood-containing structures is said patient's main pulmonary artery. 6. The method of either claim 1 or claim 2, wherein said selected characteristic of blood is the degree of saturation by oxygen of hemoglobin in said blood. 7. The method of claim 6, including measuring hemoglobin content as a second one of said at least one selected characteristic of blood. 8. The method of claim 2, wherein said selected characteristic of blood is hemoglobin content of said blood. 9. The method of claim 2, wherein said selected aspect o... | - | A61B005/00
US06978171 | B2 | 2005-12-20 | Automated impedance measurement of an implantable medical device | 2002-03-15 | The present invention automates impedance measurements between pairs of electrodes that are associated with a lead of an implanted device. Apparatus comprises a communications module and a processor, in which the processor communicates with the implanted device over a communications channel. The processor instructs the implanted device to perform a voltage measurement that is associated with a stimulation pulse between the pair of electrodes. The implanted device sends the voltage measurement to the apparatus over the communications channel from which the apparatus determines the impedance between the pair of electrodes. The apparatus initiates an impedance measurement for each pair of electrodes along the lead. With a plurality of leads, the apparatus repeats the impedance measurements for electrode pairs of the other leads. The embodiment also supports measuring a current between pairs of electrodes. The impedance measurements and current measurements are displayed to a clinician on a user interface. | 1. A method for automating impedance measurements associated with an implanted device, the method comprising: (a) configuring a parameter of a stimulation pulse of the implanted device;(b) performing a measurement process with the stimulation pulse in order to determine an impedance value corresponding to at least one electrode comprising: (i) instructing the implanted device to perform a measurement that is associated with the at least one electrode, wherein the at least one electrode is contained on a lead;(ii) calculating the impedance value utilizing the measurement, wherein (i) and (ii) are repeated for each at least one other electrode that is associated with the lead; and(iii) upon completion of (i) and (ii), restoring an original value of the parameter; and(c) automatically processing the impedance value. | 1. A method for automating impedance measurements associated with an implanted device, the method comprising: (a) configuring a parameter of a stimulation pulse of the implanted device;(b) performing a measurement process with the stimulation pulse in order to determine an impedance value corresponding to at least one electrode comprising: (i) instructing the implanted device to perform a measurement that is associated with the at least one electrode, wherein the at least one electrode is contained on a lead;(ii) calculating the impedance value utilizing the measurement, wherein (i) and (ii) are repeated for each at least one other electrode that is associated with the lead; and(iii) upon completion of (i) and (ii), restoring an original value of the parameter; and(c) automatically processing the impedance value. 2. The method of claim 1, wherein the at least one electrode comprises at least eight electrodes. 3. The method of claim 1, wherein (c) comprises: (i) displaying the impedance value on a user interface. 4. The method of claim 1, wherein (c) comprises: (i) storing the impedance value. 5. The method of claim 1, wherein the measurement corresponds to a final voltage value of a stimulation pulse. 6. The method of claim 1, wherein (b) further comprises: (iv) instructing the implanted device to activate a second lead; and(v) repeating (i) and (ii) for each of the at least one electrode that is associated with the second lead. 7. The method of claim 1, further comprising: (d) interrogating the implanted device to determine a device type. 8. The method of claim 1, further comprising: (d) instructing the implanted device to measure a value of current of the at least one electrode; and(e) processing the value of current. 9. The method of claim 8, wherein the at least one electrode comprises at least eight electrodes. 10. The method of claim 1, wherein an amplitude of a stimulation pulse is essentially zero between a repetition of (i) and (ii). 11. A method for automating impedance measurements associated with an implanted device, the method comprising the steps of: (a) configuring a parameter of a stimulation pulse of the implanted device;(b) performing a measurement process with the stimulation pulse in order to determine an impedance value corresponding to a set of electrodes, wherein step (b) comprises the steps of: (i) instructing the implanted device to perform a measurement that is associated with the set of electrodes, wherein the set of electrodes are contained on a lead;(ii) calculating the impedance value utilizing the measurement; and(iii) restoring an original value of the parameter upon completing step (i) and step (ii), wherein step (i), step (ii) and step (iii) are repeated for each set of electrodes that is associated with the lead; and(c) automatically processing the impedance value. 12. Apparatus for automating impedance measurements associated with an implanted device, the apparatus comprising: a communications module... | Medtronic, Inc. | A61B005/00
US06893405 | B2 | 2005-05-17 | Analysis of Sleep Apnea | 2003-01-22 | A non-intrusive and quantitative method and apparatus for diagnosing sleep apnea and detecting apnea events by monitoring during sleep abdominal effort and thoracic effort, determining the phase of each effort, determining the difference in phase between each type of effort, and then determining the rate of phase angle change and standard deviation over time. Also provided may be treatment when apnea events are detected to trigger therapy apparatus such as airway positive pressure apparatus. | 1. A method for diagnosing sleep apnea comprising the steps of measuring thoracic volume of a patient as a function of time and determining the phase of the change of thoracic volume of said patient, measuring abdominal volume of said patient as a function of time and determining the phase of the change of abdominal volume of said patient, calculating the phase angle between said thoracic volume phase and said abdominal volume phase as a function of time, and observing changes of the phase angle over a selected time period to indicate the extent to which the phase angle between the two effort signals fluctuates. | 1. A method for diagnosing sleep apnea comprising the steps of measuring thoracic volume of a patient as a function of time and determining the phase of the change of thoracic volume of said patient, measuring abdominal volume of said patient as a function of time and determining the phase of the change of abdominal volume of said patient, calculating the phase angle between said thoracic volume phase and said abdominal volume phase as a function of time, and observing changes of the phase angle over a selected time period to indicate the extent to which the phase angle between the two effort signals fluctuates. 2. The method according to claim 1 in which the phase angle calculation is made by cross-correlation between two signals created by the measurements, and the distance between two adjacent maxima in the cross correlation estimates the dominant period time with the distance between zero lag and the nearest maximum being a measure of phase lag with phase in degrees estimated by dividing estimated phase lag by the estimated period time. 3. The method according to claim 2 and further calculating the standard deviation of the phase angle over a certain time period to use as a statistical method to for the diagnosis. 4. The method according to claim 2 in which data from the measurements are sub-sampled at predetermined intervals of about 10 Hz, and the results of the phase calculation stored at about 1-second intervals. 5. The method according to claim 1 in which the measurements and calculations are analyzed to determine the standard deviation of the phase angle between the two effort signals over a sample period. 6. The method according to claim 5 in which the variability is measured over substantially the whole sleep period in about 10 second periods with more than about 10 standard deviations summed to calculate a flip-flop state index directly related to upper airway obstruction. 7. The method according to claim 6 in which a flip-flop state index higher than about 0.1 is used as a predictor of obstructive sleep apnea. 8. The method according to claim 1 in which a respiratory distress index by the number of times of arousal from sleep in a predetermined time is correlated with the instability of the phase angles for the diagnosis of apnea and/or hypopnea. 9. The method according to claim 1 in which the abdominal and thoracic phase information is input into a computer with a software utility that calculates phase change and determines significant change or changes in sets of data. 10. Apparatus for diagnosing sleep apnea comprising means for measuring thoracic effort and abdominal effort of a patient as a function of time and determining the phase of the change of each of the efforts, means for calculating the phase angle between the efforts as a function of time, means for plotting the measurements, determinations and calculations, and observing changes of the phase angle over a selected time period to indicate the extent to which t... | Medcare Flaga hf | A61B005/00
US06937887 | B2 | 2005-08-30 | Atrial fibrillation detection method and apparatus | 2002-10-18 | A method and apparatus to detect irregular heart activity and estimate heart rate. In one embodiment, the heart rate estimation apparatus includes a group of six RR storage structures that receives an input of six successive RR interval values. The apparatus also includes a probability engine coupled to the group of RR storage structures, the probability engine is operable to calculate and output a mean value of the six RR interval values and a median value of the mean value and the six RR interval. In one embodiment, the apparatus carries out a method including the acts of determining a plurality of interval values between successive R waves; determining a mean of a group of six interval values; determining a median interval value for a group including the group of six interval values and the mean of the group of six interval values; and determining a running average of interval values between successive R waves using the median interval value. | 1. A computer implemented method of estimating a heart rate, the method comprising: determining a plurality of interval values between successive R waves; determining a mean of a group of six interval values; determining a median interval value for a group including the group of six interval values and the mean of the group of six interval values; and determining a running average of interval values between successive R waves using the median interval value. | 1. A computer implemented method of estimating a heart rate, the method comprising: determining a plurality of interval values between successive R waves; determining a mean of a group of six interval values; determining a median interval value for a group including the group of six interval values and the mean of the group of six interval values; and determining a running average of interval values between successive R waves using the median interval value. 2. A method as claimed in claim 1, wherein the group of six interval values are successive in time. 3. A method as claimed in claim 2, further comprising: maintaining at least six interval values, one each in a memory structure; shifting interval values in the memory structures, such that the oldest interval value is pushed out of its associated memory structure; and adding a new interval value to create a new group of six interval values. 4. A heart rate estimation apparatus, comprising: a group of six RR storage structures that receives an input of six successive RR interval values; and a probability engine coupled to the group of RR storage structures, the probability engine operable to calculate and output a mean value of the six RR interval values and a median value of the mean value and the six RR interval values. 5. An apparatus as claimed in claim 4, wherein the probability engine is operable to shift the RR interval values in the six RR storage structures to push out the oldest RR interval value and to make room for a new RR interval. 6. A heart rate estimation apparatus, comprising: means for receiving a plurality of interval values between successive R waves; means for determining a mean of a group of six interval values; means for determining a median interval value for a group including the group of six interval values and the mean of the group of six interval values; and means for determining a running average of interval values between successive R waves using the median interval value. 7. A heart rate estimation apparatus as claimed in claim 6, further comprising: means for shifting the group of six interval values. | GE Medical Systems Information Technologies, Inc. | A61B005/0468
US06892093 | B2 | 2005-05-10 | Method and apparatus for real time display of filtered electrocardiogram data | 2002-08-01 | A method for displaying waveform data on a display device is disclosed. In an exemplary embodiment, the method includes apportioning a display region into a first portion and a second portion immediately adjacent to the first portion. The first portion is used to display a first segment of the waveform data including the most recently received data extending back to a determined delay period. The second portion is used to display a second segment of the waveform data, the second segment including the remainder of the waveform data. The data displayed in the first portion has a continuously varying amplitude level adjustment applied thereto for partial baseline correction thereof, while the data displayed in the second portion has a corrected baseline amplitude adjustment with no further amplitude level adjustment applied thereto. | 1. A method for displaying waveform data on a display device, the method comprising: apportioning a display region into a first portion and a second portion immediately adjacent to said first portion; said first portion being used to display a first segment of the waveform data, said first segment comprising the most recently received data extending back to a determined delay period; said second portion being used to display a second segment of the waveform data, said second segment comprising the remainder of the waveform data; wherein the data displayed in said first portion has a continuously varying amplitude level adjustment applied thereto for partial baseline correction thereof, and the data displayed in said second portion has a corrected baseline amplitude adjustment with no further amplitude level adjustment applied thereto. | 1. A method for displaying waveform data on a display device, the method comprising: apportioning a display region into a first portion and a second portion immediately adjacent to said first portion; said first portion being used to display a first segment of the waveform data, said first segment comprising the most recently received data extending back to a determined delay period; said second portion being used to display a second segment of the waveform data, said second segment comprising the remainder of the waveform data; wherein the data displayed in said first portion has a continuously varying amplitude level adjustment applied thereto for partial baseline correction thereof, and the data displayed in said second portion has a corrected baseline amplitude adjustment with no further amplitude level adjustment applied thereto. 2. The method of claim 1, wherein said baseline adjustment is applied through a symmetrical finite impulse response filter. 3. The method of claim 2, wherein said symmetrical finite impulse response filter is centered about a given data sample at a time delay of about 0.5 to about 3.0 seconds. 4. The method of claim 1, wherein: the waveform data is scrolled across said display region such that the waveform data initially appears in said first portion, thereafter scrolling through said first portion while being subjected to said baseline amplitude adjustment, and thereafter appearing in and scrolling through said second portion with no further baseline adjustment being applied thereto. 5. The method of claim 1, wherein: a wiper bar is scrolled across said display region such that said first portion of said display region adjacently follows said wiper bar along an edge thereof. 6. A method of filtering and displaying sequential waveform data samples, the method comprising: shifting a sequence of stored uncorrected data samples; receiving and storing a new uncorrected data sample; computing a baseline estimate correction using said stored uncorrected data samples and said new uncorrected data sample; shifting a sequence of stored corrected data samples and thereafter determining a new corrected data sample, wherein said new corrected data sample is determined by applying said baseline estimate correction to a specific one of said stored uncorrected data samples; creating a sequence of temporary display data samples by applying said baseline correction to each of said stored uncorrected data samples that were stored subsequent to said specific one of said stored uncorrected data samples, as well as to said new uncorrected data sample; and displaying said sequence of corrected data samples, said new corrected data sample, and said sequence of temporary display data samples. 7. The method of claim 6, wherein said baseline estimate correction is implemented by a symmetrical finite impulse response filter. 8. The method of claim 7, wherein said specific one of said stored uncorrected data samples is centrally loc... | GE Medical Systems Information Technologies Inc. | A61B005/04
US07189204 | B2 | 2007-03-13 | Sleep detection using an adjustable threshold | 2002-12-04 | Devices and methods for sleep detection involve the use of an adjustable threshold for detecting sleep onset and termination. A method for detecting sleep includes adjusting a sleep threshold associated with a first sleep-related signal using a second sleep-related signal. The first sleep-related signal is compared to the adjusted threshold and sleep is detected based on the comparison. The sleep-related signals may be derived from implantable or external sensors. Additional sleep-related signals may be used to confirm the sleep condition. A sleep detector device implementing a sleep detection method may be a component of an implantable pulse generator such as a pacemaker or defibrillator. | 1. A method of detecting sleep, comprising: adjusting a sleep threshold associated with a first sleep-related signal using a second sleep-related signal;comparing the first sleep-related signal to the adjusted threshold; anddetecting sleep based on the comparison, wherein at least one of adjusting, comparing or detecting steps is performed at least in part implantably. | 1. A method of detecting sleep, comprising: adjusting a sleep threshold associated with a first sleep-related signal using a second sleep-related signal;comparing the first sleep-related signal to the adjusted threshold; anddetecting sleep based on the comparison, wherein at least one of adjusting, comparing or detecting steps is performed at least in part implantably. 2. The method of claim 1, wherein at least one of the first or the second sleep-related signals is a physiological signal. 3. The method of claim 1, further comprising establishing the sleep threshold associated with the first sleep-related signal. 4. The method of claim 3, wherein the sleep threshold is established based on clinical data. 5. The method of claim 3, wherein the sleep threshold is established using the first sleep-related signal. 6. The method of claim 3, wherein the sleep threshold is established using data associated with historical sleep times. 7. The method of claim 1, wherein detecting sleep based on the comparison comprises: detecting sleep onset if the first sleep-related signal falls below the adjusted threshold; anddetecting sleep termination if the first sleep-related signal rises above the adjusted threshold. 8. The method of claim 1, wherein the first sleep-related signal is associated with patient activity. 9. The method of claim 1, wherein the first sleep-related signal is an accelerometer signal. 10. The method of claim 1, wherein the first sleep-related signal is sensed using an implantable sensor. 11. The method of claim 1, wherein the first sleep-related signal is sensed using an external sensor. 12. The method of claim 1, wherein the second sleep-related signal is associated with patient respiration. 13. The method of claim 1, wherein the second sleep-related signal comprises transthoracic impedance. 14. The method of claim 1, wherein the second sleep-related comprises minute ventilation. 15. The method of claim 1, wherein the second sleep-related signal comprises tidal volume. 16. The method of claim 1, wherein the second sleep-related signal is associated with cardiac function. 17. The method of claim 1, wherein the second sleep-related signal comprises a QT interval signal. 18. The method of claim 1, wherein the second sleep-related signal is sensed using an implantable sensor. 19. The method of claim 1, wherein the second sleep-related signal is sensed using an external sensor. 20. The method of claim 1, wherein detecting sleep comprises detecting sleep using an implantable device. 21. The method of claim 20, wherein the implantable device is an implantable pulse generator. 22. The method of claim 1, wherein detecting sleep comprises detecting sleep using an external device. 23. The method of claim 1, wherein detecting sleep further comprises confirming sleep using a third sleep-related signal. 24. The method of claim 1, wherein detecting sleep further comprises: comparing a third sleep-related signal to a predetermined v... | Cardiac Pacemakers, Inc. | A61B005/02 A61B005/00 A61B005/08
US06983184 | B2 | 2006-01-03 | Interactive-modified interactive event related potential (IMIERP) | 2002-05-13 | A method modifying a particular electrophysiological feature generated in response to a stimulus is disclosed. The method involves a repeated process of: sampling the brain wave state in order to apply a modifying pulse; modifying the brain wave state to a state more conducive to a required response; sampling the brain wave state in order to apply a response stimulus; applying a stimulus only when the brain wave state substantially meets the preselected response criteria; recording the brain wave activity of the subject subsequent to the application of the stimulus. | 1. A process (IMIERP process) for modifying the features of electrophysiological activity in response to a stimulus, in the brain of a subject, by interactively applying a brain-wave-state modifying stimulus, said process including the steps of: (a) initially monitoring the subject's brain-wave-state; (b) comparing the monitored brain-wave-state with pre-selected modifying criteria, associated with a requisite (post-modification) brain-wave-state; (c) applying a brain-wave-state modifying pulse to the subject when the brain wave state substantially meets the pre-selected criteria, otherwise returning to step (a); (d) monitoring the subjects modified brain-wave-state; (e) comparing the monitored brain wave state with pre-selected stimulus-response criteria, associated with a requisite (post-stimulus) response; and (f) applying a selected response-inducing stimulus to the subject when the brain wave state substantially meets the pre-selected response criteria. | 1. A process (IMIERP process) for modifying the features of electrophysiological activity in response to a stimulus, in the brain of a subject, by interactively applying a brain-wave-state modifying stimulus, said process including the steps of: (a) initially monitoring the subject's brain-wave-state; (b) comparing the monitored brain-wave-state with pre-selected modifying criteria, associated with a requisite (post-modification) brain-wave-state; (c) applying a brain-wave-state modifying pulse to the subject when the brain wave state substantially meets the pre-selected criteria, otherwise returning to step (a); (d) monitoring the subjects modified brain-wave-state; (e) comparing the monitored brain wave state with pre-selected stimulus-response criteria, associated with a requisite (post-stimulus) response; and (f) applying a selected response-inducing stimulus to the subject when the brain wave state substantially meets the pre-selected response criteria. 2. The process of claim 1, wherein the steps (a) to (e) are carried out repetitively before step (f). 3. A method for acquiring event related data, representative of physiological activity in the brain of a subject, sad method including the steps of: (a) applying the process of claim 1; (b) recording the brain wave activity of the subject subsequent to the application of the stimulus; (c) processing the recording of brain wave activity to extract the event related data for further analysis. 4. The method for acquiring data of claim 3 wherein the steps (a) and (b) are carried out repetitively before step (c). 5. The method for acquiring data of claim 3 wherein the event related data representative of physiological activity is detected by means other than event related potentials (ERP or ERF) selected from the group consisting of: (a) reaction time data; (b) questionaire data; (c) other physiological parameters. 6. A method for diagnosis of an illness (said illness being based on current diagnostic features), said illness being associated with particular electrophysiological activity in response to a stimulus, by comparing the electrophysiological activity in response to a stimulus obtained using the process of claim 1, with electrophysiological activity in response to a stimulus obtained using other stimulation processes. 7. The process of claim 1, wherein the brain-wave-state modifying pulse is one of Transcranial Magnetic Stimulation, Transcranial Electrical Stimulation, or sensory stimulus. | Gregory W Price | A61B005/0484
US07175588 | B2 | 2007-02-13 | Method and system for detecting ventricular collapse | 2003-01-08 | A pump system includes an implantable pump including a motor having a rotor and a stator. The stator includes a plurality of stator windings, and a motor controller is coupled to the motor to energize the windings so as cause the rotor to turn. A time-based system parameter of the pump is sampled and the system parameter is analyzed to calculate a suction probability index that provides an indication of the imminence of ventricle collapse. | 1. A method of detecting ventricular collapse in a patient having a blood pump implanted, the method comprising: sampling a time-based system parameter of the blood pump;analyzing the time-based system parameter; andcalculating a suction probability index in response to the analysis of the time-based system parameter that provides an indication of the imminence of ventricle collapse. | 1. A method of detecting ventricular collapse in a patient having a blood pump implanted, the method comprising: sampling a time-based system parameter of the blood pump;analyzing the time-based system parameter; andcalculating a suction probability index in response to the analysis of the time-based system parameter that provides an indication of the imminence of ventricle collapse. 2. The method of claim 1, wherein sampling a time-based system parameter of the blood pump includes sampling the pump current. 3. The method of claim 1, wherein sampling a time-based system parameter of the blood pump includes sampling the pump speed. 4. The method of claim 1, wherein the implantable pump includes a flow sensing device providing an indication of the flow rate through the pump, and wherein sampling a time-based system parameter of the blood pump includes sampling the flow rate. 5. The method of claim 1, further comprising converting the sampled time-based parameter to a digital signal. 6. The method of claim 1, further comprising computing the spectral content of the sampled time-based signal. 7. The method of claim 6, further comprising applying the computed spectral content to a spectral analysis equation to calculate the suction probability index. 8. The method of claim 7, wherein analyzing the time-based system parameter includes computing the total distortion of the time-based system parameter waveform. 9. The method of claim 8, wherein computing the spectral content includes applying a Fourier Transform, and wherein the suction probability index is calculated according tosuction probability index=[∑n=1x⁢[A⁡[f(n·ⅆF)]]2-[A⁡[f1]]2]·100A⁡[f1]wherein A represents amplitude, n indicates the position of the spectral component in the array resulting from the Fourier Transform; x is the last position in the array; dF represents the frequency resolution/interval of the Fourier Transform in Hertz; and f1 is the fundamental frequency, maximum (amplitude) spectral peak in the array. 10. The method of claim 7, wherein analyzing the time-based system parameter includes computing the harmonic distortion of the time-based system parameter waveform. 11. The method of claim 10, wherein computing the spectral content includes applying a Fourier Transform, and wherein the suction probability index is calculated according tosuction⁢⁢probability⁢⁢index=[∑n=2x⁢⁢[A⁡[fn]]2]·100A⁡[f1]wherein A represents amplitude, n indicates the nth harmonic in the array resulting from the Fourier Transform; x is the last position in the array; f1 is the fundamental frequency, the maximum (amplitude) spectral peak in the FFT resultant array; and fn represents integer multiples of the fundamental f1 from n=2 (second harmonic) to x (xth harmonic). 12. The method of claim 7, wherein analyzing the time-based system parameter includes computing the distortion of the time-based system parameter waveform below the fundamental frequency. 13. The method of claim 12, wherein com... | Micromed Technology, Inc. | A61N001/362 A61B005/02 A61M001/10 A61M037/00 A61M031/00
US07013169 | B2 | 2006-03-14 | Dual steer preshaped catheter | 2003-01-27 | An electrophysiology catheter includes a flexible shaft with one or more electrodes disposed at a distal end having a preformed curve. A first steering apparatus allows altering a curvature of the distal end and a second steering apparatus allows steering of the distal end without significantly altering the curvature of the distal end. The first and second steering apparatuses are used for maneuvering the catheter and providing positive contact between the distal end and the tissue targeted for treatment. | 1. A catheter adapted for electrophysiological therapy, comprising: a flexible shaft comprising: a preformed curve on a distal end of the flexible shaft;one or more electrodes disposed along the distal end of the flexible shaft; anda shape deflection area defined as a region on the flexible shaft between the preformed curve and a generally straight proximal portion of the flexible shaft;a first anchor member attached to a distal portion of the pre-formed curve; anda second anchor member attached to the flexible shaft distal to the shape deflection area of the flexible shaft;a first and second steering tendon coupled to the first and second anchor members, respectively, so that a force applied to a proximal end of the first steering tendon causes a change in a curvature of the preformed curve and a force applied to a proximal end of the second steering tendon causes movement of the distal end about the shape deflection area of the flexible shaft. | 1. A catheter adapted for electrophysiological therapy, comprising: a flexible shaft comprising: a preformed curve on a distal end of the flexible shaft;one or more electrodes disposed along the distal end of the flexible shaft; anda shape deflection area defined as a region on the flexible shaft between the preformed curve and a generally straight proximal portion of the flexible shaft;a first anchor member attached to a distal portion of the pre-formed curve; anda second anchor member attached to the flexible shaft distal to the shape deflection area of the flexible shaft;a first and second steering tendon coupled to the first and second anchor members, respectively, so that a force applied to a proximal end of the first steering tendon causes a change in a curvature of the preformed curve and a force applied to a proximal end of the second steering tendon causes movement of the distal end about the shape deflection area of the flexible shaft. 2. The catheter of claim 1, further comprising a stylet disposed within the distal end of the flexible shaft, the stylet substantially shaping the preformed curve on the flexible shaft. 3. The catheter of claim 2, wherein the stylet is a resilient and shape retentive member. 4. The catheter of claim 2, wherein the stylet comprises a nitinol wire. 5. The catheter of claim 1, wherein the preformed curve comprises a circular loop. 6. The catheter of claim 1, wherein the one or more electrodes comprise a plurality of band electrodes. 7. A catheter adapted for electrophysiological therapy, comprising: a flexible shaft comprising: a preformed curve on a distal end of the flexible shaft;an electrical energy delivery means disposed along the distal end of the flexible shaft; anda shape deflection area defined as a region on the shaft between the preformed curve and a generally straight proximal portion of the flexible shaft;a first steering means for changing a curvature of the preformed curve of the flexible shaft;a second steering means for changing an orientation of the distal end of the flexible shaft about the shape deflection area without substantially changing the curvature of the preformed curve of the flexible shaft. 8. The catheter of claim 7, further comprising a shape retentive means disposed within the distal end of the flexible shaft for shaping the preformed curve. 9. The catheter of claim 8, wherein the shape retentive means comprises a nitinol wire. 10. The catheter of claim 7, wherein the preformed curve comprises a circular loop. 11. The catheter of claim 7, wherein the electrical energy delivery means comprises one or more electrodes. 12. The catheter of claim 11, wherein the one or more electrodes comprise a plurality of band electrodes. 13. The catheter of claim 7, wherein the first steering means comprises a steering tendon attached to an anchor member at a distal portion of the pre-formed curve. 14. The catheter of claim 7, wherein the second steering means comprises a steerin... | Cardiac Pacemakers, Inc. | A61B005/04 A61B018/14 A61N001/05
US07016720 | B2 | 2006-03-21 | System and method for monitoring blood glucose levels using an implantable medical device | 2002-10-21 | T-wave amplitude and QT interval are derived from patient cardiac signals. Then blood glucose levels are determined based on a combination of the T-wave amplitude and the QT interval. By using a combination of both T-wave-based and QT interval-based signals, blood glucose levels can be reliably detected throughout a wide range of blood glucose levels. Once the blood glucose level has been detected, the implanted device compares the blood glucose level against upper and lower acceptable bounds and appropriate warning signals are generated if the level falls outside the bounds. In one example, wherein an implantable insulin pump is additionally provided, the pump is controlled based on the detected blood glucose level to maintain glucose levels within an acceptable range. A calibration technique is also provided for determining patient-specific parameters for use in the detection of blood glucose levels. | 1. A method performed by an implantable medical device, the method comprising: detecting T-wave amplitude based on cardiac signals;detecting QT interval based on cardiac signals; andderiving a blood glucose level based on a combination of the detected T-wave amplitude and the detected QT interval. | 1. A method performed by an implantable medical device, the method comprising: detecting T-wave amplitude based on cardiac signals;detecting QT interval based on cardiac signals; andderiving a blood glucose level based on a combination of the detected T-wave amplitude and the detected QT interval. 2. The method of claim 1 further including converting T-wave amplitude into a T-wave amplitude fraction and converting QT interval into a corrected QT (QTc) interval delta and wherein deriving blood glucose levels includes determining the blood glucose level based on a combination of the T-wave amplitude fraction and the QTc delta. 3. The method of claim 2 wherein determining the blood glucose level includes calculating: blood glucose level=A−B*(QTc delta)−C*(T-wave amplitude fraction)wherein A, B and C are predetermined coefficients. 4. The method of claim 3 wherein: A is in the range of 250–270;B is in the range of 1.25–1.50; andC is in the range of 160–180. 5. The method of claim 2 wherein converting T-wave amplitude into a T-wave amplitude fraction includes: inputting a T-wave baseline value; andconverting T-wave amplitude into a T-wave amplitude fraction using the T-wave baseline value. 6. The method of claim 2 wherein converting QT interval into QTc interval includes: inputting a QT interval baseline value;measuring heart rate based on cardiac electrical signals;converting QT interval into QTc interval based on the heart rate; andconverting the QTc interval into a QT interval delta using the QT baseline value. 7. The method of claim 1 further including: comparing the blood glucose level against an upper bound and, if the blood glucose levels exceeds the upper bound, generating a warning signal indicating the blood glucose level is too high. 8. The method of claim 7 wherein the upper bound is 120 mg/dl. 9. The method of claim 1 further including: comparing the blood glucose level against a lower bound and, if the blood glucose levels falls below the lower bound, generating a warning signal indicating the blood glucose level is too low. 10. The method of claim 9 wherein the lower bound is 60 mg/dl. 11. The method of claim 1 wherein the implantable device includes an insulin pump and wherein the method includes controlling the insulin pump based on the blood glucose level. 12. The method of claim 1 wherein the implanted device performs functions and wherein the method includes controlling the functions of the implanted device based on the blood glucose level. 13. The method of claim 1 further including recording the blood glucose level for subsequent transmission to an external device. 14. An implantable glucose monitor comprising: means for detecting cardiac signals;means for detecting T-wave amplitude based on the cardiac signals;means for detecting QT interval based on the cardiac signals; andmeans for deriving blood glucose levels based on a combination of the detected T-wave amplitude and the detected QT interval. 15. An implantabl... | Pacesetter, Inc. | A61B005/0402
US07047059 | B2 | 2006-05-16 | Simplified water-bag technique for magnetic susceptibility measurements on the human body and other specimens | 2001-03-27 | A probe instrument using room-temperature sensor(s) that can measure variations in magnetic susceptibilities. The instrument has sufficient resolution to monitor paramagnetic materials in a human body, such as iron in a human liver, by noninvasively examining patients with iron-overload diseases. The instrument includes room temperature magnetic sensors, and detects the sample, that is, the tissue response to an alternating current field applied by an applied field coil. The applied field coil dimensions are chosen so that the applied field is optimized for maximum response from the liver while minimizing the effects due to the overlying abdominal tissue and at the same time not unduly increasing the sensitivity of the instrument to the lung. To overcome variations in the sensor output due to fluctuations in the applied field, change in the ambient temperature and mechanical relaxation of the instrument, the sensor-sample distance is modulated. The detector assembly is oscillated while the examined patient remains stationary. An improved water-bag technique is employed to eliminate background tissue response. The detector assembly forms part of a probe instrument for performing noninvasively the paramagnetic concentration of a patient. | 1. A method for noninvasive determination of magnetic susceptibility variation in a patient by measuring magnetic susceptibilities of selected tissue of the patient, the method comprising: providing an instrument which includes at least one magnetic sensor, an applied field coil and a current source connected to the applied field coil and means for processing sensed signals from the at least one magnetic sensor;positioning on the patient a flexible bag substantially filled with deformable material having a magnetic susceptibility substantially similar to that of body tissue, said bag being attached to a substantially rigid barrier, the barrier being spaced a predetermined distance from the patient by the material-filled bag, which predetermined distance remains substantially constant during a measuring sequence, the amount of deformable material in the flexible bag remaining constant during a measuring sequence, the instrument being moved with respect to the barrier during the measuring sequence;positioning the instrument external to the patient in proximity to the tissue of interest and adjacent the barrier;supplying the applied field coil with current thereby applying a magnetic field to the tissue of interest;sensing a response from the tissue of interest wit the instrument; andoutputting data corresponding to the magnetic susceptibility variation in the tissue. | 1. A method for noninvasive determination of magnetic susceptibility variation in a patient by measuring magnetic susceptibilities of selected tissue of the patient, the method comprising: providing an instrument which includes at least one magnetic sensor, an applied field coil and a current source connected to the applied field coil and means for processing sensed signals from the at least one magnetic sensor;positioning on the patient a flexible bag substantially filled with deformable material having a magnetic susceptibility substantially similar to that of body tissue, said bag being attached to a substantially rigid barrier, the barrier being spaced a predetermined distance from the patient by the material-filled bag, which predetermined distance remains substantially constant during a measuring sequence, the amount of deformable material in the flexible bag remaining constant during a measuring sequence, the instrument being moved with respect to the barrier during the measuring sequence;positioning the instrument external to the patient in proximity to the tissue of interest and adjacent the barrier;supplying the applied field coil with current thereby applying a magnetic field to the tissue of interest;sensing a response from the tissue of interest wit the instrument; andoutputting data corresponding to the magnetic susceptibility variation in the tissue. 2. The method recited in claim 1, wherein an alternating current is supplied to the applied field coil. 3. The method recited in claim 1, wherein the method further comprises providing the instrument with displacement means for displacing the magnetic sensor and the applied field coil simultaneously thereby compensating for noise introduced to the sensed signals. 4. The method recited in claim 3, wherein the displacement means operates between about one to six inches. 5. The method recited in claim 3, wherein the displacement means operates between about 0.5 to 10.0 hertz. 6. The method recited in claim 1, wherein the outputting of data corresponding to the magnetic susceptibility variation in the human body comprises concentrations of paramagnetic material in the tissue of interest. 7. The method recited in claim 6, wherein the paramagnetic material is iron and the tissue of interest is a liver. 8. The method recited in claim 7, wherein the outputting of data corresponds to concentrations of iron in the liver and the resolution of the measurements corresponds to about 30 micrograms per milliliter. 9. The method recited in claim 1, wherein the positioning step functionally replaces the irregular or variable shape of the patient's body with a volume of material similar in magnetic susceptibility to body tissue, whose surface has a constant shape defined by the rigid barrier. 10. An apparatus for noninvasively measuring magnetic susceptibility variations in the body tissue of a patient to determine a compositional state in the body, the apparatus comprising: a detector assembly ... | Quantum Magnetics, Inc | A61B005/05
US07257438 | B2 | 2007-08-14 | Patient-worn medical monitoring device | 2003-01-14 | A medical monitor includes a lanyard and an electronic package supported in the manner of a pendant. The lanyard includes integral electrodes or other sensors for making physiological measurements, auxiliary components and connectors for electrically connecting the electrodes or sensors to the electronic package. The physiological measurements may be stored in the monitor for later readout, or may be transmitted, before or after processing, to a remote location. | 1. A monitoring device, comprising: a lanyard;a power source disposed on the lanyard;one or more physiological sensors attached to the lanyard; andan electronics package removably connectable to the lanyard, whereby power is supplied from the power source to the electronics package upon connection of the electronics package to the lanyard. | 1. A monitoring device, comprising: a lanyard;a power source disposed on the lanyard;one or more physiological sensors attached to the lanyard; andan electronics package removably connectable to the lanyard, whereby power is supplied from the power source to the electronics package upon connection of the electronics package to the lanyard. 2. The monitoring device as claimed in claim 1, wherein the lanyard has a pair of free ends, and the power source includes a battery positioned in one of the free ends. 3. A monitoring device, comprising: a lanyard having a pair of free ends;one or more physiological sensors connected to the lanyard; andan electronics package connectable to the free ends of the lanyard in an articulating fashion with one free end of the lanyard directly connected to the electronics package and the other free end of the lanyard directly connected to the electronics package so that, when the electronics package is hung from the lanyard, the electronics package and the lanyard form a V shape having a vertex, with the electronics package at the vertex of the V-shape. 4. A monitoring device, comprising: a lanyard;one or more physiological sensors connected to the lanyard;an electronics package removably connectable to the lanyard; andat least one conductive pathway integrated with the lanyard so that upon connection of the electronics package to the lanyard, an electrical connection is made between the sensors and the electronics package. 5. The monitoring device as claimed in claims 1 or 3 or 4, wherein the sensors comprise two or more electrocardiogram electrodes. 6. The monitoring device as claimed in claims 1 or 3 or 4, wherein the sensors include a temperature sensor. 7. The monitoring device as claimed in claims 1 or 3 or 4, wherein the sensors include a pulse oximetry sensor. 8. The monitoring device as claimed in claims 1 or 3 or 4, wherein the lanyard at least partially comprises a flat ribbon. 9. The monitoring device as claimed in claims 1 or 3 or 4, wherein the lanyard at least partially comprises a cord. 10. The monitoring device as claimed in claim 4, wherein the conductive pathway includes conductors embedded between layers of flat ribbon. 11. The monitoring device as claimed in claims 1 or 3 or 4, wherein the lanyard comprises a substantially elongated material having a longitudinal axis and one or more flaps projecting at an angle from the longitudinal axis, the sensors being directly connected to the flaps. 12. The monitoring device as claimed in claims 1 or 3 or 4, wherein the lanyard includes length adjustment means for adjusting the length of the lanyard. 13. The monitoring device as claimed in claims 1 or 3 or 4, wherein the lanyard includes means for enhancing the contact adhesion of the lanyard to a patient. 14. The monitoring device as claimed in claims 1 or 4, wherein the lanyard includes a pair of free ends and a connector, the free ends being connected to the electronics package, and the conn... | Datascope Investment Corp. | A61B005/00
US07171262 | B2 | 2007-01-30 | Vital sign display monitor | 2003-01-28 | At least an electroencephalogram (EEG) detected by a sensor is input from a signal input section as a vital sign signal. The vital sign signal is stored in a storage as vital sign data. A display includes a first display area, a second display area and a third display area. A controller reads out the vital sign data from the storage to display the vital sign signal on the display. The controller includes a first display processor, which displays the EEG in a real time manner, on the first display area, a second display processor, which displays the EEG at a time point in past, on the second display area, and a third display processor, which displays a time-varying trend of a parameter of the EEG, on the third display area. A designator designates a time point in past in the trend displayed in the third display area, so that the EEG at the designated time point is displayed on the second display area. | 1. A vital sign display monitor, comprising: a signal input section, from which at least an electroencephalogram (EEG) detected by a sensor is input as a vital sign signal;a storage, in which the vital sign signal is stored as vital sign data;a display, simultaneously displaying a first display area, a second display area and a third display area;a controller, which reads out the vital sign data from the storage to display the vital sign signal on the display, the controller including:a first display processor, which displays the EEG in a real time manner, on the first display area;a second display processor, which displays the EEG at a time point in the past, on the second display area; anda third display processor, which displays a time-varying trend of a parameter of the EEG, on the third display area; anda designator, which designates a time point in the past in the trend displayed in the third display area, so that the EEG at the designated time point is displayed on the second display area. | 1. A vital sign display monitor, comprising: a signal input section, from which at least an electroencephalogram (EEG) detected by a sensor is input as a vital sign signal;a storage, in which the vital sign signal is stored as vital sign data;a display, simultaneously displaying a first display area, a second display area and a third display area;a controller, which reads out the vital sign data from the storage to display the vital sign signal on the display, the controller including:a first display processor, which displays the EEG in a real time manner, on the first display area;a second display processor, which displays the EEG at a time point in the past, on the second display area; anda third display processor, which displays a time-varying trend of a parameter of the EEG, on the third display area; anda designator, which designates a time point in the past in the trend displayed in the third display area, so that the EEG at the designated time point is displayed on the second display area. 2. The vital sign display monitor as set forth in claim 1, wherein the controller includes a filter which eliminates noise from the electroencephalogram to be displayed on at least one of the first display area and the second display area. 3. The vital sign display monitor as set forth in claim 1, wherein the controller includes a filter which eliminates noise from the trend to be displayed on the third display area. 4. The vital sign display monitor as set forth in claim 3, wherein the noise is stemmed from a cardiac potential. 5. A vital sign display monitor, comprising: a signal input section, from which at least an electroencephalogram (EEG) detected by a sensor is input as a vital sign signal;a storage, in which the vital sign signal is stored as vital sign data;a display, simultaneously displaying a first display area, a second display area and a third display area;a controller, which reads out the vital sign data from the storage to display the vital sign signal on the display, the controller including:a first display processor, which displays the EEG in a real time manner, on the first display area;a second display processor, which displays the EEG at a time point in the past, on the second display area; anda third display processor, which displays a time-varying trend of an amplitude of the EEG, on the third display area; anda designator, which designates a time point in the past in the trend displayed in the third display area, so that the EEG at the designated time point is displayed on the second display area. 6. The vital sign display monitor as set forth in claim 5, wherein the controller includes a filter which eliminates noise from the trend to be displayed on the third display area. 7. The vital sign display monitor as set forth in claim 5, wherein the controller includes a filter which eliminates noise from the electroencephalogram to be displayed on at least one of the first display area and the second display area. 8. The vital... | Nihon Kohden Corporation | A61B005/04
US07308298 | B2 | 2007-12-11 | Magnetic resonance imaging using MT pulse of which duration is shorter | 2001-12-21 | A magnetic resonance imaging technique provides adequate MT effects by applying a shortened MT pulse and scan time. An MT pulse whose frequency is off-resonance to a region to be imaged of an object is first applied to he object, a spoiler pulse is applied to the object, and an echo signal is acquired from the region to be imaged. The duration of the MT pulse is set to a relatively shorter time than conventional. Preferably, the duration is less than 10 ms. The area of the MT pulse waveform is set to a specified value equivalent to the conventional. The MT pulse is preferably applicable to a two-dimensional scan using the multislice technique but can also be used in a three-dimensional scan. | 1. A magnetic resonance imaging system comprising: MT-pulse applying means for applying to an object an MT (magnetization transfer) pulse formed as RF (radio frequency) pulse of (i) a frequency which is off-resonance at a region to be imaged, (ii) duration less than 10 ms, and (iii) a waveform based on either a Sinc function or a Gaussian function;gradient applying means for applying, concurrently with application of the MT pulse, a gradient pulse to select an applied position of the MT pulse to the object so that the applied position of the MT pulse is different from the region to be imaged;spoiler applying means for applying a gradient spoiler pulse to the object after applying both the MT pulse and the gradient pulse to the region to be imaged; andscanning means for performing a scan to acquire an echo signal from the region to be imaged after applying the gradient spoiler pulse. | 1. A magnetic resonance imaging system comprising: MT-pulse applying means for applying to an object an MT (magnetization transfer) pulse formed as RF (radio frequency) pulse of (i) a frequency which is off-resonance at a region to be imaged, (ii) duration less than 10 ms, and (iii) a waveform based on either a Sinc function or a Gaussian function;gradient applying means for applying, concurrently with application of the MT pulse, a gradient pulse to select an applied position of the MT pulse to the object so that the applied position of the MT pulse is different from the region to be imaged;spoiler applying means for applying a gradient spoiler pulse to the object after applying both the MT pulse and the gradient pulse to the region to be imaged; andscanning means for performing a scan to acquire an echo signal from the region to be imaged after applying the gradient spoiler pulse. 2. The magnetic resonance imaging system according to claim 1 wherein the duration of the MT pulse is set to a period of time during which relaxation of a spin-lattice magnetization of magnetic spins in the object is substantially not completed. 3. The magnetic resonance imaging system according to claim 2, wherein the scan is a two-dimensional scan performed based on a multi-slice imaging technique. 4. The magnetic resonance imaging system according to claim 2, wherein the scan is a three-dimensional scan performed based on a multi-slice imaging technique. 5. A magnetic resonance imaging method for acquiring an echo signal on the basis of magnetic resonance phenomena of at least two types of nuclear pools in an object, the two types of nuclear pools being mutually coupled through a coupling relationship based on at least one of a chemical exchange phenomenon and a cross relaxation phenomenon, the method comprising: decoupling the coupling relationship between the at least two types of nuclear pools by applying to the object an MT (magnetization transfer) pulse formed as an RF (radio frequency) pulse of (i) a frequency which is off-resonance at a region to be imaged, (ii) duration less than 10 ms, and (iii) a waveform based on either a Sinc function or a Gaussian function;applying a gradient spoiler pulse to the decoupled nuclear pools applying, concurrently with application of MT pulse, a gradient pulse so as to select an applied position of the MT pulse which is different from the region to be imaged; andacquiring the echo signal from the region to be imaged after applying the gradient spoiler. 6. The magnetic resonance imaging method according to claim 5, wherein the two types of nuclear pools consist of a nuclear pool of free water and a nuclear pool of a macromolecule. 7. The magnetic resonance imaging method according to claim 6, wherein the echo signal from the region to be imaged is acquired by a two-dimensional scan performed based on a multi-slice imaging technique. 8. The magnetic resonance imaging method according to claim 6, wherein the echo signal... | Kabushiki Kaisha Toshiba | A61B005/05
US07177673 | B2 | 2007-02-13 | Medical telemetry system | 2002-04-16 | A medical telemetry system includes a first detection section for detecting a first vital sign signal using a first sensor, a transmitter 10 which modulates a signal detected by the first detection section and transmits the modulated signal by radio wave transmission, a receiver 30 which receives the detection signal that has been transmitted by radio wave transmission and demodulates the received signal, having a connector 32 for transmitting the demodulated signal to a vital sign signal input section of the vital sign signal monitor 20, wherein the connector 32 of the receiver is connectable with a commonly used vital sign signal input section of each of various types of vital sign signal monitors 20. | 1. A medical telemetry system, comprising: a first sensor, adapted to be attached on a living body to detect a first vital sign signal;a first transmitter, adapted to be connected to the first sensor and transmitting the first vital sign signal by radio wave transmission;a receiver, receiving the first vital sign signal transmitted from the first transmitter;a connector, integrally provided with the receiver and having a first terminal operable to output at least the first vital sign signal received by the receiver;at least one monitor, operable to display at least the first vital sign signal; anda connection wire, having a first end adapted to be connected to the monitor and a second end provided with a second terminal adapted to be detachably connected to the first sensor and to be detachably connected to the first terminal. | 1. A medical telemetry system, comprising: a first sensor, adapted to be attached on a living body to detect a first vital sign signal;a first transmitter, adapted to be connected to the first sensor and transmitting the first vital sign signal by radio wave transmission;a receiver, receiving the first vital sign signal transmitted from the first transmitter;a connector, integrally provided with the receiver and having a first terminal operable to output at least the first vital sign signal received by the receiver;at least one monitor, operable to display at least the first vital sign signal; anda connection wire, having a first end adapted to be connected to the monitor and a second end provided with a second terminal adapted to be detachably connected to the first sensor and to be detachably connected to the first terminal. 2. The medical telemetry system as set forth in claim 1, wherein the connector is detachably connected with the receiver. 3. The medical telemetry system as set forth in claim 1, wherein: the first sensor is operable to detect a second vital sign signal; andthe first transmitter is operable to transmit the first vital sign signal and the second vital sign signal in a multiplexing manner. 4. The medical telemetry system as set forth in claim 3, wherein the receiver is operable to demultiplex the first vital sign signal and the second vital sign signal, and is operable to transmit the first vital sign signal and the second vital sign to the first terminal in a superimposing manner. 5. The medical telemetry system as set forth in claim 4, wherein the second vital sign signal is superimposed using an impedance change technique. 6. The medical telemetry system as set forth in claim 3, wherein the first vital sign signal is an electrocardiographic waveform of the living body, and the second vital sign signal is a respiratory status of the living body. 7. The medical telemetry system as set forth in claim 1, further comprising a second sensor, adapted to be attached on the living body to detect a second vital sign signal, wherein the first transmitter is operable to transmit the first vital sign signal and the second vital sign signal in a multiplexing manner. 8. The medical telemetry system as set forth in claim 7, wherein the first vital sign signal is an electrocardiographic waveform of the living body, and the second vital sign signal is a respiratory status of the living body. 9. The medical telemetry system as set forth in claim 7, wherein the receiver is operable to demultiplex the first vital sign signal and the second vital sign signal, and is operable to transmit the first vital sign signal and the second vital sign signal to the first terminal in a superimposing manner. 10. The medical telemetry system as set forth in claim 9, wherein the second vital sign signal is superimposed using an impedance change technique. 11. The medical telemetry system as set forth in claim 1, further comprising: a second sensor, a... | Nihon Kohden Corporation | A61B005/0428 A61B005/0432
US07162289 | B2 | 2007-01-09 | Method and apparatus for enhancing the integrity of an implantable sensor device | 2002-12-31 | A method and apparatus for enhancing the integrity of an implantable sensor. Voids formed between an outer tubing and a sensor substrate or spacing element may be back-filled with a curable, implantable material, minimizing the extent to which unwanted fluids diffuse within the sensor. An enzyme or protein matrix pellet below the sensor window may be pre-treated with a reducing agent to enhance its bond stability, and to reduce undesired swelling that may cause the sensor window to detach or leak. The bonding between the enzyme pellet and a hydrogel layer may be reinforced by application of an intervening bonding layer of a protein material, such as human serum albumin (HSA). The size of the window may be minimized by minimizing the size of an underlying electrode, providing reduced flux and lengthening sensor. A coating may be deposited on the surface of the sensor leads, providing stiffening and lubrication. | 1. An implantable sensing apparatus comprising: a substrate having an electrode;an outer tubing surrounding said substrate, said outer tubing having a window opening; andan enzyme aligned between said electrode and said window opening, said enzyme treated with a reducing agent to inhibit excessive swelling. | 1. An implantable sensing apparatus comprising: a substrate having an electrode;an outer tubing surrounding said substrate, said outer tubing having a window opening; andan enzyme aligned between said electrode and said window opening, said enzyme treated with a reducing agent to inhibit excessive swelling. 2. The apparatus of claim 1, wherein said reducing agent comprises Ascorbate. 3. The apparatus of claim 1, wherein said reducing agent comprises cyanoborohydride. 4. The apparatus of claim 1, wherein the outer tubing is coated with a lubricious coating. 5. The apparatus of claim 4, wherein the lubricious coating is siloxane. 6. The apparatus of claim 1, further comprising an inner tubing disposed within an interior of the outer tubing. 7. The apparatus of claim 6, wherein the inner tubing is made from an elastomeric material. 8. The apparatus of claim 6, wherein the inner tubing is made from a metal. 9. An implantable sensing apparatus comprising: a substrate comprising an electrode;an outer tubing surrounding said substrate, said outer tubing having a window opening;an enzyme aligned between said electrode and said window opening;a hydrogel layer sealing said window opening; anda bonding layer between said enzyme and said hydrogel layer, said bonding layer being a different material from said enzyme;wherein the bonding layer comprises a protein, human serum albumin (HSA), or a polypeptide. 10. The apparatus of claim 9, wherein said bonding layer comprises a protein. 11. The apparatus of claim 9, wherein said bonding layer comprises human serum albumin (HSA). 12. The apparatus of claim 9, wherein said bonding layer comprises a polypeptide. 13. The apparatus of claim 9, wherein said bonding layer comprises an adhesive. 14. The apparatus of claim 9, wherein the outer tubing is coated with a lubricious coating. 15. The apparatus of claim 14, wherein the lubricious coating is siloxane. 16. The apparatus of claim 9, further comprising an inner tubing disposed within an interior of the outer tubing. 17. The apparatus of claim 16, wherein the inner tubing is made from an elastomeric material. 18. The apparatus of claim 16, wherein the inner tubing is made from a metal. 19. The apparatus of claim 9, wherein the enzyme comprises a protein matrix. 20. The apparatus of claim 9, wherein said bonding layer is free of said enzyme. 21. The apparatus of claim 9, wherein said bonding layer is chemically free of said enzyme. 22. The apparatus of claim 9, wherein said bonding layer is only on one side of said enzyme. 23. The apparatus of claim 9, wherein said bonding layer prevents said enzyme from directly contacting said hydrogel layer. 24. An implantable sensing apparatus comprising: a substrate comprising an electrode;a spacing element adjacent to said substrate;an outer tubing surrounding said substrate and said spacing element, said outer tubing having a window opening, wherein geometrical tolerances between said outer tubing and one... | Medtronic Minimed, Inc. | A61B005/00 A61B005/05 C12M001/00 C12Q001/00
US06916294 | B2 | 2005-07-12 | Brain retraction sensor | 2002-07-09 | An electrode grid device is disclosed comprising a deformable envelope, further comprising non-polarizable electrodes and a pressure recording port. The device is designed to allow for monitoring of brain retraction pressure and local cortical electrical activity including DC potential, as well as to redistribute the force applied during retraction and thereby diminish the chance of focal brain injury during surgery. Retraction pressure recorded is equal over the full area of contact, providing a more meaningful measurement than simply at one point on the retractor. A means is disclosed for evacuation of air from the system to improve accuracy and fidelity of the pressure measurements. It is a further aspect of the device to allow for measurement of intracranial pressure, DC potential and EEG in epileptic and severe head trauma patients for management of edema and injury, respectively. | 1. A sub-dural sensor device comprising: an elongated first end comprising a substantially inextensible cavity, wherein said first end is configured to be placed sub-durally; a plurality of non-polarizable electrodes exposed along at least one surface of said first end, and an electric conduit and a hydraulic conduit comprising an exit port at a shortened second end, wherein said first end engages said second end at said exit port, and further wherein one end of said electric conduit and one end of a remote reference electrode electrically connect to a differential DC amplifier. | 1. A sub-dural sensor device comprising: an elongated first end comprising a substantially inextensible cavity, wherein said first end is configured to be placed sub-durally; a plurality of non-polarizable electrodes exposed along at least one surface of said first end, and an electric conduit and a hydraulic conduit comprising an exit port at a shortened second end, wherein said first end engages said second end at said exit port, and further wherein one end of said electric conduit and one end of a remote reference electrode electrically connect to a differential DC amplifier. 2. The device of claim 1, wherein said first end comprises substantially flat, smooth atraumatic faces. 3. The device of claim 1, wherein said first end comprises a thin, elastically deformable, bio-compatible material that conforms to contours of the brain. 4. The device of claim 1, wherein said electric conduit comprises at least one tail at an end distal to said first end, wherein said at least one tail allows for communications between the device and at least one peripheral apparatus. 5. The device of claim 1, wherein said cavity comprises a bladder or a network of interconnected lumens internal to a flat matrix. 6. The device of claim 1, wherein said hydraulic conduit comprises a flexible, noncompliant material having at least two lumenal surfaces. 7. The device of claim 6, wherein said hydraulic conduit comprises a double lumen catheter. 8. The device of claim 7, wherein one end of each separate lumeral surface comprising said double lumen catheter engages separate fluid-flow directing connectors. 9. The device of claim 8, wherein said fluid flow connectors are female luer-lock connectors. 10. The device of claim 9, wherein said connectors are affixed in a single unit. 11. The device of claim 1, wherein said cavity comprises a fluid selected from the group consisting of normal saline, buffered salt solution, Ringer's solution, Elliott's B solution, and mock cerebrospinal fluid. 12. The device of claim 1, wherein said first end can accommodate pressures of between about one mm Hg and about 120 mm Hg. 13. The device of claim 1, wherein the outer surface flanking the cavity of said elongated first end further comprises a plurality of perforations, wherein said perforations allow liquids to flow into a cistern or lumenal surface below said outer surface for the collection of residual fluids. 14. The device of claim 13, wherein said cistern or lumenal surface channels said residual fluids to one end of said elongated first end, wherein said cistern or lumenal surface engages a conduit, and further wherein said conduit engages an apparatus for collecting said fluid. 15. The device of claim 13, wherein said residual fluids are serosanguineous wound fluid or cerebral spinal fluid (CSF). 16. The device of claim 14, wherein said residual fluid flows into said apparatus by gravity flow. 17. The device of claim 14, wherein said apparatus is a bag. 18. The meth... | George Washington University | A61B005/103 A61B005/117
US06885886 | B2 | 2005-04-26 | Method and system for visualizing a body volume and computer program product | 2002-03-27 | A method for visualizing a body volume, and a computer program product and a system for carrying out the method, wherein the method includes the steps of establishing a criterion for each pixel of a selected data set for whether or not the pixel is assigned to an image background; calculating a synthesized representation from at least two selected diagnostic data sets which are not identical, excluding the pixels assigned to an image background, the selected diagnostic data sets having a predefined spatial allocation with respect to each other; calculating each of the data values of the synthesized representation as a mathematical function of at least one data value of each of the selected data set; and displaying the synthesized representation of a data set whose data values represent the body volume two- or three-dimensionally on a display unit. | 1. A method for visualizing a body volume, wherein a data set whose data values represent said body volume is displayed on a display unit two- or three-dimensionally, in which method: a synthesized representation is calculated from at least two selected diagnostic data sets which are not identical and have a predefined spatial allocation with respect to each other, wherein each of said data values of said synthesized representation is calculated as a mathematical function of at least one data value of each of said selected data set, and said synthesized representation is displayed on the display unit, wherein it is established by way of a criterion for each pixel of a selected data set whether said pixel is assigned to an image background or not, and those pixels which are assigned to an image background are not taken into account when calculating the synthesized representation. | 1. A method for visualizing a body volume, wherein a data set whose data values represent said body volume is displayed on a display unit two- or three-dimensionally, in which method: a synthesized representation is calculated from at least two selected diagnostic data sets which are not identical and have a predefined spatial allocation with respect to each other, wherein each of said data values of said synthesized representation is calculated as a mathematical function of at least one data value of each of said selected data set, and said synthesized representation is displayed on the display unit, wherein it is established by way of a criterion for each pixel of a selected data set whether said pixel is assigned to an image background or not, and those pixels which are assigned to an image background are not taken into account when calculating the synthesized representation. 2. The method as set forth in claim 1, wherein said criterion is a threshold value and those pixels of a selected data set whose intensity is below said threshold value are assigned to an image background. 3. The method as set forth in claim 1, wherein for calculating the synthesized representation, the data value of each pixel of a selected data set is multiplied by a factor and the data value of an assigned pixel of another selected data set is multiplied by a complementary factor. 4. The method as set forth in claim 1, wherein said at least two selected data sets are selected from data sets, each of which is captured by a different method of diagnosis. 5. The method as set forth in claim 1, wherein at least two of said selected data sets are calculated by different image processing means using different image processing parameters from one and the same original data set captured by a method of diagnosis. 6. The method as set forth in claim 1, wherein said selected data sets are inputted to a data processing means. 7. The method as set forth in claim 1, wherein image processing is carried out to synthesize said synthesized representation. 8. The method as set forth in claim 1, wherein said synthesized representation is derived by at least one of subtracting and adding two selected data sets, respectively. 9. The method as set forth in claim 7, wherein image processing is carried out by means of at least one parameter influencing at least one of the color and the opacity allocation of the intensity values of said data sets. 10. The method as set forth in claim 9, wherein the result of said image processing is visualized as a predefined two-dimensional slice image through said body volume. 11. The method as set forth in claim 1, wherein, in said synthesized representation, different color values are assigned to data from different selected data sets. 12. The method as set forth in claim 11, wherein the display parameters of at least one selected data set in said synthesized representation are varied over time. 13. The method as set forth in claim 1, wherein i... | BrainLAB AG | A61B005/05 G06K009/00
US06865413 | B2 | 2005-03-08 | ECG signal processor and method | 2002-01-23 | Chest compressions are measured and prompted to facilitate the effective administration of CPR. A displacement detector determines a displacement indicative signal indicative of the displacement of the CPR recipient's chest toward the recipient's spine. A signaling mechanism provides chest compression indication signals directing a chest compression force being applied to the chest and a frequency of such compressions. An automated controller and an automated constricting device may be provided for applying CPR to the recipient in an automated fashion. The automated controller receives the chest compression indication signals from the signaling mechanism, and, in accordance with the chest compression indication signals, controls the force and frequency of constrictions. The system may be provided with a tilt compensator comprising a tilt sensor mechanism outputting a tilt compensation signal indicative of the extent of tilt of the device, and may be further provided with an adjuster for adjusting the distance value in accordance with the tilt compensation signal. An ECG signal processor may be provided which removes the CPR-induced artifact from a measured ECG signal obtained during the administration of CPR. | 1. An ECG-signal processor comprising: a first input for receiving a measured motion signal indicative of motion of a CPR recipient's chest during the administration of CPR by a rescuer to the CPR recipient; a second input for receiving a measured ECG signal during the administration of CPR; a CPR-induced artifact identifier for receiving said measured motion signal and said measured ECG signal and identifying a CPR-induced artifact contained in said measured ECG signal; and a CPR-induced artifact remover for producing a processed ECG signal by removing said CPR-induced artifact from said measured ECG signal. | 1. An ECG-signal processor comprising: a first input for receiving a measured motion signal indicative of motion of a CPR recipient's chest during the administration of CPR by a rescuer to the CPR recipient; a second input for receiving a measured ECG signal during the administration of CPR; a CPR-induced artifact identifier for receiving said measured motion signal and said measured ECG signal and identifying a CPR-induced artifact contained in said measured ECG signal; and a CPR-induced artifact remover for producing a processed ECG signal by removing said CPR-induced artifact from said measured ECG signal. 2. The ECG-signal processor of claim 1 wherein the CPR-induced artifact identifier is non-linear. 3. The ECG signal processor according to claim 1, wherein said CPR-induced artifact identifier comprises a system identifier for identifying an estimated linear system {tilde over (h)}. 4. The ECG signal processor of claim 3 wherein {tilde over (h)} is non-linear. 5. An ECG signal processor for processing a measured ECG signal from a patient undergoing CPR, said patient having a chest, comprising: a CPR-induced artifact identifier for receiving a first input and a second input, wherein said first input comprises a signal indicative of a CPR induced motion of the chest and wherein said second input comprises the measured ECG signal, and wherein the CPR-induced artifact identifier produces an output comprising an estimated CPR-induced artifact signal; and a signal adder for receiving said output and for receiving the second input, wherein said signal adder combines the output and the second input to produce an estimated true ECG signal. 6. The ECG signal processor of claim 5 wherein the artifact identifier is {tilde over (h)}non-linear. 7. An ECG signal processor for processing a measured ECG signal from a patient undergoing CPR, said patient having a chest, comprising: a first fast Fourier transform means for receiving a signal indicative of a CPR induced motion of the chest, and wherein the first fast Fourier transform means produces a frequency domain of said signal indicative of a CPR induced motion of the chest; a second fast Fourier transform means for receiving the measured ECG signal, and wherein the second fast Fourier transform means produces a frequency domain of said measured ECG signal; an autospectrum calculator for receiving the frequency domain of the signal indicative of the CPR induced motion of the chest, wherein the autospectrum calculator produces an autospectrum of the signal indicative of the CPR induced motion of the chest; a cross-spectrum calculator for receiving the frequency domain of the signal indicative of the CPR induced motion of the chest and the frequency domain of the measured ECG signal, wherein the cross-spectrum calculator produces a cross-spectrum of the signal indicative of the CPR induced motion of the chest and the measured ECG signal; a complex divider for receiving the autospectrum of the signal... | Revivant Corporation | A61B005/04
US06922577 | B2 | 2005-07-26 | Optical measuring device having d.c. component elimination | 2003-02-10 | An optical measuring device includes a sensor which emits light flashes toward a subject of detection with a light emitting element which is supplied with electric power from a rechargeable battery for a certain duration at certain intervals and receives reflected light from the subject with a light sensitive element. A subtractor subtracts a d.c. component from a sensed signal provided by the light sensitive element, thereby extracting a variation component. The subtractor has its output signal amplified with an amplifier, with the output thereof being A/D-converted with an A/D converter. | 1. An optical measuring device, comprising: a sensor having a light emitting element which emits light flashes to a subject of detection by being supplied with electric power from a battery for a prescribed duration at prescribed intervals and a light sensitive element which receives reflected light from the subject; a subtractor which subtracts a d.c. component from a sensed signal provided by the light sensitive element, thereby extracting a variation component; an amplifier which amplifies an output signal of the subtractor; an A/D converter which implements an A/D conversion for an output signal of the amplifier; and a d.c. component modifying means which modifies a d.c. component level based on an A/D-converted value of the sensed signal provided by the light sensitive element. | 1. An optical measuring device, comprising: a sensor having a light emitting element which emits light flashes to a subject of detection by being supplied with electric power from a battery for a prescribed duration at prescribed intervals and a light sensitive element which receives reflected light from the subject; a subtractor which subtracts a d.c. component from a sensed signal provided by the light sensitive element, thereby extracting a variation component; an amplifier which amplifies an output signal of the subtractor; an A/D converter which implements an A/D conversion for an output signal of the amplifier; and a d.c. component modifying means which modifies a d.c. component level based on an A/D-converted value of the sensed signal provided by the light sensitive element. 2. An optical measuring device according to claim 1, wherein the d.c. component modifying means raises the d.c. component level when the ND-converted value of the sensed signal provided by the light sensitive element becomes larger than a prescribed upper limit value, and lowers the d.c. component level when the A/D-converted value of the sensed signal provided by the light sensitive element becomes smaller than a prescribed lower limit value. 3. An optical measuring device according to claim 1, wherein the sensor is used by being put on a part of a human body, with the light emitting element being adapted to emit the light flashes toward a blood vessel as a subject of detection. 4. An optical measuring device, comprising: a sensor having a light emitting element which emits light flashes to a subject of detection by being supplied with electric power from a battery for a prescribed duration at prescribed intervals and a light sensitive element which receives reflected light from the subject; a subtractor which subtracts a d.c. component from a sensed signal provided by the light sensitive element, thereby extracting a variation component; an amplifier which amplifies an output signal of the subtractor; an A/D converter which implements an A/D conversion for an output signal of the amplifier; wherein data of the d.c. component of the sensed signal is delivered in addition to a data of A/D-converted value provided by the A/D converter, and wherein the data of the d.c. component of the sensed signal is determined from the sensed signal without being applied to the subtractor. 5. An optical measuring device according to claim 4, wherein the sensor is used by being put on a part of a human body, with the light emitting element being adapted to emit the light flashes toward a blood vessel as a subject of detection. 6. An optical measuring device comprising: a sensor having a light emitting element which emits light flashes to a subject of detection by being supplied with electric power from a battery for a prescribed duration at prescribed intervals and a light sensitive element which receives reflected light from the subject; a subtractor which subtracts a d.c. com... | Denso Corporation | A61B005/00 A61B005/02
US06907284 | B2 | 2005-06-14 | Method and apparatus for displaying a fetal heart rate signal | 2002-05-03 | A system and apparatus implementing a graphical user interface for displaying heart rate information is provided. The graphical user interface displays, in a first viewing window, a first tracing indicative of a heart rate signal. The graphical user interface displays a control allowing a user to select a portion of the first tracing. In a second viewing window displayed simultaneously with the first viewing window, the graphical user interface displays a second tracing which is a zoomed in view of the selected portion of the first tracing. In a specific implementation, the control includes a selection box having a transparent portion superposed upon the first viewing window. The portion of the first tracing viewable through the transparent portion corresponds to the second tracing. The control allows the user to displace and modify the size of the selection box to select a portion of the first tracing. | 1. A computer readable storage medium storing a program element suitable for execution by a CPU, said program element implementing a graphical user interface module for displaying fetal heart rate information, said graphical user interface module being adapted for: a) in a first viewing window, displaying a first tracing indicative of a fetal heart rate signal; b) displaying a control allowing a user to select a portion of the first tracing in the first viewing window; c) in a second viewing window displayed simultaneously with the first viewing window, displaying a second tracing which is a zoomed in view of the selected portion of the first tracing. | 1. A computer readable storage medium storing a program element suitable for execution by a CPU, said program element implementing a graphical user interface module for displaying fetal heart rate information, said graphical user interface module being adapted for: a) in a first viewing window, displaying a first tracing indicative of a fetal heart rate signal; b) displaying a control allowing a user to select a portion of the first tracing in the first viewing window; c) in a second viewing window displayed simultaneously with the first viewing window, displaying a second tracing which is a zoomed in view of the selected portion of the first tracing. 2. A computer readable storage medium as defined in claim 1, wherein said control includes a selection box. 3. A computer readable storage medium as defined in claim 2, wherein said selection box includes a transparent portion, said selection box being superposed upon said first viewing window. 4. A computer readable storage medium as defined in claim 3, wherein said control allows the user to displace the selection box along an axis associated with the first tracing in the first viewing window to select a portion of the first tracing in the first viewing window. 5. A computer readable storage medium as defined in claim 3, wherein said first tracing is indicative of a heart rate signal over a first time segment: and said second tracing is indicative of a heart rate signal over a second time segment, each of the first time segment and the second time segment having respective durations, said selection box including handles allowing a user to modify the size of the selection box to select the duration of the second time segment. 6. A computer readable storage medium as defined in claim 4, wherein said control allows the user to displace the selection box by using an input device selected from the set consisting of a mouse, keyboard, pointing device, speech recognition unit and touch sensitive screen. 7. A computer readable storage medium as defined in claim 5, wherein the control allows the user to modify the size of the selection box by using an input device selected from the set consisting of a mouse, keyboard, pointing device, speech recognition unit and touch sensitive screen. 8. A computer readable storage medium as defined in claim 3, wherein said graphical user interface module is adapted for displaying in a third display window a third tracing indicative of a uterine contraction pattern. 9. A computer readable storage medium as defined in claim 8, wherein said first tracing is indicative of a heart rate signal over a first time segment and said second tracing is indicative of a heart rate signal over a second time segment, the third tracing being indicative of a uterine contraction pattern during the second time segment. 10. A computer readable storage medium as defined in claim 3, wherein said graphical user interface module is adapted for displaying in a third display window includ... | LMS Medical Systems Ltd. | A61B005/0444
US06904311 | B2 | 2005-06-07 | Method of determining and documenting cessation of life using evoked response | 2002-12-19 | The date and time of death of a patient may be recorded in a memory of a microprocessor used in an implantable cardiac rhythm management device by detecting the failure of the heart to produce an evoked response upon the application of a pace pulse to the heart. To confirm death, other physiologic sensors used in the cardiac rhythm management device may also be sensed to determine the absence of an expected output signal. | 1. A method of sensing a death of a patient having an implanted cardiac rhythm management device, said cardiac rhythm management device having a ventricular sensing amplifier for detecting natural and evoked cardiac depolarization signals, a microprocessor-based controller coupled to receive an output from the ventricular sensing amplifier and for developing control signals, and a stimulating pulse generator for producing cardiac stimulating pulses at timed intervals determined by said control signals, comprising the steps of: (a) monitoring the ventricular sensing amplifier for an evoked response of the patient's heart to the application of the cardiac stimulating pulses; and (b) recording in the cardiac rhythm management device the date and time of death as the date and time at which no intrinsic beat is detected by the ventricular sensing amplifier and application of cardiac stimulating pulses ceases to produce an evoked response. | 1. A method of sensing a death of a patient having an implanted cardiac rhythm management device, said cardiac rhythm management device having a ventricular sensing amplifier for detecting natural and evoked cardiac depolarization signals, a microprocessor-based controller coupled to receive an output from the ventricular sensing amplifier and for developing control signals, and a stimulating pulse generator for producing cardiac stimulating pulses at timed intervals determined by said control signals, comprising the steps of: (a) monitoring the ventricular sensing amplifier for an evoked response of the patient's heart to the application of the cardiac stimulating pulses; and (b) recording in the cardiac rhythm management device the date and time of death as the date and time at which no intrinsic beat is detected by the ventricular sensing amplifier and application of cardiac stimulating pulses ceases to produce an evoked response. 2. The method as in claim 1 wherein the microprocessor-based controller further includes a buffer memory for storing electrogram signals occurring during a predetermined time interval comprising the further step of capturing the contents of the buffer upon detection that no intrinsic beats are detected and the application of cardiac stimulating pulses to the heart does not result in an evoked response. 3. The method as in claim 1 and further including the step of reading out from the cardiac rhythm management device to an external monitor the date and time of death. 4. The method as in claim 1 wherein the cardiac rhythm management device further includes a physiologic sensor for producing an output signal relating to physiologic events and comprising a further step of monitoring the output signal from the physiologic sensor. 5. The method as in claim 4 wherein the physiologic sensor comprises an accelerometer. 6. The method as in claim 4 wherein the physiologic sensor is a minute ventilation sensor. 7. The method as in claim 4 wherein the microprocessor-based controller includes a buffer memory for storing the output signals of the physiologic sensor and comprising a further step of capturing the contents of the buffer memory upon detection that no intrinsic beats are detected and the application of cardiac stimulating pulses to the heart does not result in an evoked response. 8. The method as in claim 7 and further including the step of reading out from the cardiac rhythm management device the captured contents of the buffer memory to an external monitor. | Cardiac Pacemakers, Inc. | A61B005/0402 A61N001/365 A61N001/37
US06996434 | B2 | 2006-02-07 | Method and system for verifying the integrity of normal sinus rhythm templates | 2001-08-02 | A method and system for verifying the integrity of normal sinus rhythm (NSR) templates and updating the NSR template after selected time intervals. At selected time intervals after establishing a NSR template, cardiac complexes are sensed and values for one or more cardiac parameters are measured. The values of the cardiac parameters are compared to predetermined value ranges for NSR cardiac complexes. When the values of the cardiac parameters fall within the predetermined value ranges, values for the differences between the values of the cardiac parameters for the cardiac complexes and the values for the cardiac parameters of the NSR cardiac complexes are calculated. When the values of the differences are greater than one or more threshold values, the NSR template is updated as a function of the sensed cardiac complexes. | 1. A method, comprising: comparing one or more sensed cardiac complex parameter values to one or more predetermined cardiac complex parameter values;determining whether the one or more sensed cardiac complex parameter values and the one or more predetermined cardiac complex parameter values have differences greater than a threshold; andupdating a normal sinus rhythm (NSR) template upon determining that the one or more sensed cardiac complex parameter values and the one or more predetermined cardiac complex parameter values have differences that are greater than the threshold. | 1. A method, comprising: comparing one or more sensed cardiac complex parameter values to one or more predetermined cardiac complex parameter values;determining whether the one or more sensed cardiac complex parameter values and the one or more predetermined cardiac complex parameter values have differences greater than a threshold; andupdating a normal sinus rhythm (NSR) template upon determining that the one or more sensed cardiac complex parameter values and the one or more predetermined cardiac complex parameter values have differences that are greater than the threshold. 2. The method of claim 1, further comprising retaining the NSR template upon determining that the one or more sensed cardiac complex parameter values and the one or more predetermined cardiac complex parameter values do not have differences that are greater than the threshold. 3. The method of claim 1, further comprising determining whether the one or more sensed cardiac complex parameter values are within one or more ranges of normal values prior to comparing the one or more sensed cardiac complex parameter values to the one or more predetermined cardiac complex parameter values, wherein the one or more sensed cardiac complex parameter values are compared upon determining that the one or more sensed cardiac complex parameter values are within the one or more ranges of normal values. 4. The method of claim 3, further comprising: upon determining that the one or more sensed cardiac complex parameter values are not within the one or more ranges of normal values, determining whether a maximum number of attempts has been made to evaluate the one or more sensed cardiac complex parameter values; andupon determining that the maximum number of attempts has been made, retaining the NSR template. 5. The method of claim 1, further comprising waiting a selected time period before comparing the one or more sensed cardiac complex parameter values to one or more predetermined cardiac complex parameter values and determining whether the one or more sensed cardiac complex parameter values and the one or more predetermined cardiac complex parameter values have differences greater than a threshold. 6. The method of claim 5, wherein comparing one or more sensed cardiac complex parameter values to one or more predetermined cardiac complex parameter values and determining whether the one or more sensed cardiac complex parameter values and the one or more predetermined cardiac complex parameter values have differences greater than a threshold are automatically performed after the selected time period. 7. The method of claim 1, wherein comparing one or more sensed cardiac complex parameter values and determining whether the one or more sensed cardiac complex parameter values and the one or more predetermined cardiac complex parameter values have differences greater than a threshold are performed after receiving a physician-initiated command. 8. The method of claim 1, wherein comparing one or... | Cardiac Pacemakers, Inc. | A61B005/04
US07135003 | B2 | 2006-11-14 | Feedback estimation of joint forces and joint moments | 2002-05-16 | Apparatus and methods are provided for estimating joint forces and moments in human beings. A forward dynamics module determines simulated kinematic data. An error correction controller forces tracking error between the simulated kinematic data and measured (or desired) kinematic data to approach zero. The error correction controller generates a modified acceleration for input into an inverse dynamics module. The estimated joint forces and moments track the measured (or desired) kinematics without the errors associated with computing higher order derivatives of noisy kinematic data. | 1. An apparatus for determining an estimated joint load at a joint of interest in an anatomical system, the apparatus comprising: an error correction controller configured to calculate a modified acceleration using at least kinematic data;an inverse dynamics module configured to determine the estimated joint load at the joint of interest by recursively estimating successive joint loads until the joint of interest is reached, the inverse dynamics module using at least the modified acceleration and an input force for the anatomical system; anda forward dynamics module configured to determine simulated kinematic data for the anatomical system using at least the estimated joint load. | 1. An apparatus for determining an estimated joint load at a joint of interest in an anatomical system, the apparatus comprising: an error correction controller configured to calculate a modified acceleration using at least kinematic data;an inverse dynamics module configured to determine the estimated joint load at the joint of interest by recursively estimating successive joint loads until the joint of interest is reached, the inverse dynamics module using at least the modified acceleration and an input force for the anatomical system; anda forward dynamics module configured to determine simulated kinematic data for the anatomical system using at least the estimated joint load. 2. The apparatus of claim 1 wherein the error correction controller comprises a difference module configured to generate an error value from the simulated kinematic data and the kinematic data, and wherein the error correction controller is configured to determine the modified acceleration by applying a feedback gain to the error value. 3. The apparatus of claim 2 wherein the feedback gain comprises at least one of a positional feedback gain and a velocity feedback gain. 4. The apparatus of claim 1 wherein the estimated joint load in a first recursive instance is coupled to the input force of a second recursive instance. 5. The apparatus of claim 1 wherein the error correction controller receives the simulated kinematic data as feedback. 6. The apparatus of claim 1 wherein the inverse dynamics module receives the simulated kinematic data as feedback. 7. The apparatus of claim 1 wherein the forward dynamics module receives the simulated kinematic data as feedback. 8. The apparatus of claim 1 wherein the forward dynamics module performs an integration to determine the simulated kinematic data. 9. The apparatus of claim 1 wherein the input force comprises a ground reaction force. 10. The apparatus of claim 1 wherein the input force comprises force plate measurement. 11. The apparatus of claim 1 wherein the input force is equal to zero. 12. The apparatus of claim 1 wherein the input force acts at a point of contact on a segment of the anatomical system. 13. The apparatus of claim 1 wherein the input force comprises at least one of an internal force and an internal moment. 14. The apparatus of claim 1 wherein the input force comprises a joint torque. 15. An apparatus for determining an estimated joint load at a joint in an anatomical system, the apparatus comprising: an error correction controller configured to calculate a modified acceleration using at least kinematic data, wherein the modified acceleration is computed without using accelerations estimated from measured kinematics;an inverse dynamics module configured to determine the estimated joint load using at least the modified acceleration; anda forward dynamics module configured to determine simulated kinematic data for the anatomical system using at least the estimated joint load. 16. The apparatus of ... | Honda Giken Kogyo Kabushiki Kaisha | A61B005/117 A61B005/103 G06G007/48 G06G007/58
US07198607 | B2 | 2007-04-03 | Detector unit, an arrangement and a method for measuring and evaluating forces exerted on a human body | 2002-12-18 | The invention relates an arrangement (10) for measuring and presenting forces exerted on body, describing the amount and intensity of physical exercise. The arrangement according to the invention comprises a detector unit (11), a communications device (13), a server (14) and a database (16) connected therewith, and a data presentation means (18). Furthermore, the invention relates to a detector unit (11) used in the arrangement and a method using the arrangement for measuring and analyzing the effectiveness and amount of physical exercise. | 1. A detector unit for measuring, analysing and presenting acceleration data describing forces exerted on a person's skeletal system in conjunction with the person's activity performance, which detector unit comprises: at least one acceleration; anda processing unit with a memory for carrying out by a first analysis means a first session-specific analysis based on the acceleration data and for saving the results from that analysis, characterised in that the processing unit is arranged so as to utilise measurement data from an acceleration sensor, which data comprise peak values of measured momentary acceleration values and the number of said peak acceleration values during a session, and the first analysis means for making an assessment on the development of bone mass in the first analysis. | 1. A detector unit for measuring, analysing and presenting acceleration data describing forces exerted on a person's skeletal system in conjunction with the person's activity performance, which detector unit comprises: at least one acceleration; anda processing unit with a memory for carrying out by a first analysis means a first session-specific analysis based on the acceleration data and for saving the results from that analysis, characterised in that the processing unit is arranged so as to utilise measurement data from an acceleration sensor, which data comprise peak values of measured momentary acceleration values and the number of said peak acceleration values during a session, and the first analysis means for making an assessment on the development of bone mass in the first analysis. 2. A detector unit according to claim 1, characterised in that at least one of the following is arranged to be produced on the basis of the peak values of measured momentary acceleration values and the direction of said accelerations: an evaluation of the type of exercise performed, an evaluation about a recommendation for additional exercise. 3. A detector unit according to claim 1, characterised in that it also comprises an indicator for indicating to the user of the detector unit the result from the first session-specific analysis. 4. A detector unit to claim 1, characterised in that it also comprises a communications component for enabling acceleration measurement data communication between the detector unit and an external communications device. 5. A detector unit according to claim 1, characterised in that an analysis algorithm, to be used in the processing of measured acceleration data, is arranged to be stored in the memory of the processing unit. 6. A detector unit according to claim 5, characterised in that the computing algorithm used in the first analysis is a computing algorithm for an exercise performance aimed to develop bone mass. 7. A detector unit according to claim 3, characterised in that the indicator is of ON/OFF type where ON indicates achievement of the session-specific goal of an exercise performance, and OFF indicates non-achievement of the goal of an exercise performance. 8. A detector unit according to claim 7, characterised in that the indicator is implemented using two leds of different colours so that ON state is indicated by turning on one led, and OFF state is indicated by turning on the other led. 9. A detector unit according to claim 1, characterised in that the acceleration sensor in the detector unit is arranged so as to measure accelerations in the range ±12 g. 10. A detector unit according to claim 4, characterised in that the communications component comprises a means for handling communication between a communications device and a device suitable for presenting the analysis results based on one of the following: infrared technology, time division multiple access (TDMA) technology, code division multiple access ... | Newtest Oy | A61B005/103 A61B005/117 A61B005/00
US07209777 | B2 | 2007-04-24 | Method and apparatus for automated tracking of non-linear vessel movement using MR imaging | 2000-11-30 | A system and method is disclosed for tracking a moving object using magnetic resonance imaging. The technique includes acquiring a scout image scan having a number of image frames and extracting non-linear motion parameters from the number of image frames of the scout image scan. The technique includes prospectively shifting slice location using the non-linear motion parameters between slice locations while acquiring a series of MR images. The system and method are particularly useful in tracking coronary artery movement during the cardiac cycle to acquire the non-linear components of coronary artery movement during a diastolic portion of the R—R interval. | 1. A method of imaging a moving object using MRI comprising: acquiring a scout image scan having a number of image frames;extracting non-linear motion parameters during one phase of motion of an object of interest from the number of image frames of the scout image scan; andacquiring a series of MR images while prospectively shifting a slice location using the non-linear motion parameters between slice acquisitions. | 1. A method of imaging a moving object using MRI comprising: acquiring a scout image scan having a number of image frames;extracting non-linear motion parameters during one phase of motion of an object of interest from the number of image frames of the scout image scan; andacquiring a series of MR images while prospectively shifting a slice location using the non-linear motion parameters between slice acquisitions. 2. The method of claim 1 wherein the motion parameters include linear parameters and non-linear parameters. 3. The method of claim 1 further comprising the steps of: selecting a region-of-interest encompassing a coronary artery in a reference frame;recording a reference position in the region-of-interest of the reference frame; andcross-correlating the reference position in the region-of-interest of the reference frame with such a position in a region-of-interest for subsequent image frames to extract the non-linear motion parameters. 4. The method of claim 3 further comprising the step of computing displacement of the coronary artery in each frame relative to a reference position. 5. The method of claim 4 wherein the reference position is chosen at an end of diastole. 6. The method of claim 1 further comprising the step of using a multi-phase, single slice pulse sequence to acquire the scout imaging scan, and using a gradient echo spiral pulse sequence to acquire the series of MR images. 7. The method of claim 1 further comprising the step of injecting a contrast agent prior to acquiring the series of MR images to allow use of a pulse sequence with high flip angles. 8. The method of claim 1 wherein the step of acquiring a series of MR images includes data acquisition during a portion of an R—R interval and further comprises at least one of the steps of reorienting the slice location and rotating the slice location. 9. A method of coronary MR angiography imaging comprising: tracking coronary artery movement during a cardiac cycle as the coronary artery moves between two excursion positions;determining a function of the coronary artery movement having at least a non-linear component that represents movement during a diastolic portion of an R—R interval;adjusting slice position acquisition using the determined function;acquiring MR data during each such adjustment; andreconstructing an MR image having increased imaging efficiency using the MR data acquired. 10. The method of claim 9 wherein the MR data is acquired in a single breath-held R—R interval. 11. The method of claim 9 further comprising the step of identifying a reference kernel in an ROI in one scout frame. 12. The method of claim 9 further comprising the step of acquiring a set of scout frames which display movement of the coronary artery over the cardiac cycle. 13. The method of claim 12 further comprising the step of developing a non-linear motion function based on a position of the coronary artery in each scout frame. 14. The method of claim 13 further comprisi... | General Electric Company | A61B005/05
US06975900 | B2 | 2005-12-13 | Systems and methods for determining a surface geometry | 2002-12-12 | Methods and systems for determining a surface geometry of an object, including determining a first projection matrix based on a first imaging device, determining a second projection matrix based on a second imaging device, obtaining at least one first two-dimensional (2D) image of the object using the first imaging device, obtaining at least one second 2D image of the object using the second imaging device, determining a contour of the object in the first 2D image and the second 2D image, and, based on the at least two contours, the first projection matrix, and the second projection matrix, reconstructing 3D data associated with the surface of the object. In one embodiment, the object can be a heart. | 1. A method for computing epicardial surface electric potentials based on measured body surface electric potentials, the method comprising: generating a multidimensional matrix, the matrix representing at least one geometric relationship between at least one body surface electric potential measuring system and the epicardial surface, the multidimensional matrix being based on: at least one first two-dimensional (2D) image of the epicardium using a first imaging device, at least one second 2D image of the epicardium using a second imaging device, a first projection matrix associated with the first imaging device, a second projection matrix associated with the second imaging device, a computed contour of the outer surface of the heart formed from contours identified in the at least one first 2D image and the at least one second 2D image, a computed contour of the electric potential measuring system, and, a boundary element method based on the computed contour of the outer surface of the heart and the computed contour of the electric potential measuring system, using a Generalized Minimum Residual (GMRes) method to estimate an inverse of the multidimensional matrix, and, based on the inverse matrix and the measured body surface potentials, determining the epicardial surface electric potentials. | 1. A method for computing epicardial surface electric potentials based on measured body surface electric potentials, the method comprising: generating a multidimensional matrix, the matrix representing at least one geometric relationship between at least one body surface electric potential measuring system and the epicardial surface, the multidimensional matrix being based on: at least one first two-dimensional (2D) image of the epicardium using a first imaging device, at least one second 2D image of the epicardium using a second imaging device, a first projection matrix associated with the first imaging device, a second projection matrix associated with the second imaging device, a computed contour of the outer surface of the heart formed from contours identified in the at least one first 2D image and the at least one second 2D image, a computed contour of the electric potential measuring system, and, a boundary element method based on the computed contour of the outer surface of the heart and the computed contour of the electric potential measuring system, using a Generalized Minimum Residual (GMRes) method to estimate an inverse of the multidimensional matrix, and, based on the inverse matrix and the measured body surface potentials, determining the epicardial surface electric potentials. 2. A method according to claim 1, where representing includes measuring the position of the at least one body surface electric potential measuring system. 3. A method according to claim 1, where using a GMRes method includes determining a number of iterations for the GMRes method. 4. A method according to claim 3, where using a GMRes method includes providing a maximum number of iterations for the GMRes method, and based on the data from the maximum number of iterations, determining a number of iterations for the GMRes method. 5. A method according to claim 3, where determining a number of iterations includes comparing residual error to a Hessenberg matrix condition, and computing at least one of a corner of a condition L curve and a maximum curvature of a condition L curve. 6. A method according to claim 3, where determining a number of iterations includes determining a number of iterations based on at least one of: a corner of a condition L curve, a corner of an L curve, an increase in spatial frequency of a reconstructed potential map, and an increase in amplitude of a solution norm. 7. A method according to claim 1, where using a GMRes method includes providing an initial condition of zero. 8. A method according to claim 1, where using a GMRes method includes providing an initial condition based on a Tikhonov regularization of the multidimensional matrix. | Case Western Reserve University | A61B005/044
US06945939 | B2 | 2005-09-20 | Hemodynamic analysis | 2002-10-18 | To track significant changes in a hemodynamic variable of a living body regardless of measurement noise induced by cyclical functions of the body, the hemodynamic variable is measured at substantially the same instant during several cycles and the measured values compared to identify any statistically significant changes. The measurements may be discrete or continuous. The heart may be stimulated by a cardiac device during some of the cycles and parameters of the device such as the A-V delay of a pacemaker may be adjusted in accordance with the measured change in the hemodynamic variable, which may be, among others, blood pressure or blood volume. | 1. A method of monitoring changes in a living body, the method comprising: measuring a hemodynamic variable of the living body during two different cycles of a cyclical function of the body, the hemodynamic variable being measured at substantially the same instant in each cycle; and comparing values of the hemodynamic variable measured during the two different cycles; wherein measuring the hemodynamic variable comprises measuring an average value of the hemodynamic variable for a length of time starting at substantially the same instant during each cycle. | 1. A method of monitoring changes in a living body, the method comprising: measuring a hemodynamic variable of the living body during two different cycles of a cyclical function of the body, the hemodynamic variable being measured at substantially the same instant in each cycle; and comparing values of the hemodynamic variable measured during the two different cycles; wherein measuring the hemodynamic variable comprises measuring an average value of the hemodynamic variable for a length of time starting at substantially the same instant during each cycle. 2. The method of claim 1, wherein the hemodynamic variable is blood pressure. 3. The method of claim 1, wherein a cardiac device is selected from among the group of cardiac devices comprising pace makers, defibrillators, and cardioverters. 4. The method of claim 1, wherein the cyclical function is the heart beat. 5. The method of claim 1, wherein the cyclical function is selected from among the group of cyclical functions comprising circadian rhythm, digestion, and locomotion. 6. The method of claim 1, wherein the hemodynamic variable is blood flow rate. 7. A method of monitoring changes in a living body, the method comprising: measuring a hemodynamic variable of the living body during two different cycles of a cyclical function of the body, the hemodynamic variable being measured at substantially the same instant in each cycle; and comparing values of the hemodynamic variable measured during the two different cycles; wherein the cyclical function is respiration. 8. The method of claim 7, wherein the hemodynamic variable is selected from among the group of hemodynamic variables comprising volume of vasculature, volume of cardiac chambers, stroke volume, oxygen consumption, heart sounds, respiration rate, tidal volume, blood gases, pH, and acceleration of the myocardium. 9. The method of claim 7, further comprising: stimulating a heart beat of the living body during one of the two different cycles of the cyclical function with a cardiac stimulation device. 10. The method of claim 9, wherein the cardiac stimulation device is implanted in the body. 11. The method of claim 9, further comprising: adjusting any one or more of an atrio-ventricular delay, an interchamber delay, a RV-LV delay, and a LA-LV delay of the cardiac stimulation device in accordance with a difference between the measured values. 12. The method of claim 7, further comprising: controlling a cardiac stimulation device in accordance with a difference between the measured values. 13. The method of claim 12, wherein the cardiac stimulation device is implanted in the body. 14. An apparatus for monitoring changes in a living body, comprising: a detector configured to measure a hemodynamic variable of the living body during two different cycles of a cyclical function of the body and at substantially the same instant in each cycle; and a controller configured to compare values of the hemodynamic variable measured during ... | Pacesetter, Inc. | A61B005/02
US06965792 | B2 | 2005-11-15 | Susceptometers for foreign body detection | 2001-12-18 | Methods and apparatus for minimizing the effects of temperature drift in a magnetic susceptibility measurement instrument, such as an instrument used in pre-MRI screening for the presence of ferromagnetic foreign bodies. The magnetic field source and magnetic sensors can be combined into a single, rigid unit. The stability and sensitivity required in high quality magnetic susceptibility measurements can be achieved through symmetrical design of the source-sensor unit, minimization of thermal stresses, minimization of temperature variations, use of materials with low thermal expansion coefficients, or through appropriate combinations thereof. Use of patient eye movement where an eye is being screened, use of a water bag between the patient and the instrument, or use of telemedicine to facilitate performance of the necessary computations can also be incorporated. | 1. A method for performing pre-MRI screening for ferromagnetic foreign bodies, comprising: providing a source-sensor unit including at least one applied-field generating element and at least one magnetic sensor;canceling the signal due to the applied magnetic field at the location of said at least one magnetic sensor by the relative arrangement of said at least one applied-field generating element and said at least one magnetic sensor in said source-sensor unit;generating an applied magnetic field in the eye of a patient, with said source-sensor unit;rotating said eye, thereby moving any ferromagnetic foreign body which may be present to modulate the magnetic susceptibility signal from said ferromagnetic foreign body, without substantially changing the magnetic susceptibility response of the patient's body tissues; andgenerating a magnetic susceptibility measurement relevant to the presence or absence of a ferromagnetic foreign body in said eye, with said source-sensor unit. | 1. A method for performing pre-MRI screening for ferromagnetic foreign bodies, comprising: providing a source-sensor unit including at least one applied-field generating element and at least one magnetic sensor;canceling the signal due to the applied magnetic field at the location of said at least one magnetic sensor by the relative arrangement of said at least one applied-field generating element and said at least one magnetic sensor in said source-sensor unit;generating an applied magnetic field in the eye of a patient, with said source-sensor unit;rotating said eye, thereby moving any ferromagnetic foreign body which may be present to modulate the magnetic susceptibility signal from said ferromagnetic foreign body, without substantially changing the magnetic susceptibility response of the patient's body tissues; andgenerating a magnetic susceptibility measurement relevant to the presence or absence of a ferromagnetic foreign body in said eye, with said source-sensor unit. 2. The method recited in claim 1, further comprising rotating said eye in a controlled manner. 3. The method recited in claim 1, further comprising preserving said cancellation of the applied-field signal as said entire source-sensor unit expands and contracts. | MedNovus, Inc. Quantum Magnetics, Inc. | A61B005/05
US06928142 | B2 | 2005-08-09 | Non-invasive plaque detection using combined nuclear medicine and x-ray system | 2002-10-18 | A diagnostic imaging system (20) comprising a computer workstation (26) for controlling the imaging system, interfacing with an operator and generating images. A coordinate system (100) is in data communication with the computer workstation. The coordinate system (100) is adapted to describe relative position of components in the diagnostic imaging system (20). A subject support (30) is describable within the coordinate system and an X-ray sub-system (22) is positionable around the subject support (30). Position sensors (44a) are operatively connected to the x-ray sub-system (22) and they provide signals to the workstation (26) indicative of the position of components of the x-ray sub-system (22) within the space represented by the coordinate system. A nuclear camera sub-system (24) is positionable around the subject support (30). Position sensors (44b) are operatively connected to the nuclear camera sub-system (24) and provide signals to the workstation (26) indicative of position of components of the nuclear camera sub-system within the coordinate system. The workstation (100) includes a region of interest position determination function (108) for determining the position within the imaging system of a region of interest (40) in an image generated by the x-ray sub-system (22). The region of interest (40) is precisely targeted using the control console (26) for imaging by the nuclear camera (50). | 1. A diagnostic imaging system, the system comprising: a computer workstation for controlling the imaging system, interfacing with an operator and generating images; a coordinate system in data communication with the computer workstation, the coordinate system adapted to describe relative position in the diagnostic imaging system; a subject support describable within the coordinate system; an x-ray sub-system, the x-ray sub-system positionable around the subject support; position sensors operatively connected to the x-ray sub-system and workstation, the position sensors providing signals to the workstation indicative of the position of components of the x-ray sub-system within the coordinate system; a nuclear camera sub-system, the nuclear camera sub-system positionable around the subject support; position sensors operatively connected to the nuclear camera sub-system and workstation, the position sensors providing signals to the workstation indicative of the position of components of the nuclear camera sub-system within the coordinate system; means for determining the position within the imaging coordinate system of a region of interest in an associated imaging subject from an image generated by the x-ray sub-system; and means for targeting the nuclear camera sub-system to precisely obtain an image at the region of interest. | 1. A diagnostic imaging system, the system comprising: a computer workstation for controlling the imaging system, interfacing with an operator and generating images; a coordinate system in data communication with the computer workstation, the coordinate system adapted to describe relative position in the diagnostic imaging system; a subject support describable within the coordinate system; an x-ray sub-system, the x-ray sub-system positionable around the subject support; position sensors operatively connected to the x-ray sub-system and workstation, the position sensors providing signals to the workstation indicative of the position of components of the x-ray sub-system within the coordinate system; a nuclear camera sub-system, the nuclear camera sub-system positionable around the subject support; position sensors operatively connected to the nuclear camera sub-system and workstation, the position sensors providing signals to the workstation indicative of the position of components of the nuclear camera sub-system within the coordinate system; means for determining the position within the imaging coordinate system of a region of interest in an associated imaging subject from an image generated by the x-ray sub-system; and means for targeting the nuclear camera sub-system to precisely obtain an image at the region of interest. 2. The diagnostic imaging system of claim 1 wherein the coordinate system is a two dimensional system describing the relative position of components in the imaging system in a selected plane. 3. The diagnostic imaging system of claim 1 wherein the coordinate system is a three dimensional system describing the relative position of components within a volume associated with the imaging system. 4. The diagnostic imaging system of claim 1 wherein the x-ray sub system includes and x-ray source and x-ray detector and the nuclear camera sub-system includes a nuclear camera detector head, the diagnostic imaging system including a gantry, the gantry providing a positioning support structure, the support structure commonly supporting and positioning both of the x-ray source and nuclear camera detector head. 5. The diagnostic imaging system of claim 1 including: a region of interest comparison database having a reference image; and means for determining a region of interest in an image generated by one sub-system of the imaging system by comparing the generated image and the reference image of the comparison database. 6. The diagnostic imaging system of claim 1 wherein the nuclear camera sub-system includes: a planar gamma camera head; and a pinhole collimator. 7. The diagnostic imaging system of claim 6 wherein the pinhole collimator is a multiple hole pinhole collimator. 8. The diagnostic imaging system of claim 1 including means for positioning components of the x-ray sub system and nuclear camera sub-system, the means for positioning controllably connected to the workstation. 9. A diagnostic imaging system, the system co... | Koninklijke Philips Electronics N.V. | G01N023/04 G01T001/161 A61B005/05
US07186219 | B2 | 2007-03-06 | Calibration of a doppler velocimeter for stroke volume determination | 2002-10-10 | Doppler Velocimetry is a widely used method for estimating stroke volume (SV). The accuracy and reliability of its measurement however, is dependant on a) the correct assessment of the aortic valve cross-sectional area (CSA), andb) the maximal systolic velocity integral (SVI). The invention avoids the conventional assessment of aortic valve CSA by using a calibration method: a reference stroke volume SVREF is determined by a method different from Doppler velocimetry, e.g. by thoracic electrical bioimpedance (TEB), or thoracic electrical bioadmittance, measured via surface thorax electrodes (transthoracic approach) or via electrodes located directly on an esophageal catheter/probe (esophageal approach). In the latter case, if esophageal Doppler velocimetry is used, the same catheter can be used for the placement of the electrodes and for an ultrasound transducer. In the case of esophageal Doppler velocimetry, an additional benefit of the invention is that the Doppler catheter/probe does not need to be aimed precisely at the site of maximum flow amplitude, as long as a stable catheter/probe location is maintained within the esophagus. | 1. A method of monitoring the stroke volume (SV) of a subject's heart at different time periods, comprising the steps of: determining a reference stroke volume (SVREF) of the subject's heart by a first method different from Doppler velocimetry in a first time period;simultaneously determining a reference systolic velocity integral (SVICAL) by Doppler velocimetry;calculating a constant kCAL, wherein kCAL=SVREF/SVICAL;determining a systolic velocity integral (SVI) at later time periods by Doppler velocimetry; anddetermining the stroke volume (SV) at said later periods by setting SV=kCAL·SVL. | 1. A method of monitoring the stroke volume (SV) of a subject's heart at different time periods, comprising the steps of: determining a reference stroke volume (SVREF) of the subject's heart by a first method different from Doppler velocimetry in a first time period;simultaneously determining a reference systolic velocity integral (SVICAL) by Doppler velocimetry;calculating a constant kCAL, wherein kCAL=SVREF/SVICAL;determining a systolic velocity integral (SVI) at later time periods by Doppler velocimetry; anddetermining the stroke volume (SV) at said later periods by setting SV=kCAL·SVL. 2. The method of claim 1, wherein said first method uses one of the group of transthoracic electrical impedance measurements and transthoracic electrical bioadmittance measurements. 3. The method of claim 1, wherein said first method uses one of the group of esophageal electrical bioimpedance measurements and esophageal electrical bioadmittance measurements. 4. The method of claim 1, wherein esophageal Doppler velocimetry is used. 5. The method of claim 4, wherein, prior to determining said reference stroke volume (SVREF) and said reference systolic velocity integral (SVICAL), a) a catheter comprising a device for performing Doppler velocimetry by obtaining Doppler signals is inserted into the subject's esophagus,b) said catheter is rotated until substantially maximal Doppler signal strength is obtained, andc) said catheter is fixed such that it can no longer be rotated. 6. The method of claim 4, wherein a catheter comprising a device for performing Doppler velocimetry is inserted into the subject's esophagus, and wherein it is attempted to hold the catheter fixed starting at least from said determination of said reference stroke volume (SVREF) and said reference systolic velocity integral (SVICAL), and wherein after each movement of said catheter, despite said attempt to hold the catheter fixed, said reference stroke volume (SVREF) and said reference systolic velocity integral (SVICAL) are newly determined and said constant kCAL is newly calculated, andwherein said newly calculated kCAL is used for the determination of the stroke volume (SV) until a new movement of said cathete occurs. 7. The method of claim 1, wherein during said Doppler velocimetry, a transducer is applied to the suprasternal acoustic window. 8. A method of monitoring the cardiac output (CO) of a subject's heart, comprising: determining the stroke volume (SV) of the subject's heart by using the method according to claim 1;determining the heart rate (HR) of the subject's heart; andcalculating the cardiac output as, CO=SV·HR. 9. The method according to claim 8, wherein the heart rate (HR) is determined by using an electrocardiogram. 10. A system for monitoring the stroke volume (SV) of a subject's heart at different times, comprising: a reference device for determining a reference stroke volume (SVREF) by using a first method different from Doppler velocimetry;a Doppler velocimeter fo... | Osypka Medical GmbH | A61B005/02 A61B005/04 A61B005/08
US07029443 | B2 | 2006-04-18 | System and method for monitoring blood glucose levels using an implantable medical device | 2002-10-21 | T-wave amplitude and QT interval are derived from patient cardiac signals. Then blood glucose levels are determined based on a combination of the T-wave amplitude and the QT interval. By using a combination of both T-wave-based and QT interval-based signals, blood glucose levels can be reliably detected throughout a wide range of blood glucose levels. Once the blood glucose level has been detected, the implanted device compares the blood glucose level against upper and lower acceptable bounds and appropriate warning signals are generated if the level falls outside the bounds. In one example, wherein an implantable insulin pump is additionally provided, the pump is controlled based on the detected blood glucose level to maintain glucose levels within an acceptable range. A calibration technique is also provided for determining patient-specific parameters for use in the detection of blood glucose levels. | 1. A method performed by an external programmer device in communication with an implantable medical device implanted within a patient, the method comprising: receiving signals from the implantable device representative of electrical cardiac signals;receiving signals representative of corresponding patient blood glucose levels;determining parameters relating the electrical cardiac signals to the blood glucose levels for the particular patient; andtransmitting the parameters to the implantable device such that the implantable device can thereafter derive patient blood glucose levels from subsequently detected electrical cardiac signals. | 1. A method performed by an external programmer device in communication with an implantable medical device implanted within a patient, the method comprising: receiving signals from the implantable device representative of electrical cardiac signals;receiving signals representative of corresponding patient blood glucose levels;determining parameters relating the electrical cardiac signals to the blood glucose levels for the particular patient; andtransmitting the parameters to the implantable device such that the implantable device can thereafter derive patient blood glucose levels from subsequently detected electrical cardiac signals. 2. The method of claim 1 wherein receiving signals from the implantable device representative of electrical cardiac signals comprises receiving signals representative of T-wave amplitudes and QT intervals. 3. The method of claim 2 wherein determining parameters relating T-wave amplitudes and QT intervals to blood glucose levels comprises correlating a plurality of blood glucose levels with corresponding T-wave amplitudes and QT intervals based on a pre-determined general functional relationship so as to derive parameters specifying the functional relationship for the particular patient. 4. The method of claim 3 wherein the signals representative of T-wave amplitudes are T-wave amplitude fractions and wherein the signals representative of QT intervals are QTc interval deltas and wherein the pre-determined general functional relationship is: blood glucose level =A−B*QTc−C*(T-wave amplitude fraction)wherein A, B and C are the parameters to be determined that specify the functional relationship for the particular patient. 5. The method of claim 4 wherein transmitting the parameters to the implantable device comprises: transmitting the parameters A, B and C to the implantable device. 6. The method of claim 4 wherein correlating a plurality of blood glucose levels with corresponding T-wave amplitude fractions and QTc delta values comprises performing multiple regression statistical analysis on the plurality of blood glucose levels with corresponding T-wave amplitude fractions and QTc delta values. 7. The method of claim 2 wherein receiving signals from the implantable device representative of patient T-wave amplitudes and QT intervals is performed throughout a pre-determined range of blood glucose levels. 8. The method of claim 7 wherein receiving signals from the implantable device further comprises: receiving signals from the implantable device representative of patient T-wave amplitudes and QT intervals while the blood glucose level of the patient rises to a predetermined upper threshold; andreceiving signals from the implantable device representative of patient T-wave amplitudes and QT intervals while the blood glucose level of the patient decreases from the upper threshold to a pre-determined lower threshold. 9. The method of claim 8 wherein the lower threshold is 40 mg/dl and the upper threshold is 300 mg/dl... | Pacesetter, Inc. | A61B005/145 A61B005/0452 A61N001/362
US07308293 | B2 | 2007-12-11 | Non-invasive glucose meter | 2002-08-01 | An apparatus for measuring a concentration of a substance in an eye (13) includes a measurement light source (4) producing a measurement light beam having a first wavelength at which the substance has a non-zero first absorption coefficient. The apparatus further includes a reference light source (6) producing a reference light beam having a second wavelength at which the substance has a second absorption coefficient which is substantially equal to zero. At least a portion of the measurement light beam retro-reflects from the retina (17) and at least a portion of the reference light beam retro-reflects from the retina (17). The apparatus further includes a detector (9) responsive to light having the first wavelength by generating a measurement signal and responsive to light having the second wavelength by generating a reference signal. | 1. A method of determining the concentration of a substance in an eye, the eye having a cornea, a pupil, an iris, a lens, a liquid, and a retina, the method comprising: providing a nearly parallel measurement light beam having a first wavelength at which the substance has a non-zero first absorption coefficient;providing a nearly parallel reference light beam having a second wavelength at which the substance has a second absorption coefficient which is substantially equal to zero;irradiating the retina with the measurement light beam, thereby passing the measurement light beam through the cornea, the pupil, the lens, and the liquid;irradiating the retina with the reference light beam, thereby passing the reference light beam through the cornea, the pupil, the lens, and the liquid;reflecting at least a portion of the measurement light beam from the retina and through the liquid, the lens, the pupil, and the cornea, thereby producing a measurement retro-reflected light beam having the first wavelength;reflecting at least a portion of the reference light beam from the retina and through the liquid, the lens, the pupil, and the cornea, thereby producing a reference retro-reflected light beam having the second wavelength;providing a detector adapted to generate a measurement signal in response to being irradiated by light having the first wavelength and to generate a reference signal in response to being irradiated by light having the second wavelength, the detector located at a distance of at least 100 mm from the eye;irradiating the detector with the measurement retro-reflected light beam;irradiating the detector with the reference retro-reflected light beam; anddetermining the concentration of the substance in the eye in response to the measurement signal and the reference signal from the detector. | 1. A method of determining the concentration of a substance in an eye, the eye having a cornea, a pupil, an iris, a lens, a liquid, and a retina, the method comprising: providing a nearly parallel measurement light beam having a first wavelength at which the substance has a non-zero first absorption coefficient;providing a nearly parallel reference light beam having a second wavelength at which the substance has a second absorption coefficient which is substantially equal to zero;irradiating the retina with the measurement light beam, thereby passing the measurement light beam through the cornea, the pupil, the lens, and the liquid;irradiating the retina with the reference light beam, thereby passing the reference light beam through the cornea, the pupil, the lens, and the liquid;reflecting at least a portion of the measurement light beam from the retina and through the liquid, the lens, the pupil, and the cornea, thereby producing a measurement retro-reflected light beam having the first wavelength;reflecting at least a portion of the reference light beam from the retina and through the liquid, the lens, the pupil, and the cornea, thereby producing a reference retro-reflected light beam having the second wavelength;providing a detector adapted to generate a measurement signal in response to being irradiated by light having the first wavelength and to generate a reference signal in response to being irradiated by light having the second wavelength, the detector located at a distance of at least 100 mm from the eye;irradiating the detector with the measurement retro-reflected light beam;irradiating the detector with the reference retro-reflected light beam; anddetermining the concentration of the substance in the eye in response to the measurement signal and the reference signal from the detector. 2. The method of claim 1, wherein the substance comprises glucose. 3. The method of claim 1, wherein the measurement light beam is provided by a first infrared light source and the reference light beam is provided by a second infrared light source. 4. The method of claim 3, wherein the first infrared light source comprises a laser diode or an infrared light emitting diode. 5. The method of claim 3, wherein the second infrared light source comprises a laser diode or an infrared light emitting diode. 6. The method of claim 1, wherein the measurement light beam is provided by a wideband light source and the reference light beam is provided by the wideband light source. 7. The method of claim 6, wherein the measurement light beam is transmitted through a first band filter and the reference light beam is transmitted through a second band filter. 8. The method of claim 1, wherein the measurement light beam and the measurement retro-reflected light beam are substantially parallel, and the reference light beam and the reference retro-reflected light beam are substantially parallel. 9. The method of claim 1, wherein the measurement light beam is focused a... | GlucoVista, LLC | A61B005/00
US07062313 | B2 | 2006-06-13 | Evaluation of exercise stress level dependent parameter | 2002-09-27 | A method and arrangement for estimating a stress level dependent parameter on a person that measures a person's heart rate, calculates a derivative of the measured heart rate, and estimates a value for a stress level dependent parameter by using the calculated heart rate derivative. | 1. A method for estimating a value for an exercise stress level dependent parameter in a training session performed by a person, comprising: measuring the person's heart rate;calculating a derivative of the measured heart rate; andestimating the value for an exercise stress level dependent parameter using the calculated heart rate derivative, wherein when estimating the value of the exercise stress level dependent parameter, a first dependency between the exercise stress level dependent parameter and the heart rate is used when the derivative is positive, and a second dependency between the exercise stress level dependent parameter and the heart rate is used when the derivative is negative. | 1. A method for estimating a value for an exercise stress level dependent parameter in a training session performed by a person, comprising: measuring the person's heart rate;calculating a derivative of the measured heart rate; andestimating the value for an exercise stress level dependent parameter using the calculated heart rate derivative, wherein when estimating the value of the exercise stress level dependent parameter, a first dependency between the exercise stress level dependent parameter and the heart rate is used when the derivative is positive, and a second dependency between the exercise stress level dependent parameter and the heart rate is used when the derivative is negative. 2. The method of claim 1, wherein the exercise stress level dependent parameter is the energy consumption of the person. 3. The method of claim 1, wherein the first dependency and the second dependency form a hysteresis dependency, where the first dependency provides higher values for the exercise stress level dependent parameter than does the second dependency. 4. The method of claim 1, wherein the first dependency contains a group of dependencies, and the dependency is selected on the basis of the magnitude of the derivative. 5. The method of claim 1 wherein the second dependency contains a group of dependencies, and the dependency is selected on the basis of the magnitude of the derivative. 6. The method of claim 1, wherein the value given by the first dependency and/or the second dependency depends on the magnitude of the derivative. 7. The method of claim 1 wherein the value given by first dependency and/or the second dependency depends on the physical condition of the person. 8. The method of claim 7, wherein the physical condition is characterized by one or more of the following parameters: oxygen uptake, speed, performance in a reference exercise. 9. The method of claim 1, further comprising: adjusting the first dependency and/or the second dependency depending on one or more physiological parameters of the person. 10. The method of claim 7, wherein the physiological parameter is age, gender, weight or height. 11. The method of claim 1, wherein the first dependency and/or the second dependency is formed using a neural network. 12. The method of claim 1 wherein the heart rate information is measured using a heart rate monitor during exercise, and the value for the exercise stress level dependant parameter is estimated in the heart rate monitor, and the estimated energy consumption assessment is displayed on the display of the heart rate monitor. 13. An arrangement for estimating an exercise stress level dependent parameter on a person, comprising: means for measuring the person's heart rate;means for calculating a derivative of the measured heart rate; andmeans for estimating the value for an exercise stress level dependent parameter by using the calculated heart rate derivative, wherein the estimating means is configured to use a first depe... | Polar Electro Oy | A61B005/04
US07277748 | B2 | 2007-10-02 | Spatiotemporal pattern recognition for neurological event detection and prediction in an implantable device | 2002-09-13 | A system and method for detecting and predicting neurological events with an implantable device uses a relatively low-power central processing unit in connection with signal processing circuitry to identify features (including half waves) and calculate window-based characteristics (including line lengths and areas under the curve of the waveform) in one or more electrographic signals received from a patient's brain. The features and window-based characteristics are employed within the framework of a programmable finite state machine to identify patterns and sequences in and across the electrographic signals, facilitating early and reliable detection and prediction of complex spatiotemporal neurological events in real time, and enabling responsive action by the implantable device. | 1. A method for detecting a neurological condition, the method comprising: receiving at least one electrical signal with at least one electrode, where the electrode is in electrical communication with neural tissue;processing the at least one electrical signal with a detector; anddetecting the neurological condition with the detector, where the neurological condition is determined by a sequence of detections of individual neurological events, and where the detections occur in a predetermined order. | 1. A method for detecting a neurological condition, the method comprising: receiving at least one electrical signal with at least one electrode, where the electrode is in electrical communication with neural tissue;processing the at least one electrical signal with a detector; anddetecting the neurological condition with the detector, where the neurological condition is determined by a sequence of detections of individual neurological events, and where the detections occur in a predetermined order. 2. The method for detecting a neurological condition of claim 1, wherein the sequence comprises a plurality of waveform segments. 3. The method for detecting a neurological condition of claim 2, wherein the plurality of waveform segments includes a first segment and a second segment. 4. The method for detecting a neurological condition of claim 3, wherein the first segment is observed at a first time in a first location in the patient's brain. 5. The method for detecting a neurological condition of claim 4, wherein the second segment is observed at a second time in a second location in the patient's brain. 6. The method for detecting a neurological condition of claim 5, wherein the second time follows the first time. 7. The method for detecting a neurological condition of claim 5, wherein the first location is substantially the same as the second location. 8. The method for detecting a neurological condition of claim 5, wherein the first location is substantially separate from the second location. 9. The method for detecting a neurological condition of claim 1, wherein the sequence comprises a sequence of signal characteristics. 10. The method for detecting a neurological condition of claim 9, wherein the sequence of signal characteristics is encoded in a finite state machine. 11. The method for detecting a neurological condition of claim 10, wherein the finite state machine comprises a master state machine and at least one instance. 12. The method for detecting a neurological condition of claim 11, wherein the at least one instance comprises a current state variable and an overall life remaining variable. 13. The method for detecting a neurological condition of claim 12, wherein the at least one instance further comprises a state life remaining variable. 14. The method for detecting a neurological condition of claim 11, wherein the identifying step comprises the steps of: setting an instance of the finite state machine to a first state; detecting a waveform characteristic in the sequence of signal characteristics; transitioning the instance to a destination state. 15. The method for detecting a neurological condition of claim 14, wherein the instance includes a current state, and wherein the transitioning step comprises updating the current state from the first state to the destination state. 16. The method for detecting a neurological condition of claim 14, wherein the instance includes a current state, and wherein the transitioning st... | Neuropace, Inc. | A61B005/00
US07004948 | B1 | 2006-02-28 | Cranial sealing plug | 2002-01-18 | Various embodiments of a burr hole plug assembly offer significant improvements for allowing lead and/or cannula access through a burr hole drilled through a patient's skull in connection with a Deep Brain Stimulation system, and subsequent sealing of such burr hole. The various burr hole plug assemblies described: (a) accommodate various burr hole sizes and provide a secure fit in the burr hole; (b) accommodate various locations for lead positioning and adjustment; (c) allow the lead to remain in a static position when the burr hole plug assembly is placed; (d) protect the lead from fracture at the exit location of the plug; (e) remain flush with the skull to avoid skin erosion and to eliminate unsightly flange protrusion from the patients' skull; (f) adequately hold the lead in place over time; and (g) provide a selection of various types of burr hole plug assemblies and sizes for use by the implanting surgeon, thereby eliminating the need for surgeons to resort to custom plugs and plug assemblies made in the operating room specifically to fit a given patient. | 1. A cranial sealing plug assembly adapted for insertion into a burr hole having a size D2 made in a patient's skull, comprising: a burr hole plug ring insertable into the burr hole so as to be flush with the patients skull, the plug ring having an opening through its center which allows access into the burr hole; anda sealing member that seals the opening through the center of the plug ring;wherein the burr hole plug ring has a diameter D1 that is slightly larger than the size D2 of the burr hole, thereby creating a compression force that prevents movement of a lead passing through the sealing member; andwherein the compression force further stabilizes the burr hole plug assembly in the burr hole. | 1. A cranial sealing plug assembly adapted for insertion into a burr hole having a size D2 made in a patient's skull, comprising: a burr hole plug ring insertable into the burr hole so as to be flush with the patients skull, the plug ring having an opening through its center which allows access into the burr hole; anda sealing member that seals the opening through the center of the plug ring;wherein the burr hole plug ring has a diameter D1 that is slightly larger than the size D2 of the burr hole, thereby creating a compression force that prevents movement of a lead passing through the sealing member; andwherein the compression force further stabilizes the burr hole plug assembly in the burr hole. 2. The cranial sealing plug assembly of claim 1 wherein the sealing member comprises a septum. 3. The cranial sealing plug assembly of claim 1 wherein the burr hole plug ring includes at least one segment made from a compressible biocompatible material, and at least another segment made from a noncompressible biccompatible material. 4. The cranial sealing plug assembly of claim 3 wherein the there are at least two segments made from the compressible biocompatible material, and at least three segments made from the non-compressible biocompatible material. 5. The cranial sealing plug assembly of claim 3 wherein the compressible biocompatible material compresses an elastomer selected from the group consisting of: silicone rubber and polyurethane. 6. The cranial sealing plug assembly of claim 3 wherein the non-compressible biocompatible material is selected from the group consisting of; epoxy, polyurethane, and titanium, and alloy/compounds of epoxy, polyurethane, and titanium. 7. The cranial sealing plug assembly of claim 1 wherein the sealing member and burr hole plug ring have a radial slit therein for accommodating a lead or cannula passing therethrough. 8. The cranial sealing plug assembly of claim 1 further including at least one O-ring insert adapter having an outer diameter adapted to snugly fit within the burr hole, and an inner diameter adapted to receive the plug assembly. 9. The cranial sealing plug assembly of claim 8 further including a plurality of O-ring insert adapters, each having a different outer diameter adapted to snugly fit within respective burr holes of varying sizes, and each having an inner diameter adapted to receive the plug assembly, whereby the plug assembly may be inserted into burr holes of varying sizes. | Advanced Bionics Corporation | A61B019/00 A61B005/04 A61N001/00
US07020514 | B1 | 2006-03-28 | Method of and apparatus for detecting atrial fibrillation | 2002-11-27 | A method and apparatus to determine possible atrial fibrillation that includes detecting irregular pulse rhythms from a succession of time intervals each corresponding to a respective interval of time between successive pulse beats of a sequence of the pulse beats; analyzing the detected irregular pulse rhythms to make a determination of possible atrial fibrillation; indicating the possible atrial fibrillation from the determination. If a sphygmomanometer is used, a plurality of pulse beats may be detected with it including the sequence of pulse beats; and respiratory variation in systolic pressure of the pulse beats may be compensated by designating the succession of time intervals to exclude at least an initial one of the plurality of time intervals. | 1. A method of determining possible atrial fibrillation, said method comprising the steps of: (a) detecting irregular pulse rhythms from a succession of time intervals each corresponding to a respective interval of time between successive pulse beats of a sequence of the pulse beats;(b) analyzing the detected irregular pulse rhythms to make a determination of possible atrial fibrillation; and(c) indicating the possible atrial fibrillation from the determination. | 1. A method of determining possible atrial fibrillation, said method comprising the steps of: (a) detecting irregular pulse rhythms from a succession of time intervals each corresponding to a respective interval of time between successive pulse beats of a sequence of the pulse beats;(b) analyzing the detected irregular pulse rhythms to make a determination of possible atrial fibrillation; and(c) indicating the possible atrial fibrillation from the determination. 2. A method of claim 1, wherein the analyzing includes ascertaining a mean and a standard deviation of the succession of time intervals, the determination being based upon a quotient formed by dividing said standard deviation by said mean and comparing the quotient with a threshold value. 3. A method of claim 2, wherein the determination includes determining that the sequence of the pulse beats is irregular in response to the quotient being greater than or equal to the threshold value. 4. The method of claim 2, wherein the determination includes determining that the sequence of the pulse beats is regular in response to the quotient being less than the threshold value. 5. The method of claim 1, wherein the detecting of the irregular pulse rhythms is with an inflatable cuff, and configuring the inflatable cuff to monitor the pulse rates to thereby detect the irregular pulse rhythms. 6. The apparatus of claim 5, wherein the inflatable cuff includes a sphygmomanometer. 7. The method of claim 1, wherein the detecting of the irregular pulse rhythms is by monitoring changes in light transmitted through a body appendage of the individual. 8. The method of claim 2, further comprising the step of storing at least one of said plurality of time intervals, said mean, said standard deviation, said quotient and said threshold value. 9. The method of claim 2 wherein said threshold value is within the range of 0.05 to 0.10. 10. The method of claim 1, wherein the indicating includes outputting an indication that is indicative of the possible atrial fibrillation by outputting to a device selected from a group consisting of a printer, a display, an auditory signal generator and a vibration signal generator. 11. A method as in claim 1, further comprising: detecting a plurality of pulse beats including the sequence of pulse beats using a sphygmomanometer; andcompensating for respiratory variation in systolic pressure of the pulse beats by designating the succession of time intervals to exclude at least an initial one of the plurality of time intervals. 12. An apparatus for determining possible atrial fibrillation, comprising: a detector configured to detect irregular pulse rhythms from a succession of time intervals each corresponding to a respective interval of time between successive pulse beats of a sequence of the pulse beats;a processor configured to analyze the detected irregular pulse rhythms for making a determination of possible atrial fibrillation; andan indicator configured to indicate th... | - | A61B005/04
US07139601 | B2 | 2006-11-21 | Surgical navigation systems including reference and localization frames | 2001-04-12 | A system for use during a medical or surgical procedure on a body. The system generates an image representing the position of one or more body elements during the procedure using scans generated by a scanner prior or during the procedure. The image data set has reference points for each of the body elements, the reference points of a particular body element having a fixed spatial relation to the particular body element. The system includes an apparatus for identifying, during the procedure, the relative position of each of the reference points of each of the body elements to be displayed. The system also includes a processor for modifying the image data set according to the identified relative position of each of the reference points during the procedure, as identified by the identifying apparatus, said processor generating a displaced image data set representing the position of the body elements during the procedure. The system also includes a display utilizing the displaced image data set generated by the processor, illustrating the relative position of the body elements during the procedure. Methods relating to the system are also disclosed. Also disclosed are devices for use with a surgical navigation system having a sensor array which is in communication with the device to identify its position. The device may be a reference frame for attachment of a body part of the patient, such as a cranial reference arc frame for attachment to the head or a spine reference arc frame for attachment to the spine. The device may also be a localization frame for positioning an instrument relative to a body part, such as a localization biopsy guide frame for positioning a biopsy needle, a localization drill guide assembly for positioning a drill bit, a localization drill yoke assembly for positioning a drill, or a ventriculostomy probe for positioning a catheter. | 1. A method for determining the geometry of a semi-rigid body element during a procedure, the method using an image data set having reference points for the body element, a relative position between said reference points of the semi-rigid body element being variable, the method comprising: identifying, during the procedure, the relative position of each of the reference points for the semi-rigid body element;deriving a transform relating the relative position of the reference points during the procedure to the relative position of the reference points in the image data set in order to determine the geometry of the body element during the procedure, said transform indicative of a difference between the relative position of each of the reference points for the semi-rigid body element during the procedure and the relative position of the reference points in the image data set;modifying the image data set based on the transform in order to generate a displaced image data set representing the geometry of the body element during the procedure; andgenerating a display based on the displaced image data set illustrating the geometry of the body element during the procedure. | 1. A method for determining the geometry of a semi-rigid body element during a procedure, the method using an image data set having reference points for the body element, a relative position between said reference points of the semi-rigid body element being variable, the method comprising: identifying, during the procedure, the relative position of each of the reference points for the semi-rigid body element;deriving a transform relating the relative position of the reference points during the procedure to the relative position of the reference points in the image data set in order to determine the geometry of the body element during the procedure, said transform indicative of a difference between the relative position of each of the reference points for the semi-rigid body element during the procedure and the relative position of the reference points in the image data set;modifying the image data set based on the transform in order to generate a displaced image data set representing the geometry of the body element during the procedure; andgenerating a display based on the displaced image data set illustrating the geometry of the body element during the procedure. 2. The method of claim 1 wherein the semi-rigid body element is soft tissue. 3. The method of claim 1 wherein deriving comprises deriving a transformation which allows the determination of the procedural position, orientation, and shape in surgical space of the semi-rigid body element, and wherein modifying comprises modifying the image data set according to said transfonnation to produce a displaced image data set reflecting changes in the geometry of the semi-rigid body elements during the procedure. 4. A method for determining the geometry and position of a semi-rigid body element during a procedure, the method using a system including an array, and instrument in communication with the array, and a processor in communication with the array and storing an image data set having reference points for the body element, a relative position between said reference points of the semi-rigid body element being variable, the method comprising: touching the reference points for the semi-rigid body element with the instrument during the procedure and communicating the position of the reference points to the array;communicating the position of the reference points of the body element during the procedure to the processor;determining the relative position of the reference points;deriving a transform of the relative position of the reference points during the procedure to the relative position of the reference points in the image data set in order to determine the geometry and position of the body element during the procedure, said transform indicative of a difference between the relative position of each of the reference points for the semi-rigid body element during the procedure and the relative position of the reference points in the image data set;modifying the image data set based on the tra... | Surgical Navigation Technologies, Inc. St. Louis University | A61B005/05
US07112197 | B2 | 2006-09-26 | Surgical device with pressure monitoring ability | 2003-01-21 | A device with a functional tip containing at least one active electrode capable of creating a controlled perforation in body tissue through the application of Radio Frequency (RF) energy is described. The position of the tip of the device can be determined in response to pressure sensed at the tip and determined by a monitor. The device is useful to remove or perforate unwanted tissue in a controlled manner in any location in the body, particularly in the atrial septum for controlled transseptal puncture. In this application, the device is introduced into the right atrium, and the functional tip is then positioned against the atrial septum. Energy is applied to create the perforation and pressure is monitored to determine if the perforation was created in a desired location. Other possible applications include the removal of plaque or thrombotic occlusions from diseased vessels. | 1. A surgical device for perforating material and monitoring pressure comprising: an elongate member having a distal region with a first outer diameter and a proximal region with a second outer diameter larger than said first outer diameter such that dilation of a perforation is limited while said distal region is advanced through said perforation, wherein said second outer diameter tapers down to said first outer diameter so as to provide a smooth transition between said first outer diameter and said second outer diameter;an energy delivery device associate with said elongate member at said distal region, said energy delivery device adapted for connection to an energy source and operable to effectively perforate said material by delivering energy; anda pressure sensing mechanism associated with said distal region for monitoring pressure about said distal region;wherein said pressure sensing mechanism comprises a pressure transmitting lumen extending between said proximal and distal regions, the lumen at said proximal region being adapted for fluid communication with a pressure transducer that provides a signal which varies as a function of pressure and adapted at said distal region for fluid communication with an environment about said distal region. | 1. A surgical device for perforating material and monitoring pressure comprising: an elongate member having a distal region with a first outer diameter and a proximal region with a second outer diameter larger than said first outer diameter such that dilation of a perforation is limited while said distal region is advanced through said perforation, wherein said second outer diameter tapers down to said first outer diameter so as to provide a smooth transition between said first outer diameter and said second outer diameter;an energy delivery device associate with said elongate member at said distal region, said energy delivery device adapted for connection to an energy source and operable to effectively perforate said material by delivering energy; anda pressure sensing mechanism associated with said distal region for monitoring pressure about said distal region;wherein said pressure sensing mechanism comprises a pressure transmitting lumen extending between said proximal and distal regions, the lumen at said proximal region being adapted for fluid communication with a pressure transducer that provides a signal which varies as a function of pressure and adapted at said distal region for fluid communication with an environment about said distal region. 2. The device as claimed in claim 1, wherein said energy is at least one form of energy selected from a group consisting of: electrical current; microwave; ultrasound; and laser. 3. The device as claimed in claim 2 wherein said electrical current has a frequency within the radio frequency range. 4. The device as claimed in claim 1 wherein said material comprises cellular tissue and wherein said energy delivery device is operable to deliver sufficient energy to the tissue to result in a rapid increase in the intracellular temperature causing vaporization of intracellular water and subsequent cell lysis. 5. The device as claimed in claim 1 wherein said distal region comprises at least one opening to the environment and wherein said lumen is in fluid communication with said at least one opening. 6. The device as claimed in claim 1 wherein said energy delivery device comprises a functional tip with at least one active electrode. 7. The device as claimed in claim 6 wherein said at least one electrode comprises an active electrode with an outer diameter of substantially 0.04 cm. 8. The device as claimed in claim 1 wherein said energy delivery device comprises a functional tip having two or more electrodes. 9. The device as claimed in claim 8 wherein the electrodes are configured in an arrangement where at least one of said electrodes is active and at least one is a return electrode. 10. The device as claimed in claim 1 wherein said surgical device comprises a radiopaque marker. 11. A method of surgery comprising the steps of: (i) introducing a surgical device into a body of a patient, said surgical device comprising a flexible elongate member having a distal region and a proximal region, an ener... | Baylis Medical Company Inc. | A61B018/18 A61B005/02
US06936011 | B2 | 2005-08-30 | Analysis of sleep apnea | 2003-01-22 | A non-intrusive and quantitative method and apparatus for diagnosing sleep apnea and detecting apnea events by monitoring during sleep abdominal effort and thoracic effort, determining the phase of each effort, determining the difference in phase between each type of effort, and then determining the rate of phase angle change and standard deviation over time. Also provided may be treatment when apnea events are detected to trigger therapy apparatus such as airway positive pressure apparatus. | 1. A method for diagnosing sleep apnea comprising the steps of measuring thoracic volume of a patient and determining the phase of the change of thoracic volume of said patient, measuring abdominal volume of said patient as a function of time and determining the phase of the change of abdominal volume of said patient, calculating the phase angles between said thoracic volume phase and said abdominal volume phase, and determining the standard deviation of the phase angle between the two respiratory signals for the diagnosis. | 1. A method for diagnosing sleep apnea comprising the steps of measuring thoracic volume of a patient and determining the phase of the change of thoracic volume of said patient, measuring abdominal volume of said patient as a function of time and determining the phase of the change of abdominal volume of said patient, calculating the phase angles between said thoracic volume phase and said abdominal volume phase, and determining the standard deviation of the phase angle between the two respiratory signals for the diagnosis. 2. Apparatus for diagnosing sleep apnea comprising means for measuring abdominal effort in a patient, means for measuring thoracic effort in a patient, means for determining phase angle between two phase measurements, means for determining change in the phase angle, and determining standard deviation of the phase angle. 3. Apparatus according to claim 2 and comprising a wearable garment supporting the means for detecting abdominal effort and thoracic effort. 4. The apparatus of claim 3 in which the means for measuring thoracic and abdominal effort comprise piezo belts, PDF belts, inductance or impedance measurement devices and the like. 5. Apparatus according to claim 2 and further comprising a wearable garment to which the detecting means are detachably affixed and positioned to measure the respective efforts when the garment is worn. 6. Apparatus for detecting apnea events comprising means for detecting abdominal effort and thoracic effort, means for measuring the detected abdominal effort, means for measuring the detected thoracic effort, means for gathering and storing data from said measured abdominal effort and said measured thoracic effort measurements wherein the gathered data includes measurements of the phase angles of the abdominal effort and thoracic effort, means for determining the standard deviation between the phase angles, and means for transferring said data. 7. Apparatus for detecting apnea events comprising means for detecting abdominal effort and thoracic effort, means for measuring the detected abdominal effort, means for measuring the detected thoracic effort, means for gathering and storing data from said measured abdominal effort and said measured thoracic effort measurements wherein the gathered data includes measurements of the phase angles of the abdominal effort and thoracic effort, means for determining the standard deviation between the phase angles, means for transferring said data, and a wearable garment to which the detecting means are detachably affixed and positioned to measure the respective efforts when the garment is worn. | Medcare Flags hf | A61B005/00
US06922585 | B2 | 2005-07-26 | Method and apparatus for predicting recurring ventricular arrhythmias | 2002-04-05 | An implantable medical device and method are provided for assessing autonomic tone and risk factors associated with arrhythmias and, based on this assessment, an early recurrence of ventricular tachycardia or ventricular fibrillation is predicted. Specifically, changes in R—R interval, heart rate variability, patient activity, and myocardial ischemia are measured prior to and after a detected an arrhythmia episode. A recurrence score is calculated as a weighted sum of measured parameters and compared to a prediction criterion. The prediction criterion may be a preset threshold score or an individualized episode template based on previously calculated recurrence scores associated with recurring episodes. Stored parameters and episode-related data may be downloaded for offline analyses for optimizing prediction criteria and monitoring patient status. | 1. A method for predicting a recurring arrhythmia comprising: detecting an arrhythmia episode; measuring a number of parameters based on said arrhythmia episode indicative of autonomic tone; determining whether the episode is a recurring episode; determining changes in the parameters prior to onset of the episode and subsequent to termination of the episode; calculating a recurrence score in response to the determined changes; and predicting an arrhythmia event with high probabilistic certainty based on said recurrence score. | 1. A method for predicting a recurring arrhythmia comprising: detecting an arrhythmia episode; measuring a number of parameters based on said arrhythmia episode indicative of autonomic tone; determining whether the episode is a recurring episode; determining changes in the parameters prior to onset of the episode and subsequent to termination of the episode; calculating a recurrence score in response to the determined changes; and predicting an arrhythmia event with high probabilistic certainty based on said recurrence score. 2. The method of claim 1 wherein said method of detecting includes measuring R—R intervals and P—P intervals. 3. The method of claim 1 wherein said method of measuring includes collecting R—R interval templates. 4. The method of claim 1 wherein said method includes collecting a heart rate variability (HRV) template. 5. The method of claim 1 wherein said method includes collecting an activity template. 6. The method according to claim 5 wherein a correlation between a physical activity and an R—R interval is included in said activity template. 7. The method of claim 1 wherein said collecting method includes determining indicators of autonomic tone and risk factors for VT and VF recurrence. 8. The method of claim 1 wherein said method of predicting includes comparing said recurrence score with a threshold, a VT or VF episode. 9. The method of claim 1 wherein said method of predicting includes using a range of value benchmarks to predict VT or VF episodes if the recurrence score falls within said range of value benchmarks. 10. The method of claim 1, wherein determining changes comprises: determining changes in RR intervals at onset of the episode and subsequent to termination of the episode; determining changes in activity levels at onset of the episode and subsequent to termination of the episode; and determining changes in heart rate variability at onset of the episode and subsequent to termination of the episode. 11. The method of claim 10, wherein the recurrence score is a weighted sum of the determined changes. 12. A medical device for predicting a recurring arrhythmia in a patient, comprising: a first sensor sensing cardiac signals; a second sensor sensing activity of the patient; and a microprocessor connected to the sensors determining whether an arrhythmia episode has occurred and whether the arrhythmia episode is a recurring episode, and determining parameters based on the arrhythmia episode indicative of autonomic tone, wherein the microprocessor determines changes In the parameters prior to onset of the episode and subsequent to termination of the episode, calculates a recurrence score in response to the determined changes, and predicts an arrhythmia event in response to the recurrence score. 13. The device of claim 12, wherein the microprocessor determines changes in RR intervals at onset of the episode and subsequent to termination of the episode, determines changes in activity levels at onset of th... | Medtronic, Inc. | A61B005/046
US06915151 | B2 | 2005-07-05 | Quantitative pulmonary imaging | 2002-02-08 | The present invention comprises imaging and quantitative measurement of lung ventilation, particularly in a human lung. Methods for quantitative imaging of lung ventilation, and the further provided systems and algorithmic tools therefore, comprise three primary components: the combined MRI ventilation/perfusion (V/Q) imaging techniques using hyperpolarized helium-3 (3He) gas (H3He); the three-dimensional quantitative imaging of absolute lung perfusion (Q) and collection of local magnetic resonance image data therefrom to produce an absolute lung perfusion image data; and the algorithmic co-registration of the two image data sets, (HP-3He MRI image of V/Q and MR imaging of quantitative perfusion (Q) in the lung). From the data acquired in the combined data sets and their spatial co-registration, absolute ventilation (V) is computed. | 1. A quantitative method for very high resolution, three-dimensional combined imaging of pulmonary ventilation (V) and perfusion (Q), synchronously (V/Q), wherein the method comprises delivering a predetermined volume of hyperpolarized noble gas into the conducting airways in each ventilated region of the pulmonary system, collecting local magnetic resonance image data from said hyperpolarized noble gas to produce a V/Q image data set, three-dimensionally, quantitatively imaging absolute lung perfusion (Q), and collecting local magnetic resonance image data from said perfusion imaging to produce an absolute lung perfusion image data set. | 1. A quantitative method for very high resolution, three-dimensional combined imaging of pulmonary ventilation (V) and perfusion (Q), synchronously (V/Q), wherein the method comprises delivering a predetermined volume of hyperpolarized noble gas into the conducting airways in each ventilated region of the pulmonary system, collecting local magnetic resonance image data from said hyperpolarized noble gas to produce a V/Q image data set, three-dimensionally, quantitatively imaging absolute lung perfusion (Q), and collecting local magnetic resonance image data from said perfusion imaging to produce an absolute lung perfusion image data set. 2. The method of claim 1, wherein the noble gas is hyperpolarized helium-3 gas (HP-3 He). 3. The method of claim 1, wherein the imaging of absolute lung perfusion (Q) comprises a gadolinium based method. 4. The method of claim 1, wherein the imaging of absolute lung perfusion (Q) comprises an arterial spin-tagging method. 5. The method of claim 1, further comprising the step of using a co-registration algorithm to spatially co-register the MR image V/Q data set with data set Q, comprising the quantitative lung perfusion. 6. The method of claim 5, further comprising dividing the image into as many distinct voxels as imaging resolution permits. 7. The method of claim 6, further comprising computing absolute ventilation (V) by multiplying the ventilation/perfusion ratio (V/Q) by perfusion (Q) for each point in the lung image. 8. The method of claim 7, further comprising applying the method to the pulmonary system of a mammalian subject. 9. The method of claim 8, wherein the mammalian subject is human. 10. The method of claim 9, wherein the lung is normal. 11. The method of claim 9, wherein the lung is injured or diseased. 12. The very high resolution three-dimensional images of pulmonary ventilation, produced by the method of claim 7 using hyperpolarized HP-3He gas. | Trustees of the University of Pennsylvania | A61B005/055
US06934578 | B2 | 2005-08-23 | Method and apparatus for interactive annotation and measurement of time series data with automatic marker sequencing | 2002-09-18 | A system displays time-series data, such as electrocardiographic data. The data may be displayed as a trace with markers identifying data features. The markers may be automatically sequenced by the system. | 1. A method of marking time-series medical data that are graphically displayed in the form of a set of waveform traces on a display monitor, each trace embodying a plurality of features, the method comprising the steps of: A. manually identifying a location of at least one feature of a first trace; B. displaying on the monitor in relation to the identified first trace feature location at least one marker; C. displaying on the monitor at least one label corresponding to the marker and the trace feature; and D. automatically displaying on the monitor markers and labels associated with the one feature on subsequent waveform traces. | 1. A method of marking time-series medical data that are graphically displayed in the form of a set of waveform traces on a display monitor, each trace embodying a plurality of features, the method comprising the steps of: A. manually identifying a location of at least one feature of a first trace; B. displaying on the monitor in relation to the identified first trace feature location at least one marker; C. displaying on the monitor at least one label corresponding to the marker and the trace feature; and D. automatically displaying on the monitor markers and labels associated with the one feature on subsequent waveform traces. 2. The method of claim 1 wherein the time series medical data is electrocardiogram (ECG) data and and step A comprises indicating a type of ECG sequence for which a trace feature is to be located. 3. The method of claim 2, wherein said type of sequence is a PQJT sequence. 4. The method of claim 3, wherein step C comprises displaying a label “P” in proximity to a corresponding first marker in the PQJT sequence. 5. The method of claim 2, wherein step A comprises inputting data to a graphical user interface. 6. The method of claim 2, wherein said type of seuence is a QPJT sequence. 7. The method of claim 6, wherein step C comprises displaying a label “Q” in proximity to a corresponding first marker in the QPJT sequence. 8. The method of claim 2, wherein step A comprises indicating that a QRS onset sequence is to be labeled. 9. The method of claim 8, wherein step C comprises displaying a Q onset label in proximity to a corresponding first marker in the QRS onset sequence. 10. The method of claim 9, further comprising the steps of: displaying on the monitor a second marker in relation to a P onset; and displaying a P onset label corresponding to the second marker. 11. The method of claim 1, further comprising the step of displaying a reference line between the at least one marker and the label. 12. The method of claim 1 further comprising the step of displaying the time series medical data as a waveform trace superimposed on a gridded background. 13. A method of marking time-series electrocardiogram (ECG) data comprising the steps of: A. displaying on a display monitor the time series ECG data as a waveform trace superimposed on a gridded background; B. accepting user input that indicates a type of (ECG) sequence for which a feature is to be located; C. locating a displayed trace feature of the indicated ECG sequence on the monitor; D. in response to said user input, placing a first marker on the displayed waveform trace in relation to the located displayed trace feature; E. displaying a label on the displayed waveform trace that identifies the first marker; and F. automatically displaying on the monitor markers and labels associated with the one feature on subsequent waveform traces. 14. The method of claim 13 further comprising the step of displaying the gridded background as rectangular cells that delimit 40... | Covance Inc. | A61B005/04
US06955651 | B2 | 2005-10-18 | ALGORITHMS, SYSTEMS, AND METHODS FOR ESTIMATING CARBON DIOXIDE STORES, TRANSFORMING RESPIRATORY GAS MEASUREMENTS, AND OBTAINING ACCURATE NONINVASIVE PULMONARY CAPILLARY BLOOD FLOW AND CARDIAC OUTPUT MEASUREMENTS | 2002-04-11 | Methods for estimating the volume of the carbon dioxide stores of an individual's respiratory tract include determining a carbon dioxide store volume at which a correlation between corresponding signals of carbon dioxide elimination and an indicator of the content of carbon dioxide in blood of the individual is optimized. The estimate of the volume of carbon dioxide stores, which comprises a model of the respiratory tract, or lungs, of the individual, may be used as a transformation to improve the accuracy of one or both of the carbon dioxide elimination and carbon dioxide content signals. Transformation, or filtering, algorithms are also disclosed, as are systems in which the methods and algorithms may be used. The methods, algorithms, and systems may be used to accurately and noninvasively determine one or both of the pulmonary capillary blood flow and cardiac output of the individual. | 1. A method for optimizing the accuracy of at least one of a carbon dioxide elimination measurement and a carbon dioxide excretion signal obtained from an individual based on carbon dioxide stores of a respiratory tract of the individual, comprising use of the formula: {dot over({circumflex over (V)})}BCO2(n)=(1−α){dot over (V)}MCO2(n)+α{dot over({circumflex over (V)})}BCO2(n−1), where {dot over (V)}MCO2(n) is the carbon dioxide elimination measurement, {dot over({circumflex over (V)})}BCO2(n−1) is an estimate of carbon dioxide excretion of at least one preceding or subsequent breath, {dot over({circumflex over (V)})}BCO2(n) is an estimate of carbon dioxide excretion of the same breath as that for which the carbon dioxide elimination measurement was obtained, and α is a transformation coefficient based on an estimate of the carbon dioxide stores of the respiratory tract of the individual. | 1. A method for optimizing the accuracy of at least one of a carbon dioxide elimination measurement and a carbon dioxide excretion signal obtained from an individual based on carbon dioxide stores of a respiratory tract of the individual, comprising use of the formula: {dot over({circumflex over (V)})}BCO2(n)=(1−α){dot over (V)}MCO2(n)+α{dot over({circumflex over (V)})}BCO2(n−1), where {dot over (V)}MCO2(n) is the carbon dioxide elimination measurement, {dot over({circumflex over (V)})}BCO2(n−1) is an estimate of carbon dioxide excretion of at least one preceding or subsequent breath, {dot over({circumflex over (V)})}BCO2(n) is an estimate of carbon dioxide excretion of the same breath as that for which the carbon dioxide elimination measurement was obtained, and α is a transformation coefficient based on an estimate of the carbon dioxide stores of the respiratory tract of the individual. 2. The method of claim 1, comprising estimating carbon dioxide excretion of the at least one preceding or subsequent breath ({dot over({circumflex over (V)})}BCO2(n)) as follows: {dot over({circumflex over (V)})}BCO2(n)={dot over (V)}MCO2(n)+{circumflex over (V)}A*(n)[fACO2(n)−fACO2(n−1)]RR, where {dot over (V)}MCO2(n) is the carbon dioxide elimination measurement for the at least one preceding or subsequent breath, {circumflex over (V)}A*(n) is an estimate of a volume of the carbon dioxide stores of the respiratory tract of the individual during the at least one preceding or subsequent breath, [fACO2(n)−fACO2(n−1)] is a difference between a fraction of carbon dioxide in alveoli of the respiratory tract of the individual for the at least one preceding or subsequent breath and a breath that immediately preceded the at least one preceding or subsequent breath, and RR is a respiratory rate of the individual. 3. The method of claim 1, comprising determining an optimal α value by at least one of iterative searching, rote searching, gradient searching, use of a set of predetermined equations, and adaptive filtering. 4. The method of claim 1, comprising adjusting the estimate of carbon dioxide excretion of the at least one preceding or subsequent breath to improve a correlation between the carbon dioxide elimination measurement and an indicator of the content of carbon dioxide in blood of the individual. 5. A method for improving an accuracy of a respiratory signal, comprising: obtaining the respiratory signal from respiration of an individual during at least two breaths; and applying a transformation based on carbon dioxide stores of a respiratory tract of the individual including carbon dioxide in lung tissues, to the respiratory signal; wherein applying the transformation comprises employing a carbon dioxide elimination signal in the following algorithm: {dot over({circumflex over (V)})}BCO2(n)=(1−α){dot over (V)}MCO2(n)+α{dot over({circumflex over (V)})}BCO2(n−1), where {dot over (V)}MCO2(n) is a carbon dioxide elimination measurement, {dot over({circumflex o... | Respironics, Inc. | A61B005/00
US07194295 | B2 | 2007-03-20 | Medical navigation and/or pre-operative treatment planning with the assistance of generic patient data | 2002-04-26 | The invention relates to a method for computer-assisted medical navigation and/or pre-operative treatment planning, wherein the current position of a patient or a part of a patient's body and the positions of medical treatment devices or treatment-assisting devices are detected by means of a position detection unit, and wherein said detected positional data are assigned to body structure data, in order to jointly use said body structure data in assignment with said positional data, within the context of assisting the treatment, wherein body structure data are used which are obtained based on a generic model which has been adapted by linking it with patient-characteristic detection data. | 1. A method for computer-assisted medical navigation and/or pre-operative treatment planning, comprising: linking a generic model with patient-characteristic data to obtain body structure data;detecting a current position of a patient or a part of a patient's body and positions of medical treatment devices or treatment-assisting devices by a position detection unit, wherein said detected positional data are assigned to the body structure data in order to jointly use said body structure data in assignment with said positional data within the context of assisting the treatment;obtaining additional points on the patient in the form of landmarks or randomly acquired points; andusing the additional points to further correlate the body structure data to the patient or part of the patient's body. | 1. A method for computer-assisted medical navigation and/or pre-operative treatment planning, comprising: linking a generic model with patient-characteristic data to obtain body structure data;detecting a current position of a patient or a part of a patient's body and positions of medical treatment devices or treatment-assisting devices by a position detection unit, wherein said detected positional data are assigned to the body structure data in order to jointly use said body structure data in assignment with said positional data within the context of assisting the treatment;obtaining additional points on the patient in the form of landmarks or randomly acquired points; andusing the additional points to further correlate the body structure data to the patient or part of the patient's body. 2. The method as set forth in claim 1, wherein said body structure data are provided in the form of image data. 3. The method as set forth in claim 2, wherein the image data set is a tomographic image data set. 4. The method as set forth in claim 1, wherein said generic model comprises a typical or average body structure. 5. The method as set forth in claim 1, wherein said generic model includes a statistical model of said body structure. 6. The method as set forth in claim 5, wherein the statistical model is based on statistical evaluations of a number of image data sets of the body structure. 7. The method as set forth in claim 1, wherein said generic model includes a multitude of body structures of the same type. 8. The method as set forth in claim 1, wherein said generic model comprises a two- or three-dimensional data set of a body structure, and/or a geometric model. 9. The method as set forth in claim 1, wherein said positional data obtained while determining said patient-characteristic data can be used to register said adapted body structure data in said navigation system and to visually display or use treatment devices and/or treatment-assisting devices in their registration to said adapted body structure. 10. The method as set forth in claim 9, wherein said patient-characteristic detection data includes acquiring landmark positions or x-ray imaging registered in said navigation system. 11. The method as set forth in claim 1, wherein the generic model is based on data corresponding to a weight of the patient or the part of the patient's body. 12. The method as set forth in claim 1, wherein linking includes surface matching the generic model to the patient-characteristic data. 13. A method for computer-assisted medical navigation and/or pre-operative treatment planning, comprising: detecting a current position of a patient or a part of a patient's body and positions of medical treatment devices or treatment-assisting devices by a position detection unit, wherein said detected positional data are assigned to body structure data in order to jointly use said body structure data in assignment with said positional data within the context of assis... | BrainLAB AG | A61B005/00
US07239730 | B2 | 2007-07-03 | Method and apparatus for volume scoring calcification concentrations of a CT scan | 2003-01-29 | A method for volume scoring a concentration of calcification within a region of interest includes receiving a value for a first calcium lesion volume based on the concentration of calcification within the region of interest, determining an intensity value of the concentration of calcification within the region of interest, calculating the slope of a regression line based on the determined intensity value, and calculating a second calcium lesion volume based on the first calcium lesion volume and the calculated slope of the regression line. | 1. A method for volume scoring a concentration of calcification within a region of interest, comprising: receiving a value for a first calcium lesion volume based on the concentration of calcification within the region of interest;determining an intensity value of the concentration of calcification within the region of interest;calculating the slope of a regression line based on the determined intensity value; andcalculating a second calcium lesion volume based on the first calcium lesion volume and the calculated slope of the regression line. | 1. A method for volume scoring a concentration of calcification within a region of interest, comprising: receiving a value for a first calcium lesion volume based on the concentration of calcification within the region of interest;determining an intensity value of the concentration of calcification within the region of interest;calculating the slope of a regression line based on the determined intensity value; andcalculating a second calcium lesion volume based on the first calcium lesion volume and the calculated slope of the regression line. 2. The method for volume scoring a concentration of calcification set forth in claim 1, wherein said receiving a value for a first calcium lesion volume comprises: acquiring projection data representing the calcium concentration of a CT scan of an object;defining a calcified voxel within the region of interest based on the acquired projection data;determining a slice thickness of the acquired projection data;calculating a calcified voxel volume based on the pixel area of the calcified voxel and the slice thickness of the acquired projection data; andcalculating a first calcium lesion volume based on the calculated calcified voxel volume. 3. The method for volume scoring set forth in claim 1, wherein said determining an intensity value further comprises: determining a mean intensity value. 4. The method for volume scoring set forth in claim 1, wherein said calculating a second calcium lesion volume further comprises: calculating a second calcium lesion volume having a score accuracy error equal to or greater than about −10% and equal to or less than about +10%. 5. The method for volume scoring set forth in claim 2, wherein said determining a slice thickness further comprises: determining a slice thickness equal to or greater than about 1.25 millimeters and equal to or less than about 5 millimeters. 6. The method for volume scoring set forth in claim 5, wherein said determining a slice thickness further comprises: determining a slice thickness equal to or greater than about 1.25 millimeters and equal to or less than about 2.5 millimeters. 7. The method for volume scoring set forth in claim 2, wherein said defining a calcified voxel within the region of interest further comprises: defining a calcified voxel within the region of interest based on a signal attenuation threshold and a pixel connectivity value. 8. The method for volume scoring set forth in claim 7, wherein said defining a calcified voxel within the region of interest further comprises: defining a calcified voxel within the region of interest based on a signal attenuation of greater than about 130 Hounsfield Units and a pixel connectivity value of 2. 9. The method for volume scoring set forth in claim 2, wherein said determining a slice thickness further comprises; adjusting the slice thickness by the slice spacing for overlapping slices. 10. A method for volume scoring a concentration of calcification within a region of interest, comprisi... | GE Medical Systems Global Technology Company, LLC | G06K009/00 A61B006/00 A61B005/05
US07182754 | B2 | 2007-02-27 | Containerless tissue sample collection trap | 2002-06-19 | A containerless tissue sample collection trap featuring, in the preferred embodiment, a flexible filter and an applicator having a shaft terminating in a distal end with a channel therein receiving a portion of the flexible filter, the combination of the applicator and the flexible filter receivable in a suction tube and configured such that when the applicator is removed, the flexible filter forms a removable elongated tissue trap in the suction tube. Also disclosed is a method of trapping tissue wholly within a suction tube thereby eliminating the need for a specimen trap container and a method of making a containerless tissue specimen trap. | 1. A containerless tissue sample collection trap comprising: a flexible netting; andan applicator having a shaft terminating in a distal end with a tubular channel therein continuing within said shaft, said channel receiving a portion of the flexible netting, the combination of the applicator and the flexible netting receivable in a suction tube; andsaid applicator and said flexible netting configured such that when the applicator is removed, a pocket in the flexible netting forms a removable elongated tissue trap in the suction tube. | 1. A containerless tissue sample collection trap comprising: a flexible netting; andan applicator having a shaft terminating in a distal end with a tubular channel therein continuing within said shaft, said channel receiving a portion of the flexible netting, the combination of the applicator and the flexible netting receivable in a suction tube; andsaid applicator and said flexible netting configured such that when the applicator is removed, a pocket in the flexible netting forms a removable elongated tissue trap in the suction tube. 2. The containerless tissue sample collection trap of claim 1 in which the flexible netting is made of a single ply between 2 inches and 5 inches in diameter. 3. The containerless tissue sample collection trap of claim 2 in which the netting is made of a polymer material. 4. The containerless tissue sample collection trap of claim 2 in which the netting has pores between 0.01 and 0.05 inches in perimeter. 5. The containerless tissue sample collection trap of claim 2 in which the netting has between 1,000 and 2,000 pores per square inch. 6. The containerless tissue sample collection trap of claim 1 in which the applicator is made of plastic. 7. The containerless tissue sample collection trap of claim 1 in which the applicator has a broad proximal end opposite the distal end for easier insertion of the distal end thereof into the suction tube. 8. The containerless tissue sample collection trap of claim 7 in which the distal end of the applicator is flat and the channel extends through the flat distal end and continues within the shaft. 9. The containerless tissue sample collection trap of claim 8 in which the outer diameter of the distal end is between 0.4 and 0.5 inches. 10. The containerless tissue sample collection trap of claim 8 in which the inner diameter of the distal end is between 0.1 and 0.2 inches. 11. The containerless tissue sample collection trap of claim 8 in which the shaft is between 1 and 2 inches in length. 12. The containerless tissue sample collection trap of claim 8 in which the channel is between 0.7 and 0.8 inches in length. 13. A containerless tissue sample collection trap comprising: a flexible netting; andan applicator having a shaft terminating in a distal end and a tubular channel therein extending from the distal end and continuing within said shaft, with a portion of the flexible netting received in the channel, the distal end for urging at least a portion of the flexible netting received in the channel, the distal end for urging at least a portion of the flexible netting into a suction tube, said applicator configured to be removable from said suction tube prior to operation thereof, the combination of the applicator and the flexible netting receivable in a suction tube and configured such that the flexible netting forms a removable elongated tissue trap in the suction tube. 14. A method of trapping tissue wholly within a suction tube thereby eliminating the need for a tiss... | N.M. Beale Company | B01D029/00 B01D039/10 A61B010/00 A61B005/00 A61M001/00 A61M031/00
US07031764 | B2 | 2006-04-18 | Cardiac rhythm management systems and methods using multiple morphology templates for discriminating between rhythms | 2002-11-08 | Multiple morphology templates for discliminating between rhythms have been used, such as supraventricular tachyarrhythmias (SVTs) and ventricular tachyarrhythmias (VTs), for delivering a countershock in response to a VT episode, but withholding delivery of such a countershock in response to an SVT episode. In certain examples, the particular morphology used for storing morphological features is selected at least in part using a sensor-indicated activity level of a subject, or a metabolic need of the subject. | 1. A system comprising: an electrogram sensing circuit, configured to be coupled to at least one electrode for receiving an intrinsic cardiac signal, the electrogram sensing circuit configured to provide cardiac complexes and an indication of heart rate obtained from a subject's heart;a stored first morphological template, coupled to the electrogram sensing circuit, the stored first morphological template obtained from the heart, in the absence of an arrhythmia, while the subject is resting or inactive;a stored second morphological template, coupled to the electrogram sensing circuit, the stored second morphological template obtained from the heart under a condition different from that of the first morphological template;a correlation module, coupled to the first and second morphological templates and the electrogram sensing circuit, the correlation module configured to declare an indication of a first rhythm state if the indication of heart rate exceeds a first predetermined threshold value and a morphology of at least one cardiac complex of the intrinsic cardiac signal correlates to either of the first and second morphological templates, and configured to declare an indication of a second rhythm state if the indication of heart rate exceeds the first predetermined threshold value and a morphology of the at least one cardiac complex of the intrinsic cardiac signal is uncorrelated to both the first and second morphological templates. | 1. A system comprising: an electrogram sensing circuit, configured to be coupled to at least one electrode for receiving an intrinsic cardiac signal, the electrogram sensing circuit configured to provide cardiac complexes and an indication of heart rate obtained from a subject's heart;a stored first morphological template, coupled to the electrogram sensing circuit, the stored first morphological template obtained from the heart, in the absence of an arrhythmia, while the subject is resting or inactive;a stored second morphological template, coupled to the electrogram sensing circuit, the stored second morphological template obtained from the heart under a condition different from that of the first morphological template;a correlation module, coupled to the first and second morphological templates and the electrogram sensing circuit, the correlation module configured to declare an indication of a first rhythm state if the indication of heart rate exceeds a first predetermined threshold value and a morphology of at least one cardiac complex of the intrinsic cardiac signal correlates to either of the first and second morphological templates, and configured to declare an indication of a second rhythm state if the indication of heart rate exceeds the first predetermined threshold value and a morphology of the at least one cardiac complex of the intrinsic cardiac signal is uncorrelated to both the first and second morphological templates. 2. The system of claim 1, in which the first rhythm state is a supraventricular tachyarrhythmia (SVT) and the second rhythm state is a ventricular tachyarrhythmia. 3. The system of claim 1, further comprising a sensor configured to detect at least one of an activity of the subject and a metabolic need of the subject. 4. The system of claim 3, in which the second morphological template is coupled to the sensor such that the second morphological template is obtained from the subject's heart during a period of exercise indicated by the sensor. 5. The system of claim 1, in which the second morphological template is obtained under a supraventricular tachyarrhythmia (SVT) condition and the first morphological template is obtained under a resting normal sinus rhythm condition. 6. The system of claim 1, further comprising at least one electrode. 7. The system of claim 1, further comprising a remote user interface configured to be communicatively coupled to an implantable device. 8. The system of claim 1, further comprising a template updating control module, coupled to at least one of the first and second morphological templates to occasionally update the at least one of the first and second morphological templates. 9. The system of claim 8, further comprising a sensor including a sensor output configured to provide an indication of at least one of an activity of the subject and a metabolic need of the subject, and in which the template updating control module is coupled to the sensor, and in which the template updat... | Cardiac Pacemakers, Inc. | A61B005/452 A61N001/365
US07150715 | B2 | 2006-12-19 | Network enabled biofeedback administration | 2001-02-05 | An apparatus and method for controlling a biofeedback session. The system utilizes distributed processing that includes localized processing capability allowing a user to initiate a self-controlled biofeedback session and additionally includes remote processing capabilities enabling simultaneous remote analysis and monitoring of the session. | 1. A distributed biofeedback system for managing a biofeedback session, comprising; (a) an input means for receiving physiological data from a user;(b) an output means for communicating feedback to the user;(c) an input node connected to the input means for receiving data from the input means;(d) a first-level data node connected to the input node for receiving and processing data received from the input node;(e) a remote input node connected to the first-level data node for receiving data from the first-level data node and further processing the data and preparation of the data for remote outputting;(f) a remote feedback node having remote feedback inputs connected to the remote input node for producing a remote feedback output responsive to the remote feedback inputs;(g) a first-level feedback node having first-level feedback inputs connected to the first-level data node and the remote feedback node for producing a first-level feedback output responsive to the first-level feedback inputs;(j) a primary feedback node having primary feedback inputs connected to the input node and the first-level feedback node for producing a primary feedback output for controlling the output means. | 1. A distributed biofeedback system for managing a biofeedback session, comprising; (a) an input means for receiving physiological data from a user;(b) an output means for communicating feedback to the user;(c) an input node connected to the input means for receiving data from the input means;(d) a first-level data node connected to the input node for receiving and processing data received from the input node;(e) a remote input node connected to the first-level data node for receiving data from the first-level data node and further processing the data and preparation of the data for remote outputting;(f) a remote feedback node having remote feedback inputs connected to the remote input node for producing a remote feedback output responsive to the remote feedback inputs;(g) a first-level feedback node having first-level feedback inputs connected to the first-level data node and the remote feedback node for producing a first-level feedback output responsive to the first-level feedback inputs;(j) a primary feedback node having primary feedback inputs connected to the input node and the first-level feedback node for producing a primary feedback output for controlling the output means. 2. The distributed biofeedback system for managing a biofeedback session as in claim 1, further comprising a local database for storing and retrieving data from a local node, a first-level input node, a first-level feedback node, and a primary feedback node. 3. The distributed biofeedback system for managing a biofeedback session as in claim 2, further comprising a remote database for storing and retrieving data from the remote input node and remote feedback node. 4. The distributed biofeedback system for managing a biofeedback session as in claim 3, further comprising remote output means connected to a second remote input node for communicating data to a monitor at a remote location; and remote input means for receiving input from the monitor at the remote location. 5. The distributed biofeedback system for managing a biofeedback session as in claim 1, further comprising: (a) a second remote input node connected to the remote input node for receiving data from the remote input node and further processing the data and preparation of the data for remote outputting;(b) remote output means connected to the second remote input node for communicating data to a monitor at a remote location;(c) remote input means for receiving input from the monitor at the remote location;(d) a second remote feedback node having secondary remote feedback inputs connected to the remote input means and second remote input node for producing a second remote feedback output responsive to the secondary remote feedback inputs, the second remote feedback node being connected to one of the inputs of the remote feedback node. 6. The distributed biofeedback system for managing a biofeedback session as in claim 2, wherein the remote input node is connected to a first-level data node by the Internet ... | - | A61B005/00 A61B005/04
US07144369 | B2 | 2006-12-05 | System and method for providing collection and analysis of patient information for use in automated patient care | 2001-09-06 | A system and method for providing collection and analysis of patient information for use in automated patient care. One or more physiological measures relating to individual patient information recorded on a substantially continuous basis are collected. Each of the one or more physiological measures collected at substantially the same time as a set of physiological measures are received. The physiological measures set are stored into a patient care record for an individual patient into a database. One or more of the physiological measures in the patient care record are retrieved from the database. The retrieved physiological measures are analyzed relative to one or more other physiological measures to determine a patient status indicator. | 1. A system for providing collection and analysis of patient information for use in automated patient care, comprising: an implantable medical device directly measuring and recording individual patient information on a substantially continuous basis, comprising a recordation module collecting physiological measures relating to the individual patient information;a collection module receiving the physiological measures collected at substantially the same time as a set of physiological measures directly from the implantable medical device;a database stating the physiological measures set into a patient care record for an individual patient;an analysis module retrieving a plurality of the physiological measures in the patient care record for the individual patient from the database and analyzing the retrieved physiological measures relative to another plurality of the physiological measures to determine a patient status indicator of patient wellness, to recognize a trend in patient well being for the individual patient, and to detennine whether medical intervention is necessary; anda feedback module sending automated feedback directly to the individual patient based on the patient status indicator. | 1. A system for providing collection and analysis of patient information for use in automated patient care, comprising: an implantable medical device directly measuring and recording individual patient information on a substantially continuous basis, comprising a recordation module collecting physiological measures relating to the individual patient information;a collection module receiving the physiological measures collected at substantially the same time as a set of physiological measures directly from the implantable medical device;a database stating the physiological measures set into a patient care record for an individual patient;an analysis module retrieving a plurality of the physiological measures in the patient care record for the individual patient from the database and analyzing the retrieved physiological measures relative to another plurality of the physiological measures to determine a patient status indicator of patient wellness, to recognize a trend in patient well being for the individual patient, and to detennine whether medical intervention is necessary; anda feedback module sending automated feedback directly to the individual patient based on the patient status indicator. 2. A system according to claim 1, further comprising: a database module logically grouping a plurality of the physiological measures, each relating to a same type of individual patient infonnation, into a plurality of peer measures sets; anda comparison module comparing, during analysis, a plurality of the physiological measures selected from one such peer measures set to a further plurality of the physiological measures selected from another such peer measures set. 3. A system according to claim 1, further comprising: a database module logically grouping a plurality of the physiological measures, each relating to a different type of individual patient information, into a plurality of sibling measures sets; anda comparison module comparing, during analysis, a plurality of the physiological measures selected from one such sibling measures set to a further plurality of the physiological measures selected from another such sibling measures set. 4. A system according to claim 1, further comprising: a database module logically grouping a plurality of the physiological measures, each relating to a same type of individual patient information, into a plurality of peer measures sets and logically grouping a plurality of the physiological measures, each relating to a different type of individual patient information, into a plurality of sibling measures sets; anda comparison module comparing, during analysis, a plurality of the physiological measures selected from one such peer measures set to a further plurality of the physiological measures selected from one such sibling measures set. 5. A system according to claim 1, further comprising: a database module selecting a plurality of collected physiological measures, each relating to individual patient information ... | Cardiac Intelligence Corporation | A61B005/02
US07113217 | B2 | 2006-09-26 | Multi-view imaging apparatus | 2001-07-13 | Systems and methods are provided for taking images of a sample. The sample is placed in an imaging box comprising a moveable stage that allows images of the sample to be taken from various positions and angles within the imaging box. The images are taken by a camera and sent to a processor. Structured light images obtained from one or more views within the imaging box may be used to build a structured light representations of the sample. | 1. An imaging system for capturing an image of a sample with a camera, the imaging system comprising: an imaging box having a set of walls enclosing an interior cavity;a camera mount configured to position the camera relative to a fixed datum on one of said walls for viewing by the camera;a light transmission device configured to rotate about the fixed datum; anda moveable stage apparatus including a transport mechanism and a stage configured to support the sample within the interior cavity, said stage being coupled to the transport mechanism for movement of the sample to one of a plurality of positions in said interior cavity, said transport mechanism and said light transmission device are configured to cooperatively direct light reflected or emitted from said sample to said fixed datum to capture said image by said camera. | 1. An imaging system for capturing an image of a sample with a camera, the imaging system comprising: an imaging box having a set of walls enclosing an interior cavity;a camera mount configured to position the camera relative to a fixed datum on one of said walls for viewing by the camera;a light transmission device configured to rotate about the fixed datum; anda moveable stage apparatus including a transport mechanism and a stage configured to support the sample within the interior cavity, said stage being coupled to the transport mechanism for movement of the sample to one of a plurality of positions in said interior cavity, said transport mechanism and said light transmission device are configured to cooperatively direct light reflected or emitted from said sample to said fixed datum to capture said image by said camera. 2. The system of claim 1 wherein the light transmission device comprises a mirror that reflects light emitted from the sample towards the fixed datum. 3. The system of claim 2 wherein the fixed datum is a fixed axis passing though the camera mount and perpendicular to a vertical wall of the imaging box. 4. The system of claim 3 wherein the light transmission device is configured to rotate about an axis co-axially aligned with the fixed axis. 5. The system of claim 3 wherein the moveable stage apparatus rotates the stage about the fixed axis. 6. The system of claim 5 wherein the stage is substantially horizontal during rotation about the fixed axis. 7. The system of claim 5 wherein the light transmission device rotates with the stage about the fixed axis. 8. The system of claim 1 wherein the camera mount is situated on a vertical wall of the set of walls. 9. The system of claim 1 wherein a portion of the stage that supports the sample is transparent. 10. The system of claim 9 wherein the portion comprises a transparent wire array. 11. The system of claim 1 further including a processor configured to provide control signals to a motor coupled to the moveable stage apparatus. 12. The system of claim 11 wherein the processor is further configured to perform crash protection during movement of the moveable stage apparatus. 13. The system of claim 1 further including a hardware based crash protection measure disposed on the moveable stage. 14. The system of claim 1 further including a structured light source configured to transmit structured light onto the sample. 15. The system of claim 14 wherein the processor is further configured to produce a structured light representation using a structured light image obtained from said one of a plurality of positions in said interior cavity. 16. An imaging apparatus used in capturing an image of a sample, the imaging apparatus comprising: an imaging box including an interior cavity for receiving the sample;a stage for supporting the sample;a first linear actuator attached to the imaging box and capable of positioning the moveable stage in a first direction;a second linear act... | Xenogen Corporation | H04N005/225 A61B005/05
US07137953 | B2 | 2006-11-21 | Implantable system and method for measuring left atrial pressure to detect, diagnose and treating congestive heart failure | 2002-10-11 | The invention provides improved apparatus and methods for treating congestive heart failure in a medical patient. The apparatus includes a pressure transducer permanently implantable within the left atrium of the patient's heart and operable to generate electrical signals indicative of fluid pressures within the patient's left atrium. The pressure transducer is connected to a flexible electrical lead, which is connected in turn to electrical circuitry, which in the preferred embodiment includes digital circuitry for processing electrical signals. The electrical circuitry processes the electrical signals from the pressure transducer and, based at least in part on those signals, generates a signal that indicates a desired therapeutic treatment for treating the patient's condition. That signal is then communicated to the patient via a patient signaling device, following which the patient administers to him or herself the prescribed therapeutic treatment indicated by the signal. | 1. A method for treating congestive heart failure in a medical patient, the method comprising: implanting a cardiac pacing lead combined with a sensor lead connected to a pressure transducer substantially permanently within the medical patient;operating the pressure transducer to generate a signal indicative of fluid pressure within the left atrium;communicating the signal indicative of the fluid pressure within the left atrium to a signal processing apparatus;operating the signal processing apparatus to generate an instructive signal indicative of an appropriate therapeutic treatment, wherein operating the signal processing apparatus comprises generating one of at least two distinct signals, each said signal indicative of a different therapeutic treatment; andcommunicating the instructive signal indicative of the appropriate therapeutic treatment to the patient. | 1. A method for treating congestive heart failure in a medical patient, the method comprising: implanting a cardiac pacing lead combined with a sensor lead connected to a pressure transducer substantially permanently within the medical patient;operating the pressure transducer to generate a signal indicative of fluid pressure within the left atrium;communicating the signal indicative of the fluid pressure within the left atrium to a signal processing apparatus;operating the signal processing apparatus to generate an instructive signal indicative of an appropriate therapeutic treatment, wherein operating the signal processing apparatus comprises generating one of at least two distinct signals, each said signal indicative of a different therapeutic treatment; andcommunicating the instructive signal indicative of the appropriate therapeutic treatment to the patient. 2. The method of claim 1, and further comprising implanting at least a portion of the signal processing apparatus substantially permanently within the patient's body. 3. The method of claim 1, and further comprising anchoring the pressure transducer to the patient's atrial septum. 4. The method of claim 3, wherein implanting the signal processing apparatus within the patient's body comprises implanting the signal processing apparatus in a subcutaneous pocket in the region of the patient's shoulder. 5. The method of claim 3, wherein anchoring the pressure transducer to the patient's atrial septum comprises advancing a helical screw into the patient's atrial septum. 6. The method of claim 3, wherein anchoring the pressure transducer to the patient's atrial septum comprises deploying at least one anchor by spring force to bear against the patient's atrial septum. 7. The method of claim 1, wherein communicating the signal indicative of the appropriate therapeutic treatment comprises generating an electrical shock perceptible by the patient. 8. The method of claim 1, wherein generating one of at least two distinct signals, each said signal indicative of a different therapeutic treatment, comprises generating one of at least four distinct signals. 9. The method of claim 1, wherein generating one of at least two distinct signals, each said signal indicative of a different therapeutic treatment, comprises transmitting at least one message through the patient's skin to a remote device. 10. The method of claim 9, further comprising providing written or visual instructions to the patient. 11. The method of claim 9, wherein said remote device is a hand held device. 12. The method of claim 9, wherein said remote device is a tabletop display. 13. The method of claim 1, wherein generating one of at least two distinct signals, each said signal indicative of a different therapeutic treatment, comprises transmitting at least one message through the patient's skin to a device for generating spoken instructions to the patient. 14. The method of claim 13, wherein said spoken instructions to the ... | Cedars-Sinai Medical Center | A61B005/02
US07001344 | B2 | 2006-02-21 | Blood sampling device with diaphragm actuated lancet | 2002-06-12 | A device and method to obtain a blood sample in a sealed configuration. The device includes a sampling assembly having a driver, a lancet, and a blood reservoir, and a pad having a proximal side and a distal side. The proximal side of the pad conforms to the shape of a skin area, and the distal side is attached to the sampling assembly. The pad and the sampling assembly are configured to provide a sealed configuration around the skin area for transferring blood from the skin area to the blood reservoir. | 1. A blood sampling device, comprising: a sampling assembly including a driver, a lancet, and a blood reservoir, said lancet having a proximal end and a distal end; anda pad having a proximal side and a distal side, said proximal side conforming to the shape of a skin area, and said distal side attached to the sampling assembly, said pad and said sampling assembly configured to provide a sealed configuration around the skin area for transferring blood from the skin area to said reservoir;wherein the driver is electromagnetically actuated to move the lancet from a first position to a second position to pierce the skin area;wherein the driver includes a diaphragm coupled to the lancet, said diaphragm to drive said lancet. | 1. A blood sampling device, comprising: a sampling assembly including a driver, a lancet, and a blood reservoir, said lancet having a proximal end and a distal end; anda pad having a proximal side and a distal side, said proximal side conforming to the shape of a skin area, and said distal side attached to the sampling assembly, said pad and said sampling assembly configured to provide a sealed configuration around the skin area for transferring blood from the skin area to said reservoir;wherein the driver is electromagnetically actuated to move the lancet from a first position to a second position to pierce the skin area;wherein the driver includes a diaphragm coupled to the lancet, said diaphragm to drive said lancet. 2. The blood sampling device of claim 1, wherein said pad includes relatively soft durometer type material. 3. The blood sampling device of claim 2, wherein said material includes silicone rubber. 4. The blood sampling device of claim 1, wherein said diaphragm is formed with an elastomer of harder durometer type material than said pad. 5. The blood sampling device of claim 1, wherein said diaphragm is formed with an elastomer of less permeability than said pad. 6. The blood sampling device of claim 1, wherein said diaphragm compresses and flexes toward the skin area when said diaphragm is actuated and said lancet is protracted out. 7. The blood sampling device of claim 1, wherein said diaphragm creates suction upon retraction of said lancet. 8. The blood sampling device of claim 1, wherein said blood reservoir is enclosed by said diaphragm and said sampling assembly. 9. The blood sampling device of claim 8, further comprising: an opening on a surface of said sampling assembly, said opening configured to provide a conduit for the blood sample into the reservoir. 10. The blood sampling device of claim 9, wherein said surface of said sampling assembly is coupled to the distal side of said pad. 11. The blood sampling device of claim 9, wherein said proximal end of the lancet is approximately perpendicular to said surface of the sampling assembly, and is concentric with said opening. 12. The blood sampling device of claim 9, further comprising: a first valve coupled to said opening, said first valve to regulate inflow of the blood sample into said reservoir. 13. The blood sampling device of claim 12, wherein said lancet passes through said first valve to hold said first valve open prior to piercing the skin area. 14. The blood sampling device of claim 12, wherein said first valve includes a duckbill check valve. 15. The blood sampling device of claim 8, further comprising: a second valve to enable outflow of the blood sample from said reservoir. 16. The blood sampling device of claim 9, further comprising: a sensor disposed on the surface of said sampling assembly, said sensor configured to detect when said proximal side of said pad conforms to the shape of said skin area. 17. A device, comprising: a blood sampling asse... | Pelikan Technologies, Inc. | A61B005/00
US07191006 | B2 | 2007-03-13 | Cardiac rhythm management systems and methods for rule-illustrative parameter entry | 2002-12-05 | This document describes systems, devices, and methods for programmably adjusting therapy or other operative parameter values of an implantable cardiac rhythm management device using a programmer that includes a display that graphically illustrates one or more relationship between a parameter being programmed and other parameters and/or rules governing interactions between parameters. | 1. A system comprising: a programmer device, including a communication circuit configured to be communicatively coupled to an implantable cardiac rhythm management (CRM) device, and including a user interface configured to program a plurality of parameters used by the CRM device, the user interface configured to interactively display, during a programming of at least one of the parameters: a graphical indication of a changing value of at least one of the parameters as at least one of the parameters is being programmed graphically; anda graphical illustration of how a rule constrains the value of the at least one of the parameters being programmed graphically. | 1. A system comprising: a programmer device, including a communication circuit configured to be communicatively coupled to an implantable cardiac rhythm management (CRM) device, and including a user interface configured to program a plurality of parameters used by the CRM device, the user interface configured to interactively display, during a programming of at least one of the parameters: a graphical indication of a changing value of at least one of the parameters as at least one of the parameters is being programmed graphically; anda graphical illustration of how a rule constrains the value of the at least one of the parameters being programmed graphically. 2. The system of claim 1, in which the user interface is configured to interactively flag a violation of the rule constraining the at least one of the parameters being programmed as the value of the at least one of the parameters being programmed is being changed such that it violates the rule constraining the at least one of the parameters being programmed. 3. The system of claim 2, in which the user interface is configured to provide a corrective indication, if a violation of the rule is flagged, the corrective indication informing a user how the value of at least one parameter included by the rule should be changed to correct the violation of the rule. 4. The system of claim 1, in which the graphical indication of the changing value of the at least one parameter being programmed includes a draggable bar graph icon having a bar length indicative of the value of the at least one parameter being programmed. 5. The system of claim 4, in which the graphical indication of the rule includes a bar graph icon presented in visual correspondence with the graphical indication of the changing value of the at least one parameter being programmed. 6. The system of claim 1, in which the rule constraining the value of the at least one parameter being programmed includes a limit that constrains a combination of the value of the at least one parameter being programmed and a value of at least one other parameter. 7. The system of claim 1, in which the user interface includes a graphical indication of a changing value of the at least one parameter as it is being programmed together with a graphical indication of a plurality of rules constraining the value of the at least one parameter being programmed. 8. The system of claim 1, in which the user interface includes a graphical indication of a changing value of the at least one parameter as it is being programmed together with a graphical indication of every rule constraining the value of the at least one parameter being programmed. 9. The system of claim 1, in which the graphical indication of the rule constraining the value of the at least one parameter being programmed includes an accompanying icon that, when selected, triggers a textual display of a rationale underlying the rule constraining the value of the at least one parameter being programmed. ... | Cardiac Pacemakers, Inc. | A61B005/044
US06917826 | B2 | 2005-07-12 | Automatic and quantitative assessment of image quality in SPECT myocardial perfusion imaging | 2002-02-01 | A system and method for automatic quantitative assessment of myocardial image quality includes a liver-heart artifact quantification unit for providing a sequence of images, detecting the heart position in each image, segmenting non-heart activity relative to the detected heart positions, measuring the liver-heart effect in each image based on the segmentations, fusing the measurements obtained in each image, and determining a quality number for the plurality of images responsive to the fused measurements; and also includes a heart detection unit for receiving the image sequence, generating a region of interest comprising the heart by 1-D pseudo-motion analysis of the image sequence, localizing the heart center by a training-based algorithm, fitting a curve to the localized heart centers of each image in the sequence, refining the heart center positions of localized heart centers falling outside of a tolerance distance from the fitted curve, and determining a bounding circle comprising the heart in accordance with the localized and refined heart center positions. | 1. A method for automatic quantitative assessment of myocardial image quality, the method comprising: providing a plurality of images; detecting the heart position in each provided image; segmenting non-heart activity relative to the detected heart position in each image; measuring the liver-heart effect in each image relative to the segmentations; fusing the measurements obtained in each image; determining a quality number for the plurality of images responsive to the fused measurements; and alerting a user to low image quality if the quality number is less than a threshold. | 1. A method for automatic quantitative assessment of myocardial image quality, the method comprising: providing a plurality of images; detecting the heart position in each provided image; segmenting non-heart activity relative to the detected heart position in each image; measuring the liver-heart effect in each image relative to the segmentations; fusing the measurements obtained in each image; determining a quality number for the plurality of images responsive to the fused measurements; and alerting a user to low image quality if the quality number is less than a threshold. 2. A method as defined in claim 1, further comprising: comparing the quality number with a threshold; and performing appropriate actions if the quality number is poor relative to the threshold. 3. A method as defined in claim 1 wherein detecting the heart position in each image comprises: receiving an image sequence; generating a region of interest comprising the heart by 1-D pseudo-motion analysis of the image sequence; localizing the heart center by a training-based algorithm; fitting a curve to the localized heart centers of each image in the sequence; refining the heart center positions of localized heart centers falling outside of a tolerance distance from the fitted curve; and determining a bounding circle comprising the heart in accordance with the localized and refined heart center positions. 4. A method as defined in claim 3 wherein said curve is a sine curve for compensating for the motion of an imaging device. 5. A method as defined in claim 1 wherein said images comprise single-photon emission computed tomography images. 6. A system (100) for automatic quantitative assessment of myocardial image quality, the system comprising: a heart detection unit (180) for heart detection, outlier detection and outlier correction; a liver-heart artifact quantification unit (170) responsive to the heart detection unit for determining a quality number indicative of the severity of image-quality induced liver-heart; and an output device (116. 132) for alerting a user to low image quality if the quality number is less than a threshold. 7. A system (100) as defined in claim 6 wherein the images comprise single-photon emission computed tomography images. 8. A system (100) for automatic quantitative assessment of myocardial image quality, the system comprising: a heart detection unit (180) for heart detection, outlier detection and outlier a liver-heart artifact quantification unit (170) responsive to the heart detection unit for determining a quality number indicative of the severity of image-quality induced liver-heart artifacts; and a CPU (102) in signal communication with said liver-heart artifact quantification unit (170) for processing the images. 9. A system (100) as defined in claim 8, further comprising: a display adapter (110) in signal communication with the CPU (102) for displaying the images; and an I/O adapter (112) in signal communication with the CPU (102... | Siemens Corporate Research, Inc. | A61B005/05
US06947877 | B2 | 2005-09-20 | Methods for optimizing magnetic resonance imaging systems | 2000-12-07 | A computer-implemented method for optimizing the design of an electromagnetic coil arrangement that generates a uniform magnetic field in a desired region, with the electromagnetic coil arrangement having a number of coils and a shape defined by r and z, where r is the radial coordinate of a cylindrical coordinate system having (r, z, φ) coordinates, and z is the axial coordinate of the cylindrical coordinate system. The method has an important use for designing superconductive or resistive coil arrangements for MRI systems. The core of the method is to compute a solution to a master equation (25) for a set of parameters lambda_j which define the coil arrangement and the currents needed. An important advantage is that the resultant coil arrangement has a minimum volume, which saves on material, or uses a minimum power. | 1. A computer-implemented method for designing an electromagnetic coil arrangement that generates a uniform magnetic field in a desired region, said electromagnetic coil arrangement having a number of coils and a shape defined by r and z, where r is the radial coordinate of a cylindrical coordinate system having (r, z, φ) coordinates, and z is the axial coordinate of the cylindrical coordinate system; comprising the steps: (a) computing a solution to Eq. (25) as defined in the accompanying specification for a set of parameters lambda_j under the conditions recited in Eqs. (5), (6), and (8) of the accompanying specification, (b) using a computer, numerically evaluating Eq. (14) as defined in the accompanying specification using the set of parameters lambda_j computed in step (a) to obtain an ideal current density J(r, z) for the electromagnetic coil arrangement to generate a desired magnetic field intensity, (c) plotting J as a function of r and z as computed in step (b) to obtain the number of coils and the shape of the electromagnetic coil arrangement to determine its geometry. | 1. A computer-implemented method for designing an electromagnetic coil arrangement that generates a uniform magnetic field in a desired region, said electromagnetic coil arrangement having a number of coils and a shape defined by r and z, where r is the radial coordinate of a cylindrical coordinate system having (r, z, φ) coordinates, and z is the axial coordinate of the cylindrical coordinate system; comprising the steps: (a) computing a solution to Eq. (25) as defined in the accompanying specification for a set of parameters lambda_j under the conditions recited in Eqs. (5), (6), and (8) of the accompanying specification, (b) using a computer, numerically evaluating Eq. (14) as defined in the accompanying specification using the set of parameters lambda_j computed in step (a) to obtain an ideal current density J(r, z) for the electromagnetic coil arrangement to generate a desired magnetic field intensity, (c) plotting J as a function of r and z as computed in step (b) to obtain the number of coils and the shape of the electromagnetic coil arrangement to determine its geometry. 2. A computer-implemented method for designing an electromagnetic coil arrangement that generates a uniform magnetic field in a desired region as set forth in claim 1, wherein the desired region is an imaging region of interest for a magnetic resonance imaging system, and step (b) reformulates the set of parameters lambda_j computed in step (a) into the standard form of Eq. (14) to give the ideal current density for a selected coil feasible volume producing in the imaging region a selected magnetic field intensity. 3. A computer-implemented method for designing an electromagnetic coil arrangement that generates a uniform magnetic field in a desired region as set forth in claim 2, further comprising the step: (d) numerically evaluating Eq. (9) as defined in the accompanying specification using the result obtained in step (c) to determine the minimum coil volume of the electromagnetic coil arrangement. 4. A computer-implemented method for optimizing the design of an electromagnetic coil arrangement that generates a uniform magnetic field in a desired region, said electromagnetic coil arrangement having a number of coils and a shape defined by r and z, where r is the radial coordinate of a cylindrical coordinate system having (r, z, φ) coordinates, and z is the axial coordinate of the cylindrical coordinate system; comprising the steps: (a). Select a feasible volume for the coils, (b). Select a desired current density magnitude for the coils, (c). Select a desired magnetic field level for the desired region, (d). Select the number and values of harmonics to be controlled in the desired region, (e) computing a solution to Eq. (25) as defined in the accompanying specification for a set of parameters lambda_j, (f) reformulating the solution to Eq. (25) to the standard form of Eq. (14) as defined in the accompanying specification using the set of parameters lambda_j compute... | New York University | G06F017/10 A61B005/05
US06950702 | B2 | 2005-09-27 | Use of curvature based features for beat detection | 2002-07-15 | A cardiac rhythm management system includes a sensing circuit to sense a cardiac signal and a sensing processor to detect cardiac depolarizations (beats) by utilizing certain morphological context of the sensed cardiac signal. The sensing processor samples the sensed cardiac signal, computes curvatures from the sampled cardiac signal to generate a cardiac curvature signal corresponding to the cardiac signal, derives cardiac signal features reflecting morphologically significant points along the cardiac signal from the cardiac curvature signal, and detects cardiac depolarizations based on an analysis of the cardiac signals features. | 1. A system, comprising: a sensing circuit to sense a cardiac signal; a sampling circuit, coupled to the sensing circuit, to sample the cardiac signal on a continuous basis; a feature analyzer, coupled to the sampling circuit, to generate a cardiac curvature signal based on the cardiac signal and detect cardiac signal features from the cardiac curvature signal based on predetermined detection criteria; a counter, coupled to the feature analyzer, to count the cardiac signal features over a predetermined period of time; and a comparator, coupled to the counter, to receive a first input representative of a number of the cardiac signal features counted over the predetermined period of time and compare that input to a second input representative of a predetermined threshold number, and to provide a depolarization indicating signal when the number of the cardiac signal features counted over the predetermined period of time exceeds the threshold number. | 1. A system, comprising: a sensing circuit to sense a cardiac signal; a sampling circuit, coupled to the sensing circuit, to sample the cardiac signal on a continuous basis; a feature analyzer, coupled to the sampling circuit, to generate a cardiac curvature signal based on the cardiac signal and detect cardiac signal features from the cardiac curvature signal based on predetermined detection criteria; a counter, coupled to the feature analyzer, to count the cardiac signal features over a predetermined period of time; and a comparator, coupled to the counter, to receive a first input representative of a number of the cardiac signal features counted over the predetermined period of time and compare that input to a second input representative of a predetermined threshold number, and to provide a depolarization indicating signal when the number of the cardiac signal features counted over the predetermined period of time exceeds the threshold number. 2. The system of claim 1, further comprising a therapy circuit, coupled to the comparator, to trigger or inhibit a delivery of a cardiac therapy based on the depolarization indicating signal. 3. The system of claim 1, wherein the feature analyzer comprises: a curvature generator, coupled to the sampling circuit, to generate a cardiac curvature signal corresponding to the cardiac signal by computing a curvature, after each sample is taken, based on the each sample and a predetermined number of preceding samples; and a feature detector adapted to detect the cardiac signal features from the cardiac curvature signal. 4. The system of claim 3, further comprising an intracardiac electrode coupled to the sensing circuit. 5. The system of claim 3, wherein the feature detector comprises a feature comparator adapted to compare an amplitude of the cardiac curvature signal to a predetermined threshold amplitude and generate a signal indicating that a cardiac signal feature has been detected when the amplitude of the cardiac curvature signal exceeds the predetermined threshold amplitude. 6. The system of claim 3, wherein the feature detector comprises: a zero-crossing detector to detect lobes of the cardiac curvature signal, the lobes each including a beginning time when the cardiac curvature signal reaches a first threshold and an ending time when the cardiac curvature signal reaches a second threshold, the first and second thresholds each being slightly away from or approximately equal to a baseline of the cardiac curvature signal; a size calculator to calculate an approximate area between the cardiac curvature signal and the baseline for each of the lobes, and a centroid locator to temporally locate a centroid of the calculated approximate area of the each of the lobes, the centroid indicative of an approximate time of occurrence of one of the cardiac signal features representing the each of the lobes. 7. A system comprising: a sensing circuit to sense a cardiac signal; a sampling circuit, coupled to the ... | Cardiac Pacemakers, Inc. | A61N001/362 A61B005/0452
US06884225 | B2 | 2005-04-26 | Medical guide wire | 2002-12-03 | In a medical guide wire (1), a rigid portion (4A) is located at a distal end of a helical spring (4), and the rigid portion (4A) measures 0.5 mm or less which extends from a top end of a head plug (5) to a rear end of a portion in which the helical spring (4) is secured to the head plug (5), and an intermediary location (11) is secured to an intermediary boss portion (9) of a core line (3) to provide a loosely wound portion (8) from a rear end of the head plug (5) to the intermediary location (11) to be elastically expandable due to a clearence (C) appeared between line element turns of the loosely wound portion (8). | 1. A medical guide wire comprising a flexible line wire having a main wire portion, a core line member and a helical spring provided around a distal end of said core line member, a front end of said helical spring being fixedly interfit to a head plug provided at a tip of said core line member; a rigid portion located on a front end of aid helical spring, said rigid portion measuring 0.5 mm or less which extends lengthwisely from a top end of said head plug to a rear end of a portion in which said helical spring is fixed to said head plug, and an intermediary location is determined on an intermediary portion of said helical spring in which an intermediary boss portion is secured to said core line member to define loosely wound portion as an elastically expandable portion from a rear end of said head plug to said intermediary location; a clearance between line element turns of said loosely wound portion being within a range of 10-30% of a line diameter of said helical spring; and said loosely wound portion being approximately 24 mm in length. | 1. A medical guide wire comprising a flexible line wire having a main wire portion, a core line member and a helical spring provided around a distal end of said core line member, a front end of said helical spring being fixedly interfit to a head plug provided at a tip of said core line member; a rigid portion located on a front end of aid helical spring, said rigid portion measuring 0.5 mm or less which extends lengthwisely from a top end of said head plug to a rear end of a portion in which said helical spring is fixed to said head plug, and an intermediary location is determined on an intermediary portion of said helical spring in which an intermediary boss portion is secured to said core line member to define loosely wound portion as an elastically expandable portion from a rear end of said head plug to said intermediary location; a clearance between line element turns of said loosely wound portion being within a range of 10-30% of a line diameter of said helical spring; and said loosely wound portion being approximately 24 mm in length. 2. A medical guide wire comprising a flexible line wire having a main wire portion, a core member and a helical spring provide around a distal end of said core line member, a front end of said helical spring being fixedly interfit to a head plug provided a tip of said core line member; a rigid portion located on a front end of said helical spring, said rigid portion measuring 0.5 mm or less which extends lengthwisely from a top end of said head plug to a rear end of a portion in which said helical spring is fixed to said head plug, and an intermediary location is determined on an intermediary portion of said helical spring in which an intermediary boss portion is secured to said core line member to define a loosely wound portion as an elastically expandable portion from a rear end of said head plug to said intermediary location; and said head plug forming a spherical or semi-spherical configuration, and a right and left side portions of said head plug being undercut to form flat surface portions. 3. The medical guide wire according to claim 2, wherein said core line member is rectangular in cross section, longer sides of which position in parallel with said flat surface portions of said head plug. 4. The medical guide wire according to claim 2, wherein said front end of said helical spring and said head plug are produced from TIG or laser spot welding. 5. A medical guide wire comprising a flexible line wire having a main wire portion, a core member and a helical spring provided around a distal end of said core line member, a front end of said helical spring being fixedly interfit to a head plug provided at a tip of said core line member; a rigid portion located on a front end of said helical spring, said rigid portion measuring 0.5 mm or less which extends lengthwisely from a top end of said head plug to a rear end of a portion in which said helical spring is fixed to said head plug, and an intermediary... | Asahi Intecc Co., Ltd. | A61B005/00 A61M025/00
US06950691 | B2 | 2005-09-27 | Surgery support system and surgery support method | 2002-04-10 | A first controller synthesizes image output of a CCU and a cursor created based on trigger information from a second controller, following location information. The synthesized image is displayed on a first monitor, so an endoscope image and a cursor image which moves synchronously with the cursor image on a display device can be simultaneously viewed at the surgery room side on a single monitor, and accordingly, the problem of the working space inside the surgery room becoming crowded due to providing multiple observation monitors in the surgery room does not occur. | 1. A surgery support system, for performing communication of information regarding surgery, and supporting surgery, said system comprising: an imaging device for taking images of a treatment region of a patient, and outputting picture signals; a first display device for displaying images of said treatment region, based on picture signals output from said imaging device; a first transmitting circuit for transmitting picture signals output from said imaging device to a communication line; a first receiving circuit for receiving said picture signals from said first transmitting circuit via said communication line; a second display device for displaying images of said treatment region, based on picture signals received with said first receiving circuit; a first pointer superimposing circuit for superimposing a first pointer for instructing location on said second display device display said treatment region; an operating unit for moving said first pointer displayed on said second display device by said first pointer superimposing circuit; a second transmitting circuit for transmitting location information of said first pointer to said communication line; a second receiving circuit for receiving location information of said first pointer from said second transmitting circuit via said communication line; and a second pointer superimposing circuit for superimposing a second pointer on said first display device, based on location information received with said second receiving means. | 1. A surgery support system, for performing communication of information regarding surgery, and supporting surgery, said system comprising: an imaging device for taking images of a treatment region of a patient, and outputting picture signals; a first display device for displaying images of said treatment region, based on picture signals output from said imaging device; a first transmitting circuit for transmitting picture signals output from said imaging device to a communication line; a first receiving circuit for receiving said picture signals from said first transmitting circuit via said communication line; a second display device for displaying images of said treatment region, based on picture signals received with said first receiving circuit; a first pointer superimposing circuit for superimposing a first pointer for instructing location on said second display device display said treatment region; an operating unit for moving said first pointer displayed on said second display device by said first pointer superimposing circuit; a second transmitting circuit for transmitting location information of said first pointer to said communication line; a second receiving circuit for receiving location information of said first pointer from said second transmitting circuit via said communication line; and a second pointer superimposing circuit for superimposing a second pointer on said first display device, based on location information received with said second receiving means. 2. A surgery support system according to claim 1, further comprising a transmission control circuit for controlling said second transmitting circuit, based on change in location information of said first pointer which is moved by said operating unit. 3. A surgery support system according to claim 1, further comprising a display control circuit for controlling said second pointer superimposing circuit, based on change in location information received with said second receiving circuit. 4. A surgery support system according to claim 1, wherein the imaging device is provided in an endoscope for observing within a body cavity. 5. A surgery support system according to claim 4, wherein, in the event that there is no change in the location information of said first pointer which is changed by said operating unit for a predetermined amount of time, said first pointer superimposing circuit outputs the location information of said first pointer to said second receiving circuit. 6. A surgery support system according to claim 4, wherein, in the event that there is no change in the location information of said first pointer which is received by said first receiving circuit over a predetermined amount of time, said second superimposing circuit superimposes said second pointer on said images. 7. A surgery support system according to claim 4, wherein said first pointer superimposing circuit inputs trigger information for confirming location information from said operating unit, and o... | Olympus Corporation | A61B005/05
US06856826 | B2 | 2005-02-15 | Fluoroscopic tracking and visualization system | 2002-11-15 | A method for surgical imaging and display including (i.) positioning a defined set of markers disposed in a pattern so as to be imaged in each pose or view of an imaging assembly, the set of markers being fixed in pre-determined positions in a rigid carrier, (ii.) securing a first tracking element against motion with respect to the rigid carrier so that determining a position of the first tracking element in a single measurement determines positions of all the markers of the set, and (iii.) identifying images of at least a subset of the markers in a first view. | 1. A method for surgical imaging and display of the type including an imaging assembly having an x-ray source and detector, for displaying an image for surgical guidance, wherein the method comprises the steps of: positioning a defined set of markers disposed in a pattern so as to be imaged in each pose or view of the imaging assembly, the set of markers being fixed in pre-determined positions in a rigid carrier; securing a first tracking element against motion with respect to the rigid carrier so that determining a position of the first tracking element in a single measurement determines positions of all the markers of the set; and identifying images of at least a subset of the markers in a first view. | 1. A method for surgical imaging and display of the type including an imaging assembly having an x-ray source and detector, for displaying an image for surgical guidance, wherein the method comprises the steps of: positioning a defined set of markers disposed in a pattern so as to be imaged in each pose or view of the imaging assembly, the set of markers being fixed in pre-determined positions in a rigid carrier; securing a first tracking element against motion with respect to the rigid carrier so that determining a position of the first tracking element in a single measurement determines positions of all the markers of the set; and identifying images of at least a subset of the markers in a first view. 2. The method of claim 1, wherein said positioning step comprises disposing the set of markers in a plurality of planes having a defined separation therebetween. 3. The method of claim 1, further comprising securing additional tracking elements to at least one object selected from among a patient and a medical instrument so as to determine a position of said object. 4. The method of claim 1, further comprising securing a second tracking element to the detector. 5. The method of claim 1, further comprising securing a second tracking element to a patient being imaged. 6. The method of claim 1, further comprising securing a second tracking element to a medical instrument. 7. A method for surgical imaging and display of tissue structures of a patient, comprising: affixing a fixture to an imaging side of a fluoroscope for providing patterns of an array of markers that are imaged in each fluoroscope image; deriving at least one image in a display from a fluoroscope during a surgical procedure; securing at least one of a plurality of tracking elements against motion with respect to the fixture; determining positions of the fixture and the patient through a tracking assembly having a plurality of tracking elements so that determining a position of the at least one of the plurality of tracking elements determines a position of each of the array of markers in a single measurement; and displaying images in coordination with a tool image to facilitate manipulation of a tool during the surgical procedure. 8. The method of claim 7, wherein said affixing step comprises disposing the array of markers in a plurality of planes having a defined separation therebetween. 9. The method of claim 7, further comprising providing an identification of a subset of the markers in at least one of the images. 10. The method of claim 9, further comprising determining a camera model that includes at least one distortion of each fluoroscope image. 11. The method of claim 7, further comprising automatically determining relative position data of the array of markers and the tool. 12. The method of claim 7, further comprising determining an image transformation such that a representation of the tool is correctly positioned and oriented in the fluoroscope image. 13. The ... | GE Medical Systems Global Technology Company, LLC | A61B005/05
US06859662 | B2 | 2005-02-22 | Method and device for the diagnosis and therapy of chronic heart failure | 2001-12-21 | Method and device for the diagnosis and therapy of chronic heart failure comprising continuous monitoring of the patient and continuous determination of significant decompensation parameters during a sample period of normal patient life, recording the data determined, continuously monitoring these data during therapy, comparing the memorized data with those determined during the same time span of the sample period and comparing the duration of periods in which decompensation is present with the total duration of those periods during which decompensation is absent or conforms to that determined during the sample period. | 1. A method for the diagnosis and therapy of chronic heart failure comprising continuous monitoring of the patient and continuous determination of significant decompensation parameters during a sample period of normal patient life, recording the data determined, continuously monitoring these data during therapy, comparing the memorized data with those determined during the same time span of the sample period and comparing the duration of periods in which decompensation is present with the total duration of those periods during which decompensation is absent or conforms to that determined during the sample period, wherein the memorized parameters comprise the curve of ventricular contractile force variation as a function of heart rate. | 1. A method for the diagnosis and therapy of chronic heart failure comprising continuous monitoring of the patient and continuous determination of significant decompensation parameters during a sample period of normal patient life, recording the data determined, continuously monitoring these data during therapy, comparing the memorized data with those determined during the same time span of the sample period and comparing the duration of periods in which decompensation is present with the total duration of those periods during which decompensation is absent or conforms to that determined during the sample period, wherein the memorized parameters comprise the curve of ventricular contractile force variation as a function of heart rate. 2. A method as claimed in claim 1, wherein the memorized parameters are the heart rate, the ventricular contractile force, and the curve of force variation as a function of heart rate. 3. A method as claimed in claim 2, wherein the memorized parameters are the heart rate, the ventricular contractile force, and the curve of force variation as a function of heart rate at determined moments of the sample period. 4. A method as claimed in claim 2, in which memorizing the force-frequency curve variations occurs at least during a period equal to the sample period. 5. A method as claimed in claim 2, deriving the force data from the intracardiac pressure curves. 6. A method as claimed in claim 2, deriving the force data from the peripheral pressure curves. 7. A method as claimed in claim 2, deriving the force data from the ventricular volumes. 8. A method as claimed in claim 2, deriving the force data from the pressure/volume relationship. 9. A method as claimed in claim 2, deriving the force data from cardiac tone (force of contraction and/or rate of tension development). 10. A method as claimed in claim 2, deriving the force data via impedance. 11. A method as claimed in claim 2, deriving the force data from Doppler flow measurement. 12. A method as claimed in claim 2, deriving the force data with echo-Doppler. 13. A method as claimed in claim 2, deriving the force data from a combination of all or part of the following parameters: the intracardiac pressure curves, the peripheral pressure curves, the ventricular volumes, the pressure/volume relationship, the cardiac tone, the impedance, the Doppler flow measurement, the echo-Doppler measurement. 14. A method as claimed in claim 2, expressing the ventricular force as dP/dt/EDV, which indicates the maximal rate of left ventricular pressure development divided by end-diastolic volume (EDV). 15. A method as claimed in claim 2, expressing the ventricular force as sP/ESV, which indicates the left ventricular end-systolic pressure volume ratio (end-systolic ventricular pressure divided by the end-systolic volume). | - | A61B005/04
US07081095 | B2 | 2006-07-25 | Centralized hospital monitoring system for automatically detecting upper airway instability and for preventing and aborting adverse drug reactions | 2002-05-17 | A system and method for the automatic diagnosis of obstructive sleep apnea in a centralized hospital critical care monitoring system for the monitoring of a plurality of patients in at least one of a critical care, step down, and cardiac ward by telemetry. The system includes a central processor having a display, and a plurality of telemetry units for mounting with patients, each of the telemetry units has a plurality of sensors for connection with each patient, the telemetry unit is capable of the transmission of multiple signals derived from the sensors to the central processor, in one preferred embodiment the method comprising steps of programming the system to analyze the signals and to automatically identify the presence and severity of obstructive sleep apnea and to provide an indication of the identification. | 1. A method for the automatic detection of obstructive sleep apnea in a centralized hospital critical care monitoring system for the monitoring by telemetry of a plurality of patients in at least one of a critical care ward and a hospital ward, the system including a central processor having a display, and a plurality of telemetry units for mounting with said patients, each of said telemetry units having a plurality of sensors for connection with each said patient, each of said telemetry units transmitting a plurality of signals to said central processor, the method comprising the steps of: a) programming said system to analyze said signals and to-automatically identify at least one of the presence and severity of obstructive sleep apnea,b) programming said system to output an indication of said identified obstructive sleep apnea; andc) programming said system to communicate with an intravenous infusion system to prevent the progression of said identified obstructive sleep apnea by limiting infusion of a narcotic. | 1. A method for the automatic detection of obstructive sleep apnea in a centralized hospital critical care monitoring system for the monitoring by telemetry of a plurality of patients in at least one of a critical care ward and a hospital ward, the system including a central processor having a display, and a plurality of telemetry units for mounting with said patients, each of said telemetry units having a plurality of sensors for connection with each said patient, each of said telemetry units transmitting a plurality of signals to said central processor, the method comprising the steps of: a) programming said system to analyze said signals and to-automatically identify at least one of the presence and severity of obstructive sleep apnea,b) programming said system to output an indication of said identified obstructive sleep apnea; andc) programming said system to communicate with an intravenous infusion system to prevent the progression of said identified obstructive sleep apnea by limiting infusion of a narcotic. 2. A method for providing automatic detection of upper airway instability in hospitalized patients who are being otherwise routinely monitored with a centralized hospital monitoring system, in at least one of an acute care environment or a surgical hospital environment, so that the routine monitoring provides automatic screening for upper airway instability, the method comprising the steps of: a) programming said centralized system to automatically identify the presence of upper airway instability;b) programming said centralized system to output an indication of said identified upper airway instability; andc) programming said centralized system to communicate with an intravenous infusion system to prevent the progression of said identified upper airway instability by limiting infusion of a narcotic. 3. The method of claim 2, further comprising the steps of: a) programming said centralized system to automatically identify the presence of obstructive sleep apnea; andb) outputting an indication of said identified obstructive sleep apnea. 4. The method of claim 2, wherein said method further comprises the step of: programming said centralized system to provide a graded indication based on said identified upper airway instability. 5. The method of claim 2, wherein said method further comprises the step of: programming said centralized system to communicate with a system for intravenously infusing a medication. 6. The method of claim 2, wherein said method further comprises the step of: programming said centralized system to communicate with an airway pressure delivery system for treatment of airway instability. 7. A method for facilitating case finding of obstructive sleep apnea in hospitalized patients who are being otherwise routinely monitored with a centralized hospital monitoring system in at least one of an acute care environment, a sub acute care environment and a surgical hospital environment, the method comprising: a) programm... | - | A61B005/00
US07050844 | B2 | 2006-05-23 | Method for detecting the three-dimensional position of a medical examination instrument introduced into a body region, particularly of a catheter introduced into a vessel | 2002-03-22 | In a method for detecting the three-dimensional position of a medical examination instrument introduced into an examination region of a subject and an apparatus including a C-arm system for registering radiation images, two-dimensional angiography projection images of the examination region are registered for later three-dimensional volume reconstruction of the examination region. After introduction of the instrument, at least one two-dimensional projection image pair of the examination region with image planes residing at an angle relative to one another is registered. The spatial coordinates of a selected point of the instrument shown in the projection images are identified in a common coordinate system for the angiography projection images and the projection images. Two-dimensional projection image pairs are registered in succession for the continuous presentation of the selected point in the display of the aforementioned volume reconstruction. | 1. A method for detecting a three-dimensional position of a medical instrument introduced into a body region, using a radiological imaging system having at least one C-arm system with a C-arm having a radiation source and a radiation detector mounted thereon, comprising the steps of: obtaining a plurality of two-dimensional angiography projection images relative to a coordinate system, using said radiological imaging system, of an examination region into which said instrument is to be introduced, and subsequently producing a three-dimensional volume reconstruction of said examination region from said two-dimensional angiography projection images;introducing said instrument into said examination region and, after introduction of said instrument, obtaining at least one pair of two-dimensional projection images of said examination region relative to said coordinate system, using said radiological examination system, said two-dimensional projection images in said pair being disposed at an angle relative to each other;identifying a first set of projection matrices for respective positions of said C-arm when obtaining said angiography images, and identifying a second set of projection matrices for respective positions of said C-arm when obtaining said respective pairs of said two-dimensional projection images, and calibrating said first set of projection matrices and said second set of projection matrices with respect to said coordinate system;identifying spatial coordinates of a selected point of said instrument in said pair of two-dimensional projection images in said coordinate system;representing said selected point of said instrument in a visual presentation of said three-dimensional volume reconstruction of said examination region in said coordinate system; andrepeatedly obtaining respective pairs of said two-dimensional projection images in sufficiently frequent succession to continuously represent said selected point of said examination instrument in said examination region in said visual representation. | 1. A method for detecting a three-dimensional position of a medical instrument introduced into a body region, using a radiological imaging system having at least one C-arm system with a C-arm having a radiation source and a radiation detector mounted thereon, comprising the steps of: obtaining a plurality of two-dimensional angiography projection images relative to a coordinate system, using said radiological imaging system, of an examination region into which said instrument is to be introduced, and subsequently producing a three-dimensional volume reconstruction of said examination region from said two-dimensional angiography projection images;introducing said instrument into said examination region and, after introduction of said instrument, obtaining at least one pair of two-dimensional projection images of said examination region relative to said coordinate system, using said radiological examination system, said two-dimensional projection images in said pair being disposed at an angle relative to each other;identifying a first set of projection matrices for respective positions of said C-arm when obtaining said angiography images, and identifying a second set of projection matrices for respective positions of said C-arm when obtaining said respective pairs of said two-dimensional projection images, and calibrating said first set of projection matrices and said second set of projection matrices with respect to said coordinate system;identifying spatial coordinates of a selected point of said instrument in said pair of two-dimensional projection images in said coordinate system;representing said selected point of said instrument in a visual presentation of said three-dimensional volume reconstruction of said examination region in said coordinate system; andrepeatedly obtaining respective pairs of said two-dimensional projection images in sufficiently frequent succession to continuously represent said selected point of said examination instrument in said examination region in said visual representation. 2. A method as claimed in claim 1 comprising repeatedly obtaining said respective pairs of said two-dimensional projection images continuously. 3. A method as claimed in claim 1 comprising repeatedly obtaining said respective pairs of said two-dimensional projection images discontinuously. 4. A method as claimed in claim 1 wherein said radiological imaging system comprises two C-arm systems disposed at an angle relative to each other, and wherein the step of obtaining the pair of two-dimensional projection images and the step of repeatedly obtaining respective pairs or two-dimensional projection images each comprises obtaining a first image of each pair with a first of said two C-arm system and obtaining a second image of each pair with a second of the two C-arm systems, and further comprising calibrating both of said first and second C-arm systems with respect to said coordinate system. 5. A method as claimed in claim 4 comprising visually... | Siemens Aktiengesellschaft | A61B005/05
US06985771 | B2 | 2006-01-10 | Rapid response system for the detection and treatment of cardiac events | 2002-01-22 | Disclosed is a system for detecting an acute myocardial infarction (i.e., a heart attack) at the earliest possible time and promptly warning the patient that he should immediately seek medical care. The present invention includes an implantable electronic system that can sense a change in the patient's electrogram that is indicative of a heart attack. If a heart attack is sensed, the device would then cause an implantable and/or externally located alarm to be actuated to warn the patient of his condition and a medical practitioner at a remote diagnostic center would receive the patient's electrogram for analysis. The patient or a caretaker would then be informed to self-inject medication through a subcutaneous, pass-through drug port that can be a separate device or integrated into the implanted device that is designed for the early detection of a heart attack. Since an implantable heart pacemaker or defibrillator already has within its structure many of the elements required for the device to recognize a heart attack, it would be expeditious to add a capability to these existing devices to detect a heart attack, have a pass-through drug port and provide an implantable and/or external alarm means to inform the patient to take appropriate action. | 1. An implanted medical device which is a cardiosaver device that is designed for rapid treatment of a heart attack of a human patient, the medical device including: an implanted sensor for detecting the occurrence of the heart attack of the human patient;an alarm that produces an alarm signal for indicating to the human patient that a heart attack is occurring when the sensor to which the alarm is electrically coupled detects the occurrence of a heart attack; and,a pass-through drug port implanted within the human patient, the drug port being designed to be used with an externally located drug delivery device for injecting medication into the bloodstream of the human patient at a time after the alarm indicates that a heart attack is occurring. | 1. An implanted medical device which is a cardiosaver device that is designed for rapid treatment of a heart attack of a human patient, the medical device including: an implanted sensor for detecting the occurrence of the heart attack of the human patient;an alarm that produces an alarm signal for indicating to the human patient that a heart attack is occurring when the sensor to which the alarm is electrically coupled detects the occurrence of a heart attack; and,a pass-through drug port implanted within the human patient, the drug port being designed to be used with an externally located drug delivery device for injecting medication into the bloodstream of the human patient at a time after the alarm indicates that a heart attack is occurring. 2. The cardiosaver device of claim 1 wherein the implantable sensor is two or more electrodes. 3. The cardiosaver device of claim 2 wherein one of the electrodes is a first electrode and the cardiosaver device include an electronics module that has a metal case that acts as a second electrode. 4. The cardiosaver device of claim 1 wherein the sensor has an output electrical signal which is the electrogram of the human patient and there is electronics circuitry contained within the electronics module, the electronics circuitry being designed to detect the occurrence of a heart attack from a change of the sensor's output electrical signal when compared to the sensor's output signal when no heart attack is occurring. 5. The cardiosaver device of claim 4 wherein the electronics circuitry within the electronics module includes a detector for detecting a variation of the ST segment of the electrogram as an indication of the heart attack. 6. The cardiosaver device of claim 4 wherein the electronics circuitry within the electronics module includes the capability for recording an electrogram signal from the sensor, the recorded electrogram being retained in memory for playback at a later time. 7. The cardiosaver device of claim 1 wherein the alarm of the cardiosaver system includes a subcutaneous electrical tickle which is an electrical stimulation. 8. The cardiosaver device of claim 1 wherein the alarm includes an implanted audio device having an alarm output that can be sensed by the human patient. 9. The cardiosaver device of claim 1 further including an external alarm system that is situated externally from the human patient for indicating to the human patient that he is experiencing a heart attack. 10. The cardiosaver device of claim 1 further including a self-sealing septum, drug chamber and exit port each being a part of the drug port, and the drug delivery device being a hypodermic syringe that includes a needle having a sharp point at its distal end, the sharp point being designed to be placed through the self-sealing septum of the drug port. 11. The cardiosaver device of claim 10 further including a medication delivery catheter that is joined to the exit port, the medication delivery catheter hav... | Angel Medical Systems, Inc. | A61B005/04
US07283859 | B2 | 2007-10-16 | Artifact suppression in dynamic magnetic resonance imaging | 2002-04-19 | A method for removing one or more artifacts from a time series of magnetic resonance (MR) images is provided. In one embodiment of the invention, an MR time series of data sets is acquired while changing the k-space that is sampled at each time frame. By using the acquired data, information is produced for a plurality of MR images at different time points, wherein the images contain desired components and one or more artifacts. Changing the k-space locations that are sampled from time frame to time frame makes the artifacts behave in a signature way through time, so that they can be identified as being artifacts and be removed through temporal analysis. | 1. A method for removing one or more leftover artifacts from a time series of dynamic magnetic resonance (MR) images, the method comprising: acquiring an MR data set by changing k-space locations that are sampled at each time point;separating desired components from one or more leftover artifacts by filtering the one or more leftover artifacts from the desired components in a temporal frequency domain;producing information for a first MR image at a first time point using a first MR reconstruction technique; andproducing information for a second MR image at a second time point using the first MR reconstruction technique. | 1. A method for removing one or more leftover artifacts from a time series of dynamic magnetic resonance (MR) images, the method comprising: acquiring an MR data set by changing k-space locations that are sampled at each time point;separating desired components from one or more leftover artifacts by filtering the one or more leftover artifacts from the desired components in a temporal frequency domain;producing information for a first MR image at a first time point using a first MR reconstruction technique; andproducing information for a second MR image at a second time point using the first MR reconstruction technique. 2. The method of claim 1, wherein separating the desired components from the one or more leftover artifacts comprises: (a) Fourier-transforming an image or k-space pixel over time to obtain a result of a Fourier transform;(b) filtering the result of the Fourier transform to separate the desired components from the one or more leftover artifacts; and(c) repeating steps (a) and (b) above for a plurality of image or k-space pixels. 3. A method of claim 2, wherein filtering the result of the Fourier transform comprises removing a frequency band that is located around a Nyquist frequency. 4. A method of claim 1, wherein the MR data set has one or more artifacts centered at a DC frequency in the temporal frequency domain, the method further comprising: producing information for a third MR image at the first time point using a second MR reconstruction technique on the MR data set; andproducing information for a fourth MR image at the second time point using the second MR reconstruction technique on the MR data set;wherein separating desired components from one or more leftover artifacts comprises:replacing a frequency band around the DC frequency in information treated with the first MR reconstruction technique with a corresponding frequency band from information treated with the second MR reconstruction technique. 5. The method of claim 1, wherein acquiring an MR data set by changing the k-space that is sampled at each time point comprises sampling a first set of k-space locations every odd frame and sampling a second set of k-space locations every even frame with a modified partial Fourier imaging technique. 6. The method of claim 1, wherein acquiring an MR data set by changing the k-space that is sampled at each time point comprises: sampling a first set of k-space locations every odd frame; andsampling a second set of k-space locations every even frame. 7. The method of claim 1, wherein the MR reconstruction technique is SENSE. 8. The method of claim 1, wherein the MR reconstruction technique is SMASH. 9. The method of claim 1, further comprising converting noise amplification into an artifact such that the artifact can be reduced. 10. The method according to claim 1, where the filtering in a temporal frequency domain is performed after the MR reconstruction technique. 11. The method according to claim 1, where the MR recon... | Brigham and Womens' Hospital, Inc. | A61B005/05
US07186222 | B1 | 2007-03-06 | Vascular introducer with temperature monitoring probe and systems for endovascular temperature control | 2002-09-10 | An introducer sheath/temperature probe assembly that is insertable into a blood vessel of a human or veterinary patent to measure the temperature of blood flowing through that blood vessel. The introducer sheath/temperature probe assembly may be used in conjunction with an indwelling heat exchange catheter system to warm or cool all or a portion of the patient's body to a desired target temperature and to maintain such target temperature for a desired period of time. | 1. An introducer sheath/temperature probe assembly comprising: a tubular introducer sheath having a proximal end, a distal end and a lumen which extends longitudinally therethrough, said introducer sheath being insertable into a blood vessel;an elongate temperature probe having a temperature sensor, a distal end and a proximal end, said temperature probe being insertable, distal end first, into the lumen of the introducer sheath and useable, when inserted into the lumen of the introducer sheath, to sense the temperature of blood flowing through the blood vessel into which the introducer sheath is inserted; anda tip member disposed on the distal end of the temperature probe such that the tip member protrudes beyond the distal end of the introducer sheath when the temperature probe is inserted into the lumen of the introducer sheath, said tip member being constructed so as not to cause substantial trauma to the wall of the blood vessel as the introducer sheath/temperature probe assembly is inserted into the blood vessel;said temperature sensor being located within said elongate temperature probe proximal of said tip member so that said tip member may extend substantially out of the distal end of said introducer sheath while said temperature sensor resides within said lumen of said introducer sheath;wherein said tubular introducer sheath has a proximal opening diameter on said proximal end, said temperature probe has a distal length having a first diameter and a proximal length having a second diameter, said first diameter smaller than said opening diameter and said second diameter larger than said opening diameter, and the length of the probe distal of said second diameter up to said sensor is less than the length of the lumen of said introducer sheath, and the length of said distal portion of said probe between said second diameter and the distal tip of said distal tip member is sized such that when said temperature probe is inserted through said proximal opening the portion of the temperature probe of said first diameter is inserted into the lumen of said introducer sheath but the portion of said probe having said second diameter will be blocked from entering said proximal opening, and when fully inserted into said introducer, said tip member will extend beyond the distal end of the introducer sheath but the sensor will remain within the lumen of said introducer sheath. | 1. An introducer sheath/temperature probe assembly comprising: a tubular introducer sheath having a proximal end, a distal end and a lumen which extends longitudinally therethrough, said introducer sheath being insertable into a blood vessel;an elongate temperature probe having a temperature sensor, a distal end and a proximal end, said temperature probe being insertable, distal end first, into the lumen of the introducer sheath and useable, when inserted into the lumen of the introducer sheath, to sense the temperature of blood flowing through the blood vessel into which the introducer sheath is inserted; anda tip member disposed on the distal end of the temperature probe such that the tip member protrudes beyond the distal end of the introducer sheath when the temperature probe is inserted into the lumen of the introducer sheath, said tip member being constructed so as not to cause substantial trauma to the wall of the blood vessel as the introducer sheath/temperature probe assembly is inserted into the blood vessel;said temperature sensor being located within said elongate temperature probe proximal of said tip member so that said tip member may extend substantially out of the distal end of said introducer sheath while said temperature sensor resides within said lumen of said introducer sheath;wherein said tubular introducer sheath has a proximal opening diameter on said proximal end, said temperature probe has a distal length having a first diameter and a proximal length having a second diameter, said first diameter smaller than said opening diameter and said second diameter larger than said opening diameter, and the length of the probe distal of said second diameter up to said sensor is less than the length of the lumen of said introducer sheath, and the length of said distal portion of said probe between said second diameter and the distal tip of said distal tip member is sized such that when said temperature probe is inserted through said proximal opening the portion of the temperature probe of said first diameter is inserted into the lumen of said introducer sheath but the portion of said probe having said second diameter will be blocked from entering said proximal opening, and when fully inserted into said introducer, said tip member will extend beyond the distal end of the introducer sheath but the sensor will remain within the lumen of said introducer sheath. 2. An introducer sheath/temperature probe assembly according to claim 1 further comprising a hub on the proximal end of the introducer sheath. 3. An introducer sheath/temperature probe assembly according to claim 2 wherein a side port is formed on the hub of the introducer sheath. 4. An introducer sheath/temperature probe assembly according to claim 1 wherein said tip member is coated with a lubricious coating. 5. An introducer sheath/temperature probe assembly according to claim 1 wherein said tip member is coated with a coating that is both lubricious and anti-thrombogenic. ... | Radiant Medical, Inc. | A61B005/00
US07161596 | B2 | 2007-01-09 | Display location calculation means | 2002-12-20 | Display location calculation means and methods for calculating a display location are disclosed. It has become common for users to indicate a point on a display in order to communicate with a machine. It is therefore necessary to be able to calculate the location on the display that is being indicated by the user. A display generator (123) is arranged in operation to generate a display in response to display data generated by a machine (111). A camera operable to generate image data representing at least part of the scene within the field of view of the camera (which part, includes at least a portion of the display) is carried in a pointer device (103). Computation means (111, 113, 115) are arranged in operation to receive the image data and the display data and to calculate from these data sets the position and/or orientation of the pointer device (103) relative to the display. A display location can then be calculated from the calculated position and/or orientation. | 1. Display location calculation means comprising: a display generator arranged in operation to generate a display in response to display data generated by a machine;a pointer device carrying a camera operable to generate image data representing at least part of the scene within the field of view of the camera, which part, includes an image of at least a portion of the display;computation means arranged in operation to:receive said image data;receive said display data;calculate, from said image data and said display data, the position and/or orientation of said pointer device relative to said display;calculate a display location from said calculated position and/or orientation. | 1. Display location calculation means comprising: a display generator arranged in operation to generate a display in response to display data generated by a machine;a pointer device carrying a camera operable to generate image data representing at least part of the scene within the field of view of the camera, which part, includes an image of at least a portion of the display;computation means arranged in operation to:receive said image data;receive said display data;calculate, from said image data and said display data, the position and/or orientation of said pointer device relative to said display;calculate a display location from said calculated position and/or orientation. 2. Display location calculation means according to claim 1 wherein said pointer device is elongate in shape and wherein said display location is the point where the longitudinal axis of said pointer device intersects with said display. 3. An interface arrangement for providing an interface between a user and a machine comprising: display location calculation means according to claim 2;wherein said computation means is further arranged in operation to control said machine in accordance with the position of the cursor. 4. An interface arrangement according to claim 3 wherein said machine comprises a computer having a processor and wherein said computation means comprises said processor. 5. An interface arrangement according to claim 3 wherein said pointer device further carries a computer having a processor and wherein said computation means comprises said processor. 6. An interface arrangement according to claim 3 wherein said machine comprises a first computer having a first processor and wherein said pointer device further carries a second computer having a second processor and wherein said computation means comprises said first processor and said second processor. 7. An interface arrangement according to claim 3 wherein said pointer device further comprises at least one activation means for use in controlling said machine. 8. An interface arrangement according to claim 3 wherein said display generator comprises a projector. 9. An interface arrangement according to claim 3 wherein said camera is a digital video camera. 10. An interface arrangement according to claim 3 wherein said camera is a fixed focus camera. 11. An interface arrangement according to claim 10 further comprising an indicator arranged in operation to indicate that the image captured by said camera is focussed. 12. An interface arrangement according to claim 3 wherein said camera has a fixed spatial relationship with said pointer device. 13. Display location calculation means according to claim 1 wherein a cursor is included in said display at said display location. 14. Display location calculation means according to claim 13 wherein the position of said cursor varies in accordance with said calculated position and/or orientation. 15. A method of calculating a display location said method com... | British Telecommunications public limited company | A61B005/00
US06983182 | B2 | 2006-01-03 | Time resolved computed tomography angiography | 2002-08-20 | A contrast enhanced dynamic study of a subject is performed with a CT system. A series of undersampled image data sets are acquired during the study with successive data sets acquired at interleaved projection angles. More fully sampled image data sets are formed by transforming the x-ray attenuation projection data to k-space and then sharing peripheral k-space data between undersampled k-space data sets. Artifacts due to undersampling are thus reduced without significantly affecting the time resolution of a series of reconstructed images. | 1. A method for producing an image with a computed tomography (CT) system, the steps comprising: a) positioning a subject in the CT system;b) acquiring a first undersampled image data set from a selected region of interest in the subject, the undersampled image data set comprising a plurality of projections acquired at a corresponding plurality of projection angles;c) transforming the undersampled image data set to form a corresponding first undersampled k-space data set;d) repeating steps b) and c) to produce a second undersampled k-space data set from a plurality of projections acquired at projection angles interleaved with the projection angles used to acquire the first undersampled image data set;e) combining k-space data from one of said undersampled k-space data sets with peripheral k-space data from the other undersampled k-space data set; andf) reconstructing an image of the region of interest by transforming the combined k-space data. | 1. A method for producing an image with a computed tomography (CT) system, the steps comprising: a) positioning a subject in the CT system;b) acquiring a first undersampled image data set from a selected region of interest in the subject, the undersampled image data set comprising a plurality of projections acquired at a corresponding plurality of projection angles;c) transforming the undersampled image data set to form a corresponding first undersampled k-space data set;d) repeating steps b) and c) to produce a second undersampled k-space data set from a plurality of projections acquired at projection angles interleaved with the projection angles used to acquire the first undersampled image data set;e) combining k-space data from one of said undersampled k-space data sets with peripheral k-space data from the other undersampled k-space data set; andf) reconstructing an image of the region of interest by transforming the combined k-space data. 2. The method as recited in claim 1 in which the region of interest includes vasculature and the method further includes: g) injecting a contrast agent which flows into the vasculature;h) repeating steps b), c), d), e) and f) to produce a contrast enhanced image of the region of interest; andi) subtracting the image of the region of interest from the contrast enhanced image of the region of interest to produce an angiogram. 3. The method as recited in claim 2 in which steps h) and i) are repeated to produce a time series of angiograms that depict the vasculature as the contrast agent flows into the vasculature. 4. The method as recited in claim 1 which includes repeating steps b) and c) to produce a third undersampled k-space data set from a plurality of projections acquired at projection angles interleaved with the projection angles used to acquire the first and second undersampled image data sets; and step e) is performed by combining substantially all the k-space data from the second undersampled k-space data set with peripheral k-space data from the first and third undersampled k-space data sets. 5. The method as recited in claim 1 in which step c) is performed by Fourier transforming each acquired projection. 6. The method as recited in claim 1 in which step f) includes: regridding the combined k-space and;Fourier transforming the regridded k-space data. 7. A method for producing an image of a subject with a computed tomography x-ray imaging system, the steps comprising: a) acquiring a set of projection data with the imaging system from a selected view angle with respect to the subject;b) changing the view angle by a first selected amount;c) repeating steps a) and b) to acquire a first plurality of sets of projection data; d) transforming the first plurality of sets of projection data to form a first undersampled k-space data set;e) changing the view angle by a second selected amount which is less than said first selected amount;f) repeating steps a) and b) to acquire a second plurality of sets o... | Wisconsin Alumni Research Foundation | A61B005/05
US07224472 | B2 | 2007-05-29 | Operation lamp with camera system for 3D referencing | 2002-04-29 | The invention relates to a system for the combined shadow-free illumination of a pre-definable area and for referencing three-dimensional spatial co-ordinates, and to an active or passive referencing system, each in particular for referencing surgical or medical instruments. The system is characterized in that at least two cameras and the light source (operation lamp) are held together such that the optical signals detected by the cameras, for referencing three-dimensional spatial co-ordinates in the area illuminated by the light source, can be evaluated. Since the field of view of the light source, in its conventional use, is not obscured or only negligibly obscured, this field of view can simultaneously be used for optical navigation by the cameras held together with the light source. In accordance with the invention, the cameras are automatically aligned optimally for optical navigation.;The cameras can be rigidly or flexibly connected to the light source. The position and/or orientation of the cameras can be adjustable. | 1. A system for the combined illumination of a medical workspace and referencing of three-dimensional spatial coordinates of surgical or medical instruments, comprising: an operation lamp producing shadow-free illumination in the visible spectrum for the shadow-free illumination of said medical workspace; andat least two cameras for detecting optical signals, wherein said operation lamp and said at least two cameras are mounted with respect to one another such that when the operation lamp is moved the cameras are correspondingly moved to retain the same positional relationship, for referencing three-dimensional spatial coordinates in said illuminated medical workspace. | 1. A system for the combined illumination of a medical workspace and referencing of three-dimensional spatial coordinates of surgical or medical instruments, comprising: an operation lamp producing shadow-free illumination in the visible spectrum for the shadow-free illumination of said medical workspace; andat least two cameras for detecting optical signals, wherein said operation lamp and said at least two cameras are mounted with respect to one another such that when the operation lamp is moved the cameras are correspondingly moved to retain the same positional relationship, for referencing three-dimensional spatial coordinates in said illuminated medical workspace. 2. The system as set forth in claim 1, wherein the orientation and/or position of the at least two cameras may be adjusted relative to said operation lamp. 3. The system as set forth in claim 1, wherein the at least two cameras are arranged substantially point-symmetrically with respect to said center of said operation lamp. 4. The system of claim 1, further comprising a radiation source for emitting infrared radiation, and the at least two cameras detect the infrared radiation. 5. A system for the combined illumination of a medical workspace and referencing of three-dimensional spatial coordinates of surgical or medical instruments, comprising: an operation lamp producing shadow-free illumination in the visible spectrum for the shadow-free illumination of said medical workspace; andat least two cameras for detecting optical signals, wherein said operation lamp and said at least two cameras are mounted together for common movement such that said optical signals may be evaluated, for referencing three-dimensional spatial coordinates in said illuminated medical workspace, wherein the at least two cameras are arranged outside a casing of the operation lamp and near a circumferential edge of said casing, such that the sight of said at least two cameras onto said illuminated area is not obstructed by sections of said casing. 6. The system as set forth in claim 5, wherein the at least two cameras are integrated in said casing of said operation lamp, viewing windows being respectively provided for said at least two cameras on sections of said casing where no illuminating means of said operation lamp are arranged. 7. The system as set forth in claim 5, wherein the at least two cameras are external to said casing of said operation lamp. 8. A system for the combined illumination of a medical workspace and referencing of three-dimensional spatial coordinates of surgical or medical instruments, comprising: an operation lamp producing shadow-free illumination in the visible spectrum for the shadow-free illumination of said medical workspace; andat least two cameras for detecting optical signals, wherein said operation lamp and said at least two cameras are held together such that said optical signals may be evaluated, for referencing three-dimensional spatial coordinates in said illumina... | BrainLAB AG | G01B011/24 G01B011/30 G01B011/14 A61B001/06 A61B005/05
US06959212 | B2 | 2005-10-25 | System and method for arrhythmia discrimination | 2001-10-22 | The present invention provides a system and a method for discriminating supraventricular tachyarrhythmias from ventricular arrhythmias during a tachycardia episode. First cardiac signals and second cardiac signals are sampled for cardiac complexes. A first feature on the first cardiac signal and a second feature on the second cardiac signal are utilized to determine an average time difference for a plurality of normal sinus rhythm complexes. A time difference between the first feature and the second feature is then determined for each cardiac complex of a tachycardiac rhythm. The cardiac complex is characterized as a ventricular tachycardia complex if the time difference exceeds the average time difference by a predetermined amount. Otherwise, it is classified as VT if its morphology after alignment is different from that during normal sinus rhythm. | 1. A method, comprising: sensing a first cardiac complex of a heart contraction at a first cardiac region of a patient; sensing a second cardiac complex, associated with the same heart contraction as the first cardiac complex, at a second cardiac region of said patient; selecting a first fiducial feature of the first cardiac complex; selecting a second fiducial feature of the second cardiac complex; determining a time difference between the first and second fiducial features; comparing the time difference to a template time difference between the first and second fiducial features obtained from said patient during normal sinus rhythm; and if the time difference differs from the template time difference by at least a predetermined margin, characterizing the first and second signals as arrhythmic. | 1. A method, comprising: sensing a first cardiac complex of a heart contraction at a first cardiac region of a patient; sensing a second cardiac complex, associated with the same heart contraction as the first cardiac complex, at a second cardiac region of said patient; selecting a first fiducial feature of the first cardiac complex; selecting a second fiducial feature of the second cardiac complex; determining a time difference between the first and second fiducial features; comparing the time difference to a template time difference between the first and second fiducial features obtained from said patient during normal sinus rhythm; and if the time difference differs from the template time difference by at least a predetermined margin, characterizing the first and second signals as arrhythmic. 2. The method of claim 1, further including: if the time difference exceeds the template time difference by at least a predetermined margin, characterizing the heart contraction as a tachycardia heart contraction; determining a percentage of tachycardia heart contractions in a plurality of heart contractions; and when the percentage of tachycardia heart contractions exceeds a threshold, delivering therapy for treating the tachycardia. 3. The method of claim 1, wherein the first and second fiducial features include at least one of: a predetermined signal deviation from a baseline representing a beginning of the cardiac complex; a maximum deflection of the cardiac complex; a signal return to the baseline within a predetermined time window from the beginning of the cardiac complex, the signal return representing an end of the cardiac complex; and a maximum slope portion of the cardiac complex. 4. The method of claim 1, wherein the template time difference is obtained from said patient using the time differences determined for a plurality of heart contractions during normal sinus rhythm. 5. A method, comprising: sensing a first cardiac complex at a first cardiac region during a cardiac cycle of a patient; sensing a second cardiac complex at a second cardiac region during the cardiac cycle; selecting a first fiducial feature of the first cardiac complex; selecting a second fiducial feature of the second cardiac complex; determining a first amplitude for the first fiducial feature relative to a baseline of the first cardiac complex; determining a second amplitude for the second fiducial feature relative a baseline of the second cardiac complex; comparing the first and second amplitudes to respective first and second thresholds obtained from said patient during normal sinus rhythm; and characterizing whether the first and second cardiac complexes are arrhythmic using the results of the comparison. 6. The method of claim 5, further comprising: determining the first threshold using a statistical value obtained from the first amplitudes over a plurality of cardiac cycles during normal sinus rhythm; and determining the second threshold using a statistical v... | Cardiac Pacemakers, Inc. | A61B005/0402
US06915152 | B2 | 2005-07-05 | Method for MR imaging with an array of RF coils | 2001-04-19 | A method and apparatus for producing an image of a region of interest using a Magnet Resonance Imaging (MRI) system comprises the steps of acquiring a plurality of under-sampled MR data sets from a plurality of radiofrequency (RF) coils forming an array and combining the plurality of under-sampled MR data sets to produce the image. Each of the RF coils has a non-uniform sensitivity. | 1. A method for producing an image of a region of interest using a Magnet Resonance Imaging (MRI) system comprising: acquiring a plurality of under-sampled Magnetic Resonance (MR) data sets from an array of radiofrequency (RF) coils in said MRI system, each of said RF coils having a non-uniform sensitivity profile; and, reconstructing a plurality of regional images from said acquired plurality of under-sampled MR data sets, said reconstructing of each respective regional image being performed substantially simultaneously with one another; and, weighting and summing said respective regional images to produce said image of said region of interest. | 1. A method for producing an image of a region of interest using a Magnet Resonance Imaging (MRI) system comprising: acquiring a plurality of under-sampled Magnetic Resonance (MR) data sets from an array of radiofrequency (RF) coils in said MRI system, each of said RF coils having a non-uniform sensitivity profile; and, reconstructing a plurality of regional images from said acquired plurality of under-sampled MR data sets, said reconstructing of each respective regional image being performed substantially simultaneously with one another; and, weighting and summing said respective regional images to produce said image of said region of interest. 2. The method of claim 1 wherein each of said regional images is aliased and the weighting and summing step substantially eliminates aliasing in said image of said region of interest. 3. The method of claim 1 wherein said RF coils in said array are arranged linearly relative to each other. 4. The method of claim 1 wherein said region of interest is a three-dimensional volume. 5. The method of claim 1 wherein said region of interest is a two-dimensional plane. 6. The method of claim 1 wherein the reconstructing step is performed in accordance with: M⁢(x,y,z)=∑n⁢Mn⁢(x,y,z)⁢h⁢(z-n⁢ ⁢Δz)and Mn⁢(x,y,z)=∑m⁢(∫∫Sn⁢(kx,ky,(m-ɛ)⁢Δk⁢ ⁢z)⁢ⅇj2π⁢(kx⁢x+ky⁢y)⁢ⅆkx⁢ⅆky)⁢ⅇj2π⁢(m-ɛ)⁢Δk⁢ ⁢z⁢zwhere h(z) is a reconstruction weighting function that is pre-derived from Δz, Δkz, and w(z), w(z) being the non-uniform sensitivity profile. 7. The method of claim 6 wherein h(z) is derived in accordance with singular value decomposition techniques that minimize ∥h∥ subject to ∥Ah−e1∥&lt;τ, where τ is a scalar representing a maximum level of tolerable residual aliasing. 8. An imaging apparatus for producing Magnetic Resonance (MR) images of a subject having a magnet assembly for producing a static magnetic field and a gradient coil assembly disposed within said magnet assembly for generating a magnetic field gradient for use in producing MR images, said apparatus comprising; a plurality of radiofrequency (RF) coils forming an array and said RF coils having overlapped non-uniform sensitivity profiles and acquiring in parallel under-sampled Magnetic Resonance (MR) data sets; a processor for combining said plurality of under-sampled MR data sets to produce said image by reconstructing a plurality of regional images from said plurality of under-sampled MR data sets, said reconstructing of each respective regional image being performed substantially simultaneously with one another, and weighting and summing said respective regional images to produce said image of said region of interest. 9. The apparatus of claims 8 wherein said processor reconstructs a plurality of aliased regional images corresponding to said plurality of under-sampled MR data sets. 10. The apparatus of claim 8 wherein said processor combines said aliased regional images by weighting and summing such that said image of the region of interest is substantially fr... | General Electric Company | A61B005/05 G01V003/00
US07031765 | B2 | 2006-04-18 | Algorithms for detecting atrial arrhythmias from discriminatory signatures of ventricular cycle lengths | 2002-11-11 | Detection of arrhythmias is facilitated using irregularity of ventricular beats measured by delta-RR (ΔRR) intervals that exhibit discriminatory signatures when plotted in a Lorenz scatter-plot. An “AF signature metric” is established characteristic of episodes of AF that exhibit highly scattered (sparse) distributions or formations of 2-D data points. An “AFL signature metric” is established characteristic of episodes of AFL that exhibit a highly concentrated (clustered) distribution or formation of 2-D data points. A set of heart beat interval data is quantified to generate highly scattered (sparse) formations as a first discrimination metric and highly concentrated (clustered) distributions or formations as a second discrimination metric. The first discrimination metric is compared to the AF signature metric, and/or the second discrimination metric is compared to the AFL signature metric. AF or HFL is declared if the first discrimination metric satisfies either one of the AF signature metric. | 1. A method of detecting atrial arrhythmias, comprising one or more of atrial fibrillation (AF) and atrial flutter (AFL) within a segment of ventricular heartbeats signified by intervals between successive ventricular heartbeats that exhibit discriminatory signatures when plotted as data points in a scatter-plot comprising: defining an AF signature metric for AF representative of the sparse distribution of data points during AF episodes and an AFL signature metric representative of the clustering distribution of data points during AFL episodes;determining each interval between successive ventricular heartbeats;plotting each succeeding interval as a data point in a scatter-plot;developing a discriminatory metric signifying the degree of sparseness or clustering of the data points of the scatter-plot;declaring AF if the discriminatory metric satisfies the AF signature metric representative of the distribution of data points during AF; anddeclaring AFL if the discriminatory metric satisfies the AFL signature metric representative of the distribution of data points during AFL. | 1. A method of detecting atrial arrhythmias, comprising one or more of atrial fibrillation (AF) and atrial flutter (AFL) within a segment of ventricular heartbeats signified by intervals between successive ventricular heartbeats that exhibit discriminatory signatures when plotted as data points in a scatter-plot comprising: defining an AF signature metric for AF representative of the sparse distribution of data points during AF episodes and an AFL signature metric representative of the clustering distribution of data points during AFL episodes;determining each interval between successive ventricular heartbeats;plotting each succeeding interval as a data point in a scatter-plot;developing a discriminatory metric signifying the degree of sparseness or clustering of the data points of the scatter-plot;declaring AF if the discriminatory metric satisfies the AF signature metric representative of the distribution of data points during AF; anddeclaring AFL if the discriminatory metric satisfies the AFL signature metric representative of the distribution of data points during AFL. 2. The method of claim 1, wherein the scatter plot is replaced by a 3-D plot having a first interval on the X axis, a succeeding interval on the Y axis, and a net succeeding interval on the Z axis. 3. The method of claim 2, wherein the scatter plot is replaced by a multidimensional plot having multi-dimensional axis in successive orders thereof. 4. The method of claim 1, wherein the scatter-plot is a Lorenz plot having an origin, and the plotting step further comprises assigning an abscissa value and an ordinate value referenced to the origin of the Lorenz plot to each data point. 5. The method of claim 1, wherein the step of determining the interval comprises determining the change in interval between each succeeding ventricular heart beat from a preceding ventricular heartbeat. 6. The method of claim 4 or 5, wherein rather than using RR or ΔRR, the absolute value RR or the absolute value of ΔRR is used. 7. The method of claim 4 or 5, wherein rather than using RR or ΔRR, a non-linear function RR or a no-linear function of ΔRR is used, such as the log of the absolute value multiplied by the sign of the value. 8. The method of claim 5, wherein the step of determining the interval includes one interval axis and a change between intervals axis. 9. The method of claim 5, wherein the scatter-plot is a Lorenz plot having an origin, and the plotting step further comprises assigning an abscissa value and an ordinate value referenced to the origin of the Lorenz plot to each data point, whereby the data point is a two-dimensional (2-D) data point. 10. The method of claim 9, further comprising the steps of defining 2-D bins of the Lorenz plot with reference to the origin of the Lorenz plot and correlating each 2-D data point of the Lorenz plot with a bin of the Lorenz plot that the 2-D data point falls within and wherein: the step of developing a discriminatory metric further com... | Medtronic, Inc | A61B005/046
US06988996 | B2 | 2006-01-24 | Test media cassette for bodily fluid testing device | 2002-06-07 | A bodily fluid sampling device includes a piercing device and a sensor enclosed in a housing. A cassette, which contains test media, is positioned proximal to the sensor so that the sensor is able to analyze a bodily fluid sample collected on the test media. The cassette includes a supply portion from which unused test media is supplied and a storage portion in which contaminated test media is stored after exposure to the bodily fluid. The cassette is adapted to collect a series of bodily fluid samples without requiring disposal of the test media. | 1. A bodily fluid sampling device for analyzing a bodily fluid, comprising: a test media cassette including a test media tape adapted to collect the bodily fluid, a supply portion storing an uncontaminated section of the test media tape that is uncontaminated with the bodily fluid, a storage portion storing a contaminated section of the test media tape that is contaminated with the bodily fluid, an exposure portion positioned between the supply portion and the storage portion, the exposure portion being adapted to expose a section of the test media tape to the bodily fluid; a sensor positioned between the supply portion and the storage portion to sense at least one property of the bodily fluid collected on the test media tape at the exposure portion of the cassette; a piercing device adapted to pierce skin; and the test media tape defining a throughway to allow passage of the piercing device through the test media tape during piercing. | 1. A bodily fluid sampling device for analyzing a bodily fluid, comprising: a test media cassette including a test media tape adapted to collect the bodily fluid, a supply portion storing an uncontaminated section of the test media tape that is uncontaminated with the bodily fluid, a storage portion storing a contaminated section of the test media tape that is contaminated with the bodily fluid, an exposure portion positioned between the supply portion and the storage portion, the exposure portion being adapted to expose a section of the test media tape to the bodily fluid; a sensor positioned between the supply portion and the storage portion to sense at least one property of the bodily fluid collected on the test media tape at the exposure portion of the cassette; a piercing device adapted to pierce skin; and the test media tape defining a throughway to allow passage of the piercing device through the test media tape during piercing. 2. The bodily fluid sampling device of claim 1, wherein: the supply portion includes a supply reel, wherein the uncontaminated section of the test media tape is wound on the supply reel; and the storage portion includes a storage reel, wherein the contaminated section of the test media tape is wound on the storage reel. 3. The bodily fluid sampling device of claim 2, further comprising: wherein the test media tape includes a cover tape to preserve the test media tape; and a cover reel attached to the cover tape, the cover reel being adapted to peel the cover tape from the test media tape before exposure to the bodily fluid. 4. The bodily fluid sampling device of claim 1, wherein the test media tape includes: a test pad containing chemicals for testing the bodily fluid; and a blister pack positioned proximal to the test pad, the blister pack having a raised surface adapted to draw the bodily fluid to the test pad through capillary action. 5. The bodily fluid sampling device of claim 1, further comprising: a vacuum assembly to generate a vacuum; a piercing device adapted to pierce skin and draw the bodily fluid; wherein the test media tape includes separate test sections, each test section including a vacuum passageway defined in the test media and a seal membrane, the seal membrane being adapted to be pierced and seal around the piercing device when the bodily fluid is collected, each test section further including a test pad positioned between the vacuum passageway and the seal membrane, the test pad being adapted to collect the bodily fluid; and wherein the vacuum assembly being adapted to form a vacuum in the vacuum cavity to draw the bodily fluid up the piercing device and onto the test pad during testing. 6. The bodily fluid sampling device of claim 1, further comprising: a piercing device mounted on a first pivotal arm to pierce skin; and wherein the cassette is coupled to a second pivotal arm to pivot the cassette into position to collect the bodily fluid after the piercing device pierces the skin.... | Roche Diagnostics Operatons, Inc. | A61B005/00 B65D081/00
US07146207 | B2 | 2006-12-05 | Health administration apparatus | 2001-10-19 | A health administration apparatus 10 is provided, which is capable of dispensing with a troublesome measuring operation such as to wrap a tape measure or the like around a part of concern of a human body by automatically finding the abdominal circumference (waist size) or the lumbar circumference (hip size) of a human body. The health administration apparatus 10 is configured to allow entry of body data of an examinee, including the height and weight of the examinee, and to be capable of finding the abdominal circumference of the examinee which is the circumferential size of a trunk part of the examinee by a computation based on the body data including the height and the weight of the examinee. | 1. A health administration apparatus comprising: an input device configured to receive input of body data of an examinee including the height and weight of the examinee;impedance measurement device for measuring the impedance of the body of the examinee;a computing section including a storage device configured to store a formula for computing the abdominal circumference of an examinee based on body data of the examinee including the height and weight of the examinee, and based on the measured impedance of the body of the examinee, and a processing unit configured to apply the formula to compute the abdominal circumference of the examinee based on the inputted body data including the height and weight of the examinee, the computed abdominal circumference being an estimated circumferential size of a trunk part of the examinee; anda display section configured to display the computed abdominal circumference. | 1. A health administration apparatus comprising: an input device configured to receive input of body data of an examinee including the height and weight of the examinee;impedance measurement device for measuring the impedance of the body of the examinee;a computing section including a storage device configured to store a formula for computing the abdominal circumference of an examinee based on body data of the examinee including the height and weight of the examinee, and based on the measured impedance of the body of the examinee, and a processing unit configured to apply the formula to compute the abdominal circumference of the examinee based on the inputted body data including the height and weight of the examinee, the computed abdominal circumference being an estimated circumferential size of a trunk part of the examinee; anda display section configured to display the computed abdominal circumference. 2. The health administration apparatus according to claim 1, wherein the computing section is further configured to find a visceral fat cross-sectional area by a computation based on the abdominal circumference found. 3. The health administration apparatus according to claim 2, wherein the computing section is further configured to form an obesity judgment based on the visceral fat cross-sectional area found. 4. The health administration apparatus according to claim 1, wherein the computing section is further configured to form an obesity judgment based on the abdominal circumference of the examinee. 5. The health administration apparatus according to claim 1, wherein the computing section is further configured to find the visceral fat cross-sectional area of the examinee by a computation on the abdominal circumference found with a predetermined computing formula; wherein said predetermined computing formula is created by statistical processing of correlations between the actual visceral fat cross-sectional areas, as measured by tomography, of a statistical sample of human bodies, and the abdominal circumferences of the human bodies in the sample. 6. A health administration apparatus comprising: an input device configured to receive entry of body data of an examinee including the height and weight of the examinee;impedance measurement means for measuring the impedance of the body of the examinee via electrodes brought into contact with extremities of the body of the examinee;a storage device configured to store a formula for computing the abdominal circumference of an examinee based on body data of the examinee including the height, weight, and impedance of the body of the examinee;a processing unit configured to apply this formula to compute the abdominal circumference of the examinee based on entered body data including the height and weight, and measured impedance of the body of the examinee, the computed abdominal circumference being an estimated circumferential size of a trunk part of the examinee; anda display section configured to dis... | Yamato Scale Co., Ltd. | A61B005/05
US07130676 | B2 | 2006-10-31 | Fluoroscopic image guided orthopaedic surgery system with intraoperative registration | 2002-08-29 | A fluoroscopic image guided surgery system, comprising a C-arm fluoroscope for obtaining fluoroscopic images of an object bone, the C-arm fluoroscope including at least one set of emitters; a reference bar capable of attaching to an object bone, the reference bar including emitters; a surgical instrument for performing an operation, the instrument including emitters; a digitizer system in communication with the at least one set of emitters of the C-arm fluoroscope, the emitters of the reference bar, and the emitters of the surgical instrument so that the digitizer system can determine a position of each of the C-arm fluoroscope, the reference bar, and the surgical instrument; and a single fiducial marker for attachment to an object bone, the single fiducial marker being visible in the fluoroscopic images for determining a position of an object bone relative to the digitizer system. | 1. An image guided surgery system comprising: an imaging device operable to obtain images of a region of a patient;a reference device operable to be positioned in the region of the patient, said reference device including at least one localization sensor being positioned on said reference device;a surgical instrument operable for use in performing an operation, said surgical instrument including at least one localization sensor positioned on said surgical instrument;a localization system operable to determine a position of said surgical instrument with respect to said reference device by use of said at least one localization sensor associated with said surgical instrument and said at least one localization sensor associated with said reference device; andonly one fiducial marker for attachment to the region, said only one fiducial marker being visible in the image of the region of the patient, wherein said only one fiducial marker is used for determining the position of the region relative to the localization system. | 1. An image guided surgery system comprising: an imaging device operable to obtain images of a region of a patient;a reference device operable to be positioned in the region of the patient, said reference device including at least one localization sensor being positioned on said reference device;a surgical instrument operable for use in performing an operation, said surgical instrument including at least one localization sensor positioned on said surgical instrument;a localization system operable to determine a position of said surgical instrument with respect to said reference device by use of said at least one localization sensor associated with said surgical instrument and said at least one localization sensor associated with said reference device; andonly one fiducial marker for attachment to the region, said only one fiducial marker being visible in the image of the region of the patient, wherein said only one fiducial marker is used for determining the position of the region relative to the localization system. 2. The image guided surgery system as defined in claim 1 wherein said imaging device is a fluoroscopic imaging device. 3. The image guided surgery system as defined in claim 2 wherein said fluoroscopic imaging device is a C-arm fluoroscopic imaging device. 4. The image guided surgery system as defined in claim 3 wherein said C-arm fluoroscopic imaging device includes an x-ray generator and an x-ray receiver operable to obtain fluoroscopic images of the region of the patient. 5. The image guided surgery system as defined in claim 1 wherein said localizing sensors are emitters. 6. The image guided surgery system as defined in claim 5 wherein said emitters are optical emitters. 7. The image guided surgery system as defined in claim 6 wherein said optical emitters are LED optical emitters. 8. The image guided surgery system as defined in claim 1 wherein said localization system is an optical localization system. 9. The image guided surgery system as defined in claim 8 wherein said optical localization system includes an optical digitizer camera. 10. The image guided surgery system as defined in claim 1 wherein the region of the patient is an object bone in the patient. 11. The image guided surgery system as defined in claim 10 wherein said only one fiducial marker and a shape of the object bone is used for determining position of the object bone relative to the localization system. 12. The image guided surgery system as defined in claim 1 further comprising a computer programmed to match images from said imaging device to effect a two-dimensional image registration. 13. The image guided surgery system as defined in claim 1 further comprising a computer programmed for determining a plane in which a patient body lies in relation to said reference device. 14. The image guided surgery system as defined in claim 1 wherein said imaging device further includes at least one localization sensor being positioned on said imaging device... | Sofamor Danek Holdings, Inc. | A61B005/00
US07228165 | B1 | 2007-06-05 | Apparatus and method for performing a tissue resection procedure | 2000-06-26 | An apparatus and a method for performing a medical procedure is disclosed. In an embodiment for an apparatus of the present invention, the apparatus includes a positioning system and a resection device disposed within the positioning system. The positioning system includes an imaging device, a video processor coupled to the imaging device, a computer coupled to the video processor, and a video display coupled to the computer. In an embodiment for a method of practicing the present invention, the method includes the steps of creating an image of a lesion within a patient's body on an imaging device. Data representative of the lesion image is processed by a video processor. Tissue margins around the lesion are defined by a processor based on the processed data representative of the lesion image and a resection device is operated during a resection procedure within the patient's body based upon the defined tissue margins. | 1. A tissue resectioning system, comprising: a resection head mounted at a distal end of an elongate flexible body, the resection head including a marker thereon wherein, when in an operative position, the resection head is located within a body lumen with the elongate flexible body extending through the body lumen to a naturally occurring body orifice;an imager which remains outside the patient's body, the imager generating image data of a selected region within the patient's body including a predetermined portion of tissue marked for resection;an image processing unit analyzing the image data to define a region of tissue to be resectioned and to locate the marker; anda control unit controlling the resection head based on the defined region of tissue and the location of the marker to resect the region of tissue. | 1. A tissue resectioning system, comprising: a resection head mounted at a distal end of an elongate flexible body, the resection head including a marker thereon wherein, when in an operative position, the resection head is located within a body lumen with the elongate flexible body extending through the body lumen to a naturally occurring body orifice;an imager which remains outside the patient's body, the imager generating image data of a selected region within the patient's body including a predetermined portion of tissue marked for resection;an image processing unit analyzing the image data to define a region of tissue to be resectioned and to locate the marker; anda control unit controlling the resection head based on the defined region of tissue and the location of the marker to resect the region of tissue. 2. The system according to claim 1 wherein the imager includes a fluoroscope and an x-ray imaging sensor. 3. The system according to claim 1 wherein the marker is radiopaque. 4. The system according to claim 1 wherein the defined region of tissue and the location of the marker are displayed on a video display coupled with the control unit. 5. The system according to claim 1 wherein the control unit disables the resection head if the marker indicates that the resection head is oriented outside the defined region of tissue. 6. The system according to claim 1 further comprising an alarm device wherein the control unit transmits an alarm signal to the alarm device when the marker indicates that the resection head is oriented outside the defined region of tissue. 7. The system according to claim 6 wherein the alarm device generates a visual alarm. 8. The system according to claim 1 wherein the imager is a magnetic resonance imager. 9. The system according to claim 1 wherein the image processing unit determines the defined region of tissue by an absolute measure of tissue. 10. The system according to claim 1 wherein the image processing unit determines the defined region of tissue by a percentage of a physical dimension of the lesion. | Boston Scientific Scimed, Inc. | A61B005/00
US07174202 | B2 | 2007-02-06 | Medical navigation apparatus | 2002-12-17 | Methods and apparatus are provided for locating the position, preferably in three dimensions, of a sensor by generating magnetic fields which are detected at the sensor. The magnetic fields are generated from a plurality of locations and, in one embodiment of the invention, enable both the orientation and location of a single coil sensor to be determined. The present invention thus finds application in many areas where the use of prior art sensors comprising two or more mutually perpendicular coils is inappropriate. | 1. A medical navigation apparatus for navigating an anatomical body during a medical procedure, the medical navigation apparatus comprising: a medical instrument operable to be navigated through the anatomical body during the medical procedure;at least one field sensor associated with the medical instrument and operable to be moved in association with the medical instrument;at least one magnetic field source operable to radiate at least one magnetic field into the anatomical body and operable to be detected by the at least one field sensor;a controller operable to receive from the at least one field sensor signals indicative of the at least one magnetic field and operable to continuously ascertain positions of the at least one field sensor based upon the at least one magnetic field detected by the at one least field sensor; anda display operable to continuously display representations of the position of the at least one field sensor associated with the medical instrument as it moves from a first position to a second position to form continuous representations, wherein the continuous representations include a representation of the first position, a representation of the second position, and representations of the positions between the first position and the second positionwherein the continuous representations of the position of the at least one field sensor is operable to be rotated in any direction on the display to illustrate the continuous representations from a plurality of perspectives of the continuous representations. | 1. A medical navigation apparatus for navigating an anatomical body during a medical procedure, the medical navigation apparatus comprising: a medical instrument operable to be navigated through the anatomical body during the medical procedure;at least one field sensor associated with the medical instrument and operable to be moved in association with the medical instrument;at least one magnetic field source operable to radiate at least one magnetic field into the anatomical body and operable to be detected by the at least one field sensor;a controller operable to receive from the at least one field sensor signals indicative of the at least one magnetic field and operable to continuously ascertain positions of the at least one field sensor based upon the at least one magnetic field detected by the at one least field sensor; anda display operable to continuously display representations of the position of the at least one field sensor associated with the medical instrument as it moves from a first position to a second position to form continuous representations, wherein the continuous representations include a representation of the first position, a representation of the second position, and representations of the positions between the first position and the second positionwherein the continuous representations of the position of the at least one field sensor is operable to be rotated in any direction on the display to illustrate the continuous representations from a plurality of perspectives of the continuous representations. 2. The medical navigation apparatus as defined in claim 1 wherein the medical instrument is a catheter. 3. The medical navigation apparatus as defined in claim 1 wherein the medical instrument is an endoscope. 4. The medical navigation apparatus as defined in claim 1 wherein the medical instrument is a colonoscope. 5. The medical navigation apparatus defined in claim 1 wherein the medical instrument is used for a cardiac medical procedure. 6. The medical navigation apparatus as defined in claim 1 wherein the at least one field sensor includes a single sensing coil. 7. The medical navigation apparatus as defined in claim 6 wherein the single sensing coil is an air-core coil. 8. The medical navigation apparatus as defined in claim 6 wherein the single sensing coil includes a magnetic permeable core. 9. The medical navigation apparatus as defined in claim 1 wherein the at least one field sensor is a flux gate magnetometer. 10. The medical navigation apparatus as defined in claim 1 wherein the at least one field sensor includes three sensing coils oriented orthogonal to one another about a common origin. 11. The medical navigation apparatus as defined in claim 1 wherein the at least one magnetic field is a quasi-static magnetic field. 12. The medical navigation apparatus as defined in claim 1 wherein the medical procedure is selected from a group comprising a colonoscopy, a bronchoscopy, a gastroscopy, a nasogastric pr... | British Telecommunications | A61B005/05
US07014617 | B2 | 2006-03-21 | Pivoted tensiometer for measuring tension in an intervertebral disc space | 2002-09-20 | The present invention relates to a novel pivoted tensiometer adapted for use in measuring tension in an intervertebral disc space. | 1. A tensiometer for measuring tension in an intervertebral disc space, comprising: a) a first longitudinal member having a distal end portion adapted for engaging a first vertebral surface, an intermediate portion, and a proximal handle portion having an attachment point,b) a second longitudinal member having a distal end portion adapted for engaging a second vertebral surface, an intermediate portion, and a proximal handle portion having an attachment point, the first and second longitudinal members being pivotally attached at a first pivot junction between the proximal handle and intermediate portions of each longitudinal member, andc) a load cell attached to each of the attachment points and adapted to report a load corresponding to a change in distance between the attachment points,wherein each distal end portion comprises a distal tip adapted to enter the disc space and distract the disc space when moved apart from each other, wherein the distal tips comprise two substantially parallel opposed plates adapted for distraction of the vertebral disc space. | 1. A tensiometer for measuring tension in an intervertebral disc space, comprising: a) a first longitudinal member having a distal end portion adapted for engaging a first vertebral surface, an intermediate portion, and a proximal handle portion having an attachment point,b) a second longitudinal member having a distal end portion adapted for engaging a second vertebral surface, an intermediate portion, and a proximal handle portion having an attachment point, the first and second longitudinal members being pivotally attached at a first pivot junction between the proximal handle and intermediate portions of each longitudinal member, andc) a load cell attached to each of the attachment points and adapted to report a load corresponding to a change in distance between the attachment points,wherein each distal end portion comprises a distal tip adapted to enter the disc space and distract the disc space when moved apart from each other, wherein the distal tips comprise two substantially parallel opposed plates adapted for distraction of the vertebral disc space. 2. The tensiometer of claim 1 wherein each intermediate portion and distal tip has a length, and the length of the intermediate portion is at least 10 times the length of the corresponding distal tip. 3. The tensiometer of claim 1 wherein at least one of the distal end portions comprises a proximally-positioned stop. 4. The tensiometer of claim 1 wherein each intermediate portion is substantially rigid. 5. The tensiometer of claim 1 wherein the first intermediate portion comprises a second pivot, the second intermediate portion comprises a third pivot, wherein the second and third pivots are positioned to form a double action pivot. 6. The tensiometer of claim 1 wherein the distal end portion further comprises first and second cross bars proximally pivotally attached to the respective first and second longitudinal members; pivotally attached together at a fourth pivot; and slidably distally attached to the respective second and first longitudinal members. 7. The tensiometer of claim 1 wherein the first pivot is located from a proximal handle end of the tensiometer a distance of between about 10–50% of an overall length of the tensiometer. 8. The tensiometer of claim 1 wherein the first pivot is adapted so that, when the proximal handle portions are squeezed together, the proximal portions of the intermediate portions move together. 9. The tensiometer of claim 1 wherein the first pivot is adapted so that, when the proximal handle portions are squeezed together, the proximal portions of the intermediate portions move apart. 10. The tensiometer of claim 1 wherein the proximal handle portions comprise an angled portion. 11. The tensiometer of claim 1 wherein the angled portion forms an angle of between 120 and 150 degrees. 12. The tensiometer of claim 1 wherein the load cell comprises a compression spring. 13. The tensiometer of claim 1 wherein the load cell comprises an expansion spr... | Depuy Acromed, Inc. | A61B005/103
US07118559 | B2 | 2006-10-10 | Body fluid collection apparatus | 2002-05-24 | A surgical fluid collection apparatus includes a housing member defining an internal chamber for collection of body fluids and a holder member releasably couplable to the housing member. The holder member includes an internal penetrating member defining an internal conduit in fluid communication with a fluid outlet of the housing member when the holder member is coupled to the housing member to permit passage of the body fluids therethrough. A collection member is positionable with respect to the holder member to receive the body fluids passing through the internal conduit of the internal penetrating member. | 1. A fluid collection apparatus comprising: a housing configured for receipt of fluid and having a first surface which defines a needleless first mating portion;a holder releasably couplable to the housing, the holder having a first end and a second end configured to receive an evacuated tube, the first end defining a second mating portion on an outer surface thereof which is in fluid communication with the evacuated tube,wherein the first mating portion sealingly engages the second mating portion to establish fluid communication therebetween; andwherein the first mating portion includes a male connector and the second mating portion includes a female connector. | 1. A fluid collection apparatus comprising: a housing configured for receipt of fluid and having a first surface which defines a needleless first mating portion;a holder releasably couplable to the housing, the holder having a first end and a second end configured to receive an evacuated tube, the first end defining a second mating portion on an outer surface thereof which is in fluid communication with the evacuated tube,wherein the first mating portion sealingly engages the second mating portion to establish fluid communication therebetween; andwherein the first mating portion includes a male connector and the second mating portion includes a female connector. 2. A fluid collection apparatus as recited in claim 1, wherein the first end of the holder includes a needle hub supporting a needle cannula in fluid communication with the second mating portion and extending away from the first mating portion. 3. A fluid collection apparatus as recited in claim 2, wherein the needle cannula engages the evacuated tube to establish fluid communication between the first mating portion and the evacuated tube. 4. A fluid collection apparatus as recited in claim 1, wherein the housing has a cylindrical body portion configured and dimensioned to support at least a portion of an umbilical cord. 5. A fluid collection apparatus as recited in claim 1, wherein the first surface of the housing has a funnel configuration. 6. A fluid collection apparatus comprising: a housing configured for receipt of fluid and having a first surface which defines a needleless first mating portion;a holder releasably couplable to the housing, the holder having a first end and a second end configured to receive an evacuated tube, the first end defining a second mating portion on an outer surface thereof which is in fluid communication with the evacuated tube,wherein the first mating portion sealingly engages the second mating portion to establish fluid communication therebetween; andwherein the first mating portion includes a female connector and the second mating portion includes a male connector. 7. A fluid collection apparatus comprising: a housing configured for receipt of fluid and having a first surface which defines a needleless first mating portion; anda holder releasably couplable to the housing, the holder having a first end and a second end configured to receive an evacuated tube, the first end defining a second mating portion on an outer surface thereof which is in fluid communication with the evacuated tube,wherein the first mating portion sealingly engages the second mating portion to establish fluid communication therebetween; andwherein the first mating portion engages the second mating portion in a slip interference fit. 8. An umbilical cord fluid collection apparatus comprising: a housing defining a cavity for receipt of at least a portion of an umbilical cord, the housing having a funnel surface which defines a male luer connector; anda holder having a first en... | Tyco Healthcare Group LP | A61B019/00 A61B005/00 B65D008/00
US06976967 | B2 | 2005-12-20 | Apparatus and method for sensing spatial displacement in a heart | 2002-02-19 | An electrical lead includes an elongate body having a proximal end, and a sensing unit capable of resolving a change in a spatial configuration of the electrical lead. A medical device includes a control unit, an elongate body having a proximal end coupled with the control unit, and a sensing unit capable of resolving a change in a spatial configuration of the electrical lead and relating it to the amount of blood ejected from the heart. A method includes receiving a signal transmitted from a lead disposed within a heart and determining a change in a dimension of the heart, due to the heart beating, based upon the signal. | 1. An electrical lead adapted to be routed into a heart through a venous system of a body, comprising: an elongate body adapted to be anchored to a portion of the heart to retain the lead in position; anda sensing unit capable of resolving a change in a spatial configuration of the electrical lead, wherein the sensing unit of the electrical lead is capable of resolving a change in a spatial configuration of the electrical lead that is functionally related to a change in one of the following: (1) the distance between a tricuspid valve in the heart and an apex of the heart and (2) the distance between a mitral valve of the heart and the apex of the heart. | 1. An electrical lead adapted to be routed into a heart through a venous system of a body, comprising: an elongate body adapted to be anchored to a portion of the heart to retain the lead in position; anda sensing unit capable of resolving a change in a spatial configuration of the electrical lead, wherein the sensing unit of the electrical lead is capable of resolving a change in a spatial configuration of the electrical lead that is functionally related to a change in one of the following: (1) the distance between a tricuspid valve in the heart and an apex of the heart and (2) the distance between a mitral valve of the heart and the apex of the heart. 2. An electrical lead, according to claim 1, wherein the sensing unit further comprises: a substrate;a strain gauge mounted to the substrate and being capable of sensing a flexural stress induced in the electrical lead; anda plurality of conductors extending from the strain gauge toward the proximal end of the electrical lead. 3. An electrical lead, according to claim 1, wherein the sensing unit further comprises: a substrate;a plurality of strain gauges mounted to the substrate in a pattern and being capable of sensing a flexural stress induced in the electrical lead; anda plurality of conductors extending from the plurality of strain gauges toward the proximal end of the electrical lead. 4. An electrical lead, according to claim 1, wherein the sensing unit further comprises: a substrate;a piezoresistive element mounted to the substrate and being capable of sensing a bending stress induced in the electrical lead; anda plurality of conductors extending from the piezoresistive element toward the proximal end of the electrical lead. 5. An electrical lead, according to claim 1, wherein the electrical lead is adapted to be anchored in a portion of the heart selected from a group consisting of the superior vena cava, the right apex, the coronary sinus, and the middle cardiac vein. 6. An electrical lead, according to claim 1 wherein the sensing unit further comprises: a first ultrasonic element disposed proximate and adjoined to a distal end portion of the elongate body;a second ultrasonic element adjoined to the elongate body and disposed intermediate the first ultrasonic element and the proximal end portion of the elongate body; anda plurality of conductors extending from each of the first ultrasonic element and the second ultrasonic element toward the proximal end of the electrical lead. 7. An electrical lead, according to claim 1, wherein the sensing unit further comprises: a first ultrasonic element disposed proximate and adjoined to a distal end portion of the elongate body and being capable of emitting ultrasonic energy;a second ultrasonic element adjoined to the elongate body and disposed intermediate the first ultrasonic element and the proximal end portion of the elongate body, wherein the second ultrasonic element is capable of receiving at least a portion of the ultrasonic energy emitt... | Medtronic, Inc. | A61B005/04 A61N001/365
US06850787 | B2 | 2005-02-01 | Signal component processor | 2002-06-26 | A signal processor generates basis functions and identifies at least one basis function component of a sensor signal in order to calculate a physiological measurement. The signal processor is advantageously applied to pulse oximetry so as to directly calculate oxygen saturation and so as to calculate a robust oxygen saturation measurement. In particular, a signal component transform may be calculated within a window around a derived pulse rate estimate. A signal component transform may also utilize sinusoidal basis functions, and an optimization of a signal component transform may occur at a particular frequency or a set of frequencies. A sinusoidal waveform or waveforms at that frequency or set of frequencies is generated to select associated red and infrared components of the sensor signal, and an oxygen saturation is calculated from a magnitude ratio of these components. | 1. A pulse oximeter system comprising: an input capable of accepting a physiological signal; and a component processor capable of communicating with said input and comprising: a first calculator capable of determining a basis function index from said signal; and a second calculator capable of generating a basis function waveform according to said index, deriving a component from said sensor signal and said waveform, and generating a physiological measurement responsive to said component, wherein second calculator derives said component using at least the inner product of said sensor signal and said waveform. | 1. A pulse oximeter system comprising: an input capable of accepting a physiological signal; and a component processor capable of communicating with said input and comprising: a first calculator capable of determining a basis function index from said signal; and a second calculator capable of generating a basis function waveform according to said index, deriving a component from said sensor signal and said waveform, and generating a physiological measurement responsive to said component, wherein second calculator derives said component using at least the inner product of said sensor signal and said waveform. 2. A pulse oximeter system comprising: an input capable of accepting a physiological signal; and a component processor capable of communicating with said input and comprising: a first calculator capable of determining a basis function index from said signal; and a second calculator capable of generating a basis function waveform according to said index, deriving a component from said sensor signal and said waveform, and generating a physiological measurement responsive to said component, wherein said first calculator determines a basis function frequency as said basis function index, and said second calculator generates a sinusoid at said frequency as said basis function waveform. 3. The pulse oximeter system according to claim 2, wherein said first calculator derives a pulse rate estimate from said signal selects said frequency from a window including said pulse rate estimate. 4. The pulse oximeter system according to claim 2, wherein said second calculator generates an oxygen saturation value responsive to a magnitude of said component as said physiological measurement. 5. A signal processing method comprising the steps of: receiving a sensor signal; calculating an estimated pulse rate; determining an optimization of said sensor signal proximate said estimated pulse rate; defining a frequency corresponding to said optimization; and outputting a physiological measurement responsive to a component of said sensor signal at said frequency, and wherein said outputting step comprises the substeps of: inputting a red (RD) portion and an infrared (IR) portion of said sensor signal; deriving a RD component of said RD portion and an IR component of said IR portion corresponding to said frequency; and computing an oxygen saturation based upon a magnitude ratio of said RD component and said IR component; and wherein said deriving step comprises the substeps of: generating a sinusoidal waveform at said frequency; and selecting said RD component and said IR component utilizing said waveform. 6. The signal processing method according to claim 5, wherein said selecting step comprises the substep of calculating the inner product between said waveform and said RD portion and the inner product between said waveform and said IR portion. 7. The signal processing method according to claim 6, wherein said selecting step comprises the substeps of: can... | Masimo Laboratories, Inc. | A61B005/00
US06936012 | B2 | 2005-08-30 | Method and apparatus for identifying constituent signal components from a plurality of evoked physiological composite signals | 2002-06-18 | A novel application of independent component analysis (ICA) to data acquired by a single sensor. The technique exploits the unique relationship between multiple physiologic (source) and electronic (artifact) components in surface recorded sensory nerve action potential (SNAP) waveforms that are evoked by different activating magnitudes. A forward model of the SNAP is developed and used to test the approach on a simplified data simulation. The method is applied to experimental data and shown to be effective at separating artifact and source components and reconstructing artifact-free traces. A method of automated reconstruction for use within an expert system is also disclosed. | 1. A method for identifying constituent signal components from a plurality of evoked physiological composite signals, said method comprising: activating a subject's anatomy so as to generate the plurality of composite signals, wherein at least two of said composite signals comprise distinct mixtures of the same constituent signal components; recording, at a single anatomical site, the composite signals; and performing blind source separation on the composite signals so as to identify the constituent signal components. | 1. A method for identifying constituent signal components from a plurality of evoked physiological composite signals, said method comprising: activating a subject's anatomy so as to generate the plurality of composite signals, wherein at least two of said composite signals comprise distinct mixtures of the same constituent signal components; recording, at a single anatomical site, the composite signals; and performing blind source separation on the composite signals so as to identify the constituent signal components. 2. A method according to claim 1 wherein said blind source separation is performed so as to identify the amount that each constituent signal component contributes to the composite signals. 3. A method according to claim 1 wherein said blind source separation is performed using independent component analysis. 4. A method according to claim 1 wherein at least one of the constituent signal components is an evoked neuromuscular response. 5. A method according to claim 4 wherein the evoked neuromuscular response is a compound muscle action potential (CMAP). 6. A method according to claim 4 wherein the evoked neuromuscular response is a sensory nerve action potential (SNAP). 7. A method according to claim 1 wherein at lease one of the constituent signal components is a directly recorded artifact of activating a subject's anatomy. 8. A method according to claim 1 wherein activating a subject's anatomy involves activating neuromuscular tissue. 9. A method according to claim 8 wherein the neuromuscular tissue comprises a peripheral nerve. 10. A method according to claim 1 wherein a subject's anatomy is activated with different levels of the same type of stimuli. 11. A method according to claim 10 wherein the stimuli comprises electrical stimuli. 12. A method according to claim 10 wherein the stimuli comprises mechanical stimuli. 13. A method according to claim 10 wherein the stimuli comprises magnetic stimuli. 14. A method according to claim 1 wherein the method comprises the additional step of removing a constituent signal component from a composite signal so as to yield a calculated composite signal. 15. A method according to claim 14 wherein the removed constituent signal component comprises a directly recorded artifact of activating the subject's anatomy. 16. A method according to claim 1 wherein the method comprises the additional step of removing a plurality of constituent signal components from a composite signal so as to yield a calculated composite signal. 17. A method according to claim 1 wherein the method comprises the additional step of constructing a calculated composite signal from at least one of the identified constituent signal components. 18. A method according to claim 17 wherein said calculated composite signal reduces the contribution of directly recorded artifacts of activating the subject's anatomy vis-à-vis the original composite signal. 19. A method according to claim 1 wherein the method comprise... | NeuroMetrix, Inc. | A61B005/05
US06854334 | B2 | 2005-02-15 | Negative inspiratory force manometer apparatus | 2003-01-28 | The apparatus is connected by tubing to a patient. The apparatus has a manometer on one side in communication with an entrance port conveying ambient air to the patient. A spring loaded cap is depressed to block air inspired by the patient and cause such air to move a diaphragm in the manometer so that the inspired negative air pressure of the patient can be recorded. As soon as the cap is released a spring opens the entrance port to continue normal flow of ambient air to the patient. | 1. An apparatus for measuring and recording maximum patient inspiratory air pressure, the apparatus comprising: a housing containing a patient port for connection to the patient and an entrance port open to ambient air; a manometer mounted within the housing; a means mounted over entry port for blocking the entrance port and causing the patient inspiratory air pressure to create a vacuum in the manometer; a channel within the housing connecting the patient port and entrance port for permitting flow of patient inspiratory air when the entrance port is not blocked; and a means for recording the patient inspiratory air pressure and viewing the said recording from outside the housing. | 1. An apparatus for measuring and recording maximum patient inspiratory air pressure, the apparatus comprising: a housing containing a patient port for connection to the patient and an entrance port open to ambient air; a manometer mounted within the housing; a means mounted over entry port for blocking the entrance port and causing the patient inspiratory air pressure to create a vacuum in the manometer; a channel within the housing connecting the patient port and entrance port for permitting flow of patient inspiratory air when the entrance port is not blocked; and a means for recording the patient inspiratory air pressure and viewing the said recording from outside the housing. 2. The apparatus according to claim 1 wherein the housing has a top, intermediate and bottom portion joined together, the top portion containing the entrance port and the bottom portion containing the patient port. 3. The apparatus according to claim 1 wherein the means for blocking the entrance port is a depressable cap mounted over the entrance port with a spacer bar descending downwardly from the cap through the entrance port into the channel to depress a spring, the downward depression of the cap sealing off the entrance port. 4. The apparatus according to claim 1 wherein the means for recording the patient inspiratory air pressure is a primary pointer rotating on a shaft below a clear cap over the manometer, the shaft engaged by a helix structure on an outer surface of the shaft to a notch in a stem mounted in a central portion of a diaphragm, the diaphragm converting negative air pressure from the patient on a surface of the diaphragm to rotary motion on the shaft when the entrance port is blocked. 5. The apparatus according to claim 2 wherein the top portion of the housing contains a support for a clear cover containing indicia, the support having a central aperture for passage of a rotating shaft, the shaft having a primary indicator pointer at a top portion, the intermediate portion having a right side opening promoting the flow of the patient inspiratory air and a left side opening supporting a diaphragm on a top annular edge and the bottom portion containing an interior side wall and a depression distal from the input port containing a spring, the spring depressed by a stem attached to a middle portion of the diaphragm. 6. The apparatus according to claim 5 wherein the intermediate housing portion has a restricted orifice in a side wall for passage of air. 7. An apparatus having a patient port connected to a patient, the apparatus comprising: a housing containing the patient port and an entrance port for flow of inspired air to the patient, and a manometer mounted within the housing for measuring and recording maximum patient inspiratory air pressure; a cap mounted over the entrance port so that, it can block entrance of ambient air to the patient when depressed and cause the inspired air to move a diaphragm in the manometer, a first and a second side ... | Mercury Enterprises, Inc. | A61B005/00
US07169109 | B2 | 2007-01-30 | Methods and apparatus for controlling heart assist devices | 2001-10-01 | An apparatus for a heart assist device, comprising a processing unit for computing the blood flow rate from the arterial pressure curve and for predicting at every heartbeat the closing time of the heart valve from the curve of the blood flow rate. The processing unit is adapted to deliver a signal for controlling a heart assist device at a point in time, a period ahead in time of the closing time of the heart valve, wherein the mechanical properties of the said heart assist device are taken into account in determining the period. The apparatus adapts itself to changes in a patient's heart frequency and aortic pressure. | 1. An apparatus for controlling a heart assist device, comprising a processing unit for computing the blood flow rate (q(t)) from the arterial pressure curve and for predicting at every heartbeat the closing time of the heart valve from the curve of the blood flow rate, wherein the processing unit is adapted to deliver a signal for controlling a heart assist device at a point in time, a period ahead in time of the closing time of the heart valve, wherein the mechanical properties of said heart assist device are taken into account in determining said period, wherein the processing unit is adapted to determine said period by determining the maximum of the blood flow rate curve, taking a percentage of said maximum as a threshold value, and comparing the computed blood flow rate with said threshold value to determine said point in time, and wherein said threshold value is greater than q(t)=0. | 1. An apparatus for controlling a heart assist device, comprising a processing unit for computing the blood flow rate (q(t)) from the arterial pressure curve and for predicting at every heartbeat the closing time of the heart valve from the curve of the blood flow rate, wherein the processing unit is adapted to deliver a signal for controlling a heart assist device at a point in time, a period ahead in time of the closing time of the heart valve, wherein the mechanical properties of said heart assist device are taken into account in determining said period, wherein the processing unit is adapted to determine said period by determining the maximum of the blood flow rate curve, taking a percentage of said maximum as a threshold value, and comparing the computed blood flow rate with said threshold value to determine said point in time, and wherein said threshold value is greater than q(t)=0. 2. The apparatus according to claim 1, wherein said heart assist device is an intra-aortic balloon pump having a balloon, wherein said signal initiates the inflation of the balloon. 3. The apparatus according to claim 1, characterized in that the processing unit is arranged for automatically controlling an intra-aortic balloon pump, wherein the deflation of the balloon is initiated on the basis of a derivative from the electrocardiogram with every heartbeat. 4. The apparatus according to claim 1, wherein computation of the curve of the blood flow rate from the arterial blood pressure is based on a Windkessel model that comprises a characteristic input resistance, an arterial compliance, and a peripheral resistance. 5. A method for controlling a heart assist device, comprising computing the blood flow rate (q(t)) from the arterial pressure curve and predicting at every heartbeat the closing time of the heart valve from the curve of the blood flow rate, wherein a signal for controlling a heart assist device is delivered at a point in time, a period ahead in time of the closing time of the heart valve, wherein the mechanical properties of said heart assist device are taken into account in determining said period, wherein said period is determined by determining the maximum of the blood flow rate curve, taking a percentage of said maximum as a threshold value, and comparing the computed blood flow rate with said threshold value to determine said point in time, and wherein said threshold value is greater than q(t)=0. 6. The method according to claim 5, wherein an intra-aortic balloon pump is controlled automatically, wherein the deflation of the balloon is initiated on the basis of a derivative from the electrocardiogram with every heartbeat. 7. The method according to claim 5, wherein computation of the curve of the blood flow rate from the arterial blood pressure is based on a Windkessel model that comprises a characteristic input resistance, an arterial compliance, and a peripheral resistance. 8. The method according to claim 5, wherein said heart assist de... | Arrow International, Inc. | A61B005/02 A61M001/10 A61N001/362
US07039462 | B2 | 2006-05-02 | Method and apparatus for detecting oscillations in cardiac rhythm | 2002-06-14 | A cardiac rhythm management device is configured to detect mechanical oscillations in cardiac rhythm. Upon detection of such oscillations, the device may adjust its operating behavior to compensate for the deleterious effects of the condition. | 1. An implantable cardiac rhythm management device, comprising: an exciter for generating an excitation current between excitation electrodes;voltage sense electrodes for generating an impedance signal corresponding to cardiac volume when the excitation current is applied; andcircuitry for processing the impedance signal to detect oscillations in cardiac volume during successive heart beats and to detect volumetric alternans when the measured oscillations exceed a specified magnitude for a specified number of beats. | 1. An implantable cardiac rhythm management device, comprising: an exciter for generating an excitation current between excitation electrodes;voltage sense electrodes for generating an impedance signal corresponding to cardiac volume when the excitation current is applied; andcircuitry for processing the impedance signal to detect oscillations in cardiac volume during successive heart beats and to detect volumetric alternans when the measured oscillations exceed a specified magnitude for a specified number of beats. 2. The device of claim 1 further comprising a device controller and wherein the controller is programmed to adjust the operating behavior of the device when volumetric alternans is detected. 3. The device of claim 2 further comprising: a first atrial pacing channel;a second atrial pacing channel; and,wherein the controller is programmed to deliver pacing therapy in accordance with a bradycardia mode, and further wherein the controller is programmed to initiate atrial resynchronization pacing by delivering paces through both atrial channels when volumetric alternans is detected. 4. The device of claim 2 further comprising: a sensing/pacing channel; andwherein the controller is programmed to deliver pacing therapy in accordance with a bradycardia mode, and further wherein the controller is programmed to decrease a programmed pacing rate when volumetric alternans is detected. 5. The device of claim 2 further comprising: a sensing/pacing channel; andwherein the controller is programmed to deliver anti-tachycardia pacing therapy when a measured heart rate exceeds a specified tachycardia threshold rate, and further wherein the controller is programmed to decrease the tachycardia threshold rate when volumetric alternans is detected. 6. The device of claim 2 further comprising: a first ventricular pacing channel;a second ventricular pacing channel; andwherein the controller is programmed to deliver pacing therapy in accordance with a bradycardia mode, and further wherein the controller is programmed to initiate ventricular resynchronization pacing by delivering paces through both ventricular pacing channels when volumetric alternans is detected. 7. The device of claim 2 further comprising: a first ventricular pacing channel;a second ventricular pacing channel; andwherein the controller is programmed to deliver ventricular resynchronization pacing by delivering paces through both ventricular channels, and further programmed to adjust a resynchronization operating parameter when volumetric alternans is detected. 8. The device of claim 7 wherein the resynchronization operating parameter adjusted by the controller is an atrio-ventricular interval. 9. The device of claim 7 wherein the resynchronization operating parameter adjusted by the controller is a biventricular delay interval. 10. The device of claim 2 further comprising: a ventricular pacing channel; andwherein the controller is programmed to initiate ventricular resynchronization ... | Cardiac Pacemakers, Inc. | A61N001/39 A61B005/05
US07226415 | B2 | 2007-06-05 | Microwave hemorrhagic stroke detector | 2002-09-23 | The microwave hemorrhagic stroke detector includes a low power pulsed microwave transmitter with a broad-band antenna for producing a directional beam of microwaves, an index of refraction matching cap placed over the patients head, and an array of broad-band microwave receivers with collection antennae. The system of microwave transmitter and receivers are scanned around, and can also be positioned up and down the axis of the patients head. The microwave hemorrhagic stroke detector is a completely non-invasive device designed to detect and localize blood pooling and clots or to measure blood flow within the head or body. The device is based on low power pulsed microwave technology combined with specialized antennas and tomographic methods. The system can be used for rapid, non-invasive detection of blood pooling such as occurs with hemorrhagic stoke in human or animal patients as well as for the detection of hemorrhage within a patient's body. | 1. An apparatus for detecting blood flow, comprising: means for producing a microwave signal, wherein said microwave signal is transmitted through a medium selected from a group consisting of a portion of a human body and a portion of an animal body to produce a transmitted signal;means for detecting said transmitted signal to produce a detected signal; andmeans for analyzing said microwave signal and said detected signal to determine whether blood is flowing within said medium, wherein said means for detecting said transmitted signal comprise at least one microwave receiver antenna, wherein said at least one microwave receiver antenna comprises a parabolic reflector antenna, wherein said at least one microwave receiver antenna comprises an array of at least 4 receiver antennas mounted with a gap of 3 to 5 mm between them, wherein said gap comprises microwave absorbing material to help reduce any residual crosstalk between antennas. | 1. An apparatus for detecting blood flow, comprising: means for producing a microwave signal, wherein said microwave signal is transmitted through a medium selected from a group consisting of a portion of a human body and a portion of an animal body to produce a transmitted signal;means for detecting said transmitted signal to produce a detected signal; andmeans for analyzing said microwave signal and said detected signal to determine whether blood is flowing within said medium, wherein said means for detecting said transmitted signal comprise at least one microwave receiver antenna, wherein said at least one microwave receiver antenna comprises a parabolic reflector antenna, wherein said at least one microwave receiver antenna comprises an array of at least 4 receiver antennas mounted with a gap of 3 to 5 mm between them, wherein said gap comprises microwave absorbing material to help reduce any residual crosstalk between antennas. 2. An apparatus for detecting blood flow, comprising: means for producing a microwave signal, wherein said microwave signal is transmitted through a medium selected from a group consisting of a portion of a human body and a portion of an animal body to produce a transmitted signal;means for detecting said transmitted signal to produce a detected signal; andmeans for analyzing said microwave signal and said detected signal to determine whether blood is flowing within said medium, wherein said means for detecting said transmitted signal comprise at least one microwave receiver antenna, wherein said at least one microwave receiver antenna comprises a parabolic reflector antenna, wherein said housing comprises microwave absorber material, wherein said microwave absorber material provides shielding, prevents entry of signals from outside objects, reduces residual crosstalk within said means for producing a microwave signal and reduces residual crosstalk within said means for detecting said transmitted signal, wherein said housing further comprises an element insert, wherein said element insert comprises a bag designed to fit within said housing and over a patient's head. 3. The apparatus for detecting blood flow of claim 2, wherein said bag is filled with dielectric material. 4. The apparatus for detecting blood flow of claim 3, wherein said bag comprises a material which can stretch and expand to fill said housing and conform to said patient's head. 5. The apparatus for detecting blood flow of claim 4, wherein said element insert further comprises an inlet quick disconnect to enable said element insert to be connected to a dielectric liquid fill tube through which dielectric filling liquid may be input to said element insert. 6. The apparatus for detecting blood flow of claim 5, wherein said element insert comprises material having about the same refractive index, in the microwave band being used, as said dielectric liquid. 7. The apparatus for detecting blood flow of claim 6, wherein said element insert comprises s... | The Regents of the University of California | A61B005/00 A61B005/05 A61B005/02 A61B005/08
US07233823 | B2 | 2007-06-19 | Apparatus for measuring body composition | 2002-10-11 | The invention concerns an apparatus for measuring a subject's body composition comprising a first electronic module (6) for measuring a bioelectric impedance and at least an electric power source (10) for generating a variable electric signal (11) which passes through the body or part of the body of the subject when the latter is connected to the measuring apparatus (2). The invention is characterized in that said power source is a voltage generator (10) designed to supply a rectangular voltage signal (11). | 1. Apparatus for measuring the body composition of an individual having a first electronic module (6) for measuring a bioelectric impedance and at least one electrical energy source intended to generate a variable electrical signal that passes through the body or a part of the body of the individual when the body or the part of the body of the individual is connected to said at least one electrical energy source, characterized in that said at least one energy source is a voltage generator (10) adapted to apply a voltage signal (11) of rectangular waveform across the body or the part of the body of the individual, and said module comprises means for obtaining voltage measurements across the body or the part of the body of the individual based on the voltage signal of rectangular waveform and for determining, on the basis of the voltage measurements, at least one of: the extracellular resistance Re of the body or of the part of the body of the individual; the intracellular resistance Ri of the body or of the part of the body of the individual; and a value that is a combination of the extracellular resistance and the intracellular resistance of the body or of the part of the body of the individual. | 1. Apparatus for measuring the body composition of an individual having a first electronic module (6) for measuring a bioelectric impedance and at least one electrical energy source intended to generate a variable electrical signal that passes through the body or a part of the body of the individual when the body or the part of the body of the individual is connected to said at least one electrical energy source, characterized in that said at least one energy source is a voltage generator (10) adapted to apply a voltage signal (11) of rectangular waveform across the body or the part of the body of the individual, and said module comprises means for obtaining voltage measurements across the body or the part of the body of the individual based on the voltage signal of rectangular waveform and for determining, on the basis of the voltage measurements, at least one of: the extracellular resistance Re of the body or of the part of the body of the individual; the intracellular resistance Ri of the body or of the part of the body of the individual; and a value that is a combination of the extracellular resistance and the intracellular resistance of the body or of the part of the body of the individual. 2. Apparatus according to claim 1, characterized in that said apparatus comprises means permitting voltage measurement to be effectuated at at least two different predetermined moments in order to directly determine the equivalent resistance Req of the body or of the part of the body of the individual at two of the moments and to deduce therefrom the extracellular resistance Re and the intracellular resistance Ri of the body or of the part of the body. 3. Apparatus according to claim 1, characterized in that said apparatus comprises means permitting a first voltage measurement to be effectuated, at the end of a plateau (32) of said voltage signal (11) of rectangular waveform, to determine the equivalent resistance Req equal to the extracellular resistance Re of the body or of the part of the body, as well as a second voltage measurement, at the end of a rising edge (31) of said signal (11) to determine the equivalent resistance Req equal to Re in parallel with Riand to determine, on the basis of the first voltage measurement, the intracellular resistance Ri of the body or of the part of the body of the individual. 4. Apparatus according to claim 1, characterized in that said voltage generator (10) is intended to furnish a voltage signal (11) that presents a wave form whose period, amplitude and duty cycle are adjustable. 5. Apparatus according to claim 1, characterized in that said voltage generator (10) is adapted to furnish a voltage signal (11) that is a low voltage signal with an amplitude comprised between 3 and 6V. 6. Apparatus according to claim 1, characterized in that said apparatus has a second electronic module (7) adapted to measure the weight of the individual. 7. Apparatus according to claim 6, characterized in that the first electroni... | SEB S.A. | A61B005/05
US07142922 | B2 | 2006-11-28 | Method and apparatus for predicting acute response to cardiac resynchronization therapy at a given stimulation site | 2002-12-20 | Response to cardiac resynchronization therapy is predicted for a given stimulation site so that an atrioventricular delay of an implantable device administering cardiac resynchronization therapy may be set to a proper amount. The first deflection of ventricular depolarization is measured, such as through a surface electrocardiogram or through an intracardiac electrogram measured by a lead positioned in the heart at the stimulation site. The maximum deflection of the ventricular depolarization is then measured by the lead positioned at the stimulation site. The interval of time between the first deflection and the maximum deflection of the ventricular depolarization is compared to a threshold to determine whether the stimulation site is a responder site. If the interval is larger than the threshold, then the site is a responder and the atrioventricular delay of the implantable device may be set to less than the intrinsic atrioventricular delay of the patient. Otherwise, the atrioventricular may be set to approximately equal the intrinsic atrioventricular delay. | 1. A method for predicting acute response to cardiac resynchronization therapy at a stimulation site of a patient, comprising the steps of: inserting a lead to the heart of the patient such that an electrode of the lead is positioned at the stimulation site;detecting a first deflection of an intrinsic ventricular depolarization;detecting a maximum deflection of the intrinsic ventricular depolarization at the electrode; andcomparing an interval of time between the first deflection and the maximum deflection to a threshold to determine whether the stimulation site is a responder site. | 1. A method for predicting acute response to cardiac resynchronization therapy at a stimulation site of a patient, comprising the steps of: inserting a lead to the heart of the patient such that an electrode of the lead is positioned at the stimulation site;detecting a first deflection of an intrinsic ventricular depolarization;detecting a maximum deflection of the intrinsic ventricular depolarization at the electrode; andcomparing an interval of time between the first deflection and the maximum deflection to a threshold to determine whether the stimulation site is a responder site. 2. The method of claim 1, wherein the lead is electrically connected to an implantable device that delivers cardiac resynchronization therapy, and wherein the method further comprises the step of setting the atrioventricular delay of the implantable device to substantially less than an intrinsic atrioventricular delay of the patient when the stimulation site is a responder site. 3. The method of claim 2, wherein the method further comprises the step of setting the atrioventricular delay of the implantable device to approximately the intrinsic atrioventricular delay of the patient when the stimulation site is a non-responder site. 4. The method of claim 1, wherein detecting the first deflection comprises detecting the onset of the first deflection detected at the electrode of the lead. 5. The method of claim 1, wherein the detecting a first deflection of an intrinsic ventricular depolarization comprises measuring an onset of the first deflection at the electrode. 6. The method of claim 5, wherein the threshold is 100 milliseconds. 7. The method of claim 1, wherein the step of detecting a first deflection of an intrinsic ventricular depolarization comprises measuring an onset of the first deflection detected by a surface electrocardiogram. 8. The method of claim 7, wherein the threshold is 80 milliseconds. 9. A system for predicting acute response to cardiac resynchronization therapy at a stimulation site of a patient, comprising: a lead having an electrode placed at the stimulation site that detects an intrinsic ventricular depolarization;a processing device that finds an interval of time between a first deflection and a maximum deflection of the intrinsic ventricular depolarization as detected by the electrode and compares the interval to a threshold to determine whether the stimulation site is a responder; and,wherein the processing device is further configured to communicate with an implantable device to adjust the atrioventricular delay of the implantable device in accordance with whether the interval is less than or greater than the threshold. 10. The system of claim 9, wherein the processing device is further configured to communicate with an implantable device to set the atrioventricular delay of the implantable device to substantially less than an intrinsic atrioventricular delay of the patient when the interval is greater than the threshold. 11. The ... | Cardiac Pacemakers, Inc. | A61B005/04
US07076287 | B2 | 2006-07-11 | System and method for detecting new left bundle branch block for accelerating treatment of acute myocardial infarction | 2000-12-29 | A system and a method for automatically detecting a new left bundle branch block (LBBB) in an ECG series and then issuing an alert for the purpose of accelerating treatment for acute myocardial infarction. A serial comparison program is used to compare each current ECG with a previous ECG for the same patient. Diagnostic statements, measurements and waveforms are compared and based on the comparisons, the system is able to automatically determine that the patient has a new LBBB. The system then automatically determines whether the patient belongs to the category of patients having a high probability of acute myocardial infarction. If acute myocardial infarction is suspected, the system generates a diagnostic statement stating that the new left bundle branch block may be due to acute myocardial infarction. The current ECG exhibiting a new LBBB is then identified with a special tag and sent to a central database server. The special tag enables the central database server to perform special routing of that record, e.g., via facsimile or digital pager, to alert on-call medical personnel to the need for immediate treatment. | 1. A method for monitoring a patient for the presence of a new bundle branch block, comprising the steps of: acquiring current and previous electrocardiogram records at different times for a particular patient;determining that a new bundle branch block is present in said particular patient based on an automated comparison of data in said current and previous electrocardiogram records;automatically detecting that said particular patient belongs to a category of patients comprising patients identified to have a high probability of acute myocardial infarction; andautomatically generating a diagnostic statement stating that the new bundle branch block may be due to acute myocardial infarction,wherein said automated comparison comprises the following steps:determining that said current electrocardiogram record comprises a diagnostic statement indicating the presence of a bundle branch block and determining that said previous electrocardiogram record does not comprise a diagnostic statement indicating the presence of a bundle branch block;computing a percentage change in a measurement relating to a bundle branch block criterion in said current and previous electrocardiogram records; anddetermining that said percentage change exceeds a first predetermined threshold. | 1. A method for monitoring a patient for the presence of a new bundle branch block, comprising the steps of: acquiring current and previous electrocardiogram records at different times for a particular patient;determining that a new bundle branch block is present in said particular patient based on an automated comparison of data in said current and previous electrocardiogram records;automatically detecting that said particular patient belongs to a category of patients comprising patients identified to have a high probability of acute myocardial infarction; andautomatically generating a diagnostic statement stating that the new bundle branch block may be due to acute myocardial infarction,wherein said automated comparison comprises the following steps:determining that said current electrocardiogram record comprises a diagnostic statement indicating the presence of a bundle branch block and determining that said previous electrocardiogram record does not comprise a diagnostic statement indicating the presence of a bundle branch block;computing a percentage change in a measurement relating to a bundle branch block criterion in said current and previous electrocardiogram records; anddetermining that said percentage change exceeds a first predetermined threshold. 2. The method as recited in claim 1, further comprising the step of automatically tagging said current electrocardiogram record with a special tag. 3. The method as recited in claim 2, wherein said tagging step is performed by a satellite system connected to a network, further comprising the step of sending said specially tagged current electrocardiogram record to a central database via said network. 4. The method as recited in claim 1, wherein said automated comparison further comprises the steps of: computing a correlation value for respective waveforms in said current and previous electrocardiogram records using cross correlation; anddetermining that said correlation value exceeds a second predetermined threshold. 5. A system for monitoring a patient for the presence of a new bundle branch block, comprising: means for acquiring current and previous electrocardiogram records at different times for a particular patient;means for determining whether a new bundle branch block is present in said particular patient based on a comparison of data in said current and previous electrocardiogram records;means for detecting whether said particular patient belongs to a category of patients comprising patients identified to have a high probability of acute myocardial infarction, said detecting means being operative if said determining means have determined that a new bundle branch block is present; andmeans for generating a diagnostic statement stating that the new bundle branchblock may be due to acute myocardial infarction, said generating means being operative if said detecting means have detected that said particular patient belongs to said category,wherein said determining means comprise a com... | GE Medical Systems Information Technologies, Inc. | A61B005/04
US07177682 | B2 | 2007-02-13 | Highly specific technique for discriminating atrial fibrillation from atrial flutter | 2002-08-13 | A cardiac rhythm management system including a highly sensitive technique for discriminating AF from AFL. An electrode disposed in or about a heart senses a cardiac signal. A controller through a sensing circuit receives the sensed cardiac signal from the electrode and processes the sensed cardiac signal to compute an average cycle length-to-cycle length variation based on serial interval relationships from the sensed cardiac signal, and then comparing the computed average cycle length-to-cycle length variation to one or more pre-determined threshold values to discriminate AF from AFL. As a result of substantial difference in cycle length-to-cycle length variation between AF and AFL (despite similar average cycle lengths), the sequence-based measure of cycle length-to-cycle length variability proves to be a highly specific and reliable discriminator of AF from AFL. | 1. A method including: sensing a cardiac signal; anddiscriminating a cardiac arrhythmia based on comparison of two or more pre-determined interval sequences of the sensed cardiac signal by:computing a central tendency cycle length-to-cycle length variation for a pre-determined interval sequence of the sensed cardiac signal;comparing the computed central tendency cycle length-to-cycle length variation with one or more pre-determined threshold values; anddetermining atrial fibrillation from atrial flutter based on the outcome of the comparison. | 1. A method including: sensing a cardiac signal; anddiscriminating a cardiac arrhythmia based on comparison of two or more pre-determined interval sequences of the sensed cardiac signal by:computing a central tendency cycle length-to-cycle length variation for a pre-determined interval sequence of the sensed cardiac signal;comparing the computed central tendency cycle length-to-cycle length variation with one or more pre-determined threshold values; anddetermining atrial fibrillation from atrial flutter based on the outcome of the comparison. 2. The method of claim 1, where discriminating further comprises: computing an average cycle length-to-cycle length variation for a pre-determined interval sequence;comparing the computed average cycle length-to-cycle length variation with one or more pre-determined threshold values; andissuing a command signal based on the outcome of the comparison. 3. The method of claim 2, where computing of average cycle length-to-cycle length variation further includes: computing cycle lengths (Δt1, Δt2 . . . ΔtN) from the sensed cardiac signal for a pre-determined interval sequence; andcomputing average cycle length-to-cycle length variation ((1/(N−1)) (Σ(¦Δt1−Δt2¦+¦Δt2−Δt3¦. . . +¦ΔtN−1−ΔtN¦))) from the computed cycle lengths for the pre-determined interval sequence. 4. The method of claim 2, further includes providing electrical energy to at least one electrode upon receiving the command signal. 5. The method of claim 2, further includes providing a cardiac therapy to a heart upon receiving the command signal. 6. The method of claim 2, where sensing a cardiac signal includes sensing atrial signals. 7. The method of claim 2, further including activating an implanted device to administer a drug therapy to the heart upon receiving the command signal. 8. A method including: sensing a cardiac signal;computing cycle lengths (Δt1, Δt2 . . . ΔtN) from the detected cardiac signal for a pre-determined interval sequence;computing average cycle length-to-cycle length variation from the computed cycle lengths;comparing one or more predetermined threshold values with the computed average cycle length-to-cycle length;introducing a time delay to provide therapy to a heart based on the outcome of the comparison; andproviding the therapy to a heart. 9. A method comprising: sensing a cardiac signal;computing cycle lengths from the sensed cardiac signal for the pre-determined number of sequential cardiac cycles;computing cycle length-to-cycle length variations from the computed cycle lengths; andcomputing a central tendency cycle length-to-cycle length variation from the computed cycle length-to-cycle length variationsdiscriminating between ventricular fibrillation and ventricular tachycardia based on the outcome of the comparison. 10. The method of claim 9, wherein computing the central tendency cycle length-to-cycle length variation comprises computing an average cycle length-to-cycle length variation. 11. The method of clai... | Cardiac Pacemakers, Inc. | A61N001/362 A61B005/046
US07207946 | B2 | 2007-04-24 | Automated provision of information related to air evacuation from a chest cavity | 2002-09-26 | A device according to an aspect of the invention includes an air escapement conduit having an inlet port providing air communication with a chest cavity and an outlet port providing air communication with a vacuum source, the conduit allowing an air flow from the inlet port to the outlet port in response to a pressure differential between the ports, and a detector responsive to the air flow that provides a signal related to the air flow. The device may include an indicator device operable to provide evacuation information in response to the signal. The air escapement conduit may include a bubble chamber or a valving mechanism. The detector may be a pressure differential sensor operable to detect a difference in air pressure between the inlet port and the outlet port, and generate a signal related to the difference. | 1. A method of providing a signal indicating information related to air evacuation from a chest cavity, the method comprising the steps of: inserting an air escapement conduit into an air evacuation pathway between the chest cavity and a vacuum source, the air escapement conduit allowing an air flow in response to a pressure differential between an inlet port and an outlet port;generating a signal related to the air flow; andindicating air evacuation information in response to the signal,wherein the air escapement conduit includes a bubble chamber having a fluid disposed between the inlet port and the outlet port, the bubble chamber arranged so that air flowing between the inlet port and the outlet port flows through the fluid and forms bubbles. | 1. A method of providing a signal indicating information related to air evacuation from a chest cavity, the method comprising the steps of: inserting an air escapement conduit into an air evacuation pathway between the chest cavity and a vacuum source, the air escapement conduit allowing an air flow in response to a pressure differential between an inlet port and an outlet port;generating a signal related to the air flow; andindicating air evacuation information in response to the signal,wherein the air escapement conduit includes a bubble chamber having a fluid disposed between the inlet port and the outlet port, the bubble chamber arranged so that air flowing between the inlet port and the outlet port flows through the fluid and forms bubbles. 2. The method of claim 1, wherein the air escapement conduit includes a mechanical valving mechanism. 3. The method of claim 1, wherein the step of indicating further includes a flash of light. 4. The method of claim 3, wherein the flash of light is provided by a light emitting diode. 5. The method of claim 1, wherein the step of indicating includes a sound. | Spiration, Inc. | A61B005/08
US06920347 | B2 | 2005-07-19 | Trajectory storage apparatus and method for surgical navigation systems | 2002-06-21 | Apparatus and methods are disclosed for use within an image-guided surgical navigation system for the storage and measurement of trajectories for surgical instruments. An icon representing the real-time trajectory of a tracked instrument is overlaid on one or more pre-acquired images of the patient. At the surgeon's command, the navigation system can store multiple trajectories of the instrument and create a static icon representing each saved trajectory for display. The surgeon may also measure a planar angle between any two trajectories. The angle is computed in the plane of the image, and therefore will be computed separately for each image displayed. Furthermore, the surgeon has the option of computing and displaying the three-dimensional distance between two points defined by any two trajectories. | 1. A method for storing a trajectory of at least one object during a surgical procedure, the method comprising: providing image data of a patient; tracking a three-dimensional trajectory of the object; displaying a first representation of the object's first trajectory superimposed onto the image data of the patient; receiving a command to store the object's first trajectory; storing the first representation of the object's first trajectory; and superimposing a second representation of the object's second trajectory onto the image data of the patient. | 1. A method for storing a trajectory of at least one object during a surgical procedure, the method comprising: providing image data of a patient; tracking a three-dimensional trajectory of the object; displaying a first representation of the object's first trajectory superimposed onto the image data of the patient; receiving a command to store the object's first trajectory; storing the first representation of the object's first trajectory; and superimposing a second representation of the object's second trajectory onto the image data of the patient. 2. The method as defined in claim 1 wherein tracking the three-dimensional trajectory of the object further includes tracking a three-dimensional trajectory of a surgical instrument. 3. The method as defined in claim 1 wherein tracking the three-dimensional trajectory of the object further includes tracking a three-dimensional trajectory of an implant. 4. The method as defined in claim 1 wherein providing the image data of the patient further includes generating the image data from a fluoroscopic x-ray imager. 5. The method as defined in claim 4 further comprising generating the image data from a C-arm fluoroscopic x-ray imager. 6. The method as defined in claim 1 wherein providing the image data of the patient further includes generating the image data from computer tomography imaging. 7. The method as defined in claim 1 wherein providing the image data of the patient further includes generating the image data from ultrasound imaging. 8. The method as defined in claim 1 wherein providing the image data of the patient further includes generating the image data from magnetic resonance imaging. 9. The method as defined in claim 1 wherein providing the image data of the patient further includes generating the image data from PET imaging. 10. The method as defined in claim 1 wherein providing the image data of the patient further includes providing two-dimensional image data of the patient. 11. The method as defined in claim 1 wherein providing the image data of the patient further includes providing three-dimensional image data of the patient. 12. The method as defined in claim 1 further comprising performing a medical procedure selected from a group comprising spinal implant placement, bone fragment alignment, bone fracture fixation, implant placement and bone removal. 13. The method as defined in claim 1 further comprising: selecting the first and second trajectory; projecting the first trajectory into a plane of the image data to form a first set of projected points; projecting the second trajectory into the plane of the image data to form a second set of projected points; calculating a first segment connecting the first set of projected points; calculating a second segment connecting the second set of projected points; computing an angle between the first and second segments in the plane of the image data; and displaying the computed angle. 14. The method as defined in claim 13 furthe... | Surgical Navigation Technologies, Inc. | A61B005/05
US06934572 | B2 | 2005-08-23 | Implantable enzyme-based monitoring system having improved longevity due to improved exterior surfaces | 2002-10-31 | An implantable enzyme-based monitoring system suitable for long term in vivo use to measure the concentration of prescribed substances such as glucose is provided. In one embodiment, the implantable enzyme-based monitoring system includes at least one sensor assembly, an outer membrane surrounding the sensor assembly and having a window therein, and a polymeric window cover affixed to the outer membrane and covering the window. Preferably, the outer membrane of the monitoring system is silicone and the window cover is a polymer of 2-hydroxyethyl methacrylate (HEMA), N,N,-dimethylaminoethyl methacrylate (DMAEMA) and methacrylic acid (MA). Also provided herein is an implantable enzyme-based monitoring system having at least one sensor assembly, an outer membrane surrounding the sensor assembly and a coating affixed to the exterior surface of the outer membrane, wherein the coating resists blood coagulation and protein binding to the exterior surface of the outer membrane. Preferably, the coating is polyethylene glycol (PEG) and heparin in an 80:20 molar ratio. Finally, provided herein is a method of coating the exterior surface of the outer membrane of an implantable enzyme-based monitoring system comprising the steps of forming hydroxyl groups on the silicone surface by plasma etching; reacting the silicone surface with amino functionalized silane, thereby forming amino groups on the silicone surface; simultaneously, covalently binding polyethylene glycol (PEG) and heparin to the amino groups; and ionically binding heparin to the monitoring system surface. | 1. An implantable monitoring system comprising: a plurality of sensor assemblies coupled in a daisy chain, each sensor assembly having a sensing area for sensing a parameter of an implant environment; and a membrane for housing the plurality of sensor assemblies; wherein the configuration of the monitoring system is non-linear. | 1. An implantable monitoring system comprising: a plurality of sensor assemblies coupled in a daisy chain, each sensor assembly having a sensing area for sensing a parameter of an implant environment; and a membrane for housing the plurality of sensor assemblies; wherein the configuration of the monitoring system is non-linear. 2. An implantable monitoring system according to claim 1, wherein at least three sensing areas are facing three distinct directions. 3. An implantable monitoring system according to claim 1, wherein at least two sensing areas are rotationally displaced approximately 180 degrees relative to a longitudinal dimension of the membrane. 4. An implantable monitoring system according to claim 1, wherein the membrane is curved along a longitudinal dimension of the membrane. 5. An implantable monitoring system according to claim 4, wherein each sensing area is located on a concave portion of the membrane. 6. An implantable monitoring system according to claim 5, wherein the membrane is spiral along the longitudinal dimension of the membrane. 7. An implantable monitoring system according to claim 5, wherein the membrane is acruate along the longitudinal dimension of the membrane. 8. An implantable monitoring system according to claim 5, wherein the membrane is S-shaped along the longitudinal dimension of the membrane. 9. An implantable monitoring system comprising: a plurality of sensor assemblies coupled in a daisy chain, each sensor assembly having a sensing area for sensing a parameter of an implant environment; a membrane for housing the plurality of sensor assemblies; wherein the membrane is curved along a longitudinal dimension; and wherein the sensor assemblies are disposed along the longitudinal dimension of the membrane. 10. An implantable monitoring system to claim 9, wherein each sensing area is located on a concave portion of the membrane. 11. An implantable monitoring system according to claim 10, wherein the membrane is spiral along the longitudinal dimension of the membrane. 12. An implantable monitoring system according to claim 10, wherein the membrane is arcuate along the longitudinal dimension of the membrane. 13. An implantable monitoring system according to claim 10, wherein the membrane is S-shaped along the longitudinal dimension of the membrane. | Alfred E. Mann Foundation for Scientific Research | A61B005/00
US06920345 | B2 | 2005-07-19 | Optical sensor including disposable and reusable elements | 2003-01-24 | A pulse oximetry sensor includes reusable and disposable elements. To assemble the sensor, the reusable element is snapped or slid within assembly mechanisms of the disposable element. In one embodiment, the reusable element includes protrusions while the disposable element includes tabs which receive the protrusions. In another embodiment, the reusable element includes neck portions and the disposable element includes straps which receive the neck portions. In any event, the assembly provides longitudinal movement between elements. | 1. A pulse oximetry sensor for outputting a signal indicative of one or more physiological parameters of tissue of a patient, the sensor comprising: a reusable element comprising an energy emission device adapted to transmit energy into tissue of a patient, an energy detection device adapted to detect attenuated energy from the tissue and output a signal indicative of one or more physiological parameters of the tissue, and a plurality of protrusions extending from the reusable element; and a disposable element comprising a tape adapted to removably attach to the tissue and a plurality of tabs each including a receiving slot, wherein each of the slots is adapted to receive one of the plurality of protrusions and adapted to permit the one protrusion to slide within a corresponding slot. | 1. A pulse oximetry sensor for outputting a signal indicative of one or more physiological parameters of tissue of a patient, the sensor comprising: a reusable element comprising an energy emission device adapted to transmit energy into tissue of a patient, an energy detection device adapted to detect attenuated energy from the tissue and output a signal indicative of one or more physiological parameters of the tissue, and a plurality of protrusions extending from the reusable element; and a disposable element comprising a tape adapted to removably attach to the tissue and a plurality of tabs each including a receiving slot, wherein each of the slots is adapted to receive one of the plurality of protrusions and adapted to permit the one protrusion to slide within a corresponding slot. 2. The sensor of claim 1, wherein at least one of the plurality of tabs includes a sloped surface that displaces the tab allowing receipt of the protrusion. 3. The sensor of claim 1, wherein the plurality of protrusions include protrusions that share a common substrate. 4. The sensor of claim 1, wherein the reusable element is encapsulated with a pliable material. 5. The sensor of claim 1, wherein the plurality of tabs include tabs that share a common substrate. 6. The sensor of claim 1, wherein the disposable element further comprises an information element adapted to provide information to a signal processor. 7. The sensor of claim 1, wherein some of the plurality of tabs each include an aperture, wherein each aperture is adapted to receive one of the plurality of protrusions and adapted to substantially prevent the one protrusion from sliding within a corresponding aperture. 8. A method of attaching a pulse oximetry sensor comprising a reusable element and a disposable element around a measurement site of a patient, the method comprising: inserting protrusions of a reusable element respectively into tabs of a disposable element; and attaching the disposable element proximate a measurement site, wherein the tabs permit slidable movements of at least a portion of the reusable element with respect to the disposable element. 9. The method of claim 8, wherein a plurality of the tabs each include a receiving slot respectively sized to receive a plurality of the protrusions, and wherein another plurality of the tabs each include an aperture. 10. A method of attaching a pulse oximetry sensor comprising a reusable element and a disposable element to a measurement site of a patient, the reusable element including a first neck section and a second neck section, each neck section being narrower than adjacent sections of the reusable element, the disposable element including a first strap and a second strap, wherein each of the straps includes a gap adapted to receive one of the neck sections, the method comprising: placing a first neck section of a reusable sensor element within a gap of a first strap of a disposable sensor element; placing a second neck section o... | Masimo Corporation | A61B005/00
US06856823 | B2 | 2005-02-15 | Spiral magnetic transmitter for position measurement system | 2002-06-18 | A transmitter consists of a plurality of magnetic transmitters, each of which is substantially planar and made of a spiral shaped conductor. The result is a transmitter having a substantially uniform cross-sectional current density along the radial direction of the spiral from the center to the periphery. Near the plane of a given spiral, magnetic vectors produced by such a conductor arrangement have improved angular characteristics as compared to prior art systems. This results in a larger region with useful vector crossing angles and operation of the system is enhanced as compared to prior art techniques. The transmitters produce magnetic fields which have a monotonically increasing intensity as one approaches the center of the transmitter spiral from any given direction. This feature simplifies and increases the accuracy of sensor position determination. If desired, the turns of the spiral transmitter may become closer together as one goes from the periphery to the center thereof to thereby concentrate the magnetic field more centrally of the transmitter. Transmitters having one or two spiral coils are also contemplated. | 1. In a position measurement system, the improvement comprising a transmitter including a first conductor arranged in a first planar spiral having a plurality of turns, a second conductor arranged in a second planar spiral having a plurality of turns, and sensing means attached to an object for sensing magnetic fields emanating from said spirals and therefrom calculating position of said object in a space, said first and second spirals being mounted on respective opposed parallel faces of a substrate, one spiral directly overlying another spiral, said system measuring position in a number of degrees of freedom at most equal to a number of transmitter axes times a number of sensor means axes. | 1. In a position measurement system, the improvement comprising a transmitter including a first conductor arranged in a first planar spiral having a plurality of turns, a second conductor arranged in a second planar spiral having a plurality of turns, and sensing means attached to an object for sensing magnetic fields emanating from said spirals and therefrom calculating position of said object in a space, said first and second spirals being mounted on respective opposed parallel faces of a substrate, one spiral directly overlying another spiral, said system measuring position in a number of degrees of freedom at most equal to a number of transmitter axes times a number of sensor means axes. 2. The system of claim 1, wherein each spiral has adjacent conductor rings that are substantially concentric. 3. The system of claim 1, wherein each spiral has adjacent conductor rings, spacing between adjacent conductor rings becoming smaller toward a center of each spiral. 4. The system of claim 1, wherein said spirals are electically connected through said substrate. 5. The system of claim 1, wherein each of said first and second spirals consists of a first spiral half located on one face of said substrate and a second spiral half located on another face of said substrate aligned with said first spiral half. 6. The system of claim 1, further including a third planar spiral, said spirals overlapping one another. 7. The system of claim 6, wherein each of said spirals consists of a first spiral half located on one face of a substrate and a second spiral half located on another face of said substrate aligned with said first spiral half. 8. The system of claim 7, wherein spiral halves on each face of said substrate are co-planar. 9. The system of claim 6, wherein each spiral has adjacent conductor rings that are substantially concentric. 10. The system of claim 6, wherein each spiral has adjacent conductor rings, spacing between adjacent conductor rings becoming smaller toward a center of each spiral. 11. In a position measurement system, the improvement comprising a transmitter including a first conductor arranged in a first planar spiral having a plurality of turns, a second conductor arranged in a second planar spiral having a plurality of turns, and a third conductor arranged in a third planar spiral having a plurality of turns, said spirals being mounted on a substrate having opposed parallel faces, and sensing means attached to an object for sensing magnetic fields emanating from said spirals and therefrom calculating position of said object in a space, said system measuring position in a number of degrees of freedom at most equal to a number of transmitter axes times a number of sensor means axes, said spirals overlapping one another and lying in a common plane. 12. The system of claim 11, wherein each spiral has adjacent conductor rings that are substantially concentric. 13. The system of claim 11, wherein each spiral has adjacent conductor ... | Ascension Technology Corporation | A61B005/05
US07089048 | B2 | 2006-08-08 | Biological signal transmission apparatus | 2002-04-02 | An electrode 4 for detecting a biological signal and a loop antenna 3 are integrally mounted on a support 2 placed on the surface of a living body and a transmitter 5 is placed on the support 2. A biological signal detected on the electrode 4 is input through a connector 11 to electric circuitry 10 of the transmitter 5 and an electric signal processed by the electric circuitry 10 is output through connectors 12 and 13 to both ends of the loop antenna 3 from which the biological signal is emitted to a receiver. At this time, the opening face of the loop antenna 3 is in a direction almost perpendicular to the surface of a living body for improving sensitivity. | 1. A method of detecting an ECG wave of a subject, comprising: providing only one electrode group consisting of two electrodes:positioning the two electrodes within 2.5 cm of an area defined between a xiphoid process and a navel perpendicular to a midsternal line of the subject; anddetecting the ECG wave from signals from the two electrodes according to a standard limb lead method. | 1. A method of detecting an ECG wave of a subject, comprising: providing only one electrode group consisting of two electrodes:positioning the two electrodes within 2.5 cm of an area defined between a xiphoid process and a navel perpendicular to a midsternal line of the subject; anddetecting the ECG wave from signals from the two electrodes according to a standard limb lead method. 2. A method according to claim 1, wherein said electrodes are spaced from each other within a range of 1 to 7.5 cm. 3. A method according to claim 1, wherein said electrodes are spaced from each other in a range of 2 to 7.5 cm. 4. A method of detecting an ECG wave of a subject, comprising: providing an ECG signal transmission apparatus which comprises:only one electrode group consisting of two electrodes each of which is adapted to detect an ECG signal of the subject;a support, supporting the electrodes, and adapted to be placed on the subject; anda transmitter, having an electric circuit for processing the ECG signal detected by the electrodes;positioning the two electrodes within 2.5 cm of an area defined between a xiphoid process and a navel perpendicular to a midsternal line of the subject; anddetecting the ECG wave from signals from the two electrodes according to a standard limb lead method. 5. A method according to claim 4, wherein said electrodes are spaced from each other within a range of 1 to 7.5 cm. 6. A method according to claim 4, wherein said electrodes are spaced from each other within a range of 2 to 7.5 cm. | Nihon Kohden Corporation | A61B005/04
US07117037 | B2 | 2006-10-03 | Event marker alignment by inclusion of event marker transmission latency in the real-time data stream | 2003-01-27 | A cardiac rhythm management system transmits event markers from an implantable device to a remote user interface such as an external programmer. The event markers include timestamp information. A printed/displayed event marker representation is aligned to a printed/displayed heart signal representation based at least in part on the timestamp information. Proper alignment of the printed/displayed heart signal representation to the printed/displayed event marker, and proper alignment of the printed/displayed event markers with respect to each other, provides more accurate diagnostic information to the caregiver. This assists the caregiver in using the cardiac rhythm management system to provide a patient with proper cardiac rhythm management therapy. | 1. An implantable cardiac rhythm management (CRM) device comprising: at least one sensing circuit, configured for receiving at least one heart signal from a heart;a processor, configured to process information obtained from the at least one heart signal to generate a substantially real-time event marker and a substantially real-time electrogram; andat least one communication circuit, configured to communicate the substantially real-time event marker and the substantially real-time electrogram from the implantable CRM device, in which the event marker includes a timestamp relating at least in part to an approximate time between an event associated with the event marker and another event associated with another event marker for aligning an indication of the event marker to a substantially real-time representation of a heart signal on one of a display or a recording that is remote from the implantable CRM device. | 1. An implantable cardiac rhythm management (CRM) device comprising: at least one sensing circuit, configured for receiving at least one heart signal from a heart;a processor, configured to process information obtained from the at least one heart signal to generate a substantially real-time event marker and a substantially real-time electrogram; andat least one communication circuit, configured to communicate the substantially real-time event marker and the substantially real-time electrogram from the implantable CRM device, in which the event marker includes a timestamp relating at least in part to an approximate time between an event associated with the event marker and another event associated with another event marker for aligning an indication of the event marker to a substantially real-time representation of a heart signal on one of a display or a recording that is remote from the implantable CRM device. 2. The implantable CRM device of claim 1, further including at least one electrode adapted for coupling the at least one head signal from the heart to the at least one sensing circuit. 3. The implantable CRM device of claim 1, in which the communication circuit transmits data from the CRM device in frames, in which each data frame includes a plurality of data slots including: a first data slot associated with a first event marker associated with a first electrode; anda second data slot associated with a second event marker associated with a second electrode different from the first electrode. 4. The implantable CRM device of claim 3, in which at least one event marker requires more than one data frame for transmission of the event marker from the CRM device to a user interface (UI) remote from the implantable CRM device. 5. The implantable CRM device of claim 4, in which a first type of event marker includes an express event marker including at least two bytes of data, and a second type of event marker includes a full event marker including more bytes of data than the express event marker. 6. The implantable CRM device of claim 5, in which the express event marker includes a Hamming code error correction field. 7. The implantable CRM device of claim 1, in which the timestamp relates at least in part to an approximate time between an event associated with the event marker and another event associated with a previous event marker associated with the same electrode. 8. The implantable CRM device of claim 1, in which the timestamp relates at least in part to an approximate time between an event associated with the event marker and another event associated with a previous event marker associated with the same heart chamber. 9. The implantable CRM device of claim 1, in which the timestamp relates at least in part to an approximate time between an event associated with the event marker and the transmission completion time associated with a previous event marker associated with the same heart chamber. 10. A remote user interface (UI) for a... | Cardiac Pacemakers, Inc. | A61N001/362 A61B005/044
US06999813 | B2 | 2006-02-14 | Biomedical signal processor and method of controlling biomedical signal processor | 2002-10-25 | To disclose a biomedical signal processor with a simple configuration and a minimum space capable of confirming an operating condition and a signal collection state of the biomedical signal processor such as an electrocardiograph. As an example of the biomedical signal processor according to the present invention, a Holter electrocardiograph supplies biomedical signals to a light-emitting element (LED) 30 and applies a bias supply voltage from a power supply section 80 to the LED 30. By setting a bias voltage to a voltage so that a level of a minimum supply voltage to the LED 30 becomes a positive potential when electrocardiogram signals are input to an input amplification section 10, the LED 30 turns OFF when power is turned OFF, becomes half-lit when power is turned ON and its luminescence intensity is variable when signals are collected normally, making it possible to easily check a state of the apparatus. | 1. A biomedical signal processor comprising: a power supply section that supplies drive power to the apparatus;an input amplification section that takes in and amplifies biomedical signals from a detection section that detects those biomedical signals;bias means for applying a predetermined bias voltage to an output signal from said input amplification section; andlight-emitting means for emitting light with intensity corresponding to the biomeidcal signals to which said bias means applies the predetermined bias voltage,wherein it is possible to simultaneously recognize an operating condition of said power supply section and a detection condition of said detection section from a luminescence intensity of said light-emitting means. | 1. A biomedical signal processor comprising: a power supply section that supplies drive power to the apparatus;an input amplification section that takes in and amplifies biomedical signals from a detection section that detects those biomedical signals;bias means for applying a predetermined bias voltage to an output signal from said input amplification section; andlight-emitting means for emitting light with intensity corresponding to the biomeidcal signals to which said bias means applies the predetermined bias voltage,wherein it is possible to simultaneously recognize an operating condition of said power supply section and a detection condition of said detection section from a luminescence intensity of said light-emitting means. 2. The biomedical signal processor according to claim 1, wherein said biomedical signals are used as electrocardiogram signals and said detection section is used as an electrocardiogram electrode, and said bias means applies a bias voltage so that the luminescence intensity of said light-emitting means becomes intermediate luminescence intensity when no electrocardiogram signal is detected from said input amplification section. 3. A method of controlling a biomedical signal processor provided with an input amplification section that takes in and amplifies biomedical signals from a detection section that detects those biomedical signals and light-emitting means, comprising the steps of: amplifying an input signal at said input amplification section;adding a predetermined bias voltage thereto; andletting said light-emitting means emit light with intensity corresponding to the biomedical signals with said bias voltage added,wherein it is possible to simultaneously recognize the operating condition of the power supply section and the detection condition of the detection section from the light-emitting state of said light-emitting means. 4. The method of controlling a biomedical signal processor according to claim 3, wherein said biomedical signals are used as electrocardiogram signals and said detection section is used as an electrocardiogram electrode, and a bias voltage is applied so that the luminescence intensity of said light-emitting means becomes intermediate luminescence intensity when no electrocardiogram signal is detected from said input amplification section. | Fukuda Denshi Co., Ltd. | A61B005/0402
US07167738 | B2 | 2007-01-23 | Method for navigating in the interior of the body using three-dimensionally visualized structures | 2001-07-31 | A method is described for navigating in the interior of the body using three-dimensionally visualized structures. In a first step of the method, at least two two-dimensional images of the same anatomical object are provided from different perspectives, and also items of information that make it possible to draw conclusions about the respective spatial position of an imaging system relative to the anatomical object. The projections of a geometrical structure to be visualized are then created in every two-dimensional image, wherein a geometrical structure to be visualized is created in each two-dimensional image, wherein the geometrical structure to be visualized is different from the anatomical object. A cone surface is then generated in space for each image wherein the spatial positions of the cone vertex and cone directrix are determined from the respective spatial position of the imaging system and the shape of the cone directrix is determined from the shape of the projection of the geometrical structure to be visualized on the image. Finally, a spatial intersection of the individual cone surfaces is formed to determine the geometrical structure and the geometrical structure determined and/or an intersection of a plurality of geometrical structures determined are/is represented and the representation is used for navigation. | 1. Method for navigating in the interior of the body using three-dimensionally visualized structures, comprising the following steps: providing at least two two-dimensional images (50, 50′) of the same anatomical object (12) from different perspectives and also of information which makes it possible to draw a conclusion about the respective spatial position of an imaging system (26) relative to the anatomical object (12);defining a projection (76, 76′, 78, 78′) of a geometrical structure (72, 80, 82) to be visualized or a part thereof in each two-dimensional image, wherein the geometrical structure (72, 80, 82) to be visualized is different from the anatomical object (12);generating a conical surface (40, 40′) in space for every image (50, 50′), wherein the spatial positions of cone vertex (46, 46′) and cone directrix are determined from the respective spatial position of the imaging system (26) and the shape of the cone directrix is determined from the shape of the projection (76, 76′, 78, 78′) of the geometrical structure (72, 80, 82) to be visualized;forming a spatial intersection of the individual conical surfaces (40, 40′) to determine the geometrical structure (72, 80, 82); anddisplaying the geometrical structure (72, 80, 82) determined and/or an intersection (74) of a plurality of geometrical structures (72, 80, 82) determined and using the representation for navigation. | 1. Method for navigating in the interior of the body using three-dimensionally visualized structures, comprising the following steps: providing at least two two-dimensional images (50, 50′) of the same anatomical object (12) from different perspectives and also of information which makes it possible to draw a conclusion about the respective spatial position of an imaging system (26) relative to the anatomical object (12);defining a projection (76, 76′, 78, 78′) of a geometrical structure (72, 80, 82) to be visualized or a part thereof in each two-dimensional image, wherein the geometrical structure (72, 80, 82) to be visualized is different from the anatomical object (12);generating a conical surface (40, 40′) in space for every image (50, 50′), wherein the spatial positions of cone vertex (46, 46′) and cone directrix are determined from the respective spatial position of the imaging system (26) and the shape of the cone directrix is determined from the shape of the projection (76, 76′, 78, 78′) of the geometrical structure (72, 80, 82) to be visualized;forming a spatial intersection of the individual conical surfaces (40, 40′) to determine the geometrical structure (72, 80, 82); anddisplaying the geometrical structure (72, 80, 82) determined and/or an intersection (74) of a plurality of geometrical structures (72, 80, 82) determined and using the representation for navigation. 2. Method according to claim 1, characterized in that the created projection of the geometrical structure (72, 80, 82) is a point, a straight line (76, 76′), a circular segment (78, 78′) or another structure having the form of a line. 3. Method according to claim 1 or 2, characterized in that the geometrical structure to be visualized is a point (72), a straight line (82), a plane, a sphere (80) or another two-dimensional or three-dimensional structure. 4. Method according to claim 1, 2 or 3, characterized in that the two-dimensional images (50, 50′) are generated by X-ray methods and/or magnetoresonance methods. 5. Method according to one of claims 1 to 4, characterized in that the spatial intersection is revised using at least one further data set of the anatomical object (12). 6. Method according to claim 5, characterized in that a two-dimensional or three-dimensional image or a generic model of the anatomical object (12) is used as further data set. 7. Method according to one of claims 1 to 6, characterized in that suitable perspectives are determined for further two-dimensional images by inverse calculations. 8. Method according to one of claims 1 to 7, characterized in that, additionally, the spatial position of an instrument (14) to be navigated is shown graphically relative to the geometrical structure (72, 80, 82) or to the intersection (74). 9. Method according to one of claims 1 to 8, characterized in that the effective axis (60, 60′, 60″, 60′″) of a surgical instrument (14) is shown graphically. 10. Method according to one of claims 1 to 9, characteriz... | Stryker Leibinger GmbH & Co., KG | A61B005/05 G06K009/00
US07072706 | B2 | 2006-07-04 | Equilibration method for high resolution imaging of lung compliance and distribution of functional residual capacity | 2002-02-08 | The present invention offers novel equilibration methods for very high resolution, three-dimensional imaging of imaging of lung compliance and distribution of functional residual capacity (FRC) in the lung using hyperpolarized helium-3 (3He) gas (H3He), and collecting local magnetic resonance image data therefrom. Using the present methods permits many functions that have been performed on a regional level for the whole lung using non-polarized helium, to be calculated for the first time from the local MRI measurements of local H3He, such as measuring volume or compliance. | 1. An equilibration method for very high resolution, three-dimensional imaging and determination of pulmonary compliance, wherein the method comprises: delivering and distributing a predetermined volume of hyperpolarized noble gas to conducting airways in each ventilated lung region of a pulmonary system, wherein the lung is inflated and sealed, or during a lung-inflating held breath (tidal volume);measuring distributed volume of local concentrations of the hyperpolarized noble gas in the airways by high resolution, three-dimensional local magnetic resonance imaging, wherein the lung image is segmented using scale-based fuzzy connectedness; andcalculating local pulmonary compliance from said imaged distribution. | 1. An equilibration method for very high resolution, three-dimensional imaging and determination of pulmonary compliance, wherein the method comprises: delivering and distributing a predetermined volume of hyperpolarized noble gas to conducting airways in each ventilated lung region of a pulmonary system, wherein the lung is inflated and sealed, or during a lung-inflating held breath (tidal volume);measuring distributed volume of local concentrations of the hyperpolarized noble gas in the airways by high resolution, three-dimensional local magnetic resonance imaging, wherein the lung image is segmented using scale-based fuzzy connectedness; andcalculating local pulmonary compliance from said imaged distribution. 2. The method of claim 1, wherein the noble gas is hyperpolarized helium-3 gas (H3He). 3. The method of claim 2, further comprising dividing the lung images into as many distinct voxels as imaging resolution permits. 4. The method of claim 3, further comprising calculating local lung volume by dividing average signal intensity in each voxel by tracheal signal intensity. 5. The method of claim 4, further comprising calculating average concentration of H3He in each voxel by dividing an amount of H3He in each voxel by volume of the voxel, and calculating amount of H3He in each voxel by multiplying concentration of H3He in the gas space of the voxel by volume of gas space in the voxel. 6. The method of claim 5, further comprising calculating compliance of the voxel by calculating pressure difference between alveolar gas inside the voxel and pleural space outside of the lung, and calculating local compliance in the voxel by dividing the volume of the voxel by transmural pressure gradient (tracheal minus esophageal pressures). 7. An equilibration method for very high resolution, three-dimensional imaging and determination of lung functional residual capacity (FRC) using hyperpolarized noble gas, wherein the method comprises: delivering and distributing a predetermined volume of hyperpolarized noble gas to conducting airways in each ventilated lung region of a pulmonary system;measuring distributed volume of local concentrations of the hyperpolarized noble gas to provide local FRC by high resolution, three-dimensional local magnetic resonance imaging, wherein the lung image is segmented using scale-based fuzzy connectedness; anddetermining whole lung FRC based upon summation of the local FRC measurements. 8. The method of claim 7, wherein the noble gas is hyperpolarized helium-3 gas (H3He). 9. The method of claim 8, further comprising dividing the lung images into as many distinct voxels as imaging resolution permits. 10. The method of claim 9, further comprising, calculating local lung volume by dividing average signal intensity in each voxel by tracheal signal intensity. 11. The method of claim 10, further comprising calculating average concentration of H3He in each voxel by dividing an amount of H3He in each voxel by volume of the v... | The Trustees of the University of Pennsylvania | A61B005/55
US07286866 | B2 | 2007-10-23 | Method, system and computer product for cardiac interventional procedure planning | 2002-03-15 | A method of creating 3D models to be used for cardiac interventional procedure planning. Acquisition data is obtained from a medical imaging system and cardiac image data is created in response to the acquisition data. A 3D model is created in response to the cardiac image data and three anatomical landmarks are identified on the 3D model. The 3D model is sent to an interventional system where the 3D model is in a format that can be imported and registered with the interventional system. | 1. A method of creating 3D models to be used for cardiac interventional procedure planning, the method comprising: obtaining acquisition data from a medical imaging system;creating cardiac image data from the acquisition data using a protocol configured for imaging a cardiac chamber;creating a 3D model of the cardiac chamber from the cardiac image data;identifying three anatomical landmarks on said 3D model;inserting three geometric markers at the three anatomical landmarks on the 3D model;exporting said 3D model to an interventional system, the 3D model including the geometric markers;importing said 3D model with the geometric markers at said interventional system;identifying the three anatomical landmarks in a coordinate system of the interventional system; andregistering said 3D model with the interventional system by transforming the 3D model with the geometric markers to the coordinate system of the interventional system using the three anatomical landmarks. | 1. A method of creating 3D models to be used for cardiac interventional procedure planning, the method comprising: obtaining acquisition data from a medical imaging system;creating cardiac image data from the acquisition data using a protocol configured for imaging a cardiac chamber;creating a 3D model of the cardiac chamber from the cardiac image data;identifying three anatomical landmarks on said 3D model;inserting three geometric markers at the three anatomical landmarks on the 3D model;exporting said 3D model to an interventional system, the 3D model including the geometric markers;importing said 3D model with the geometric markers at said interventional system;identifying the three anatomical landmarks in a coordinate system of the interventional system; andregistering said 3D model with the interventional system by transforming the 3D model with the geometric markers to the coordinate system of the interventional system using the three anatomical landmarks. 2. The method of claim 1 further comprising receiving from said interventional system electrical signals calculated before an ablation treatment and electrical signals calculated after the ablation treatment. 3. The method of claim 2 further comprising storing the electrical signals calculated before the ablation treatment and the electrical signals calculated after the ablation treatment in a DICOM format. 4. The method of claim 1 further comprising: creating 3D renderings for visual reference in response to said cardiac image data; andsending said 3D renderings to said interventional system. 5. The method of claim 4 wherein said 3D renderings are blended with a projection image on a fluoroscopy system. 6. The method of claim 1 wherein: said creating cardiac image data is performed using a tool to step an operator through said creating cardiac image data;said creating a 3D model is performed using a tool to step an operator through said creating a 3D model; andsaid identifying three anatomical landmarks on said 3D model is performed using a tool to step an operator through said identifying. 7. The method of claim 1 wherein said creating cardiac image data includes: segmenting said acquisition data using post processing software that includes a 3D protocol and instructions for extracting an inner surface of the area of interest; andvisualizing said area of interest including creating an immersible view. 8. The method of claim 7 wherein said post processing software further includes instructions for applying advanced vessel analysis. 9. The method of claim 7 wherein said post processing software further includes instructions for depositing seeds. 10. The method of claim 7 wherein said post processing software further includes instructions for using connectivity, thresholding and morphological operators. 11. The method of claim 7 wherein said post processing software further includes instructions for performing region growing techniques. 12. The method of claim 7 wherein said im... | GE Medical Systems Global Technology Company, LLC Jasbir S. Sra | A61B005/00
US07147604 | B1 | 2006-12-12 | High Q factor sensor | 2002-08-07 | A sensor for wirelessly determining a physical property within a defined space comprises an electrical resonance and has a high quality factor Q. The quality factor Q is sufficiently high that a signal generated by the sensor can be received outside the defined space. The sensor may optimally have a dielectric coating. | 1. A sensor for in vivo use within the body of a patient, comprising: a first substrate of polyimide, said first substrate having first and second sides;a first inductor comprised of a conductive material and disposed on said first side of said first substrate;a second substrate of polyimide, said second substrate having first and second sides;a second inductor comprised of a conductive material and disposed on said first side of said second substrate;said first and second substrates being disposed in parallel, spaced apart relation with said first sides of said substrates in apposition; andan intermediate layer disposed between said first and second substrates and maintaining said first and second inductors in spaced-apart relation, said intermediate layer having an opening formed therein such that at least a portion of said first and second inductors are movable toward one another through said opening in response to a change in an external condition. | 1. A sensor for in vivo use within the body of a patient, comprising: a first substrate of polyimide, said first substrate having first and second sides;a first inductor comprised of a conductive material and disposed on said first side of said first substrate;a second substrate of polyimide, said second substrate having first and second sides;a second inductor comprised of a conductive material and disposed on said first side of said second substrate;said first and second substrates being disposed in parallel, spaced apart relation with said first sides of said substrates in apposition; andan intermediate layer disposed between said first and second substrates and maintaining said first and second inductors in spaced-apart relation, said intermediate layer having an opening formed therein such that at least a portion of said first and second inductors are movable toward one another through said opening in response to a change in an external condition. 2. The sensor of claim 1, wherein said intermediate layer comprises an adhesive. 3. The sensor of claim 1, wherein said sensor is enveloped in silicone. 4. The sensor of claim 3, wherein said silicone is approximately 200 microns in thickness. 5. A sensor for in vivo use within the body of a patient, comprising: a first substrate of liquid crystal polymer, said first substrate having first and second sides;a first inductor comprised of a conductive material and disposed on said first side of said first substrate;a second substrate of liquid crystal polymer, said second substrate having first and second sides;a second inductor comprised of a conductive material and disposed on said first side of said second substrate;said first and second substrates being disposed in parallel, spaced apart relation with said first sides of said substrates in apposition; andan intermediate layer disposed between said first and second substrates and maintaining said first and second inductors in spaced-apart relation, said intermediate layer having an opening formed therein such that at least a portion of said first and second inductors are movable toward one another through said opening in response to a change in an external condition. 6. The sensor of claim 5, wherein said intermediate layer comprises a liquid crystal polymer. 7. The sensor of claim 5, wherein said sensor is enveloped in silicone. 8. The sensor of claim 7, wherein said silicone is approximately 200 microns in thickness. | CardioMEMS, Inc. | A61B005/00
US07220231 | B2 | 2007-05-22 | Walking condition determining device and method | 2002-08-23 | A device and a method that permit easy and accurate determination of a walking condition regardless of differences in foot landing points on soles or lengths of legs of a walker are provided. According to this device, a measuring device measures a parameter indicating a displacement amount of a bottom end portion of a leg of the walker. A first storage device stores patterns of plots in a determination space corresponding to parameters and walking conditions of the walker such that they are associated with each other. Furthermore, a generating device generates plots defined in the determination space by parameters measured by the measuring device. Then, a determining device determines walking conditions of the walker on the basis of the patterns of the plots stored in conjunction with the walking conditions by the first storage device and the patterns of the plots generated by the generating device. | 1. A walking condition determining device for determining a walking condition of a walker having a plurality of legs, comprising: a measuring means for measuring a parameter that indicates a displacement amount of a bottom end portion of a leg of the walker;a first storage means for storing a pattern of a plot in a determination space corresponding to the parameter and a walking condition of the walker such that they are associated with each other;a generating means for generating a plot defined in the determination space by the parameter measured by the measuring means; anda determining means for determining the walking condition of the walker on the basis of comparison between the pattern of the plot stored in association with the walking condition by the first storage means and a pattern of the plot generated by the generating means,wherein the first storage means stores an existence pattern of a plot in the determination space as the plot pattern, and the determining means determines the walking condition of the walker on the basis of the determination whether or not the existence pattern stored by the first storage means is identical or similar to an existence pattern of the plot generated by the generating means,wherein the measuring means comprises a first measuring means for measuring, as a first parameter, a difference between the length of a leg of the walker and a vertical interval between a top end portion and a bottom end portion of the leg, andthe determining means determines that the walker is in a normal walking condition if an existence pattern indicates that a plot in the determination space is located in a low region wherein the first parameter is below a predetermined threshold value, whereas it determines that the walker is in a slope walking condition if an existence pattern indicates that a plot is located in a high region wherein the first parameter is the predetermined threshold value or more. | 1. A walking condition determining device for determining a walking condition of a walker having a plurality of legs, comprising: a measuring means for measuring a parameter that indicates a displacement amount of a bottom end portion of a leg of the walker;a first storage means for storing a pattern of a plot in a determination space corresponding to the parameter and a walking condition of the walker such that they are associated with each other;a generating means for generating a plot defined in the determination space by the parameter measured by the measuring means; anda determining means for determining the walking condition of the walker on the basis of comparison between the pattern of the plot stored in association with the walking condition by the first storage means and a pattern of the plot generated by the generating means,wherein the first storage means stores an existence pattern of a plot in the determination space as the plot pattern, and the determining means determines the walking condition of the walker on the basis of the determination whether or not the existence pattern stored by the first storage means is identical or similar to an existence pattern of the plot generated by the generating means,wherein the measuring means comprises a first measuring means for measuring, as a first parameter, a difference between the length of a leg of the walker and a vertical interval between a top end portion and a bottom end portion of the leg, andthe determining means determines that the walker is in a normal walking condition if an existence pattern indicates that a plot in the determination space is located in a low region wherein the first parameter is below a predetermined threshold value, whereas it determines that the walker is in a slope walking condition if an existence pattern indicates that a plot is located in a high region wherein the first parameter is the predetermined threshold value or more. 2. The walking condition determining device according to claim 1, wherein the measuring means comprises a second measuring means for measuring, as a second parameter, a longitudinal position of a bottom end portion of a leg with respect to a top end portion of the leg of the walker, and when the determining means determines that the walker is in the slope walking condition, the determining means determines that the walker is in an ascending walking condition if the plot in the determination space indicates an existence pattern in which the plot is located in a predetermined positive region wherein the second parameter takes a positive threshold value or more, whereas the determining means determines that the walker is in a descending walking condition if the plot indicates an existence pattern in which the plot is in a predetermined negative region wherein the second parameter takes a negative threshold value or less. 3. The walking condition determining device according to claim 1, comprising a second storage means for storing jo... | Honda Giken Kogyo Kabushiki Kaisha | A61B005/103 A61B005/117
US06942622 | B1 | 2005-09-13 | Method for monitoring autonomic tone | 2002-07-29 | A method for monitoring the progression of the hemodynamic status of a patient by tracking autonomic tone. For example, the method may be applied to patients suffering from heart failure, diabetic neuropathy, cardiac ischemia, sleep apnea and hypertension. An implantable or other ambulatory monitor senses a pulse amplitude signal such as a vascular plethysmography signal. Variations of the signal amplitude on a scale greater than the heartbeat to heartbeat scale are indicative of variations in autonomic tone. A significant reduction in pulse amplitude and pulse amplitude variability are indicative of a heart failure exacerbation or other disease state change. This information may be used to warn the patient or healthcare providers of changes in the patient's condition warranting attention. | 1. A method for monitoring the progression of the disease state of a patient comprising the steps of; (a) sensing a pulse amplitude signal indicative of cardiac contraction of said patient's heart;(b) identifying amplitude changes within said sensed signal over time;(c) calculating a measure of autonomic tone based on said identified amplitude changes;(d) repeating steps (a) through (c) to acquire a plurality of autonomic tone measures; and(e) analyzing said plurality of autonomic tone measures to provide an indication of the progression of the disease state of said patient. | 1. A method for monitoring the progression of the disease state of a patient comprising the steps of; (a) sensing a pulse amplitude signal indicative of cardiac contraction of said patient's heart;(b) identifying amplitude changes within said sensed signal over time;(c) calculating a measure of autonomic tone based on said identified amplitude changes;(d) repeating steps (a) through (c) to acquire a plurality of autonomic tone measures; and(e) analyzing said plurality of autonomic tone measures to provide an indication of the progression of the disease state of said patient. 2. The method of claim 1 wherein said disease state comprises one or more of heart failure, diabetic neuropathy, cardiac ischemia, sleep apnea and hypertension. 3. The method of claim 1 wherein said step of sensing a pulse amplitude signal comprises sensing a vascular plethysmography signal. 4. The method of claim 1 and further including the step of storing a prescribed duration of the pulse amplitude signals in a long term memory. 5. The method of claim 1 wherein said step of sensing a pulse amplitude signal comprises sensing a measure of variability of said signal amplitude. 6. The method of claim 5 wherein said step (c) includes calculating a standard deviation of the pulse amplitudes. 7. The method of claim 1 wherein said step (e) includes the steps of calculating an average and a standard deviation of said plurality of autonomic tone measures. 8. The method of claim 7 and further including the step of determining a number n of a autonomic tone measures that differ by more than a prescribed amount from said average. 9. The method of claim 8 wherein said prescribed amount is a number greater than 1.5 time the calculated standard deviation. 10. The method of claim 1 and further including the step of using said output from step (e) to trigger a warning of the patient's changing condition. 11. The method of claim 10 wherein said step of triggering a warning includes inducing a physically perceptible stimulation of the patient. 12. The method of claim 1 and further including the step of periodically automatically telemetering some or all of the data collected and calculated in steps (a) through (e) to a receiver external to said patient. 13. A method for monitoring the progression of the disease state of a patient comprising the steps of: (a) sensing a pulse amplitude signal indicative of cardiac contractions of said patient's heart;(b) calculating a measure of autonomic tone based on amplitude variations of said sensed pulse amplitude signal;(c) repeating steps (a) and (b) to acquire a plurality of autonomic tone measures;(d) analyzing said plurality of autonomic tone measures to provide an indication of the progression of the disease state of said patient. 14. The method of claim 13 wherein said measure of autonomic tone is based on a number of said sensed pulse amplitude signals occurring in each of a plurality of predetermined counting intervals. 15. The meth... | Pacesetter, Inc. | A61B005/00 A61B005/021 A61B005/0452
US06974420 | B2 | 2005-12-13 | Method and apparatus for reducing noise and detecting electrode faults in medical equipment | 2002-07-22 | A method and apparatus for reducing noise and detecting electrode faults in a physiological activity acquisition system. The method includes the act of delivering a carrier signal through an electrode connected to a subject. Once the carrier signal is delivered, a combined signal having an electrical-activity portion and a carrier-signal portion is sensed by at least one signal sensing electrode attached to the subject. A low-pass, finite impulse response filter, having a first zero point frequency substantially the same as the carrier signal, separates the carrier signal portion from the electrical activity portion. An impedance value for the sensing electrode is calculated using the carrier signal portion. The calculated impedance value is compared against a known value to determine whether an electrode fault exists. | 1. A method of identifying an electrode fault in a system designed to measure electrical activity of a subject, the method comprising: delivering a carrier signal to the subject by connecting a first electrode to the subject;sensing a combined signal having an electrical-activity portion and a carrier-signal portion by attaching at least one signal sensing electrode to the subject;processing the combined signal detected by the sensing electrode by dividing the electrical activity portion from the carrier signal portion;calculating an impedance value for the at least one signal sensing electrode; anddetermining whether the calculated impedance for the at least one sensing electrode exceeds a predetermined value. | 1. A method of identifying an electrode fault in a system designed to measure electrical activity of a subject, the method comprising: delivering a carrier signal to the subject by connecting a first electrode to the subject;sensing a combined signal having an electrical-activity portion and a carrier-signal portion by attaching at least one signal sensing electrode to the subject;processing the combined signal detected by the sensing electrode by dividing the electrical activity portion from the carrier signal portion;calculating an impedance value for the at least one signal sensing electrode; anddetermining whether the calculated impedance for the at least one sensing electrode exceeds a predetermined value. 2. A method as claimed in claim 1, further comprising reducing high-frequency noise of the electrical activity portion of the combined signal. 3. A method as claimed in claim 1, wherein processing the combined signal detected by the sensing electrode includes filtering the combined signal in a low-pass filter with a first zero point frequency, and wherein the carrier signal has a frequency substantially the same as the first zero point frequency. 4. A method as claimed in claim 3, wherein the low-pass filter is a finite-impulse response filter. 5. A method as claimed in claim 4, wherein the low-pass filter has a cut-off frequency of about 150 Hz. 6. A method as claimed in claim 4, wherein calculating an impedance value involves using the carrier-signal portion. 7. A method as claimed in claim 1, wherein the combined signals are sensed by attaching multiple sensing electrodes to the subject. 8. A method of identifying an electrode fault in a system designed to measure electrical activity of a subject, the method comprising: delivering a carrier signal to the subject by connecting a first electrode to the subject;sensing a combined signal having an electrical-activity portion and a carrier-signal portion by attaching at least one signal sensing electrode to the subject;filtering the combined signal detected by the at least one signal sensing electrode with a low-pass filter having a zero point frequency; andcalculating the frequency of the filtered carrier-signal portion based on the zero-point frequency of the low-pass filter. 9. A method as claimed in claim 8, further comprising: identifying an impedance value for the at least one signal sensing electrode; anddetermining whether the calculated impedance for the at least one sensing electrode exceeds a predetermined value. 10. An apparatus for detecting an electrode fault, the apparatus comprising: a signal processor;a carrier signal generator;at least one physiological activity electrode to sense a signal and coupled to the signal processor;a carrier signal electrode coupled to the signal processor and the carrier signal generator; anda filter to remove a carrier signal from the signal sensed by the at least one physiological electrode, wherein the filter is a low-pass filter, and... | GE Medical Systems Information Technologies, Inc. | A61B005/05 A61N005/04
US07090638 | B2 | 2006-08-15 | System and method for optimized dietary menu planning | 2002-06-03 | A dietary menu planning system. The dietary menu planning system receives personal characteristics and food preferences for an individual. The personal characteristics includes a desired physiological rate of change for the individual. Based on the physiological rate of change and other personal characteristics, a set of dietary constraints is developed for the individual. A dietary menu plan is created for the individual including food items selected by the individual and with quantities determined to maximize the food preferences of the individual while satisfying the dietary constraints. The creation of the dietary menu plan includes applying an optimization function to maximize the food preferences. The dietary menu planning system is implemented as a stand-alone application for a computer or as a networked-system, for example, as a fee-for-service basis through the Internet. | 1. A method of creating an optimized dietary menu for an individual using a computer having a processor, said method comprising the steps of: receiving personal characteristics and a desired physiological change rate from the individual;generating a set of dietary constraints based upon said personal characteristics using the processor, said desired physiological change rate, and nutritional profile constraints, wherein said set of dietary constraints comprises a system of inequalities;receiving palatability scores for a plurality of food items, wherein each palatability score defines the palatability to the individual of the corresponding food item;generating the optimized dietary menu comprising a set of food items selected from said plurality of food items using the processor, wherein said set of food items is selected so as to satisfy said system of inequalities and maximize a resultant palatability score associated with said set of food items; andoutputting said optimized dietary menu,wherein said system of inequalities defines a relationship between said selected food items and one or more nutritional constraints, and wherein said step of generating the optimized dietary menu includes determining a quantity for each selected food item and maximizing an optimization function, wherein the optimization function comprises a sum of terms, and wherein each term comprises the product of the quantity and the palatability score associated with one of said selected food items, and wherein the result of said sum of terms comprises said resultant palatability score. | 1. A method of creating an optimized dietary menu for an individual using a computer having a processor, said method comprising the steps of: receiving personal characteristics and a desired physiological change rate from the individual;generating a set of dietary constraints based upon said personal characteristics using the processor, said desired physiological change rate, and nutritional profile constraints, wherein said set of dietary constraints comprises a system of inequalities;receiving palatability scores for a plurality of food items, wherein each palatability score defines the palatability to the individual of the corresponding food item;generating the optimized dietary menu comprising a set of food items selected from said plurality of food items using the processor, wherein said set of food items is selected so as to satisfy said system of inequalities and maximize a resultant palatability score associated with said set of food items; andoutputting said optimized dietary menu,wherein said system of inequalities defines a relationship between said selected food items and one or more nutritional constraints, and wherein said step of generating the optimized dietary menu includes determining a quantity for each selected food item and maximizing an optimization function, wherein the optimization function comprises a sum of terms, and wherein each term comprises the product of the quantity and the palatability score associated with one of said selected food items, and wherein the result of said sum of terms comprises said resultant palatability score. 2. The method as claimed in claim 1, wherein, for each of said selected food items, said quantity comprises a serving weight for said selected food item. 3. The method as claimed in claim 1, wherein said system of inequalities includes two or more inequalities, and wherein each inequality comprises a sum of nutritional terms and a relation between the sum of nutritional terms and one of said nutritional constraints. 4. The method as claimed in claim 3, wherein said sum of nutritional terms comprises a sum of contributions of said selected food items to a nutritional variable for which said one of said nutritional constraints prescribes a maximum or minimum. 5. The method as claimed in claim 4, wherein said one of said nutritional constraints comprises a maximum total caloric intake. 6. The method as claimed in claim 4, wherein said one of said nutritional constraints comprises a minimum dietary fibre. 7. The method as claimed in claim 3, wherein said step of generating said set of dietary constraints includes obtaining one or more recommended daily maximums or minimums of particular nutrients from a nutritional profile database, and wherein said one or more recommended daily maximums or minimums comprise said nutritional constraints within said system of inequalities. 8. A system for creating an optimized dietary menu plan for an individual, said system comprising: a food items datab... | - | A61B005/00
US07283863 | B2 | 2007-10-16 | Method and apparatus for identifying cardiac and non-cardiac oversensing using intracardiac electrograms | 2002-04-29 | A method and apparatus for automatically identifying various types of cardiac and non-cardiac oversensing is provided. EGM data, including time intervals between sensed and paced events and signal morphologies, are analyzed for patterns indicative of various types of oversensing, including oversensing of far-field R-waves, R-waves, T-waves, or noise associated with electromagnetic interference, non-cardiac myopotentials, a lead fracture, or a poor lead connection. Identification of oversensing and its suspected cause are reported so that corrective action may be taken. | 1. An implantable medical device system for identifying cardiac and non-cardiac oversensing using intracardiac electrograms, the system comprising: an implantable medical device (IMD) containing EGM data;means for accessing said EGM data;means for evaluating interval differences:means for evaluating interval patterns:means for determining the presence of short and long intervals;means for identifying oversensing based on the evaluated interval differences, the evaluated interval patterns, and the determined presence of short and long intervals; andmeans for recommending a corrective action based on at least one characteristic of an identified oversensing to implement said corrective action in said IMD. | 1. An implantable medical device system for identifying cardiac and non-cardiac oversensing using intracardiac electrograms, the system comprising: an implantable medical device (IMD) containing EGM data;means for accessing said EGM data;means for evaluating interval differences:means for evaluating interval patterns:means for determining the presence of short and long intervals;means for identifying oversensing based on the evaluated interval differences, the evaluated interval patterns, and the determined presence of short and long intervals; andmeans for recommending a corrective action based on at least one characteristic of an identified oversensing to implement said corrective action in said IMD. 2. The system of claim 1 wherein said EGM data relates to data collected by an arrhythmia detection means. 3. The system of claim 1 wherein said EGM data relates to data collected by monitoring means. 4. The system of claim 1 wherein said means for accessing includes one of a programmer, a personal computer and an external device in data communication with the said IMD. 5. The system of claim 1 wherein said EGM includes EGM signal, sensed and/or paced interval data, and annotations of sensed and/or paced events. 6. The system of claim 4 wherein one of said programmer, said personal computer and said external device is implemented to perform said analysis of said EGM data. 7. The system of claim 1 wherein said means for accessing is implemented in said IMD. 8. The system of claim 1 wherein said means for performing analysis is implemented in said IMD. 9. The system of claim 1 wherein said means for recommending a corrective action is implemented in said IMD. 10. The system of claim 9 wherein said corrective action includes warning a patient via a notification means. 11. The system of claim 10 wherein said corrective action includes one of reprogramming a sensitivity setting, reprogramming a sensing electrode configuration, tightening set screws in the connector block, investigating for a likely lead fracture to thereby repair or replace, and implement investigation to eliminate oversensing. | Medtronic, Inc. | A61B005/04
US07286871 | B2 | 2007-10-23 | Method and apparatus for reducing contamination of an electrical signal | 2001-08-15 | The method of reducing contamination of electrical signals recorded in the presence of repeated interference contamination comprises obtaining an electrical signal recorded in the presence of a contaminating signal, and detecting a timing signal that occurs at a fixed time point during the electrical signal relative to the onset of the contaminating signal. The electrical signal is digitized, wherein the digitizing begins with the timing signal. A plurality of digitized electrical signals is analyzed, wherein the electrical signals are synchronized with respect to the timing signal, to obtain an estimated contaminating signal that is subtracted from the digitized electrical signal. This method can be used with electrophysiological signals, such as EEG, ECG, EMG and galvanic skin response, and for elimination of noise associated with concurrently used methods such as MRI. The method of noise reduction is applicable to recordings of other electrical signals, including audio recordings. | 1. A method of reducing contamination of electrical signals recorded in the presence of repeated interference contamination comprising: (a) obtaining an electrical signal, wherein the electrical signal was recorded in the presence of a contaminating signal;(b) detecting a timing signal that occurs at a fixed time point during the electrical signal relative to the onset of the contaminating signal;(c) digitizing the electrical signal, wherein the digitizing begins with the timing signal;(d) analyzing a plurality of digitized electrical signals, wherein the electrical signals are synchronized with respect to the timing signal, to obtain an estimated contaminating signal; and(e) subtracting the estimated contaminating signal from the digitized electrical signal, thereby reducing contamination of the electrical signal;wherein the analyzing to obtain the estimate of the contaminating signal comprises calculating a weighted average of the contaminating signals; and wherein the estimate of the contaminating signal is biased towards recent events, and comprises a first sum consisting of the nth electrical signal added to a scalar multiple, w, of a prior estimate of the contaminating signal, divided by a second sum obtained by adding a series 1+w2+w3+w4 . . . + . . . wn. | 1. A method of reducing contamination of electrical signals recorded in the presence of repeated interference contamination comprising: (a) obtaining an electrical signal, wherein the electrical signal was recorded in the presence of a contaminating signal;(b) detecting a timing signal that occurs at a fixed time point during the electrical signal relative to the onset of the contaminating signal;(c) digitizing the electrical signal, wherein the digitizing begins with the timing signal;(d) analyzing a plurality of digitized electrical signals, wherein the electrical signals are synchronized with respect to the timing signal, to obtain an estimated contaminating signal; and(e) subtracting the estimated contaminating signal from the digitized electrical signal, thereby reducing contamination of the electrical signal;wherein the analyzing to obtain the estimate of the contaminating signal comprises calculating a weighted average of the contaminating signals; and wherein the estimate of the contaminating signal is biased towards recent events, and comprises a first sum consisting of the nth electrical signal added to a scalar multiple, w, of a prior estimate of the contaminating signal, divided by a second sum obtained by adding a series 1+w2+w3+w4 . . . + . . . wn. 2. The method of claim 1, wherein the electrical recording comprises an electrophysiological signal. 3. The method of claim 1, wherein the estimate of the contaminating signal is multiplied by a scalar prior to the subtracting of step (e). 4. The method of claim 2, wherein the electrophysiological signal comprises an electroencephalographic recording, an electromyelographic recording, an electrocardiographic recording or a measure of galvanic skin resistance. 5. The method of claim 1, wherein the interference comprises interference arising from inductively coupled magnetic fields. 6. The method of claim 1, wherein the interference comprises interference arising from alternating current (AC) line noise. 7. A method of reducing contamination of electrical signals recorded in the presence of repeated interference contamination comprising: (a) obtaining an electrical signal, wherein the electrical signal was recorded in the presence of a contaminating signal;(b) detecting a timing signal that occurs at a fixed time point during the electrical signal relative to the onset of the contaminating signal;(c) digitizing the electrical signal, wherein the digitizing begins with the timing signal;(d) analyzing a plurality of digitized electrical signals, wherein the electrical signals are synchronized with respect to the timing signal, to obtain an estimated contaminating signal; and(e) subtracting the estimated contaminating signal from the digitized electrical signal, thereby reducing contamination of the electrical signal,wherein the digitizing is performed at a sampling rate below the Nyquist rate for the contaminating signal. 8. The method of claim 1, wherein the electrical signal obtained... | The Regents of the University of California | A61B005/04
US07261713 | B2 | 2007-08-28 | Adjustable fixator | 2002-10-07 | A bone fixator for securing a first bone portion in a position relative to second bone portion is disclosed, including at least two clamping assemblies each for receiving at least one bone pin and a main body disposed between the clamping assemblies. The main body includes at least one joint for orienting the clamping assemblies with respect to each other. The at least one joint may permit pivoting or rotation. Some joints permit the articulation of several components to be simultaneously and releasably lockable. | 1. A bone fixator comprising: at least two clamping assemblies each for receiving at least one bone fastener;a main body disposed between the clamping assemblies, the main body having first and second ends;a first coupling member being pivotably coupled to the first end of the main body about a first axis;a second coupling member being pivotably coupled to the second end of the main body about a second axis;a wedge disposed within the main body having a borehole disposed along a translation axis;a locking member disposed along the translation axis and operatively associated with the wedge;a first moveable element disposed between the first coupling member and the wedge;a second moveable element disposed between the second coupling member and the wedge, wherein the wedge is engageable with the moveable elements to arrest pivoting of the coupling members and disengageable from the moveable elements to permit pivoting thereof with respect to the main body; anda joint assembly comprising: a joint assembly body with a borehole extending therethrough along a borehole axis and a fastener hole extending along a fastener axis transverse to the borehole axis, the body having an outer surface and an inner borehole surface;a slit extending along the borehole axis and across the fastener hole from the outer surface to the inner borehole surface, the slit defining opposed slit surfaces having a separation width; anda fastener configured and dimensioned to be received in the fastener hole,wherein the size of the borehole is adjustable by changing the separation width of the opposed slit surfaces. | 1. A bone fixator comprising: at least two clamping assemblies each for receiving at least one bone fastener;a main body disposed between the clamping assemblies, the main body having first and second ends;a first coupling member being pivotably coupled to the first end of the main body about a first axis;a second coupling member being pivotably coupled to the second end of the main body about a second axis;a wedge disposed within the main body having a borehole disposed along a translation axis;a locking member disposed along the translation axis and operatively associated with the wedge;a first moveable element disposed between the first coupling member and the wedge;a second moveable element disposed between the second coupling member and the wedge, wherein the wedge is engageable with the moveable elements to arrest pivoting of the coupling members and disengageable from the moveable elements to permit pivoting thereof with respect to the main body; anda joint assembly comprising: a joint assembly body with a borehole extending therethrough along a borehole axis and a fastener hole extending along a fastener axis transverse to the borehole axis, the body having an outer surface and an inner borehole surface;a slit extending along the borehole axis and across the fastener hole from the outer surface to the inner borehole surface, the slit defining opposed slit surfaces having a separation width; anda fastener configured and dimensioned to be received in the fastener hole,wherein the size of the borehole is adjustable by changing the separation width of the opposed slit surfaces. 2. The bone fixator of claim 1, wherein the borehole and locking member are threadably associated with each other. 3. The bone fixator of claim 1, wherein the locking member is fixed in position along the translation axis and rotatable thereabout. 4. The bone fixator of claim 1, wherein the moveable elements are slidably associated with the wedge. 5. The bone fixator of claim 1, wherein at least one of the coupling members and at least one of the moveable elements each have textured portions, the textured portion of the at least one coupling member being positively lockable with the textured portion of the at least one moveable element. 6. The bone fixator of claim 5, wherein the textured portions include textures selected from serrations and facets. 7. The bone fixator of claim 1, wherein at least a portion of the main body has an inner cylindrical surface and the moveable elements are configured and dimensioned to be slidably associated with the cylindrical surface. 8. The bone fixator of claim 1, wherein the moveable elements are resiliently biased toward each other. 9. The bone fixator of claim 1, wherein the wedge comprises first and second outer engagement surfaces disposed transverse to each other. 10. The bone fixator of claim 9, wherein the first and second outer engagement surfaces are disposed at an angle of between about 10° and about 70° with resp... | Synthes (USA) | A61B005/04
US07270638 | B2 | 2007-09-18 | Method and device for diagnosis using an oscillating airflow | 2003-01-31 | It has surprisingly been found that nasal NO release is greatly increased by the presence of a an oscillating airflow in healthy subjects. Accordingly, a device comprising means for generating and/or maintaining an oscillating airflow is described and suggested for use in the diagnosis of a condition of the upper airways of a human. The invention also encompasses method wherein the concentration and/or flow of at least one gaseous component in nasally exhaled air is determined and recorded in the presence and absence of an oscillating airflow acting on the upper airways or parts thereof. | 1. A device for use in diagnosis of a condition of the upper airways in a human, wherein said device comprises means for generating and/or maintaining an oscillating airflow in said upper airways or parts thereof; said device being suitable for connection to means for determining and recording the concentration and/or flow of nitric oxide (NO) present in exhaled air of said human; and wherein the condition of the upper airways is an absence or presence of a condition selected from the group consisting of an inflammatory condition, an airway infection, a common cold, a tumor, a drug related effect, an anatomical abnormality, a patency of the sinus ostia, a size of the sinuses, a biochemical status of the sinuses, a sinusitis affecting one or more sinuses, a location of said sinusitis, a risk for developing sinusitis, a bacteriological status of the sinuses, and a combination thereof. | 1. A device for use in diagnosis of a condition of the upper airways in a human, wherein said device comprises means for generating and/or maintaining an oscillating airflow in said upper airways or parts thereof; said device being suitable for connection to means for determining and recording the concentration and/or flow of nitric oxide (NO) present in exhaled air of said human; and wherein the condition of the upper airways is an absence or presence of a condition selected from the group consisting of an inflammatory condition, an airway infection, a common cold, a tumor, a drug related effect, an anatomical abnormality, a patency of the sinus ostia, a size of the sinuses, a biochemical status of the sinuses, a sinusitis affecting one or more sinuses, a location of said sinusitis, a risk for developing sinusitis, a bacteriological status of the sinuses, and a combination thereof. 2. A device according to claim 1, wherein the means for generating and/or maintaining an oscillating airflow comprises means for electronically or mechanically generating said airflow. 3. A device according to claim 1, wherein the means for generating and/or maintaining an oscillating airflow comprises means for registering a duration, frequency and/or volume of an oscillating airflow produced by a patient coupled to a means for indicating to said patient that a pre-set necessary duration, frequency and/or volume has been reached. 4. A device according to claim 1, wherein the oscillating airflow oscillates at a frequency between 1 to 1000 Hz. 5. A device according to claim 1, wherein the oscillating airflow oscillates at a frequency between 10 to 1000 Hz. 6. A device according to claim 1, wherein the oscillating airflow oscillates at a frequency between 100 to 1000 Hz. 7. A device according to claim 1, wherein the oscillating airflow oscillates at a frequency between 100 to 500 Hz. 8. A device according to claim 1, wherein the oscillating airflow oscillates at a frequency close to a resonance frequency of the paranasal sinuses. 9. A device according to claim 1, wherein said device comprises means for unilaterally diagnosing an abnormal state affecting only the right or left hand side of the nasal airways including paranasal sinuses. 10. A device for use in diagnosis of a condition of the upper airways in a human, wherein said device comprises means for generating and/or maintaining an oscillating airflow in said upper airways or parts thereof and means for determining and recording a concentration and/or flow of nitric oxide (NO) present in exhaled air of said human, and means for supplying a breathing gas devoid of or containing only minute and/or known amounts of NO to said human. 11. A device according to claim 10, wherein said breathing gas is a nitric oxide-free gas suitable for inhalation. 12. A device according to claim 10, wherein said device is suitable for connection to means for determining and recording a concentration and/or flow of a tracer ga... | Aerocrine AB | A61B005/08
US06999809 | B2 | 2006-02-14 | Central venous catheter having a soft tip and fiber optics | 2002-07-16 | A multiple lumen catheter having a soft, tapered multiple lumen distal tip for insertion into a body vessel. One of the lumens is sized to pass over a guidewire such that the catheter can be inserted into the body vessel using the Seldinger technique. At least one medical implement lumen is used for placement or positioning of a biomedical sensor or other medical implement. For example, at least one optical fiber passing through the medical implement lumen may transmit and receive light at the distal tip for measuring oxygen saturation of the blood. The catheter may have a cylindrical catheter body to which the soft distal tip attaches. The soft tip reduces the possibility of vessel or tissue puncture and abrasion. The tip is constructed of a soft plastic or pliable material that yields easily when force is applied. For example, the tip may be made of a softer material than the catheter body, or if made of the same material, the tip can be configured with thinner walls or a higher air-to-material ratio cross-section. Various geometrical configurations and combinations of materials can be used to decrease the resistance of the tip to an applied load. One particular useful application for the catheter of the present invention is as a central venous catheter equipped with fibers for measuring oximetry. The fibers extend to the distal end of the tip and are preferably secured therein with minimal adhesive so as to limit the stiffness added to the tip. One particular useful construction is to secure the fibers within the medical implement lumen using adhesive only along a length of between about 0.5–3.5 mm. | 1. A multiple lumen central venous catheter having a soft tip, comprising: a generally tubular catheter body having a proximal end and a distal end, the catheter body defining therein at least one primary lumen and at least one medical implement lumen;a soft, tapered distal catheter tip having a distal end and a proximal end abutting the distal end of the catheter body, the catheter tip defining therein at least one primary lumen aligned with the at least one primary lumen of the catheter body and at least one medical implement lumen aligned with the at least one medical implement lumen of the catheter body, wherein the catheter tip has a proximal region and a distal region, and wherein the proximal region is more flexible than the distal region and more flexible than the catheter body and forms a bending portion of the catheter tip; andat least one optical fiber extending through the at least one medical implement lumen of the catheter body and the at least one medical implement lumen of the catheter tip, and being secured thereto only at the distal end of the catheter tip. | 1. A multiple lumen central venous catheter having a soft tip, comprising: a generally tubular catheter body having a proximal end and a distal end, the catheter body defining therein at least one primary lumen and at least one medical implement lumen;a soft, tapered distal catheter tip having a distal end and a proximal end abutting the distal end of the catheter body, the catheter tip defining therein at least one primary lumen aligned with the at least one primary lumen of the catheter body and at least one medical implement lumen aligned with the at least one medical implement lumen of the catheter body, wherein the catheter tip has a proximal region and a distal region, and wherein the proximal region is more flexible than the distal region and more flexible than the catheter body and forms a bending portion of the catheter tip; andat least one optical fiber extending through the at least one medical implement lumen of the catheter body and the at least one medical implement lumen of the catheter tip, and being secured thereto only at the distal end of the catheter tip. 2. The central venous catheter of claim 1, wherein the at least one optical fiber extends to a distal face of the catheter tip and is secured thereto. 3. The central venous catheter of claim 1, wherein the optical fiber is secured using adhesive that is only applied along a short portion of the medical implement lumen of the catheter tip having an axial length of between 0.5–3.5 mm. 4. The central venous catheter of claim 3, wherein the adhesive is cured by ultraviolet light, and wherein a portion of the catheter tip permits passage of ultraviolet light therethrough into the medical implement lumen. 5. The central venous catheter of claim 4, wherein the catheter tip is formed of two materials, one material comprising the portion that permits passage of ultraviolet light, and another material that provides a radiopaque marker. 6. The central venous catheter of claim 1, wherein the catheter tip is tapered to have a smaller diameter on its distal end than on its proximal end and has a proximal region and a distal region, the catheter tip having at least two exterior taper angles, and wherein the proximal region has a greater taper angle than the distal region. 7. The central venous catheter of claim 1, wherein the proximal region is made of a material that has a durometer of less than or equal to 100 Shore A hardness. 8. The central venous catheter of claim 1, wherein the proximal region has a larger air-to-material ratio in transverse cross-section than the distal region. 9. The central venous catheter of claim 8, wherein, in transverse cross-section, both the at least one primary lumen and the at least one medical implement lumen are enlarged in the proximal region relative to the surrounding material in comparison to the relative size of the lumens and material in the distal region, such that the proximal region has a larger air-to-material ratio than the distal regio... | Edwards Lifesciences Corporation | A61B005/00 A61M025/00 A61N001/04
US07016719 | B2 | 2006-03-21 | System and methods for noninvasive electrocardiographic imaging (ECGI) using generalized minimum residual (GMRes) | 2002-10-04 | Methods and systems for computing epicardial surface electric potentials based on measured body surface electric potentials, where the methods and systems include representing at least one geometric relationship between at least one body surface electric potential measuring system and the epicardial surface as a multidimensional matrix, estimating an inverse of the multidimensional matrix based on a Generalized Minimum Residual (GMRes) method, and, based on the inverse matrix and the measured body surface potentials, determining the epicardial surface electric potentials. | 1. A method for computing epicardial surface electric potentials based on measured body surface electric potentials, the method comprising: representing at least one geometric relationship between at least one body surface electric potential measuring system and the epicardial surface, as a multidimensional matrix,using a Generalized Minimum Residual (GMRes) method to estimate an inverse of the multidimensional matrix, and,based on the inverse matrix and the measured body surface potentials, determining the epicardial surface electric potentials. | 1. A method for computing epicardial surface electric potentials based on measured body surface electric potentials, the method comprising: representing at least one geometric relationship between at least one body surface electric potential measuring system and the epicardial surface, as a multidimensional matrix,using a Generalized Minimum Residual (GMRes) method to estimate an inverse of the multidimensional matrix, and,based on the inverse matrix and the measured body surface potentials, determining the epicardial surface electric potentials. 2. A method according to claim 1, where representing includes measuring the position of the at least one body surface electric potential measuring system. 3. A method according to claim 1, where representing includes measuring the position of at least one electrode. 4. A method according to claim 1, where representing includes providing at least one of a CT scan, a MRI and an X-ray. 5. A method according to claim 1, where representing includes determining an epicardial envelope. 6. A method according to claim 1, where representing includes employing a boundary element method. 7. A method according to claim 1, where using a GMRes method includes determining a number of iterations for the GMRes method. 8. A method according to claim 7, where determining a number of iterations includes comparing residual error to a Hessenberg matrix condition, and computing at least one of a corner of a condition L curve and a maximum curvature of a condition L curve. 9. A method according to claim 7, where determining a number of iterations includes determining a number of iterations based on at least one of: a corner of a condition L curve, a corner of an L curve, an increase in spatial frequency of a reconstructed potential map, and an increase in amplitude of a solution norm. 10. A method according to claim 1, where using a GMRes method includes providing a maximum number of iterations for the GMRes method, and based on the data from the maximum number of iterations, determining a number of iterations for the GMRes method. 11. A method according to claim 1, where using a GMRes method includes providing an initial condition of zero. 12. A method according to claim 1, where using a GMRes method includes providing an initial condition based on a Tikhonov regularization of the multidimensional matrix. 13. A method for computing electric potentials on an epicardial surface of a patient, the method comprising: measuring electric potentials on the patient's body surface,expressing a geometrical relationship between the patient's body surface and the epicardial surface as a multi-dimensional matrix,determining an approximation of the two-dimensional matrix based on a Generalized Minimum Residual (GMRes) method, and,computing the electric potentials on the epicardial surface based on the approximated inverse and the measured electric potentials. 14. A method according to claim 13, where measuring electric potentials... | Case Western Reserve University | A61B005/04
US07142905 | B2 | 2006-11-28 | Visual imaging system for ultrasonic probe | 2003-02-04 | A non-invasive visual imaging system is provided, wherein the imaging system procures an image of a transducer position during diagnostic or therapeutic treatment. In addition, the system suitably provides for the transducer to capture patient information, such as acoustic, temperature, or ultrasonic images. For example, an ultrasonic image captured by the transducer can be correlated, fused or otherwise combined with the corresponding positional transducer image, such that the corresponding images represent not only the location of the transducer with respect to the patient, but also the ultrasonic image of the region of interest being scanned. Further, a system is provided wherein the information relating to the transducer position on a single patient may be used to capture similar imaging planes on the same patient, or with subsequent patients. Moreover, the imaging information can be effectively utilized as a training tool for medical practitioners. | 1. A non-invasive treatment and imaging system comprising: a control system configured for control of said non-invasive treatment and imaging system;an extracorporeal probe assembly configured for treatment of a patient, said probe assembly comprising a transducer configured for obtaining patient information and for facilitating treatment of the patient; andan imaging device configured for capturing positional information of said transducer with respect to a patient, said imaging device being directly attached to said transducer to move along with said transducer and to facilitate tracking any movement of said transducer. | 1. A non-invasive treatment and imaging system comprising: a control system configured for control of said non-invasive treatment and imaging system;an extracorporeal probe assembly configured for treatment of a patient, said probe assembly comprising a transducer configured for obtaining patient information and for facilitating treatment of the patient; andan imaging device configured for capturing positional information of said transducer with respect to a patient, said imaging device being directly attached to said transducer to move along with said transducer and to facilitate tracking any movement of said transducer. 2. A system according to claim 1, wherein said control system is configured for assessing the patient information and the positional information to facilitate use of said transducer at a desired position with respect to the patient. 3. A system according to claim 1, wherein said transducer comprises an ultrasonic transducer, said ultrasonic transducer being adapted to transmit imaging information comprising ultrasonic images of a treatment region of the patient to said control system. 4. A system according to claim 3, wherein said control system further comprises a central processing unit, said central processing unit comprising an algorithm configured for processing the imaging information and the positional information. 5. A system according to claim 4, wherein said control system further comprises a memory unit for storing the imaging information and the positional information. 6. A system according to claim 4, wherein said control system further comprises a display unit for displaying the imaging information and the positional information. 7. A system according to claim 6, wherein the positional information is combined via said algorithm within said control system with at least one of an acoustic image, a thermal image, and a therapy treatment image to provide a combined image. 8. A system according to claim 6, wherein the positional information is combined via said algorithm and said control system with the imaging information to provide a fused image. 9. A system according to claim 8, wherein said fused image comprises the positional information configured to display a position of said transducer relative to the patient in fusion with the imaging information configured to display an ultrasonic image of a treatment region. 10. A system according to claim 9, wherein said fused image provides an image similar to a cross-sectional view of the patient and the treatment region. 11. A system according to claim 9, wherein said display unit is configured to display a primary image comprising said fused image, and is configured to display a secondary image comprising one of the positional information configured to display the position of said transducer relative to the patient and the imaging information configured to display the ultrasonic image of the treatment region. 12. A system according to claim 8, wherein said pro... | Guided Therapy Systems, Inc. | A61B005/00
US06993374 | B2 | 2006-01-31 | Instrumentation and method for mounting a surgical navigation reference device to a patient | 2002-04-17 | Instrumentation and methods are provided for mounting a surgical navigation reference frame to a patient for performing an image-guided surgical procedure on an anatomical component. In one embodiment of the invention, a trocar is positioned within a cannula to form an insertion device adapted for percutaneous introduction into the patient. A bone anchor having a bone engaging portion is inserted through the cannula and is anchored to bone. The bone anchor cooperates with the cannula to form a mounting device adapted for coupling with the surgical navigation reference frame. In one embodiment of the invention, an image-guided surgical procedure is performed at a location remote from the anchoring location. In a specific embodiment, an image-guided surgical procedure is performed adjacent the spinal column, with the bone engaging portion of the bone anchor anchored to the iliac region of the pelvic bone, and more specifically to the iliac crest of the pelvic bone. | 1. A method of mounting a surgical navigation reference device to a patient for performing an image-guided surgical procedure said surgical procedure being performed on the patient's spinal column, comprising: providing a bone anchor having a bone engaging portion;anchoring the bone engaging portion of the bone anchor to bane at a location remote from the patient's spinal column; andcoupling the reference device to the bone anchor to mount the reference device in a substantially fixed position relative to the patient's spinal column. | 1. A method of mounting a surgical navigation reference device to a patient for performing an image-guided surgical procedure said surgical procedure being performed on the patient's spinal column, comprising: providing a bone anchor having a bone engaging portion;anchoring the bone engaging portion of the bone anchor to bane at a location remote from the patient's spinal column; andcoupling the reference device to the bone anchor to mount the reference device in a substantially fixed position relative to the patient's spinal column. 2. The method of claim 1, wherein the anchoring comprises anchoring the bone engaging portion of the bone anchor to a portion of the patient's pelvic bone. 3. The method of claim 2, wherein the portion of the pelvic bone is the iliac region of the pelvic bone. 4. The method of claim 3, wherein the portion of the pelvic bone is the iliac crest. 5. The method of claim 1, further comprising percutaneously inserting the bone engaging portion of the bone anchor into the patient. 6. The method of claim 5, wherein the percutaneously inserting comprises: providing a cannula having a proximal end and a distal end;providing a trocar having a distal end portion;positioning the trocar within the cannula with the distal end portion of the trocar extending beyond the distal end of the cannula;percutaneously introducing the cannula into the patient;removing the trocar from the cannula; andinserting the bone engaging portion of the bone anchor through the cannula and into engagement with the bone. 7. The method of claim 1, further comprising: forming a surgical incision at a location remote from the anchoring location; andperforming an image-guided surgical procedure on the patient through the surgical incision. 8. The method of claim 7, wherein the surgical incision is formed subsequent to the anchoring. 9. The method of claim 1, further comprising performing a percutaneous image-guided surgical procedure at a location remote from the anchoring location. 10. The method of claim 1 wherein said bone engaging portion of the bone anchor includes treads and the anchoring is accomplished by use of said threads. 11. A method of mounting a surgical navigation reference device to a patient for performing an image-guided surgical procedure, comprising: providing a cannula having a proximal end and a distal end, a trocar having a distal end portion, and a bone anchor having a bone engaging portion, said cannula and bone anchor being separate elements not connected to one another;positioning the trocar within the cannula with the distal end portion of the trocar extending beyond the distal end of the cannula;percutaneously introducing the cannula into the patient;removing the trocar from the cannula;inserting the bone engaging portion of the bone anchor through the cannula;anchoring the bone engaging portion of the bone anchor to bone; andcoupling the surgical navigation reference device to the bone anchor. 12. The method of claim 1... | - | A61B005/05
US07108666 | B2 | 2006-09-19 | Method and apparatus for performing a forced expiratory maneuver in an infant | 2003-01-07 | Disclosed is a method and device for performing forced expiratory maneuver in an infant to assess the infant's pulmonary function. Under this method, the infant's lungs are synchronously inflated to super-atmospheric levels synchronous with the infant's natural tidal inspiration for a plurality of consecutive respiratory cycles. The end-expiratory CO2 levels in the infant's respiration are measured during the test. When the end-expiratory CO2 concentration decreases from the baseline by a pre-defined amount, the infant lungs are rapidly inflated to substantially total lung volume and rapidly deflated to produce a maximum forced expiration. The pre-defined amount of change in CO2 concentration is usually determined by the testing clinician. Typical concentration drop in CO2 levels, for example, ranges between 4 and 8 mmHg. But the decrease may also be as little as 2 mmHg or as much as 15 mmHg, depending on the testing clinician. The decrease in the end-expiratory CO2 level of the infant indicates that the infant's respiratory center is sufficiently modified to allow for the measurement. | 1. A method of performing a forced expiratory maneuver in an infant comprising the steps of: a) inflating the lungs of the infant with air synchronously with natural tidal inspiration to a lung volume greater than that reached at the end tidal inspiration for a plurality of consecutive respiratory cycles;b) monitoring end-tidal CO2 concentration of the infant's respiration during each respiratory cycle;c) determining whether the end-tidal CO2 concentration decreases from a baseline concentration by a pre-defined amount;d) inflating the infant's lungs to substantially total lung volume after a determination that the end-tidal CO2 concentration decreases from the baseline concentration by the pre-defined amount; ande) immediately deflating the infant's lungs to produce a maximum forced expiration. | 1. A method of performing a forced expiratory maneuver in an infant comprising the steps of: a) inflating the lungs of the infant with air synchronously with natural tidal inspiration to a lung volume greater than that reached at the end tidal inspiration for a plurality of consecutive respiratory cycles;b) monitoring end-tidal CO2 concentration of the infant's respiration during each respiratory cycle;c) determining whether the end-tidal CO2 concentration decreases from a baseline concentration by a pre-defined amount;d) inflating the infant's lungs to substantially total lung volume after a determination that the end-tidal CO2 concentration decreases from the baseline concentration by the pre-defined amount; ande) immediately deflating the infant's lungs to produce a maximum forced expiration. 2. The method of claim 1 wherein the pre-defined amount ranges between 2 mmHg to 15 mmHg. 3. The method of claim 2 wherein the pre-defined amount ranges between 4 mmHg to 8 mmHg. 4. The method of claim 1 wherein the step of determining whether the end-tidal CO2 concentration decreases from the baseline concentration is performed by a computer. 5. The method of claim 1 wherein the inflation of the infant's lungs to substantially total lung volume comprises the step of inflating the infant's lungs to a pressure of about 30 cm H2O. 6. The method of claim 1 wherein said deflating step (e) is performed by compressing the infant's chest and abdomen with an inflatable vest. 7. The method of claim 1 wherein the deflating step (e) is performed by sucking the air out of the infant's lungs using a negative pressure reservoir. 8. A method of performing a forced expiratory maneuver in an infant comprising the steps of: a) inflating the lungs of the infant with air synchronously with natural tidal inspiration to a lung volume greater than that reached at the end tidal inspiration for a plurality of consecutive respiratory cycles;b) monitoring end-tidal CO2 concentration of the infant's respiration during each respiratory cycle;c) determining whether the end-tidal CO2 concentration decreases from a baseline concentration by a pre-defined amount;d) inflating the infant's lungs to substantially total lung volume after a determination that the end-tidal CO2 concentration decreases from the baseline concentration by the pre-defined amount;e) immediately deflating the infant's lungs to produce a maximum forced expiration; and wherein steps (a)–(e) are performed by an apparatus in an automated manner. 9. An apparatus for performing a maximum forced expiration maneuver in an infant comprising: a mask connected to a source of breathable air that delivers the breathable air at super-atmospheric pressure;a flow sensor in fluid communication with the mask;a CO2 sensor in fluid communication with the mask,an inflatable vest for compressing the infant's thoracic-abdominal region;a controller configured to receive a first signal from the flow sensor and a second signal from t... | Sensormedics Corporation | A61H031/00 A61B005/087 A61M016/00 A62B009/02 A62B018/02
US07060030 | B2 | 2006-06-13 | Two-hop telemetry interface for medical device | 2002-01-08 | This document discusses a system that includes an intermediary telemetry interface device for communicating between a cardiac rhythm management system or other implantable medical device and a programmer or other remote device. One example provides an inductive near-field communication link between the telemetry interface and the implantable medical device, and a radio-frequency (RF) far-field communication link between the telemetry interface device and the remote device. The telemetry interface device provides data buffering. In another example, the telemetry interface device includes a data processing module for compressing and/or decompressing data, or for extracting information from data. Such information extraction may include obtaining heart rate, interval, and/or depolarization morphology information from an electrogram signal received from the implantable medical device. | 1. An external telemetry interface device to assist in communication between an implantable medical device to be implanted in a patient the implantable medical device carrying an inductive element and an external remote first far-field radio-frequency (RF) transceiver, the telemetry interface device including: an inductive transceiver, configured to be communicatively coupled to the inductive element in the implantable medical device via a mutual inductance therebetween;a second far-field RF transceiver, configured to be communicatively coupled to the remote first far-field RF transceiver;a data buffer, coupled to the inductive and second far-field RF transceivers, to store data from the inductive and second far-field RF transceivers; anda controller, coupled to the inductive and second far-field RF transceivers and the data buffer, to issue control signals directing communication by the inductive and second far-field RF transceivers using data from the data buffer, the controller including a data nrocessing module to alter or remove a first header from data received from the implantable medical device and to provide a different second header to data to be provided to the remote first far-field transceiver. | 1. An external telemetry interface device to assist in communication between an implantable medical device to be implanted in a patient the implantable medical device carrying an inductive element and an external remote first far-field radio-frequency (RF) transceiver, the telemetry interface device including: an inductive transceiver, configured to be communicatively coupled to the inductive element in the implantable medical device via a mutual inductance therebetween;a second far-field RF transceiver, configured to be communicatively coupled to the remote first far-field RF transceiver;a data buffer, coupled to the inductive and second far-field RF transceivers, to store data from the inductive and second far-field RF transceivers; anda controller, coupled to the inductive and second far-field RF transceivers and the data buffer, to issue control signals directing communication by the inductive and second far-field RF transceivers using data from the data buffer, the controller including a data nrocessing module to alter or remove a first header from data received from the implantable medical device and to provide a different second header to data to be provided to the remote first far-field transceiver. 2. The device of claim 1, further including a power source. 3. The device of claim 2, in which the power source includes a battery. 4. The device of claim 3, in which the battery is rechargeable. 5. The device of claim 1, in which the controller includes a data processing module. 6. The device of claim 5, in which the data processing module is configured to execute instructions compressing data in the data buffer. 7. The device of claim 5, in which the data processing module is configured to execute instructions to obtain extracted information, from electrogram data received from the implantable device, for transmission byte second far-field RF transceiver to the remote first far-field RF transceiver. 8. The device of claim 7, in which the extracted information includes heart rate information. 9. The device of claim 7, in which the extracted information includes cardiac signal morphology information. 10. The device of claim 7, in which the extracted information includes times associated with intervals between heart depolarizations in the electrogram data. 11. The device of claim 7, in which the extracted information includes an alert condition relating to the patient's condition. 12. The device of claim 1, in which at least one of the inductive transceiver, the second far-field RF transceiver, and a third far-field RF transceiver is configured to be communicatively coupled to an external device associated with the patient. 13. An external telemetry interface device to assist in communication between an implantable medical device to be implanted in a patient and a remote far-field radio-frequency (RF) transceiver, the telemetry interface device including: a first transceiver, configured to be communicatively coupled to the implantab... | Cardiac Pacemakers, Inc. | A61B005/00
US06856827 | B2 | 2005-02-15 | Fluoroscopic tracking and visualization system | 2002-12-02 | A system for surgical imaging and display of tissue structures of a patient, including a display and an image processor for displaying such images in coordination with a tool image to facilitate manipulation of the tool during the surgical procedure. The system is configured for use with a fluoroscope such that at least one image in the display is derived from the fluoroscope at the time of surgery. A fixture is affixed to an imaging side of the fluoroscope for providing patterns of an of array markers that are imaged in each fluoroscope image. A tracking assembly having a plurality of tracking elements is operative to determine positions of said fixture and the patient. One of the tracking elements is secured against motion with respect to the fixture so that determining a position of the tracking element determines a position of all the markers in a single measurement. | 1. A system for surgical imaging and display of tissue structures of a patient, said system including a display and an image processor for displaying such images in coordination with a tool image to facilitate manipulation of the tool during the surgical procedure, the system being configured for use with a fluoroscope such that at least one image in the display is derived from the fluoroscope at the time of surgery, said system comprising: a fixture affixed to an imaging side of said fluoroscope for providing patterns of an array of markers that are imaged in each fluoroscope image; and a tracking assembly having a plurality of tracking elements operative to determine positions of said fixture and the patient, one of said plurality of tracking elements being secured against motion with respect to said fixture so that determining a position of said one of said plurality of tracking elements determines a position of all the markers in a single measurement. | 1. A system for surgical imaging and display of tissue structures of a patient, said system including a display and an image processor for displaying such images in coordination with a tool image to facilitate manipulation of the tool during the surgical procedure, the system being configured for use with a fluoroscope such that at least one image in the display is derived from the fluoroscope at the time of surgery, said system comprising: a fixture affixed to an imaging side of said fluoroscope for providing patterns of an array of markers that are imaged in each fluoroscope image; and a tracking assembly having a plurality of tracking elements operative to determine positions of said fixture and the patient, one of said plurality of tracking elements being secured against motion with respect to said fixture so that determining a position of said one of said plurality of tracking elements determines a position of all the markers in a single measurement. 2. The system of claim 1, wherein said array of markers are disposed in a plurality of planes having a defined separation therebetween. 3. The system of claim 2, further comprising: a marker determination unit operative on the fluoroscope image to provide an identification of a subset of said markers in the image; and a processor operative with said identification to calibrate the system, wherein said processor determines a camera model that includes at least one distortion of said fluoroscope image. 4. The system of claim 2, further comprising a correction unit, and wherein said tracking assembly automatically determines relative position data of array of markers and the tool, wherein said correction unit is responsive to a system calibration and said position data to determine an image transformation such that a representation of the tool is correctly positioned and oriented in the fluoroscope image. 5. The system of claim 1, wherein said fluoroscope includes a C-arm having a source and a detector, wherein a first tracking element is secured to said detector, wherein a second tracking element is secured to a patient being imaged, and wherein a third tracking element is secured the tool. 6. The system of claim 1, wherein individual tracking elements of said plurality of tracking elements are each secured to at least one of said fluoroscope, a patient being imaged, and said tool. 7. A system for use with a fluoroscope that includes a source and an image detector, said system comprising: a fixture including a three dimensional array of markers disposed over a region in a rigid body, said body being configured for disposition between said source and said detector such that said array of markers is automatically imaged as an image is acquired by said fluoroscope; a tracking assembly including a first tracking element associated with a tool, a second tracking element associated with a patient, a third tracking element associated with said image detector, and a fourth tracking element associa... | GE Medical Systems Global Technology Company, LLC | A61B005/05
US06878112 | B2 | 2005-04-12 | Virtual remote monitor, alert, diagnostics and programming for implantable medical device systems | 2002-10-28 | A plurality of co-operative and complementary software programs are implemented in a web-enabled high speed computer system to remotely monitor, manage and modify the operational and functional parameters of a plurality of implanted medical devices (IMDs). The system utilizes virtual electrophysiologist module (VEM), chronic monitoring module (CMM) and prescription program module (PPM) programs to effect specific therapeutic and diagnostic methods for managing the IMDs, remotely on a conditions and real-time basis. The modules enable remote and continuous monitoring, management and maintenance of the IMDs by identifying critical medical events, determining optimal clinical settings and upgrading performance parameters based on prescriptive data. The modules are implemented in a data center having high-speed computers operating in a web-enabled environment. The modules and the IMDs communicate through wireless communications system via a programmer or an interface medical unit (IMD). | 1. A method of providing programming information to at least one implantable medical device located within a patient, the method comprising: initiating an interface between the at least one implantable medical device and a data center positioned at a distal location relative to the implantable medical device; transmitting device information regarding the at least one implantable medical device and patient information regarding the patient from the implantable medical device to the data center; determining an optimum therapeutic setting for the at least one implantable medical device based upon the device information and the patient information; and remotely installing software from the data center in the at least one implantable medical device to provide the optimum therapeutic setting. | 1. A method of providing programming information to at least one implantable medical device located within a patient, the method comprising: initiating an interface between the at least one implantable medical device and a data center positioned at a distal location relative to the implantable medical device; transmitting device information regarding the at least one implantable medical device and patient information regarding the patient from the implantable medical device to the data center; determining an optimum therapeutic setting for the at least one implantable medical device based upon the device information and the patient information; and remotely installing software from the data center in the at least one implantable medical device to provide the optimum therapeutic setting. 2. The method of claim 1, further comprising the steps of: automatically transmitting device identification information identifying the at least one implantable medical device from the implantable medical to the data center; and authenticating a user prior to transmitting device information and patient information. 3. The method of claim 1, and further comprising steps of: performing a sequence of tests based upon the device information and the patient information; and performing a diagnostic routine based upon the device information and the patient information. 4. The method of claim 1, and further comprising the step of: suggesting an optimum therapeutic device setting to a component in the at least one implantable medical device. 5. The method of claim 1, further comprising the step of: verifying that the optimum therapeutic setting is implemented in the at least one implantable device for a real-time basis. 6. The method of claim 1, and further comprising steps of: distributing a report including the optimum therapeutic selling to a designated location; and storing patient session information in the data center. 7. The method of claim 1, and further comprising: implementing a software program which determines the optimum therapeutic selling of the at least one implantable medical device based upon the device information and patient information. 8. The method of claim 1, and further comprising the step of: transmitting a programming sequence to the at least one implantable medical device via a telemetry device and an RF head, the programming sequence providing the optimum therapeutic setting. 9. A method or providing programming information to at least one implantable medical device located within a patient, the method comprising: initiating an interface between the implantable medical device and an expert data center comprising a virtual elecrophysiologist module (VEM), a chronic monitor (CMM), and a prescriptive program module (PPM) positioned in a data center at a distal location relative to the implantable medical device; transmitting device information regarding the at least one implantable medical device from the implantable medical device... | Medtronic, Inc. | A61B005/00
US06950063 | B2 | 2005-09-27 | Intraluminal MRI probe | 2002-07-03 | The present invention relates to an intraluminal magnetic resonance imaging (MRI) probe which may be used for intraluminal MRI. The present invention is sized sufficiently small to be inserted into a patient intraluminally. The probe of the present invention comprises a substrate having a dielectric constant in the range of 1-1000, and first and second conducting layers on each side of the dielectric. The present invention is configured to result in a distributed capacitance. | 1. An intraluminal MRI probe comprising: a. a first rectangular dielectric sandwich comprising two opposing long sides, two opposing short sides, a dielectric substrate having a thickness in the range of 1-3000 microns, a dielectric constant in the range of 1-1000, a first surface having a surface area in the range of 0.5-10,000 square millimeters, and a second surface having a surface area in the range of 0.5-10,000 square millimeters, a first conducting layer deposited on said first surface, and a second conducting layer deposited on said second surface; a second rectangular dielectric sandwich positioned to face said first sandwich and comprising two opposing long sides, two opposing short sides, a dielectric substrate having a thickness in the range of 1-3000 microns, a dielectric constant in the range of 1-1,000, a first surface having a surface area in the range of 0.5-10,000 square millimeters, and a second surface having a surface area in the range of 0.5-10,000 square millimeters, a first conducting layer deposited on said first surface, a first conducting layer deposited on said first surface, and a second conducting layer deposited on said second surface; b. a third rectangular dielectric sandwich comprising two opposing long sides, two opposing short sides, a dielectric substrate having a thickness in the range of 1-3000 microns, a dielectric constant in the range of 1-1,000, a first surface having a surface area in the range of 0.5-10,000 square millimeters, and a second surface having a surface area in the range of 0.5-10,000 square millimeters, a first conducting layer deposited on said first surface, a first conducting layer deposited on said first surface, and a second conducting layer deposited on said second surface, said third rectangular dielectric sandwich positioned such that one of its long sides is contiguous with a long side of said first sandwich and the other of its long sides is contiguous with a long side of said second sandwich; and c. a fourth rectangular dielectric sandwich comprising two opposing long sides, two opposing short sides, a dielectric substrate having a thickness in the range of 1-3000 microns, a dielectric constant in the range of 1-1,000, a first surface having a surface area in the range of 0.5-10,000 square millimeters, and a second surface having a surface area in the range of 0.5-10,000 square millimeters, a first conducting layer deposited on said first surface, a first conducting layer deposited on said first surface, and a second conducting layer deposited on said second surface, said fourth rectangular dielectric sandwich positioned to face said third sandwich such that one of the fourth sandwich's long sides is contiguous with a long side of said first sandwich and the other of its long sides is contiguous with a long side of said second sandwich and further such that the four sandwiches define a rectangular solid internal volume and the first conducting layer of each sandwich faces said in... | 1. An intraluminal MRI probe comprising: a. a first rectangular dielectric sandwich comprising two opposing long sides, two opposing short sides, a dielectric substrate having a thickness in the range of 1-3000 microns, a dielectric constant in the range of 1-1000, a first surface having a surface area in the range of 0.5-10,000 square millimeters, and a second surface having a surface area in the range of 0.5-10,000 square millimeters, a first conducting layer deposited on said first surface, and a second conducting layer deposited on said second surface; a second rectangular dielectric sandwich positioned to face said first sandwich and comprising two opposing long sides, two opposing short sides, a dielectric substrate having a thickness in the range of 1-3000 microns, a dielectric constant in the range of 1-1,000, a first surface having a surface area in the range of 0.5-10,000 square millimeters, and a second surface having a surface area in the range of 0.5-10,000 square millimeters, a first conducting layer deposited on said first surface, a first conducting layer deposited on said first surface, and a second conducting layer deposited on said second surface; b. a third rectangular dielectric sandwich comprising two opposing long sides, two opposing short sides, a dielectric substrate having a thickness in the range of 1-3000 microns, a dielectric constant in the range of 1-1,000, a first surface having a surface area in the range of 0.5-10,000 square millimeters, and a second surface having a surface area in the range of 0.5-10,000 square millimeters, a first conducting layer deposited on said first surface, a first conducting layer deposited on said first surface, and a second conducting layer deposited on said second surface, said third rectangular dielectric sandwich positioned such that one of its long sides is contiguous with a long side of said first sandwich and the other of its long sides is contiguous with a long side of said second sandwich; and c. a fourth rectangular dielectric sandwich comprising two opposing long sides, two opposing short sides, a dielectric substrate having a thickness in the range of 1-3000 microns, a dielectric constant in the range of 1-1,000, a first surface having a surface area in the range of 0.5-10,000 square millimeters, and a second surface having a surface area in the range of 0.5-10,000 square millimeters, a first conducting layer deposited on said first surface, a first conducting layer deposited on said first surface, and a second conducting layer deposited on said second surface, said fourth rectangular dielectric sandwich positioned to face said third sandwich such that one of the fourth sandwich's long sides is contiguous with a long side of said first sandwich and the other of its long sides is contiguous with a long side of said second sandwich and further such that the four sandwiches define a rectangular solid internal volume and the first conducting layer of each sandwich faces said in... | The Board of Regents of the University of Texas System The University of Houston | G01R003/327 A61B005/05
US06885884 | B2 | 2005-04-26 | Method for simulating the use of a system option for a technical apparatus | 2002-05-15 | In a method for simulating the use of a system option for a technical apparatus, data from technical apparatuses that relate to the configuration and/or the usage of the apparatuses are collected, are stored in a central data bank, and relationships between collected, stored data and/or between collected, stored and additionally prescribable data are produced. The use of a system option for a technical apparatus is simulated based on the data placed in relationship to one another. | 1. A method for simulating a use of a system option for a technical apparatus, comprising the steps of: operating at least one technical apparatus and collecting data as a result of said operation relating to at least one of configuration and usage of said apparatus; storing the collected data in a central data bank as stored data; prescribing additional data and storing said additional data in said central data bank as additional stored data; identifying at least one relationship from relationships in the group consisting of relationships among said stored data and relationships between said stored data and said additionally stored data; and simulating operation of said at least one technical apparatus using a system option that is not physically present in said technical apparatus dependent on the data in said at least one relationship. | 1. A method for simulating a use of a system option for a technical apparatus, comprising the steps of: operating at least one technical apparatus and collecting data as a result of said operation relating to at least one of configuration and usage of said apparatus; storing the collected data in a central data bank as stored data; prescribing additional data and storing said additional data in said central data bank as additional stored data; identifying at least one relationship from relationships in the group consisting of relationships among said stored data and relationships between said stored data and said additionally stored data; and simulating operation of said at least one technical apparatus using a system option that is not physically present in said technical apparatus dependent on the data in said at least one relationship. 2. A method as claimed in claim 1 comprising filtering said collected data with regard to different data types, to obtain filtered data, and forwarding only said filtered data to said central data bank. 3. A method as claimed in claim 1 comprising automatically collecting said data. 4. A method as claimed in claim 1 comprising collecting said data via a communication network in communication with said technical apparatus. 5. A method as claimed in claim 1 wherein the step of collecting data during operation of a technical apparatus comprises collecting said data during operation of a medical technical apparatus. 6. A method as claimed in claim 5 comprising collecting said data from a medical technical apparatus selected from the group consisting of medical technical installations and medical technical devices. 7. A method as claimed in claim 5 comprising collecting said data from a medical technical apparatus selected from the group consisting of magnetic resonance apparatuses, computed tomography apparatuses, C-arm x-ray apparatuses, and ultrasound apparatuses. | Siemens Aktiengesellschaft | A61B005/05
US06949073 | B2 | 2005-09-27 | Dyspnea monitor, and telemedicine system and method | 2002-10-03 | Breathing interval measurement apparatus for measuring breathing of a subject, the apparatus comprising: a breathing interval beginning determinator, a breathing interval end determinator, and a timer associated with said breathing interval beginning determinator and said breathing interval end determinator to measure an interval between activation of said determinators, and a processor operable with said timer to assign levels of importance to said measured time interval. | 1. Breathing interval measurement apparatus for measuring breathing of a subject, the apparatus comprising a breathing interval beginning determinator for using sound detection to determine a start of breath sound, breathing interval end determinator for using sound detection to determine an end of breath sound, a timer associated with said breathing interval beginning determinator and said breathing interval end determinator to measure an interval between activation of said determinators, said interval being a breathing length between a start and an end of breathing, a processor operable with said timer to assign levels of importance to said measured time interval, storage for storing a previous measurement, and a comparator for comparing said interval with said previous measurement. | 1. Breathing interval measurement apparatus for measuring breathing of a subject, the apparatus comprising a breathing interval beginning determinator for using sound detection to determine a start of breath sound, breathing interval end determinator for using sound detection to determine an end of breath sound, a timer associated with said breathing interval beginning determinator and said breathing interval end determinator to measure an interval between activation of said determinators, said interval being a breathing length between a start and an end of breathing, a processor operable with said timer to assign levels of importance to said measured time interval, storage for storing a previous measurement, and a comparator for comparing said interval with said previous measurement. 2. Breathing interval measurement apparatus according to claim 1, further operable to base a plurality of recommendations on substantially an identical measurement set. 3. Breathing interval measuring apparatus according to claim 1, further comprising a communication unit for sending processed data of said breathing measuring to a remote location. 4. Breathing interval measuring apparatus according to claim 1, further comprising an instruction unit for providing interactive instructions to said subject. 5. Breathing interval measuring apparatus according to claim 4, wherein said instruction unit is remotely located. 6. Breathing interval measurement apparatus according to claim 1, wherein said breathing interval end determinator comprises a sound analyzer operable to identify intake of breath. 7. Breathing interval measurement apparatus according to claim 6, wherein said sound analyzer is operable to make said determination from sound envelope measurement. 8. Breathing interval measurement apparatus according to claim 6, wherein said sound analyzer is operable to make said determination from identification of sounds indicative of breathing end. 9. Breathing interval measurement apparatus according to claim 1, wherein said breathing interval beginning determinator comprises a sound analyzer operable to identify intake of breath. 10. Breathing interval measurement apparatus according to claim 9, wherein said sound analyzer is operable to make said determination from sound envelope measurement. 11. Breathing interval measurement apparatus according to claim 9, wherein said sound analyzer is operable to make said determination from identification of sounds indicative of breathing start. 12. Breathing interval measurement apparatus according to claim 6, wherein said breathing interval beginning determinator comprises a sound analyzer operable to identify intake of breath. 13. Breathing interval measurement apparatus according to claim 6, wherein said breathing interval end determinator is operable to determine a breathing interval end over a remote connection. 14. Breathing interval measurement apparatus according to claim 1, wherein said breathing interval... | Home-Medicine.com, Inc. | A61B005/08
US06926667 | B2 | 2005-08-09 | Patient compliance monitor | 2003-02-03 | A patient monitoring device includes a microprocessor controller having a clock circuit and memory coupled to one or more sensors physically carried by a medical appliance such that as the medical appliance is properly fitted to and worn by the patient, the sensors provide an electrical signal confirming that which may be then timed to provide data which confirms a patient's compliance with a recommended protocol. By combining and correlating the sensor data with the clock or timer provided as part of the controller, a time chart of data may be created indicating when and for how long the patient actually wears the device. Additionally, other sensors may be used to collect data relating to various operational parameters of the device including the amount of negative pressure presented under a vacuum dome or other parameters related to the operation of the medical appliance itself, as desired. | 1. A medical appliance having a patient monitoring device comprising at least a first sensor to sense a pressure, said sensed pressure being directly indicative of the patient's contemporaneous use of the medical appliance, said first sensor being connected to and supplying first sensor output data to a controller, said first sensor output data corresponding to said sensed pressure, the controller being configured to correlate said first sensor output data with time to thereby generate compliance data, and said controller including a download port for outputting said compliance data. | 1. A medical appliance having a patient monitoring device comprising at least a first sensor to sense a pressure, said sensed pressure being directly indicative of the patient's contemporaneous use of the medical appliance, said first sensor being connected to and supplying first sensor output data to a controller, said first sensor output data corresponding to said sensed pressure, the controller being configured to correlate said first sensor output data with time to thereby generate compliance data, and said controller including a download port for outputting said compliance data. 2. The medical appliance of claim 1 further comprising a timer associated with the controller for providing a time correlation to said first sensor output data. 3. The medical appliance of claim 2 wherein said medical appliance is worn by said patient, and wherein said medical appliance further comprises a second sensor, wherein said second sensor comprises a temperature sensor positioned to sense the patient's temperature as the medical appliance is worn to thereby provide another indication of contemporaneous patient use. 4. The medical appliance of claim 3 further comprising a memory associated with said controller for storing said compliance data prior to download through said download port. 5. The medical appliance of claim 4 wherein said electromechanical device comprises a pump for creating a negative pressure within a dome, and wherein said first sensor comprises a pressure sensor for sensing the negative pressure created by said pump. 6. The medical appliance of claim 5 further comprising a third sensor, said third sensor being adapted and positioned to sense another device parameter and being connected to said controller for transmitting third sensor output data thereto for correlation with time to thereby generate compliance data. 7. A medical device comprising: a clinical device adapted to be worn by a patient in order to administer a prescribed treatment protocol, a pressure sensor positioned to sense pressure data indicative of whether said patient is wearing said clinical device, and a controller in communication with said pressure sensor, said controller being configured to determine if and for how long a said patient wears said clinical device as prescribed by correlating said sensed pressure data with time. 8. The medical device of claim 7 wherein said clinical device comprises a soft tissue enlargement device having a tensioning applicator for applying a tension to a patient's soft tissue, and said medical device further comprising a temperature sensor positioned for being heated by said patient as the clinical device is worn to thereby provide an indication that said clinical device is being worn. 9. The medical device of claim 8 further comprising a timer for correlation with said sensed pressure data and temperature sensor data to thereby provide data corresponding to how long said patient is wearing the clinical device. 10. The medica... | Brava, LLC | A61B005/00
US07214197 | B2 | 2007-05-08 | Intraoperative neurophysiological monitoring system | 2001-11-06 | An intraoperative neurophysiological monitoring system includes an adaptive threshold detection circuit adapted for use in monitoring with a plurality of electrodes placed in muscles which are enervated by a selected nerve and muscles not enervated by the nerve. Nerve monitoring controller algorithms permit the rapid and reliable discrimination between non-repetitive electromyographic (EMG) events repetitive EMG events, thus allowing the surgeon to evaluate whether nerve fatigue is rendering the monitoring results less reliable and whether anesthesia is wearing off. The intraoperative monitoring system is designed as a “surgeon's monitor,” and does not require a neurophysiologist or technician to be in attendance during surgery. The advanced features of the intraoperative monitoring system will greatly assist neurophysiological research toward the general advancement of the field intraoperative EMG monitoring through post-surgical analysis. The intraoperative monitoring system is preferably modular, in order to allow for differential system pricing and upgrading as well as to allow for advances in computer technology; modularity can also aid in execution of the design. | 1. A method for intraoperative neurophysiological monitoring with at least one electrical stimulus probe as an intraoperative aid in defining the course of a nerve structure by monitoring electromyographic activity within a muscle enervated by the nerve structure, comprising: (a) contacting anatomical tissue with the stimulus probe in an area of the nerve structure;(b) detecting a change in impedance of the stimulus probe resulting from contact of the stimulus probe with the anatomical tissue;(c) triggering a sequence of pre-programmed intraoperative neurophysiological monitoring algorithm steps in response to the detection of the change in impedance of the stimulus probe. | 1. A method for intraoperative neurophysiological monitoring with at least one electrical stimulus probe as an intraoperative aid in defining the course of a nerve structure by monitoring electromyographic activity within a muscle enervated by the nerve structure, comprising: (a) contacting anatomical tissue with the stimulus probe in an area of the nerve structure;(b) detecting a change in impedance of the stimulus probe resulting from contact of the stimulus probe with the anatomical tissue;(c) triggering a sequence of pre-programmed intraoperative neurophysiological monitoring algorithm steps in response to the detection of the change in impedance of the stimulus probe. 2. The method for intraoperative neurophysiological monitoring of claim 1, wherein step (c) comprises closing a circuit between a current source and the stimulus probe to provide stimulus current to the nerve structure. 3. The method for intraoperative neurophysiological monitoring of claim 1, wherein step (c) comprises: (c)(1) generating a visible or audible indication of appropriate impedance of the stimulus probe resulting from contact of the stimulus probe with the anatomical tissue. 4. The method for intraoperative neurophysiological monitoring of claim 1, wherein step (c) comprises: (c)(1) initiating generation of a pre-programmed sequence of stimulus pulses to be supplied to the nerve structure via the stimulus probe; and(c)(2) storing measured responses to the stimulus pulses collected from an electrode connected to the enervated muscle. 5. The method for intraoperative neurophysiological monitoring of claim 4, further comprising: (c)(3) analyzing the stored, measured responses collected from an electrode connected to the enervated muscle, to determine the average response amplitude. 6. The method for intraoperative neurophysiological monitoring of claim 4, further comprising: (c)(3) analyzing the stored, measured responses collected from an electrode connected to the enervated muscle, to determine the peak-to-peak response amplitude. 7. The method for intraoperative neurophysiological monitoring of claim 4, further comprising: (c)(3) analyzing the stored, measured responses collected from an electrode connected to the enervated muscle, to determine the response threshold. 8. The method for intraoperative neurophysiological monitoring of claim 4, further comprising: (c)(3) analyzing the stored, measured responses collected from an electrode connected to the enervated muscle, to determine the response amplitude as a function of stimulus intensity. 9. A method for intraoperative neurophysiological monitoring with at least one electrical stimulus probe as an intraoperative aid in defining the course of a nerve structure by monitoring electromyographic activity within the nerve structure, comprising: (a) contacting the nerve structure with the stimulus probe;(b) detecting a stimulus probe impedance change resulting from said stimulus probe nerve contact;(c) triggeri... | - | A61B005/05 A61B005/04 A61N001/00
US07199100 | B2 | 2007-04-03 | Cyclic peptides, method for preparing and use as angiogenesis inhibitors or activator | 2001-10-02 | The invention relates to cyclopeptides, the method of their preparation and their utilization as inhibitors or activators of angiogenesis.;These cyclopeptides comprise the following peptide sequence: -Arg-Ile-Lys-Pro-His-Gln-Gly-;They can be used in systems for inhibition of angiogenesis that comprises a support (1), to which the cyclopeptide is affixed by means of an organic spacer arm (3) that may be provided with a moiety (4) capable of being spliced by an enzyme system. | 1. A cyclopeptide chosen from among the following compounds:SEQ ID NO:5P11:cyclo(DPhe-Pro-Gln-Ile-Met-Arg-Ile-Lys-Pro-His-Gln-Gly-Gln-His-Ile-Gly-Glu)SEQ ID NO:8P16:cyclo(Arg-Ile-Lys-Pro-His-Gln-Gly)SEQ ID NO:9P17:cyclo(Pro-Arg-Ile-Lys-Pro-His-Gln-Gly)SEQ ID NO:10P19:cyclo(Gln-Ile-Met-Arg-Ile-Lys-Pro-His-Gln-Gly-Gln-His-Ile-Gly-Glu)SEQ ID NO:11P20:cyclo(DPhe-Pro-Gln-Ile-Met-Arg-Ile-Lys-Pro-His-Gln-Gly-Gln-His-Ile-Gly)SEQ ID NO:13P23:cyclo(DPhe-Pro-Arg-Ile-Lys-Pro-His-Gln)SEQ ID NO:25P24:cyclo(Gly-Arg-Ile-Lys-Pro-His)and the compounds P11 and P20, wherein DPhe is re-placed by DTyr. | 1. A cyclopeptide chosen from among the following compounds:SEQ ID NO:5P11:cyclo(DPhe-Pro-Gln-Ile-Met-Arg-Ile-Lys-Pro-His-Gln-Gly-Gln-His-Ile-Gly-Glu)SEQ ID NO:8P16:cyclo(Arg-Ile-Lys-Pro-His-Gln-Gly)SEQ ID NO:9P17:cyclo(Pro-Arg-Ile-Lys-Pro-His-Gln-Gly)SEQ ID NO:10P19:cyclo(Gln-Ile-Met-Arg-Ile-Lys-Pro-His-Gln-Gly-Gln-His-Ile-Gly-Glu)SEQ ID NO:11P20:cyclo(DPhe-Pro-Gln-Ile-Met-Arg-Ile-Lys-Pro-His-Gln-Gly-Gln-His-Ile-Gly)SEQ ID NO:13P23:cyclo(DPhe-Pro-Arg-Ile-Lys-Pro-His-Gln)SEQ ID NO:25P24:cyclo(Gly-Arg-Ile-Lys-Pro-His)and the compounds P11 and P20, wherein DPhe is re-placed by DTyr. 2. A pharmaceutical composition for inhibiting angiogenesis comprising a cyclopeptide chosen among the following cyclopeptides:P11: cyclo(DPhe-Pro-Gln-Ile-Met-Arg-SEQ ID NO:5Ile-Lys-Pro-His-Gln-Gly-Gln-His-Ile-Gly-Glu)P16: cyclo(Arg-Ile-Lys-Pro-His-Gln-SEQ ID NO:8Gly)P17: cyclo(Pro-Arg-Ile-Lys-Pro-His-SEQ ID NO:9Gln-Gly)P19: cyclo(Gln-Ile-Met-Arg-Ile-Lys-SEQ ID NO:10Pro-His-Gln-Gly-Gln-His-Ile-Gly-Glu)P20: cyclo(DPhe-Pro-Gln-Ile-Met-Arg-SEQ ID NO:11Ile-Lys-Pro-His-Gln-Gly-Gln-His-Ile-Gly)P23: cyclo(DPhe-Pro-Arg-Ile-Lys-Pro-SEQ ID NO:13His-Gln)P24: cyclo(Gly-Arg-Ile-Lys-Pro-His)SEQ ID NO:25and the cyclopeptides P11 and P20, wherein DPhe is replaced by DTyr. 3. A pharmaceutical composition for activating angiogenesis comprising two identical or different cyclopeptides, coupled with a pharmaceutically acceptable organic compound, the cyclopeptides being chosen from among the following cyclopeptides:P11:cyclo(DPhe-Pro-Gln-Ile-Met-Arg-Ile-Lys-Pro-His-Gln-Gly-Gln-His-Ile-Gly-Glu)SEQ ID NO: 5P16:cyclo(Arg-Ile-Lys-Pro-His-Gln-Gly)SEQ ID NO: 8P17:cyclo(Pro-Arg-Ile-Lys-Pro-His-Gln-Gly)SEQ ID NO: 9P19:cyclo(Gln-Ile-Met-Arg-Ile-Lys-Pro-His-Gln-Gly-Gln-His-Ile-Gly-Glu)SEQ ID NO: 10P20:cyclo(DPhe-Pro-Gln-Ile-Met-Arg-Ile-Lys-Pro-His-Gln-Gly-Gln-His-Ile-Gly)SEQ ID NO: 11P23:cyclo(DPhe-Pro-Arg-Ile-Lys-Pro-His-Gln)SEQ ID NO: 13P24:cyclo(Gly-Arg-Ile-Lys-Pro-His)SEQ ID NO: 25and the cyclopeptides P11 and P20, wherein DPhe is replaced by DTyr. 4. The system comprising a cyclopeptide according to claim 1, covalently bound to an organic spacer arm. 5. The system according to claim 4, wherein the organic spacer arm is covalently bound to a support. 6. A system comprising two cyclopeptides according to 1 covalently bound to an organic compound. 7. The system according to claim 6, wherein the distance between the two cyclopeptides is such that, when the system is brought into contact with cells expressing vascular endothelium growth factor (VEGF) receptors, it allows dimerization of said receptors. 8. The system according to claim 6, wherein the organic compound is covalently bound to a support with the aid of an organic spacer arm. 9. The system according to claim 7, wherein the organic compound is covalently bound to a support with the aid of an organic spacer arm. 10. A system comprising a solid support upon where upon cyclopeptides according to claim 1, are attached by covalent bond... | Commissariat a l'Energie Atomique Universite Victor Segalen Bordeaux 2 Universite de Bordeaux I | A61K038/00 A61B005/055
US07250056 | B2 | 2007-07-31 | Lancet-integrated mounter and method of making the same | 2002-06-10 | A mounter (3) integrally provided with a lancing element (47B) and analyzer includes a lancing member (47) provided with the lancing element (47B), a first member (4) including an accommodation portion (40) for holding the lancing member (47), and a second member (5) formed separately from the first member (4) for holding the analyzer (6). The accommodation portion (40) includes an accommodation space (43) communicable with the outside through a first opening (43a) and a second opening (43b). The lancing member (47) is held and hermetically sealed in the accommodation space (43) by closing the first opening (43a) and the second opening (43b) with a first seal member (44a) and a second seal member (44b). | 1. A lancet-integrated mounter comprising: a lancing member including a lancing element and being forwardly movable for lancing, a first member including an accommodation portion for holding the lancing member, an analyzer to which body fluid is supplied, and a second member spaced from the accommodation portion in a forward direction of the lancing member and formed separately from the first member for holding the analyzer; the accommodation portion including an accommodation space communicatable with an outside through a first opening and a second opening;the lancing member being held and hermetically sealed in the accommodation space by closing the first opening and the second opening with a first seal member and a second seal member, the analyzer being held by the second member outside the sealed accommodation space. | 1. A lancet-integrated mounter comprising: a lancing member including a lancing element and being forwardly movable for lancing, a first member including an accommodation portion for holding the lancing member, an analyzer to which body fluid is supplied, and a second member spaced from the accommodation portion in a forward direction of the lancing member and formed separately from the first member for holding the analyzer; the accommodation portion including an accommodation space communicatable with an outside through a first opening and a second opening;the lancing member being held and hermetically sealed in the accommodation space by closing the first opening and the second opening with a first seal member and a second seal member, the analyzer being held by the second member outside the sealed accommodation space. 2. The lancet-integrated mounter according to claim 1, wherein the lancing member moves from an idling position to a lancing position upon receiving an external force to cause the lancing element to project from the accommodation portion. 3. The lancet-integrated mounter according to claim 2, wherein the first seal member is adapted to be formed with a hole in exerting the external force to the lancing member, and the lancing element forms a hole in the second seal member during the movement of the lancing member from the idling position to the lancing position. 4. The lancet-integrated mounter according to claim 2, wherein the accommodation portion is provided with an engagement portion for holding the lancing member at the idling position. 5. The lancet-integrated mounter according to claim 4, wherein the accommodation portion is provided with a stopper for controlling movement of the lancing member at the lancing position. 6. The lancet-integrated mounter according to claim 5, wherein the lancing member at the lancing position is secured between the engagement portion and the stopper. 7. The lancet-integrated mounter according to claim 1, wherein the first member further includes a holder portion for holding the second member, and a support portion for supporting the accommodation portion at a higher position than the holder portion. 8. The lancet-integrated mounter according to claim 7, wherein the accommodation portion, the holder portion and the support portion are integrally formed as one piece. 9. The lancet-integrated mounter according to claim 7, wherein the accommodation portion is formed separately from the holder portion and the support portion, the accommodation portion being removable from the holder portion and the support portion. 10. The lancet-integrated mounter according to claim 7, wherein the support portion is elastic, and the support portion flexes to move the accommodation portion downward when an external force in the downward direction is exerted to the accommodation portion and elastically restores the accommodation portion to an original position when the external force is removed. 11. The l... | Arkray, Inc. | A61B017/32 A61B005/151
US07081096 | B2 | 2006-07-25 | Temperature mapping balloon | 2003-01-24 | The temperature mapping balloon of the present invention provides a device for locating inflammation and increased metabolic activity associated with conditions such as vulnerable plaque, by mapping temperature of a body lumen, such as an artery or blood vessel. The temperature mapping balloon comprises a balloon with a thermal mapping coating disposed on the inside or outside of the balloon. The thermal mapping coating can be a thermochromic coating that changes color in response to temperature or a sensor coating comprising a plurality of temperature sensors. | 1. A temperature mapping balloon for use in a vessel having a vessel wall comprising: a balloon, the balloon being conformable over the vessel wall; anda thermal mapping coating, the thermal mapping coating disposed on the balloon and comprising a sensor coating, the sensor coating comprising a plurality of temperature sensors. | 1. A temperature mapping balloon for use in a vessel having a vessel wall comprising: a balloon, the balloon being conformable over the vessel wall; anda thermal mapping coating, the thermal mapping coating disposed on the balloon and comprising a sensor coating, the sensor coating comprising a plurality of temperature sensors. 2. The temperature mapping balloon of claim 1 wherein the thermal mapping coating is disposed in a location selected from the group consisting of outside the balloon and inside the balloon. 3. The temperature mapping balloon of claim 1 further comprising a position indicator disposed on the balloon. 4. The temperature mapping balloon of claim 3 wherein the position indicator is selected from the group consisting of radiopaque markers, radio-frequency (RF) coils, and ultrasonic marker transponders. 5. The temperature mapping balloon of claim 1 further comprising an insulating layer disposed on the sensor coating. 6. The temperature mapping balloon of claim 5 wherein the insulating layer is a material selected from the group consisting of parylene, silicone rubber, polytetrafluoroethylene, polyurethane, and plastic. 7. The temperature mapping balloon of claim 1 wherein the sensor coating is made of materials selected from the group consisting of nickel, copper, platinum, chromium, iron, rhodium, aluminum, conducting graphite polymers, and alloys and combinations thereof. 8. The temperature mapping balloon of claim 1 wherein the temperature sensors are selected from the group consisting of thermocouples, thermisters, resistance thermometers, resistive temperature detectors (RTDs), PN junctions, and thermal piles. 9. The temperature mapping balloon of claim 1 wherein each of the plurality of temperature sensors provides a temperature signal, and the temperature signals from the plurality of temperature sensors are multiplexed. 10. A method of using a temperature mapping balloon with a patient comprising: providing a thermochromic temperature mapping balloon, the thermochromic temperature mapping balloon comprising a plurality of arrays, each of the plurality of arrays comprises a plurality of bistable dots, the bistable dots change from one color to another color in response to a threshold temperature, and the threshold temperatures for the bistable dots in each of the plurality of arrays are distributed over a temperature range;inserting the thermochromic temperature mapping balloon into the patient;inflating the thermochromic temperature mapping balloon;allowing thermochromic temperature mapping balloon temperature to stabilize;deflating the thermochromic temperature mapping balloon;removing the thermochromic temperature mapping balloon from the patient; andreading a vessel temperature map from the temperature mapping balloon. 11. The method of claim 10 wherein reading a vessel temperature map comprises reading the vessel temperature map visually. 12. The method of claim 10 wherein reading a vessel temperature map... | Medtronic Vascular, Inc. | A61B005/00
US07239910 | B2 | 2007-07-03 | Methods and devices for transcranial magnetic stimulation and cortical cartography | 2002-04-29 | The invention relates to: a method for stimulating and/or inhibiting at least one point or area of a brain using at least one stimulation device, wherein: the spatial structure of the head or brain is recorded; a three-dimensional simulation model of the surface of the brain is generated from the recording of the spatial structure of the brain; and the stimulation device is arranged relative to the head or brain using the three-dimensional simulation model of the surface of the brain, such that the at least one point or area of the brain can be stimulated using the stimulation device; a device for stimulating and/or inhibiting at least one point or area of a brain, comprising a recording device for detecting the spatial structure of the brain, a computational device for generating a simulation model of the surface of the brain and at least one stimulation or induction device, in particular a coil; a method for determining the function of a particular area of the brain, wherein at least one particular area is stimulated using a stimulation device and the stimulus response is measured at least two different positions, and a device for determining the function of a particular area of the brain, comprising at least one stimulation device and at least two stimulus detection devices. | 1. A method for stimulating and/or inhibiting at least one point or area of a brain using at least one stimulation device, comprising: a) recording a spatial structure of the head or brain;b) generating a three-dimensional simulation model of the surface of the brain from the recording of the spatial structure of the brain; andc) arranging the stimulation device relative to the head or brain using the three-dimensional simulation model of the surface of the brain, whereby the stimulation device is accurately positioned relative to the surface of the brain, wherein the stimulation device is positioned such that a magnetic field generated or generable by the stimulation device is normal to the simulation model of the surface of the brain on or over the area to be stimulated or exhibits a predetermined spatial angle relative to the simulation model of the surface of the brain. | 1. A method for stimulating and/or inhibiting at least one point or area of a brain using at least one stimulation device, comprising: a) recording a spatial structure of the head or brain;b) generating a three-dimensional simulation model of the surface of the brain from the recording of the spatial structure of the brain; andc) arranging the stimulation device relative to the head or brain using the three-dimensional simulation model of the surface of the brain, whereby the stimulation device is accurately positioned relative to the surface of the brain, wherein the stimulation device is positioned such that a magnetic field generated or generable by the stimulation device is normal to the simulation model of the surface of the brain on or over the area to be stimulated or exhibits a predetermined spatial angle relative to the simulation model of the surface of the brain. 2. The method as set forth in claim 1, wherein the spatial structure of the head or brain is recorded using a nuclear spin resonance, computer tomography, x-ray or ultrasound method. 3. The method as set forth in claim 1, wherein the simulation model of the surface of the brain is formed by a plurality of polygons. 4. The method as set forth in claim 1, wherein one or more coils are used as the stimulation device. 5. The method as set forth in claim 1, wherein the stimulation device is navigated to a calculated position for the desired stimulation. 6. The method as set forth in claim 1, wherein a stimulation signal is only released or generated when the deviation of a calculated position of the stimulation device from the actual position of the stimulation device is smaller than a predetermined value or keeps within a predetermined angular range. 7. The method as set forth in claim 1, wherein positional data for the stimulation device and/or the brain or head and/or the electrical field which may be induced on the surface of the brain are displayed. 8. The method as set forth in claim 1, wherein a simulation model of the stimulation device and/or the head and/or the brain is used. 9. The method as set forth in claim 8, wherein the simulation model is a multi-shell model. 10. A computer program embodied on a computer readable medium for stimulating and/or inhibiting at least one point or area of a brain using at least one stimulation device, comprising: a) code that records a spatial structure of the head or brain;b) code that generates a three-dimensional simulation model of the surface of the brain from the recording of the spatial structure of the brain; andc) code that arranges the stimulation device relative to the head or brain using the three-dimensional simulation model of the surface of the brain, whereby the stimulation device is accurately positioned relative to the surface of the brain, wherein the stimulation device is positioned such that a magnetic field generated or generable by the stimulation device is normal to the simulation model of the surface of ... | BrainLAB AG | A61B005/0484
US06904301 | B2 | 2005-06-07 | Micro-invasive method for painless detection of analytes in extracellular space | 2002-12-31 | A method of detecting at least one analyte in extra-cellular spaces includes the step of inserting a microprobe through the stratum corneum toward the stratum basale of the skin of a subject into extra-cellular spaces containing interstitial fluid having at least one analyte to be detected, said microprobe having a diameter at its tip no larger than approximately 10-50 microns. The method further includes optically testing for a predetermined analyte in the extra-cellular space adjacent the distal end of the microprobe without drawing a sample of the interstitial fluid. Preferably the microprobe body includes a sensor layer covering the distal optical tip of the microprobe body, the sensor layer being adapted to interact with a predetermined analyte to be detected in the interstitial fluid, and an optical detector responsive to interaction of the sensor layer with the predetermined analyte to signal detection of said predetermined analyte. | 1. A method of detecting at least one analyte in extra-cellular spaces comprising: inserting a non-porous distal end of a microprobe through the stratum corneum toward but not through the stratum basale of the skin of a subject into extra-cellular spaces containing interstitial fluid having at least one analyte to be detected; optically testing for a predetermined analyte in the extra-cellular space adjacent the distal end of the microprobe without drawing a sample of the interstitial fluid. | 1. A method of detecting at least one analyte in extra-cellular spaces comprising: inserting a non-porous distal end of a microprobe through the stratum corneum toward but not through the stratum basale of the skin of a subject into extra-cellular spaces containing interstitial fluid having at least one analyte to be detected; optically testing for a predetermined analyte in the extra-cellular space adjacent the distal end of the microprobe without drawing a sample of the interstitial fluid. 2. The method as set forth in claim 1 wherein said microprobe has a diameter no larger than approximately 10-50 microns. 3. The method as set forth in claim 1 wherein the microprobe is held in place with the distal end thereof in the interstitial fluid for a predetermined time, further including the step of monitoring the concentration of the analyte a plurality of times while the microprobe is held in place. 4. The method as set forth in claim 3 wherein the concentration of the analyte is continuously monitored during the time the microprobe is held in place with the distal end thereof in the interstitial fluid. | Becton. Dickinson and Company | A61B005/00
US07260429 | B2 | 2007-08-21 | Method and apparatus for phonocardiographic image acquisition and presentation | 2002-12-02 | A cardiac rhythm management system provides a phonocardiographic image indicative of a heart's mechanical events related to hemodynamic performance. The phonocardiographic image includes a stack of acoustic sensor signal segments representing multiple cardiac cycles. Each acoustic sensor signal segment includes heart sounds indicative of the heart's mechanical events and representations of the heart's electrical events. The stack of acoustic sensor signal segments are aligned by a selected type of the heart's mechanical or electrical events and are grouped by a cardiac timing parameter for presentation. | 1. A system, comprising: a signal input to receive an acoustic sensor signal indicative of heart sounds;a marker input to receive event markers indicative of cardiac events, the event markers temporally associated with the acoustic sensor signal;a signal segmenting module coupled to the signal input and the marker input, the signal segmenting module adapted to segment the acoustic sensor signal into acoustic sensor signal segments based on a selected first type of the event markers representing a first type of the cardiac events; anda signal alignment module coupled to the signal segmenting module, the signal alignment module adapted to temporally align the acoustic sensor signal segments based on the first type of the event markers. | 1. A system, comprising: a signal input to receive an acoustic sensor signal indicative of heart sounds;a marker input to receive event markers indicative of cardiac events, the event markers temporally associated with the acoustic sensor signal;a signal segmenting module coupled to the signal input and the marker input, the signal segmenting module adapted to segment the acoustic sensor signal into acoustic sensor signal segments based on a selected first type of the event markers representing a first type of the cardiac events; anda signal alignment module coupled to the signal segmenting module, the signal alignment module adapted to temporally align the acoustic sensor signal segments based on the first type of the event markers. 2. The system of claim 1, further comprising an accelerometer, coupled to the sensor signal input, to sense an acceleration signal indicative of the heart sounds. 3. The system of claim 1, further comprising a microphone to sense the heart sounds. 4. The system of claim 1, further comprising a cardiac sensing circuit, coupled to the marker input, to sense a cardiac signal, detect the cardiac events from the cardiac signal, and generate the event markers to identify the cardiac events. 5. The system of claim 4, further comprising an intracardiac electrode system, coupled to the cardiac sensing circuit, the intracardiac electrode system including at least one electrode adapted to sense an intracardiac electrogram. 6. The system of claim 4, further comprising a surface electrode system coupled to the cardiac sensing circuit, the surface electrode system including two or more electrodes adapted to sense a surface electrocardiogram. 7. The system of claim 1, further comprising a signal grouping module coupled to the signal alignment module, the signal grouping module adapted to sort and group the acoustic sensor signal segments according to a grouping instruction. 8. The system of claim 7, wherein the signal grouping module comprises a signal segment averaging module adapted to average selected segments of the acoustic sensor signal segments according to the grouping instruction. 9. The system of claim 7, further comprising a video presentation module coupled to the signal grouping module, the video presentation module adapted to convert the acoustic sensor signal segments into image pixels. 10. The system of claim 9, wherein the video presentation module is adapted to average acoustic sensor signal segments associated with a predetermined number of cardiac cycles and to convert the averaged acoustic sensor signal segments into the image pixels. 11. The system of claim 9, further comprising a display, coupled to the signal grouping module, to present a phonocardiographic image including at least the aligned and grouped acoustic sensor signal segments and selected representations of the event markers. 12. The system of claim 11, further comprising a mapping module, coupled to the display, to map amplitudes of th... | Cardiac Pacemakers, Inc. | A61B005/04
US07062315 | B2 | 2006-06-13 | Automated template generation algorithm for implantable device | 2002-04-25 | A method of generating a template in an implantable medical device for implantation within a patient, and a processor readable medium for performing the method, that includes generating a template from collected events corresponding to the patient, delaying the generation of the template for a first predetermined time period in response to the template not being generated within a predetermined number of collected events, determining whether the template is valid, and monitoring the template to determine whether the template is an accurate representation of the patient. | 1. A method of generating a template in an implantable medical device for implantation within a patient, comprising the steps of: generating a template from collected events corresponding to the patient;delaying the step of generating a template for a first predetermined time period in response to the template not being generated within a predetermined number of collected events;determining whether the template is valid; andmonitoring the template to determine whether the template is an accurate representation of the patient, wherein the step of generating a template comprises the steps of:monitoring a heart rate of the patient to generate the collected events;determining whether beats corresponding to the collected events are normal beats; anddetermining whether a predetermined number of normal beats has been collected within the predetermined number of collected events, wherein the step of delaying the template generation includes delaying the template generation in response to the predetermined number of normal beats not being collected within the predetermined number of collected events, and wherein the step of generating a template further comprises:determining whether the predetermined number of normal beats have been collected and computing cross matches between the predetermined number of collected normal beats to form corresponding computed cross matches;determining whether a predetermined number of the computed cross matches exceed a threshold;determining whether a predetermined number of cross matching attempts have failed;delaying the template generation for a second predetermined time period in response to the predetermined number of failed cross matching attempts; andforming the template from the predetermined number of computed cross matches in response to the predetermined number of the computed cross matches exceeding the threshold. | 1. A method of generating a template in an implantable medical device for implantation within a patient, comprising the steps of: generating a template from collected events corresponding to the patient;delaying the step of generating a template for a first predetermined time period in response to the template not being generated within a predetermined number of collected events;determining whether the template is valid; andmonitoring the template to determine whether the template is an accurate representation of the patient, wherein the step of generating a template comprises the steps of:monitoring a heart rate of the patient to generate the collected events;determining whether beats corresponding to the collected events are normal beats; anddetermining whether a predetermined number of normal beats has been collected within the predetermined number of collected events, wherein the step of delaying the template generation includes delaying the template generation in response to the predetermined number of normal beats not being collected within the predetermined number of collected events, and wherein the step of generating a template further comprises:determining whether the predetermined number of normal beats have been collected and computing cross matches between the predetermined number of collected normal beats to form corresponding computed cross matches;determining whether a predetermined number of the computed cross matches exceed a threshold;determining whether a predetermined number of cross matching attempts have failed;delaying the template generation for a second predetermined time period in response to the predetermined number of failed cross matching attempts; andforming the template from the predetermined number of computed cross matches in response to the predetermined number of the computed cross matches exceeding the threshold. 2. The method of claim 1, wherein the predetermined number of cross matches is four and the second predetermined time period is approximately equal to 156 minutes. 3. The method of claim 1, wherein the predetermined number of beats is six and the predetermined number of computed cross matches is four. 4. The method of claim 1, wherein the step of monitoring the template comprises the steps of: (a) collecting subsequent normal beats within a second predetermined time period;(b) computing a match between a subsequently collected beat and the template;(c) determining whether the match is within a predetermined threshold to form matched beats and other than matched beats;(d) determining whether x out of the last y subsequently collected beats are other than matched beats; and(e) repeating steps (a)–(d) in response to x out of the last y beats not being other than matched beats. 5. The method of claim 4, wherein step (a) comprises collecting one normal beat every 1000 seconds. 6. The method of claim 4, wherein the step of computing a match comprises computing the match for one out of every one thousan... | Medtronic, Inc. | A61B005/04
US07146218 | B2 | 2006-12-05 | Adaptive method and apparatus for forecasting and controlling neurological disturbances under a multi-level control | 2003-02-04 | A method and apparatus for forecasting and controlling neurological abnormalities in humans such as seizures or other brain disturbances. The system is based on a multi-level control strategy. Using as inputs one or more types of physiological measures such as brain electrical, chemical or magnetic activity, heart rate, pupil dilation, eye movement, temperature, chemical concentration of certain substances, a feature set is selected off-line from a pre-programmed feature library contained in a high level controller within a supervisory control architecture. This high level controller stores the feature library within a notebook or external PC. The supervisory control also contains a knowledge base that is continuously updated at discrete steps with the feedback information coming from an implantable device where the selected feature set (feature vector) is implemented. This high level controller also establishes the initial system settings (off-line) and subsequent settings (on-line) or tunings through an outer control loop by an intelligent procedure that incorporates knowledge as it arises. The subsequent adaptive settings for the system are determined in conjunction with a low-level controller that resides within the implantable device. The device has the capabilities of forecasting brain disturbances, controlling the disturbances, or both. Forecasting is achieved by indicating the probability of an oncoming seizure within one or more time frames, which is accomplished through an inner-loop control law and a feedback necessary to prevent or control the neurological event by either electrical, chemical, cognitive, sensory, and/or magnetic stimulation. | 1. A system for predicting and controlling the electrographic and clinical onset of a seizure and other neurological disturbances in an individual, comprising: a data generation component to acquire physiological signals from the individual;an intelligent data processing unit to preprocess the physiological signals, to extract and select a plurality of features, and to provide an estimation of the probability of seizure for at least one specific time frame; anda low level controller connected to the intelligent data processing unit to automatically activate a preventive therapeutic intervention measure to control the onset of a seizure in the individual in response to the probability of seizure for at least one specific time frame exceeding a threshold. | 1. A system for predicting and controlling the electrographic and clinical onset of a seizure and other neurological disturbances in an individual, comprising: a data generation component to acquire physiological signals from the individual;an intelligent data processing unit to preprocess the physiological signals, to extract and select a plurality of features, and to provide an estimation of the probability of seizure for at least one specific time frame; anda low level controller connected to the intelligent data processing unit to automatically activate a preventive therapeutic intervention measure to control the onset of a seizure in the individual in response to the probability of seizure for at least one specific time frame exceeding a threshold. 2. The system for predicting and controlling the electrographic onset of a seizure of claim 1 further comprising at least one sensor for detecting physiological signals that indicate the state of activity in the brain of the individual. 3. The system for predicting and controlling the electrographic onset of a seizure of claim 2 wherein the sensor is at least one of an implanted intracranial electrode, an epidural electrode, a scalp electrode, a sphenoidal electrode, a foramen ovale electrode, an intravascular electrode, a chemical sensor, a pupil dilation sensing device, an eye movement sensor, a heart rate sensor, and a body temperature sensor. 4. The system for predicting and controlling the electrographic onset of a seizure of claim 1 further comprising a high level controller that communicates with the intelligent data processing unit to retune at least one parameter that is used to extract and select a feature. 5. The system for predicting and controlling the electrographic onset of a seizure of claim 1 further comprising an external portable module including an external communications unit that enables the transfer of physiological data that is sensed in the individual to the external portable module for analysis and storage. 6. The system for predicting and controlling the electrographic onset of a seizure of claim 5 wherein the external portable module further comprises a display device that shows the probability output from the intelligent data processing unit for having a seizure in at least one time frame. 7. The system for predicting and controlling the electrographic onset of a seizure of claim 5 wherein the external portable module further comprises an alarm device which is activated to alert the individual of an oncoming seizure when the probability of having a seizure in at least one time frame exceeds an adaptive threshold. 8. The system for predicting and controlling the electro graphic onset of a seizure of claim 5 wherein the external portable module further comprises a battery recharger. 9. The system for predicting and controlling the electrographic onset of a seizure of claim 5 wherein the external portable module further comprises at least one of a microprocessor, ... | The Trustees of the University of Pennsylvania | A61N001/18 A61B005/02
US07194298 | B2 | 2007-03-20 | Method and apparatus for trend detection in an electrocardiogram monitoring signal | 2002-10-02 | The present invention can be used in conjunction with a wearable digital wireless ECG monitoring system. A full ECG curve is received by a central module. The central module is worn on the belt like a cellular or a pager. It is made of four different devices operating together: a handheld computer, a GPS, a cellular board and a multiplexing device. The system wirelessly receives the complete cardiac curve from the ECG and is able to distinguish not only the beat rate, but also to analyze any abnormal heart contractions. In fact, most common heart diseases are not related to the acceleration or deceleration of the heart rate. In case of problem detected by the central unit, the system automatically calls a central station and can send the GPS positioning and ECG monitoring of the patient with the detected anomaly data. Voice communication with the patient or the passers by is also possible. | 1. A trend detection method for detecting trends in an electrocardiogram monitoring signal of a patient, comprising: extracting locations of base points from said electrocardiogram monitoring signal;for each said location of base points, the location of the base point is compared against a predetermined threshold range based on a slow adaptative learning process using said patient's own cardiac movement, said range having an upper and lower threshold value for both an horizontal and vertical adjusted value for each of said base points;if all locations are within the predetermined, threshold range, performing a proximity analysis to determine improbable locations;if all locations are possible, determining an occurrence of a trend in said electrocardiogram monitoring signal from said locations. | 1. A trend detection method for detecting trends in an electrocardiogram monitoring signal of a patient, comprising: extracting locations of base points from said electrocardiogram monitoring signal;for each said location of base points, the location of the base point is compared against a predetermined threshold range based on a slow adaptative learning process using said patient's own cardiac movement, said range having an upper and lower threshold value for both an horizontal and vertical adjusted value for each of said base points;if all locations are within the predetermined, threshold range, performing a proximity analysis to determine improbable locations;if all locations are possible, determining an occurrence of a trend in said electrocardiogram monitoring signal from said locations. 2. A trend detection method as claimed in claim 1, further comprising providing a portable electrocardiogram monitor generating said electrocardiogram monitoring signal, wherein said portable electrocardiogram can be worn during day-to-day activities and for long periods. 3. A trend detection method as claimed in claim 1, further comprising extracting and storing trend data corresponding to said occurrence. 4. A trend detection method as claimed in claim 1, further comprising providing a relative position between two base points, which refers to a normal PQRST morphology learned from past normal cardiac movements of said patient;and wherein said determining comprises determining if a position of one of said base points with respect to another one of said base points matches said relative position using said locations;wherein said trend is detected if said position does not match said relative position over a predetermined period of time. 5. A trend detection method as claimed in claim 1, further comprising providing a distance tolerance;and wherein said determining comprises determining if a distance between one of said base points and another one of said base points is acceptable using said locations and said distance tolerance, based on possible rhythm changes;wherein said trend is detected if said distance is not acceptable over a predetermined period of time. 6. A trend detection method as claimed in claim 1, further comprising providing a base model for said locations of base points;and wherein said determining comprises determining if said locations correspond to said base model;wherein said trend is detected if said locations do not correspond to said base model over a predetermined period of time. 7. A trend detection method as claimed in claim 1, wherein said electrocardiogram monitoring signal is a digital signal. 8. A trend detection method as claimed in claim 1, further comprising sending an alarm message when said trend is determined to have occurred. 9. A trend detection method as claimed in claim 8, further comprising extracting and storing trend data corresponding to said occurrence and wherein said alarm message comprises said trend ... | Medicale Intelligence Inc. | A61B005/402
US06997878 | B2 | 2006-02-14 | Cuff for blood pressure monitor | 2003-02-03 | The invention provides a variety of cuff design for blood pressure monitor, which assures an easy and firm fitting to an arm of a patient. Specifically, the curled elastic member of the cuff is overlapped at both end portions for easy fitting. The rigidity and thickness of the curled elastic member are designed to follow the shape of the upper arm of the patient, which significantly varies in size depending on the location of the arm. | 1. A cuff for a blood pressure monitor, comprising a bladder and a normally curled elastic member disposed on the outside of the bladder to hold a ring shape of the cuff, wherein the normally curled elastic member comprises a material having at least two different rigidities in an axial direction of the ring shape,the portion of the normally curled elastic member having the lower rigidity including holes formed therein. | 1. A cuff for a blood pressure monitor, comprising a bladder and a normally curled elastic member disposed on the outside of the bladder to hold a ring shape of the cuff, wherein the normally curled elastic member comprises a material having at least two different rigidities in an axial direction of the ring shape,the portion of the normally curled elastic member having the lower rigidity including holes formed therein. 2. The cuff for a blood pressure monitor of claim 1, wherein the cuff is configured to be attached on an arm, and the normally curled elastic member has a higher rigidity of the different rigidities in a portion of the normally curled elastic member that is positioned at a thin portion of the arm when the cuff is attached on the arm and has a lower rigidity of the different rigidities in a portion of the normally curled elastic member that is positioned at a thick portion of the arm when the cuff is attached on the arm. 3. The cuff for a blood pressure monitor of claim 1, wherein the normally curled elastic member has at least two different rigidities also in a circumferential direction of the ring shape. 4. A cuff for a blood pressure monitor, comprising a bladder and a normally curled elastic member disposed on the outside of the bladder to hold a ring shape of the cuff, wherein the normally curled elastic member is made of a material having at least two different rigidities in a circumferential direction of the ring shape, andwherein the cuff is configured to be attached on an arm, and the normally curled elastic member has a higher rigidity of the different rigidities in a portion of the normally curled elastic member that is positioned at a thin portion of the arm when the cuff is attached on the arm and has a lower rigidity of the different rigidities in a portion of the normally curled elastic member that is positioned at a thick portion of the arm when the cuff is attached on the armthe portion of the normally curled elastic member having the lower rigidity including holes formed therein. 5. A cuff for a blood pressure monitor, comprising a bladder and a curled elastic member disposed on the outside of the bladder to hold a ring shape of the cuff, wherein the curled elastic member has at least two different rigidities in a circumferential direction of the ring shape,wherein the cuff is configured to be attached on an arm, and the curled elastic member has a higher rigidity of the different rigidities in a portion of the curled elastic member that is positioned at a thin portion of the arm when the cuff is attached on the arm and has a lower rigidity of the different rigidities in a portion of the curled elastic member that is positioned at a thick portion of the arm when the cuff is attached on the arm, andwherein the curled elastic member has the lower rigidity of the different rigidities at an end portion of the curled elastic member along the circumferential direction, the end portion being configured to be curled i... | Omron Corporation | A61B005/02
US07039455 | B1 | 2006-05-02 | Apparatus and method for removing magnetic resonance imaging-induced noise from ECG signals | 2002-09-27 | An apparatus and method is provided for improving the quality of electrocardiogram (ECG) signals obtained from a patient undergoing magnetic resonance imaging (MRI) wherein the ECG signal has relatively high levels of noise or interference voltages induced on it by changing magnetic fields. The apparatus includes the arrangement of a differential amplifier, a prefilter, a signal limiter (SL) circuit and an intermediate amplifier with an integral low pas filter. The prefilter limits the rise time or high frequency component of the noise or interfering voltages induced on the ECG that are presented to the signal limiter. | 1. An apparatus for processing an ECG signal having MRI noise induced thereon, the apparatus comprising: (a) a differential amplifier;(b) a prefilter electrically connected to the differential amplifier;(c) a signal limiter circuit electrically connected to the prefilter; and(d) an intermediate amplifier electrically connected to the signal limiter, the intermediate amplifier comprising a low pass filter. | 1. An apparatus for processing an ECG signal having MRI noise induced thereon, the apparatus comprising: (a) a differential amplifier;(b) a prefilter electrically connected to the differential amplifier;(c) a signal limiter circuit electrically connected to the prefilter; and(d) an intermediate amplifier electrically connected to the signal limiter, the intermediate amplifier comprising a low pass filter. 2. The apparatus of claim 1 wherein the differential amplifier has a gain of 7. 3. The apparatus of claim 1 wherein the prefilter is a single pole, low pass filter. 4. The apparatus of claim 1 wherein the prefilter has a cutoff frequency of approximately 80 Hz. 5. The apparatus of claim 1 wherein the low pass filter has a cutoff frequency of approximately 40 Hz. 6. The apparatus of claim 1 wherein the intermediate amplifier has a gain of 64. 7. A method of processing an ECG signal having MRI noise including RF magnetic field noise and gradient field noise induced thereon, the method comprising: (a) amplifying the ECG signal having MRI induced noise thereon;(b) prefiltering the ECG signal to remove at least a portion of the RF magnetic field noise therefrom;(c) signal limiting the ECG signal to remove at least a portion of the gradient field noise therefrom;(d) amplifying the ECG signal and further filtering therefrom frequencies above those of interest in the ECG signal. 8. A noise reduction circuit for use with an ECG monitor, the noise reduction circuit comprising: (a) a differential amplifier for amplifying a difference in electrical signals indicative of cardiac function to yield an ECG signal therefrom;(b) a prefilter electrically connected to the differential amplifier for filtering the ECG signal;(c) a signal limiter circuit electrically connected to the prefilter for removing high frequency noise from the ECG signal; and(d) an intermediate amplifier electrically connected to the signal limiter for amplification of the ECG signal received therefrom. 9. The noise reduction circuit of claim 8 wherein the differential amplifier has a gain of 7. 10. The noise reduction circuit of claim 8 wherein the prefilter is a single pole, low pass filter. 11. The noise reduction circuit of claim 8 wherein the prefilter has a cutoff frequency of approximately 500 Hz. 12. The noise reduction circuit of claim 8 wherein the intermediate amplifier further comprises a low pass filter for further removal of noise from the ECG signal. 13. The noise reduction circuit of claim 12 wherein the low pass filter has a cutoff frequency of approximately 500 Hz. 14. The noise reduction circuit of claim 1 wherein the intermediate amplifier has a gain of 64. 15. A method of adapting an ECG monitor equipped with a conventional noise reduction circuit, the method comprising the steps of: (a) providing a second noise reduction circuit, the second noise reduction circuit comprising (i) a differential amplifier, (ii) a prefilter electrically connected to the differ... | Medrad, Inc. | A61B005/0402 A61N005/0425
US06980850 | B1 | 2005-12-27 | System and method for emulating a surface EKG using an implantable cardiac stimulation device | 2002-12-30 | A surface electrocardiogram (EKG) is emulated using signals detected by the internal leads of an implanted device. The emulation is performed using a matrix-based technique that separately emulates each of the individual signals of a multiple-lead EKG, rather than merely emulating a single combined EKG. In one example, each of the twelve signals of a standard 12-lead EKG are individually emulated, allowing for separate processing and display. The emulation technique takes into account factors affecting the relative locations of the internal leads, such as respiration and posture, to thereby provide a more accurate emulation. A calibration technique is provided for calibrating the EKG emulation for use with a particular patient and a verification technique is provided for automatically verifying the reliability of the emulation. Any significant loss in emulation reliability is likely caused by lead dislodgment and so automatic detection of possible lead dislodgment is also achieved. | 1. A method for emulating a multiple-lead surface electrocardiogram (EKG) of a patient in which an implantable cardiac stimulation device is implanted, the method comprising: inputting electrical cardiac signals sensed from a combination of pairs of sensing electrodes implanted within the patient;inputting a signal representative of factors affecting the positions of the sensing electrodes; andgenerating an emulated surface electrocardiogram (EKG) based on the electrical cardiac signals while taking into account the factors affecting the positions of the sensing electrodes. | 1. A method for emulating a multiple-lead surface electrocardiogram (EKG) of a patient in which an implantable cardiac stimulation device is implanted, the method comprising: inputting electrical cardiac signals sensed from a combination of pairs of sensing electrodes implanted within the patient;inputting a signal representative of factors affecting the positions of the sensing electrodes; andgenerating an emulated surface electrocardiogram (EKG) based on the electrical cardiac signals while taking into account the factors affecting the positions of the sensing electrodes. 2. The method of claim 1 wherein the factors affecting the positions of the sensing electrodes include one or more of: phase of a respiratory cycle of the patient, phase of a cardiac cycle of the patient, and posture of the patient. 3. The method of claim 2wherein the signal representative of the factors affecting the positions of the sensing electrodes specifies the phase of the cardiac cycle of the patient; andwherein the step of generating an emulated surface EKG distinguishes among, at least, P-waves, QRS-complexes and T-waves. 4. The method of claim 2wherein the signal representative of the factors affecting the positions of the sensing electrodes specifies the phase of the respiratory cycle of the patient; andwherein the step of generating an emulated surface EKG distinguishes between, at least, inhalation and exhalation. 5. The method of claim 2wherein the signal representative of the factors affecting the positions of the sensing electrodes specifies the posture of the patient; andwherein the step of generating an emulated surface EKG distinguishes among, at least, standing, sitting and lying down. 6. The method of claim 1 wherein generating an emulated surface electrocardiogram is performed to emulate each of a plurality of separate signals associated with the multiple-lead surface EKG. 7. The method of claim 6 wherein emulating each of the plurality of separate signals is performed to emulate each of the twelve signals associated with a 12-lead surface EKG. 8. The method of claim 6 wherein emulating each of the plurality of separate signals associated with the multiple-lead surface EKG comprises converting the input cardiac signals into a time-varying vector F(t) having individual elements corresponding to the pairs of sensing electrodes;inputting a conversion matrix M of weighting factors representative of the relative extent to which the signals derived from the pairs of sensing electrodes influence surface voltages at locations corresponding to the multiple-lead surface EKG being emulated; andgenerating a time-varying vector E(t), representative of the surface voltages at the locations corresponding to the multiple-lead surface EKG being emulated, by calculating: E(t)=M*F(t). 9. The method of claim 8wherein the combination of sensing electrodes includes N pairs of electrodes and the surface EKG to be emulated includes M surface leads;wherein F(t) has N ele... | Pacesetter, Inc. | A61B005/04
US06963768 | B2 | 2005-11-08 | Whole body MRI scanning with moving table and interactive control | 2002-05-16 | The present invention includes a method and apparatus for high sensitivity whole body scanning using MR imaging. The invention includes acquiring MR data as the patient moves through the iso-center of the magnet while providing interactive control for the operator to change scan parameters and table motion and direction. The technique allows efficient whole body scanning for fast screening of abnormalities while allowing operator control during the screening process to interrupt table motion and redirect the speed and direction of the table while also allowing control over the acquisition plane, number of sections imaged, inter-section spacing, and the scan location. | 1. A method of MR image acquisition comprising: positioning a subject on a movable table; entering initial scan parameters; automatically moving the movable table during MR data acquisition; acquiring MR data based on the scan parameters entered; and allowing modification of at scan parameters while automatically moving the table and acquiring MR data. | 1. A method of MR image acquisition comprising: positioning a subject on a movable table; entering initial scan parameters; automatically moving the movable table during MR data acquisition; acquiring MR data based on the scan parameters entered; and allowing modification of at scan parameters while automatically moving the table and acquiring MR data. 2. The method of claim 1 further comprising providing interactive control of table motion and scan parameters that include control over and adjustment of at least one of: speed of the movable table, direction of table motion, and pulse sequence for MR data acquisition. 3. The method of claim 1 wherein the step of acquiring MR data includes ability to continuously scan while the movable table is in translation using a multi-planar fast imaging pulse sequence. 4. The method of claim 3 further comprising the step of initially selecting a desired plane for MR data acquisition through an anatomic region of interest, and as the subject traverses through a magnet iso-center, allowing operator adjustment of at least one of a number of sections imaged, section spacing, section scan locations, and imaging plane. 5. The method of claim 1 further comprising providing an ability to interrupt scanning after identifying an abnormality of interest, reversing the movable table and acquiring high spatial resolution image data. 6. The method of claim 5 wherein the abnormality of interest is a tumor and further comprises the steps of acquiring functional images and characterizing the tumor using one of contrast media uptake, diffusion, and multi-parametric imaging. 7. The method of claim 1 further comprising tailoring acquisition plane and spatial coverage to each anatomical region desired during MR image acquisition in real-time. 8. The method of claim 1 further comprising allowing continuous variation of imaging parameters including transmitter/receive gains and localized shimming. 9. The method of claim 1 further comprising obtaining multiple images of a same anatomical region to decrease false positive possibilities. 10. An MRI apparatus with sensitive whole body screening ability comprising: a magnetic resonance imaging (MRI) system having a plurality of gradient coils positioned about a bore of a magnet to impress a polarizing magnetic field and an RF transceiver system and an RF switch controlled by a pulse module to transmit RF signals to an RF coil assembly to acquire MR images; a table movable under computer control within the bore of the magnet; a computer programmed to: receive initial scan parameters and table translation parameters; translate the table; acquire MR data while the table translates; allow reception of user input during table translation and if so received, modify translation in response thereto; and allow reception of user input of scan parameters during table translation and if so received, modify MR data acquisition in response thereto. 11. The MRI apparatus of claim 10 where... | General Electric Company | A61B005/05
US06907282 | B2 | 2005-06-14 | Intensity map resampling for multi-leaf collimator compatibility | 2001-11-09 | A method for converting a source intensity map into a target intensity map having a geometry compatible with a desired multi-leaf collimator configuration. The source intensity map and the target intensity map each have a geometry defining sampling points of cells within the maps. The method includes defining a field on an object for radiation delivery. The field includes a plurality of cells defining the source intensity map, each of the cells having a treatment intensity level. An intermediate intensity map geometry is created such that the intermediate map contains sampling points of the source intensity map and the target intensity map. The method further includes defining treatment intensity levels for cells of the intermediate map and calculating treatment intensity levels for cells of the target intensity map based on the intensity level of the intermediate map cells. | 1. A method for converting a source intensity map into a target intensity map having a geometry compatible with a desired multi-leaf collimator configuration, each of the source intensity map and the target intensity map having a geometry defining sampling points of cells within the maps, the source intensity map having a geometry that is incompatible with the desired multi-leaf collimator configuration, the method comprising: defining a field on an object for radiation delivery, said field including a plurality of cells defining the source intensity map, each of said plurality of cells having a treatment intensity level; creating an intermediate intensity map geometry such that the intermediate map contains sampling points of the source intensity map and the target intensity map; defining treatment intensity levels for cells of the intermediate map; and calculating treatment intensity levels for cells of the target intensity map based on the intensity level of the intermediate map cells. | 1. A method for converting a source intensity map into a target intensity map having a geometry compatible with a desired multi-leaf collimator configuration, each of the source intensity map and the target intensity map having a geometry defining sampling points of cells within the maps, the source intensity map having a geometry that is incompatible with the desired multi-leaf collimator configuration, the method comprising: defining a field on an object for radiation delivery, said field including a plurality of cells defining the source intensity map, each of said plurality of cells having a treatment intensity level; creating an intermediate intensity map geometry such that the intermediate map contains sampling points of the source intensity map and the target intensity map; defining treatment intensity levels for cells of the intermediate map; and calculating treatment intensity levels for cells of the target intensity map based on the intensity level of the intermediate map cells. 2. The method of claim 1 wherein defining treatment intensity levels for the intermediate map cells comprises retrieving the treatment intensity levels from a treatment planning system and wherein the source intensity map and the intermediate intensity map have the same geometry and intensity levels. 3. The method of claim 1 wherein defining treatment intensity levels for the intermediate map cells comprises calculating the intermediate map treatment intensity levels based on the source map treatment intensity levels. 4. The method of claim 3 wherein calculating the intermediate map treatment intensity levels comprises interpolating or extrapolating using source map treatment intensity levels. 5. The method of claim 1 wherein calculating treatment intensity levels for cells of the target intensity map comprises down sampling values from the intermediate intensity map. 6. The method of claim 5 wherein down sampling values comprises performing an interpolation or extrapolation process. 7. The method of claim 1 further comprising delivering a radiation treatment with the desired multi-leaf collimator configuration from the target intensity map. 8. The method of claim 1 further comprising creating an intensity map for virtual micro intensity modulation radiation therapy based on the target intensity map. 9. A system for converting a source intensity map into a target intensity map having a geometry compatible with a desired multi-leaf collimator configuration, each of the source intensity map and the target intensity map having a geometry defining sampling points of cells within the map, the source intensity map having a geometry that is incompatible with the desired multi-leaf collimator configuration, the system comprising: a processor operable to receive a source intensity map geometry, target intensity map geometry, and treatment intensity levels for cells of the source intensity map, create an intermediate map geometry containing sampling points of the... | Siemens Medical Solutions USA, Inc. | A61B005/05
US06869402 | B2 | 2005-03-22 | Method and apparatus for measuring pulsus paradoxus | 2002-08-27 | A method and apparatus are disclosed of utilizing a source of arterial and/or arteriolar pulse waveform data from a patient for the purpose of measuring pulsus paradoxus. The arterial pulse waveform data source described is a pulse oximeter plethysmograph but can be any similar waveform data source, including intra-arterial transducer, blood pressure transducer, or plethysmograph. Through incorporation of the measurements of values, such as the area under the pulse waveform curve, that are time-domain functions of a change in height of the pulse waveform over at least a partial duration of the waveform, embodiments of the present invention represent a significant improvement upon previously described methods of measuring pulsus paradoxus and, further, produce improved accuracy in measurement of the multiple contributing variables generating pulsus paradoxus, in particular the contribution of diastolic events, to the extent that the physical signs so measured can be regarded as “Revised Pulsus Paradoxus.” | 1. A method for measuring pulsus paradoxus in a mammal, the method comprising: obtaining data indicative of pulsatile cardiovascular behavior from the mammal, the data comprising at least a first and a second pulse waveform; determining a first AUC from at least one component of the first pulse waveform; determining a second AUC from at least one component of the second pulse waveform; comparing the first AUC and the second AUC to measure pulsus paradoxus. | 1. A method for measuring pulsus paradoxus in a mammal, the method comprising: obtaining data indicative of pulsatile cardiovascular behavior from the mammal, the data comprising at least a first and a second pulse waveform; determining a first AUC from at least one component of the first pulse waveform; determining a second AUC from at least one component of the second pulse waveform; comparing the first AUC and the second AUC to measure pulsus paradoxus. 2. The method of claim 1, further comprising determining the presence or absence of pulsus paradoxus. 3. The method of claim 2, further comprising correlating a magnitude of the difference between the first AUC and second AUC with a magnitude of pulsus paradoxus. 4. The method of claim 1, wherein the comparing step is performed by a computer. 5. The method of claim 1, wherein the data indicative of pulsatile cardiovascular behavior is obtained from plethysmography. 6. The method of claim 5, wherein the data indicative of pulsatile cardiovascular behavior is obtained from pulse oximetry. 7. The method of claim 6, wherein the data indicative of pulsatile cardiovascular behavior is obtained from reflective pulse oximetry. 8. The method of claim 6, wherein the data indicative of pulsatile cardiovascular behavior is obtained from transmissive pulse oximetry. 9. The method of claim 1, wherein the at least one component of the first pulse waveform comprises a systolic portion of the first pulse waveform, and the at least one component of the second pulse waveform comprises a systolic portion of the second pulse waveform. 10. The method of claim 9, wherein the at least one component of the first pulse waveform further comprises a diastolic portion of the first pulse waveform, and the at least one component of the second pulse waveform further comprises a diastolic portion of the second pulse waveform. 11. The method of claim 1, wherein the comparing step comprises determining a difference between the first AUC and second AUC. 12. The method of claim 11, wherein an absolute value of the difference between the first AUC and second AUC is maximal for any two of a set of pulse waveforms obtained over a duration of at least one respiratory cycle. 13. The method of claim 11, further comprising providing a determination of the presence or absence of pulsus paradoxus in the mammal, the determination based upon whether a magnitude of the difference between the first AUC and second AUC is greater than a threshold value. 14. The method of claim 13, wherein pulsus paradoxus is determined to exist when the magnitude of the difference between the first AUC and second AUC is greater than 10 percent. 15. The method of claim 14, wherein pulsus paradoxus is determined to exist when the magnitude of the difference between the first AUC and second AUC is greater than 20 percent. 16. The method of claim 15, wherein pulsus paradoxus is determined to exist when the magnitude of the difference between the firs... | Precision Pulsus, Inc. | A61B005/00
US07003343 | B2 | 2006-02-21 | Method and apparatus for anatomically tailored k-space sampling and recessed elliptical view ordering for bolus-enhanced 3D MR angiography | 2002-04-01 | Current bolus chase magnetic resonance angiography is limited by the imaging time for each station. Tailoring the density of k-space sampling along the anterior-posterior direction of the coronal station allows a substantial decrease in scan time that leads to greater contrast bolus sharing among stations and consequently a significant improvement in image quality. Fast arterial-venous transit in the carotid arteries requires accurate, reliable timing of the acquisition to the bolus transit to maximize arterial signal and minimize venous artifacts. The rising edge of the bolus is not utilized in conventional elliptical-centric view ordering because the critical k-space center must be acquired with full arterial enhancement. The invention provides a recessed elliptical-centric view ordering scheme is introduced in which the k-space center is acquired a few seconds following scan initiation. The recessed view ordering is shown to be more robust to timing errors in a patient studies. | 1. A method of accelerating volumetric MR imaging of arteries or other structures comprising the following: a) an image volume having x, y and z orthogonal coordinates and a thickness along z=Tv;b) for every x, y location within the volume the structures to be imaged have a thickness Ts<Tv;c) three-dimensional k space sampling is accelerated by a factor F≦Tv/Ts by increasing the spacing of k space sampling along kz by a factor of F while maintaining the sampling spatial frequency ranges unchanged;d) reconstruct with 3-D Fourier Transformation utilizing a field of view in the z direction which is equal to Tv/F;e) create 1/F duplicate copies of the reconstructed volumetric data and stacked together adjacent to each other in the z direction; andf) post-process the combined multiple copies of volumetric data to select a subset of the total combined volume containing a single complete data set of the structures of interest. | 1. A method of accelerating volumetric MR imaging of arteries or other structures comprising the following: a) an image volume having x, y and z orthogonal coordinates and a thickness along z=Tv;b) for every x, y location within the volume the structures to be imaged have a thickness Ts&lt;Tv;c) three-dimensional k space sampling is accelerated by a factor F≦Tv/Ts by increasing the spacing of k space sampling along kz by a factor of F while maintaining the sampling spatial frequency ranges unchanged;d) reconstruct with 3-D Fourier Transformation utilizing a field of view in the z direction which is equal to Tv/F;e) create 1/F duplicate copies of the reconstructed volumetric data and stacked together adjacent to each other in the z direction; andf) post-process the combined multiple copies of volumetric data to select a subset of the total combined volume containing a single complete data set of the structures of interest. 2. The method of claim 1 in which the data is shifted in z prior to creating multiple copies so as to make step f) easier to perform. 3. A method of increasing the resolution of volumetric MR imaging of arteries or other structures comprising the following: a) an image volume having x, y and z orthogonal coordinates and a thickness along z=Tvb) for every x, y location within the volume the structures to be imaged have a thickness Ts&lt;Tvc) increase the maximum spatial frequency sampled by a factor F≦Tv/Ts by increasing the spacing of k space sampling along kz by a factor of F while maintaining the same total acquisition timed) reconstruct with 3-D Fourier Transformation utilizing a field of view in the z direction which is equal to Tv/Fe) create 1/F duplicate copies of the reconstructed volumetric data and stacked together adjacent to each other in the z directionf) post-process the combined multiple copies of volumetric data to select a subset of the total combined volume containing a single complete data set of the structures of interest. 4. A method of 3-D gradient echo MR imaging of the abdomen, pelvis and leg arteries of a patient lying on a moveable table in an MR imaging apparatus during a single injection of contrast agent comprising the following: a first station positioned to image abdominal and pelvis arteries with k space sampling starting at the edge of k space and the center of k space sampled during the second half of the acquisition; sampling of the center of k space is timed to coincide with contrast passing through the abdomen and pelvis arteries;a second station positioned to image thigh arteries accelerated with the method of any one of claims 1–2 with table movement and timing of the sampling of the center of k-space to coincide with contrast passing through the thigh arteries;a third station positioned to image calf arteries in which the center of k space is sampled near the beginning of the acquisition with table movement and timing of the sampling of the center of k-space to coincide with contrast pas... | Cornell Research Foundation, Inc. | A61B005/05 G01V003/00
US07174201 | B2 | 2007-02-06 | Position sensing system with integral location pad and position display | 2002-06-17 | Apparatus for performing a medical procedure on a tissue within a body of a subject includes a wireless tag which is fixed to the tissue and includes a first sensor coil. A second sensor coil is fixed to a medical device for use in performing the procedure. An integral processing and display unit includes a plurality of radiator coils, along with processing circuitry and a display. The radiator coils generate electromagnetic fields in a vicinity of the tissue, thereby causing currents to flow in the sensor coils. The processing circuitry processes the currents so as to determine coordinates of the tag relative to the medical device. The display is driven by the processing circuitry so as to present a visual indication to an operator of the medical device of an orientation of the device relative to the tag. | 1. Apparatus for performing a medical procedure on a tissue within a body of a subject, comprising: a wireless tag configured to be fixed to the tissue and comprising a first sensor coil;a medical device having a second sensor coil fixed to the medical device for use in performing the procedure; anda processing and display unit, comprising:a plurality of radiator coils, which generate electromagnetic fields in a vicinity of the tissue, thereby causing a first current and a second currents to flow, respectively, in the first sensor coil of the wireless tag and the second sensor coil of the medical device;processing circuitry for receiving a first signal and a second signal from the wireless tag and from the medical device, respectively, indicative of the first current and the second current, and to process the signals so as to determine position and orientation coordinates of the tag relative to the medical device;a display, coupled to the processing circuitry so as to present a visual indication to an operator of the medical device of a position and orientation of the medical device relative to the tag; anda case, containing the plurality of radiator coils, processing circuitry and display as an integral unit. | 1. Apparatus for performing a medical procedure on a tissue within a body of a subject, comprising: a wireless tag configured to be fixed to the tissue and comprising a first sensor coil;a medical device having a second sensor coil fixed to the medical device for use in performing the procedure; anda processing and display unit, comprising:a plurality of radiator coils, which generate electromagnetic fields in a vicinity of the tissue, thereby causing a first current and a second currents to flow, respectively, in the first sensor coil of the wireless tag and the second sensor coil of the medical device;processing circuitry for receiving a first signal and a second signal from the wireless tag and from the medical device, respectively, indicative of the first current and the second current, and to process the signals so as to determine position and orientation coordinates of the tag relative to the medical device;a display, coupled to the processing circuitry so as to present a visual indication to an operator of the medical device of a position and orientation of the medical device relative to the tag; anda case, containing the plurality of radiator coils, processing circuitry and display as an integral unit. 2. Apparatus according to claim 1, wherein the case is positionable by the operator in a location adjacent to the body and in proximity to the tissue. 3. Apparatus according to claim 1, wherein the tag further comprises a radio-frequency (RF) transmitter, which is adapted to transmit the first signal over the air, and wherein the processing and display unit comprises a RF receiver, which is adapted to receive the first signal over the air. 4. Apparatus according to claim 3, and comprising one or more acoustic transmitters, which are adapted to transmit acoustic energy into the body in a vicinity of the tissue, and wherein the tag is adapted to receive and use the acoustic energy in generating the first signal. 5. Apparatus according to claim 1, wherein the display is further adapted to present a visual indication of a distance from the probe to the tag. 6. Apparatus according to claim 1, wherein the medical device comprises an invasive medical tool, which is adapted to penetrate into the body so as to reach the tissue, and wherein the display is adapted to present the visual indication of the orientation of the tool within the body. 7. Apparatus according to claim 6, wherein the display is adapted to present a further visual indication of a distance from the tool to the tag. 8. Apparatus according to claim 6, wherein the invasive medical tool is adapted to perform a surgical procedure on the tissue. 9. Apparatus according to claim 6, wherein the invasive medical tool comprises an endoscope. 10. A method for performing a medical procedure on a tissue within a body of a subject, comprising: fixing a wireless tag to the tissue, the tag comprising a first sensor coil;coupling a second sensor coil to a medical device for use in performi... | Biosense, Inc. | A61B005/05
US07054449 | B2 | 2006-05-30 | Method for adjusting a transmission characteristic of an electronic circuit | 2001-09-25 | A method for adjusting a transmission characteristic of an electronic circuit, in a hearing aid. The hearing aid is connected with a computer. The current operating parameters of the hearing aid are transmitted to the computer. The computer from the operating parameters, calculates the current transmission characteristic and represents it on a monitor screen as a three-dimensional graphic diagram. The acoustics specialist, with the help of a mouse, selects a point of a segment of the graphic diagram and modifies the graphically depicted transmission characteristic. As a result the current operating parameters are modified as well. From these, a new current transmission characteristic is calculated and represented on the monitor screen as a three-dimensional graphic. This adaptation process is repeated until the current transmission characteristic coincides with the desired response. Then the current operating parameters are transmitted from the computer to the hearing aid and there stored in memory. | 1. A method for the adjustment of a transmission characteristic of an electronic circuit installed in a hearing aid (1), comprising the steps of: providing a computing means (2) with a display device (3) and a pointing and/or selecting device (4);providing a data transmission connection (6) between the electronic circuit to be adjusted and the computer (2);a current transmission characteristic of the electronic circuit is represented by a graphic diagram on the display device (3);with the help of the pointing and/or selecting device (4), if and when required, a point (10.1, 10.2, . . .) or a segment of the graphic diagram is selected on the display device (3) and the graphically represented transmission characteristic is changed with the pointing and/or selecting device (4), as a result of which also prescription for modifying certain operating parameters is evaluated, andtransmitting the modified operating parameters to the electronic circuit and storing the modified operating parameters in memory therein; and, whereinthe transmission characteristic of the electronic circuit is represented as a structure (8) in a three-dimensional graphic diagram representing at least one of amplification and output sound level as a function of frequency and of input sound level on the display device (3). | 1. A method for the adjustment of a transmission characteristic of an electronic circuit installed in a hearing aid (1), comprising the steps of: providing a computing means (2) with a display device (3) and a pointing and/or selecting device (4);providing a data transmission connection (6) between the electronic circuit to be adjusted and the computer (2);a current transmission characteristic of the electronic circuit is represented by a graphic diagram on the display device (3);with the help of the pointing and/or selecting device (4), if and when required, a point (10.1, 10.2, . . .) or a segment of the graphic diagram is selected on the display device (3) and the graphically represented transmission characteristic is changed with the pointing and/or selecting device (4), as a result of which also prescription for modifying certain operating parameters is evaluated, andtransmitting the modified operating parameters to the electronic circuit and storing the modified operating parameters in memory therein; and, whereinthe transmission characteristic of the electronic circuit is represented as a structure (8) in a three-dimensional graphic diagram representing at least one of amplification and output sound level as a function of frequency and of input sound level on the display device (3). 2. The method in accordance with claim 1, wherein the electronic circuit processes an input signal into an output signal and wherein, in the three-dimensional graphic diagram, the amplification and the output sound level are parameters of the output signal represented as a function of frequency and input sound level of the input signal. 3. The method according to claim 1, wherein a desired response is predefined and the current transmission characteristic is adapted to the desired response step by step. 4. The method in accordance with claim 1, wherein, for the purpose of the selection of a suitable view, the graphic diagram is rotated. 5. The method according to claim 2, wherein a desired response is predefined and the current transmission characteristic is adapted to the desired response step by step. 6. The method in accordance with claim 2, wherein, for the purpose of the selection of a suitable view, the graphic diagram is rotated. 7. The method in accordance with claim 3, wherein, for the purpose of the selection of a suitable view, the graphic diagram is rotated. 8. The method in accordance with claim 5, wherein, for the purpose of the selection of a suitable view, the graphic diagram is rotated. 9. The method according to claim 2, wherein the electric circuit is installed in a hearing aid. 10. The method according to claim 3, wherein the electric circuit is installed in a hearing aid. 11. The method according to claim 4, wherein the electric circuit is installed in a hearing aid. | Bernafon AG | H04R029/00 A61B005/00
US07018339 | B2 | 2006-03-28 | Method for processing heart rate information in a portable computer | 2002-02-25 | An apparatus and method for processing heart rate information by wirelessly transferring the information between a heart rate monitor and a personal computer device (PCD). The information includes programming information and heart rate or training information collected during use of the heart rate monitor. The PCD may include a personal digital assistant (PDA), a pocket personal computer (pocket PC), or a laptop computer. The information is transferred over an infrared, sonic, audible, ultrasonic, magnetic, radio frequency, IEEE 802.11, or Bluetooth-based communication link. | 1. A method of processing heart rate information in a portable computer device comprising monitoring by a wearable heart rate monitor and wirelessly transferring the heart rate information between the wearable heart rate monitor and the portable computer device, the method comprising the steps of: initiating a connect mode in the wearable heart rate monitor, the connect mode enabling a wireless communication link to be established with the wearable heart rate monitor;initiating transfer software in the portable computer device, the transfer software controlling the transfer of information between the wearable heart rate monitor and the portable computer device through the wireless communication link, the information including setting information adapted to program the wearable heart rate monitor;transferring the setting information bidirectionally between the wearable heart rate monitor and the portable computer device through the communication link comprising: initiating a batch heart rate/training information transfer mode in the portable computer device;transferring batch heart rate/training information from the heart rate monitor to the portable computer device comprising initiating wireless communication with the heart rate monitor by the portable computer device; requesting heart rate monitor information from the heart rate monitor by the portable computer device; transferring the heart rate monitor information from the heart rate monitor to the portable computer device; requesting file information from the heart rate monitor by the portable computer device; transferring the file information from the heart rate monitor to the portable computer device; requesting heart rate information from the heart rate monitor by the portable computer device comprising requesting a quantity of samples, information associated with at least one of the samples, a quantity of laps, and information associated with at least one of the laps; transferring the heart rate information from the heart rate monitor to the portable computer device; storing the heart rate information in the portable computer device;returning the heart rate monitor to a normal mode in response to completion of the batch transfer of the heart rate/training information; andprocessing the heart rate information in the portable computer device. | 1. A method of processing heart rate information in a portable computer device comprising monitoring by a wearable heart rate monitor and wirelessly transferring the heart rate information between the wearable heart rate monitor and the portable computer device, the method comprising the steps of: initiating a connect mode in the wearable heart rate monitor, the connect mode enabling a wireless communication link to be established with the wearable heart rate monitor;initiating transfer software in the portable computer device, the transfer software controlling the transfer of information between the wearable heart rate monitor and the portable computer device through the wireless communication link, the information including setting information adapted to program the wearable heart rate monitor;transferring the setting information bidirectionally between the wearable heart rate monitor and the portable computer device through the communication link comprising: initiating a batch heart rate/training information transfer mode in the portable computer device;transferring batch heart rate/training information from the heart rate monitor to the portable computer device comprising initiating wireless communication with the heart rate monitor by the portable computer device; requesting heart rate monitor information from the heart rate monitor by the portable computer device; transferring the heart rate monitor information from the heart rate monitor to the portable computer device; requesting file information from the heart rate monitor by the portable computer device; transferring the file information from the heart rate monitor to the portable computer device; requesting heart rate information from the heart rate monitor by the portable computer device comprising requesting a quantity of samples, information associated with at least one of the samples, a quantity of laps, and information associated with at least one of the laps; transferring the heart rate information from the heart rate monitor to the portable computer device; storing the heart rate information in the portable computer device;returning the heart rate monitor to a normal mode in response to completion of the batch transfer of the heart rate/training information; andprocessing the heart rate information in the portable computer device. 2. A method of processing heart rate information in a portable computer device comprising monitoring by a heart rate monitor and wirelessly transferring the heart rate information between the heart rate monitor and the portable computer device as defined by claim 1, wherein the step of processing further includes the step of displaying the processed information on the portable computer device. 3. A method of processing heart rate information in a portable computer device comprising monitoring by a heart rate monitor and wirelessly transferring the heart rate information between the heart rate monitor and the portable computer device as defined by cla... | Polar Electro Oy | A61B005/024
US07031763 | B1 | 2006-04-18 | MRI shoulder coil | 2002-04-18 | A MRI quadrature coil for imaging a human shoulder has a first coil shaped to conform to the top of the shoulder and a second coil adapted to encircle the shoulder. The second coil has a magnetic axis generally orthogonal to that of the first coil. | 1. A MRI quadrature coil for imaging a human shoulder, said shoulder having a superior portion, said quadrature coil comprising: a first coil arcuately shaped to conform to and extend over said superior portion of said shoulder from a chest portion to a shoulder blade portion, said first coil having a magnetic axis; anda second coil adapted to encircle said shoulder, said second coil having a magnetic axis generally orthogonal to said first coil magnetic axis. | 1. A MRI quadrature coil for imaging a human shoulder, said shoulder having a superior portion, said quadrature coil comprising: a first coil arcuately shaped to conform to and extend over said superior portion of said shoulder from a chest portion to a shoulder blade portion, said first coil having a magnetic axis; anda second coil adapted to encircle said shoulder, said second coil having a magnetic axis generally orthogonal to said first coil magnetic axis. 2. A MRI quadrature coil according to claim 1, wherein said first coil includes a loop coil element and said second coil includes a loop coil element. 3. A MRI quadrature coil according to claim 1, wherein said first coil includes a Helmholtz coil element and said second coil includes a loop coil element. 4. A MRI quadrature coil according to claim 3, wherein said first coil further includes another Helmholtz coil element, said first coil being an array coil. 5. A MRI quadrature coil according to claim 1, wherein said coils are flexible. 6. A MRI quadrature coil according to claim 1, further comprising a rigid housing, said first and second coils extending therefrom. 7. A MRI quadrature coil according to claim 1, wherein said first and second coils are mounted on a common flexible frame. 8. An imaging system for imaging a shoulder of a body, said imaging system comprising: a first coil configured to conform to and extend over a superior portion of the shoulder from a chest portion to a shoulder blade portion; anda second coil separate from the first coil and configured to encircle the shoulder. 9. An imaging system according to claim 8, wherein said first coil has a magnetic axis orthogonal to a magnetic axis of said second coil. 10. An imaging system according to claim 8, wherein said imaging system comprises a magnetic resonance imaging system. 11. An imaging system according to claim 8, wherein said first coil comprises Helmholtz coil elements. 12. An imaging system according to claim 8, wherein said second coil encircles the shoulder when an arm of the body is inserted through the second coil. 13. An imaging system according to claim 8, wherein said first coil extends from an end of a clavicle bone of the body to an end of a humerus bone of the body. 14. An imaging system according to claim 8, wherein said first coil comprises Helmholtz coil elements, and the second coil is located between a portion of a first of the Helmholtz coil elements and a portion of a second of the Helmholtz coil elements. 15. A method for imaging, said method comprising: providing a first coil;providing a second coil separate from the first coil;conforming the first coil to extend over a superior portion of a shoulder of a body and partially down a front and a back of the shoulder across an axis from a neck end of a clavicle bone to a head of a humerus bone; andconfiguring the second coil to encircle the shoulder. 16. A method according to claim 15, further comprising configuring the first coil t... | General Electric Company | A61B005/055
US07039446 | B2 | 2006-05-02 | Indirect measurement of tissue analytes through tissue properties | 2003-01-22 | Methods and system for noninvasive determination of tissue analytes utilize tissue properties as reflected in key features of an analytical signal to improve measurement accuracy and precision. Physiological conditions such as changes in water distribution among tissue compartments lead to complex alterations in the measured analytical signal of skin, leading to a biased noninvasive analyte measurement. Changes in the tissue properties are detected by identifying key features in the analytical signal responsive to physiological variations. Conditions not conducive to the noninvasive measurement are detected. Noninvasive measurements that are biased by physiological changes in tissue are compensated. In an alternate embodiment, the analyte is measured indirectly based on natural physiological response of tissue to changes in analyte concentration. A system capable of such measurements is provided. | 1. A method for noninvasive measurement of a target analyte in a tissue, comprising the steps of: collecting an analytical signal from the tissue, said collected signal comprising a tissue measurement;extracting at least one spectral feature from the analytical signal indicative of the effect of the target analyte on the probed tissue; andeither correcting a direct analyte measurement based on said at least one feature; orcalculating concentration of said analyte indirectly by application of a calibration model to said at least one feature. | 1. A method for noninvasive measurement of a target analyte in a tissue, comprising the steps of: collecting an analytical signal from the tissue, said collected signal comprising a tissue measurement;extracting at least one spectral feature from the analytical signal indicative of the effect of the target analyte on the probed tissue; andeither correcting a direct analyte measurement based on said at least one feature; orcalculating concentration of said analyte indirectly by application of a calibration model to said at least one feature. 2. The method of claim 1, wherein said analytical signal is generated using any of: fluorescence spectroscopy;ultraviolet spectroscopy from 200 to 400 nm;visible spectroscopy from 400 to 700 nm;infrared and Fourier transform infrared (FTIR) spectroscopy;passive infrared spectroscopy;mid infrared spectroscopy from 4000–700 cm−1 in any of diffuse reflectance and transmission;attenuated total reflectance (ATR) spectroscopy;far infrared radiation spectroscopy;Kromoscopy in reflection or transmission;emission spectroscopy;Raman spectroscopy;photoacoustic and pulse laser photoacoustic spectroscopy;photon scattering from 400 to 2500 nm;bioelectrical response spectroscopy;oscillating thermal gradient spectroscopy;polarimetry;ultrasound spectroscopy;near infrared spectroscopy from 700 to 2500 nm in any of diffuse reflectance, transflectance, and transmission mode; andradio wave spectroscopy. 3. The method of claim 2, further comprising the step of: preprocessing said tissue measurement. 4. The method of claim 3, wherein said step of preprocessing said tissue measurement comprises any of the steps of: correcting said signal utilizing a reference;filtering said signal;calculating any of a first and second derivative of said signal;normalizing said signal;selecting portions of said signal;scatter correcting said signal; andtranslating said signal. 5. The method of claim 3, further comprising the step of: determining difference between a tissue template and either the preprocessed tissue measurement or the at least one extracted feature according to: z=x−(cxt+d);wherein x comprises either the pre-processed measurement or a set of at least one extracted features, xt comprises a tissue template associated with a measurement period, and c and d are slope and intercept adjustments to the tissue template. 6. The method of claim 5, wherein said tissue template is determined through one or more tissue measurements combined according to a predetermined data selection criterion during each measurement period. 7. The method of claim 6, wherein a measurement period comprises a time period over which accuracy of the tissue measurement remains within desired specifications. 8. The method of claim 6, further comprising the step of: providing an associated set of reference values combined according to said predetermined data selection criterion to form a measurement bias adjustment. 9. The method of claim 5, wherein the tissue te... | Sensys Medical, Inc. | A61B005/00
US07100439 | B2 | 2006-09-05 | Balance control system for weight scales | 2002-12-02 | The present invention is directed to a weight scale having a balance control system that detects unequal load distribution to the scale's load sensors and produces an output signal that enables the user to re-distribute weight until a predetermined level of balance is achieved. Various load sensors are positioned in an array within the scale and are in communication with a controller that receives signals from the sensors indicative of relative load applied to each. The controller generates a signal that is displayed or emitted to the user indicative of the position of the user's center of gravity relative to the sensors. | 1. A weight position sensor system for use with a scale for measuring the weight of an object, said scale comprising: a weight supporting surface for supporting an object to be weighed;three or more load detection sensors adapted to detect the presence of said object on said supporting surface;a controller adapted to receive a signal from said three or more load detection sensors indicative of the relative position of the center of gravity of said object with respect to the location of said three or more load detection sensors;an output signal generator adapted to generate a signal indicative of said relative position; andan electronic visual indicator adapted to display said signal indicative of said relative position in the form of an array of lights arranged symmetrically,wherein said electronic visual indicator further comprises a central light within said array of lights. | 1. A weight position sensor system for use with a scale for measuring the weight of an object, said scale comprising: a weight supporting surface for supporting an object to be weighed;three or more load detection sensors adapted to detect the presence of said object on said supporting surface;a controller adapted to receive a signal from said three or more load detection sensors indicative of the relative position of the center of gravity of said object with respect to the location of said three or more load detection sensors;an output signal generator adapted to generate a signal indicative of said relative position; andan electronic visual indicator adapted to display said signal indicative of said relative position in the form of an array of lights arranged symmetrically,wherein said electronic visual indicator further comprises a central light within said array of lights. 2. A system according to claim 1, wherein said output signal generator generates signals and displays corresponding visual signals in response to successive re-positioning of said object's center of gravity relative to said three or more load detection sensors. 3. A system according to claim 1, wherein said output signal generator generates an audible signal indicative of said relative position. 4. A scale for measuring the weight of an object, said scale comprising a weight supporting surface for supporting an object to be weighed;three or more load detection sensors adapted to detect the presence and weight of said object on said supporting surface;a controller adapted to receive a signal from said three or more load detection sensors indicative of the relative position of the center of gravity of said object with respect to the location of said three or more load detection sensors, and one or more signals indicative of the weight of said object;an output signal generator adapted to generate a signal indicative of the relative position and to generate a signal indicative of the weight of said object; andan electronic visual indicator adapted to display said signal indicative of said relative position of said object in the form of an array of lights arranged symmetrically and a central light within said array of lights. | Conair Corporation | A61B005/00 G01G019/00
US06931276 | B2 | 2005-08-16 | Method and device for use in micturition studies | 2002-04-29 | Method and device for the sensing of the electromyographic activity of an individual's bladder detrusor muscle, where an elongated electrode is inserted through the individual's urethra, so that the proximal end of the electrode becomes positioned in the bladder, and the bladder is charged with a fluid having electrical conductivity, and the electrical potential of the detrusor muscle, mediated by the fluid, is sensed by the electrode. | 1. Method for the sensing of the electromyographic activity of an individual's bladder detrusor muscle, wherein an elongated electrode is inserted through the individual's urethra, so that the proximal end of the electrode becomes positioned in the bladder, and the bladder is charged with a fluid having electrical conductivity, and the electrical potential of the detrusor muscle, mediated by the fluid, is sensed by the electrode. | 1. Method for the sensing of the electromyographic activity of an individual's bladder detrusor muscle, wherein an elongated electrode is inserted through the individual's urethra, so that the proximal end of the electrode becomes positioned in the bladder, and the bladder is charged with a fluid having electrical conductivity, and the electrical potential of the detrusor muscle, mediated by the fluid, is sensed by the electrode. 2. The method according to claim 1, wherein the electrode is inserted by use of a protective tube. 3. The method according to claim 2, wherein the proximal end of the electrode is protected by the protective tube during the insertion step, and said proximal end is released from the protective tube when located in the desired position. 4. The method according to claim 2, wherein the bladder is charged with the fluid through the channel of the protective tube. 5. The method according to claim 4, wherein the fluid is charged by use of a protective tube having an additional channel for the discharge. 6. A method according to claim 2, wherein the bladder pressure is sensed simultaneously by a pressure probe, located on the protective tube at a position so as to become positioned inside the bladder. 7. A method according to claim 2, wherein the electromyographic activity of the urethral sphincter is sensed simultaneously by one or more electrodes arranged on the outer side of the protective tube at a position so as to come in contact with the external urethral sphincter when the protective tube is properly inserted. 8. The method according to claim 7, wherein the bladder pressure is sensed simultaneously by a pressure probe located on the protective tube at a position so as to become positioned inside the bladder. 9. The method according to claim 1, wherein the fluid is an aqueous electrolyte solution. 10. A device for the recording of the electromyographic activity of an individual's bladder detrusor muscle, wherein said device is intended to be inserted through the individual's urethra, said device comprising an elongated electrode adapted to sense electromyographic activity of the detrusor muscle, said electromyographic activity being mediated by an electrically conductive fluid in the bladder, said electrode being an elongated threadlike body made of a metal or metal alloy, the proximal end of said electrode being adapted to to be received in the bladder, and a distal end, and means for recording the electrical activity sensed by the electrode. 11. The device according to claim 10, comprising a protective tube surrounding the electrode. 12. The device according to claim 11, comprising a plunger, arranged to be longitudinally movable in the protective tube, said plunger having a proximal end adapted to receive the electrode. 13. The device according to claim 11, wherein the protective tube has at least two parallel channels. 14. The device according to claim 11, wherein a pressure probe, intended for sensing ... | - | A61B005/04
US07014615 | B2 | 2006-03-21 | Interstitial fluid collection and constituent measurement | 2003-01-29 | An apparatus and method is disclosed for obtaining and measuring constituents in a sample of body fluid. The apparatus includes a member which is sized to penetrate into at least the dermal layer of skin to collect a sample of body fluid located within the dermal layer. | 1. A method for testing a sample of body fluid, comprising: forming an artificial opening into a patient's skin but not through a dermal layer of said skin with a penetration member, wherein said opening is sized to access a constituent-containing, substantially blood-free fluid in the skin;drawing a sample of said blood-free fluid into a first portion of said penetration member that is outside of said dermal layer while a second portion of said penetration member is within said dermal layer; andtesting said sample for said constituent. | 1. A method for testing a sample of body fluid, comprising: forming an artificial opening into a patient's skin but not through a dermal layer of said skin with a penetration member, wherein said opening is sized to access a constituent-containing, substantially blood-free fluid in the skin;drawing a sample of said blood-free fluid into a first portion of said penetration member that is outside of said dermal layer while a second portion of said penetration member is within said dermal layer; andtesting said sample for said constituent. 2. A method according to claim 1 wherein said constituent is glucose. 3. A method according to claim 1 wherein said sample is tested by passing light through said sample and measuring an absorption of predetermined wavelengths. 4. A method according to claim 1 wherein said sample is tested colorimetrically. 5. A method according to claim 1 wherein said sample is tested electro-chemically. 6. A method according to claim 1 wherein said sample is tested by attenuated total reflection. 7. A method according to claim 1 comprising compressing said skin in a region surrounding said opening. 8. A method for testing a sample of body fluid, comprising: inserting a penetration member into a patient's skin but not through a dermal layer of the skin wherein said penetration member is sized to access a constituent-containing, substantially blood-free fluid in the skin and wherein a first portion of said penetration member is outside of said dermal layer while a second portion of said penetration member is within said dermal layer; and testing said accessed fluid for said constituent wherein said fluid is drawn into the first portion of said penetration member prior to said testing. 9. A method according to claim 8 wherein said constituent is glucose. 10. A method according to claim 8 wherein said sample is tested colorimetrically. 11. A method according to claim 8 wherein said sample is tested electrochemically. 12. A method according to claim 8 wherein said sample is tested by attenuated total reflection. 13. A method according to claim 8 wherein said penetration member is a needle. | Integ, Inc. | A61B005/00
US07108662 | B2 | 2006-09-19 | Device and method for sample collection | 2002-02-28 | A collection device for a testing system includes a central volume and a trough, adjacent the central volume. The trough is configured with a pour spout such that, as fluid is poured from the collection device, a controlled volume remains in the trough. The trough is configured to receive a portion of a test kit and to apply the controlled volume of the sample to test strips within the test kit. The trough can also include apertures in the wall between the trough and the central volume which allow excess sample to flow back into the cup when the cup is placed upright after pouring and when the test kit is inserted into the trough. | 1. A collection device for use with a testing system, said device comprising: a central chamber defining an open volume;a trough defining an open volume, smaller than said open volume of said central chamber, said trough being positioned adjacent said central chamber, and said trough comprising a surface that includes at least one aperture positioned to allow excess sample stored in said trough to flow into said central chamber when said collection device is placed in an upright position after pouring;an interior pour spout positioned for flow of a sample from said central chamber into said trough; andan exterior pour spout positioned for flow of a sample from said trough and out from said device;wherein said central chamber is adapted to hold a sample, said trough is adapted to receive a controlled volume of the sample as the sample is poured from said device through said interior pour spout and said exterior pour spout, and said trough is configured to receive a portion of a test kit for submersion in said controlled volume of the sample. | 1. A collection device for use with a testing system, said device comprising: a central chamber defining an open volume;a trough defining an open volume, smaller than said open volume of said central chamber, said trough being positioned adjacent said central chamber, and said trough comprising a surface that includes at least one aperture positioned to allow excess sample stored in said trough to flow into said central chamber when said collection device is placed in an upright position after pouring;an interior pour spout positioned for flow of a sample from said central chamber into said trough; andan exterior pour spout positioned for flow of a sample from said trough and out from said device;wherein said central chamber is adapted to hold a sample, said trough is adapted to receive a controlled volume of the sample as the sample is poured from said device through said interior pour spout and said exterior pour spout, and said trough is configured to receive a portion of a test kit for submersion in said controlled volume of the sample. 2. A collection device according to claim 1, further including an engagement mechanism configured to hold a test kit in said trough. 3. A collection device according to claim 2, wherein said engagement mechanism includes locating fingers configured to engage the test kit. 4. A collection device according to claim 3, wherein said trough holds a predetermined volume sample when the test kit is inserted into said trough. 5. A collection device according to claim 4, wherein said predetermined volume is in a range of 1000 microliters to 2000 microliters. 6. A collection device according to claim 1, wherein said central chamber has an asymmetrical cross-sectional shape. 7. A collection device according to claim 6, wherein said central chamber defines a “D” shaped cross-section. 8. A collection device according to claim 1, further comprising a cup defining said central chamber and a lid configured to at least partially cover said central chamber. 9. A collection device according to claim 8, wherein said trough is formed in said cup. 10. A collection device according to claim 9, said trough being formed on an exterior surface of a side surface of said central chamber. 11. A collection device for use with a testing system, said device comprising: a cup defining a central chamber, said chamber defining an open volume;a lid configured to at least partially cover said central chamber;a trough formed in said lid, said trough defining an open volume smaller than said open volume of said central chamber, said trough being positioned adjacent said central chamber;an interior pour snout positioned for flow of a sample from said central chamber into said trough; andan exterior pour spout positioned for flow of a sample from said trough and out from said device,wherein said central chamber is adapted to hold a sample, said trough is adapted to receive a controlled volume of the sample as the sample is poured from said ... | Quest Diagnostics Incorporated | A61B005/00 B65D081/00
US07006856 | B2 | 2006-02-28 | Signal quality metrics design for qualifying data for a physiological monitor | 2003-01-10 | A method and a device for determining the quality of signal used for measuring a physiological parameter. One embodiment of the present invention is directed towards a pulse oximeter, where the measured physiological parameter includes a patient's pulse rate and blood oxygen saturation. The signal quality, which is indicative of the accuracy and reliability of the measured physiological parameter, is calculated by combining a plurality of signal quality indicators, each of which is an indicator of a quality of the measured signal. The value of the signal quality metric is compared to a threshold and based on this comparison various decisions are made by the medical device. One decision is directed towards deciding whether or not to display the measured physiological parameter, to ensure that only accurate measured values are displayed. Another decision is directed towards providing feedback to guide the clinician to adjust the location of the sensor to a more suitable tissue location. | 1. A method for displaying a physiological parameter, comprising: processing a signal quality metric associated with a measured physiological parameter to temper changes in the signal quality metric over time;comparing an attribute of the processed signal quality metric to a threshold; anddisplaying the measured physiological parameter for an interval determined by the attribute of the processed signal quality metric if the attribute is greater than the threshold. | 1. A method for displaying a physiological parameter, comprising: processing a signal quality metric associated with a measured physiological parameter to temper changes in the signal quality metric over time;comparing an attribute of the processed signal quality metric to a threshold; anddisplaying the measured physiological parameter for an interval determined by the attribute of the processed signal quality metric if the attribute is greater than the threshold. 2. The method of claim 1, comprising ceasing to display the measured physiological parameter if the attribute is less than the threshold. 3. The method of claim 1, comprising deriving the signal quality metric from one or more signal quality indicators associated with the measured physiological parameter. 4. The method of claim 1, wherein processing the signal quality metric comprises applying an integral control process to the signal quality metric. 5. The method of claim 1, wherein processing the signal quality metric comprises applying a derivative control process to the signal quality metric. 6. A method of indicating signal quality for a measured physiological parameter, comprising: performing at least two of a proportional control operation, an integral control operation, or a derivative control operation to a signal quality metric associated with a measured physiological parameter to generate at least two of a proportional signal quality metric, an integral signal quality metric, or a derivative signal quality metric, respectively;generating a processed signal quality metric based upon the at least two of the proportional signal quality metric, the integral signal quality metric, or the derivative signal quality metric; andproviding an indication of the signal quality of the measured physiological parameter based upon the processed signal quality metric. 7. The method of claim 6, wherein providing the indication comprises adjusting the display duration of the measured physiological parameter based upon the processed signal quality metric. 8. The method of claim 6, wherein providing the indication comprises ceasing to display the measured physiological parameter if the processed signal quality metric is less than a threshold value. 9. The method of claim 6, wherein providing the indication comprises displaying an alphanumeric message if the processed signal quality metric is less than a threshold value. 10. The method of claim 6, wherein providing the indication comprises displaying a visual alarm if the processed signal quality metric is less than a threshold value. 11. The method of claim 6, wherein providing the indication comprises sounding an audible alarm if the processed signal quality metric is less than a threshold value. 12. The method of claim 6, wherein providing the indication comprises displaying a numerical scale of signal quality. 13. A tangible machine readable medium, comprising: code adapted to process a signal quality metric associated with a measur... | Nellcor Puritan Bennett Incorporated | A61B005/00
US07037273 | B2 | 2006-05-02 | Core body temperature monitoring in heart failure patients | 2002-05-22 | An implanted heart monitor includes sensors that measure various aspects of the heart failure patient's heart. A remote heart monitoring system connects the implanted heart monitor to a care provider, such as a physician. The data provided by the implanted heart monitor permits the care provider to obtain valuable data on the heart in order to make health care decisions affecting the heart failure patient's treatment. In many cases, the measurement of core body temperature and other patient data will enable the care provider to alter the patient's treatment to address the patient's condition. The implanted heart monitor can communicate over a wireless communication link with an external monitor. The implanted heart monitor may be implemented as part of a pacing device (i.e., pace maker) or may be a separate unit devoted to monitoring functions. The external monitor communicates with a monitoring station over a communication link. The monitoring station can operate as a centralized data collection unit, collecting data from multiple external monitors and multiple implanted heart monitors. Various other aspects of a heart failure patient's heart and/or body can be monitored, such as heart rate, blood pH levels, blood CO2 levels, and any other indications of the heart failure patient's activity. Various predetermined thresholds may be set to trigger alarms and/or data reports. | 1. A heart monitoring system, comprising: a temperature sensor for determining a core body temperature of a patient;a bodily activity sensor for determining a bodily activity level of the patient; anda heart monitor operatively connected to the temperature sensor and bodily activity sensor and configured to receive data representative of the core body temperature and the bodily activity level;the heart monitor further being configured to generate an alarm message upon determination that a relationship between the core body temperature and the bodily activity level is outside of a predefined range. | 1. A heart monitoring system, comprising: a temperature sensor for determining a core body temperature of a patient;a bodily activity sensor for determining a bodily activity level of the patient; anda heart monitor operatively connected to the temperature sensor and bodily activity sensor and configured to receive data representative of the core body temperature and the bodily activity level;the heart monitor further being configured to generate an alarm message upon determination that a relationship between the core body temperature and the bodily activity level is outside of a predefined range. 2. The heart monitoring system of claim 1, wherein the temperature sensor is of a configuration that allows the temperature sensor to be implanted in the patient's body. 3. The heart monitoring system of claim 1, wherein the bodily activity sensor is of a configuration that allows the bodily activity sensor to be implanted in the patient's body. 4. The heart monitoring system of claim 1, wherein the heart monitor is of a configuration that allows the external monitor to be implanted in the patient's body. 5. The heart monitoring system of claim 1, wherein the heart monitor is of a configuration that allows the heart monitor to be outside the patient's body and wherein the heart monitoring system includes a wireless communication link that allows wireless communications between the temperature sensor, bodily activity sensor and heart monitor. 6. The heart monitoring system of claim 1, wherein the heart is operative to determine a relationship between the core body temperature and the bodily activity level. 7. The heart monitoring system of claim 6, wherein the heart monitor is further operative to determine the slope of a two-dimensional representation of the relationship between the core body temperature and the bodily activity level. 8. The heart monitoring system of claim 6, wherein the heart monitor is further operative to generate an alarm when the relationship between the core body temperature and the bodily activity level is outside of a predetermined acceptable range. 9. The heart monitoring system of claim 6, wherein the heart monitor is further operative to transmit an alarm to a remote monitoring station, when the relationship between the core body temperature and the bodily activity level is outside of a predetermined acceptable range. 10. The heart monitoring system of claim 1, wherein the heart monitor is further operative transmit the core body temperature and the bodily activity level to a remote monitoring station. 11. A method for diagnosing a performance level of a heart in a patient, comprising the steps of: determining a core body temperature;determining a bodily activity level;determining a relationship between the core body temperature and the bodily activity level; andgenerating a first alarm message, in response to determining that the relationship between the core body temperature and the bodily activity level is outsi... | Cardiac Pacemakers, Inc. | A61B005/00
US07011630 | B2 | 2006-03-14 | Methods for computing rolling analyte measurement values, microprocessors comprising programming to control performance of the methods, and analyte monitoring devices employing the methods | 2002-06-21 | The present invention relates to methods to increase the number of analyte-related signals used to provide analyte measurement values, e.g., when two or more analyte-related signals are used to obtain a single analyte measurement value a “rolling” value based on the two or more signals can be employed. In another aspect, interpolation and/or extrapolation methods are used to estimate unusable, missing or error-associated analyte-related signals. Further, interpolation and extrapolation of values are employed in another aspect of the invention that reduces the incident of failed calibrations. Further, the invention relates to methods, which employ gradients and/or predictive algorithms, to provide an alert related to analyte values exceeding predetermined thresholds. The invention includes the above-described methods, one or more microprocessors programmed to execute the methods, one or more microprocessors programmed to execute the methods and control at least one sensing and/or sampling device, and monitoring systems employing the methods described herein. | 1. A method of increasing the number of analyte measurement values related to the amount or concentration of an analyte in a subject as measured using an analyte monitoring device, said method comprising providing a series of analyte-related signals obtained from the analyte monitoring device over time, wherein (a) two or more contiguous analyte-related signals are used to obtain a single analyte measurement value (M), (b) analyte-related signals are not used to calculate more than one analyte measurement value, and (c) said two or more contiguous analyte-related signals, used to obtain the single analyte measurement value, comprise first and last analyte-related signals of the series;mathematically computing rolling analyte measurement values, wherein (i) each rolling analyte measurement value is calculated based on two or more contiguous analyte-related signals from the series of analyte-related signals obtained from the analyte monitoring device, (ii) a subsequent rolling analyte measurement value is mathematically computed by dropping said first analyte-related signal and including an analyte-related signal contiguous and subsequent to the last analyte-related signal, (iii) further rolling analyte measurement values are obtained by repeating the dropping of the first analyte-related signal used to calculate the previous rolling analyte measurement and including an analyte-related signal contiguous and subsequent to the last analyte-related signal used to calculate the previous rolling analyte measurement, and (iv) each rolling analyte measurement value provides a measurement related to the amount or concentration of analyte in the subject; andincreasing the number of analyte measurement values derived from the analyte-related signals in the series of analyte-related signals obtained from the analyte monitoring device by serially calculating rolling analyte measurement values, thereby increasing the number of analyte measurement values relative to the number of analyte measurement values provided when two or more contiguous analyte-related signals are used to obtain a single analyte measurement value (M) and analyte-related signals are not used to calculate more than one analyte measurement value. | 1. A method of increasing the number of analyte measurement values related to the amount or concentration of an analyte in a subject as measured using an analyte monitoring device, said method comprising providing a series of analyte-related signals obtained from the analyte monitoring device over time, wherein (a) two or more contiguous analyte-related signals are used to obtain a single analyte measurement value (M), (b) analyte-related signals are not used to calculate more than one analyte measurement value, and (c) said two or more contiguous analyte-related signals, used to obtain the single analyte measurement value, comprise first and last analyte-related signals of the series;mathematically computing rolling analyte measurement values, wherein (i) each rolling analyte measurement value is calculated based on two or more contiguous analyte-related signals from the series of analyte-related signals obtained from the analyte monitoring device, (ii) a subsequent rolling analyte measurement value is mathematically computed by dropping said first analyte-related signal and including an analyte-related signal contiguous and subsequent to the last analyte-related signal, (iii) further rolling analyte measurement values are obtained by repeating the dropping of the first analyte-related signal used to calculate the previous rolling analyte measurement and including an analyte-related signal contiguous and subsequent to the last analyte-related signal used to calculate the previous rolling analyte measurement, and (iv) each rolling analyte measurement value provides a measurement related to the amount or concentration of analyte in the subject; andincreasing the number of analyte measurement values derived from the analyte-related signals in the series of analyte-related signals obtained from the analyte monitoring device by serially calculating rolling analyte measurement values, thereby increasing the number of analyte measurement values relative to the number of analyte measurement values provided when two or more contiguous analyte-related signals are used to obtain a single analyte measurement value (M) and analyte-related signals are not used to calculate more than one analyte measurement value. 2. The method of claim 1, wherein said rolling analyte measurement value is an average of two or more analyte-related signals. 3. The method of claim 1, wherein said rolling analyte measurement value is a sum of two or more analyte-related signals. 4. The method of claim 1, wherein each analyte-related signal is represented by an integral over time, and said rolling analyte measurement value is obtained by integral splitting. 5. The method of claim 1, wherein said monitoring device comprises a sampling device and a sensing device, and wherein said providing the series of analyte-related signals obtained from an analyte monitoring device comprises extracting a sample from the subject alternately into a first collection reservoir and then into a se... | Animas Technologies, LLC | A61B005/00
US07147602 | B2 | 2006-12-12 | Blood rheology measuring apparatus | 2002-07-08 | A blood rheology measuring apparatus has a measuring portion for measuring blood circulation information inside of a living body from outside of the living body. The blood circulation information is a maximum blood flow velocity for one pulse. An information processing portion processes the measured information from the measuring portion to obtain information concerning a blood rheology of the living body. | 1. A blood rheology measuring apparatus comprising: a measuring portion for measuring blood circulation information inside of a living body from outside of the living body, the blood circulation information being a maximum blood flow velocity for one pulse;an information processing portion for processing the measured information from the measuring portion to obtain information concerning a blood rheology of the living body; anda data inputting portion for inputting individual information of a person to be measured, the individual information being information about a food. | 1. A blood rheology measuring apparatus comprising: a measuring portion for measuring blood circulation information inside of a living body from outside of the living body, the blood circulation information being a maximum blood flow velocity for one pulse;an information processing portion for processing the measured information from the measuring portion to obtain information concerning a blood rheology of the living body; anda data inputting portion for inputting individual information of a person to be measured, the individual information being information about a food. 2. A blood rheology measuring apparatus comprising: a measuring portion for measuring blood circulation information inside of a living body from outside of the living body, the blood circulation information being a maximum blood flow velocity for one pulse; andan information processing portion for processing the measured information from the measuring portion to obtain information concerning a blood rheology of the living body;wherein the measuring portion measures the information by a predetermined timing and when a predetermined value of the blood rheology is reached, the value is detected by the person to be measured. 3. A blood rheology measuring apparatus comprising: a measuring portion for measuring blood circulation information inside of a living body from outside of the living body; andan information processing portion for processing the measured information from the measuring portion to obtain information concerning a blood rheology of the living body;wherein the blood circulation information is a maximum blood flow velocity for one pulse, and the blood rheology is analyzed from a temporal change of the blood flow velocity value by dividing a maximum blood flow velocity of one pulse by an integrated value of a pulse velocity. 4. A blood rheology measuring apparatus according to claim 3; wherein the measuring portion and the information processing portion are integrally or individually made portable to thereby enable the information concerning the blood rheology to be obtained continuously or daily. 5. A blood rheology measuring apparatus according to claim 3; wherein the blood circulation information is a blood flow velocity based on a Doppler effect by a blood flow. 6. A blood rheology measuring apparatus according to claim 3; wherein the measuring portion comprises an ultrasonic wave sensor for transmitting and receiving an ultrasonic wave. 7. A rheology measuring apparatus according to claim 3; further comprising a data inputting portion for inputting individual information of a person to be measured. 8. A blood rheology measuring apparatus according to claim 7; wherein the individual information is information about a food. 9. A blood rheology measuring apparatus according to claim 3; wherein life habit information suitable for a person to be measured is informed to a subject based on information concerning the blood rheology of the person to be measured. ... | Seiko Instruments Inc. | A61B005/026 A61B008/06
US07228315 | B2 | 2007-06-05 | Computer-automated implementation of user-definable decision rules for medical diagnostic or screening interpretations | 2002-03-01 | Software is utilized for allowing a physician to modify the rules of a decision tree so that the individual judgment of the doctor is best utilized. A technician can then carry out tests on a patient, obtain test data in the form of digital data, and apply the test data to the software. The patent test data is then processed through the decision tree to thereby obtain a diagnosis of the patent using the professional judgment without the presence of the physician. | 1. A method of processing data representing results of a diagnostic or screening test, the method comprising the steps of: performing a first procedure to customize decision making criteria, said first procedure including the steps of:defining one or more features characteristic of said data; defining a plurality of decision levels;defining, for each decision level, at least one rule based on at least one of said features; anddefining, for each rule, a criteria; andperforming a second procedure to process said data, said second procedure including the steps of: receiving said data to be processed, said data having values corresponding to said features;comparing said values of said data with said rules according to said criteria; andoutputting a conclusion based on said comparison, said conclusion comprising an indication of the outcome of the diagnostic or screening test. | 1. A method of processing data representing results of a diagnostic or screening test, the method comprising the steps of: performing a first procedure to customize decision making criteria, said first procedure including the steps of:defining one or more features characteristic of said data; defining a plurality of decision levels;defining, for each decision level, at least one rule based on at least one of said features; anddefining, for each rule, a criteria; andperforming a second procedure to process said data, said second procedure including the steps of: receiving said data to be processed, said data having values corresponding to said features;comparing said values of said data with said rules according to said criteria; andoutputting a conclusion based on said comparison, said conclusion comprising an indication of the outcome of the diagnostic or screening test. 2. A method in accordance with claim 1, wherein said first procedure further comprises a step of designating, for each decision level, a quantity of rules effective to influence a result for that level. 3. A method in accordance with claim 1, wherein said first procedure further comprises a step of designating, for each decision level, whether that level represents a condition selected from the group consisting of a technical fail condition, a necessary condition, a sufficient condition, a contributing condition and a refer condition, and wherein said step of comparing said values with said rules according to said criteria includes performing a comparison of said values with said rules according to said criteria for each decision level, andwherein, in said step of outputting a conclusion, the conclusion is output based on a combination of at least a comparison result for said necessary condition and a comparison result for said sufficient condition. 4. A method in accordance with claim 3, wherein, in said step of outputting a conclusion, the conclusion is output based on a combination of at least a comparison result for said necessary condition, a comparison result for said sufficient condition, and a comparison result for said contributing condition. 5. A method in accordance with claim 3, wherein said first procedure further comprises a step of designating, for each decision level, a quantity of rules effective to influence a result for that level. 6. A method in accordance with claim 1, wherein: said step of comparing said values with said rules according to said criteria includes performing a comparison of said values with said rules according to said criteria for each decision level, andin said step of outputting a conclusion, the conclusion is output based on a combination of comparison results for at least two levels. 7. A method in accordance with claim 1, wherein said step of comparing said values with said rules according to said criteria includes performing a comparison of said values with said rules according to said criteria for each decision level. 8. A soft... | - | G06F017/00 A61B005/00 G06F007/00
US07108863 | B2 | 2006-09-19 | Liposome composition for improved intracellular delivery of a therapeutic agent | 2002-03-26 | A liposomal composition and a method of using the same for achieving intracellular delivery of a liposome-entrapped agent is described. The liposomes are composed of a pH sensitive lipid and include a targeting ligand to direct the liposomes to a target cell. The liposomes also include a stabilizing component, such a polymer-derivatized lipid, where the polymer is attached to the lipid by a releasable linkage. Administration of the liposomes results in cellular internalization and destabilization of the liposome for intracellular delivery of the entrapped agent. | 1. A liposome composition for intracellular delivery of a therapeutic agent, comprising liposomes consisting essentially of (i) a pH-sensitive lipid; (ii) between 1–20 mole percent of a lipid derivatized with a hydrophilic polymer, said polymer attached to said lipid by a bond effective to release the hydrophilic polymer in response to an existing or an induced physiologic condition; (iii) a targeting ligand; and (iv) an entrapped therapeutic agent;wherein said composition is adapted to bind to a target cell and to release said entrapped agent to achieve at least a two-fold increase in intracellular concentration of the agent, when compared to intracellular concentration of the agent delivered by similar liposomes lacking said releasable bond and/or said targeting ligand wherein, said releasable bond is a dithiobenzyl bond. | 1. A liposome composition for intracellular delivery of a therapeutic agent, comprising liposomes consisting essentially of (i) a pH-sensitive lipid; (ii) between 1–20 mole percent of a lipid derivatized with a hydrophilic polymer, said polymer attached to said lipid by a bond effective to release the hydrophilic polymer in response to an existing or an induced physiologic condition; (iii) a targeting ligand; and (iv) an entrapped therapeutic agent;wherein said composition is adapted to bind to a target cell and to release said entrapped agent to achieve at least a two-fold increase in intracellular concentration of the agent, when compared to intracellular concentration of the agent delivered by similar liposomes lacking said releasable bond and/or said targeting ligand wherein, said releasable bond is a dithiobenzyl bond. 2. The composition of claim 1, wherein said pH-sensitive lipid is dioleoylphosphatidylethanolmaine (DOPE). 3. The composition of claim 1, wherein said liposomes further include a stabilizing component. 4. The composition of claim 3, wherein said stabilizing component is cholesteryl hemisuccinate (CHEMS). 5. The composition of claim 1, wherein said lipid derivatized polymer is phosphatidylethanolamine derivatized with polyethyleneglycol. 6. The composition of claim 1, wherein said targeting ligand an antibody or an antibody fragment. 7. The composition of claim 6, wherein said antibody fragment is selected from the group consisting of anti-CD19, anti-CD20, anti-CD22. 8. A method for increasing intracellular cytotoxicity of a liposome-entrapped agent, comprising providing liposomes consisting essentially of (i) a pH-sensitive lipid; (ii) between 1–20 mole percent of a lipid derivatized with a hydrophilic polymer, said polymer attached to said lipid by a bond effective to release the hydrophilic polymer in response to an existing or an induced physiologic condition; (iii) a targeting ligand; and (iv) an entrapped therapeutic agent;administering said liposomes to achieve (i) cleavage of the releasable bond, thereby releasing said hydrophilic polymer chain; (ii) binding of said ligand to a target cell, where said binding occurs prior to or subsequent to said cleavage; and (iii) internalization of said liposome by said target cell;whereby said administering is effective to achieve at least a two-fold higher intracellular cytotoxicity of said agent, relative to intracellular concentration of the agent delivered by similar liposomes lacking said releasable bond and/or said targeting ligand wherein, said releasable bond is a dithiobenzyl bond. 9. The method of claim 8, wherein cleavage of said releasable bond after said administering is achieved by naturally-occurring blood components. 10. The method of claim 8, wherein said pH-sensitive lipid is dioleoylphosphatidylethanolmaine (DOPE). 11. The method of claim 8, wherein said liposomes further include a stabilizing component. 12. The method of claim 11, wherein said stabiliz... | Alza Corporation | A61B005/055 A61B008/00 A61K009/127
US06984373 | B2 | 2006-01-10 | Fibrin binding moieties useful as imaging agents | 2001-12-21 | The present invention provides binding moieties for fibrin which have a variety of uses wherever detecting, isolating or localizing fibrin, and particularly fibrin as opposed to fibrinogen, is advantageous. Particularly disclosed are synthetic, isolated polypeptides capable of binding fibrin and recognizing the form of polymerized fibrin found in thrombi. In addition, the polypeptides have a slow dissociation rate from fibrin, which improves their ability to form a contrast image at the site of a fibrin clot, making the disclosed binding moieties particularly useful as imaging agents for thrombi. | 1. An isolated polypeptide having the ability to bind to fibrin, the polypeptide comprising the amino acid sequence: Cys-X2-X3-X4-X5-X6-X7-X8-Cys (SEQ ID NO: 2), wherein X2 is Pro, Arg, Asn, Asp, Gln, Gly, Phe, Ser, Thr, or Tyr; X3 is Ala, Asn, Asp, Gln, Glu, Gly, Ile, Leu, Met, Phe, Pro, Ser, Thr, Trp, Tyr, or Val; X4 is Glu, Gly, Lys, Ser, or Tyr; X5 is Pro, Asp, Glu, Asn, Gln, Gly, Leu, Lys, Ser, Thr, or Tyr; X6 is Arg, Gly, or Trp; X7 is Leu, Ile, Lys, Met, Asn, Gln, Pro, Ser, Thr, or Val; and X8 is Ile, Leu, Phe, Trp, or Tyr. | 1. An isolated polypeptide having the ability to bind to fibrin, the polypeptide comprising the amino acid sequence: Cys-X2-X3-X4-X5-X6-X7-X8-Cys (SEQ ID NO: 2), wherein X2 is Pro, Arg, Asn, Asp, Gln, Gly, Phe, Ser, Thr, or Tyr; X3 is Ala, Asn, Asp, Gln, Glu, Gly, Ile, Leu, Met, Phe, Pro, Ser, Thr, Trp, Tyr, or Val; X4 is Glu, Gly, Lys, Ser, or Tyr; X5 is Pro, Asp, Glu, Asn, Gln, Gly, Leu, Lys, Ser, Thr, or Tyr; X6 is Arg, Gly, or Trp; X7 is Leu, Ile, Lys, Met, Asn, Gln, Pro, Ser, Thr, or Val; and X8 is Ile, Leu, Phe, Trp, or Tyr. 2. The polypeptide according to claim 1, comprising the amino acid sequence: Cys-X2-X3-X4-X5-Trp-X7-X8-Cys (SEQ ID NO: 42), wherein X2 is Pro, Asn, Gln, Ser, or Thr; X3 is Ala, Asn, Asp, Gln, Glu, Gly, Ile, Leu, Met, Phe, Pro, Ser, Thr, Trp, Tyr, or Val; X4 is Glu or Ser; X5 is Pro, Asp, Glu, Asn, Gln, Ser, Thr, or Tyr; X7 is Leu, Ile, Met, Asn, Gln, Ser, Thr, or Val; and X8 is Phe, Trp, or Tyr. 3. The polypeptide according to claim 1, wherein X2 is Pro; X3 is Asp, Glu, Gly, Met, or Trp; X4 is Glu; X5 is Asn, Asp, Glu, Pro, or Ser; X6 is Trp; X7 is Leu or Thr; X8 is Phe. 4. The polypeptide according to claim 1, wherein the polypeptide comprises an amino acid sequence selected from the group consisting of: Cys-Ser-Asp-Glu-Asn-Trp-Leu-Trp-Cys (SEQ ID NO:21); Cys-Pro-Met-Ser-Glu-Trp-Leu-Tyr-Cys (SEQ ID NO: 22); Cys-Pro-Trp-Glu-Ser-Trp-Thr-Phe-Cys (SEQ ID NO: 23); Cys-Gln-Glu-Glu-Pro-Trp-Leu-Phe-Cys (SEQ ID NO: 24); Cys-Pro-Gly-Glu-Asp-Trp-Leu-Phe-Cys (SEQ ID NO: 25); Cys-Tyr-Gly-Glu-Ser-Gly-Ile-Phe-Cys (SEQ ID NO:43); Cys-Thr-Gly-Glu-Pro-Gly-Pro-Ile-Cys (SEQ ID NO:44); Cys-Gln-Leu-Gly-Tyr-Arg-Thr-Tyr-Cys (SEQ ID NO:45); Cys-Asp-Gly-Glu-Pro-Trp-Leu-Phe-Cys (SEQ ID NO:46); Cys-Gly-Trp-Gly-Ser-Trp-Lys-Phe-Cys (SEQ ID NO:47); Cys-Gly-Trp-Gly-Ser-Gly-Lys-Leu-Cys (SEQ ID NO:48); Cys-Pro-Gly-Glu-Pro-Trp-Thr-Phe-Cys (SEQ ID NO:49); Cys-Pro-Gly-Glu-Pro-Trp-Thr-Phe-Cys (SEQ ID NO:50); Cys-Pro-Gly-Tyr-Leu-Arg-Ser-Leu-Cys (SEQ ID NO: 51); Cys-Pro-Gly-Glu-Pro-Trp-Ser-Phe-Cys (SEQ ID NO:52); Cys-Arg-Gly-Glu-Ser-Trp-Pro-Tyr-Cys (SEQ ID NO:53); Cys-Pro-Gly-Tyr-Lys-Arg-Gln-Phe-Cys (SEQ ID NO:54); Cys-Gly-Gln-Glu-Ser-Arg-Thr-Phe-Cys (SEQ ID NO:55); and Cys-Phe-Gln-Lys-Gly-Gly-Thr-Leu-Cys (SEQ ID NO:56). 5. An isolated polypeptide having the ability to bind to fibrin, the polypeptide comprising the amino acid sequence: X1-X2-X3-X4-X5-X6-X7-X8-X9-X10-X11-X12-X13-X14-X15 (SEQ ID NO:1), wherein X1 is Cys, Pro, or Trp; X2 is Ala, Arg, Asn, Asp, Gln, Glu, Gly, His, Ile, Leu, Lys, Met, Phe, Pro, Ser, Thr, Trp, Tyr or Val, or if X4 and X12 are not Cys, then X2 may be Cys; X3 is Ala, Asn, Gln, Gly, Ile, Leu, Met, Phe, Pro, or Thr; X4 is Cys or another amino acid capable of forming a covalent cross-link to X12; X5 is Pro, Arg, Asn, Asp, Gln, Gly, Phe, Ser, Thr or Tyr; X6 is Ala, Asn, Asp, Gln, Glu, Gly, Ile, Leu, Met, Phe, Pro, Ser, Thr, Thr, Tyr, or Val; X7 is Glu, Gly, Lys, Ser, or Tyr; X8 is Pro, Asp, Glu, Asn, Gln, Glu, Gly, Leu, Lys, Ser, T... | Dyax Corp. | A61B005/055 A61B008/00 A61B008/15 A61K049/14 A61K051/08
US06974422 | B1 | 2005-12-13 | Catheter pressure transducer with pressure sensing module | 2002-04-19 | The present invention relates to a catheter pressure transducer having a pressure sensing module operable to provide a signal representative of pressure. The pressure sensor includes a sensing side and a pressure reference side. At least one signal transmission media is coupled to the sensor for transmitting the signal representative of pressure. A convex-shaped carrier is provided having a slot formed to expose the sensing side of the sensor. The convex-shaped carrier is adapted for receiving the sensor such that the sensing side of the sensor is adjacent the slot. A bonding material is provided for mounting the sensor to the carrier. The pressure sensing module provides ease of installation and manufacture because the module may easily be bonded and installed within a catheter. | 1. A catheter pressure transducer for insertion into a living body, said catheter pressure transducer comprising: a catheter body having an internal lumen and an opening formed through sidewalls of the catheter body for providing access to a reference pressure within the internal lumen;a tube having an open channel and an opening formed through sidewalls of the tube for providing access to the open channel, wherein said tube is inserted within and extending along al least a portion of the internal lumen; anda pressure sensor module comprising a carrier having a generally curved surface with an opening formed there through and a pressure sensor coupled to the carrier, wherein a sensing side of the pressure sensor is exposed to an external pressure through the opening formed within the carrier, and wherein the carrier is fitted within the opening formed through the catheter body and coupled at both ends to sidewall surfaces of the same opening. | 1. A catheter pressure transducer for insertion into a living body, said catheter pressure transducer comprising: a catheter body having an internal lumen and an opening formed through sidewalls of the catheter body for providing access to a reference pressure within the internal lumen;a tube having an open channel and an opening formed through sidewalls of the tube for providing access to the open channel, wherein said tube is inserted within and extending along al least a portion of the internal lumen; anda pressure sensor module comprising a carrier having a generally curved surface with an opening formed there through and a pressure sensor coupled to the carrier, wherein a sensing side of the pressure sensor is exposed to an external pressure through the opening formed within the carrier, and wherein the carrier is fitted within the opening formed through the catheter body and coupled at both ends to sidewall surfaces of the same opening. 2. The catheter pressure transducer of claim 1, wherein said tube is constructed of metal. 3. The catheter pressure transducer of claim 1, wherein said tube is constructed of ceramic. 4. The catheter pressure transducer of claim 1, wherein said pressure sensor module is coupled to the sidewall surfaces of the opening formed through said catheter body using epoxy resin. 5. The catheter pressure transducer of claim 1, wherein said pressure sensor module is coupled to the sidewall surfaces of the opening formed through said catheter body using flexible RTV silicone rubber. 6. The catheter pressure transducer of claim 1, wherein said pressure sensor is selected from a group, comprising a resistive-type sensor, a capacitive-type sensor, a photoelectric-type sensor anti a strain-gauge type pressure sensor. 7. The catheter pressure transducer of claim 1, further comprising a first signal transmission media extending along the internal lumen of said catheter. 8. The catheter pressure transducer of claim 7, further comprising a second transmission media coupled to the pressure sensor for transmitting a signal representative of pressure, said second signal transmission media coupled to said first transmission media in order to transmit the signal representative of pressure. 9. The catheter pressure transducer of claim 1, wherein the pressure sensor module is fitted within the opening formed through the catheter body, such that an outer surface of the pressure sensor module is substantially flush with an outer surface of the catheter body. 10. The catheter pressure transducer of claim 1, wherein said pressure sensor is operable to provide a signal that is representative of cardiac pressure. 11. The catheter pressure transducer of claim 1, wherein said pressure sensor is operable to provide a signal that is representative of intracranial pressure. 12. The catheter pressure transducer of claim 1, wherein said pressure sensor is operable to provide a signal that is representative of rectal pressure. 13. The c... | Millar Instruments, Inc. | A61B005/00
US06958045 | B2 | 2005-10-25 | Gait detection system, gait detection apparatus, device, and gait detection method | 2001-12-26 | A gait detection apparatus includes a microphone for picking up low-frequency-band sounds which are transmitted through the body of a pedestrian while walking and an analyzer for performing analysis. Accordingly, the gait of the pedestrian is detected. It is also possible to distinguish the gait pattern on the basis of the stance-phase time of a foot sole, the signal intensity, etc. The gait detection apparatus can accurately estimate the step length on the basis of a detected gait cycle, the height of the pedestrian, and signals detected during walking. On the basis of low-frequency-region sounds picked up by the microphone while the pedestrian is walking, the pedestrian can be identified. | 1. A gait detection system comprising: a microphone for picking up vibrations generated by a pedestrian while walking and for converting the vibrations into electrical signals; analysis means for analyzing variations corresponding to a frequency less than or equal to a predetermined frequency of the electrical signals converted by the microphone, for detecting a gait of the pedestrian, and for generating information on the detected gait; and output means for outputting the information on the detected gait. | 1. A gait detection system comprising: a microphone for picking up vibrations generated by a pedestrian while walking and for converting the vibrations into electrical signals; analysis means for analyzing variations corresponding to a frequency less than or equal to a predetermined frequency of the electrical signals converted by the microphone, for detecting a gait of the pedestrian, and for generating information on the detected gait; and output means for outputting the information on the detected gait. 2. The gait detection system according to claim 1, wherein the analysis means determines a pattern of the gait of the pedestrian based on at least one of a duration and a frequency intensity of a signal in a frequency band less than or equal to 100 Hz. 3. The gait detection system according to claim 1, wherein the analysis means estimates a stride length of the pedestrian based on a gait cycle detected by the analysis means and a pre-input height of the pedestrian. 4. The gait detection system according to claim 3, wherein the analysis means estimates a distance traveled by the pedestrian based on the stride length and a number of steps detected by the analysis means. 5. A gait detection apparatus comprising: analysis means for analyzing a frequency component of a signal based on a vibration transmitted through the body of a pedestrian while walking and for detecting a gait of the pedestrian; and output means for outputting information on the gait detected by the analysis means. 6. The gait detection apparatus according to claim 5, further comprising a filter for passing only a signal in a predetermined frequency band, wherein the analysis means detects the gait of the pedestrian based on a presence or absence of the signal having passed through the filter. 7. The gait detection apparatus according to claim 5, further comprising data storage means for storing signal data corresponding to a gait model pattern, wherein the analysis means analyzes a signal by comparing the signal with the signal data stored in the data storage means and by determining whether a pattern of the signal detected by the analysis means matches the signal data. 8. A device to be mounted on a user, comprising: a microphone for picking up ambient sounds and for converting the ambient sounds into electrical signals: analysis means for analyzing variations in signals corresponding to a frequency less than or equal to a predetermined frequency based on the electrical signals converted by the microphone and for detecting a gait of the user; and display means for outputting information concerning the gait detected by the analysis means using characters. 9. The device according to claim 8, wherein the analysis means detects a number of steps walked by the user. 10. A gait detection method comprising the steps of: detecting components in a frequency band less than or equal to 100 Hz from vibrations transmitted through the body of a pedestrian while walking; and conv... | Sony Corporation | A61B005/11 G01C022/00
US07123762 | B2 | 2006-10-17 | Method and system for risk-modulated diagnosis of disease | 2003-02-10 | A method of calculating a disease assessment by analyzing a medical image, comprising (1) extracting at least one lesion feature value from the medical image; (2) extracting at least one risk feature value from the medical image; and (3) determining the disease assessment based on the at least one lesion feature value and the at least one risk feature value. The method employs lesion characterization for characterizing the lesion, and risk assessment based on the lesion's surroundings, i.e., the environment local and distal to the lesion. Computerized methods both characterize mammographic lesions and assess the breast parenchymal pattern on mammograms, resulting in improved characterization of lesions for specific subpopulations, combining the benefits of both techniques. | 1. A method of calculating a disease assessment by analyzing a medical image, comprising: extracting at least one lesion feature value from the medical image;extracting at least one risk feature value from the medical image; anddetermining the disease assessment based on the at least one lesion feature value and the at least one risk feature value. | 1. A method of calculating a disease assessment by analyzing a medical image, comprising: extracting at least one lesion feature value from the medical image;extracting at least one risk feature value from the medical image; anddetermining the disease assessment based on the at least one lesion feature value and the at least one risk feature value. 2. The method of claim 1, wherein the step of extracting the at least one lesion feature value comprises: locating a lesion in the medical image;segmenting lesion image data corresponding to the located lesion; andextracting the at least one lesion feature value from the segmented lesion image data. 3. The method of claim 2, wherein the step of extracting the at least one lesion feature value comprises: extracting at least one feature value selected from the group consisting of margin sharpness, degree of spiculation, density, homogeneity, texture, asymmetry, shape, and temporal stability of the lesion. 4. The method of claim 2, wherein the determining step comprises: determining, based on the at least one lesion feature value and the at least one risk feature value, at least one of (1) a likelihood that the lesion is malignant, (3) a stage of disease of the lesion, and (4) a likelihood that a future malignancy will develop, as the disease assessment. 5. The method of claim 1, wherein the step of extracting at least one risk feature value comprises: locating a region of interest (ROI) in the medical image corresponding to a parenchymal region; andextracting the at least one risk feature value from the ROI. 6. The method of claim 5, wherein the step of extracting the at least one risk feature value comprises: extracting at least one feature value indicating a maximum gray level of the ROI, a minimum gray level of the ROI, an average gray level of the ROI, a skewness of the ROI, a coarseness of the ROI, a contrast of the ROI, a root mean square variation of a power spectrum of the ROI, and a first moment of the power spectrum of the ROI. 7. The method of claim 1, wherein the determining step comprises: calculating a quantitative measure of malignancy as the disease assessment by applying the at least one lesion feature value and the at least one risk feature value to a classifier. 8. The method of claim 7, wherein the calculating step comprises: calculating the quantitative measure of malignancy as the disease assessment by applying the at least one lesion feature value and the at least one risk feature value to a linear discriminant. 9. The method of claim 7, wherein the calculating step comprises: calculating the quantitative measure of malignancy as the disease assessment by applying the at least one lesion feature value and the at least one risk feature value to an artificial neural network. 10. The method of claim 7, wherein the calculating step comprises: training the classifier in relation to the at least one lesion feature value and the at least one risk feature value obtained from a set... | University of Chicago | G06K009/00 A61B005/00
US07197355 | B2 | 2007-03-27 | Variable-motion optical tomography of small objects | 2002-04-19 | Motion of an object of interest, such as a cell, has a variable velocity that can be varied on a cell-by-cell basis. Cell velocity is controlled in one example by packing cells into a capillary tube, or any other linear substrate that provides optically equivalent 360 degree viewing access, so that the cells are stationary within the capillary tube, but the capillary tube is translated and rotated mechanically through a variable motion optical tomography reconstruction cylinder. The capillary tube motion may advantageously be controlled in a start-and-stop fashion and translated and rotated at any velocity for any motion interval, under the control of a computer program. As such, there are several configurations of the optical tomography system that take advantage of this controlled motion capability. Additionally, the use of polarization filters and phase plates to reduce light scatter and diffraction background noise is described. | 1. A method for variable-motion optical tomography of an object of interest, comprising the steps of: (a) packing a plurality of objects of interest into a tubular container so that the plurality of objects of interest remain stationary within the tubular container, the tubular container having a central axis;(b) illuminating the plurality of objects of interest with at least one optical projection beam;(c) translating the tubular container until at least one object of interest is located within a region of the at least one optical projection beam;(d) centering the at least one object of interest as necessary;(e) rotating the tubular container holding the at least one object of interest through a plurality of radial angles that are perpendicular to the central axis;(f) generating a set of projection images onto at least one detector surface at each radial angle of the plurality of angles, where spaces between the at least one optical projection beam and the at least one object of interest and spaces between the at least one object of interest and the at least one detector surface are filled with a material whose index of refraction is matched with that of the tubular container; and(g) repeating the steps (b) through (f) until the plurality of objects of interest have been radially scanned. | 1. A method for variable-motion optical tomography of an object of interest, comprising the steps of: (a) packing a plurality of objects of interest into a tubular container so that the plurality of objects of interest remain stationary within the tubular container, the tubular container having a central axis;(b) illuminating the plurality of objects of interest with at least one optical projection beam;(c) translating the tubular container until at least one object of interest is located within a region of the at least one optical projection beam;(d) centering the at least one object of interest as necessary;(e) rotating the tubular container holding the at least one object of interest through a plurality of radial angles that are perpendicular to the central axis;(f) generating a set of projection images onto at least one detector surface at each radial angle of the plurality of angles, where spaces between the at least one optical projection beam and the at least one object of interest and spaces between the at least one object of interest and the at least one detector surface are filled with a material whose index of refraction is matched with that of the tubular container; and(g) repeating the steps (b) through (f) until the plurality of objects of interest have been radially scanned. 2. The method of claim 1, wherein the at least one object of interest is a cell or a cell nucleus. 3. The method of claim 1 wherein the optical projection beam is selected from the group consisting of fan beam projections and cone beam projections. 4. A method for three dimensional (3D) reconstruction of an object of interest, the method comprising the steps of: (a) packing a plurality of objects of interest into a tubular container so that the plurality of objects of interest remain stationary within the tube, where an axis of rotation is aligned with a central axis of the tube;(b) illuminating the plurality of objects of interest with a plurality of optical projection beams;(c) translating the linear container until at least one object of interest is located within a region of the plurality of optical projection beams;(d) centering the at least one object of interest as necessary;(c) generating a set of projection images onto at least one detector surface at a plurality of radial angles that are perpendicular to the axis of rotation, where spaces between the plurality of optical projection beams and the at least one object of interest and spaces between the object and the at least one detector surface are filled with a material whose index of refraction is matched with that of the tabular container;(f) repeating the steps (b) through (e) until the objects of interest have been radially scanned; and(g) reconstructing a 3D image from the set of projection images. 5. A variable-motion optical tomography system comprising: a tubular container for holding objects of interest having a tube wall and a tube axis, wherein the objects of interest are packed in the ... | VisionGate, Inc. | A61B005/05
US07273457 | B2 | 2007-09-25 | Barometric pressure correction based on remote sources of information | 2002-05-20 | The invention includes systems and methods for converting absolute pressure data acquired in anatomical environments to gauge pressure data using an implant configured to monitor pressure. The implant is configured to communicate with an external controller, which is configured to communicate with a remote microprocessor that includes real-time barometric pressure data for one or more geographic locations. | 1. A system for measuring pressure in a patient's body, comprising: an implant device configured for measuring absolute pressure in a body, the implant device further configured to communicate measured absolute pressure information outside of the body using telemetric signals; andan external monitor configured to receive telemetric signals from the implant device,receive barometric pressure information from a remote source, the barometric pressure information associated with a geographic location of the body, andderive gauge pressure from the received absolute pressure information and barometric pressure information, factoring in a correction factor based on the actual altitude of the implant device at the geographic location. | 1. A system for measuring pressure in a patient's body, comprising: an implant device configured for measuring absolute pressure in a body, the implant device further configured to communicate measured absolute pressure information outside of the body using telemetric signals; andan external monitor configured to receive telemetric signals from the implant device,receive barometric pressure information from a remote source, the barometric pressure information associated with a geographic location of the body, andderive gauge pressure from the received absolute pressure information and barometric pressure information, factoring in a correction factor based on the actual altitude of the implant device at the geographic location. 2. The system of claim 1, wherein the remote source comprises real-time barometric pressure information for a plurality of geographic locations. 3. The system of claim 1, further comprising a global positioning system (GPS) signal receiver coupled to at least one of the external monitor and the implant device. 4. The system of claim 3, wherein the external monitor is configured to receive position information from the GPS signal receiver and communicate the position information to the remote source, and wherein the received barometric pressure information corresponds to the position information. 5. A system for measuring pressure in a patient's body, comprising: an implant device configured for measuring absolute pressure in a body, the implant device further configured to communicate measured absolute pressure information outside of the body using telemetric signals; andan external monitor configured to receive telemetric signals comprising absolute pressure information from the implant device,transmit received absolute pressure information to a remote source, the remote source comprising real-time barometric pressure information corresponding with a geographic location of the body, andreceive gauge pressure information from the remote source, the gauge pressure information derived from the absolute pressure information and barometric pressure information, factoring in a correction factor based on the actual altitude of the implant device at the geographic location. 6. The system of claim 5, wherein the remote source comprises real-time barometric pressure information for a plurality of geographic locations. 7. A system for measuring pressure in a patient's body, comprising: an external monitor configured to receive barometric pressure information from a remote source, the barometric pressure information associated with a geographic location of the body; andan implant device configured to receive barometric pressure information from the external monitor via acoustic telemetric signals,measure absolute pressure in a body, andderive gauge pressure from the received absolute pressure information and measured barometric pressure information, factoring in a correction factor based on the actual altitude of the implant dev... | Remon Medical Technologies, Ltd. | A61B005/00
US06996431 | B2 | 2006-02-07 | Method for alignment of bone using position sensors | 2002-12-09 | A method of measuring the size of an area (22) within a body, including bringing at least one position sensor (28) to each of a plurality of points in the area, determining position coordinates of the plurality of points, using the sensor (28) and calculating distances between the plurality of points. | 1. A method for alignment of pieces of bone comprising the steps of: fixing a sensor to each piece of bone;producing an image of the pieces of bone;determining the position of each sensor using a position determining system;associating the determined position of each sensor with the image; andaligning the pieces of bone based on the determined position of each sensor associated with the image. | 1. A method for alignment of pieces of bone comprising the steps of: fixing a sensor to each piece of bone;producing an image of the pieces of bone;determining the position of each sensor using a position determining system;associating the determined position of each sensor with the image; andaligning the pieces of bone based on the determined position of each sensor associated with the image. 2. The method according to claim 1, further comprising implanting an object at each piece of bone. 3. The method according to claim 2, further comprising mounting a sensor to the object. 4. The method according to claim 3, wherein the object is a prosthesis. 5. The method according to claim 1, wherein the sensor is fixed to each piece of bone using a connection mechanism. 6. The method according to claim 5, wherein the connection mechanism is a clamp. 7. The method according to claim 5, wherein the connection mechanism is a staple. 8. The method according to claim 5, wherein the connection mechanism is an adhesive. 9. The method according to claim 1, wherein the image is a three dimensional image. 10. The method according to claim 9, wherein the image is obtained using fluoroscopy. 11. The method according to claim 9, wherein the image is obtained using MRI. 12. The method according to claim 9, wherein the image is obtained using X-Ray. 13. The method according to claim 9, wherein the image is obtained using CT. 14. The method according to claim 9, wherein the image is displayed on a screen. 15. The method according to claim 1, wherein the sensor saves as an anchor point. 16. The method according to claim 1, wherein the position determining system uses non-ionizing fields. 17. The method according to claim 16, further comprising using electromagnetic waves to determine the position of each sensor. 18. The method according to claim 17, wherein the sensor comprises at least one coil. 19. The method according to claim 16, further comprising using ultrasound to determine the position of each sensor. 20. The method according to claim 19, wherein the sensor comprises a transponder. 21. The method according to claim 16, further comprising using infrared to determine the position of each sensor. 22. The method according to claim 1, wherein the pieces of bone comprise rigid fragments. 23. The method according to claim 22, wherein the rigid fragments are vertebrae of the spine. 24. The method according to claim 1, further comprising fixing the sensor to each piece of bone by screwing the sensor into the piece of bone. 25. The method according to claim 1, wherein the pieces of bone are pieces of fractured bone. 26. A method for alignment of pieces of bone comprising the steps of: fixing at least one sensor to each piece of bone;determining position coordinates of the at least one sensor on each piece of bone;determining position coordinates of a plurality of points for each piece of bone;determining a shape and position of each piece of bone b... | - | A61B005/00
US07280864 | B2 | 2007-10-09 | Method and apparatus for automated selection of correct image for quantitative analysis | 2002-11-27 | A method for selecting images for coronary analysis using time stamps to correlate cardiovascular images to corresponding physiological or hemodynamic monitoring, e.g., ECG, data. The method uses time stamps that are intrinsic to a distributed network clock synchronization protocol for correlation of images and data. A coronary analysis system is employed to correlate images with physiologic data, e.g., using time stamp data as well as offset data derived using the distributed network clock synchronization protocol. | 1. A method comprising the following steps: (a) acquiring frames of imaging data representing a succession of images of a patient's heart during a study using an imaging system;(b) automatically time stamping each acquired frame with respective times generated by a first clock;(c) automatically calculating respective offsets of said first clock relative to a reference clock for each time stamp associated with said acquired frames;(d) storing said acquired frames and said time stamps and said offsets associated with said acquired frames;(e) acquiring physiologic data from said patient during said study using a hemodynamic monitoring system;(f) automatically time stamping data in said acquired physiologic data representing one or more predetermined cardiac events with respective times generated by a second clock;(g) automatically calculating respective offsets of said second clock relative to said reference clock for each time stamp associated with said physiologic data;(h) storing said acquired physiologic data and said time stamps and said offsets associated with said acquired physiologic data;(i) computing which one of said acquired frames was acquired in substantial synchronism with a first predetermined cardiac event based on said time stamps and said offsets; and(j) associating a first acquired frame identified by step (i) with said first predetermined cardiac event,further comprising the step of displaying said first acquired frame of imaging data concurrently with acquired physiologic data having a visual indicator indicating a physiologic datum corresponding to said first predetermined cardiac event. | 1. A method comprising the following steps: (a) acquiring frames of imaging data representing a succession of images of a patient's heart during a study using an imaging system;(b) automatically time stamping each acquired frame with respective times generated by a first clock;(c) automatically calculating respective offsets of said first clock relative to a reference clock for each time stamp associated with said acquired frames;(d) storing said acquired frames and said time stamps and said offsets associated with said acquired frames;(e) acquiring physiologic data from said patient during said study using a hemodynamic monitoring system;(f) automatically time stamping data in said acquired physiologic data representing one or more predetermined cardiac events with respective times generated by a second clock;(g) automatically calculating respective offsets of said second clock relative to said reference clock for each time stamp associated with said physiologic data;(h) storing said acquired physiologic data and said time stamps and said offsets associated with said acquired physiologic data;(i) computing which one of said acquired frames was acquired in substantial synchronism with a first predetermined cardiac event based on said time stamps and said offsets; and(j) associating a first acquired frame identified by step (i) with said first predetermined cardiac event,further comprising the step of displaying said first acquired frame of imaging data concurrently with acquired physiologic data having a visual indicator indicating a physiologic datum corresponding to said first predetermined cardiac event. 2. A method comprising the following steps: (a) acquiring frames of imaging data representing a succession of images of a patient's heart during a study using an imaging system;(b) automatically time stamping each acquired frame with respective times generated by a first clock;(c) automatically calculating respective offsets of said first clock relative to a reference clock for each time stamp associated with said acquired frames;(d) storing said acquired frames and said time stamps and said offsets associated with said acquired frames;(e) acquiring physiologic data from said patient during said study using a hemodynamic monitoring system;(f) automatically time stamping data in said acquired physiologic data representing one or more predetermined cardiac events with respective times generated by a second clock;(g) automatically calculating respective offsets of said second clock relative to said reference clock for each time stamp associated with said physiologic data;(h) storing said acquired physiologic data and said time stamps and said offsets associated with said acquired physiologic data;(i) computing which one of said acquired frames was acquired in substantial synchronism with a first predetermined cardiac event based on said time stamps and said offsets; and(j) associating a first acquired frame identified by step (i) with said first... | GE Medical Systems Information Technologies, Inc. | A61B005/05
US06889830 | B2 | 2005-05-10 | Device for separating the connecting end of a hypodermic needle from the tip of an injection instrument | 2002-11-07 | The connecting end of the hypodermic needle is so designed that it can be attached to and then disconnected from the tip of an injection instrument by pressure, that is, an axial tractional force exerted between the said needle and the said injection instrument. The extraction device has an opening delimited by locking elements forming one piece with elastic pieces so as to allow its diameter to vary between a minimum diameter and a maximum diameter at least equal to the diameter of the said connecting end and at least one piece associated with the said locking elements for the purpose of converting an axial force exerted on the said piece into at least one radial component capable of being applied to the said elastic pieces so as to deform them radially in order to increase the diameter of the said opening when axial pressure is exerted by the said connecting end. | 1. Device for separating a connecting end of a hypodermic needle from a tip of an injection instrument, the connecting end of the hypodermic needle having connectors complementary to connectors on the tip, one of the connectors having radial elasticity and devices for converting an axial force exerted between the needle and the injection instrument into at least one radial component capable of deforming the connectors, comprising an opening delimited by locking elements forming one piece with elastic pieces to allow its diameter to vary between a minimum diameter and a maximum diameter at least equal to the diameter of the connecting end and at least one piece associated with the locking elements for the purpose of converting an axial force exerted on the one piece with elastic pieces into at least one radial component capable of being applied to the elastic pieces to deform them radially so as to increase the diameter of the opening when the connecting end is displaced along its axis with its needle pointing forward through the opening and to restore its initial diameter after the connecting end has passed through it and allow the locking elements to engage with the rear face of the connecting end, wherein the device has a guiding surface concentric with the opening, wherein the device has a cylindrical housing divided longitudinally into two parts whose respective lengths are approximately equal, matching the length of the connecting end, by the opening delimited by the locking elements which form one piece with elastic arms attached to one end of the cylindrical housing which communicates with the exterior at its other end through an access opening having a diameter approximately equal to that of the connecting end of the hypodermic needle, a ring, having an internal diameter approximately equal to that of the access opening, positioned to slide in the cylindrical housing, being capable of occupying two limiting positions within the housing, in one of which positions it is adjacent to the access opening, the ring having passages in such an arrangement and of such dimensions as to allow the locking elements to project on the inner face of this ring when the latter is adjacent to the access opening, while in its other limiting position it holds the locking elements forming one piece with the elastic arms back from the connecting end to which it adheres by friction. | 1. Device for separating a connecting end of a hypodermic needle from a tip of an injection instrument, the connecting end of the hypodermic needle having connectors complementary to connectors on the tip, one of the connectors having radial elasticity and devices for converting an axial force exerted between the needle and the injection instrument into at least one radial component capable of deforming the connectors, comprising an opening delimited by locking elements forming one piece with elastic pieces to allow its diameter to vary between a minimum diameter and a maximum diameter at least equal to the diameter of the connecting end and at least one piece associated with the locking elements for the purpose of converting an axial force exerted on the one piece with elastic pieces into at least one radial component capable of being applied to the elastic pieces to deform them radially so as to increase the diameter of the opening when the connecting end is displaced along its axis with its needle pointing forward through the opening and to restore its initial diameter after the connecting end has passed through it and allow the locking elements to engage with the rear face of the connecting end, wherein the device has a guiding surface concentric with the opening, wherein the device has a cylindrical housing divided longitudinally into two parts whose respective lengths are approximately equal, matching the length of the connecting end, by the opening delimited by the locking elements which form one piece with elastic arms attached to one end of the cylindrical housing which communicates with the exterior at its other end through an access opening having a diameter approximately equal to that of the connecting end of the hypodermic needle, a ring, having an internal diameter approximately equal to that of the access opening, positioned to slide in the cylindrical housing, being capable of occupying two limiting positions within the housing, in one of which positions it is adjacent to the access opening, the ring having passages in such an arrangement and of such dimensions as to allow the locking elements to project on the inner face of this ring when the latter is adjacent to the access opening, while in its other limiting position it holds the locking elements forming one piece with the elastic arms back from the connecting end to which it adheres by friction. 2. Device according to claim 1, wherein the said opening serves as the access opening in a closed housing. 3. Device according to claim 1, wherein one rim of the said guiding surface is linked by a horizontal surface to a storage container for a hypodermic needle. 4. Device according to claim 1, wherein it has the form of a tubular body in which a number of hypodermic needles are successively placed in corresponding positions so as to be accessible to the injection instrument from one end of the tubular body, while the other end has the radially elastic locking elements to allow it... | Ares Trading S.A. | B65D083/10 A61B005/00 A61M005/32
US06895267 | B2 | 2005-05-17 | Systems and methods for guiding and locating functional elements on medical devices positioned in a body | 2001-10-24 | A system for performing an invasive diagnostic or therapeutic medical procedure includes a first probe having a distal portion carrying a plurality of functional elements and a second probe having a distal portion carrying a functional element and a location sensor. The system includes a proximity processing subsystem in communication with the functional elements carried by each of the first and second probes and configured to determine a proximity of the functional element carried by the second probe relative to a first functional element carried by the first probe. The systems further includes a location processing subsystem configured to determine an absolute position and orientation of the location sensor carried by the second probe with respect to a three-dimensional reference coordinate system that can be internal or external to the three-dimensional space in which the first and second probes are positioned. | 1. A system for performing a medical procedure, comprising: a first probe having a distal portion carrying a first set of functional elements comprising a first functional element; a second probe having a distal portion carrying a second set of functional elements comprising a second functional element; a proximity determination subsystem in communication with the first and second sets of functional elements, and configured to determine a proximity of the first functional element relative to the second functional element; and a location determination subsystem in communication with the second set of functional elements and configured to determine an absolute position of the second functional element in a three-dimensional coordinate system. | 1. A system for performing a medical procedure, comprising: a first probe having a distal portion carrying a first set of functional elements comprising a first functional element; a second probe having a distal portion carrying a second set of functional elements comprising a second functional element; a proximity determination subsystem in communication with the first and second sets of functional elements, and configured to determine a proximity of the first functional element relative to the second functional element; and a location determination subsystem in communication with the second set of functional elements and configured to determine an absolute position of the second functional element in a three-dimensional coordinate system. 2. The system of claim 1, wherein the first set of functional elements comprises a plurality of functional elements, and the proximity determination subsystem is configured to determine a proximity of each of the plurality of functional elements relative to the second functional element. 3. The system of claim 1, wherein the second set of functional elements comprises only the second functional element, and both the proximity determination subsystem and location determination subsystem are in communication with the second functional element. 4. The system of claim 1, wherein the second functional element comprises a proximity element, the second set of functional elements further comprises a location element, the proximity determination subsystem is in communication with the first functional element and the proximity element and is configured to determine a proximity of the first functional element relative to the proximity element, the location determination subsystem is in communication with the location element and is configured to determine the absolute location of the location element. 5. The system of claim 1, wherein the distal portion of the second probe carries a third functional element that is offset from the second functional element a known distance, and the location determination subsystem is configured to determine an absolute location of the third functional element in the coordinate system based on the absolute location of the second functional element and the known offset distance between the second and third functional elements. 6. The system of claim 5, wherein the third functional element is an ablation electrode. 7. The system of claim 1, wherein the second functional element is an ablation electrode. 8. The system of claim 1, further comprising a location-proximity association subsystem configured to determine an absolute location of the first functional element in the coordinate system based on the absolute location of the second functional element and the proximity of the second functional element to the first functional element. 9. The system of claim 8, wherein the location-proximity association subsystem is configured to determine the absolute location of the first function... | SciMed Life Systems, Inc. | A61B005/05
US06915156 | B2 | 2005-07-05 | Intracardiac detection and control of repolarization alternans | 2001-10-23 | Nonlinear-dynamical control, also known as chaos control, can modulate human cardiac electrophysiological dynamics by stabilizing an unstable target rhythm. Intracardiac, repolarization alternans detection methods are disclosed that operates on the basis of amplitude or repolarization duration detection. Also disclosed are control methods for stabilizing repolarization alternans. | 1. A method for intracardiac detection of repolarization alternans, comprising the steps of: a) placing an electrode within the body and within the vicinity of the heart; b) sensing at the location of the electrode the electrical voltage amplitude at a predetermined time relative to a fiducial point in the heart rhythm of a series of beats, the series of beats including a predetermined number of pairs of adjacent beats, each of the pairs of adjacent beats including an even numbered beat and an odd numbered beat; c) subtracting the voltage amplitude of the even numbered beat from the odd numbered beat in the adjacent beats, respectively, to obtain for each pair of adjacent beats a sign of the difference in voltage amplitude; and d) indicating the presence of repolarization alternans if the sign obtained in step (c) for a predetermined number of adjacent pairs of beats is consistent. | 1. A method for intracardiac detection of repolarization alternans, comprising the steps of: a) placing an electrode within the body and within the vicinity of the heart; b) sensing at the location of the electrode the electrical voltage amplitude at a predetermined time relative to a fiducial point in the heart rhythm of a series of beats, the series of beats including a predetermined number of pairs of adjacent beats, each of the pairs of adjacent beats including an even numbered beat and an odd numbered beat; c) subtracting the voltage amplitude of the even numbered beat from the odd numbered beat in the adjacent beats, respectively, to obtain for each pair of adjacent beats a sign of the difference in voltage amplitude; and d) indicating the presence of repolarization alternans if the sign obtained in step (c) for a predetermined number of adjacent pairs of beats is consistent. 2. The method as in claim 1, wherein the electrode placement step comprises contacting the electrode to cardiac tissue. 3. The method as in claim 1, wherein the electrode placement step comprises implanting the electrode into cardiac tissue. 4. The method as in claim 1, wherein the predetermined time relative to a fiducial point occurs during the ventricular repolarization phase of the cardiac cycle. 5. The method as in claim 1, wherein steps (b) through (d) are repeated multiple different predetermined times relative to the fiducial point in the heart rhythm, whereby the detection method indicates the presence of repolarization alternans at any of a number of different parts of the cardiac beat cycle corresponding to said the multiple different predetermined times. 6. A method for stabilizing repolarization alternans, comprising the steps of: a) contacting at least one electrode to cardiac tissue; b) sensing at the location of one of the electrodes the electrical voltage amplitude at a predetermined time relative to a fiducial point in the heart rhythm of a series of beats, the series of beats including a predetermined number of pairs of adjacent beats, each of the pairs of adjacent beats including an even numbered beat and an odd numbered beat; c) subtracting the voltage amplitude of the even numbered beat from the odd numbered beat in at least one pair of adjacent beats to obtain for each said pair of adjacent beats a magnitude of the difference in voltage amplitude and a sign of said difference; d) dynamically defining an electrical stimuli for delivery to the cardiac tissue, the electrical stimuli including a timing and an amplitude selected in response to the values obtained for the difference in voltage amplitude and the sign of said difference; e) conditionally delivering the electrical stimuli to the cardiac tissue if the obtained sign of a predetermined number of adjacent pairs of beats is consistent. 7. The method as in claim 6, including the additional step of repeating steps (b) through (e). 8. The method as in claim 6, including the additional st... | Cornell Research Foundation, Inc. | A61B005/04
US06973338 | B2 | 2005-12-06 | Conjunctival monitor | 2002-12-09 | A method for monitoring blood/tissue gases, blood pH, and other elements of blood and tissue chemistry is described that uses a sensor tip that is placed adjacent to the conjunctiva of a patient to be monitored. Monitors and sensor using optical fiber sensors for performing the monitoring are also described. | 1. A method for monitoring blood and tissue gas or pH of a patient, comprising placing an optical blood and tissue gas monitor sensor tip adjacent to the palpebral or bulbar conjunctiva of said patient, wherein said sensor tip comprises at least one optical sensor fiber for detecting at least one blood and tissue gas or blood and tissue pH, and said sensor tip is connected with analysis and display components of a blood monitor system; andsaid sensor tip is adapted and configured to be held in place adjacent the palpebral or bulbar conjunctiva of the eye;determining patient blood and tissue gas or pH values or both with said blood monitor system; andrecording or displaying said patient blood and tissue gas or pH values or both. | 1. A method for monitoring blood and tissue gas or pH of a patient, comprising placing an optical blood and tissue gas monitor sensor tip adjacent to the palpebral or bulbar conjunctiva of said patient, wherein said sensor tip comprises at least one optical sensor fiber for detecting at least one blood and tissue gas or blood and tissue pH, and said sensor tip is connected with analysis and display components of a blood monitor system; andsaid sensor tip is adapted and configured to be held in place adjacent the palpebral or bulbar conjunctiva of the eye;determining patient blood and tissue gas or pH values or both with said blood monitor system; andrecording or displaying said patient blood and tissue gas or pH values or both. 2. The method of claim 1, wherein said sensor tip has a detection length of 2 cm or less. 3. The method of claim 1, wherein said sensor tip further comprises an exterior biocompatible, microporous membrane containing at least two of a pH optical sensor fiber, a PCO2 optical sensor fiber, and a PO2 optical sensor fiber, wherein fluid flow trough said microporous membrane is inhibited. 4. The method of claim 3, wherein said sensor tip further comprises at least one distal terminal mirror that returns light through one or more of said optical sensor fibers. 5. The method of claim 4, wherein said sensor tip is 1.5 cm or less in length. 6. The method of claim 4, wherein said sensor tip is 1.0 cm or less in length. 7. The method of claim 3, wherein said fluid flow is inhibited by a gel. 8. The method of claim 7, wherein said gel comprises polyacrylamide gel. 9. The method of claim 3, wherein said sensor tip is coated with heparin. 10. The method of claim 3, wherein said optical sensor fibers comprise acrylic optical fibers. 11. The method of claim 1, wherein said monitoring is conducted during treatment of said patient. 12. The method of claim 11, wherein said treatment comprises surgery. 13. The method of claim 12, wherein said surgery comprises heart surgery. 14. The method of claim 12, wherein said surgery comprises lung surgery. 15. The method of claim 11, wherein said treatment comprises injection or infusion of a drug. 16. The method of claim 11, wherein said monitoring is performed for a period of at least one hour. 17. The method of claim 1, wherein said sensor tip further comprises a pH optical sensor fiber, a PCO2 optical sensor fiber, and a PO2 optical sensor fiber. 18. The method of claim 17, wherein said sensor tip further comprises a temperature sensor. 19. The method of claim 18, wherein said temperature sensor comprises a thermocouple. 20. The method of claim 19, wherein temperature readings from said temperature sensor are used to normalize readings for one or more of pH, CO2, and O2. 21. The method of claim 17, wherein said pH optical sensor fiber comprises phenol red in polyacrylamide gel, and said fiber is perforated. 22. The method of claim 17, wherein said PCO2 optical sensor fiber comp... | Los Angeles Biomedical Research Institute at Harbor-UCLA Medical Center | A61B005/00
US07209778 | B2 | 2007-04-24 | Embedded thermal control system for high field MR scanners | 2002-09-27 | A thermal control system having an initial bore condition and a thermal boundary condition for an MRI includes a thermal controller adapted to set at least one dynamic limit on power input into an MRI gradient coil. The dynamic limit is determined for a first commanded prescription as a function of the initial bore condition and the thermal boundary condition. | 1. A thermal control system having an initial bore condition and a thermal boundary condition for an MRI comprising: a thermal controller adapted to set at least one dynamic limit on a power input into at least one gradient coil of the MRI while maintaining at least some power to said at least one gradient coil said at least one dynamic limit being a function of the initial bore condition and the thermal boundary condition. | 1. A thermal control system having an initial bore condition and a thermal boundary condition for an MRI comprising: a thermal controller adapted to set at least one dynamic limit on a power input into at least one gradient coil of the MRI while maintaining at least some power to said at least one gradient coil said at least one dynamic limit being a function of the initial bore condition and the thermal boundary condition. 2. The system of claim 1 wherein at least one temperature sensor is adapted to sense the initial bore condition and therefrom generate an initial condition signal and said at least one temperature sensor further adapted to sense the thermal boundary condition and therefrom generate a thermal boundary signal. 3. The system of claim 2 wherein said at least one temperature sensor comprises a first temperature sensor adapted to sense the initial condition and therefrom generate said initial condition signal and a second temperature sensor adapted to sense the thermal boundary and therefrom generate said thermal boundary signal. 4. The system of claim 1 wherein said thermal controller sets said at least one dynamic limit by receiving at least one gradient field; converting said at least one gradient field command into a power level signal;generating a model of said power level signal;generating a model change in a gradient temperature from said model of said power level signal;generating an initial bore condition signal from a temperature sensor coupled to the MRI;generating a gradient temperature signal from a sum of said initial bore condition signal and said model change in said gradient temperature;generating a change in bore temperature signal through summing said model of said power level signal and a model of said current bore temperature generated from said gradient temperature signal summed with said bore temperature signal;summing a second boundary condition with said change in bore temperature to generate said current bore temperature; andadjusting a gradient field. 5. The system of claim 1 wherein said thermal controller is adapted to set said at least one dynamic limit through a thermal predictor software module, which includes a computational algorithm for modeling a series of coupled first order dynamic subsystems designed to simulate at least one actual thermal characteristic of the MRI. 6. The system of claim 1 further comprising a gradient coil control actuator adapted to adjust a gradient field of said at least one gradient coil in response to an activation signal from said thermal controller generated in response to said dynamic limit. 7. A method for controlling thermal behavior in an MRI system comprising: inputting at least one gradient field command into an embedded thermal controller;converting said at least one gradient field command into a power level signal;generating a model of said power level signal;generating a model change in a gradient temperature from said model of said power level signal;ge... | General Electric Company | A61B005/05
US07297109 | B2 | 2007-11-20 | Method and system for improving adherence with a diet program or other medical regimen | 2002-12-12 | A method, system and computer program product for remotely measuring one's adherence to a diet program while making it convenient and easy to place food orders, and learn about and try new foods that are acceptable to the diet program. The system includes a user system with a processor, memory and at least one user interface device, a server system with a processor and memory, and a food delivery system coupled to the server system over the network. | 1. A method comprising: generating a query at a server system; sending the generated query to a user system coupled to the server system over a network; responding to the query; sending the responses to the query from the user system to the server system; determining a user's diet program information based upon the query responses, user profile information and a signal from a medical monitoring device measuring a medical condition of a user; generating a food delivery request at a food delivery system coupled to the server system over the network according to the determined diet program information; preparing a food order according to the generated food delivery request; and delivering the food order to the user; and sending a record of the food order to the server system. | 1. A method comprising: generating a query at a server system; sending the generated query to a user system coupled to the server system over a network; responding to the query; sending the responses to the query from the user system to the server system; determining a user's diet program information based upon the query responses, user profile information and a signal from a medical monitoring device measuring a medical condition of a user; generating a food delivery request at a food delivery system coupled to the server system over the network according to the determined diet program information; preparing a food order according to the generated food delivery request; and delivering the food order to the user; and sending a record of the food order to the server system. 2. The method of claim 1, further comprising storing the record of the food order at the server system. 3. The method of claim 1, further comprising suggesting one or more alternate food items for one or more of the selected food items. 4. The method of claim 1, entering point values for food items consumed by the user, sending the entered point values to the server system, comparing at the server system the received point values to predetermined meal plan limits and creating a meal plan according to the comparison. 5. The method of claim 1, further comprising scanning a bar code present on packages of user consumed food items, sending the scanned bar code information from the user system to the server system, automatically ordering refills according to program information, recording the scanned food item as being consumed, and updating the program information according to the recorded scanned food item. 6. The method of claim 1, further comprising generating a program adherence value based on the program information, any predefined program goals and the responses to the query. 7. A system comprising: a user system comprising a processor, memory and at least one medical monitoring device, the medical monitor device configured to send a signal measuring a medical condition of a user to the processor; a server system comprising a processor and memory, the server system being coupled to the user system over a network; and a food delivery system coupled to the server system over the network, wherein the server system generates and sends a query to the user system, the server system determines diet program information for the user associated with the user system based upon responses to the query sent from the user system to the server system, the user profile information, the signal measuring the medical condition of the user, and generates a food delivery request according to the determined diet program information, and wherein the food delivery system prepares a food order according to a food list, delivers the food order to the user, and sends a record of the food order to the server system. 8. The system of claim 7, wherein the medical monitoring device includes a weigh... | Health Hero Network, Inc. | A61B005/00
US07192405 | B2 | 2007-03-20 | Integrated lancet and bodily fluid sensor | 2002-09-30 | A lancing device integrated with a sensing strip and method of making and using the same are provided. A lance and analyte sensing strip are combined in a single, preferably disposable, device. In a preferred embodiment, a slot is provided in a support member so the lance can be engaged by a lance driver and driven into the skin producing a bodily fluid sample from a wound. The sensing area of the sensing strip is adjacent to the wound location reducing dexterity requirements for patients. | 1. A bodily fluid sample acquisition and testing device comprising: a lancet adapted to be engaged by a lancing device;a bodily fluid chemistry sensing strip comprising a sensing area; anda primary support member adapted to fixedly engage said sensing strip, and to slidably engage said lancet, said primary support member further comprising a slot adapted to allow contact between said lancet and said lancing device, said lancing device being adapted to slide said lancet in the direction of said slot between an unexposed position and an exposed position. | 1. A bodily fluid sample acquisition and testing device comprising: a lancet adapted to be engaged by a lancing device;a bodily fluid chemistry sensing strip comprising a sensing area; anda primary support member adapted to fixedly engage said sensing strip, and to slidably engage said lancet, said primary support member further comprising a slot adapted to allow contact between said lancet and said lancing device, said lancing device being adapted to slide said lancet in the direction of said slot between an unexposed position and an exposed position. 2. The bodily fluid sample acquisition and testing device of claim 1, wherein said sensing area is in electrical contact with two electrodes. 3. The bodily fluid sample acquisition and testing device of claim 1, wherein said sensing area is adjacent to said exposed position of said lancet. 4. A bodily fluid sample acquisition and testing device as in claim 1, wherein said sensing strip comprises a wick. 5. A bodily fluid sample acquisition and testing device as in claim 1, wherein said sensing strip is adapted to sense bodily fluid having a volume less than one microliter. 6. A bodily fluid sample acquisition and testing device as in claim 1, wherein said sensing strip is adapted to sense bodily fluid having a volume less than one half microliter. 7. A bodily fluid sample acquisition and testing device comprising: a lancet adapted to be engaged by a lancing device; anda primary support member adapted to fixedly engage a bodily fluid chemistry sensing strip, and to slidably engage said lancet, said primary support member further comprising a slot adapted to allow contact between said lancet and said lancing device, said lancing device adapted to slide said lancet between an unexposed position and an exposed position. 8. A bodily fluid sample acquisition and testing device comprising: a primary support member having a bodily fluid chemistry sensing strip attached thereto, said sensing strip comprising a sensing area in proximity to a sensing end of said primary support member;a cantilever member comprising a first end hinged to said support member;a lancing member comprising a lancet support member and a lancet having a pointed end attached to said lancet support member, said lancet support member being hinged to a second end of said cantilever member;said primary support member, cantilever member, and lancing member adapted to adopt a first orientation being substantially coplanar, and to adopt a second orientation with said lancing member defining an angle with said primary support member such that said pointed end of said lancet is adjacent to said sensing area of said sensing strip. 9. The bodily fluid sample acquisition and testing device of claim 8, wherein said sensing area is in electrical communication with two electrodes. 10. The bodily fluid sample acquisition and testing device of claim 8, wherein said primary support member has a substantially flat orientation. 11. The bodily f... | Becton, Dickinson and Company | A61B005/00
US07313434 | B2 | 2007-12-25 | Impedance monitoring for detecting pulmonary edema and thoracic congestion | 2002-11-25 | An implantable medical device includes a housing for the implantable device that is sized for implantation in a chest region of a patient and includes a housing electrode. A left ventricular lead port is connectable to a proximal end of a lead having first and second insulated conductors that extend from the proximal end of the lead to corresponding first and second electrodes near a distal end of the lead. An electrical impedance measurement circuit is electrically connected to the left ventricular lead port and the housing electrode. The circuit may inject current between the first lead electrode and the housing electrode, and measure a voltage potential, created by the current injection, between the second lead electrode and the housing electrode. The ratio of the measured voltage to injected current may be used for assessing pulmonary edema. The lead electrodes may be located in a left ventricular coronary vein. | 1. A method of monitoring pulmonary edema in a human being, the method comprising: injecting an electrical current between a first electrode located in a coronary vein of a left ventricle and a housing of a medical device implanted in a chest region;measuring a voltage potential between a second electrode located in the coronary vein of the left ventricle and the medical device housing; andassessing pulmonary edema based on an impedance value calculated from the electrical current and the voltage potential and a stored edema threshold impedance value. | 1. A method of monitoring pulmonary edema in a human being, the method comprising: injecting an electrical current between a first electrode located in a coronary vein of a left ventricle and a housing of a medical device implanted in a chest region;measuring a voltage potential between a second electrode located in the coronary vein of the left ventricle and the medical device housing; andassessing pulmonary edema based on an impedance value calculated from the electrical current and the voltage potential and a stored edema threshold impedance value. 2. The method of claim 1 further comprising calculating an impedance value by taking a ratio of the measured voltage potential to the injected current. 3. A method of monitoring pulmonary edema in a human being, the method comprising: injecting a first electrical current between a first electrode located in a coronary vein of a left ventricle and a housing of a medical device implanted in a chest region;measuring a first voltage potential between a second electrode located in the coronary vein of the left ventricle and the medical device housing, the first voltage potential created by the first current injection;injecting a second electrical current between a third electrode located in a right atrium and the medical device housing;measuring a second voltage potential between a fourth electrode located in the right atrium and the medical device housing, the second voltage potential created by the second current injection; andassessing pulmonary edema based on an impedance value calculated from the electrical currents and the voltage potentials and a stored edema threshold impedance value. 4. The method of claim 3 further comprising: calculating a first impedance value by taking a first ratio of the measured first voltage potential to the injected first current; and calculating a second impedance value by taking a second ratio of the measured second voltage potential to the injected second current. 5. The method of claim 4 wherein a weighted combination of the first and second impedance values is used to assess pulmonary edema. | Regents of the University of Minnesota | A61B005/05
US07153277 | B2 | 2006-12-26 | Composite medical device with markers | 2002-12-03 | Medical devices and methods of manufacturing medical devices. The medical devices include an elongate member having a proximal portion and a distal portion and a connector assembly disposed adjacent the elongated member to connect the proximal and distal portions, the connector assembly including a radiopaque marker. | 1. A medical device comprising: a first elongated member having a proximal portion and a distal portion;a second elongated member having a proximal portion and a distal portion;wherein the first elongated member comprises a first material and the second elongated member comprises a second material different from the first material;wherein the second material comprises a nickel-titanium alloy;a tubular connector disposed adjacent the distal portion of the first elongated member and adjacent the proximal portion of the second elongated member to connect the first and second elongated members, wherein the connector has a proximal-most end and a distal-most end; andone or more non-coiled radiopaque marker bands disposed adjacent the connector for imaging the device, wherein at least one of the one or more non-coiled radiopaque marker bands contacts either the proximal-most end or the distal-most end of the connector. | 1. A medical device comprising: a first elongated member having a proximal portion and a distal portion;a second elongated member having a proximal portion and a distal portion;wherein the first elongated member comprises a first material and the second elongated member comprises a second material different from the first material;wherein the second material comprises a nickel-titanium alloy;a tubular connector disposed adjacent the distal portion of the first elongated member and adjacent the proximal portion of the second elongated member to connect the first and second elongated members, wherein the connector has a proximal-most end and a distal-most end; andone or more non-coiled radiopaque marker bands disposed adjacent the connector for imaging the device, wherein at least one of the one or more non-coiled radiopaque marker bands contacts either the proximal-most end or the distal-most end of the connector. 2. The medical device of claim 1, wherein the connector is disposed about the distal portion of the first elongated member and about the proximal portion of the second elongated member. 3. The medical device of claim 1, wherein at least one of the one or more non-coiled radiopaque marker bands is disposed about the distal portion of the first elongated member adjacent the connector. 4. The medical device of claim 1, wherein at least one of the one or more non-coiled radiopaque marker bands is disposed about the proximal portion of the second elongated member adjacent the connector. 5. The medical device of claim 1, wherein the device is a guidewire. 6. The medical device of claim 1, wherein the first material comprises a stainless steel alloy. 7. The medical device of claim 1, wherein two non-coiled radiopaque marker bands are disposed adjacent the connector for imaging the device. 8. The medical device of claim 7, further comprising a third radiopaque marker disposed adjacent the connector for imaging the device. 9. The medical device of claim 8, wherein the connector is disposed about the distal portion of the first elongated member and about the proximal portion of the second elongated member, and wherein the third radiopaque marker is disposed within the tubular connector between the distal portion of the first elongated member and the proximal portion of the second elongated member. 10. A medical device, comprising: a first elongated member having a proximal portion and a distal portion;a second elongated member having a proximal portion and a distal portion;wherein the first elongated member comprises a stainless steel alloy, and the second elongated member comprises a nickel-titanium alloy; anda tubular connector assembly including a non-coiled radiopaque marker band that is in contact with either a proximal-most end or a distal-most end of the connector assembly, the connector assembly disposed adjacent the distal portion of the first elongated member and adjacent the proximal portion of the second elongate member to con... | Scimed Life Systems, Inc. | A61B005/00 A61M025/00
US07177674 | B2 | 2007-02-13 | Patient-specific parameter selection for neurological event detection | 2002-05-13 | An epileptiform activity patient-specific template creation system permits a user to efficiently develop an optimized set of patient-specific parameters for epileptiform activity detection algorithms. The epileptiform activity patient template creation system is primarily directed for use with an implantable neurostimulator system having EEG storage capability, in conjunction with a computer software program operating within a computer workstation having a processor, disk storage and input/output facilities for storing, processing and displaying patient EEG signals. The implantable neurostimulator is operative to store records of EEG data when neurological events are detected, when it receives external commands to record, or at preset or arbitrary times. The computer workstation operates on stored and uploaded records of EEG data to derive the patient-specific templates via a single local minimum variant of a multidimensional greedy line search process and a feature overlay process. | 1. A method for creating a patient-specific template from at least one record of electrographic signal data stored by an implantable device, the method comprising the steps of: uploading the at least one record of electrographic signal data from the implantable device to an external apparatus; processing the at least one record to generate at least one parameter; and forming the patient-specific template from at the at least one parameter by performing a single local minimum variant of a greedy line search. | 1. A method for creating a patient-specific template from at least one record of electrographic signal data stored by an implantable device, the method comprising the steps of: uploading the at least one record of electrographic signal data from the implantable device to an external apparatus; processing the at least one record to generate at least one parameter; and forming the patient-specific template from at the at least one parameter by performing a single local minimum variant of a greedy line search. 2. The method for developing a patient-specific template of claim 1, wherein performing a single local minimum variant of a greedy line search comprises the steps of: selecting a variable parameter from a working parameter set; adjusting a value of the variable parameter in a first direction and checking a result value until the result value increases or becomes incomputable, or an end of range is reached; determining whether a local minimum has been identifield, and if not, adjusting a value of the variable parameter in a second direction and checking a result value until the result value increases or becomes incomputable, or an end of range is reached; and identifying a desired value of the variable parameter at which the result value is at a minimum. 3. The method for developing a patient-specific template of claim 2, further comprising the step of selecting the at least one test record as a subset of the at least one record of electrographic signal data. 4. The method for developing a patient-specific template of claim 2, further comprising the step of incorporating the desired value of the variable parameter into the working parameter set. 5. The method for developing a patient-specific template of claim 2, further comprising the steps of: selecting an initial parameter set; and deriving the working parameter set from the initial parameter set. 6. The method for developing a patient-specific template of claim 5, wherein selecting an initial parameter set comprises the steps of: employing at least one heuristic algorithm to generate at least one initial parameter; and forming an initial parameter set from the at least one initial parameter. 7. The method or developing a patient-specific template of claim 2, further comprising the step of selecting a termination criterion. 8. The method for developing a patient-specific template of claim 7, further comprising the step of comparing the working parameter set to a stored parameter set. 9. The mothod for developing a patient-specific template of claim 8, wherein the step of comparing the working parameter set to a stored parameter set comprises the steps of: identifying a plurality of performance factors on the at least one test record; combining the plurality of performance factors into a working performance metric; and comparing the working performance metric to a stored performance metric corresponding to the stored parameter set. 10. The method for developing a patient-specifi... | - | A61B005/0476
US07062308 | B1 | 2006-06-13 | Remote physiological monitoring with the reticulum of livestock | 2002-07-03 | A system for monitoring physiological parameters within the stomachs of ruminant animals and including two primary interdependent components. Both components are linked by radio frequency communication and software designed specifically for this application. The first component is a sensing transceiver, battery power and electric components hereafter referred to as the device that is tethered by a cord to a weight in the form of a conventional cow magnet. The second component includes both stationary and portable readers/transmitters capable of initiating radio frequency communication with the device(s) for receiving data from them and sending data to them. Software residing in both components utilizes digitized data and operating commands for transmission between the components, for storing selected data and program commands and manipulating the data for viewing, hearing and related analysis. | 1. A device for monitoring at least one physiological parameter from within the stomach of ruminant animals comprising: a) a sensor component for obtaining physiological data and providing the data in a form accessible from an external location by radio frequency communication;b) a weight component of a size and shape for introduction to the stomach of a ruminant animal, the weight component comprising a cow magnet; andc) a cord connecting the sensor component and the weight component to each other;d) so that as the weight component settles to a location in the stomach the cord tethers the sensor component to float in spaced relation to the weight component. | 1. A device for monitoring at least one physiological parameter from within the stomach of ruminant animals comprising: a) a sensor component for obtaining physiological data and providing the data in a form accessible from an external location by radio frequency communication;b) a weight component of a size and shape for introduction to the stomach of a ruminant animal, the weight component comprising a cow magnet; andc) a cord connecting the sensor component and the weight component to each other;d) so that as the weight component settles to a location in the stomach the cord tethers the sensor component to float in spaced relation to the weight component. 2. The device according to claim 1, wherein the sensor component comprises a pH sensor. 3. The device according to claim 1, wherein the sensor component comprises a temperature sensor. 4. The device according to claim 1, wherein the sensor component comprises an acoustic sensor to monitor heartbeat, respiration and rumen contraction vibrations. 5. The device according to claim 2, wherein the pH sensor comprises: a) a sensor housing:b) a chamber defined in the housing;c) a pH sensing device in communication with the chamber; and;d) a diffusion barrier to allow slow flow of ions between the chamber and a medium being monitored for pH. 6. The device according to claim 5, further including anode and cathode electrodes in communication with the chamber for establishing a reaction in the chamber for measurement of pH. 7. A system for monitoring at least one physiological parameter within the stomach of at least one ruminant animal comprising: a) a device for introduction into the stomach of the animal and comprising a sensor for obtaining physiological data and providing the data in a form accessible by radio frequency communication, a weight component of a size and shape for introduction to the stomach of the animal and for settling by gravity to a low region of the stomach, the weight component comprising a cow magnet, and a cord tethering the sensor to the weight component; andb) a radio frequency communication device located external to the animal for communicating with the device to obtain the physiological data. 8. The system according to claim 7, wherein the radio frequency communication device comprises a transceiver. 9. The system according to claim 7, wherein the device includes a transceiver operatively associated with the sensor. 10. The system according to claim 7, wherein the radio frequency communications device is fixed at a location so as to be stationary. 11. The system according to claim 7, wherein the radio frequency communications device is portable. 12. The system according to claim 7, including a plurality of devices for introduction into the stomachs of a corresponding plurality of animals and wherein the sensor of each device is provided with a unique identification code for access by the radio frequency communications device. 13. The system according to claim 7... | - | A61B005/05 A61B005/00 A61B007/00
US07209779 | B2 | 2007-04-24 | Methods and software for retrospectively gating a set of images | 2002-05-30 | Methods, devices, and software for selecting image slices. The image slices are preferably taken during a quiet part of the patient's heart cycle in which the heart is relatively motionless so as to reduce blurring and the introduction of artifacts into the resultant reconstructed projection image of the heart. The methods of the present invention can analyze information regarding the duration of the patient's R-R cycles and allow the user to choose from a plurality of selection criteria, typically an absolute time or a percentage of the R-R cycle, so as to select slice images for inclusion in a projection image. The ability to choose from a plurality of selection criteria provides for flexibility to account for irregular heartbeats or a change in the duration of the patient's R-R cycles. In exemplary embodiments, the resultant projection image can be used for coronary calcium scoring or 3-D rendering of the heart. | 1. A method of selecting slices from an image scan, the method comprising: acquiring a plurality of overlapping slice images of a patient with a CT scanner;acquiring an ECG signal of the patient during the acquiring of plurality of overlapping slice images;choosing a first set of slice images from the plurality of overlapping slice images based on a timing of the slice image relative to the EGG signal;applying a maximum intensity projection algorithm to each slice image of the first set of slices and slices immediately adjacent the slice images in the first set to generate a maximum intensity projection slice image that contains, on a pixel by pixel basis, information from a pixel with the highest intensity in the three slice images; andgenerating a second set of slice images from the maximum intensity projection slices. | 1. A method of selecting slices from an image scan, the method comprising: acquiring a plurality of overlapping slice images of a patient with a CT scanner;acquiring an ECG signal of the patient during the acquiring of plurality of overlapping slice images;choosing a first set of slice images from the plurality of overlapping slice images based on a timing of the slice image relative to the EGG signal;applying a maximum intensity projection algorithm to each slice image of the first set of slices and slices immediately adjacent the slice images in the first set to generate a maximum intensity projection slice image that contains, on a pixel by pixel basis, information from a pixel with the highest intensity in the three slice images; andgenerating a second set of slice images from the maximum intensity projection slices. 2. The method of claim 1 comprising generating a coronal or sagittal reconstruction with the second set of slice images. 3. The method of claim 1 wherein the EGG signal comprises a plurality of cardiac cycles, wherein each of the cardiac cycles includes an R-wave, wherein choosing comprises selecting slices that are at an absolute time period before or after the R-wave. 4. The method of claim 1, wherein the EGG signal comprises a plurality of cardiac cycles, wherein each of the cardiac cycles includes an R-wave, wherein choosing comprises selecting slices that are a percentage of the cardiac cycle before or after the R-wave. 5. The method of claim 1 comprising calcium scaring the second set of slices. 6. The method of claim 1 comprising 3-D rendering the second set of slices. | Accuimage Diagnostics Corp. | A61B005/05
US06913594 | B2 | 2005-07-05 | Dual-function catheter handle | 2001-12-31 | A dual function catheter handle is provided for simultaneous movement of two control wires. The catheter handle has a handle housing having proximal and distal ends and a generally hollow interior. A core is mounted in the interior of the handle housing, the core having a longitudinal slot therethrough. A first moveable member is provided with a proximal end mounted in the interior of the handle housing and a distal end extending outside the handle housing. The first moveable member is longitudinally moveable relative to the handle housing. A second moveable member is mounted in the longitudinal slot of the core and is longitudinally moveable relative to the core and handle housing. The second moveable member has a threaded surface. A rotatable member is mounted on the handle housing. The rotatable member has a threaded inner surface that mates with the threaded surface of the second moveable member so that rotation of the rotatable member causes longitudinal movement of the second moveable member. The first moveable member and second moveable member are capable of simultaneously moving proximally relative to the handle housing. | 1. A catheter handle comprising: a handle housing having proximal and distal ends and a generally hollow interior; a first moveable member having a proximal end mounted in the interior of the handle housing and a distal end extending outside the handle housing, the first moveable member being longitudinally moveable relative to the handle housing; a second moveable member mounted in the interior of the handle housing and longitudinally moveable relative to the handle housing; a rotatable member mounted on the handle housing, whereby rotation of the rotatable member causes longitudinal movement of the second moveable member; and a core mounted in the interior of the handle housing, the core having a longitudinal slot therethrough in which the second moveable member is slidably mounted. | 1. A catheter handle comprising: a handle housing having proximal and distal ends and a generally hollow interior; a first moveable member having a proximal end mounted in the interior of the handle housing and a distal end extending outside the handle housing, the first moveable member being longitudinally moveable relative to the handle housing; a second moveable member mounted in the interior of the handle housing and longitudinally moveable relative to the handle housing; a rotatable member mounted on the handle housing, whereby rotation of the rotatable member causes longitudinal movement of the second moveable member; and a core mounted in the interior of the handle housing, the core having a longitudinal slot therethrough in which the second moveable member is slidably mounted. 2. The catheter handle of claim 1, wherein the first moveable member and second moveable member are capable of simultaneously moving proximally relative to the handle housing. 3. The catheter handle of claim 1, wherein the second moveable member is not rotatably moveable relative to the handle housing. 4. The catheter handle of claim 1, further comprising a thumb control mounted at or near the distal end of the first moveable member. 5. The catheter handle of claim 1, wherein the rotatable member has a threaded inner surface that mates with a threaded surface of the second moveable member. 6. The catheter handle of claim 1, wherein the core has a lumen extending through at least a portion of the length of the core. 7. The catheter of claim 1, wherein the first moveable member is generally tubular and is mounted in surrounding relation to a portion of the core. 8. The catheter handle of claim 1, wherein the handle housing is generally tubular and the rotatable member extends around the entire circumference of the handle housing. 9. A catheter comprising: an elongated, flexible catheter body having proximal and distal ends and a lumen extending therethrough; an intermediate section at the distal end of the catheter body having proximal and distal ends and first and second off-axis lumens extending therethrough; a handle according to claim 1 mounted at the proximal end of the catheter body, whereby the catheter body is attached to the handle housing; a first puller wire extending through the catheter body and first off-axis lumen of the intermediate section, having a proximal end anchored to the first moveable member of the handle, and having a distal end anchored at or near the distal end of the catheter; and a second puller wire extending through the catheter body and second off-axis lumen of the intermediate section, having a proximal end anchored to the second moveable member of the handle, and having a distal end anchored at or near the distal end of the catheter. 10. The catheter of claim 9, wherein the distal end of the first puller wire is anchored to the intermediate section. 11. The catheter of claim 9, further comprising a generally circular mapp... | Biosense Webster, Inc. | A61M037/00 A61M025/00 A61M025/08 A61M025/09 A61B005/00
US06842636 | B2 | 2005-01-11 | Medical electrode | 2002-09-27 | An electrode providing electrical contact with a patient's skin includes a conductive member adapted for connection to an external electrical apparatus and a non-liquid water containing film for electrically interfacing to said patient's skin, the non-liquid film being electrically and mechanically connected to said conductive member. The non-liquid water containing film includes an electrically conductive organic polymer plasticized with a polyhydric alcohol with said organic polymer being derived from a monomeric mixture comprising from about 2 to 30 pph acrylic acid, 2 to 30 pph of a glycolvinylether and 0.01 to 1.5 pph of a crosslinking agent. Preferably the polyhydric alcohol is glycerol. | 1. An electrode providing electrical contact with a patient's skin comprising: a conductive member including means for connection to an external electrical apparatus; and means for electrically interfacing to said patient's skin, said interfacing means being electrically and mechanically connected to said conductive member, said interfacing means being a non-liquid water containing film and which comprises an electrically conductive organic polymer plasticized with a polyhydric alcohol with said organic polymer being derived from a monomeric mixture comprising from about 2 to 30 pph acrylic acid, 2 to 30 pph of a glycolvinylether, and 0.01 to 1.5 pph of a cross linking agent. | 1. An electrode providing electrical contact with a patient's skin comprising: a conductive member including means for connection to an external electrical apparatus; and means for electrically interfacing to said patient's skin, said interfacing means being electrically and mechanically connected to said conductive member, said interfacing means being a non-liquid water containing film and which comprises an electrically conductive organic polymer plasticized with a polyhydric alcohol with said organic polymer being derived from a monomeric mixture comprising from about 2 to 30 pph acrylic acid, 2 to 30 pph of a glycolvinylether, and 0.01 to 1.5 pph of a cross linking agent. 2. The electrode of claim 1 wherein said polyhydric alcohol is glycerol. 3. The electrode of claim 2 wherein said crosslinking agent is selected from the group consisting of methylene bis-acrylamide and polyethylene glycol diacrylates and dimethacrylates having a molecular weight from about 200 to about 600. 4. The electrode of claim 3 wherein said cross linking agent is ETMPTA. 5. The electrode of claim 1 wherein said glycolvinylether is selected from the group consisting of hydroxybutylvinylether, ethyleneglycolvinylether, diethylene glycol vinylether, and triethyleneglycolmethylvinylether. 6. The electrode of claim 5 wherein said glycolvinylether is diethylene glycol vinyl ether. 7. The electrode of claim 5 wherein said monomeric mixture further comprises from about 35 to 60 pph of glycerol. 8. The electrode of claim 7 wherein at least a portion of said acrylic acid is replaced with 2-acrylamido propane sulfonic acid (AMPS) or a cationic acrylic ester. 9. The electrode of claim 1 wherein said monomeric mixture further comprises from 0.01 to 3 pph of ultraviolet sensitive curing agents. 10. The electrode of claim 9 wherein said ultraviolet sensitive curing agent is selected from the group consisting of 2-hydroxy-2-methyl-1-phenylpropan-2-one, 4-(2-hydroxyethoxy)phenyl(2-hydroxy-2-phenyl-2-hydroxy-2-propyl)ketone, 2,2-dimethoxy-2-phenyl-acetophenone 1-[4-(2-Hydroxyethoxy)-phenyl]-2-hydroxy-2-methyl-1-propane-1-one, 1-hydroxycyclohexylphenyl ketone, trimethyl benzoyl diphenyl phosphine oxide and mixtures thereof. 11. The electrode of claim 10 wherein said mixture comprises from 5 to 25 pph acrylic acid and from 5 to 25 pph of a glycolvinylether. 12. The electrode of claim 11 wherein said glycolvinylether is selected from the group consisting of hydroxybutylvinylether, ethyleneglycolvinylether, diethylene glycol vinylether, and triethyleneglycolmethylvinylether. 13. The electrode of claim 12 wherein said glycol vinylether is diethylene glycol vinyl ether. 14. The electrode of claim 1 wherein said monomeric mixture further comprises an aldehyde reactant or neutralizer. 15. The electrode of claim 14 wherein said aldehyde reactant or neutralizer is selected from the group consisting of hydrogen peroxide, hydroxyethylethylene urea and L-arginine hydrochloride. 16.... | Axelgaard Manufacturing Co., Ltd. | A61B005/04 A61N001/04
US06934570 | B2 | 2005-08-23 | Physiological sensor combination | 2002-12-19 | A physiological sensor combination has a flexible substrate configured to attach to a tissue site. Multiple sensors are disposed on the substrate, which generate physiological signals. Each of the signals is responsive to a different physiological parameter. Conductors are carried on the substrate and routed between the sensors and at least one connector. The connector is configured to communicate the physiological signals to at least one monitor, which derives measurements of the parameters. | 1. A physiological sensor combination comprising: a flexible substrate configured to attach to a tissue site; a plurality of sensors disposed on said substrate and adapted to provide a corresponding plurality of physiological signals, each of said signals responsive to at least one of a plurality of physiological parameters; a plurality of conductors disposed on said substrate between said sensors and at least one connector, and said at least one connector configured to communicate said signals to at least one monitor so as to derive a plurality of measurements of said parameters, wherein said sensors comprise: a plurality of electrodes disposed on said substrate, each of said electrodes adapted to be in electrical communication with said tissue site and electrically connected to at least one of said conductors; and an emitter and a detector mounted to said substrate and electrically connected to at least one of said conductors, said emitter adapted to transmit light into said tissue site and said detector adapted to receive reflected light from said tissue site, wherein: said substrate has a first side adapted to face toward said tissue site and a second side adapted to face away from said tissue site; said conductors and said electrodes disposed on said first side; and said emitter and said detector mounted to said first side wherein said substrate comprises a fold-over portion having a circuit side corresponding to said first side, and said fold-over portion adapted to fold so that said circuit side is proximate said second side. | 1. A physiological sensor combination comprising: a flexible substrate configured to attach to a tissue site; a plurality of sensors disposed on said substrate and adapted to provide a corresponding plurality of physiological signals, each of said signals responsive to at least one of a plurality of physiological parameters; a plurality of conductors disposed on said substrate between said sensors and at least one connector, and said at least one connector configured to communicate said signals to at least one monitor so as to derive a plurality of measurements of said parameters, wherein said sensors comprise: a plurality of electrodes disposed on said substrate, each of said electrodes adapted to be in electrical communication with said tissue site and electrically connected to at least one of said conductors; and an emitter and a detector mounted to said substrate and electrically connected to at least one of said conductors, said emitter adapted to transmit light into said tissue site and said detector adapted to receive reflected light from said tissue site, wherein: said substrate has a first side adapted to face toward said tissue site and a second side adapted to face away from said tissue site; said conductors and said electrodes disposed on said first side; and said emitter and said detector mounted to said first side wherein said substrate comprises a fold-over portion having a circuit side corresponding to said first side, and said fold-over portion adapted to fold so that said circuit side is proximate said second side. 2. The physiological sensor combination according to claim 1 wherein said emitter and said detector are mounted to said a fold-over portion. 3. The physiological sensor combination according to claim 2 wherein said substrate defines at least one aperture configured so that said emitter and said detector each align with a corresponding one of said at least one aperture when said fold-over is in a folded position. 4. The physiological sensor combination according to claim 3 further comprising: a plurality of biopotential sensor pinouts corresponding to said electrodes; a plurality of optical sensor pinouts corresponding to said emitter and said detector; and a common connector extending from said substrate, said biopotential sensor pinouts and said optical sensor pinouts each disposed on said common connector. 5. The physiological sensor combination comprising: a substrate means for combining a first sensor and a second sensor; a connector means for communicating signals from said first sensor and said second sensor to at least one monitor; an identifying means of conveying information about each of said first sensor and said second sensor to said monitor; and fold-over means for positioning sensor components so as to extend away from a tissue site. 6. The physiological sensor combination according to claim 5 further comprising an aperture means for providing light communications between sensor components and ... | Masimo Corporation | A61B005/00 A61B005/0205
US07283862 | B1 | 2007-10-16 | Rapid multi-slice MR perfusion imaging with large dynamic range | 2002-05-16 | The present invention includes a method and apparatus to perform rapid multi-slice MR imaging without ECG gating or requiring breath-holding that is capable of renal profusion analysis and angiographic screening. An ungated interleaved pulse sequence is applied in rapid succession, followed by a delay interval. The ungated interleaved pulse sequence is repeatedly played out with the predefined delay interval between each application of the pulse sequence. The resulting images provide not only a series of temporal phases of contrast-enhanced blood uptake for renal profusion analysis, but also provide enough dynamic range to include angiographic screening of the renal arteries. | 1. A method of acquiring MR images comprising: (A) applying an ungated pulse sequence and acquiring multi-slice MR data in rapid succession of a first selected anatomy;(B) providing a delay interval of sufficient length to allow for subject free-breathing, wherein the delay interval is user-prescribed and is in the range of 10-20 seconds to allow for the subject free-breathing;(C) repeating steps (A) and (B) a prescribed number of times;(D) transmitting a notched RF saturation pulse having a stop-band between a pair of pass-bands; and(E) acquiring MR data for the slice location in the stop-band of the notched RF saturation pulse. | 1. A method of acquiring MR images comprising: (A) applying an ungated pulse sequence and acquiring multi-slice MR data in rapid succession of a first selected anatomy;(B) providing a delay interval of sufficient length to allow for subject free-breathing, wherein the delay interval is user-prescribed and is in the range of 10-20 seconds to allow for the subject free-breathing;(C) repeating steps (A) and (B) a prescribed number of times;(D) transmitting a notched RF saturation pulse having a stop-band between a pair of pass-bands; and(E) acquiring MR data for the slice location in the stop-band of the notched RF saturation pulse. 2. The method of claim 1 further comprising acquiring MR data of a second selected anatomy simultaneously with the first selected anatomy. 3. The method of claim 2 wherein the second selected anatomy is a cardiac region of the patient to acquire angiographic information about the patient. 4. The method of claim 1 further comprising the step of injecting a contrast agent in a patient for perfusion analysis imaging. 5. The method of claim 4 wherein the first selected anatomy are kidneys of the patient and further comprising reconstructing the MR data acquired into a series of images depicting temporal phases of contrast-enhanced blood uptake for renal function and perfusion evaluation. 6. The method of claim 1 further comprising reconstructing ungated multi-shot echo-planar images and acquiring the MR data during free-breathing. 7. The method of claim 1 further comprising applying an interleaved saturation recovery sequence for magnetization preparation before each pulse sequence is played out. 8. The method of claim 1 wherein the notched RF saturation pulse is designed to saturate all slice locations in a volume of slice locations selected except a slice location in which MR data is to be acquired after the transmission of the notched RF saturation pulse. 9. The method of claim 1 wherein a width of the stop-band of the notched RF saturation pulse is greater than that of a slice location thickness. 10. The method of claim 9 wherein a width of the notched RF saturation pulse is a user-selectable parameter. 11. A computerized system configured to play out a pulse sequence for perfusion imaging of an anatomy of interest, the computerized system comprising a computer programmed to apply: an interleaved preparation/acquisition sequence having a number of preparation pulses and a number of acquisition slices, wherein each preparation pulse corresponds to a given acquisition slice and is timed to occur such that at least one other preparation pulse and one other acquisition slice occur therebetween;a dead space to provide a delay time sufficient to allow patient free-breathing;a repetition time defining a time of the interleaved preparation/acquisition sequence followed by the dead space; anda scanning session defining a prescribed number of times the pulse sequence with the repetition time is repeated; andwherein the c... | General Electric Company | A61B005/05
US07190992 | B2 | 2007-03-13 | Magnetic resonance imaging using technique of positioning multi-slabs to be imaged | 2002-08-30 | A magnetic resonance imaging system is provided for obtaining MR images by scanning a region of the object previously located on an object's positioning image. The system comprises a displaying unit, inputting unit, approximating unit, and locating unit. The displaying unit displays a plurality of tomographic images of the object as the positioning image, each of the tomographic images including an indication of a target of interest thereon. The inputting unit enables information about a running state of the target in a direction along the target to be supplied toward each of the tomographic images. The approximating unit calculates three-dimensionally an approximated curve indicating the running state of the target in the direction on the basis of the supplied information about the running state. The locating unit locates the region substantially perpendicular to the approximated curve. | 1. A magnetic resonance imaging system in which an object to be imaged is subjected to scanning to obtain an MR image, the scanning being applied to a region of the object, a sub-region to be imaged being previously located on a positioning image of the object, the system comprising: a displaying unit configured to display the positioning image;a locating unit configured to repetitively allow an operator to locate two points at desired points on the positioning image; anda setting unit configured to set a further sub-region substantially perpendicular to the positioning image on the basis of the located two points, the further sub-region thus located being set by the setting unit being a saturation sub-region to be saturated with an MR signal before scanning the region for the MR image of the sub-region to be imaged. | 1. A magnetic resonance imaging system in which an object to be imaged is subjected to scanning to obtain an MR image, the scanning being applied to a region of the object, a sub-region to be imaged being previously located on a positioning image of the object, the system comprising: a displaying unit configured to display the positioning image;a locating unit configured to repetitively allow an operator to locate two points at desired points on the positioning image; anda setting unit configured to set a further sub-region substantially perpendicular to the positioning image on the basis of the located two points, the further sub-region thus located being set by the setting unit being a saturation sub-region to be saturated with an MR signal before scanning the region for the MR image of the sub-region to be imaged. 2. The magnetic resonance imaging system of claim 1, wherein the sub-region to be imaged is also set by a setting unit by locating an additional pair of points defining a slab consisting of at least one slice to be scanned for the MR image. 3. The magnetic resonance imaging system of claim 2, further comprising a specifying unit configured to specify the number of slices composing a single slab serving as the sub-region to be imaged, wherein the setting unit is configured to set at least two slices adjacently and in parallel with each other on the basis of the located additional pair of points and the specified number of slices. 4. A magnetic resonance imaging system in which an object to be imaged is subjected to scanning to obtain an MR image, the scanning being applied to a sub-region of the object, the sub-region being previously located on a positioning image of the object, the system comprising: a displaying unit configured to display the positioning image;a locating unit configured to repetitively allow an operator to locate two points at desired points on the positioning image; anda setting unit configured to set the sub-region to be imaged substantially perpendicular to the positioning image on the basis of the located two points. 5. The magnetic resonance imaging system of claim 4, wherein the sub-region to be imaged set by the setting unit is a slab consisting of at least one slice to be scanned for obtaining the MR image. 6. The magnetic resonance imaging system of claim 5, further comprising a specifying unit configured to specify the number of slices composing a single slab serving as the sub-region to be imaged, wherein the setting unit is configured to set at least two slices adjacently and in parallel with each other on the basis of the located two points and the specified number of slices. 7. A method for planning a sub-region to be scanned by magnetic resonance imaging, the sub-region of the object being previously located on a positioning image previously acquired from the object, the method comprising: displaying the positioning image;interactively allowing two points to be located at desired points on the ... | Kabushiki Kaisha Toshiba | A61B005/055
US07018344 | B2 | 2006-03-28 | Retractable needle shielding device | 2003-01-21 | The present invention is directed to a shieldable needle device. The shieldable needle device includes a hub having a lever extending from an external surface thereof, and a needle cannula with a puncture tip extending from said hub. The hub is interengaged with a housing having a main body portion with the needle cannula extending through a forward opening therein, and a rearward extension with a longitudinal slot extending axially therealong between a forward slot end and a rearward slot end, such that the lever of the hub extends through the longitudinal slot and is locked at the forward slot end thereof. The housing also includes a tip guard within the main body portion which is axially movable along the needle cannula. The lever is radially movable with respect to a longitudinal axis of the needle device to unlock the lever from the forward slot end, and is axially movable toward the rearward slot end to cause axial movement and to thereby retract the needle cannula, with the tip guard protectively covering the puncture tip of the needle cannula. | 1. A shieldable needle device comprising: a hub including a lever extending from an external surface thereof;a needle cannula extending from said hub, said needle cannula having a puncture tip; anda housing including a main body portion having a forward opening with said needle cannula extending through said main body portion from said forward opening, and a rearward extension having a longitudinal slot extending axially therealong between a forward slot end and a rearward slot end with said lever of said hub extending through said longitudinal slot and locked at said forward slot end thereof, said housing further including a tip guard within said main body portion;wherein said lever is radially movable with respect to a longitudinal axis of the needle device to unlock said lever from the forward slot end, and said lever is axially movable toward the rearward slot end causing axial movement of said needle cannula through said forward opening of said housing, thereby retracting said needle cannula within said main body portion of said housing to a position with said tip guard protectively covering said puncture tip of said needle cannula. | 1. A shieldable needle device comprising: a hub including a lever extending from an external surface thereof;a needle cannula extending from said hub, said needle cannula having a puncture tip; anda housing including a main body portion having a forward opening with said needle cannula extending through said main body portion from said forward opening, and a rearward extension having a longitudinal slot extending axially therealong between a forward slot end and a rearward slot end with said lever of said hub extending through said longitudinal slot and locked at said forward slot end thereof, said housing further including a tip guard within said main body portion;wherein said lever is radially movable with respect to a longitudinal axis of the needle device to unlock said lever from the forward slot end, and said lever is axially movable toward the rearward slot end causing axial movement of said needle cannula through said forward opening of said housing, thereby retracting said needle cannula within said main body portion of said housing to a position with said tip guard protectively covering said puncture tip of said needle cannula. 2. A shieldable needle device as in claim 1, wherein said lever and said longitudinal slot are in frictional engagement along said longitudinal slot, thereby requiring a force to be exerted on said lever toward said rearward slot end to cause axial movement of said lever toward said rearward slot end. 3. A shieldable needle device as in claim 1, wherein said lever includes a finger surface having a profile for accommodating a user's finger. 4. A shieldable needle device as in claim 3, wherein said lever is axially movable toward said rearward slot end through pressure exerted on said finger surface. 5. A shieldable needle device as in claim 1, wherein said tip guard includes a portion for accommodating said needle cannula during axial movement of said needle cannula for retraction within said main body portion. 6. A shieldable needle device as in claim 1, wherein said tip guard comprises a resiliently deflectable material. 7. A shieldable needle device as in claim 6, wherein said resiliently deflectable material includes a forward end and a rearward end with a spring leg extending therebetween, said forward end including a bend to form a lockout leg, said rearward end including a bend forming a rearward extent, each of said rearward end and said rearward extent including openings for accommodating said needle cannula therethrough, wherein said spring leg is biased against said needle cannula during axial movement of said needle cannula and said lockout leg is resiliently moved over said puncture tip of said needle cannula when said needle cannula is retracted within said main body portion. 8. A shieldable needle device as in claim 1, wherein said longitudinal slot is free of a rearward locking mechanism. 9. A shieldable needle device as in claim 1, wherein said longitudinal slot is of substantially unifor... | Becton, Dickinson and Company | A61B005/00 B65D081/00
US07063673 | B2 | 2006-06-20 | Coupling device for blood collection assembly | 2003-01-23 | The present invention is directed to a coupler for providing fluid connection between a needle holder and an intravenous needle assembly. The coupler includes a coupler housing and a coupler needle cannula. The coupler housing includes a forward end and a rearward end with the forward end having a piercable septum. The coupler housing also includes at least one lug on its external surface. The coupler is capable of forming a blood collection assembly to provide for fluid connection between a conventional needle holder and an intravenous needle assembly which includes guide slots traditionally provided for engagement with lugs on specific specimen tube arrangements. | 1. A coupler for engaging a needle assembly with a needle holder comprising: a coupler housing having a forward end and a rearward end, said forward end including a piercable septum therein, said coupler housing including at least one lug on an external surface thereof; anda coupler needle cannula extending from said rearward end of said coupler housing;wherein the at least one lug is configured to removably engage the coupler with the needle assembly. | 1. A coupler for engaging a needle assembly with a needle holder comprising: a coupler housing having a forward end and a rearward end, said forward end including a piercable septum therein, said coupler housing including at least one lug on an external surface thereof; anda coupler needle cannula extending from said rearward end of said coupler housing;wherein the at least one lug is configured to removably engage the coupler with the needle assembly. 2. The coupler according to claim 1, further comprising an elastomeric sheath encompassing the coupler needle cannula which extends from said rearward end of said coupler housing. 3. The coupler according to claim 1, wherein the coupler has a plurality of lugs. 4. The coupler according to claim 3, wherein the coupler has at least three lugs. 5. The coupler according to claim 1, wherein the rearward end of the coupler housing includes means for engagement with a needle holder. 6. The coupler according to claim 1, wherein the rearward end of the coupler housing includes external threads for engagement with a needle holder. 7. The coupler according to claim 6, wherein the rearward end of the coupler housing includes a profile for snap-fit engagement with a needle holder. 8. A blood collection assembly comprising a needle holder and an intravenous needle assembly which are coupled together through the coupler of claim 1, wherein the intravenous needle assembly includes slots for corresponding engagement with said lugs of said coupler. 9. A sampling container needle holder assembly for engagement with a needle assembly having an intravenous puncture end, a non-patient puncture tip and a guide sleeve housing with at least one slot extending along said housing adjacent said non-patient puncture tip, said needle holder assembly comprising: a needle holder comprising a hollow body having a forward end and a rearward end having an internal opening, said needle holder capable of accommodating a plurality of sampling containers therein; anda coupler extending from said forward end of said needle holder for providing fluid connection between the needle holder and the needle assembly, said coupler comprising a coupler housing having a forward end including a piercable septum therein and a rearward end having a coupler needle cannula extending therefrom, said forward end of said coupler housing accommodating said non-patient puncture tip of said needle assembly puncturing said piercable septum, said rearward end of said coupler housing connected with said needle holder with said coupler needle cannula extending into said hollow body of said needle holder through said forward end thereof, said coupler housing including at least one lug on an external surface thereof for removable engagement with said at least one slot of said guide sleeve housing. 10. The needle holder assembly according to claim 9, further including an elastomeric sleeve extending about said coupler needle cannula. 11. The needle holder... | Becton Dickinson and Company | A61B005/00 B65D081/00
US07280859 | B2 | 2007-10-09 | Biological measurement probe, biological optical measurement instrument using the same, and brain function measurement instrument | 2002-05-24 | The present invention provides a biological measurement probe readily wearable on a biological surface having a large curvature and a measurement technique of simultaneously measuring the blood circulation dynamic change accompanying the auditory/language/visual function activities of a subject (especially baby) and brain waves. A biological measurement probe having illuminating means for illuminating a subject through a first optical waveguide and light collecting means for collecting light illuminated from the illuminating means and propagated through the subject through a second optical waveguide, has first and second guide members having passages through which the first and second optical waveguides are passed and fixing the first and second optical waveguides, and a base member connecting the first guide member to the second one, holding them, and forming a portion brought into contact with the subject. | 1. A biological measurement probe comprising illuminating means for illuminating the head of a subject through a plurality of optical waveguides and light collecting means for collecting light illuminated from said illuminating means and propagated through the subject through a plurality of other optical waveguides; a plurality of optical waveguide fixing parts each having a plurality of optical waveguide fixing members for fixing the plurality of optical waveguides;a plurality of holding parts holding and covering the plurality of optical waveguide fixing parts, said plurality of holding parts being constructed with extensible belts or strings, and pressing the plurality of optical waveguides to the head of the subject when the probe is installed on the head of the subject; anda plurality of connecting parts connecting respective holding parts. | 1. A biological measurement probe comprising illuminating means for illuminating the head of a subject through a plurality of optical waveguides and light collecting means for collecting light illuminated from said illuminating means and propagated through the subject through a plurality of other optical waveguides; a plurality of optical waveguide fixing parts each having a plurality of optical waveguide fixing members for fixing the plurality of optical waveguides;a plurality of holding parts holding and covering the plurality of optical waveguide fixing parts, said plurality of holding parts being constructed with extensible belts or strings, and pressing the plurality of optical waveguides to the head of the subject when the probe is installed on the head of the subject; anda plurality of connecting parts connecting respective holding parts. 2. The biological measurement probe according to claim 1, wherein the optical waveguide fixing members are arranged on the optical waveguide fixing parts in such manner that the optical waveguides for illumination and the optical waveguides for collection are alternately disposed on top positions of a square grid shape or a rhombus grid shape. 3. The biological measurement probe according to claim 1, wherein the illuminating means includes a light source for illumination and a light source control circuit controlling an injected current to the light source, and the light collecting means includes an optical detector and a detection signal circuit processing a detection signal obtained by the original detector. 4. The biological measurement probe according to claim 1, wherein the optical waveguide fixing members are each provided with an internally disposed electrode for brain wave measurement together with the optical waveguide so as to simultaneously make the measurement with light and the brain wave measurement. 5. The biological measurement probe according to claim 1, wherein the cross-section area of the optical waveguides at the side contacting to the head of the subject broadens as approaching the end surface of the optical waveguides. 6. A biological optical measurement instrument, comprising: a biological measurement probe having illuminating means for illuminating the head of a subject through a plurality of optical waveguides and light collecting means for collecting light illuminated from said illuminating means and propagated through the subject through a plurality of other optical waveguides, comprising:a plurality of optical waveguide fixing parts each having a plurality of optical waveguide fixing members for fixing the plurality of optical waveguides, anda plurality of holding parts holding and covering the plurality of optical waveguide fixing parts, said plurality of holding parts being constructed with extensible belts or strings, and pressing the plurality of optical waveguides to the head of the subject when the probe is installed on the head of the subject, anda plurality of con... | Hitachi, Ltd. Hitachi Medical Corporation | A61B005/00
US07258665 | B2 | 2007-08-21 | High availability deployment of an off-site management system for digital cardiac electrocardiograms operating in an application service provider model | 2001-09-21 | A preferred embodiment of the present invention provides a method and system for high availability remote centralized data management. In a preferred embodiment, the system includes a primary remote centralized data storage storing medical data, and a redundant remote centralized data storage storing a redundant copy of the medical data. In a preferred embodiment, the system also includes a primary data server for storing the medical data in the primary remote centralized data storage. In a preferred embodiment, the primary data server also stores the medical data in the redundant remote centralized data storage. Preferably, the system also includes a redundant data server for storing the medical data in the redundant remote centralized data storage and also the primary remote centralized data storage. Additionally, in a preferred embodiment, the system includes an interface allowing access to the system. Preferably, the interface is a redundant interface providing high availability access to the remote centralized data management system. | 1. A centralized remote medical data management system, said system comprising: a primary remotely accessible centralized data storage storing medical data, wherein said medical data comprises at least electrocardiogram (ECG) data and wherein said medical data is collected from at least one medical device and is stored and managed at a location remote from where said medical data is collected;a redundant remotely accessible centralized data storage storing a redundant copy of said medical data;primary and redundant medical device connections connecting said at least one medical device to said primary remotely accessible centralized data storage and said redundant remotely accessible centralized data storage to collect said medical data from said at least one medical device; andprimary and redundant user interfaces accessing said medical data at said primary remotely accessible centralized data storage and said redundant remotely accessible centralized data storage, said access including providing said at least ECG data for at least one of viewing and processing. | 1. A centralized remote medical data management system, said system comprising: a primary remotely accessible centralized data storage storing medical data, wherein said medical data comprises at least electrocardiogram (ECG) data and wherein said medical data is collected from at least one medical device and is stored and managed at a location remote from where said medical data is collected;a redundant remotely accessible centralized data storage storing a redundant copy of said medical data;primary and redundant medical device connections connecting said at least one medical device to said primary remotely accessible centralized data storage and said redundant remotely accessible centralized data storage to collect said medical data from said at least one medical device; andprimary and redundant user interfaces accessing said medical data at said primary remotely accessible centralized data storage and said redundant remotely accessible centralized data storage, said access including providing said at least ECG data for at least one of viewing and processing. 2. The system of claim 1, further comprising a data server for storing said medical data at said remotely accessible centralized data storage. 3. The system of claim 1, wherein said data server is capable of storing said medical data at said remotely accessible centralized data storage. 4. The system of claim 1, wherein said primary and redundant user interfaces comprise at least one web server. 5. The system of claim 1, wherein said primary and redundant user interfaces comprise at least one modem. 6. The system of claim 1, further comprising a web browser for accessing said medical data. 7. The system of claim 1, further comprising a workstation for accessing said medical data. 8. The system of claim 1, wherein said at least one medical device comprises an electrocardiograph capable of recording electrocardiogram data. 9. The system of claim 1, further comprising an authentication unit for authenticating access to said centralized remote medical data management system. 10. The system of claim 1, wherein said medical data comprises at least one of waveform data, measurements, and interpretations of signal data. 11. A method for remotely storing medical data, said method comprising: receiving said medical data at a remotely accessible centralized data storage via a primary medical device connection, wherein a redundant medical device connection is used to store said medical data upon a failure of said primary medical device connection, said medical data comprising at least electrocardiogram (EGG) data;storing said medical data at said remotely accessible centralized data storage, wherein said medical data is collected from at least one medical device and is stored and managed at a location remote from where said medical data is collected; andstoring said medical data in a redundant remotely accessible centralized data storage,wherein said medical data is stored in said remotely ... | GE Medical Systems Global Technology Company, LLC | A61B005/00
US07236816 | B2 | 2007-06-26 | Biopsy and sampling needle antennas for magnetic resonance imaging-guided biopsies | 2002-04-24 | Herein is disclosed a magnetic resonance imaging antenna, including an inner conductor, an outer shield slideably displaceable with respect to the inner conductor, and an insulator electrically insulating the inner conductor from the outer shield. Herein is disclosed a biopsy needle antenna, including a magnetic resonance imaging antenna, having an outer shield, and an inner conductor electrically insulated from the outer shield by a dielectric; and a biopsy needle electrically connected to the inner conductor and electrically insulated from the outer shield by the dielectric. Herein is disclosed a method of obtaining a sample with magnetic resonance imaging guidance, including providing a sampling needle magnetic resonance imaging antenna, advancing the antenna to a structure from which the sample is to be taken, detecting magnetic resonance data by the antenna, and coupling the sample to the antenna. | 1. A biopsy needle antenna, comprising: a magnetic resonance imaging antenna, having:an outer shield; andan inner conductor electrically insulated from the outer shield by a dielectric; anda biopsy needle electrically connected to the inner conductor and electrically insulated from the outer shield by the dielectric. | 1. A biopsy needle antenna, comprising: a magnetic resonance imaging antenna, having:an outer shield; andan inner conductor electrically insulated from the outer shield by a dielectric; anda biopsy needle electrically connected to the inner conductor and electrically insulated from the outer shield by the dielectric. 2. The biopsy needle antenna of claim 1, further comprising a sheath, the biopsy needle being slideably displaceable within the sheath. 3. The biopsy needle antenna of claim 2, wherein the sheath is defined by the outer shield. 4. The biopsy needle antenna of claim 1, wherein at least one of the outer shield, the inner conductor, and the biopsy needle comprise a magnetic resonance compatible material. 5. The biopsy needle antenna of claim 1, wherein the biopsy needle antenna is configured with circuitry that receives magnetic resonance spectroscopy information from a sample. 6. The biopsy needle antenna of claim 1, wherein the outer shield and the inner conductor form a coaxial cable. 7. The biopsy needle antenna of claim 6, wherein the coaxial cable is electrically interconnected to an impedance matching circuit. 8. The biopsy needle antenna of claim 1, wherein at least one of the outer shield, the inner conductor, and the biopsy needle comprise a superelastic material. 9. The biopsy needle antenna of claim 1, wherein the outer shield is slideably displaceable with respect to the inner conductor. 10. The biopsy needle antenna of claim 1, wherein the inner conductor comprises an obturator and the outer shield comprises a cannula slideably displaceable over the obturator. 11. The biopsy needle antenna of claim 10, wherein the obturator further comprises a side-slit. 12. The biopsy needle antenna of claim 11, wherein the cannula comprises a distal end having a cutting edge, the cutting edge slideably displaceable over the side-slit. 13. The biopsy needle antenna of claim 11, wherein the cannula covers at least the side-slit. 14. The biopsy needle antenna of claim 10, wherein the cannula is spring-loaded. 15. The biopsy needle antenna of claim 10, wherein at least a portion of the obturator protrudes from a distal end of the cannula. 16. The biopsy needle antenna of claim 1, further comprising a spring coupled to the outer shield. 17. The biopsy needle of claim 16, wherein the spring is electrically coupled to the outer shield. 18. The biopsy needle antenna of claim 1, wherein the dielectric comprises at least one of fluroethylene polymer, tetrafluoroethylene, polyester, polyethylene, silicone, metal oxide, glass, and polyethylene terephthalate. 19. The biopsy needle antenna of claim 1, wherein the dielectric is covered by a lubricious coating. 20. The biopsy needle antenna of claim 19, wherein the lubricious coating comprises at least one of polyvinylpyrrolidone, polyacrylic acid, and silicone. 21. The biopsy needle antenna of claim 1, wherein the inner conductor and outer shield are electrically coupled to an inter... | Johns Hopkins University | A61B005/055
US07052468 | B2 | 2006-05-30 | Method and apparatus for detecting environmental smoke exposure | 2002-05-22 | The present invention includes a method and apparatus for detecting exposure to environmental tobacco smoke by analyzing a sample of breath using electronic sensor technology, including surface acoustic-wave gas sensor technology. The method determines the presence and concentration of substance(s) (or a class of substances) indicative of environmental smoke exposure. Diagnostic software is used to identify substances where a stored library of signatures is compared to the signature obtained from the system. Signal processing and neural networks are preferably utilized in the analysis. | 1. A method of determining smoke exposure, comprising: obtaining a sample of exhaled breath; andanalyzing said sample of breath with sensor technology to determine the presence of smoke related components thereby determining smoke exposure, wherein the analysis of said sample includes comparing the results sensed in said sample with a predetermined signature profile of a specific smoke related component,wherein the predetermined signature profile of said specific smoke related component is associated with a chemical selected from the group consisting of: acetaldehyde, acetic acid, acetone, acrolein, 4-Amniobiphenol, ammonia, anatabine, aniline, benzene, benz[a]anthracene, benzo[a]pryrene, benzoic acid, gamma-butyrolactone, cadmium, carbonyl sulfide, catechol, cotinine, cyanide derivatives, dimethylamine, formaldehyde, formic acid, glycolic acid, Harman, hydrazine, hydrogen cyanide, hydroquinone, isoprene, methylamine, methyl chloride, 3-methylpyridine 2-naphthylamino, nickel, nicotine, nitrogen oxides, N-Nitrosodiethanolamine, N-Nitrosodiethylamine, N-Nitrosodimethylamine, NNK, phenol, polonium-210, propionaldehyde, pyridine, quinoline, succinic acid, toluene, 2-Toluidine, 3-Vinylpyridine, xylene, and Zinc. | 1. A method of determining smoke exposure, comprising: obtaining a sample of exhaled breath; andanalyzing said sample of breath with sensor technology to determine the presence of smoke related components thereby determining smoke exposure, wherein the analysis of said sample includes comparing the results sensed in said sample with a predetermined signature profile of a specific smoke related component,wherein the predetermined signature profile of said specific smoke related component is associated with a chemical selected from the group consisting of: acetaldehyde, acetic acid, acetone, acrolein, 4-Amniobiphenol, ammonia, anatabine, aniline, benzene, benz[a]anthracene, benzo[a]pryrene, benzoic acid, gamma-butyrolactone, cadmium, carbonyl sulfide, catechol, cotinine, cyanide derivatives, dimethylamine, formaldehyde, formic acid, glycolic acid, Harman, hydrazine, hydrogen cyanide, hydroquinone, isoprene, methylamine, methyl chloride, 3-methylpyridine 2-naphthylamino, nickel, nicotine, nitrogen oxides, N-Nitrosodiethanolamine, N-Nitrosodiethylamine, N-Nitrosodimethylamine, NNK, phenol, polonium-210, propionaldehyde, pyridine, quinoline, succinic acid, toluene, 2-Toluidine, 3-Vinylpyridine, xylene, and Zinc. 2. The method of claim 1 wherein said sample is analyzed to determine the presence of said smoke related components by sensor technology selected from the group consisting of: semiconductor gas sensor technology; conductive polymer gas sensor technology; aptamer sensor technology; and amplifying fluorescent polymer (AFP) sensor technology. 3. The method of claim 1 wherein said sample is analyzed to determine the presence of said smoke related components by at least one surface acoustic wave gas sensor produced by technology selected from the group consisting of pulsed laser deposition, matrix assisted pulsed laser evaporation, and pulsed laser assisted surface functionalization. 4. The method of claim 1 wherein the sensor technology produces a unique electronic fingerprint to characterize the smoke exposure such that the presence and concentration of the smoke related components is determined. 5. The method of claim 1 wherein said sample is analyzed to confirm the presence of said smoke related components by a spectrophotometer. 6. The method of claim 1 wherein said sample is analyzed to confirm the presence of said smoke related components by a mass spectrometer. 7. The method of claim 1 further comprising the step of recording data resulting from analysis of said sample. 8. The method of claim 1 further comprising the step of communicating data resulting from analysis of said sample to a remote system. 9. The method of claim 1 further comprising the step of capturing exhaled breath in a vessel prior to analysis. 10. The method of claim 1 further comprising the step of dehumidifying said sample prior to analysis. 11. The method of claim 1 wherein said analysis further includes detecting exhalation with a sensor. 12. The method of c... | University of Florida Research Foundation, Inc. | A61B005/08 G01N001/22
US07075537 | B2 | 2006-07-11 | Apparatus for measuring a biological parameter equipped with a graphic representation display | 2002-11-27 | The invention concerns an apparatus for measuring a biological parameter of an individual comprising measuring means, means for storing and/or for computing values of said parameter connected to a display (3) for graphic representation of the measured parameter relative to a representation axis. The invention is characterized in that said display (3) is designed to adjust the measured parameter relative to a reference zone delimited on either side on said axis by a fixed mark combined with at least an adjacent digital indication. | 1. Apparatus for measuring a biological parameter of an individual, said apparatus comprising: measuring means (4) for measuring the biological parameter; computing means (7) for computing a range of normality values of said parameter as a function of data specific to the individual; and a display (3) that is integrated into said apparatus and is connected to said computing means for graphic representation (14) of the measured parameter relative to a representation axis (15), characterized in that said computing means is operative to adjust the position of the graphic representation of said measured parameter on said display with respect to a reference zone on said display representing the computed normality range, the reference zone being delimited to one side and the other of said axis by a fixed mark (m, M) in combination with at least one adjacent digital indication (30,31,32). | 1. Apparatus for measuring a biological parameter of an individual, said apparatus comprising: measuring means (4) for measuring the biological parameter; computing means (7) for computing a range of normality values of said parameter as a function of data specific to the individual; and a display (3) that is integrated into said apparatus and is connected to said computing means for graphic representation (14) of the measured parameter relative to a representation axis (15), characterized in that said computing means is operative to adjust the position of the graphic representation of said measured parameter on said display with respect to a reference zone on said display representing the computed normality range, the reference zone being delimited to one side and the other of said axis by a fixed mark (m, M) in combination with at least one adjacent digital indication (30,31,32). 2. Apparatus according to claim 1, characterized in that said computing means (7) effectuate a calibration of the graphic representation scale as a function of the value attributed to the width of the zone comprised between the two fixed marks (m,M). 3. Apparatus according to claim 1, characterized in that said fixed marks (m,M) delimit, to one side and the other of the center of said axis (15), an interval of normality (24) of the measured parameter. 4. Apparatus according to claim 1, characterized in that said display (3) is adapted to graphically represent at least two distinct parameters simultaneously. 5. Apparatus according to claim 4, characterized in that said display (3) is adapted to display digital indications of one parameter at a time, and that it comprises means for switching over to the display of digital indications of another parameter. 6. Apparatus according to claim 5, characterized in that said display (3) is adapted to display intermittently the digital indications of each parameter and that it has a cursor linking the graphic representation to the digital indication of said parameter. 7. Apparatus according to claim 4, characterized in that said at least two parameters are represented on parallel axes (15,16,17). 8. Apparatus according to claim 7, characterized in that said marks (m,M) are constituted by two parallel lines (22,23) that cross said axes (15,16,17) and define a zone of normality (24) for the measured parameters. 9. Apparatus according to claim 8, characterized in that said parallel lines (22,23) are vertical. 10. Apparatus according to claim 4, characterized in that said parameters are the weight, the quantity of fat mass, and the quantity of lean mass of the individual. 11. Apparatus according to claim 4, characterized in that said parameters are the height, the weight, and the body mass index. 12. Apparatus according to claim 1, characterized in that the graphic representation axis (15,16,17) of said display (3) is constituted by a succession of 10 to 20 points. 13. Apparatus according to claim 1, characterized in that ... | SEB S.A. | G09G005/22 A61B005/05 G01G023/36 G01G019/414 G01G019/50
US07010339 | B2 | 2006-03-07 | Hybrid lidar-radar for medical diagnostics | 2002-07-29 | A hybrid lidar-radar system for detecting the presence of objects, such as cancerous tumors, within tissues by detecting reflected signals from the tissue and discriminating the information related to the cancerous tumor from the undesirable backscattering of light created by the tissue itself. The hybrid lidar-radar system utilizes continuous wave light that is preferably modulated at frequencies up to 60 GHz. The present invention filters the return signals from the tissue at a subcarrier modulation frequency so as to reject erroneous information contained in scattered lights, while at the same time retaining the coherent, unscattered and modulated light information so as to provide for an accuracy detection of tumors within tissues. | 1. A hybrid lidar-radar system for detecting the presence of tissue inhomogeneities within tissue by detection of reflected and scattered light signals from the tissue and discriminating information related to cancerous tumors from scattering of light created by the tissue itself, comprising: a source of microwave energy for providing a microwave subcarrier signal;a source of continuous wave light for providing and for transmitting toward the tissue an optical carrier signal that is modulated by the microwave subcarrier signal, said source of light transports the microwave subcarrier signal through the tissue and is reflected and scattered by the tissue, said optical carrier signal has a frequency and said source of continuous wave light comprises: first and second lasers each having an operating wavelength of about 1064 nm, said first and second lasers having an optical frequency difference in a range from 0.01 GHz to 60 GHz, each of said first and second lasers having an output;a 50/50 optical splitter receiving and combining the outputs of said first and second lasers, said optical splitter having first and second outputs with the first output of the optical splitter transmitted toward the tissue;a light source optical detector for receiving the second output of said optical splitter and a mixer for mixing it with a reference signal having a frequency slightly offset from the desired modulation frequency, said mixer providing an output serving as an offset and having an intermediate frequency; anda laser offset locking assembly receiving the output of said mixer and providing an output signal to each of said first and second lasers that provides coherence between said first and second lasers to said offset intermediate frequency;a detector for detecting the scattered and reflected light and for producing electrical signals indicative thereof; andan optical detector that provides an output signal to a network analyzer that receives said produced electrical signals and measures phase and amplitude differences between the light scattered from the tissue and said optical carrier signal for the detection of said tissue inhomogeneities. | 1. A hybrid lidar-radar system for detecting the presence of tissue inhomogeneities within tissue by detection of reflected and scattered light signals from the tissue and discriminating information related to cancerous tumors from scattering of light created by the tissue itself, comprising: a source of microwave energy for providing a microwave subcarrier signal;a source of continuous wave light for providing and for transmitting toward the tissue an optical carrier signal that is modulated by the microwave subcarrier signal, said source of light transports the microwave subcarrier signal through the tissue and is reflected and scattered by the tissue, said optical carrier signal has a frequency and said source of continuous wave light comprises: first and second lasers each having an operating wavelength of about 1064 nm, said first and second lasers having an optical frequency difference in a range from 0.01 GHz to 60 GHz, each of said first and second lasers having an output;a 50/50 optical splitter receiving and combining the outputs of said first and second lasers, said optical splitter having first and second outputs with the first output of the optical splitter transmitted toward the tissue;a light source optical detector for receiving the second output of said optical splitter and a mixer for mixing it with a reference signal having a frequency slightly offset from the desired modulation frequency, said mixer providing an output serving as an offset and having an intermediate frequency; anda laser offset locking assembly receiving the output of said mixer and providing an output signal to each of said first and second lasers that provides coherence between said first and second lasers to said offset intermediate frequency;a detector for detecting the scattered and reflected light and for producing electrical signals indicative thereof; andan optical detector that provides an output signal to a network analyzer that receives said produced electrical signals and measures phase and amplitude differences between the light scattered from the tissue and said optical carrier signal for the detection of said tissue inhomogeneities. 2. The hybrid lidar-radar system according to claim 1, wherein each of said first and second lasers provide coherent light having a power level of about 700 mW. 3. A method using a hybrid lidar-radar system for detecting the presence of tissue inhomogeneities within tissue by detecting reflected and scattered light signals from the tissue and discriminating information related to cancerous tumors from scattering of light created by the tissue itself comprising the steps: providing a microwave subcarrier signal;providing a modulator having said microwave subcarrier signal applied thereto;providing a source of light that provides an optical carrier signal comprising a reference signal that is transmitted toward the tissue;modulating said optical carrier signal with said microwave subcarrier signal;detecting said opti... | The United States of America as represented by the Secretary of the Navy | A61B005/05
US07059060 | B1 | 2006-06-13 | Device for recording events and measuring growth in an individual's life | 2002-09-13 | A device for recording both chronological events and physical growth events of an individual or group of individuals' lives. The device includes linear measuring device having at least first and second opposed sides. The first side has a linear measurement scale inscribed thereon and at least one recording surface associated therewith. The second side has a chronological scale, preferably in months and years, inscribed thereon and also has at least one recording surface associated with the chronological scale. Thus, physical growth can be measured on the first side and recorded on the recording surface associated therewith as the individual or several individuals grow physically. Significant events can be recorded on the other side. | 1. A device for measuring and recording both chronological events and physical growth events in an individual's life comprising: an elongated rod-like structure comprising a measuring and recording device having a front and reverse side;said elongated rod-like structure includes an extended linear measuring device having at least a front side and a reverse side,said front side of said extended linear measuring device having a linear scale measurement for measuring physical growth inscribed thereon and at least one recording surface associated therewith; andsaid reverse side of said extended linear measuring device having a chronological scale inscribed thereon wherein the chronological scale is comprised of calendar year markings and at least one recording surface associated therewith. | 1. A device for measuring and recording both chronological events and physical growth events in an individual's life comprising: an elongated rod-like structure comprising a measuring and recording device having a front and reverse side;said elongated rod-like structure includes an extended linear measuring device having at least a front side and a reverse side,said front side of said extended linear measuring device having a linear scale measurement for measuring physical growth inscribed thereon and at least one recording surface associated therewith; andsaid reverse side of said extended linear measuring device having a chronological scale inscribed thereon wherein the chronological scale is comprised of calendar year markings and at least one recording surface associated therewith. 2. The device as defined in claim 1 wherein said elongate rod-like structure is a unitary pole. 3. The device as defined in claim 2 wherein said recording surfaces are each recessed. 4. The device as defined in claim 2 wherein each said linear scale and chronological scale are inscribed on a surface separate from the recording surface associated therewith. 5. The device as defined in claim 2 wherein said front and reverse sides of said extended linear measuring device are opposite each other. 6. The device as defined in claim 1 wherein the surfaces of the front and reverse sides of said extended linear measuring device are curvilinear. 7. The device as defined in claim 1 wherein said front and reverse sides of said extended linear measuring device include flat surfaces. 8. The device as defined in claim 1 wherein the shape of the rod-like structure in cross-section includes a polyhedral shape. 9. The device as defined in claim 1 wherein the device is at least seven feet tall. 10. The device as defined in claim 1 wherein the linear scale is in English or metric units. 11. The device as defined in claim 1 wherein the chronological scale is in Julian or Gregorian units. 12. The device according to claim 1 having a cross-sectional configuration selected from the group consisting of polyhedrons, hourglass, modified hourglass, round, and non-cordal round shape. 13. A device for measuring and recording both chronological events and physical growth events in an individual's life comprising: an elongated rod-like structure comprising a measuring and recording device having a front and reverse side;said elongated rod-like structure includes an extended linear measuring device having at least a front side and a reverse side,said front side of said extended linear measuring device having a linear scale measurement for measuring physical growth inscribed thereon and at least one recording surface associated therewith;said reverse side of said extended linear measuring device having a chronological scale inscribed thereon and at least one recording surface associated therewith; andwherein said rod-like structure includes first and second independently mounted rods. ... | - | A61B005/107 B43L007/00 G01B005/02
US07006862 | B2 | 2006-02-28 | Graphical user interfaces and methods for retrospectively gating a set of images | 2002-05-30 | Graphical user interfaces and methods for displaying ECG, slice images and a reconstruction projection image of the slices. In exemplary embodiments, the graphical user interfaces and methods facilitate the selection of slice images that are taken during a quiet part of the patient's heart cycle in which the heart is relatively motionless so as to reduce blurring and the introduction of artifacts into the resultant reconstructed projection image of the heart. In exemplary embodiments, the resultant projection image can be used for coronary calcium detection and scoring or 3-D rendering. | 1. A graphical user interface for gating an image scan, the interface comprising: a display device; a first screen portion on the display device that displays at least one of a plurality of selected slice images of the image scan; a second screen portion on the display device that displays at least a portion of a patient's ECG signal that was obtained during the image scan; and a third screen portion on the display device that displays at least one of a coronal and sagittal projection image of the plurality of selected slices, scaled so as to appear with their correct aspect ratios. | 1. A graphical user interface for gating an image scan, the interface comprising: a display device; a first screen portion on the display device that displays at least one of a plurality of selected slice images of the image scan; a second screen portion on the display device that displays at least a portion of a patient's ECG signal that was obtained during the image scan; and a third screen portion on the display device that displays at least one of a coronal and sagittal projection image of the plurality of selected slices, scaled so as to appear with their correct aspect ratios. 2. The graphical user interface of claim 1 wherein the third screen portion further displays a projection of all slices of the image scan so as to indicate a quality of the coronal or sagittal projection of the plurality of selected slices. 3. The graphical user interface of claim 1 wherein the third screen portion comprises a marker that indicates a position of the displayed, selected slice image displayed in the first screen portion. 4. The graphical user interface of claim 1 comprising a fourth screen portion that displays a selected ECG interval criteria that was used for slice selection, wherein the selected ECG interval comprises an absolute position before or after an R-wave of the ECG or a percentage before or after the R-wave of the ECG. 5. The graphical user interface of claim 1 comprising a fourth screen portion that displays R—R cycles duration information of the ECG. 6. The graphical user interface of claim 5 wherein the R—R cycle information comprises a graph. 7. The graphical user interface of claim 5 wherein the fourth screen portion displays the average length of an R—R cycle for different portions of the image scan. 8. The graphical user interface of claim 7 wherein the different portions comprise a first ten cycles, a middle ten cycles, and a last ten cycles of the ECG. 9. The graphical user interface of claim 1 wherein selected intervals of the ECG displayed in the second screen portion are highlighted to illustrate the ECG signal and corresponding slices that are used in the projection image in the third screen portion. 10. The graphical user interface of claim 9 wherein a portion of the ECG that corresponds to the image displayed in the first screen portion is highlighted differently from the other highlighted portions of the ECG signal. 11. The graphical user interface of claim 1 wherein the first screen portion further displays images of slices that are adjacent to the selected slice image. 12. The graphical user interface of claim 11 wherein the displayed adjacent slice images are smaller in size than the selected slice image. 13. The graphical user interface of claim 1 wherein the first screen portion comprises a user actuatable button that allows a user to scroll through slices of an image scan. 14. The graphical user interface of claim 1 comprising a user actuatable button that when activated displays a stretched coronal or sa... | Accuimage Diagnostics Corp. | A61B005/05 A61B005/044
US07229409 | B2 | 2007-06-12 | Method of providing a home health care service and system for providing a home health care service | 2002-12-23 | A system for providing a home health care service includes a health care center, a service recipient device, and the Internet that connects these components. The health care center: is constructed on the Internet; includes a database for registering personal data that are necessary for the health care of registered at-home patients as basic health care data and storing measurement data that are provided from at-home patients over the course of time; and has the functions of inferring the health conditions of at-home patients based on the measurement data and the basic health care data, and offering appropriate services necessary for inferred health conditions. The service recipient device includes a biosensor and a signal processor. The biosensor detects chemical components contained in a substance that is discharged from the human body and converts detected values to electrical signals at the point of discharge. | 1. A method of providing a home health care service, comprising steps of: providing access to a health care center on the Internet, said health care center having personal data necessary for the health care of registered at-home patients as basic health care data and also having the capability of inferring the health conditions of the at-home patients from measurement data provided by the at-home home patients and said basic health care data, and providing services necessary for health care or services appropriate to the inferred health conditions to said at-home patients;distributing biosensor devices to at-home patients who wish to receive a home health care service, said biosensor device having a capability of detecting predetermined chemical components contained in a substance discharged from a human body and automatically generating electrical measurement data at the position of discharge as occasion arises;collecting the measurement data of each registered at-home patient to said health care center from at-home patients by way of said Internet; andmanaging time-varying health conditions of the at-home patients over a period of time based on the collected measurement data as well as said basic health care data, and providing necessary services for maintaining the health of the at-home patients,wherein said biosensor device includes a urine testing device comprising:a urine collector for collecting urine inside a toilet bowl;a biosensor exchangeably attached inside said urine collector for detecting predetermined chemical component values contained in urine and converting said detected values into electrical signals;a support member for supporting said urine collector away from said toilet bowl during collection and also supporting a transmission path of said electrical signals, said support member being arranged so that said urine collector may be automatically or manually moved from a housing location to a prescribed use-position in the toilet bowl and vice versa; anda signal processor for processing electrical signals supplied from said biosensor through said transmission path. | 1. A method of providing a home health care service, comprising steps of: providing access to a health care center on the Internet, said health care center having personal data necessary for the health care of registered at-home patients as basic health care data and also having the capability of inferring the health conditions of the at-home patients from measurement data provided by the at-home home patients and said basic health care data, and providing services necessary for health care or services appropriate to the inferred health conditions to said at-home patients;distributing biosensor devices to at-home patients who wish to receive a home health care service, said biosensor device having a capability of detecting predetermined chemical components contained in a substance discharged from a human body and automatically generating electrical measurement data at the position of discharge as occasion arises;collecting the measurement data of each registered at-home patient to said health care center from at-home patients by way of said Internet; andmanaging time-varying health conditions of the at-home patients over a period of time based on the collected measurement data as well as said basic health care data, and providing necessary services for maintaining the health of the at-home patients,wherein said biosensor device includes a urine testing device comprising:a urine collector for collecting urine inside a toilet bowl;a biosensor exchangeably attached inside said urine collector for detecting predetermined chemical component values contained in urine and converting said detected values into electrical signals;a support member for supporting said urine collector away from said toilet bowl during collection and also supporting a transmission path of said electrical signals, said support member being arranged so that said urine collector may be automatically or manually moved from a housing location to a prescribed use-position in the toilet bowl and vice versa; anda signal processor for processing electrical signals supplied from said biosensor through said transmission path. 2. The method of providing home health care service according to claim 1, comprising: measuring those components among the chemical components of urine that are associated with a morbidity and that change according to the daily health condition, over a period of time. 3. The method of providing a home health care service according to claim 2, wherein said health care center provides, as a basic service to all registered at-home patients, information that has been processed and arranged to enable an extraction of knowledge relating to a specific said morbidity; and provides, as selective services that can be selected by at-home patients, services that are appropriate to the inferred health conditions of said specific morbidity. 4. The method of providing a home health care service according to claim 2, wherein said method employs a biosensor for measuring glucose,... | NEC Corporation | A61B005/05
US07096057 | B2 | 2006-08-22 | Method and apparatus for intracorporeal medical imaging using a self-tuned coil | 2002-08-02 | Disclosed herein is an RF probe for use with a medical imaging apparatus, the probe comprising an intracorporeal self-tuned resonator coil for receiving a signal indicative of an image of an interior portion of a body. The resonator coil is preferably coupled to a transmission medium having a characteristic impedance such that the resonator coil is substantially self-matching with the transmission medium's characteristic impedance. Preferably the resonator coil comprises an open wound conductor having a plurality of turns. The length of the resonator coil can be used to tune the resonator coil to a desired frequency, such as the Larmour frequency. Further, a return lead of the transmission medium is preferably coupled to an end of the conductor, and a signal lead of the transmission medium is preferably coupled to a point on the winding, thereby defining a turns ratio for the resonator coil. The turns ratio of the resonator coil can be used to match the resonator coil to the characteristic impedance of the transmission medium. Because the resonator coil of the present invention is self-tuning and self-matching, it avoids the use of bulky and relatively expensive tuning and matching circuits. | 1. An RF probe for use with a medical imaging apparatus, said RF probe comprising an intracorporeal self-tuned resonator coil and wherein the resonator coil is self-tuned to a frequency of substantially the Larmour frequency. | 1. An RF probe for use with a medical imaging apparatus, said RF probe comprising an intracorporeal self-tuned resonator coil and wherein the resonator coil is self-tuned to a frequency of substantially the Larmour frequency. 2. The probe of claim 1 wherein the resonator coil is an open coil. 3. The probe of claim 2 wherein the resonator coil is adapted for coupling with a transmission medium for passing a signal from the resonator coil to a processor, the transmission medium having a characteristic impedance, and wherein the resonator coil is substantially sell-matching with the transmission medium's characteristic impedance. 4. The probe of claim 3 wherein the resonator coil comprises an open wound conductor having at least two turns, and wherein the resonator coil includes a first pre-selected coupling point for coupling to a signal lead and a second pre-selected coupling point for coupling to a return lead, said coupling points being located on different turns. 5. The probe of claim 4 wherein the pre-selected coupling points define a turns ratio for the resonator coil, and wherein the resonator coil turns ratio is sufficient for the resonator coil to substantially match the transmission medium's characteristic impedance. 6. The probe of claim 5 wherein the resonator coil has a length of approximately 4⅛ inches. 7. The probe of claim 6 wherein the resonator coil turns ratio is approximately 5:1. 8. The probe of claim 4 wherein the resonator coil is substantially self-matching with the transmission medium's characteristic impedance such that the voltage standing wave ratio (VSWR) for the resonator coil is not greater than approximately 2:1. 9. The probe of claim 4 wherein the conductor is 30 gauge wire. 10. The probe of claim 4 wherein the conductor is 36 gauge wire. 11. The probe of claim 4 further comprising an insulating sheath within which the resonator coil is disposed. 12. The probe of claim 4 further comprising a processor coupled to a transmission medium, the transmission medium being connected to the resonator coil for receiving the resonator coil signal, and the processor being configured to process the resonator coil signal to generate therefrom an image of an interior portion of the body. 13. The probe of claim 4 wherein the resonator coil has a cross-sectional diameter in a range of approximately 0.25 mm to approximately 2 mm. 14. The probe of claim 13 wherein the resonator coil diameter is a range of approximately 1 mm to approximately 2 mm. 15. The probe of claim 14 wherein the resonator coil is an intravascular resonator coil. 16. The probe of claim 4 further comprising an intracorporeal catheter within which the resonator coil is disposed. 17. The probe of claim 16 wherein the catheter includes a channel for delivering a substance to an interior portion of the body. 18. An RF probe for use in analyzing an interior portion of a body, the probe comprising: a resonator coil comprising a conductor formed into an op... | Barnes Jewish Hospital | A61B005/05
US07039456 | B2 | 2006-05-02 | Conductance path mechanism for wireless heartbeat transmitter | 2003-01-31 | This invention relates to a conductance path mechanism of a wireless heartbeat transmitter comprising a circuit board, a spiral spring, a U-shaped contact rod integrally molded to both side of a bottom cover and a rubber conductible sheet. A horizontal portion of said U-shaped contact rod is totally inserted inside a foundation of a bottom cover to form a steady mechanism of the present invention. | 1. A conductance path mechanism of a wireless heartbeat transmitter comprising an upper cover, a bottom cover provided with a front wall rib and a rear wall rib, a circuit board, two spiral springs, two U-shaped contact rods integrally molded on both left and right sides of said bottom cover, two fixing bands, two rubber conductible sheets and two pressing plate; and wherein said bottom cover further comprising a left wall rib and a right wall rib for connecting and strengthening said front wall rib and said rear wall rib, and two rib walls positioned around a contact positioning pin and located on both sides of a central district of said bottom cover to form a spring detent groove. | 1. A conductance path mechanism of a wireless heartbeat transmitter comprising an upper cover, a bottom cover provided with a front wall rib and a rear wall rib, a circuit board, two spiral springs, two U-shaped contact rods integrally molded on both left and right sides of said bottom cover, two fixing bands, two rubber conductible sheets and two pressing plate; and wherein said bottom cover further comprising a left wall rib and a right wall rib for connecting and strengthening said front wall rib and said rear wall rib, and two rib walls positioned around a contact positioning pin and located on both sides of a central district of said bottom cover to form a spring detent groove. 2. The conductance path mechanism according to claim 1, wherein said U-shaped contact rod has one horizontal portion totally inserted inside a foundation of said bottom cover, and has two vertical ends; and wherein one said vertical end is provided with three upward vertical pins protruding out of the foundation of said bottom cover and inserted into a rubber conductible sheet at the bottom of a fixing band to form a conductible, insert-connecting mechanism and the other vertical end of said U-shaped contact rods forms a contact positioning pin, which extend within said spring dent groove and places spiral springs on both sides of the bottom of said circuit board inside a collar of said contact positioning pin. 3. The conductance path mechanism according to claim 2, wherein lower sections of said spiral spring is tightly inserted inside said spring detent groove to make said spiral spring and said bottom cover form a tight and steady mechanism, and to make said spiral spring and said contact positioning pin form a superior electrical contact and an excellent conductance path. 4. The conductance path mechanism according to claim 3, further comprising a pressing plate sheet positioned at a front end surface of each fixing band to make said three upward vertical pins of said U-shaped contact rod completely inserted into the interior of said rubber conductible sheet. | - | A61B005/04
US07083579 | B2 | 2006-08-01 | Encapsulated medical device and method of examining, curing, and treating internal region of body cavity using encapsulated medical device | 2002-09-20 | An encapsulated medical device includes a radio antenna through which the encapsulated medical device transmits or receives a radio wave to or from an extracorporeal device. The encapsulated medical device is passed through the lumen of a body cavity in order to examine, cure, or treat an internal region of the body cavity under the control of the extracorporeal device. A capsule body included in the encapsulated medical device has an extended portion formed at an upper end of one axis of a cross section of the capsule body cut at right angle to the longitudinal axis thereof. The capsule body has a linkage hole bored in the extended portion thereof to allow a fluid such as a gas or humor to flow into forward and backward parts of the lumen through the hole. | 1. An encapsulated medical device comprising: a capsule body capable of passing through the lumen of a body cavity to permit examination or treatment; andfluid passage forming means, included in said capsule body, for when said capsule body passes through the lumen of a body cavity, if said capsule body comes into close contact with the internal surface of the body cavity to block the lumen thereof, forming a fluid passage that links the forward and backward parts of the lumen of the body cavity. | 1. An encapsulated medical device comprising: a capsule body capable of passing through the lumen of a body cavity to permit examination or treatment; andfluid passage forming means, included in said capsule body, for when said capsule body passes through the lumen of a body cavity, if said capsule body comes into close contact with the internal surface of the body cavity to block the lumen thereof, forming a fluid passage that links the forward and backward parts of the lumen of the body cavity. 2. The encapsulated medical device according to claim 1, wherein said fluid passage forming means is formed directly in said capsule body or in an armor member that is freely detachably attached to the periphery of said capsule body. 3. The encapsulated medical device according to claim 1, wherein: said fluid passage forming means is realized with an extended portion of said capsule body, of which cross section cut at right angle to the longitudinal axis thereof is non-circular, formed at an end of one axis of the cross section of said capsule body; provided that said capsule body does not have the extended portion, the cross section of said capsule body is substantially circular; andsaid fluid passage forming means forms a fluid passage. 4. The encapsulated medical device according to claim 1, wherein a plurality of fluid passages are formed in one direction with respect to the longitudinal axis of said capsule body. 5. The encapsulated medical device according to claim 1, wherein said capsule body includes at least observing means. 6. The encapsulated medical device according to claim 1, wherein a plurality of fluid passages are formed substantially symmetrically to one another with respect to the longitudinal axis of said capsule body. 7. The encapsulated medical device according to claim 1, wherein said capsule body has a plurality of hard members linked by a flexible string-like member. 8. The encapsulated medical device according to claim 1, wherein said capsule body has a switch, which is used to control a battery that supplies power and to control power supply from the battery, incorporated therein. 9. The encapsulated medical device according to claim 1, wherein said capsule body includes at least one medical-purpose examining means such as arm optical sensor, a pH sensor, a temperature sensor, a pressure sensor, or an ultrasonic probe. 10. The encapsulated medical device according to claim 1, wherein said capsule body includes at least one opening means for a treating or examining means, such as, a medicine administration port or a humor infusion port. 11. The encapsulated medical device according to claim 1, wherein said capsule body includes at least one treating means such as an injection needle. 12. The encapsulated medical device according to claim 1, wherein said fluid passage is realized with a spiral groove. 13. The encapsulated medical device according to claim 1, wherein said capsule body has at least a magnet incorporated... | Olympus Corporation | A61B005/117 A61B005/103
US07113820 | B2 | 2006-09-26 | Real-time, high frequency QRS electrocardiograph | 2001-07-12 | Real time cardiac electrical data are received from a patient, manipulated to determine various useful aspects of the ECG signal, and displayed in real time in a useful form on a computer screen or monitor. The monitor displays the high frequency data from the QRS complex in units of microvolts, juxtaposed with a display of conventional ECG data in units of millivolts or microvolts. The high frequency data are analyzed for their root mean square (RMS) voltage values and the discrete RMS values and related parameters are displayed in real time. The high frequency data from the QRS complex are analyzed with imbedded algorithms to determine the presence or absence of reduced amplitude zones, referred to herein as “RAZs”. RAZs are displayed as “go, no-go” signals on the computer monitor. The RMS and related values of the high frequency components are displayed as time varying signals, and the presence or absence of RAZs may be similarly displayed over time. | 1. A method of monitoring cardiac function, comprising the steps of: a. sensing an electrocardiograph signal, the signal comprising a succession of waves; b. independently identifying the QRS complex with each successive wave continuously, in real time, and in a beat-by-beat manner; c. independently analyzing data from the succession QRS complexes relating to the cardiac function continuously, in real-time, and in a beat-by-beat manner, wherein the step of analyzing is comprised of band-pass filtering the data in a high frequency range between about 150 Hz and about 250 Hz; and d. displaying predetermined characteristics of all data contiuously, in real-time, and in beat-by-beat manner. | 1. A method of monitoring cardiac function, comprising the steps of: a. sensing an electrocardiograph signal, the signal comprising a succession of waves; b. independently identifying the QRS complex with each successive wave continuously, in real time, and in a beat-by-beat manner; c. independently analyzing data from the succession QRS complexes relating to the cardiac function continuously, in real-time, and in a beat-by-beat manner, wherein the step of analyzing is comprised of band-pass filtering the data in a high frequency range between about 150 Hz and about 250 Hz; and d. displaying predetermined characteristics of all data contiuously, in real-time, and in beat-by-beat manner. 2. The method of claim 1, further comprising the steps of developing measures related to the data from the succession QRS complexes and wherein the step of displaying is comprised of displaying the measures on a monitor. 3. The method of claim 1, wherein the step of analyzing is further comprised of: generating an initial template of the signal using the first QRS complex from the first wave of the signal; cross-correlating each subsequent independently identified QRS complex continuously and in a beat-by-beat manner aginst predetermined criteria wherein accepted QRS complexes are incorporated into the initial template to create a running signal average template in real-time; aligning each independently identified QRS complex with the running signal average template continuously, in real-time, and in a beat-by-beat manner; and calculating predetermined, selectable high frequency characteristics from the running signal average template in a continuous, beat-by-beatm, real-time manner. 4. The method of claim 3, wherein the step of calculating is comprised of calculating a reduced amplitude zone, wherein the reduced amplitude zone occurs when at least two local maxima or at least two local minima are present in the filtered data; and wherein each local maximum or each local minimum is a peak or a trough, respectively and wherein the absolute value of each peak or trough's voltage exceeds that of a predetermined number of sample points within the filtered data immediately preceding and following each peak or trough. 5. The method of claim 4, wherein the absolute value of the amplitude of the smallest of the at least two local maxima or the smallest of the at least two minima is at least a predetermined percentage relative to the largest of the at least two maxima or the largest fo the at least two minima, respectively. 6. The method of claim 4, wherein the absolute value of the amplitude of the smallest of the at least two local maxima or the smallest of the at least two local minima is at least a predetermined multiple of the high frequency noise level of a predetermined iso-electric portion. 7. The method of claim 3, wherein the step of calculating is comprised of calculating a skewness, kurtosis, or both of the filtered data. 8. The method of claim 1, whe... | The United States of America as represented by the Administration of the National Aeronautics and Space Administration | A61B005/0472
US07013175 | B2 | 2006-03-14 | Systems and methods for assessing and modifying an individual's physiological condition | 2002-05-21 | Systems and methods for assessing an individual's physiological condition are provided. Cycle and shape parameters are derived from a recorded time trace containing heart rate data collected while an individual performs a cyclic exercise routine. Individually tailored exercise regimens that are based on these parameters are generated and modified as desired. | 1. A computer-readable medium containing instructions for controlling a computing device for analyzing an electronic file containing a time trace of a heart rate of an individual obtained during an exercise regimen that includes at least one exercise session that itself comprises at least one exercise cycle, said instructions comprising: instructions to determine at least one cycle parameter from said time trace;instructions to compute a value based on said at least one cycle parameter; andinstructions to assign said value to a heart wave index, wherein said instructions to determine comprises instructions for statistically fitting a parabola through a portion of said time trace, said fit characterized by at least one goodness of fit parameter. | 1. A computer-readable medium containing instructions for controlling a computing device for analyzing an electronic file containing a time trace of a heart rate of an individual obtained during an exercise regimen that includes at least one exercise session that itself comprises at least one exercise cycle, said instructions comprising: instructions to determine at least one cycle parameter from said time trace;instructions to compute a value based on said at least one cycle parameter; andinstructions to assign said value to a heart wave index, wherein said instructions to determine comprises instructions for statistically fitting a parabola through a portion of said time trace, said fit characterized by at least one goodness of fit parameter. 2. The computer-readable medium of claim 1 further comprising instructions to diagnose abnormal physiological conditions by comparing said goodness of fit parameter with information on one or more of said conditions using an electronic catalog, said catalog relating said goodness of fit parameter to at least one physiological condition based on said comparison. 3. The computer-readable medium of claim 1 wherein said instructions to assign comprises instructions to assign to said heart wave index a value based on said at least one cycle parameter selected from a group consisting of a maximum heart rate, a peak heart rate, an upward slope, a downward slope, a minimum heart rate, a resting heart rate, a baseline slope, and a set of goodness of fit parameters. 4. The computer-readable medium of containing instructions for controlling a computing device for analyzing an electronic file containing a time trace of a heart rate of an individual obtained during an exercise regimen that includes at least one exercise session that itself comprises at least one exercise cycle, said instructions comprising: instructions to determine at least one cycle parameter from said time trace;instructions to compute a value based on said at least one cycle parameter;instructions to assign said value to a heart wave index; andinstructions to modify said exercise regimen for said individual, said regimen having a target heart wave index and said at least one exercise cycle having a target heart rate, further comprising: instructions to compare said heart wave index to said target heart wave index; andinstructions to change said target heart wave index and said target heart rate in response to said comparison. 5. The computer-readable medium of claim 4 wherein said instructions to determine comprises determining at least one parameter selected from a group consisting of a maximum heart rate, a peak heart rate, an upward slope, a downward slope, a resting heart rate, a minimum heart rate, and a baseline slope. | LifeWaves International, Inc. | A61B005/024
US07200431 | B2 | 2007-04-03 | Implantable blood flow monitoring system | 2002-08-02 | An implantable blood flow monitoring system for providing synchronized blood vessel flow and myocardial wall contractility data to an external monitor independent of transcutaneous leads. Synchronized electrocardiogram data allows and provides comprehensive monitoring. A means is provided for transmitting synchronized cardiac function data and blood flow data to a distant remote location to facilitate continual physician monitoring. | 1. An improved method of monitoring blood flow in coronary vessels and ventricular function with an implanted blood flow transit sensor and contractility piezoelectric crystal sensors having electrical leads extending external to a patient body and an external electrocardiogram for measuring a cardiac function signal wherein the improvement lies in: communicating signals of the implanted blood flow transit sensor and at least two implanted piezoelectric crystal contractility sensors to an external monitor by an implanted radio frequency transceiver independent of transcutaneous leads. | 1. An improved method of monitoring blood flow in coronary vessels and ventricular function with an implanted blood flow transit sensor and contractility piezoelectric crystal sensors having electrical leads extending external to a patient body and an external electrocardiogram for measuring a cardiac function signal wherein the improvement lies in: communicating signals of the implanted blood flow transit sensor and at least two implanted piezoelectric crystal contractility sensors to an external monitor by an implanted radio frequency transceiver independent of transcutaneous leads. 2. The improved method of claim 1 further comprising synchronizing signals of at least two sensors associated with blood flow, contractility and cardiac function. 3. The improved method of claim 2 further comprising communicating the at least two synchronized sensor signals to a distant radio frequency receiver. 4. A method of screening drugs for vascular effects comprising: implanting a blood flow monitoring system operating according to the method of claim 1;collecting blood flow data from a patient independent of a drug;collecting blood flow data from said patient in the presence of said drug; andcomparing blood flow data collected from said patient independent of said drug and blood flow data collected from said patient in the presence of said drug, for the purpose of screening drugs for vascular effects. 5. A method of determining the effects of activity on blood flow parameters comprising: implanting a blood flow monitoring system operating according to the method of claim 1;collecting blood flow data from a patient prior to an activity;collecting blood flow data from said patient during or subsequent to said activity; andcomparing blood flow data collected from said patient prior to said activity and blood flow data collected from said patient during or subsequent to said activity, to determine the effects of activity on blood flow parameters. | - | A61B005/02
US07039452 | B2 | 2006-05-02 | Method and apparatus for Raman imaging of macular pigments | 2002-01-07 | A method and apparatus are provided for Raman imaging of carotenoids and related chemical substances in biological tissue, such as macular pigments. The method and apparatus utilize the technique of resonance Raman spectroscopy to produce an image of the levels of carotenoids and similar substances in tissue. In this technique, light is directed upon the area of tissue which is of interest such as the retina of an eye. A small fraction of the scattered light is scattered inelastically, producing a carotenoid Raman signal which is at a different frequency than the incident light. The Raman signal is collected, filtered, and analyzed to determine the spatial position and intensity of the Raman signals in the inelastically scattered light. An image of the Raman signals is then produced on an output device, with the image representing the spatial distribution and concentration level of carotenoids in the tissue. | 1. A method for imaging the spatial distribution and concentration level of macular carotenoids, the method comprising: obtaining a light source that generates light at a wavelength that produces a Raman response with a wavelength shift for one or more macular carotenoids to be detected;directing light from the light source onto macular tissue of an eye for which macular carotenoid levels are to be measured, wherein the light directed onto the macular tissue of the eye has an intensity that does not destroy the macular tissue;collecting light scattered from the macular tissue, the scattered light including elastically and inelastically scattered light, the inelastically scattered light having a plurality of Raman signals corresponding to the one or more macular carotenoids;selectively removing the elastically scattered light;analyzing the spatial position and intensity of the Raman signals in the inelastically scattered light; andproducing an image of the Raman signals, the image representing the spatial distribution and concentration level of the one or more macular carotenoids in the macular tissue at a plurality of data points. | 1. A method for imaging the spatial distribution and concentration level of macular carotenoids, the method comprising: obtaining a light source that generates light at a wavelength that produces a Raman response with a wavelength shift for one or more macular carotenoids to be detected;directing light from the light source onto macular tissue of an eye for which macular carotenoid levels are to be measured, wherein the light directed onto the macular tissue of the eye has an intensity that does not destroy the macular tissue;collecting light scattered from the macular tissue, the scattered light including elastically and inelastically scattered light, the inelastically scattered light having a plurality of Raman signals corresponding to the one or more macular carotenoids;selectively removing the elastically scattered light;analyzing the spatial position and intensity of the Raman signals in the inelastically scattered light; andproducing an image of the Raman signals, the image representing the spatial distribution and concentration level of the one or more macular carotenoids in the macular tissue at a plurality of data points. 2. The method of claim 1, wherein the light source generates light at a wavelength that overlaps the absorption bands of the one or more macular carotenoids to be detected. 3. The method of claim 1, wherein the light source generates light in a wavelength range from about 350 nm to about 550 nm. 4. The method of claim 1, wherein the light from the light source has an intensity that does not destroy the macular tissue and does not substantially alter carotenoid levels in the macular tissue. 5. The method of claim 1, wherein the light source generates light at an exposure spot size of about 5 microns to about 10 mm. 6. The method of claim 1, wherein the light source generates light with an exposure time of about 0.001 to about 100 seconds. 7. The method of claim 1, wherein the macular tissue resides in a live subject. 8. The method of claim 1, wherein the inelastically scattered light is analyzed at frequencies characteristic of macular carotenoids. 9. The method of claim 1, wherein the image of the Raman signals is an en face map. 10. The method of claim 1, wherein the image of the Raman signals is a topographical surface plot. 11. A method for imaging the spatial distribution and concentration level of carotenoids in biological tissue, the method comprising: obtaining a light source that generates light at a wavelength that produces a Raman response with a wavelength shift for one or more carotenoids to be detected;directing light from the light source onto biological tissue for which carotenoid levels are to be measured, wherein the light directed onto the biological tissue has an intensity that does not destroy the biological tissue;collecting light scattered from the biological tissue, the scattered light including elastically and inelastically scattered light, the inelastically scattered light having a plu... | The University of Utah Research Foundation | A61B005/05
US07117029 | B2 | 2006-10-03 | Method of and apparatus for deriving indices characterizing atrial arrhythmias | 2002-09-30 | In a method of and apparatus for deriving indices characterizing atrial arrhythmias, a temporally isolated portion of a digital representation of a time-domain electrical signal originating from the atria of a patient is frequency analyzed to provide a corresponding frequency spectrum. The obtained frequency spectrum is compared with a template frequency spectrum, selectively adapted dependent on previously obtained frequency spectra from different temporally isolated portions, to determine a frequency shift value and an amplitude value that optimize a correlation between the two spectra. An output representing the two determined values is used, possibly together with parameters describing the template frequency spectrum, to classify an atrial arrhythmia manifest in the digital representation. | 1. A computer-readable medium encoded with data and being loadable in a digital data processing system for deriving indices characterizing atrial arrhythmias, said data programming said digital signal processing system to: acquire a digital representation of a time-domain electrical signal originating from atria of a heart;temporally isolate a portion of the digital representation of the time-domain signal within a variable location time window;frequency analyze the temporally isolated portion to obtain a frequency spectrum having a fundamental frequency peak and at least one harmonic thereof;compare the obtained frequency spectrum with a template frequency spectrum having a template fundamental frequency peak and at least one template harmonic thereof to determine a frequency shift value and an amplitude value that optimize a correlation between the obtained frequency spectrum and the template frequency spectrum; andemit an output representing the frequency shift value and the amplitude value as respective indices characterizing atrial arrhythmias in said electrical signal. | 1. A computer-readable medium encoded with data and being loadable in a digital data processing system for deriving indices characterizing atrial arrhythmias, said data programming said digital signal processing system to: acquire a digital representation of a time-domain electrical signal originating from atria of a heart;temporally isolate a portion of the digital representation of the time-domain signal within a variable location time window;frequency analyze the temporally isolated portion to obtain a frequency spectrum having a fundamental frequency peak and at least one harmonic thereof;compare the obtained frequency spectrum with a template frequency spectrum having a template fundamental frequency peak and at least one template harmonic thereof to determine a frequency shift value and an amplitude value that optimize a correlation between the obtained frequency spectrum and the template frequency spectrum; andemit an output representing the frequency shift value and the amplitude value as respective indices characterizing atrial arrhythmias in said electrical signal. 2. A computer-readable medium as claimed in claim 1 wherein said data further program the digital signal processing system to: adapt the template frequency spectrum by selective averaging with the obtained frequency spectrum frequency shifted to place the fundamental frequency peak coincident with the template fundamental frequency peak. 3. A computer-readable medium as claimed in claim 2 wherein said data further program the digital signal processing system to: generate an averaged frequency spectrum derived from the frequency spectrum and previous frequency spectra obtained from previous temporally isolated portions with differing timewindow locations within the digital representation of the time-domain signal; andadapt the template frequency spectrum by shifting the spectrum to make the template fundamental frequency coincide with a fundamental peak position within the averaged frequency spectrum. 4. A computer-readable medium as claimed in claim 2 wherein said data further program the digital signal processing system to: analyze the template frequency spectrum before adaptation to determine values for at least one parameter selected from the group, consisting of fundamental frequency peak amplitude, exponential decay, and signal-to-noise ratio. 5. A digital data processing system for determining indices characterizing atrial arrhythmias comprising: a storage unit for receiving and holding, for subsequent data processing a digital representation of a time-domain electrical signal originating from atria of a heart;a time window generator operably connected to the storage unit for temporally isolating a portion of the held digital representation within a moveable time window;a spectrum analyzer which frequency analyzes the temporally isolated portion to obtain a frequency spectrum having a fundamental frequency peak and one or more harmonics thereof; anda comparator for ... | Siemens Aktiengesellschaft | A61B005/04
US07123956 | B2 | 2006-10-17 | Visceral fat estimation apparatus | 2002-10-25 | Disclosed is a visceral fat estimation apparatus, comprising: an input device: an estimation device; a target calorie intake determination device; a calorie consumption determination device; an allowable calorie intake update device; and a calorie consumption device. According to the present invention the input device enters at least height, body weight and body fat rate of a person to be measured, and the estimation device estimates a visceral fat area based in the input value entered by the input device. Furthermore, the target calorie intake determination device determines a target calorie intake for a day, and the calorie consumption determination device determines a calorie to be consumed due to an exercise for a day. The allowable calorie intake update device updates an allowable calorie intake for a day, and the calorie consumption update device updates a calorie to be consumed due to an exercise for a day. | 1. A visceral fat estimation apparatus, comprising: an input device; an estimation device; a calorie consumption determination device; an exercise item selection and determination device; an exercise time duration determination device; a calorie consumption update device; a target calorie intake determination device; a food name selection and determination device; an intake amount determination device; an allowable calorie intake update device; and a display device, wherein said input device is for entering at least height, body weight and body fat rate of a person to be measured; said estimation device estimates a visceral fat area based on the input value entered by said input device; said calorie consumption determination device determines a calorie amount to be consumed by exercise for a day based on a visceral fat area to be reduced from said visceral fat area estimated by said estimation device; said exercise item selection and determination device comprises exercise item altering and decision keys and selects and determines an exercise item to be done using the exercise item altering and decision keys; said exercise time duration determination device comprises exercise time altering and decision keys and determines an exercise time duration using the exercise time altering and decision keys; said calorie consumption update device determines a calorie consumption based on said exercise item determined by said exercise item selection and determination device and said exercise time duration determined by said exercise time duration determination device and updates said calorie amount to be consumed by exercise for a day by subtracting said determined calorie consumption from said calorie amount to be consumed by exercise for a day determined by said calorie consumption determination device; said target calorie intake determination device comprises target calorie intake altering and decision keys determines a target calorie intake for a day using the target calorie intake altering and decision keys, based on the input value entered using said input device; said food name selection and determination device comprises food name altering and decision keys and selects and determines a food name to be taken using the food name altering and decision keys; said intake amount determination device comprises intake amount altering and decision keys and determines an intake amount using the intake amount altering and decision keys; said allowable calorie intake update device determines a calorie intake based on said food name determined by said food name selection and determination device and said intake amount determined by said intake amount determination device and updates said allowable calorie intake for a day by subtracting said determined calorie intake from said allowable calorie intake for a day determined by said target calorie intake determination device; and said display device simultaneously displays said calorie amount to be consumed by exer... | 1. A visceral fat estimation apparatus, comprising: an input device; an estimation device; a calorie consumption determination device; an exercise item selection and determination device; an exercise time duration determination device; a calorie consumption update device; a target calorie intake determination device; a food name selection and determination device; an intake amount determination device; an allowable calorie intake update device; and a display device, wherein said input device is for entering at least height, body weight and body fat rate of a person to be measured; said estimation device estimates a visceral fat area based on the input value entered by said input device; said calorie consumption determination device determines a calorie amount to be consumed by exercise for a day based on a visceral fat area to be reduced from said visceral fat area estimated by said estimation device; said exercise item selection and determination device comprises exercise item altering and decision keys and selects and determines an exercise item to be done using the exercise item altering and decision keys; said exercise time duration determination device comprises exercise time altering and decision keys and determines an exercise time duration using the exercise time altering and decision keys; said calorie consumption update device determines a calorie consumption based on said exercise item determined by said exercise item selection and determination device and said exercise time duration determined by said exercise time duration determination device and updates said calorie amount to be consumed by exercise for a day by subtracting said determined calorie consumption from said calorie amount to be consumed by exercise for a day determined by said calorie consumption determination device; said target calorie intake determination device comprises target calorie intake altering and decision keys determines a target calorie intake for a day using the target calorie intake altering and decision keys, based on the input value entered using said input device; said food name selection and determination device comprises food name altering and decision keys and selects and determines a food name to be taken using the food name altering and decision keys; said intake amount determination device comprises intake amount altering and decision keys and determines an intake amount using the intake amount altering and decision keys; said allowable calorie intake update device determines a calorie intake based on said food name determined by said food name selection and determination device and said intake amount determined by said intake amount determination device and updates said allowable calorie intake for a day by subtracting said determined calorie intake from said allowable calorie intake for a day determined by said target calorie intake determination device; and said display device simultaneously displays said calorie amount to be consumed by exer... | Tanita Corporation | A61B005/05
US07187962 | B2 | 2007-03-06 | Biological optical measuring instrument | 2002-06-28 | In a biological optical measuring instrument, when displaying a biological variation in a predetermined region of a subject through irradiating light to the subject and processing light signals obtained by detecting the transmitting light, the light signals of a plurality of sections having a predetermined time interval are extracted, section data are prepared from the respective extracted light signals, the plurality of the prepared section data are added and averaged, prior to preparing averaged section data for the predetermined time interval a light signal section containing noise components is specified from differential values of the extracted light signals and the section data of the light signal containing the noise components is excluded from the averaging operation, thereby, a biological optical measuring instrument is provided which permits to remove the section data containing noise components are automatically removed through computation without relying upon visual observation, to obtain highly reliable measurement result quickly and to display the same. | 1. A biological optical measuring instrument comprising a light generating unit which irradiates light having a predetermined wavelength to a subject, a light detecting unit which detects light which is irradiated from the light generating unit and is transmitted through the subject, and a signal processing unit which analyzes signals detected by the light detecting unit and prepares biological information including blood stream information in the subject, wherein the signal processing unit includes means for detecting spike shaped noise components contained in the detected signals in a form of steep peaks which are due to body movement by the subject, and means for removing a signal portion containing the detected spike shaped noise components from the signals. | 1. A biological optical measuring instrument comprising a light generating unit which irradiates light having a predetermined wavelength to a subject, a light detecting unit which detects light which is irradiated from the light generating unit and is transmitted through the subject, and a signal processing unit which analyzes signals detected by the light detecting unit and prepares biological information including blood stream information in the subject, wherein the signal processing unit includes means for detecting spike shaped noise components contained in the detected signals in a form of steep peaks which are due to body movement by the subject, and means for removing a signal portion containing the detected spike shaped noise components from the signals. 2. A biological optical measuring instrument according to claim 1, wherein the means for detecting the spiked shaped noise components judges that the spiked shape noise components are contained in a corresponding portion of the signals, when any of differential values of the signals obtained in a time course, are more than a predetermined threshold value. 3. A biological optical measuring instrument according to claim 1 or 2, wherein the signal processing unit includes means for extracting signals in a plurality of sections having a predetermined time interval from signals obtained along the time axis and for preparing section data from the respective extracted signals, means for detecting and removing section data containing noise components among the extracted and prepared section data, and means for performing adding and averaging process using the not removed section data among the plurality of the section data and for preparing averaged section data having a predetermined time interval. 4. A biological optical measuring instrument according to claim 3, wherein the section having a predetermined time interval includes a time interval, in which a task is executed by the subject. 5. A biological optical measuring instrument according to claim 4, wherein the task executed by the subject is motion and language. 6. A biological optical measuring instrument according to claim 3, wherein the light having a predetermined wavelength is near-infrared light. 7. A biological optical measuring instrument according to claim 1 or 2, wherein the light having a predetermined wavelength is near-infrared light. 8. A biological optical measuring method comprising: a first step of irradiating light having a predetermined wavelength to a subject from a first position of the subject and measuring the transmitting light at a second position of the subject as transmitting light signal data containing biological information in a plurality of times for a predetermined interval;a second step of adding the plurality of measured transmitting light signal data for the predetermined time interval and determining an average value of the transmitting light signal data;a third step of determining differences of ... | Hitachi Medical Corporation | A61B005/00
US07035683 | B2 | 2006-04-25 | Method for respiration-dependent triggering of a medical image exposure | 2002-07-31 | In a method for respiration-dependent triggering of a data acquisition event in an imaging procedure, a navigation system is employed in order to reproduce the position of a body part moving due to respiration on a scale, particularly a linear scale. Triggering for the exposure of an image ensues at prescribed positions of a pointer relative to the scale, i.e. at prescribed scale values. | 1. A method for respiration-dependent triggering of production of a medical image, comprising the steps of: designating a position relative to a scale of a body part of a subject, said body part moving due to respiration of the subject by securing a marker, separate from said scale, to said body part and indicating a position of said marker on said scale; andtriggering at least one data acquisition event for obtaining an image of said subject when said designation of said body part is at a prescribed position relative to said scale. | 1. A method for respiration-dependent triggering of production of a medical image, comprising the steps of: designating a position relative to a scale of a body part of a subject, said body part moving due to respiration of the subject by securing a marker, separate from said scale, to said body part and indicating a position of said marker on said scale; andtriggering at least one data acquisition event for obtaining an image of said subject when said designation of said body part is at a prescribed position relative to said scale. 2. A method as claimed in claim 1 comprising employing a linear scale as said scale. 3. A method as claimed in claim 1 comprising employing a scale having a plurality of scale values, and wherein the step of triggering said at least one data acquisition event comprises triggering said at least one data acquisition event when said designation of said movement of said body part is at a predetermined scale value. 4. A method for respiration-dependent trigging of production of a medical image, comprising: designating a position relative to a scale of a body part of a subject, said body part moving due to respiration of the subject;triggering at least one data acquisition event for obtaining an image of said subject when said designation of said body part is at a prescribed position relative to said scale;introducing an intervention instrument into said subject toward a target point at a fixed location in said image; andtriggering said introduction of said intervention instrument, in a subsequent motion cycle of said subject, when said designation of said body part in said subsequent cycle is at said prescribed position relative to said scale. 5. A method as claimed in claim 4 comprising introducing said intervention instrument into said subject in a plurality of successive steps, and initiating each of said plurality of successive steps, in respective subsequent motion cycles of said subject, when said designation of the position of said body part is at said prescribed position relative to said scale. 6. A method as claimed in claim 4 wherein said intervention instrument has a tip, and identifying an angle of said tip in said image relative to said starting position. 7. A method as claimed in claim 6 comprising additionally determining a distance of said tip from said known starting position. 8. A method as claimed in claim 6 comprising continuously monitoring and identifying a position of said tip of said intervention instrument using a position acquisition system and mixing a designation of said tip into said image and, in respective subsequent motion cycles of said subject, briefly displacing said intervention instrument into said subject when said designation of said body part is at said prescribed position relative to said scale. 9. A method as claimed in claim 1 wherein the step of indicating a position of said marker on said scale comprises detecting, with a position acquisition system, a physical position ... | Siemens Aktiengesellschaft | A61B005/05
US07139602 | B2 | 2006-11-21 | Shape analysis of surfaces | 2002-04-26 | A method and apparatus are provided to map the shape or thickness of a surface. Thickness estimates across a distance are received, and Fourier analysis is performed to generate Fourier amplitude coefficients and phase values for a number of frequency components. Linear discriminant functions are used to characterize the resulting shape as normal or not. Discriminant functions are provided that aid in the diagnosis of glaucoma when the thickness measured is the retinal nerve fiber layer thickness. | 1. A method comprising: receiving waveform data representing a shape of a surface;determining a plurality of Fourier coefficient amplitudes and phase values corresponding to the waveform data;entering at least one phase value into a discriminant measure; andcomparing a result of the discriminant measure against a threshold. | 1. A method comprising: receiving waveform data representing a shape of a surface;determining a plurality of Fourier coefficient amplitudes and phase values corresponding to the waveform data;entering at least one phase value into a discriminant measure; andcomparing a result of the discriminant measure against a threshold. 2. The method of claim 1 wherein the discriminant measure comprises an equation that includes the at least one phase value, and also includes at least one Fourier coefficient amplitude. 3. The method of claim 1 wherein the waveform data represents the shape of a nerve fiber layer in a hemiretina. 4. The method of claim 1 wherein the waveform data represents the shape of a nerve fiber layer in a ring around an entire retina. 5. The method of claim 1 wherein determining a plurality of Fourier coefficient amplitudes and phase values comprises performing a fast Fourier transform (FFT). 6. The method of claim 1 wherein determining a plurality of Fourier coefficient amplitudes and phase values comprises performing a discrete Fourier transform (DFT). 7. The method of claim 1 further comprising generating interocular asymmetry measures by taking a ratio of each Fourier coefficient amplitude from one eye to a corresponding Fourier coefficient amplitude from another eye. 8. The method of claim 7 wherein the discriminant measure comprises an equation that includes the at least one phase value, and also includes at least one interocular asymmetry measure. 9. The method of claim 1 further comprising generating intraocular asymmetry measures by comparing a sum of Fourier coefficient amplitudes from one hemiretina to a sum of Fourier coefficient amplitudes from another hemiretina. 10. The method of claim 9 wherein the discriminant measure comprises an equation that includes the at least one phase value, and also includes at least one intraocular asymmetry measure. | University of Louisville Research Foundation, Inc. | A61B005/05
US07221973 | B2 | 2007-05-22 | Magnetic resonance tomography method and apparatus with motion correction | 2002-04-09 | In a magnetic resonance tomography apparatus and method with motion correction in an angiography examination with magnetic resonance-monitored vessel intervention, a contrast agent-supported exposure of the vessel system (road map) is registered, with one exposure ensuing at the end of expiration and one exposure ensuing at the end of inspiration of the patient, and a magnetic resonance-visible medical intervention device introduced into the vessel system also is registered. A corrected road map of the vessel system is interpolated from the exposures at the end of inspiration and the end of expiration and the exposure of the intervention device is superimposed on the corrected road map. | 1. A magnetic resonance tomography apparatus comprising: a magnetic resonance data acquisition system which registers a contrast agent-supported exposure of a vessel system of a subject in an angiography examination, including a first exposure at an end of expiration of said subject and a second exposure at an end of inspiration of said subject, and real time exposures of a magnetic resonance-visible medical intervention device introduced into said vessel system; anda motion-correction unit connected to said magnetic resonance data acquisition system which interpolates a motion-corrected exposure of said vessel system from said first and second exposures and superimposes a current exposure of said medical intervention device and said motion-corrected exposure of said vessel system. | 1. A magnetic resonance tomography apparatus comprising: a magnetic resonance data acquisition system which registers a contrast agent-supported exposure of a vessel system of a subject in an angiography examination, including a first exposure at an end of expiration of said subject and a second exposure at an end of inspiration of said subject, and real time exposures of a magnetic resonance-visible medical intervention device introduced into said vessel system; anda motion-correction unit connected to said magnetic resonance data acquisition system which interpolates a motion-corrected exposure of said vessel system from said first and second exposures and superimposes a current exposure of said medical intervention device and said motion-corrected exposure of said vessel system. 2. A magnetic resonance tomography apparatus as claimed in claim 1 further comprising a respiratory belt adapted for application to said subject for acquiring a respiratory signal from said subject, and wherein said motion-correction unit is supplied with said respiratory signal and generates a curve therefrom representing respiratory motion of said subject and determines, from said curve, a current position for said motion-corrected exposure of said vessel system. 3. A magnetic resonance tomography apparatus as claimed in claim 2 wherein said motion-correction unit generates said motion-corrected exposure of said vessel system by linear interpolation between a time of said first exposure, determined from said curve, and a time of said second exposure, determined from said curve. 4. A magnetic resonance tomography apparatus as claimed in claim 2 wherein said motion-correction unit generates said motion-corrected exposure of said vessel system by interpolating between a time of said first exposure, determined from said curve, and a time of said second exposure, determined by said curve, weighted according to said curve. 5. A magnetic resonance tomography apparatus as claimed in claim 1 wherein said magnetic resonance data acquisition system acquires magnetic resonance signals from a navigator rod substantially perpendicularly intersecting a diaphragm of said subject, and wherein said motion-correction unit generates a curve from said magnetic resonance signals from said navigator rod representing respiratory motion of said subject, and determines a current position for said motion-corrected exposure of said vessel system along said curve. 6. A magnetic resonance tomography apparatus as claimed in claim 5 wherein said motion-correction unit generates said motion-corrected exposure of said vessel system by linear interpolation between a time of said first exposure, determined from said curve, and a time of said second exposure, determined from said curve. 7. A magnetic resonance tomography apparatus as claimed in claim 5 wherein said motion-correction unit generates said motion-corrected exposure of said vessel system by interpolating between a time of said first ex... | Siemens Aktiengesellschaft | A61B005/055
US07120484 | B2 | 2006-10-10 | Methods and apparatus for filtering EGM signals detected by an implantable medical device | 2002-01-14 | An analog physiologic signal, e.g., the cardiac EGM, sensed by an IMD is filtered with a high pass filter (HPF), the cut-off frequency of the HPF being within a predetermined frequency bandwidth, wherein a low-band portion of the predetermined frequency bandwidth is attenuated in the filtered physiologic signal. The filtered physiologic signal is digitized in real time order, and the digital data set is filtered in reverse time order employing a digital IIR filter having characteristics substantially matching the cut-off frequency and filter characteristics of the HPF. When the system is implemented within an IMD, the filtered digital data set is compressed by lossy compression algorithm, and the compressed data set is filtered in reverse time order. In certain embodiments, the filtered digital data set is uplink telemetry transmitted to an external medical device, and the uplink telemetered data set is filtered in reverse time order employing a digital backward IIR filter resident in the external medical device. | 1. A method for use by an Implantable Medical Device (IMD) of filtering an analog cardiac electrogram (EGM) signal having a predetermined frequency bandwidth, the method comprising: obtaining an analog cardiac electrogram (EGM) signal from an implantable sensor coupled to the IMD;filtering the EGM signal with a high pass filter (HPF) having a cut-of frequency within a predetermined frequency bandwidth, and wherein a low-frequency band portion of the predetermined frequency bandwidth is attenuated in the filtered EGM signal;sampling and digitizing the filtered EGM signal as a digital data set in a time order and optionally, generating a compressed digital data set having at least a 2:1 data compression factor;filtering one of said digital data set or said compressed data set in reverse time order employing a digital Infinite Impulse Response (IIR) filter having characteristics substantially matching the cut-off frequency and filter characteristics of the HPF to substantially remove distortions of the filtered analog EGM signal introduced by the HPF; andanalyzing a reconstituted EGM to determine whether cardiac ischemia is present. | 1. A method for use by an Implantable Medical Device (IMD) of filtering an analog cardiac electrogram (EGM) signal having a predetermined frequency bandwidth, the method comprising: obtaining an analog cardiac electrogram (EGM) signal from an implantable sensor coupled to the IMD;filtering the EGM signal with a high pass filter (HPF) having a cut-of frequency within a predetermined frequency bandwidth, and wherein a low-frequency band portion of the predetermined frequency bandwidth is attenuated in the filtered EGM signal;sampling and digitizing the filtered EGM signal as a digital data set in a time order and optionally, generating a compressed digital data set having at least a 2:1 data compression factor;filtering one of said digital data set or said compressed data set in reverse time order employing a digital Infinite Impulse Response (IIR) filter having characteristics substantially matching the cut-off frequency and filter characteristics of the HPF to substantially remove distortions of the filtered analog EGM signal introduced by the HPF; andanalyzing a reconstituted EGM to determine whether cardiac ischemia is present. 2. The method of claim 1, wherein the method is implemented using programmed instructions. 3. The method of claim 2, wherein the programmed instructions are executed on a processing circuit included within the IMD. 4. The method of claim 2, wherein the programmed instructions are executed on a processing circuit located external to the IMD. 5. The method of claim 1, wherein the EGM signal is generated using at least one of: an endocardial electrode, an epicardial electrode, a subcutaneously implanted electrode, an electrode mechanically coupled to a portion of a housing for the IMD. 6. The method of claim 1, further comprising the step of storing the reverse time order filtered digital data set in a memory structure of the IMD. 7. The method of claim 1, further comprising the steps of: storing the digital data set in memory of the IMD; anduplink telemetry transmitting the stored digital data set to an external medical device (EMD); andwherein the reverse time order filtering step of filtering the digital data set in reverse time order employing a digital IIR filter having characteristics substantially matching the cut-off frequency and filter characteristics of the HPF is conducted by the EMD. 8. Apparatus for use by an Implantable Medical Device (IMD) for filtering a physiological signal having a predetermined frequency bandwidth, the method comprising: a sensor coupled to the IMD for sensing an analog cardiac electrogram (EGM) signal;a high pass filter (HPF) for filtering the analog EGM signal, the HPF having a the cut-off frequency within the predetermined frequency bandwidth, wherein a low-band portion of the predetermined frequency bandwidth is attenuated in the filtered analog EGM signal;means for sampling and digitizing the filtered analog EGM signal as a digital data set In a real time order;means for comp... | Medtronic, Inc. | A61B005/04
US07087017 | B2 | 2006-08-08 | Atraumatic sensor lead assemblies | 2002-10-31 | The invention is directed to lead configurations for sensors that allow for less invasive sensor replacement procedures. In one configuration, a sensor lead assembly includes an outer lead body and an inner lead including a sensor such as an electrochemical glucose sensor. The inner lead can be positioned in an inner conduit of the outer lead body. The outer lead body may be substantially permanently implanted in the patient, and the inner lead can be implanted through the inner conduit of the outer lead body. Once the sensor of the inner lead has worn out or otherwise exhausted its useful life, the inner lead can be removed, and a new inner lead can be implanted in place of the old inner lead. | 1. A method comprising: coupling a first end of a sealing structure to a first lead;coupling a second end of the sealing structure to a connector module;removing the first lead including a first sensor from an implanted outer lead body through a channel in a connector module of an implantable medical device housing; andinserting a second lead including a second sensor into the implanted outer lead body. | 1. A method comprising: coupling a first end of a sealing structure to a first lead;coupling a second end of the sealing structure to a connector module;removing the first lead including a first sensor from an implanted outer lead body through a channel in a connector module of an implantable medical device housing; andinserting a second lead including a second sensor into the implanted outer lead body. | Medtronic, Inc. | A61B005/00
US07206629 | B2 | 2007-04-17 | Method for implementing a dynamic magnetic resonance measurement with the application of a contrast agent | 2003-01-08 | In a method for implementing a dynamic magnetic resonance measurement with contrast agent administration, at least one imaging magnetic resonance measurement of a subject region is implemented using an inversion recovery sequence at a first and possibly further points in time after administration of the contrast agent, to obtain a magnetic resonance image of the subject region. Immediately before the first and, if used, the further points in time, a pre-measurement with a magnetic resonance sequence is implemented wherein the T1 time of a tissue type, or a substance type that is of no interest in the magnetic resonance image of the subject region, is determined, and an inversion time TI used for the following imaging magnetic resonance measurement is calculated as TI=In2 ×T1. | 1. A method for implementing a dynamic magnetic resonance measurement with application of a contrast agent, comprising the steps of: administering a contrast agent to an examination subject for subsequently obtaining a first contrast agent-enhanced image of a region of said subject containing a substance type not enhanced by said contrast agent having a T1 time associated therewith;conducting a pre-measurement of said region, after administration of said contrast agent, with a magnetic resonance sequence to obtain said first contrast agent-enhanced image and identifying said T1 time of said substance type from said first contrast agent-enhanced image of said pre-measurement;determining an inversion time TI as In2×T1; andimmediately after said pre-measurement, obtaining a second contrast agent-enhanced magnetic resonance image of said region using an inversion recovery sequence with said inversion TI determined from said pre-measurement. | 1. A method for implementing a dynamic magnetic resonance measurement with application of a contrast agent, comprising the steps of: administering a contrast agent to an examination subject for subsequently obtaining a first contrast agent-enhanced image of a region of said subject containing a substance type not enhanced by said contrast agent having a T1 time associated therewith;conducting a pre-measurement of said region, after administration of said contrast agent, with a magnetic resonance sequence to obtain said first contrast agent-enhanced image and identifying said T1 time of said substance type from said first contrast agent-enhanced image of said pre-measurement;determining an inversion time TI as In2×T1; andimmediately after said pre-measurement, obtaining a second contrast agent-enhanced magnetic resonance image of said region using an inversion recovery sequence with said inversion TI determined from said pre-measurement. 2. A method as claimed in claim 1 comprising obtaining a plurality of contrast agent-enhanced magnetic resonance images of said region, using an inversion recovery sequence, following application of said contrast agent and conducting a respective pre-measurement immediately preceding each of said contrast agent-enhanced magnetic resonance measurements and determining an inversion time T1 after each pre-measurement and using the respective inversion times T1 in the inversion recovery sequences used to obtain the plurality of contrast agent-enhanced magnetic resonance images. 3. A method as claimed in claim 1 comprising the additional step of, from said contrast agent-enhanced magnetic resonance measurement, determining myocardium vitality using the late enhanced effect, and wherein said substance type for which said T1 time is determined is vital myocardium tissue. 4. A method as claimed in claim 1 comprising marking said substance type in said magnetic resonance image obtained in said pre-measurement and determining said T1 time only for the marked substance type. 5. A method as claimed in claim 1 comprising conducting said pre-measurement as a single-shot magnetic resonance measurement. 6. A method as claimed in claim 5 comprising employing an inversion recovery TrueFisp sequence for said pre-measurement. | Siemens Aktiengesellschaft | A61B005/05
US07188151 | B2 | 2007-03-06 | System and method for real-time monitoring, assessment, analysis, retrieval, and storage of physiological data over a wide area network | 2002-03-28 | A system and method are provided for network-based monitoring of physiological data. At least one patient-side device collects physiological data from a patient. A provider-side device receives the data from at least one patient-side device via a wide area network. An engine communicates with at least one patient-side device and the provider-side device. The engine manages transmission of the data from the patient-side device to the provider-side device. | 1. A system for network-based transmission of data from a patient-side physiological collecting device, comprising: at least one patient-side device for collecting physiological data from a patient, the patient-side device being operatively coupled to a wide area network;multiple provider-side devices coupled to the network, the at least one patient-side device sending the data to the provider-side devices;an engine coupled to the network and resident on a device intermediate the at least one patient-side device and the provider-side devices, the engine communicating with the at least one patient-side device and the provider-side devices, the engine managing transmission of the data from the patient-side device to the provider-side devices;wherein the data is monitored for an event, wherein an alert is output by at least one of the patient-side device and the provider-side devices upon occurrence of the event,wherein the at least one patient-side device encrypts the data prior to sending the data;wherein the at least one patient-side device compresses the data prior to sending the data;wherein a fee is charged for coordinating the sending of the data to the provider-side devices,wherein a fee is charged by a human provider viewing the data on one of the provider-side devices,wherein the at least one patient-side device includes a browser-based interface for allowing the patient to send human language communications to the provider-side devices,wherein the provider-side devices each include a browser-based interface for allowing one or more providers to send human language communications to the patient-side device,wherein one or more providers is allowed to remotely control the at least one patient-side device,wherein one or more providers is allowed to remotely control at least a portion of the browser based interface of the at least one patient-side device,wherein clinical notes are stored on the device intermediate the at least one patient-side device and the provider-side devices,wherein the at least one patient-side device automatically detects medical devices coupled thereto,wherein the at least one patient-side device is capable of connecting to any medical device having communications capabilities. | 1. A system for network-based transmission of data from a patient-side physiological collecting device, comprising: at least one patient-side device for collecting physiological data from a patient, the patient-side device being operatively coupled to a wide area network;multiple provider-side devices coupled to the network, the at least one patient-side device sending the data to the provider-side devices;an engine coupled to the network and resident on a device intermediate the at least one patient-side device and the provider-side devices, the engine communicating with the at least one patient-side device and the provider-side devices, the engine managing transmission of the data from the patient-side device to the provider-side devices;wherein the data is monitored for an event, wherein an alert is output by at least one of the patient-side device and the provider-side devices upon occurrence of the event,wherein the at least one patient-side device encrypts the data prior to sending the data;wherein the at least one patient-side device compresses the data prior to sending the data;wherein a fee is charged for coordinating the sending of the data to the provider-side devices,wherein a fee is charged by a human provider viewing the data on one of the provider-side devices,wherein the at least one patient-side device includes a browser-based interface for allowing the patient to send human language communications to the provider-side devices,wherein the provider-side devices each include a browser-based interface for allowing one or more providers to send human language communications to the patient-side device,wherein one or more providers is allowed to remotely control the at least one patient-side device,wherein one or more providers is allowed to remotely control at least a portion of the browser based interface of the at least one patient-side device,wherein clinical notes are stored on the device intermediate the at least one patient-side device and the provider-side devices,wherein the at least one patient-side device automatically detects medical devices coupled thereto,wherein the at least one patient-side device is capable of connecting to any medical device having communications capabilities. 2. The system as recited in claim 1, further comprising a storage device for storing the data. 3. The system as recited in claim 1, wherein the patient-side device includes at least one of: a medical device for collecting biomedical data from the patient and a biofeedback device for collecting biofeedback data from the patient. 4. The system as recited in claim 3, further comprising a computing device interfacing with the patient-side device, the computing device receiving a software module for managing communication with the patient-side device from the engine. 5. The system as recited in claim 1, wherein the human language communications include at least one of electronic mail, and instant messaging. 6. The system as recited in claim 1, w... | Televital, Inc. | A61B005/00 G06F015/16
US07172561 | B2 | 2007-02-06 | Spreader tensiometer for measuring tension in an intervertebral disc space | 2002-09-20 | The present invention relates to a novel spreader tensiometer adapted for use in measuring tension in an intervertebral disc space. | 1. A tensiometer for measuring tension in an intervertebral disc space between opposed vertebral surfaces, comprising: a) a disc space spreader adapted to enter a disc space and distract the disc space upon rotation, wherein the spreader comprises: i) a distal pedal portion having upper and lower surfaces defining a thickness T therebetween, and lateral surfaces defining a width W therebetween, the upper and lower surfaces being adapted to enter the disc space, the lateral surfaces being adapted for bearing against the opposed vertebral surfaces, the width being greater than the thickness, wherein the lateral surfaces have a convex shape, andii) a proximal portion, andb) a torque gauge adapted to engage the disc space spreader and report a torque required for the rotation of the spreader. | 1. A tensiometer for measuring tension in an intervertebral disc space between opposed vertebral surfaces, comprising: a) a disc space spreader adapted to enter a disc space and distract the disc space upon rotation, wherein the spreader comprises: i) a distal pedal portion having upper and lower surfaces defining a thickness T therebetween, and lateral surfaces defining a width W therebetween, the upper and lower surfaces being adapted to enter the disc space, the lateral surfaces being adapted for bearing against the opposed vertebral surfaces, the width being greater than the thickness, wherein the lateral surfaces have a convex shape, andii) a proximal portion, andb) a torque gauge adapted to engage the disc space spreader and report a torque required for the rotation of the spreader. 2. The tensiometer of claim 1 wherein the width defined by the lateral surfaces of the pedal is between 2 and 4 times the thickness defined by the upper and lower surfaces of the pedal. 3. The tensiometer of claim 2 wherein the width of the pedal is between 6 mm and 20 mm. 4. The tensiometer of claim 1 wherein at least one of the lateral surfaces and the upper surface meet at a junction, wherein the junction has a radius of between 0.5 and 1.5 mm. 5. The tensiometer of claim 1 wherein the pedal portion further comprises a tapered leading edge. 6. The tensiometer of claim 1 wherein the spreader further comprises an intermediate portion sized to provide a device working length of about 200 mm. 7. The tensiometer of claim 1 wherein the proximal portion of the spreader is adapted to engage a distal portion of the torque gauge. 8. The tensiometer of claim 1 wherein the proximal portion of the pedal portion comprises a radially-extending stop. 9. The tensiometer of claim 1 wherein the torque gauge comprises: i) a housing,ii) a twistable longitudinal beam having a distal end fixedly attached to the housing and a proximal end, andiii) a means for measuring a deflection of the beam relative to the housing when a torque is applied to the beam. 10. The tensiometer of claim 9 wherein the beam has a shape adapted to twist about the longitudinal axis in relation to the distal end when a torque is applied to the proximal end. 11. The tensiometer of claim 9 wherein the beam has a length and a greatest cross-sectional dimension and the length is at least 10 times greater than the greatest cross-sectional dimension. 12. The tensiometer of claim 9 wherein the distal end of the beam comprises an attachment flange. 13. The tensiometer of claim 9 wherein the housing comprises a body fixedly attached to the distal portion of the beam. 14. The tensiometer of claim 13 wherein the body forms an outer shell around the beam. 15. The tensiometer of claim 13 wherein the body further includes a flange extending outwardly from the body and forming a proximal face. 16. The tensiometer of claim 15 wherein the flange is located on a proximal portion of an outer shell. 17. The tens... | Depuy Acromed | A61B005/103
US07101339 | B2 | 2006-09-05 | Respiration signal measurement apparatus, systems, and methods | 2002-12-13 | A cardiac rhythm management apparatus and system may include a first sensor to sense an atrial respiration signal, a second sensor to sense a ventricular respiration signal, and a measurement module capable of being communicatively coupled to the first and second sensors to monitor the atrial and ventricular respiration signals and to select a resulting signal as an indication of respiration. An article may cause a machine to implement a method which includes measuring a second signal responsive to a first signal, measuring a third signal responsive to a first signal, determining a respiration-to-cardiac ratio associated with each one of the second and third signals, and providing a resulting signal including a selected portion of a selected one of the second and third signals as an indication of respiration. | 1. An apparatus, comprising: a first port to receive a first respiration signal comprising an atrial respiration signal;a second port to receive a second respiration signal comprising a ventricular respiration signal; anda measurement module capable of being communicatively coupled to the first and second ports to monitor the first and second respiration signals and to select a resulting signal based on the atrial and/or ventricular respiration signal as an indication of respiration. | 1. An apparatus, comprising: a first port to receive a first respiration signal comprising an atrial respiration signal;a second port to receive a second respiration signal comprising a ventricular respiration signal; anda measurement module capable of being communicatively coupled to the first and second ports to monitor the first and second respiration signals and to select a resulting signal based on the atrial and/or ventricular respiration signal as an indication of respiration. 2. The apparatus of claim 1, wherein the measurement module comprises a quadrapolar impedance measurement device capable of being communicatively coupled to the first and second ports. 3. The apparatus of claim 1, wherein the resulting signal includes a selected portion of the first respiration signal. 4. The apparatus of claim 1, wherein the resulting signal includes the first respiration signal. 5. The apparatus of claim 1, wherein the resulting signal comprises a combination of selected portions of the first and second respiration signals. 6. The apparatus of claim 1, further comprising: a first pair of electrodes coupled to the first port, the first pair of electrodes to be disposed within an atrium; anda second pair of electrodes coupled to the second port, the second pair of electrodes to be disposed within a ventricle. 7. The apparatus of claim 6, wherein the first and second pairs of electrodes are each capable of propagating an injected current and sensing at least one resulting voltage. 8. A system, comprising: an apparatus, including a first port to receive a first respiration signal comprising an atrial respiration signal, a second port to receive a second respiration signal comprising a ventricular respiration signal, and a measurement module capable of being communicatively coupled to the first and second ports to monitor the first and second respiration signals and to select a resulting signal based on the atrial and/or ventricular respiration signal as an indication of respiration; anda processor capable of being communicatively coupled to the apparatus. 9. The system of claim 8, wherein the first respiration signal is derived from a voltage measured between a first lead electrode coupled to the first port and one of two pacer electrodes. 10. The system of claim 9, wherein the second respiration signal is derived from a voltage measured between a second lead electrode coupled to the second port and one of the two pacer electrodes. 11. The system of claim 8, further comprising a current injection device capable of being communicatively coupled to the processor. 12. The system of claim 8, further comprising a filter module capable of being communicatively coupled to the measurement module. 13. The system of claim 8, further comprising: a quadrapolar impedance measurement device included in the measurement module, the quadrapolar measurement device capable of measuring a voltage existing between an electrode coupled to the first port and one ... | Cardiac Pacemakers, Inc. | A61B005/08
US06986747 | B2 | 2006-01-17 | Method of measuring the stress or relaxation level of a mammal | 2003-01-29 | A method of measuring the stress or relaxation level of a mammal and a method of measuring the activity of the sympathetic nervous system of a mammal by measuring quantititative levels of deoxyhemoglobin and oxyhemoglobin are disclosed. Preferably, the levels of deoxyhemoglobin and hemoglobin are measured by a noninvasive technique, such as spectroscopy. A method of changing the activity of the sympathetic nervous system of a mammal is also disclosed, wherein the method includes a step of administering an effective amount of sensory regimen to the mammal. The method is useful for humans who are operating vehicles or machinery, who are suffering from cardiovascular disease or related complications, who are pregnant, or who are preparing for sleep. In addition, a method of improving the complexion of the skin of a mammal is disclosed. | 1. A method of measuring the stress or relaxation level of a mammal, comprising the step of: a. measuring the level of oxyhemoglobin and the level of deoxyhemoglobin in the cutaneous blood supply of said mammal prior to exposing said mammal to an acute stressor or relaxor; b. exposing said mammal to said acute stressor or relaxor; c. measuring the level of oxyhemoglobin and the level of deoxyhemoglobin in the cutaneous blood supply of said mammal during or after said exposing step; and d. for step a and step c, calculating and comparing at least one ratio selected from the group consisting of: i. the ratio of the level of deoxyhemoglobin to the level of oxyhemoglobin of said mammal; and ii. the ratio of the level of deoxyhemoglobin to the sum of the level of deoxyhemoglobin and the level of oxyhemoglobin of said mammal. | 1. A method of measuring the stress or relaxation level of a mammal, comprising the step of: a. measuring the level of oxyhemoglobin and the level of deoxyhemoglobin in the cutaneous blood supply of said mammal prior to exposing said mammal to an acute stressor or relaxor; b. exposing said mammal to said acute stressor or relaxor; c. measuring the level of oxyhemoglobin and the level of deoxyhemoglobin in the cutaneous blood supply of said mammal during or after said exposing step; and d. for step a and step c, calculating and comparing at least one ratio selected from the group consisting of: i. the ratio of the level of deoxyhemoglobin to the level of oxyhemoglobin of said mammal; and ii. the ratio of the level of deoxyhemoglobin to the sum of the level of deoxyhemoglobin and the level of oxyhemoglobin of said mammal. 2. A method of measuring the activity of the sympathetic nervous system of a mammal, comprising the steps of: a. measuring the level of oxyhemoglobin and the level of deoxyhemoglobin in the cutaneous blood supply of said mammal prior to exposing said mammal to an acute stressor or relaxor; b. exposing said mammal to said acute stressor or relaxor; c. measuring the level of oxyhemoglobin and the level of deoxyhemoglobin in the cutaneous blood supply of said mammal during or after said exposing step; and d. for step a and step c, calculating and comparing at least one ratio selected from the group consisting of: i. the ratio of the level of deoxyhemoglobin to the level of oxyhemoglobin of said mammal; and ii. the ratio of the level of deoxyhemoglobin to the sum of the level of deoxyhemoglobin and the level of oxyhemoglobin of said mammal. 3. The method of claim 1 or 2, wherein said level of deoxyhemoglobin and said level of oxyhemoglobin are measured by at least one noninvasive technique. 4. The method of claim 3, wherein said technique is spectroscopy. 5. The method of claim 4, wherein said spectroscopy is at least one technique selected from the group consisting of diffuse reflectance spectroscopy, near infrared spectroscopy and ultraviolet-visible spectroscopy. 6. A method of changing the stress or relaxation level of a mammal, comprising the steps of: a. measuring the level of oxyhemoglobin and the level of deoxyhemoglobin in the cutaneous blood supply of said mammal prior to exposing said mammal to an acute stressor or relaxor; b. exposing said mammal to said acute stressor or relaxor; c. measuring the level of oxyhemoglobin and the level of deoxyhemoglobin in the cutaneous blood supply of said mammal during or after said exposing step; d. for step a and step c, calculating and comparing at least one ratio selected from the group consisting of: i. the ratio of the level of deoxyhemoglobin to the level of oxyhemoglobin of said mammal; and ii. the ratio of the level of deoxyhemoglobin to the sum of the level of deoxyhemoglobin and the level of oxyhemoglobin of said mammal; and e. changing at least one of said ratios bef... | Johnson & Johnson Consumer Companies | A61B005/00 A61M021/00
US07207939 | B2 | 2007-04-24 | Apparatus and method for analyzing a liquid in a capillary tube of a hematology instrument | 2002-10-03 | An apparatus and method for determining the density and fluid-type of a fluid flowing in a capillary tube, the velocity and viscosity of a blood sample flowing in a capillary tube, the erythrocyte sedimentation rate (ESR) of a blood sample after flow has been brought to an abrupt stop in a capillary tube, and/or the zeta sedimentation rate (ZSR) of a blood sample after flow has been brought to an abrupt stop in a capillary tube. These measurements are accomplished by directing a waveform pulse, such as an ultrasound pulse, at a pre-determined frequency transversely across the capillary tube and sample fluid, and by determining the flight of time of the pulse through the capillary tube and sample fluid and/or the Doppler shift of the echo signals reflecting off cells moving forwardly or transversely within a flowing, or stationary, blood sample. | 1. A device for analyzing a blood sample, comprising: a capillary tube defining a path of travel for flow of a blood sample in vitro; anda sensor assembly located about said capillary tube and having a launch sensor for emitting an ultrasound pulse into said capillary tube in a direction perpendicular to said path of travel, anda second sensor located forward along said path of travel relative to said launch sensor and canted at an angle relative to said path of travel so that said second sensor is aligned to receive echo signals of said ultrasound pulse that reflect off cells moving forward within said path of travel, wherein said device is adapted for measuring erythrocyte sedimentation rate (ESR) or zeta sedimentation rate (ZSR) of a blood sample. | 1. A device for analyzing a blood sample, comprising: a capillary tube defining a path of travel for flow of a blood sample in vitro; anda sensor assembly located about said capillary tube and having a launch sensor for emitting an ultrasound pulse into said capillary tube in a direction perpendicular to said path of travel, anda second sensor located forward along said path of travel relative to said launch sensor and canted at an angle relative to said path of travel so that said second sensor is aligned to receive echo signals of said ultrasound pulse that reflect off cells moving forward within said path of travel, wherein said device is adapted for measuring erythrocyte sedimentation rate (ESR) or zeta sedimentation rate (ZSR) of a blood sample. 2. A device according to claim 1, wherein said launch sensor is a piezo crystal transducer. 3. A device according to claim 1, wherein said capillary tube defines a straight path of travel adjacent said sensor assembly. 4. The device according to claim 1, wherein said launch sensor emits said ultrasound pulse in a direction perpendicular to said path of travel and along a diameter of said capillary tube. 5. The device according to claim 4, wherein a single launch sensor is utilized to emit said ultrasound pulse and receive an echo signal of said ultrasound pulse that propagates into said capillary tube in a direction away from said launch sensor and perpendicular to said path of travel and that is reflected back toward said launch sensor. 6. The device according to claim 1, further comprising: said sensor assembly having at least one sensor for emitting a waveform pulse at a predetermined frequency into said capillary tube perpendicularly across said path of travel and for receiving at least one of said pulse and an echo of said pulse after said pulse propagates at least partially through said capillary tube and said path of travel; anda signal processing circuit in communication with said sensor assembly. 7. The device according to claim 6, wherein, when air alone is located in said capillary tube, said signal processing circuit is adapted to determine wall thickness and inner diameter of said capillary tube based on echo signals received by said sensor assembly. 8. The device according to claim 1, further comprising a signal processing circuit interconnected to said second sensor and adapted to measure Doppler shift of echo signals received by said second sensor whereby cell velocity of the cells moving forward along said path of travel can be determined. 9. The device according to claim 8, wherein said signal processing circuit is adapted to capture Doppler shift measurements from echo signals received by said second sensor reflecting from cells located at two or more different spatial locations within said capillary tube so that a curve of velocity of cells across a cross section of said capillary tube can be generated. 10. The device according to claim 9, wherein said two or more spatial... | Coulter International Corp. | A61B005/00 A61B008/00
US07020509 | B2 | 2006-03-28 | Magnetic resonance tomography apparatus and method employing a true FISP sequence with improved off-resonant behavior of two spin ensembles | 2001-08-14 | In a magnetic resonance tomography apparatus employing a FISP pulse sequence, the pulse sequence is repeated with a repetition time TR with different phase-coding gradient directions and with an alternating operational sign of the flip angle α. The gradient pulse trains are thereby completely balanced. A phase increment ΔΦ=β is generated in addition to the alternating operational sign of the flip angle α between successive excitation pulses, so that the steady state signals for a first and a second spin ensemble optionally have either identical or reversed signal polarities. A first dataset on the basis of identical signal polarities and a second dataset on the basis of reversed signal polarities are obtained by means of the free selection of the mutual signal polarities. A pure image of the first and the second spin ensembles is thus obtained by the addition and/or subtraction of the first and second datasets. | 1. A magnetic resonance tomography apparatus comprising: a basic magnetic field system which generates a static basic magnetic field having a volume adapted to receive a subject containing different nuclei which, when excited to produce precessing nuclear spins, respectively represent a first spin ensemble and a second spin ensemble having respective steady state signals associated therewith;a gradient coil system which emits gradient pulse trains respectively for a slice selection gradient, a phase coding gradient having a phase coding direction, and a readout gradient;a radio-frequency system which emits RF excitation pulses to excite said nuclear spins and which, under said readout gradient, receives said respective steady state signals; anda sequence controller which operates said radio-frequency system and said gradient coil system to generate a FISP pulse sequence in said volume, said FISP pulse sequence including a plurality of RF excitation pulses separated by respective repetition times, with each repetition time being incremented relative to an immediately preceding repetition time by a phase angle increment, and with respective flip angles ≦90° that alternate in sign from excitation pulse-to-excitation pulse, so that said steady state signals have respective signal polarities with a relationship selected from the group consisting of identical signal polarities and opposite signal polarities, and wherein said FISP pulse sequence further includes respective phase coding gradients that alternate as to said phase coding direction from repetition time-to-repetition time, and completely balanced gradient pulse trains within each repetition time. | 1. A magnetic resonance tomography apparatus comprising: a basic magnetic field system which generates a static basic magnetic field having a volume adapted to receive a subject containing different nuclei which, when excited to produce precessing nuclear spins, respectively represent a first spin ensemble and a second spin ensemble having respective steady state signals associated therewith;a gradient coil system which emits gradient pulse trains respectively for a slice selection gradient, a phase coding gradient having a phase coding direction, and a readout gradient;a radio-frequency system which emits RF excitation pulses to excite said nuclear spins and which, under said readout gradient, receives said respective steady state signals; anda sequence controller which operates said radio-frequency system and said gradient coil system to generate a FISP pulse sequence in said volume, said FISP pulse sequence including a plurality of RF excitation pulses separated by respective repetition times, with each repetition time being incremented relative to an immediately preceding repetition time by a phase angle increment, and with respective flip angles ≦90° that alternate in sign from excitation pulse-to-excitation pulse, so that said steady state signals have respective signal polarities with a relationship selected from the group consisting of identical signal polarities and opposite signal polarities, and wherein said FISP pulse sequence further includes respective phase coding gradients that alternate as to said phase coding direction from repetition time-to-repetition time, and completely balanced gradient pulse trains within each repetition time. 2. A magnetic resonance tomography apparatus as claimed in claim 1 wherein said first spin ensemble has a first precession angle associated therewith and wherein said second spin ensemble has a second precession angle associated therewith, and wherein said sequence controller selects said phase angle increment so that said first and second precession angles satisfy respective mathematical conditions for a selected one of said signal polarity relationships. 3. A magnetic resonance tomography apparatus as claimed in claim 1 wherein each of said steady state signals is referenced relative to a zero-axis, and wherein each steady state signal exhibits zero-axis crossings, and wherein said sequence controller selects said phase angle increment to maximize a difference angle between closest respective zero-axis crossings of the respective steady state signals. 4. A magnetic resonance tomography apparatus as claimed in claim 1 wherein said sequence controller operates said radio-frequency system to excite water nuclei as said first spin ensemble and fat nuclei as said second spin ensemble. 5. A magnetic resonance tomography apparatus as claimed in claim 1 further comprising an image computer connected to said radio-frequency system for obtaining and storing a first image dataset when said polarities are ... | Siemens Aktiengesellschaft | A61B005/00
US07029441 | B2 | 2006-04-18 | Animal healthcare, well-being and nutrition | 2003-02-03 | Dynamic management of the health care, nutrition and well-being of pet animal, such as a dogs and cats. Data relating to the species of the animal, a selected group of the species, and the animal, including diagnostic laboratory test data of the animal are related by a computer. A management regimen for the animal is based on this. Updating the data picks up drifts in different populations of the animals, groups and species over time to enhance the management. | 1. A method of analyzing the health profile of a non-human animal comprising: a) providing a phenotypic and a genotypic database to the species of the non-human animal subject or a selected group of the species;obtaining animal data;c) correlating the databases of a) with the data of b) to determine a relationship between the database of a) and the data of b); andd) determining the health profile of the animal based on the correlating step. | 1. A method of analyzing the health profile of a non-human animal comprising: a) providing a phenotypic and a genotypic database to the species of the non-human animal subject or a selected group of the species;obtaining animal data;c) correlating the databases of a) with the data of b) to determine a relationship between the database of a) and the data of b); andd) determining the health profile of the animal based on the correlating step. 2. The method of claim 1, wherein the data to the animal are data obtained from analyzing bodily fluid or tissue. 3. The method of claim 2, wherein the data are metabolic components. 4. The method of claim 3, wherein the metabolic components are amino acids, carbohydrates, or lipids. 5. The method of claim 1 including the step of reporting the determination on a communications network including the Internet. 6. The method of claim 1 including the step of reporting the determination on a communications network including the Internet, and obtaining payment for the report through the Internet. 7. A method of determining a nutritional regime for a non-human animal comprising the steps of: a) providing a phenotypic and a genotypic database to the species of the non-human animal subject or a selected group of the species;obtaining animal data;c) correlating the databases of a) with the data of b) to determine a relationship between the database of a) and the data of b); andd) determining the nutritional regimen for the animal based on the correlating step. 8. The method of claim 7, wherein the data to the animal are data obtained from analyzing bodily fluid or tissue. 9. The method of claim 8, wherein the data are metabolic components. 10. The method of claim 9, wherein the metabolic components are amino acids, carbohydrates, or lipids. 11. A method of determining a nutritional regime for a non-human animal comprising the steps of: a) providing a phenotypic, a genotypic database, and a laboratory database to the species of the non-human animal subject or a selected group of the species;b) obtaining animal datac) correlating the databases of a) with the data of b) to determine a relationship between the database of a) and the data of b); andd) determining the nutritional regimen for the animal based on the correlating step. 12. The method of claim 11, wherein the laboratory database is metabolic components. 13. The method of claim 12, wherein the metabolic components are amino acids, carbohydrates, or lipids. | Hemopet | A61B005/00
US07218968 | B2 | 2007-05-15 | User interface for programming rate response technical field | 2002-10-31 | A programmer including a user interface, and techniques for programming a rate responsive implantable medical device using such a programmer and user interface are presented. The programmer may retrieve currently programmed rate response parameters and optimization target values of an implantable medical device. The programmer may display the currently programmed rate response parameters and optimization target values via a user interface. The programmer may also generate a current rate response curve and a current target rate histogram, and display the current rate response curve and current target rate histogram via the user interface. The programmer may receive changes to the displayed currently programmed rate response parameters and/or target values made by a user via the user interface, generate a pending rate response curve and/or a pending target rate histogram based on these changes, and display the pending rate response curve and/or pending target rate histogram via the user interface. | 1. A method of programming a rate response field for an implantable medical device without subjecting a patient to an exercise test or the like, comprising: receiving currently programmed rate response parameters and at least one currently programmed optimization target value from an implantable medical device;displaying the currently programmed parameters and at least one target value via a user interface;generating a current rate response curve which represents sensor indicated rate as a function of activity level based on the currently programmed parameters, and a current target rate histogram which represents an estimate of percentage of time as a function of sensor indicated rate based solely upon the currently programmed parameters and at least one target value without subjecting a patient receiving therapy from the implantable medical device to either an exercise test or a predetermined regimen of exertion and without first reprogramming the implantable medical device; anddisplaying both the current rate response curve and the current target rate histogram via the user interface. | 1. A method of programming a rate response field for an implantable medical device without subjecting a patient to an exercise test or the like, comprising: receiving currently programmed rate response parameters and at least one currently programmed optimization target value from an implantable medical device;displaying the currently programmed parameters and at least one target value via a user interface;generating a current rate response curve which represents sensor indicated rate as a function of activity level based on the currently programmed parameters, and a current target rate histogram which represents an estimate of percentage of time as a function of sensor indicated rate based solely upon the currently programmed parameters and at least one target value without subjecting a patient receiving therapy from the implantable medical device to either an exercise test or a predetermined regimen of exertion and without first reprogramming the implantable medical device; anddisplaying both the current rate response curve and the current target rate histogram via the user interface. 2. The method of claim 1, wherein receiving currently programmed rate response parameters comprises receiving currently programmed rate limits and setpoint values, and displaying at least some of the currently programmed parameters comprises: displaying the currently programmed rate limits;determining current curve values based on the currently programmed setpoint values; anddisplaying the current curve values. 3. The method of claim 2, wherein generating a current rate response curve comprises generating a current rate response curve based on the currently programmed rate limits and setpoint values. 4. The method of claim 2, wherein determining the current curve values comprises identifying the current curve values in a look-up table based on the currently programmed setpoint values. 5. The method of claim 1, wherein receiving at least one currently programmed optimization target value comprises receiving at least one target percentage of time for a sensor indicated rate to be within a range of rates. 6. The method of claim 5, wherein generating a current target rate histogram comprises estimating the percentage of time that the implantable medical device causes the sensor indicated rate to be within a number of rate range bins based on the at least one currently programmed optimization target value. 7. The method of claim 1, further comprising: receiving a change made by a user to at least one of the currently programmed parameters via a user interface;generating a pending rate response curve and a pending target rate histogram based on the change; anddisplaying the pending rate response curve and the pending target rate histogram via the user interface. 8. The method of claim 7, wherein displaying the pending rate response curve and the pending target rate histogram comprises: displaying the current rate response curve and the pending rate response curve... | Medtronic, Inc. | A61N001/05 A61B005/04
US07277745 | B2 | 2007-10-02 | Implantable myocardial ischemia detection, indication and action technology | 2002-09-15 | One embodiment enables detection of MI/I and emerging infarction in an implantable system. A plurality of devices may be used to gather and interpret data from within the heart, from the heart surface, and/or from the thoracic cavity. The apparatus may further alert the patient and/or communicate the condition to an external device or medical caregiver. Additionally, the implanted apparatus may initiate therapy of MI/I and emerging infarction. | 1. A method for monitoring electrogram signals in a subject, comprising: acquiring electrogram signals within the subject by positioning a plurality of sensors transveneously within the subject, the electrogram signals including a depolariztion wave region and a repolarization wave region, the sensors electrically coupled to an implanted device including electronics therein;analyzing the depolarization wave region of the electrogram signals using the implanted device to obtain information relating to changes in the depolarization wave region; anddetermining whether MI/I has occurred, using the information relating to changes in the depolarization wave region. | 1. A method for monitoring electrogram signals in a subject, comprising: acquiring electrogram signals within the subject by positioning a plurality of sensors transveneously within the subject, the electrogram signals including a depolariztion wave region and a repolarization wave region, the sensors electrically coupled to an implanted device including electronics therein;analyzing the depolarization wave region of the electrogram signals using the implanted device to obtain information relating to changes in the depolarization wave region; anddetermining whether MI/I has occurred, using the information relating to changes in the depolarization wave region. 2. A method as in claim 1, further comprising analyzing the repolarization wave region of the electrogram signals using the implanted device to obtain additional information relating to changes in the repolarization wave region, and wherein the determining whether MI/I has occurred also includes using the additional information related to changes in the repolarization wave region. 3. A method for monitoring electrogram signals in a subject, comprising: acquiring electrogram signals from the subject's heart, the electrogram signals including a depolariztion wave region and a repolarization wave region;analyzing the depolarization wave region of the electrogram signals using electronics positioned within the subject, wherein the analyzing includes obtaining depolarization information relating to changes in the depolarization wave region due to MI/I;analyzing the repolarization wave region of the electrogram signals using the electronics positioned within the subject, wherein the analyzing includes obtaining repolarization information relating to changes in the repolarization wave region due to MI/I; anddetermining, in the subject, whether MI/I has occurred, using the depolarization information and the repolarization information. 4. A method of detecting myocardial ischemia in a subject, comprising analyzing an electrogram signal by determining its temporal features synchronous to the Q, R, S, and T waves in a surface ECG signal, wherein determining its temporal features synchronous to the Q, R, S, and T waves includes providing a plurality of sensors adapted to deliver signals to a device implanted in the subject, storing an electrogram signal from each lead into a digital memory in the device, and comparing an acquired electrogram signal to the stored electrogram signal. 5. A method as in claim 4, wherein the analyzing includes: identifying changes in the morphology of depolarization upstroke synchronous with the Q and R waves;identifying changes in the morphology of depolarization downstroke synchronous with the R and S waves; andidentifying changes in the morphology of repolarization synchronous with the S and T waves. 6. A method as in claim 5, further comprising obtaining the electrogram signal from at least one lead positioned in a location selected from the group consisting of endo... | Infinite Biomedical Technologies, LLC | A61B005/04
US07155268 | B2 | 2006-12-26 | Non-invasive measurement of blood glucose level | 2002-10-31 | A non-invasive device for measuring blood glucose levels comprises an enclosure with a cover and dual openings, wherein the enclosure is covered with a transparent material (e.g., glass or plastic); at least two test cards arranged at specified distances and positioned within the enclosure; an illuminator and a conical light reflecting surface for casting a beam on the test cards. Light from one test card is reflected through one of the dual openings to sensing receptors of one eye. In order to enable the part of a diffuse beam reflected from the second test card made parallel within 10° to fall on the sensing receptors of the second eye, the enclosure is provided with a tubular prism covered with a closing plate having an opening at a first end while having an opening covered preferably with transparent material at a second end. | 1. A device for measuring blood glucose level comprising: a box having a cover that can be opened and closed and a side with two openings each covered with a transparent material further including at least two test cards arranged at specified distances and specified positions within the box, and at least one illuminator in the box configured to generate a diffuse beam and positioned therein to illuminate the two test cards wherein light reflected from one of the test cards is directed to a first of said openings and a diffuse beam reflected from the other test card is directed through a tubular prism configured to make diffuse light parallel within 10° that is covered with a closing plate with a first opening at one end while having a second opening aligned with a second of said openings at the other end. | 1. A device for measuring blood glucose level comprising: a box having a cover that can be opened and closed and a side with two openings each covered with a transparent material further including at least two test cards arranged at specified distances and specified positions within the box, and at least one illuminator in the box configured to generate a diffuse beam and positioned therein to illuminate the two test cards wherein light reflected from one of the test cards is directed to a first of said openings and a diffuse beam reflected from the other test card is directed through a tubular prism configured to make diffuse light parallel within 10° that is covered with a closing plate with a first opening at one end while having a second opening aligned with a second of said openings at the other end. 2. Blood glucose level measuring device as in claim 1, characterized by that a plurality of pairs of test cards are fastened on sides of a regular polygonal prism that can be rotated within the blood glucose level measuring device by a turn knob; and further a positioning spring holds the regular polygonal prism in position and a corresponding blood glucose level value is indicated on a dial of the turn-knob. 3. Blood glucose level measuring device as in claim 2, characterized by that only a single pair of test cards is viewable through the openings in the box and the other pairs of test cards are covered by two cover plates in order to eliminate their disturbing effects. 4. Blood glucose level measuring device as in claim 1, characterized by that one of said openings in the box is on a portion adjustable relative to the other opening. 5. Blood glucose level measuring device as in claim 1, characterized by that the illuminator is actuated by a micro switch mounted between two springs fastened to a base plate of the box. 6. Blood glucose level measuring device as in claim 1, characterized by that the diameter D1 of the first opening of the polygonal tubular prism is up to 2 mm, the diameter D2 of the second opening of the polygonal tubular prism and the diameter D4 of the first opening in the box are less than 10 mm each, and further the diameter D3 of the tubular prism is greater than or equal to the diameter D2. 7. Blood glucose level measuring device as in claim 1, characterized by that the distances between each opening in the box and the corresponding test card are equal and the length of the tubular prism is longer than 10mm. 8. Blood glucose level measuring device as in claim 1, characterized by that the illuminator is a glow lamp controlled by a push-button switch or a rotating switch and the illumination is active for a pre-determined time followed by a period where the illumination is disabled for a specified time; the readiness of the device to be switched on again is indicated by an indicator lamp; and further comprising a general switch to turn the device OFF and ON. 9. Blood glucose level measuring device as in claim 1, char... | Iris Medical Technologies International Ltd. | A61B005/00
US07077806 | B2 | 2006-07-18 | Adaptive interactive preceptored teaching system incorporating remote image monitoring | 2001-12-19 | A method for providing a patient with an analysis result which the patient uses to determine a dosage of a drug for self-administration when the patient is unable to obtain the analysis result themself. The method involves gathering by the patient of information which affects the dosage, forwarding the information to an advisor who analyzes the information to obtain the analysis result, sending the analysis result to the patient, and determining by the patient of the dosage using the analysis result. The information can be gathered by the patient by using a digital camera to photograph an item of food, and the analysis result can be an amount of food. Preferably, the information and analysis result are exchanged via the Internet. | 1. A method or providing a patient with an analysis result which the patient can use to determine a dosage of a drug for self-administration comprising: (a) gathering information which affects the dosage comprising information related to an item of food that may be eaten by a patient, the gathering being performed by the patient; (b) automatically forwarding the information to an advisor, computer system, or combination thereof, which analyzes the information to obtain an analysis result; and (c) sending the analysis result to the patient, whereby the patient can determine the dosage using the analysis result and wherein the analysis result optionally comprises a portion for the item of food. | 1. A method or providing a patient with an analysis result which the patient can use to determine a dosage of a drug for self-administration comprising: (a) gathering information which affects the dosage comprising information related to an item of food that may be eaten by a patient, the gathering being performed by the patient; (b) automatically forwarding the information to an advisor, computer system, or combination thereof, which analyzes the information to obtain an analysis result; and (c) sending the analysis result to the patient, whereby the patient can determine the dosage using the analysis result and wherein the analysis result optionally comprises a portion for the item of food. 2. The method of claim 1, wherein the step of gathering information comprises obtaining a photograph of the item of food. 3. The method of claim 2, wherein the photograph is a digital image. 4. The method of claim 2 wherein the photograph shows at least a part of a meal to be eaten by the patient. 5. The method of claim 2, wherein the step of gathering information further comprises obtaining the patient's estimate of the size of the item of food. 6. The method of claim 2, wherein the step of gathering information is performed using a digital camera. 7. The method of claim 6, wherein the digital camera is a part of a device used by the patient for self-monitoring one or more physical conditions relevant to the patient's health. 8. The method of claim 7, wherein the device is a blood glucose monitor. 9. The method of claim 1, wherein the analysis result comprises a portion size for the item. 10. The method of claim 9, wherein the analysis result indicates that the patient should not eat the item of food. 11. The method of claim 1 wherein the patient suffers from diabetes and the dosage reflects an amount of insulin. 12. The method of claim 1, further comprising the step of storing the analysis result for use in evaluation of the patient, education of the patient, or both. 13. The method of claim 1, wherein at least one of the forwarding of the information and the sending of the analysis result is performed using the Internet. 14. A method for providing a patient with an analysis result which the patient can use to determine a dosage of a drug for self-administration comprising performing steps (a)-(c) of claim 1 for more than one item of food in a meal. 15. A computer-readable storage medium comprising a program for performing the method of claim 1. 16. The method of claim 1, wherein the patient performs a self-analysis and the method comprises comparing the analysis result with the patient's self-analysis. 17. The method of claim 1, wherein the method comprises forwarding the information to an advisor that analyzes the information to obtain an analysis result. 18. The method of claim 1, wherein the analysis result indicates that the patient should take no dosage of the drug. 19. A system for providing a patient with an analysis result which... | Novo Nordisk A/S Mayo Medical Ventures | A61B005/00
US07244235 | B2 | 2007-07-17 | Split-night sleep diagnostic system | 2002-10-30 | A split-night sleep diagnostic system is provided for diagnosing and treating sleep apnea in a single night. The diagnostic system includes a blower, a respiratory interface and a conduit connected between the blower and the respiratory interface. Also included is a valve disposed along the conduit. The valve includes an aperture formed in the conduit between the blower and the respiratory interface, and a valve member movable between a first position partially occluding the conduit so that air flows through the aperture and a second position blocking the aperture so that continuous positive airway pressure is provided to a patient. Both diagnosis and titration can be performed using a single system at a single location over the course of a single night. | 1. A breathing assistance system, comprising: a blower that provides a gas flow and can provide a continuous positive airway pressure;a respiratory interface;a conduit connected between said blower and said respiratory interface; anda valve disposed along said conduit between a first conduit portion a second conduit portion, said valve including an aperture formed in said conduit between said blower and said respiratory interface, and a valve member that pivots between:(a) a first, fully open position wherein the valve member partially occludes said conduit and allows gas flow through said aperture, the fully open position of said valve member providing a gas flow path directing said gas flow provided by said blower to pass from said first conduit portion into said second conduit portion to purge expired respiratory gases located in said second conduit portion through said aperture, and(b) a second position blocking said aperture so that continuous positive airway pressure is provided to a patient. | 1. A breathing assistance system, comprising: a blower that provides a gas flow and can provide a continuous positive airway pressure;a respiratory interface;a conduit connected between said blower and said respiratory interface; anda valve disposed along said conduit between a first conduit portion a second conduit portion, said valve including an aperture formed in said conduit between said blower and said respiratory interface, and a valve member that pivots between:(a) a first, fully open position wherein the valve member partially occludes said conduit and allows gas flow through said aperture, the fully open position of said valve member providing a gas flow path directing said gas flow provided by said blower to pass from said first conduit portion into said second conduit portion to purge expired respiratory gases located in said second conduit portion through said aperture, and(b) a second position blocking said aperture so that continuous positive airway pressure is provided to a patient. 2. The system of claim 1, wherein said valve member is pivoted to said first position for diagnosing sleep apnea and said valve member is pivoted to said second position for treating sleep apnea. 3. The system of claim 1, further comprising a controller. 4. The system of claim 3, further comprising a signal line connected between said controller and said valve for conveying signals therebetween and allowing said controller to control the operation of said valve. 5. The system of claim 3, wherein said controller and said valve are positioned along the length of said conduit at spaced intervals, said valve located downstream of said controller. 6. The system of claim 3, wherein said controller and said valve are positioned between said blower and said respiratory interface at any location. 7. The system of claim 1, wherein said valve member is movable within said conduit between open and closed positions by means of an actuator. 8. The system of claim 1, wherein said valve member is movable by pneumatic actuation. 9. The system of claim 1, wherein said valve member is movable by solenoid actuation. 10. The system of claim 1, wherein said valve member is movable by motor actuation. 11. The system of claim 1, wherein said valve member is movable by electromagnetic actuation. 12. The system of claim 1, wherein said respiratory interface comprises a vent for purging expired respiratory gases. 13. A device for a breathing assistance system, said device comprising: a valve disposed along a conduit between a first conduit portion and a second conduit portion, said valve including:an aperture formed in said conduit between a blower and a respiratorya valve member that pivots between:(a) a first, fully open position partially occluding said conduit and allowing wherein the valve member partially occludes said conduit and allows gas flow through said aperture, the fully open position of said valve member providing a gas flow path directing gas to flow ... | Mallinckrodt, Inc. | A61B005/08 A61M016/00 A62B007/00 F16K031/02
US07089047 | B2 | 2006-08-08 | Fat depth sensor | 2002-06-13 | A non-invasive sensor suitable for measuring the depth of fat in animal tissue is described, the sensor generating and transmitting focussed discrete narrow band microwave signals into the tissue, receiving reflected microwave signals from the tissue and providing magnitude and phase information of the reflected microwave signals for a processor to compute a value of the fat depth within the tissue. | 1. A non-invasive sensor suitable for measuring the depth of fat in tissue, the sensor including: microwave signal generation and transmission means adapted to generate and transmit microwave signals into the tissue;focussing means adapted to focus the microwave signals onto a predetermined focal point;control means adapted to control the microwave signal generation and transmission means to transmit, in use, at least three discrete narrow band microwave signals into the tissue;receiving means adapted to receive reflected microwave signals from the tissue and provide magnitude and phase information of the reflected microwave signals; andprocessing means for receiving the magnitude and phase information provided by the receiving means and computing a value suitable for identification of the fat depth within the tissue. | 1. A non-invasive sensor suitable for measuring the depth of fat in tissue, the sensor including: microwave signal generation and transmission means adapted to generate and transmit microwave signals into the tissue;focussing means adapted to focus the microwave signals onto a predetermined focal point;control means adapted to control the microwave signal generation and transmission means to transmit, in use, at least three discrete narrow band microwave signals into the tissue;receiving means adapted to receive reflected microwave signals from the tissue and provide magnitude and phase information of the reflected microwave signals; andprocessing means for receiving the magnitude and phase information provided by the receiving means and computing a value suitable for identification of the fat depth within the tissue. 2. A non-invasive sensor as claimed in claim 1, wherein the microwave generation and transmission means generates and transmits microwave signals in the X-band. 3. A non-invasive sensor as claimed in claim 1, wherein the receiving means includes a six port reflectometer. 4. A non-invasive sensor as claimed in claim 1, wherein the focussing means includes a lens system. 5. A non-invasive sensor as claimed in claim 1, wherein the processing means is programmed to compute the reflection coefficient of received microwave signals. 6. A non-invasive sensor as claimed in claim 5, wherein the processing means is programmed to compute the phase of the received microwave signals from the reflection coefficient. 7. A non-invasive sensor as claimed in claim 1 wherein the processing means is programmed to compute the depth of fat by first computing the electrical length of the fat and computing the permittivity of the tissue at the focal point, and computing the depth of fat from the electrical length divided by the square root of the permittivity. 8. A non-invasive sensor as claimed in claim 7, wherein the processing means is programmed to compute the electrical length (de) of the fat and permittivity (εr) of the tissue by computing the values of electrical length and permittivity that minimise the error function∑i=1N⁢⁢[tan2⁢θi2-1ɛr⁢tan2⁡[2⁢π⁢⁢f1·dec]]=0or is programmed to perform an equivalent computation, wherein θi is the phase of the received microwave signals, f1 is the frequency of the received signals or frequency of the transmitted signal, c is speed of the microwave signals and N is the number of discrete narrow band signals transmitted into the tissue for the purposes of the measurement. 9. A non-invasive sensor as claimed in claim 8, wherein the processing means is programmed to compute the reflection coefficient of reflected microwave signals from the magnitude information from the receiving means and compute the phase of the received microwave signals from the reflection coefficient. 10. A non-invasive sensor as claimed in claim 8, wherein the processing means is programmed to use at least one iterative regression analysis... | Industrial Research Limited | A61B005/00
US07033325 | B1 | 2006-04-25 | Guidewire with multiple radiopaque marker sections | 2000-10-30 | An over-the-wire balloon catheter has a radiopaque marker thereon at a known distance from a distal end of the catheter. An elongated flexible guide wire has a plurality of longitudinally spaced radiopaque markers on a distal portion thereof, with adjacent markers on the guide wire being longitudinally spaced at a distance equal to the known distance between the distal end of the catheter and its radiopaque marker. The guide wire is advanced through an artery until one of its markers is positioned at a desired location relative to a stenosis. Subsequently, the balloon catheter is advanced over the guide wire until the radiopaque marker of the catheter is in a predetermined position relationship to the radiopaque marker of the guide wire which has aligned on the stenosis. The cooperating radiopaque markers of the guide wire and catheter thus allow the balloon member of the catheter to be positioned relative to the stenosis in a timely fashion without the need to constantly inject fluoroscopic dye to view the stenosis. | 1. An intravascular guidewire, comprising: an elongate shaft having a proximal portion and a distal portion, the proximal portion comprising a solid wire which tapers distally, the distal portion having distal end portion including a radiopaque coil tip disposed at the distal end portion; anda plurality of radiopaque markers disposed on the shaft proximal of the coil tip, the markers defining a plurality of 1.5 cm longitudinal spaces therebetween. | 1. An intravascular guidewire, comprising: an elongate shaft having a proximal portion and a distal portion, the proximal portion comprising a solid wire which tapers distally, the distal portion having distal end portion including a radiopaque coil tip disposed at the distal end portion; anda plurality of radiopaque markers disposed on the shaft proximal of the coil tip, the markers defining a plurality of 1.5 cm longitudinal spaces therebetween. 2. The intravascular guidewire of claim 1, wherein the radiopaque markers are 1 mm wide. | SciMed Life Systems, Inc. | A61B005/00 A61M025/00 A61B005/103
US07149571 | B2 | 2006-12-12 | Portable biological data measuring apparatus | 2002-11-15 | To provide a portable biological data measuring apparatus capable of measuring, storing and displaying biological data of a human being, in particular. The apparatus includes a GND electrode (12B) and a minus electrode (12C) which are provided on the opposite side faces of a roughly rectangular parallelopiped-shaped chassis (11), and a plus electrode (12A) provided on the face curved into an elliptical arc shape in cross section and adjacent to the opposite faces on which the GND electrode (12B) and the minus electrode (12C) are provided, in which biological data is measured using the plus electrode (12A), the GND electrode (12B) and the minus electrode (12C). | 1. A portable biological data measuring apparatus, comprising: a GND electrode (12B) and a minus electrode (12C) which are provided on the opposite side faces of a roughly rectangular parallelopiped-shaped chassis (11); a plus electrode (12A) provided on the face adjacent to the opposite faces on which said GND electrode (12B) and said minus electrode (12C) are provided, and display means (17) for displaying thereon the biological data, wherein an electrocardiographic waveform is displayed on said display means with the waveform being compressed in the direction of a time base, wherein said plus electrode (12A), said GND electrode (12B) and said minus electrode (12C) have irregularities in their surface and biological data is measured using said plus electrode (12A), said GND electrode (12B) and said minus electrode (12C) and said plus electrode is provided on an end face of said chassis that is curved to have an ellipse shape and is adjacent to the opposite faces on which the GND electrode and the minus electrode are disposed. | 1. A portable biological data measuring apparatus, comprising: a GND electrode (12B) and a minus electrode (12C) which are provided on the opposite side faces of a roughly rectangular parallelopiped-shaped chassis (11); a plus electrode (12A) provided on the face adjacent to the opposite faces on which said GND electrode (12B) and said minus electrode (12C) are provided, and display means (17) for displaying thereon the biological data, wherein an electrocardiographic waveform is displayed on said display means with the waveform being compressed in the direction of a time base, wherein said plus electrode (12A), said GND electrode (12B) and said minus electrode (12C) have irregularities in their surface and biological data is measured using said plus electrode (12A), said GND electrode (12B) and said minus electrode (12C) and said plus electrode is provided on an end face of said chassis that is curved to have an ellipse shape and is adjacent to the opposite faces on which the GND electrode and the minus electrode are disposed. 2. A portable biological data measuring apparatus according to claim 1, further comprising display means (17) for displaying thereon the biological data, wherein a first R wave and a second R wave contained in an electrocardiographic waveform are displayed on said display means, and wherein an axis (17a) for reading-out of the number of heartbeats on which data representing the number of heartbeats is plotted is displayed on the axis of abscissa indicating a time base with the peak portion of the first R wave being set as a reference point. 3. A portable biological data measuring apparatus according to claim 1, further comprising display means (17) for displaying thereon the biological data, wherein the data representing the time transition of the number of heartbeats between the adjacent R waves is displayed on said display means. 4. A portable biological data measuring apparatus, comprising: a GND electrode (12B) and a minus electrode (12C) which are provided on the opposite side faces of a roughly rectangular parallelopiped-shaped chassis (11); and a plus electrode (12A) provided on the face adjacent to the opposite faces on which said GND electrode (12B) and said minus electrode (12C) are provided, wherein said plus electrode (12A), said GND electrode (12B) and said minus electrode (12C) have irregularities in their surface and biological data is measured using said plus electrode (12A), said GND electrode (12B) and said minus electrode (12C) and said plus electrode is provided on an end face of said chassis that is curved to have an ellipse shape and is adjacent to the opposite faces on which the GND electrode and the minus electrode are diposed. | Parama Tech Co., Ltd. | A61B005/04
US07299192 | B2 | 2007-11-20 | Process, system, and computer executable program on a storage medium for recording patient treatment by progress toward identified goals | 2002-02-28 | A process for tracking skilled-service rendered a patient by a provider at a treatment location including: generating, at the treatment location, at least one record of progress made toward each treatment goal identified for the patient and addressed by the provider during a session; receiving an input comprising each record of progress so generated, for storage as treatment record data at a host location. Data is stored in searchable data fields for access and retrieval upon command. Using the treatment record data and a host processor, a patient report is automatically generated for submitting to an entity selected from the group consisting of a health care facility, a health care provider, an insurer, the patient, an entity authorized to receive records of the patient, and a record-keeping entity for the provider. Each treatment goal identified is preferably based upon an evaluation of the patient's condition. The evaluation may take place prior to, or during, the current treatment session, by the same or another provider. A system and computer executable program code are also included. | 1. A process for tracking skilled-service rendered to a patient at a physical therapy treatment location by a physical therapy provider trained in a physical therapy discipline, comprising the steps of: treating said patient based on at least one physical therapy treatment goal at said physical therapy treatment location by said provider;generating at least one record of incremental provider observed progress development, said record of incremental provider observed progress development comprising an indication of incremental patient advances made toward a composite physical therapy treatment goal as compared to a preferred result;evaluating development of said patient during at least one treatment modality based on said at least one physical therapy treatment goal during a treatment session;dynamically modifying said composite physical therapy treatment goal to create at least one edited treatment goal identified for the patient based upon a development evaluation of said patient during at least one treatment modality;receiving an input comprising at least one of said record of incremental provider observed progress development so generated, for storage as treatment record data at a host location;using said treatment record data and a host processor, automatically generating a patient report; andsubmitting said patient report to an entity selected from the group consisting of a health care facility, a health care provider, an insurer, the patient, an entity authorized to receive records of the patient, and a record-keeping entity for the provider. | 1. A process for tracking skilled-service rendered to a patient at a physical therapy treatment location by a physical therapy provider trained in a physical therapy discipline, comprising the steps of: treating said patient based on at least one physical therapy treatment goal at said physical therapy treatment location by said provider;generating at least one record of incremental provider observed progress development, said record of incremental provider observed progress development comprising an indication of incremental patient advances made toward a composite physical therapy treatment goal as compared to a preferred result;evaluating development of said patient during at least one treatment modality based on said at least one physical therapy treatment goal during a treatment session;dynamically modifying said composite physical therapy treatment goal to create at least one edited treatment goal identified for the patient based upon a development evaluation of said patient during at least one treatment modality;receiving an input comprising at least one of said record of incremental provider observed progress development so generated, for storage as treatment record data at a host location;using said treatment record data and a host processor, automatically generating a patient report; andsubmitting said patient report to an entity selected from the group consisting of a health care facility, a health care provider, an insurer, the patient, an entity authorized to receive records of the patient, and a record-keeping entity for the provider. 2. The process of claim 1 wherein: said at least one record of incremental provider observed progress development further comprises a description of said composite physical therapy treatment goal, a date of a treatment session, and the provider's identity; said step of receiving an input further comprises receiving said description and said provider's identity into respective data fields; and said step of submitting further comprises submitting said report for a payment from said entity for the skilled-service rendered. 3. The process of claim 1 wherein: said at least one record of incremental provider observed progress development further comprises a description of said composite physical therapy treatment goal and a date of a treatment session; said step of generating further comprises entering said incremental provider observed progress development made and said date into respective input fields of a treatment record display; said step of submitting further comprises transmitting said report as electronic data over a data transfer pathway to said entity for a payment for the skilled-service rendered. 4. The process of claim 1 wherein said physical therapy treatment goal was identified based upon an evaluation of the patient's condition, and said step of generating further comprises entering said at least one record of incremental provider observed progress development onto a hardcopy composite re... | - | A61B005/00
US07198602 | B2 | 2007-04-03 | Device, method and system for monitoring pressure in body cavities | 2002-10-30 | A method for analysing pressure-signals derivable from pressure measurements on or in a body of a human being or animal, comprising the steps of identifying from said digital data features related to single pressure waves, and determining at least one parameter of the single wave parameters elected from the group of: pressure amplitude ΔP, latency (ΔT), rise time coefficient ΔP/ΔT, and wavelength of the single wave, as well as determining numbers of single pressure waves with pre-selected combinations of two or more of said single pressure wave parameters during said given time sequence. In another aspect of the invention, the method is capable of identifying from said digital data features related to absolute pressures relative to atmospheric pressure a number of different pressure levels and duration thereof, and presenting the numbers of levels of various time durations in said matrix format. | 1. A method for analysing pressure-signals derivable from pressure measurements on or in a body of a human being or animal, comprising sampling said signals at specific intervals, converting the pressure signals into pressure-related digital data with a time reference, identifying from said digital data features related to single pressure waves in said pressure signals,said identifying step including determination of a minimum pressure value related to diastolic minimum value and a maximum pressure value related to systolic maximum value, anddetermining at least one parameter of the single wave parameters elected from the group of: pressure amplitude=ΔP=[(maximum pressure value)−(minimum pressure value)], latency (ΔT), rise time or rise time coefficient=ΔP/ΔT, and wavelength of the single wave, anddetermining numbers of said single pressure waves occurring during a given time sequence,wherein said determining of numbers includes:determining numbers of single pressure waves with pre-selected values of one or more of said single pressure wave parameters during said given time sequence, andfurther includes determining numbers of single pressure waves with pre-selected combinations of two or more of said single pressure wave parameters during said given time sequence. | 1. A method for analysing pressure-signals derivable from pressure measurements on or in a body of a human being or animal, comprising sampling said signals at specific intervals, converting the pressure signals into pressure-related digital data with a time reference, identifying from said digital data features related to single pressure waves in said pressure signals,said identifying step including determination of a minimum pressure value related to diastolic minimum value and a maximum pressure value related to systolic maximum value, anddetermining at least one parameter of the single wave parameters elected from the group of: pressure amplitude=ΔP=[(maximum pressure value)−(minimum pressure value)], latency (ΔT), rise time or rise time coefficient=ΔP/ΔT, and wavelength of the single wave, anddetermining numbers of said single pressure waves occurring during a given time sequence,wherein said determining of numbers includes:determining numbers of single pressure waves with pre-selected values of one or more of said single pressure wave parameters during said given time sequence, andfurther includes determining numbers of single pressure waves with pre-selected combinations of two or more of said single pressure wave parameters during said given time sequence. 2. A method according to claim 1, wherein the latency ΔT is defined as the time elapsed from occurrence of the minimum pressure value to the maximum pressure value. 3. The method according to claim 1, wherein the amplitude ΔP is defined as the pressure difference between the systolic maximum pressure and the diastolic minimum pressure during a series of increasing pressures of single pressure wave. 4. The method according to claim 1, wherein said maximum pressure value is one of three peak values occurring in said single pressure wave. 5. The method according to claim 4, wherein a first (P1) of said three peak values in said single pressure wave has an amplitude ΔP1 representing relative pressure difference between systolic maximum pressure and diastolic minimum pressure,a second (P2) of said three peak values has an amplitude ΔP2 related to a tidal wave portion of said single pressure wave, anda third (P3) of said three peak values has an amplitude ΔP3 related a dichrotic wave portion of said single pressure wave. 6. The method according to claim 3, further comprising calculation of one or more rise time coefficients between said amplitude and latency values. 7. The method according to claim 1, wherein in said single pressure wave said amplitude and latency values are relative values only, respectively. 8. The method according to claim 1, wherein said single pressure wave amplitudes and latency values are relative values not related to a zero pressure level. 9. The method according to claim 1, wherein numbers of single pressure waves during a said time sequence corresponds to the heart rate during said time sequence of said human being or animal. 10. The method according to cl... | Sensometrics AS | A61B005/02
US07065393 | B2 | 2006-06-20 | Apparatus, system and method of calibrating medical imaging systems | 2002-07-11 | A C-arm imaging device method, apparatus and system are provided. The apparatus and system includes a module for determining the positions of an imaging source and an imaging receptor and a local gravity vector to determine an imaging source focal point displacement. The apparatus and system may include a module for determining magnetic field distortion. Data obtained by the apparatus and system can be stored for use in subsequent interpolation of imaging source focal point displacement for the C-arm imaging device. | 1. A method of calibrating a C-arm imaging device comprising the steps of: determining a position of an imaging source for each of a plurality of orientations of the C-arm, including determining a reference position of the imaging source for a reference orientation;determining a position of an imaging receptor for each of the plurality of orientations, including determining a reference position of the imaging receptor for the reference orientation;determining a local gravity vector for each of the plurality of orientations;determining an imaging source focal point displacement due to bending for each of the plurality of orientations using the reference position of the imaging source, the reference position of the imaging receptor, the position of the imaging source and the position of the imaging receptor in correspondence to the local gravity vector; andstoring information on the position of the imaging source, the position of the imaging receptor, the gravity vector, and the imaging source focal point displacement for each of the plurality of orientations, for use in subsequent interpolation of imaging source focal point displacement based on a specific position of an imaging receptor and a specific local gravity vector. | 1. A method of calibrating a C-arm imaging device comprising the steps of: determining a position of an imaging source for each of a plurality of orientations of the C-arm, including determining a reference position of the imaging source for a reference orientation;determining a position of an imaging receptor for each of the plurality of orientations, including determining a reference position of the imaging receptor for the reference orientation;determining a local gravity vector for each of the plurality of orientations;determining an imaging source focal point displacement due to bending for each of the plurality of orientations using the reference position of the imaging source, the reference position of the imaging receptor, the position of the imaging source and the position of the imaging receptor in correspondence to the local gravity vector; andstoring information on the position of the imaging source, the position of the imaging receptor, the gravity vector, and the imaging source focal point displacement for each of the plurality of orientations, for use in subsequent interpolation of imaging source focal point displacement based on a specific position of an imaging receptor and a specific local gravity vector. 2. A method of claim 1 wherein the plurality of orientations is obtained by rotating the C-arm about a plurality of angles around a first axis and a plurality of angles around a second axis. 3. A method of claim 2 wherein the plurality of angles are spaced 10 degrees apart. 4. A method of claim 3 wherein the plurality of orientations comprise about 100 or more orientations. 5. A method of claim 1, wherein the position of the imaging source is determined by obtaining a coordinate position in three dimensional space. 6. A method of claim 1, wherein the position of the imaging receptor is determined by obtaining a coordinate position in three dimensional space. 7. A method of claim 1 wherein the local gravity vector is determined with reference to a moveable part of the C-arm. 8. A method of claim 7 wherein the local gravity vector is determined by obtaining the angular displacement of the moveable part from alignment with the local gravitational field. 9. A method of claim 7 wherein the moveable part of the C-arm is the imaging receptor. 10. A method of claim 1 wherein the imaging source focal point displacement is determined in conjunction with reference coordinates for the imaging source. 11. A method of claim 10 wherein the imaging source focal point displacement is derived from the change in position of the imaging source from the reference coordinates for the imaging source. 12. A method of claim 1 wherein the step of storing comprises saving information in a look-up table. 13. A method of claim 1 further comprising the steps of: determining a specific position of the imaging receptor;determining a specific gravity vector; andinterpolating imaging source focal point displacement using the stored information. 14.... | Cedara Software Corp. | A61B005/05
US07089046 | B2 | 2006-08-08 | Removable cap for tissue-insertable connections | 2002-07-02 | A helical element for insertion into tissue comprises a helical element having an insertion end, a protruding end and an open central area within the wire, rods, filaments, cables or the like that form the helix. The helical element has at least its insertion end covered by a cap of a water-soluble or water-dispersible composition. The cap is provided with a surface shape in a cross-section in which surface variations are present in the cross-section which create a surface orientation where a line from the center of the cross-section can intersect the surface, and a line perpendicular to said radius at a point of intersection with said surface forms four quadrants, three of said quadrants containing water-soluble or water-dispersible cap material. There is either a hollow area within the composition within the open central area or the material is more porous than the remaining material. The helical element preferably comprises an electrical lead, such as a positive endocardial lead, with an electrode at the protruding or distal end of the lead.;The helical element may comprise any biocompatable material with sufficient structural integrity to provide a secure attachment to tissue in a patient. Where the helical element is also to provide an active (electrically active) function, the composition of the helical element should also be electrically conductive. | 1. A process for the manufacture of a device for insertion into tissue comprising a helical element with an open central area and a water-soluble or water-dispersible cap for said helical element, the cap for said helical element having a hollow area which overlaps at least a part of said open central area, said process comprising forming said cap with a hollow core extending at least partially through said cap and pressing said helical element into said hollow core, and molding indentations in a surface of said cap which will reduce the time necessary to dissolve or disperse said cap in water. | 1. A process for the manufacture of a device for insertion into tissue comprising a helical element with an open central area and a water-soluble or water-dispersible cap for said helical element, the cap for said helical element having a hollow area which overlaps at least a part of said open central area, said process comprising forming said cap with a hollow core extending at least partially through said cap and pressing said helical element into said hollow core, and molding indentations in a surface of said cap which will reduce the time necessary to dissolve or disperse said cap in water. 2. The process of claim 1 wherein said pressing is accompanied by a rotation of said helical element. 3. The process of claim 2 wherein said cap comprises a water-soluble or water-dispersible material selected from the group consisting of mannitol, iditol, glucitol, rabitol, heptitol, octitol, arabinitol, betitol, bomesitiol, dambonitol, inositol, laminitol, ononitol, pinitol, sorbitol, and organic polymers. 4. The process of claim 2 wherein said cap comprises mannitol. 5. The process of claim 1 wherein said cap comprises a water-soluble or water-dispersible material selected from the group consisting of mannitol, iditol, glucitol, rabitol, heptitol, octitol, arabinitol, betitol, bomesitiol, dambonitol, inositol, laminitol, ononitol, pinitol, sorbitol, and organic polymers. 6. The process of claim 1 wherein said cap comprises mannitol. 7. A method of forming a cap for a helical element of an implantable lead, the method comprising: selecting a water-soluble or water-dispersible material; andforming a cap from the water-soluble or water-dispersible material, the cap formed to include a hollow area which overlaps at least a part of an open central area of the helical element of the implantable lead, the cap also formed to include indentations in a surface of said cap which will reduce the time necessary to dissolve or disperse said cap in water. 8. The method of claim 7, wherein the water-soluble or water-dispersible material includes a material selected from the group consisting of mannitol, iditol, glucitol, rabitol, heptitol, octitol, arabinitol, betitol, bomesitiol, dambonitol, inositol, laminitol, ononitol, pinitol, sorbitol, and aqueous-soluble organic polymers. 9. The method of claim 7, wherein the water-soluble or water-dispersible material includes mannitol. 10. The method of claim 7, wherein forming the cap includes molding the cap. 11. The method of claim 7, wherein forming the cap includes dip-coating the cap and forming the indentations by sculpting. 12. A method of forming a cap for a helical element of an implantable lead, the method comprising: selecting a water-soluble or water-dispersible mannitol material; andforming a cap from the water-soluble or water-dispersible mannitol material, including forming grooves along a surface of the cap which decrease the time for dissolution of the cap in an aqueous solution as compared to a ca... | Cardiac Pacemakers, Inc. | A61B005/04
US07164940 | B2 | 2007-01-16 | Therapeutic system | 2001-06-28 | A therapeutic system comprising a therapeutic applicator for treating living tissues and an observation unit for determining the position of the applicator. The therapeutic energy applied from the applicator to the living tissues is controlled in accordance with data representing the position of the applicator, determined by the observation unit. It is therefore easy to operate the system even if the applicator and the observation unit (e.g., an MRI apparatus) are used at the same time. The image of the living tissues remains clear during the use of the therapeutic applicator. | 1. A therapeutic system comprising: an energy source which generates energy to treat an internal affected region; an applicator having a distal end portion and a proximal end portion and including a longitudinal member, wherein at least the distal end portion of the applicator is inserted in a living body and arranged at a position such that the affected region is enabled to be treated by the energy, and the energy is supplied to the affected region through the distal end portion of the applicator; a data acquisition device to acquire tomogram data of the living body including the affected region; a display to display a tomogram of living tissues including the affected region based on the tomogram data; a setting section to set at least one of a plurality of reference points and at least one measuring line on the tomogram displayed on the display before the energy is supplied to the affected region; and an observation unit including the display and the setting section, wherein the observation unit is adapted to observe a condition of the affected region treated by the energy based on a change of the tomogram data corresponding to the at least one of the plurality of reference points and the at least one measuring line which varies as the energy is supplied to the affected region; wherein the setting section sets the at least one of the plurality of reference points and the at least one measuring line based on the tomogram data acquired by the data acquisition device. | 1. A therapeutic system comprising: an energy source which generates energy to treat an internal affected region; an applicator having a distal end portion and a proximal end portion and including a longitudinal member, wherein at least the distal end portion of the applicator is inserted in a living body and arranged at a position such that the affected region is enabled to be treated by the energy, and the energy is supplied to the affected region through the distal end portion of the applicator; a data acquisition device to acquire tomogram data of the living body including the affected region; a display to display a tomogram of living tissues including the affected region based on the tomogram data; a setting section to set at least one of a plurality of reference points and at least one measuring line on the tomogram displayed on the display before the energy is supplied to the affected region; and an observation unit including the display and the setting section, wherein the observation unit is adapted to observe a condition of the affected region treated by the energy based on a change of the tomogram data corresponding to the at least one of the plurality of reference points and the at least one measuring line which varies as the energy is supplied to the affected region; wherein the setting section sets the at least one of the plurality of reference points and the at least one measuring line based on the tomogram data acquired by the data acquisition device. 2. A therapeutic system according to claim 1, wherein the observation unit one of (i) inhibits the energy generating means from generating the energy and (ii) decreases the energy generated by the energy generating means. 3. A therapeutic system according to claim 1, wherein the observation unit has manual setting means for setting the at least one of the plurality of reference points and the at least one measuring line when manually operated. 4. A therapeutic system according to claim 1, wherein the tomogram data is acquired by supplying, to the affected region, energy that is generated on a principle different from that of the energy source. 5. A therapeutic system comprising: an energy source which generates energy to treat an internal affected region; an applicator having a distal end portion and a proximal end portion and including a longitudinal member, wherein at least the distal end portion of the applicator is inserter in a living body and arranged at a position such that the affected region is enabled to be treated by the energy, and energy is supplied to the affected region through the distal end portion of the applicator; a data acquisition device to acquire tomogram data of the living body including the affected region; a display to display a tomogram of living tissues including the affected region based on the tomogram data; a setting section to set at least one of a plurality of reference points and at least one measuring line on the tomogram displayed on the disp... | Olympus Optical Co., Ltd. | A61B005/05
US07136696 | B2 | 2006-11-14 | Neuron signal analysis system and method | 2002-05-03 | Electrical signal data from a target location are recorded and collected. A signal analyzer analyzes the signal data and determines whether characteristics of the signal data resemble noise data or whether a departure from noise is present. If the characteristics show a departure from noise, neuronal activity at the target location is indicated. | 1. A system for positioning a microelectrode within a tissue, the system comprising: a microelectrode for collecting electrical signals;a computer readable medium having: distribution logic that causes a computer to generate a statistical distribution of the electrical signals; andidentification logic that that causes a computer to identify whether the statistical distribution resembles a noise distribution without neuronal discharge; anda position controller for controlling the position of microelectrode, the position controller moving or not moving the microelectrode based on the identification of whether the statistical distribution resembles a noise distribution without neuronal discharge. | 1. A system for positioning a microelectrode within a tissue, the system comprising: a microelectrode for collecting electrical signals;a computer readable medium having: distribution logic that causes a computer to generate a statistical distribution of the electrical signals; andidentification logic that that causes a computer to identify whether the statistical distribution resembles a noise distribution without neuronal discharge; anda position controller for controlling the position of microelectrode, the position controller moving or not moving the microelectrode based on the identification of whether the statistical distribution resembles a noise distribution without neuronal discharge. 2. The system as set forth in claim 1 wherein the statistical distribution is based on voltage amplitude values of the electrical signals. 3. The system as set forth in claim 1 wherein the noise distribution without neuronal discharge is a bell-shaped curve. 4. The system as set forth in claim 1 wherein the identification logic (i) includes logic for determining when the statistical distribution differs from the noise distribution without neuronal discharge, and (ii) generates a signal indicating possible neuron cell discharges are present in the electrical signals. 5. The system as set forth in claim 4 wherein the position controller moves the microelectrode within the brain until receipt of the signal indicating possible neuron cell discharge. 6. The system as set forth in claim 4 further including count logic for analyzing tails of the statistical distribution and indicating a quantity of neuron cell discharges in the electrical signals. 7. The system as set forth in claim 1 wherein the identification logic includes means for analyzing the statistical distribution. 8. The system as set forth in claim 1 wherein the identification logic includes logic for determining a kurtosis value of the statistical distribution. 9. The system as set forth in claim 1 further including correlation logic that correlates the statistical distribution with an applied stimulus. 10. A software product for analyzing electrical signal data collected from within tissue, the software product comprising: a computer readable medium including:logic that causes a computer to obtain electrical signal data from an microelectrode configured to be positioned within tissue;first computer executable instructions that cause a computer to generate a statistical distribution of the electrical signal data where the statistical distribution has a first shape which represents noise signal data without neuronal activity or a second shape that represents noise data with neuronal activity; andsecond computer executable instructions that cause a computer to analyze the statistical distribution and indicate whether neuronal activity is present in the electrical signal data based on the shape of the statistical distribution. 11. The software product as set forth in claim 10 wherein the comput... | The Cleveland Clinic Foundation | A61B005/04
US06986349 | B2 | 2006-01-17 | Systems and methods for enhancing blood circulation | 2002-08-19 | A method for increasing circulation in a breathing person utilizes a valve system that is interfaced to the person's airway and is configured to decrease or prevent respiratory gas flow to the person's lungs during at least a portion of an inhalation event. The person is permitted to inhale and exhale through the valve system. During inhalation, the valve system functions to produce a vacuum within the thorax to increase blood flow back to the right heart of the person, thereby increasing cardiac output and blood circulation. | 1. A method for increasing circulation in a breathing person, the method comprising: interfacing a valve system to the person's airway, the valve system being configured to decrease or prevent respiratory gas flow to the person's lungs during at least a portion of an inhalation event, wherein the valve system is interfaced for a time period in the range from about 30 seconds to about 24 hours; permitting the person to inhale and exhale through the valve system, wherein during inhalation the valve system functions to produce a vacuum within the thorax to increase blood flow back to the right heart of the person, thereby increasing cardiac output and blood circulation. | 1. A method for increasing circulation in a breathing person, the method comprising: interfacing a valve system to the person's airway, the valve system being configured to decrease or prevent respiratory gas flow to the person's lungs during at least a portion of an inhalation event, wherein the valve system is interfaced for a time period in the range from about 30 seconds to about 24 hours; permitting the person to inhale and exhale through the valve system, wherein during inhalation the valve system functions to produce a vacuum within the thorax to increase blood flow back to the right heart of the person, thereby increasing cardiac output and blood circulation. 2. A method as in claim 1, wherein the valve system is incorporated into a facial mask, and further comprising coupling the facial mask to the person's face. 3. A method as in claim 1, wherein the valve system includes a pressure responsive inflow valve, and further comprising setting an actuating pressure of the valve to be in the range from about 0 cm H2O to about −40 cm H2O. 4. A method as in claim 1, further comprising varying the actuating pressure of the valve over time. 5. A method as in claim 1, wherein the valve system is further configured to prevent or decrease exhaled gases from exiting the person's lungs during at least a portion of an exhalation event to further increase blood circulation. 6. A method as in claim 1, further comprising adding a supply of oxygen through the valve system to supplement oxygen delivery to the person. 7. A method as in claim 1, wherein the person is suffering from a disease that is related to impaired venous blood flow. 8. A method as in claim 1, wherein the person is suffering from a disease resulting at least in part from poor blood circulation, and wherein the valve system functions to increase blood circulation to treat the disease. 9. A method as in claim 8, wherein the disease is selected from a group consisting of venous stasis ulcers, deep vein thrombosis, wound healing, and lymphedema. 10. A method as in claim 8, wherein the disease comprises renal failure. 11. A method for increasing circulation in a breathing person, the method comprising: interfacing a valve system to the person's airway, the valve system being configured to decrease or prevent respiratory gas flow to the person's lungs during at least a portion of an inhalation event; permitting the person to inhale and exhale through the valve system, wherein during inhalation the valve system functions to produce a vacuum within the thorax to increase blood flow back to the right heart of the person, thereby increasing cardiac output and blood circulation, and wherein the person is suffering from a disease that is related to impaired venous blood flow. 12. A method for increasing circulation in a breathing person, the method comprising: interfacing a valve system to the person's airway, the valve system being configured to decrease or prevent respiratory gas flow t... | Advanced Circulatory Systems, Inc. | A61M016/00 A61B005/02
US07062316 | B2 | 2006-06-13 | Cross chamber interval correlation | 2003-01-21 | A system and method for discriminating cardiac rhythms occurring in an antegrade direction from cardiac rhythms occurring in a retrograde direction. Atrial and ventricular contractions are sensed, from which atrial and ventricular cycle lengths are determined. Ventricular contractions are also analyzed to determine the occurrence of a tachycardia episode that has a one-to-one association of atrial contractions to ventricular contractions. During a tachycardia episode having a one-to-one association of atrial contractions to ventricular contractions, the atrial cycle lengths are paired with the ventricular cycle lengths, where for each of the atrial cycle lengths the atrial cycle length is paired with at least one ventricular cycle length started before the first atrial contraction of each of the atrial cycle lengths and paired with at least one ventricular cycle length started after the first atrial contraction of each of the atrial cycle lengths. A retrograde correlation coefficient is then determined for the atrial cycle lengths paired with the ventricular cycle lengths started before the first atrial contraction, and an antegrade correlation coefficient is determined for the atrial cycle lengths paired with the ventricular cycle lengths started after the first atrial contraction. The tachycardiac episode is then classified based on a comparison of the antegrade correlation coefficient and the retrograde correlation coefficient. | 1. A system for determining a direction of conduction of an arrhythmia episode between a first cardiac region and a second cardiac region, the system comprising: a first electrode to sense a first signal indicative of first contractions in the first cardiac region; a second electrode to sense a second signal indicative of second contractions in the second cardiac region; a first contraction detector circuit, coupled to the first electrode, to detect the first contractions; a second contraction detector circuit, coupled to the second electrode, to detect the second contractions; a cardiac data analyzing circuit coupled to the first contraction detector circuit and the second contraction detector circuit, the cardiac data analyzing circuit adapted to: determine first cycle lengths each being a time interval between two successively sensed first contractions; determine second cycle lengths each being a time interval between two successively sensed second contractions; and analyze the first and second contractions to determine an occurrence of an arrhythmia episode having a one-to-one association of the first contractions to the second contractions; a cycle length pairing circuit coupled to the cardiac data analyzing circuit, the cycle length pairing circuit adapted to pair the first cycle lengths with the second cycle lengths in first and second directions during the arrhythmia episode having the one-to-one association of the first contractions to the second contractions, the first direction corresponding to a direction of conduction of the arrhythmia episode from the first cardiac region to the second cardiac region, the second direction corresponding to a direction of conduction of the arrhythmia episode from the second cardiac region to the first cardiac region; and a processor coupled to the cycle length pairing circuit, the processor adapted to: determine correlation coefficients each for the first cycle lengths paired with the second cycle lengths in one of the first and second directions; and determine the direction of the conduction of the arrhythmia episode based on the correlation coefficients. | 1. A system for determining a direction of conduction of an arrhythmia episode between a first cardiac region and a second cardiac region, the system comprising: a first electrode to sense a first signal indicative of first contractions in the first cardiac region; a second electrode to sense a second signal indicative of second contractions in the second cardiac region; a first contraction detector circuit, coupled to the first electrode, to detect the first contractions; a second contraction detector circuit, coupled to the second electrode, to detect the second contractions; a cardiac data analyzing circuit coupled to the first contraction detector circuit and the second contraction detector circuit, the cardiac data analyzing circuit adapted to: determine first cycle lengths each being a time interval between two successively sensed first contractions; determine second cycle lengths each being a time interval between two successively sensed second contractions; and analyze the first and second contractions to determine an occurrence of an arrhythmia episode having a one-to-one association of the first contractions to the second contractions; a cycle length pairing circuit coupled to the cardiac data analyzing circuit, the cycle length pairing circuit adapted to pair the first cycle lengths with the second cycle lengths in first and second directions during the arrhythmia episode having the one-to-one association of the first contractions to the second contractions, the first direction corresponding to a direction of conduction of the arrhythmia episode from the first cardiac region to the second cardiac region, the second direction corresponding to a direction of conduction of the arrhythmia episode from the second cardiac region to the first cardiac region; and a processor coupled to the cycle length pairing circuit, the processor adapted to: determine correlation coefficients each for the first cycle lengths paired with the second cycle lengths in one of the first and second directions; and determine the direction of the conduction of the arrhythmia episode based on the correlation coefficients. 2. The system of claim 1, wherein the first electrode is an atrial electrode, and the second electrode is a ventricular electrode. 3. The system of claim 2, further comprising a pacing output circuit coupled to the atrial electrode. 4. The system of claim 2, further comprising a pacing output circuit coupled to the ventricular electrode. 5. The system of claim 1, wherein the cycle length pairing circuit is adapted to: pair the first cycle lengths with the second cycle lengths in the first direction, where for each of the paired first and second cycle lengths, the first cycle lengths starts before the second cycle length starts; and pair the first cycle lengths with the second cycle lengths in the second direction, where for each of the paired first and second cycle lengths, the first cycle lengths starts after the second cycle length starts. ... | Cardiac Pacemakers, Inc. | A61B005/0402
US07248910 | B2 | 2007-07-24 | Physiological parameter monitoring system and sensor assembly for same | 2002-10-22 | A detachable sensor assembly for use in monitoring a physiological parameter of a subject includes an information containing circuit and a timer. The sensor assembly stores time varying information in the information containing circuit. The information may include cumulative use times for the detachable sensor assembly and/or for components in the detachable sensor assembly. | 1. A system for determining a value of a physiological parameter, the system comprising a detachable sensor assembly for supplying a signal to a device for monitoring a physiological parameter of a subject, the detachable sensor assembly comprising: a sensor;an information containing circuit;a timer; and,a connector comprising one or more signal conductors connected to carry information from the sensor and information from the information containing circuit to a device for monitoring a physiological parameter;wherein the information containing circuit is configured to store time-varying information in response to timing signals from the timer and to transmit the time varying information by way of the connector;wherein the time-varying information comprises a cumulative use time for a component comprising a light-emitting element in the detachable sensor assemblythe sensor assembly connected by way of the connector to a monitoring device;wherein the monitoring device is configured to:retrieve a first instance of a time-varying value from the information containing circuit at a first time;store the first instance of the time-varying value in a memory;retrieve a second instance of the time-varying value from the information containing circuit at a second time later than the first time;compare the first and second instances of the time-varying value to a difference between the first and second times;energize the component at or before the first time;retrieve from the information containing circuit information identifying a date of manufacturing the detachable sensor assembly;compare the date of manufacture to a current date maintained in the monitoring device; and,signal an error if the date of manufacture of the sensor assembly is longer before the current date than a threshold time;retrieve from the information containing circuit information identifying a cumulative use time of the detachable sensor assembly;signal an error if the cumulative use time exceeds a threshold;retrieve from the information containing circuit information identifying a cumulative on time of the component in the detachable sensor assembly;signal an error if the cumulative on time exceeds a threshold;retrieve information from the information containing circuit;compare the retrieved information to information stored in the monitoring device which was previously retrieved from the information containing circuit; and,signal an error if the retrieved information is not consistent with the stored information;wherein the monitoring device comprises a data store holding stored information identifying one or more acceptable sensor combinations, each of the acceptable sensor combinations comprising a plurality of sensor locations required for the determination of a physiological parameter; and,a plurality of sensors including said sensor, each detachably connected to the monitoring device, each of the sensors intended for application to a different location on a subject's anatomy, eac... | 1. A system for determining a value of a physiological parameter, the system comprising a detachable sensor assembly for supplying a signal to a device for monitoring a physiological parameter of a subject, the detachable sensor assembly comprising: a sensor;an information containing circuit;a timer; and,a connector comprising one or more signal conductors connected to carry information from the sensor and information from the information containing circuit to a device for monitoring a physiological parameter;wherein the information containing circuit is configured to store time-varying information in response to timing signals from the timer and to transmit the time varying information by way of the connector;wherein the time-varying information comprises a cumulative use time for a component comprising a light-emitting element in the detachable sensor assemblythe sensor assembly connected by way of the connector to a monitoring device;wherein the monitoring device is configured to:retrieve a first instance of a time-varying value from the information containing circuit at a first time;store the first instance of the time-varying value in a memory;retrieve a second instance of the time-varying value from the information containing circuit at a second time later than the first time;compare the first and second instances of the time-varying value to a difference between the first and second times;energize the component at or before the first time;retrieve from the information containing circuit information identifying a date of manufacturing the detachable sensor assembly;compare the date of manufacture to a current date maintained in the monitoring device; and,signal an error if the date of manufacture of the sensor assembly is longer before the current date than a threshold time;retrieve from the information containing circuit information identifying a cumulative use time of the detachable sensor assembly;signal an error if the cumulative use time exceeds a threshold;retrieve from the information containing circuit information identifying a cumulative on time of the component in the detachable sensor assembly;signal an error if the cumulative on time exceeds a threshold;retrieve information from the information containing circuit;compare the retrieved information to information stored in the monitoring device which was previously retrieved from the information containing circuit; and,signal an error if the retrieved information is not consistent with the stored information;wherein the monitoring device comprises a data store holding stored information identifying one or more acceptable sensor combinations, each of the acceptable sensor combinations comprising a plurality of sensor locations required for the determination of a physiological parameter; and,a plurality of sensors including said sensor, each detachably connected to the monitoring device, each of the sensors intended for application to a different location on a subject's anatomy, eac... | CardioDigital Limited | A61B005/00 A61B005/02
US07245962 | B2 | 2007-07-17 | System and method for determining reentrant ventricular tachycardia isthmus location and shape for catheter ablation | 2002-07-30 | A method and system for identifying and localizing a reentrant circuit isthmus in a heart of a subject during sinus rhythm is provided. The method may include (a) receiving electrogram signals from the heart during sinus rhythm via electrodes, (b) creating a map based on the electrogram signals, (c) determining, based on the map, a location of the reentrant circuit isthmus in the heart, and (d) displaying the location of the reentrant circuit isthmus. | 1. A system for identifying and localizing a reentrant circuit isthmus in a heart of a subject during sinus rhythm, comprising: (a) an interface for receiving electrogram signals from the heart during sinus rhythm via electrodes; (b) processing means for creating a map based on the received electrogram signals, and determining, based on the map, a location of the reentrant circuit isthmus in the heart; and (c) a display adapted to display the location of the reentrant circuit isthmus, wherein activation times of the received electrogram signals are arranged based on a position of the respective electrodes. | 1. A system for identifying and localizing a reentrant circuit isthmus in a heart of a subject during sinus rhythm, comprising: (a) an interface for receiving electrogram signals from the heart during sinus rhythm via electrodes; (b) processing means for creating a map based on the received electrogram signals, and determining, based on the map, a location of the reentrant circuit isthmus in the heart; and (c) a display adapted to display the location of the reentrant circuit isthmus, wherein activation times of the received electrogram signals are arranged based on a position of the respective electrodes. 2. The system of claim 1, wherein the activation times are measured from a predetermined start time until reception of a predetermined electrogram signal. 3. The system of claim 1, wherein the map includes isochrones for identifying electrogram signals having activation times within a predetermined range. 4. The system of claim 1, wherein a center reference activation location on the map is determined by averaging an electrode coordinate position of a predetermined number of electrogram signals selected based on an activation time. 5. The system of claim 4, wherein measurement vectors originating from the center reference activation location and extending outward on the map are defined and are used to designate the electrodes located along the measurement vectors. 6. The system of claim 5, wherein the electrodes assigned to a measurement vector are chosen according to a distance from the measurement vector. 7. The system of claim 5, wherein the electrodes assigned to a measurement vector are a subset of the electrodes chosen according to a distance from the measurement vector. 8. The system of claim 5, wherein a primary axis vector having one of an activation gradient value within a predetermined range and a highest activation uniformity value within a predetermined range is selected from the measurement vectors, and the primary axis vector indicates a location of the reentrant circuit isthmus. 9. The system of claim 8, wherein the activation uniformity value is a coefficient of linear regression. 10. The system of claim 8, wherein the activation uniformity value is a coefficient of non-linear regression. 11. The system of claim 8, wherein the activation uniformity value is a variance in activation times along a selected measurement vector. 12. The system of claim 8, wherein the activation uniformity value is a measure of variability along a selected measurement vector. 13. The system of claim 8, wherein the activation gradient value is a slope of a linear regression line. 14. The system of claim 8, wherein the activation gradient value is a slope of a non-linear regression line. 15. The system of claim 8, wherein the activation gradient value is a mean absolute difference in activation times along a selected measurement vector. 16. The system of claim 8, wherein the activation gradient value is a difference along the measurement... | The Trustees of Columbia University in the City of New York | A61B005/04
US07186223 | B2 | 2007-03-06 | Guide wire | 2002-04-05 | A guide wire, which has excellent follow-up characteristics, sufficients contrasting characteristics and flexibility at a distal end portion thereof and high elasticity at a proximal end portion thereof, includes a core wire, and a coil wire provided on the distal end portion of this core wire coaxially therewith. The coil wire is formed by winding a plurality of wires, which are formed of different constituent materials and placed side by side on a plane, uniformly in an axial direction of the core wire so that the wires are adjacent to each other on the same plane. | 1. A guide wire having a distal end portion provided with radiation impermeability and flexibility, comprising a core wire having a distal end portion and a proximal end portion and a circular cross-section decreasing in diameter at the distal end portion toward a distal end of the core wire, and a coil wire having a constant diameter provided coaxially with the core wire and provided on the distal end portion of the core wire, wherein the coil wire consists of two wires which are adjacent to each other in an axial direction of the core wire and which are alternately wound around the core wire uniformly in an axial direction of the core wire,and wherein a distal end portion of one of said two wires is formed of a radiation impermeable material selected from the group consisting of platinum, gold and iridium and a remaining portion of which is formed of a radiation permeable and elastic material selected from the group consisting of stainless steel, piano wire and an amorphous alloy, and a whole of the other of said two wires is formed of a radiation permeable and elastic material selected from the group consisting of stainless steel, piano wire and an amorphous alloy or a radiation impermeable and elastic material selected from the group consisting of tungsten, tantalum, an alloy containing tungsten as a main component, and an alloy containing tantalum as a main component. | 1. A guide wire having a distal end portion provided with radiation impermeability and flexibility, comprising a core wire having a distal end portion and a proximal end portion and a circular cross-section decreasing in diameter at the distal end portion toward a distal end of the core wire, and a coil wire having a constant diameter provided coaxially with the core wire and provided on the distal end portion of the core wire, wherein the coil wire consists of two wires which are adjacent to each other in an axial direction of the core wire and which are alternately wound around the core wire uniformly in an axial direction of the core wire,and wherein a distal end portion of one of said two wires is formed of a radiation impermeable material selected from the group consisting of platinum, gold and iridium and a remaining portion of which is formed of a radiation permeable and elastic material selected from the group consisting of stainless steel, piano wire and an amorphous alloy, and a whole of the other of said two wires is formed of a radiation permeable and elastic material selected from the group consisting of stainless steel, piano wire and an amorphous alloy or a radiation impermeable and elastic material selected from the group consisting of tungsten, tantalum, an alloy containing tungsten as a main component, and an alloy containing tantalum as a main component. 2. A guide wire having a distal end portion provided with radiation impermeability and flexibility, comprising a core wire having a distal end portion and a proximal end portion and a circular cross-section decreasing in diameter at the distal end portion toward a distal end of the core wire, and a coil wire having a constant diameter provided coaxially with the core wire and provided on the distal end portion of the core wire, wherein the coil wire consists of two wires which are adjacent to each other in an axial direction of the core wire and which are alternately wound around the core wire uniformly in an axial direction of the core wire, and wherein a distal end portion of each of said two wires is formed of a radiation impermeable material selected from the group consisting of platinum, gold and iridium and a remaining portion of each of which is formed of a radiation permeable and elastic material selected from the group consisting of stainless steel, piano wire arid an amorphous alloy, the distal end portion formed of a radiation impermeable material of one of said two wires being longer than that of the other wire. 3. A guide wire having a distal end portion provided with radiation impermeability and flexibility, comprising a core wire having a distal end portion and a proximal end portion and a circular cross-section decreasing in diameter at the distal end portion toward a distal end of the core wire, and a coil wire having a constant diameter provided coaxially with the core wire and provided on the distal end portion of the core wire, wherein the coil wire consis... | Nipro Corporation | A61B005/00 A61M025/00
US07065398 | B2 | 2006-06-20 | Atrial tachycardia termination sensing | 2002-09-30 | Techniques for preventing inadvertent detection of AT termination may involve predicting potential undersensing of atrial events, and delaying detection of AT termination when undersensing of atrial events is predicted. For example, an implantable medical device may measure ventricular-atrial (VA) intervals of a heart rhythm using tracked atrial and ventricular events. The device calculates a slope of the series of the VA intervals, and estimates a prospective VA interval using the calculated slope and the most recently measured VA interval. The device determines whether the estimated VA interval is within a threshold range. When the estimated VA interval is outside of the threshold range, the device predicts possible undersensing of atrial events and delays detection of the AT episode termination. The delay detection of AT episode termination improves the accuracy of information collected by event counters. | 1. A method comprising: tracking atrial events and ventricular events of a patient to monitor an atrial tachyarrhythmic (AT) episode;predicting possible undersensing of an atrial event; anddelaying detection of AT episode termination when possible undersensing is predicted. | 1. A method comprising: tracking atrial events and ventricular events of a patient to monitor an atrial tachyarrhythmic (AT) episode;predicting possible undersensing of an atrial event; anddelaying detection of AT episode termination when possible undersensing is predicted. 2. The method of claim 1, wherein tracking atrial and ventricular events includes: tracking ventricular-atrial (VA) time intervals; andtracking ventricular-ventricular (VV) time intervals. 3. The method of claim 1, wherein predicting possible undersensing of an atrial event includes: tracking ventricular-atrial (VA) time intervals;calculating a slope of a series of the VA intervals;estimating a prospective VA time interval based on a previous VA time interval and the calculated slope; andpredicting possible undersensing of an atrial event when the estimated prospective VA time interval is outside of a range. 4. The method of claim 3, wherein calculating the slope of a series of the VA time intervals includes calculating the slope of a series of the VA time intervals using the equation m=[(VAn−VAn−2)+(VAn−2−VAn−4)+(VAn−4−VAn−6)]/3,wherein m represents the slope, VAn represents the most recent measured VA time interval of the series of VA time intervals, VAn−2 represents the third most recent measured VA time interval, VAn−4 represents the fifth most recent measured VA time interval, and VAn−6 represents the seventh most recent measured VA time interval. 5. The method of claim 4, wherein estimating a prospective VA time interval based on a previous VA time interval and the calculated slope includes estimating a prospective VA time interval according to the equation VAn+2=VAn+m,wherein VAn+2 is the prospective VA time interval, VAn is the most recent previous VA time interval, and m is the calculated slope. 6. The method of claim 3, further comprising predicting possible undersensing of an atrial event when the estimated prospective VA interval exceeds or falls below a threshold. 7. The method of claim 6, further comprising: tracking ventricular-ventricular (VV) time intervals; and predicting possible undersensing of an atrial event when the estimated prospective VA interval exceeds a current VV interval. 8. The method of claim 6, further comprising predicting possible undersensing of an atrial event when the estimated prospective VA interval is less than a blanking interval. 9. The method of claim 8, wherein the blanking interval is approximately 25 milliseconds. 10. The method of claim 3, further comprising: calculating a difference between previous VA intervals; andpredicting possible undersensing of an atrial event when the calculated difference between previous VA intervals is less than a threshold value. 11. The method of claim 10, wherein the threshold value is approximately 60. 12. The method of claim 3, further comprising: comparing the slope of the VA intervals with a threshold slope; andpredicting possible undersensing of an atrial event when the slope is le... | Medtronic, Inc. | A61B005/452
US07014612 | B2 | 2006-03-21 | Method for diagnosis of | 2002-11-13 | A rapid, non-invasive breath-test method and device for diagnosing the presence or absence of H. pylori in a subject without administration of isotopic tracers is described. The device consists of a highly sensitive colorimetric ammonia sensor placed in contact with sampled subject breath. The sensor is measured using appropriate reflection spectroscopy instrumentation. The breath-test method consists of measuring a basal ammonia level with the device, administering non-isotopic urea and continuing measurement of the ammonia content in a plurality of consecutive breaths. Diagnostic differences in breath ammonia are identified between H. pylori infected and uninfected individuals. | 1. A method for detecting the presence or absence of H. pylori infection in an individual comprising: a) exposing at least one ammonia sensitive sensor to expiration of an individual, wherein where more than one sensor is exposed, the sensors are exposed to identical samples of respiration;b) deriving, based on response of the at least one sensor to the exposure, a basal ammonia measure of expired ammonia over a basal measurement period; andc) comparing the basal ammonia measure against a normative value that reflects at least one of a H. pylori positive population measure and a H. pylori negative population measure, wherein the H. pylori status of the individual is determined. | 1. A method for detecting the presence or absence of H. pylori infection in an individual comprising: a) exposing at least one ammonia sensitive sensor to expiration of an individual, wherein where more than one sensor is exposed, the sensors are exposed to identical samples of respiration;b) deriving, based on response of the at least one sensor to the exposure, a basal ammonia measure of expired ammonia over a basal measurement period; andc) comparing the basal ammonia measure against a normative value that reflects at least one of a H. pylori positive population measure and a H. pylori negative population measure, wherein the H. pylori status of the individual is determined. 2. The method of claim 1, further comprising: a) administering a H. pylori urease enzyme substrate following determination of the basal ammonia measure;b) deriving a post-substrate ammonia measure of expired ammonia over a post-urea measurement period of about 10 minutes to about 90 minutes after the administration of the substrate; andc) comparing at least one measure of expired ammonia selected from the group consisting of the absolute measure of the post-substrate period, the absolute difference between the measures of the basal period and post-substrate period, the relative change between the measures of the basal period and post-substrate period, and, the rate of change in expired ammonia measures during the post-substrate period, against a normative value that reflects at least one of a H. pylori positive and a H. pylori negative population measures, wherein the H. pylori status of the individual is determined. 3. The method of claim 2, further comprising: a) administering an agent intended to increase expired ammonia either coincident or subsequent to administration of the H. pylori urease enzyme substrate;b) deriving a post-substrate and post-agent ammonia measure of expired ammonia over a post-agent measurement period of about 10 minutes to about 90 minutes after ingestion of the substrate and agent; andc) comparing at least one measure of expired ammonia selected from the group consisting of the absolute measure of the post-agent period, the absolute difference between the measure of the basal period, post-urea period and post-agent period, the relative change between the measures of the basal period, post-urea period, and post-agent period, and, the rate of change in expired ammonia measures during the basal period, post-substrate period, and post-agent period, against a normative value that reflects at least one of a H. pylori positive and H. pylori negative population values, wherein the H. pylori status of the individual is determined. 4. The method of claim 3, wherein the agent intended to increase expired ammonia either coincident or subsequent to administration of the H. pylori urease enzyme substrate is an antacid. 5. The method of claim 3, wherein the agent intended to increase expired ammonia either coincident or subsequent to administration of the... | Photonic Biosystems, Inc. | A61B005/08 G01N001/22
US07048687 | B1 | 2006-05-23 | Limited use medical probe | 2003-02-06 | A limited use medical probe is disclosed, including a memory for maintaining a use value. The medical probe is coupled to a medical device that inhibits its function when the use value reaches a predetermined threshold value, preventing improper use of the probe. The probe memory may also store identification, usage, and clinical data. A probe auto-identification function, a probe re-identification function and a probe functional test sequence are disclosed for the medical probe. After use, a reprocessing step may reset the probe memory, permitting further probe use. | 1. A method for limiting the use of a medical probe according to a set of usage criteria using a security function, the method comprising the following steps: (a) coupling a probe memory to a medical probe, the probe memory having an identifying data about the medical probe, therein;(b) coupling the medical probe to a medical device;(c) coupling the medical probe to a patient;(d) retrieving a set of use values associated with use of the medical probe from the probe memory;(e) using a security function to verify that the contents of the probe memory have not been corrupted or tampered with;(f) applying a set of usage criteria periodically during use of the medical probe, comprising repetition of the following steps: (i) updating the set of use values of the medical probe during use and creating an updated set of use values;(ii) testing the updated set of use values against the set of usage criteria related to the medical probe;(iii) performing an expiration response when one or a combination of the usage criteria are violated;(iv) storing the updated set of use values on a periodic basis in the probe memory;(g) retrieving and storing the identifying data about the medical probe from the probe memory, and storing the data in a memory component of the medical device;(h) storing a time and date stamp of the use of the medical probe memory in a memory component of the medical device;(i) comparing the identifying data about the medical probe, read from the probe memory, with the identifying data and time and date stamp of the use(s) corresponding to a previously coupled medical probe(s) as stored in the medical device memory; and including determining whether an interrupted usage of the medical probe has resumed, or a new usage of the medical probe has begun. | 1. A method for limiting the use of a medical probe according to a set of usage criteria using a security function, the method comprising the following steps: (a) coupling a probe memory to a medical probe, the probe memory having an identifying data about the medical probe, therein;(b) coupling the medical probe to a medical device;(c) coupling the medical probe to a patient;(d) retrieving a set of use values associated with use of the medical probe from the probe memory;(e) using a security function to verify that the contents of the probe memory have not been corrupted or tampered with;(f) applying a set of usage criteria periodically during use of the medical probe, comprising repetition of the following steps: (i) updating the set of use values of the medical probe during use and creating an updated set of use values;(ii) testing the updated set of use values against the set of usage criteria related to the medical probe;(iii) performing an expiration response when one or a combination of the usage criteria are violated;(iv) storing the updated set of use values on a periodic basis in the probe memory;(g) retrieving and storing the identifying data about the medical probe from the probe memory, and storing the data in a memory component of the medical device;(h) storing a time and date stamp of the use of the medical probe memory in a memory component of the medical device;(i) comparing the identifying data about the medical probe, read from the probe memory, with the identifying data and time and date stamp of the use(s) corresponding to a previously coupled medical probe(s) as stored in the medical device memory; and including determining whether an interrupted usage of the medical probe has resumed, or a new usage of the medical probe has begun. 2. The method as defined in claim 1, wherein Step (e), the security function includes a physical security method, or an algorithmic security method, or both. 3. The method as defined in claim 2, wherein the physical security method includes a portion of the probe memory that is a write-once memory device. 4. The method as defined in claim 2, wherein the algorithmic security function includes a data encoding system with encryption: in the medical device for encrypting data before it is stored in the probe memory, and for decrypting data read from the probe memory in the medical device; orin the medical probe for encrypting data from the medical device in the probe memory when it is stored therein, and for decrypting data stored in the probe memory when it is read by the medical device;or in both. 5. The method as defined in claim 4, wherein the data encoding system incorporates symmetrical encryption or wherein the data encoding system incorporates asymmetrical encryption. 6. The method as defined in claim 5, wherein the data encoding system incorporates public key encryption. 7. The method as defined in claim 1, wherein the probe memory contains the set of usage criteria. 8. The method as d... | OB Scientific, Inc. | A61B005/00
US07004908 | B2 | 2006-02-28 | Method and apparatus useful in the diagnosis of obstructive sleep apnea of a patient | 2002-01-18 | Patients may operate a CPAP system to deliver appropriate airway pressure at their home. A patient's apnea problem ca be diagnosed at home without supervision with a CPAP device which delivers a continuously minimum appropriate pressure for substantially the entire period of therapy. | 1. A method of diagnosis including applying to a patient before sleep a nose piece containing a pressure sensor connected to a pressure signal recording apparatus, and identifying snoring patterns or other abnormal breathing patterns where indicated by the recorded signal so as to diagnose occurrences of apneic episodes or other respiratory disorders. | 1. A method of diagnosis including applying to a patient before sleep a nose piece containing a pressure sensor connected to a pressure signal recording apparatus, and identifying snoring patterns or other abnormal breathing patterns where indicated by the recorded signal so as to diagnose occurrences of apneic episodes or other respiratory disorders. 2. A method as defined in claim 1 wherein the pressure signals are processed by an electronic processor programmed so as to distinguish predefined snoring patterns or other abnormal breathing patterns and provide signals to the recording apparatus indicative of the occurrence of the predefined snoring pattern or other abnormal breathing pattern. 3. A method of CPAP therapy including: monitoring a patient's snoring and/or breathing patterns using a pressure sensor in pressure wave communication with the patient's respiratory system; identifying predetermined characteristic snoring and/or breathing patterns by electronic analysis of signals from the sensor; and automatically increasing CPAP pressure delivered to the patient by a predetermined amount responsive to each identified said characteristic snoring or abnormal breathing pattern. 4. A method as defined in claim 3 wherein CPAP pressure delivered to the patient is automatically gradually decreased in the absence of said characteristic snoring or abnormal breathing patterns. 5. A method as defined in claim 3 wherein CPAP therapy is commenced at a low CPAP pressure delivered to the patient. 6. A method of CPAP therapy including commencing therapy at a preset minimum CPAP pressure delivered to the patient and then gradually increasing said pressure over a preset period of time to a preset therapeutic pressure. 7. A method of CPAP therapy as defined in claim 3 and further comprising recording pressure levels, against time, delivered to the patient. 8. A method of CPAP therapy as defined in claim 3 and further comprising recording snoring or abnormal breathing patterns against time so as to confirm correct operation of the CPAP apparatus. 9. A method useful in the diagnosis of obstructive sleep apnea of a patient, the method comprising: providing a signal indicative of the patient's snoring and flow rate; sensing the patient's snoring and flow rate; and recording and storing the signal over a period of time. 10. A method according to claim 9, further comprising processing the signal to diagnose obstructive sleep apnea of the patient. 11. A method according to claim 9, further comprising: providing a predetermined snoring and flow rate pattern; comparing the patient's snoring and flow rate with the predetermined snoring and flow rate pattern; and controlling pressure within the nasal mask based on the comparison between the patient's snoring and flow rate and the predetermined snoring and flow rate pattern. 12. A method according to claim 9, further comprising connecting a pressure sensor to the nasal mask adjacent an air inlet of the nas... | ResMed Limited | A61B005/08 A61B019/00
US07043295 | B2 | 2006-05-09 | Methods and apparatus for delivering a drug influencing appetite for treatment of eating disorders | 2002-04-26 | Methods and systems for treating patients suffering from eating disorders, e.g. obesity, through the dispensation of a drug by an implantable infusion pump (IIP) delivering drug into the cerebral spinal fluid (CSF) at a site of the intrathecal space in amounts and at times effective to suppress the patient's appetite through interaction of the drug transported through the CSF with receptors in the brain. Delivery of a programmed drug dosage is preferably by one of time-out of programmed time(s) of day, a command received from the patient, or a trigger signal developed from a sensed GI tract signal accompanying peristalsis. | 1. A method of delivering drug from an implantable infusion pump to receptors of the brain to therapeutically treat an eating disorder of a patient by suppressing or enhancing appetite, the method comprising the steps of: surgically implanting an intrathecal drug infusion catheter so that a distal drug delivery portion lies in a predetermined intrathecal space site of the spinal column in cerebral spinal fluid (CSF) communication with receptors in the brain;coupling the proximal end of the drug infusion catheter to an drug infusion pump;implanting the drug infusion pump in the patient's body; andoperating the pump to discharge a predetermined dosage of drug into the CSF sufficient to suppress or enhance appetite through interaction of drug transported through the CSF with receptors in the brain. | 1. A method of delivering drug from an implantable infusion pump to receptors of the brain to therapeutically treat an eating disorder of a patient by suppressing or enhancing appetite, the method comprising the steps of: surgically implanting an intrathecal drug infusion catheter so that a distal drug delivery portion lies in a predetermined intrathecal space site of the spinal column in cerebral spinal fluid (CSF) communication with receptors in the brain;coupling the proximal end of the drug infusion catheter to an drug infusion pump;implanting the drug infusion pump in the patient's body; andoperating the pump to discharge a predetermined dosage of drug into the CSF sufficient to suppress or enhance appetite through interaction of drug transported through the CSF with receptors in the brain. 2. The method of claim 1, wherein the operating step further comprises: providing the patient with an external medical device for communicating with the pump; andutilizing the external medical device to communicate a command to the pump to discharge the predetermined dosage of drug into the CSF sufficient to suppress or enhance appetite through interaction with receptors in the brain. 3. The method of claim 2, wherein the operating step further comprises: timing out a delay time from each discharge; andinhibiting the pump from responding to the command if the delay time has not timed out. 4. The method of claim 1, wherein the drug is insulin acting to suppress appetite when detected by the receptors in the brain, and the operating step further comprises: providing the patient with an external medical device for communicating with the pump; andutilizing the external medical device to communicate a command to the pump to discharge a dosage of insulin in the range of 1×10−5–15×10−2 Units/Kg/day into the CSF sufficient to suppress appetite through interaction with receptors in the brain. 5. The method of claim 1, wherein the drug is insulin acting to suppress appetite when detected by the receptors in the brain, and the operating step further comprises: providing the patient with an external medical device for communicating with the pump; andutilizing the external medical device to communicate a command to the pump to discharge the predetermined dosage of insulin into the CSF sufficient to suppress or enhance appetite through interaction with receptors in the brain. 6. The method of claim 1, wherein the operating step further comprises: timing out the time of day; andoperating the drug infusion pump at one or more predetermined time of day to discharge the predetermined dosage of drug into the CSF sufficient to suppress or enhance appetite through interaction with receptors in the brain. 7. The method of claim 1, wherein the drug is insulin acting to suppress appetite when detected by the receptors in the brain, and the operating step further comprises: timing out the time of day; andoperating the pump at one or more predetermined time of day to dischar... | Medtronic, Inc. | A61B005/04 A61K009/22 A61M031/00
US07054676 | B2 | 2006-05-30 | MR-compatible methods and systems for cardiac monitoring and gating | 2001-04-24 | Noninvasive, MR-compatible methods and systems optically detect mechanical cardiac activity by anatomic (e.g., esophageal) movements. Most preferably, esophageal motion is detected optically and is indicative rhythmic cardiac activities. This esophageal motion may then be detected and used to provide a signal indicative of periods of cardiac activity and inactivity. The signal may be further processed so as to generate a trigger signal that may be input to a MR scanner. In such a manner, MR microscopy may be accomplished to acquire information at a specific phase of the cardiac cycle, for example, in synchrony with periods of cardiac inactivity. Moreover, since mechanical cardiac activity is detected and employed, instead of electrical activity as is employed in conventional techniques, the present invention is immune to electromagnetic interference during MR microscopy. As a result, robust cardiac signals may be monitored and gated during 2-dimensional and 3-dimensional in vivo microscopy. The present invention is therefore especially well suited for MR microscopy of small animals, such as laboratory mice and rats. | 1. A method for detecting a phase in a cardiac cycle comprising the steps of: (a) optically detecting movements of an anatomic structure affected by cardiac activity;(b) deriving a cardiac signal in response to said optically detected movements which is indicative of a phase in a cardiac cycle; and(c) generating a trigger signal in response to said derived cardiac signal which is indicative of said phase of the cardiac cycle. | 1. A method for detecting a phase in a cardiac cycle comprising the steps of: (a) optically detecting movements of an anatomic structure affected by cardiac activity;(b) deriving a cardiac signal in response to said optically detected movements which is indicative of a phase in a cardiac cycle; and(c) generating a trigger signal in response to said derived cardiac signal which is indicative of said phase of the cardiac cycle. 2. The method of claim 1, wherein step (a) is practiced by inserting an optical fiber esophageal probe into an esophagus. 3. The method of claim 2, which comprises directing laser light onto a site within the esophagus through at least one optical fiber in the probe, and receiving reflected light from said esophagus site through at least one other optical fiber in the probe. 4. The method of claim 3, which comprises directing said reflected light to a photodetector which derives said cardiac signal therefrom. 5. The method of claim 4, which comprises receiving the derived cardiac signal from said photodetector and generating said trigger signal therefrom. 6. A method of conducting magnetic resonance (MR) microscopy comprising the steps of: (a) optically detecting internal anatomic physical movement in response to rhythmic periods of cardiac activity and inactivity during a cardiac cycle;(b) conducting MR microscopy during said periods of cardiac inactivity; and(c) using the derived cardiac signal to produce cardiac images at specific phases of the cardiac cycle. 7. The method of claim 6, wherein step (a) includes deriving a cardiac signal indicative of said rhythmic periods of cardiac activity and inactivity. 8. The method of claim 7, wherein step (a) further includes generating a trigger signal in response to said derived cardiac signal of said periods of cardiac inactivity. 9. The method of claim 6, wherein step (a) includes generating a trigger signal in response to said optically detected periods of cardiac inactivity. 10. The method of claim 6, wherein step (a) is practiced by inserting an optical fiber esophageal probe into the esophagus. 11. The method of claim 10, wherein step (a) is practiced by optically detecting esophageal movements indicative of said periods of cardiac activity and inactivity. 12. The method of claim 11, comprising directing laser light onto a site within the esophagus through at least one optical fiber in the probe, and receiving reflected light from said esophagus site through at least one other optical fiber in the probe. 13. The method of claim 12, which comprises directing said reflected light to a photodetector which derives said cardiac signal therefrom. 14. The method of claim 13, which comprises processing the derived cardiac signal by receiving the derived cardiac signal from said photodetector and generating said trigger signal therefrom. 15. A magnetic resonance (MR) microscopy method comprising the steps of: (a) inserting an optical fiber probe into a vertebrate esopha... | Duke University | A61B005/05
US07008387 | B2 | 2006-03-07 | Portable device for collecting information about living body | 2002-07-26 | A portable device for collecting information about a human body is adapted to be worn on the human body and is capable of periodically performing cognitive faculty tests during daily life to discern fluctuations in the memory faculty of a wearer. The device displays a problem of cognitive tests on a display unit, and a wearer inputs an answer with an input unit. The wearer's response time is clocked with a clock unit, and the answer and the response time are stored in a memory unit in time sequence by control of a control circuit. The device can include a detection unit for detecting a living body signal to enable research on the relationship between behavioral patterns or temperamental trends and memory faculties, and the like. | 1. A portable device for collecting information about a user, comprising: a case mountable on the user's wrist; storage means disposed in the case for storing a plurality of characters; character selection means for selecting a character from the storage means; a display provided on the case for displaying the selected character; input means for allowing user input of information about the character displayed on the display; measuring means disposed in the case for measuring elapsed time from the display of the character until occurrence of the user input; memory means disposed in the case for storing the information input by the user and the time measured by the measuring means; control means for controlling operations of the character selecting means and the memory means control means for controlling operations of the character selecting means and the memory means, the control means having a timer, and, at predetermined time intervals measured by the timer, causing the memory means to store the information input by the user with the input means and the time measured by the measuring means; and delaying means for delaying the start of operation of the memory means to cause the memory means to store the information input by the user with the input means and the time measured by the measuring means at a different time delayed from the predetermined time intervals. | 1. A portable device for collecting information about a user, comprising: a case mountable on the user's wrist; storage means disposed in the case for storing a plurality of characters; character selection means for selecting a character from the storage means; a display provided on the case for displaying the selected character; input means for allowing user input of information about the character displayed on the display; measuring means disposed in the case for measuring elapsed time from the display of the character until occurrence of the user input; memory means disposed in the case for storing the information input by the user and the time measured by the measuring means; control means for controlling operations of the character selecting means and the memory means control means for controlling operations of the character selecting means and the memory means, the control means having a timer, and, at predetermined time intervals measured by the timer, causing the memory means to store the information input by the user with the input means and the time measured by the measuring means; and delaying means for delaying the start of operation of the memory means to cause the memory means to store the information input by the user with the input means and the time measured by the measuring means at a different time delayed from the predetermined time intervals. 2. A portable device for collecting information about a living body according to claim 1; wherein the delaying means includes a switch mounted on the case. 3. A portable device for collecting information about a user according to claim 1; wherein the memory means stores the information input by the user with the input means and the time measured by the measuring means for a plurality of sets of sequentially displayed characters that are displayed in time sequence and for a given duration of time. 4. A portable device for collecting information about a user according to claim 1; wherein the input means comprises a cursor key mounted on the case. 5. A portable device for collecting information about a user according to claim 1; further comprising detection means for detecting bodily movements of the user. 6. A portable device for collecting information about a user according to claim 5; wherein the detection means is an acceleration sensor. 7. A portable device for collecting information about a user, comprising: a case mountable on the user's wrist; storage means disposed in the case for storing a plurality of characters; character selection means for sequentially selecting the plurality of characters from the storage means; a display provided on the case for sequentially displaying the selected plurality of characters; input means provided on the case for allowing user input of information about sequentially displayed characters; measuring means disposed in the case for measuring elapsed time from the display of the sequentially displayed characters until occurrence of the user inp... | Seiko Instruments Inc. University of Medicine and Dentistry of New Jersey | A61B005/103 A61B005/117
US07248914 | B2 | 2007-07-24 | Method of navigating medical devices in the presence of radiopaque material | 2002-06-28 | A method of navigating the distal end of a medical device through an operating region in a subject's body includes displaying an x-ray image of the operating region, including the distal end of the medical device; determining the location of the distal end of the medical device in a reference frame translatable to the displayed x-ray image; and displaying an enhanced indication of the distal end of the medical device on the x-ray image to facilitate the navigation of the distal end of the device in the operating region. | 1. A method of navigating the distal end of a medical device through an operating region in a subject's body, the method comprising: introducing the distal end of an elongate medical device having at least one magnetically responsive element fixed thereon into the subiect's body;navigating the distal end of the medical device to the desired site of delivery;expelling a flowable radiopague substance from the medical device at the desired site of delivery;displaying an X-ray image of the operating region, including the distal end of the medical device;determining the location of a localization coil in the distal end of the medical device relative to at least two transmitters at known reference locations, to establish a distal end location in a reference frame translatable to the displayed image; anddisplaying an enhanced indication of the distal end of the medical device on the X-ray image of the operating region in which the distal end of the medical device is in the presence of the expelled flowable radiopague substance, to facilitate the navigation of the distal end of the device in the operating region. | 1. A method of navigating the distal end of a medical device through an operating region in a subject's body, the method comprising: introducing the distal end of an elongate medical device having at least one magnetically responsive element fixed thereon into the subiect's body;navigating the distal end of the medical device to the desired site of delivery;expelling a flowable radiopague substance from the medical device at the desired site of delivery;displaying an X-ray image of the operating region, including the distal end of the medical device;determining the location of a localization coil in the distal end of the medical device relative to at least two transmitters at known reference locations, to establish a distal end location in a reference frame translatable to the displayed image; anddisplaying an enhanced indication of the distal end of the medical device on the X-ray image of the operating region in which the distal end of the medical device is in the presence of the expelled flowable radiopague substance, to facilitate the navigation of the distal end of the device in the operating region. 2. The method according to claim 1, wherein the step of determining the location of the distal end of the medical device comprises transmitting signals between at least one reference location and the distal end of the medical device. 3. The method according to claim 2 wherein the step of transmitting signals between at least one reference location and the distal end of the device comprises transmitting signals from a plurality of transmitters in more than one reference relocation, and a receiver on the magnetic medical device. 4. The method according to claim 2 wherein the medical device has a transmitter or receiver coil made from a radio opaque material. 5. The method according to claim 4 wherein the coil is made from at least one of gold or platinum, or an alloy containing gold or platinum. 6. The method according to claim 1, wherein the step of determining the location of the distal end of the medical device comprises moving the distal end of the medical device, and processing x-ray images before and after the movement to determine the location of the distal tip. 7. The method according to claim 6, wherein the distal end of the medical device includes a magnetically responsive element, and wherein the step of moving the distal end of the medical device comprises applying a magnetic field or gradient from a magnet outside the subject's body to move the magnetically responsive element and thus the distal end of the medical device. 8. The method according to claim 7 wherein the medical device has a magnetically responsive element, wherein the medical device is oriented by the applying a magnetic field from a magnetic surgery system outside the subject's body, and wherein the step of displacing the distal tip comprises changing the magnetic field applied by the magnetic surgery system to the magnetically responsive element. 9. The method... | Stereotaxis, Inc. | A61B005/05
US06993377 | B2 | 2006-01-31 | Method for diagnosing heart disease, predicting sudden death, and analyzing treatment response using multifractal analysis | 2002-02-22 | A method of analyzing electrocardiogram (EKG) data for use in the diagnosis of heart disease, prognosis of cardiac conditions, and the monitoring of heart disease therapies is disclosed. The method utilizes a wavelet-based multifractal analysis with one or more of (1) a discrete wavelet smoothing step to remove the effects of abnormal beats; (2) “Levy flight” analysis to detect the frequency of abnormal beats known to adversely affect the multifractal (MF) analysis; and (3) MF alpha analysis, a multifractal extension of monofractal short term (ST) alpha analysis. The invention further comprises an EKG test battery comprising Levy flight anomalous beat/beat cluster screening, followed by (ST) MF alpha analysis and MF Holder analysis (when validated by the Levy flight analysis). The wavelet smoothing step can also be used to classify human EKGs by observing the effect of sequential smoothing on the MF Holder coefficient. Alternative choices to the wavelet smoothing approach to removal of abnormal beat effects include probability distribution function analysis to determine the MF Holder coefficient directly, abnormal beat ridge skeleton removal to remove the offending beats based on a direct multifractal spectrum calculation, and the calculation of various types of entropy coefficients for the EKG time series. | 1. A method of predicting a patient's risk of sudden death from cardiac disease, comprising the steps of: (a) collecting electrocardiogram data; (b) smoothing the data by applying a wavelet transform filter to the data; and (c) applying wavelet-based multifractal analysis to the data. | 1. A method of predicting a patient's risk of sudden death from cardiac disease, comprising the steps of: (a) collecting electrocardiogram data; (b) smoothing the data by applying a wavelet transform filter to the data; and (c) applying wavelet-based multifractal analysis to the data. 2. The method of claim 1, wherein the step of applying multifractal analysis to the data comprises the step of calculating a multifractal predictive coefficient, wherein the predictive coefficient has a numerical value that correlates to the patient's risk of sudden death from cardiac disease. 3. The method of claim 2, wherein said step of calculating a multifractal coefficient comprises the step of calculating a multifractal Holder coefficient. 4. The method of claim 1, wherein said step of applying a wavelet transform filter comprises the step of applying a MODWT wavelet transform filter. 5. The method of claim 1, further comprising the step of performing a probability distribution on the data prior to the step of applying wavelet-based multifractal analysis. 6. The method of claim 1, further comprising the step of performing a cluster analysis to the data. 7. The method of claim 6, wherein said step of performing a cluster analysis comprises the step of performing a Levy flight cluster analysis. 8. The method of claim 1, further comprising the step of calculating a short-scale multifractal coefficient, wherein said short-scale multifractal coefficient characterizes the short-scale beat region of the data. 9. The method of claim 8, wherein said step of calculating a short-scale coefficient comprises the step of calculating a multifractal alpha coefficient. 10. The method of claim 9, wherein said multifractal alpha coefficient characterizes a scale of one of 1-4 to 1-16 beats. 11. The method of claim 1, further comprising the step of quantifying the multifractality of the data by generating a multifractal cascade coefficient. 12. The method of claim 11, wherein said step of quantifying the multifractality of the data includes the step of generating a quadratic coefficient, wherein the value of said quadratic coefficient is indicative of the degree of multifractality of the data. 13. The method of claim 12, wherein said quadratic coefficient is indicative of multifractal data when the quadratic coefficient has a value greater than a threshold value. 14. A method of determining the severity of cardiac disease, comprising the steps of: (a) collecting electrocardiogram data; (b) performing a probability distribution on the data; and (c) applying wavelet-based multifractal analysis to the data. 15. The method of claim 14, wherein the step of applying multifractal analysis to the data comprises the step of calculating a multifractal predictive coefficient, wherein the predictive coefficient has a numerical value that correlates to the patient's risk of sudden death from cardiac disease. 16. The method of claim 15, wherein said step of calculating a multi... | The Board of Trustees of the University of Arkansas | A61B005/0452
US07035685 | B2 | 2006-04-25 | Apparatus and method for measuring electroencephalogram | 2002-06-06 | An electroencephalogram measuring apparatus and method. The electroencephalogram measuring method includes detecting electroencephalogram data of an experimentee, amplifying the detected electroencephalogram data, and converting the data into a digital signal. Also, provided is a high-speed Fourier transforming the converted digital electroencephalogram data, calculating an output value for each frequency of the Fourier transformed electroencephalogram, calculating an output value for each frequency of the electroencephalogram for a predetermined time interval and calculating a relative output value for the ground state of the experimentee. Further provided is comparing the calculated relative output value for each frequency for each predetermined time interval for the ground state of the experimentee, with a predetermined value. | 1. An electroencephalogram measuring apparatus comprising: an electroencephalogram detection unit which detects electroencephalogram data; an electroencephalogram amplifying unit which amplifies the electroencephalogram data measured in the electroencephalogram detection unit; an A/D converting unit which converts the amplified electroencephalogram data from an analog signal to a digital signal; an electroencephalogram processing unit which, using the electroencephalogram data which is converted into a digital signal, outputs an output value for each frequency at each time interval for the ground state of an experimentee, and determines whether or not noise waves are combined with the electroencephalogram; and a display unit which displays the result of the electroencephalogram processing unit. | 1. An electroencephalogram measuring apparatus comprising: an electroencephalogram detection unit which detects electroencephalogram data; an electroencephalogram amplifying unit which amplifies the electroencephalogram data measured in the electroencephalogram detection unit; an A/D converting unit which converts the amplified electroencephalogram data from an analog signal to a digital signal; an electroencephalogram processing unit which, using the electroencephalogram data which is converted into a digital signal, outputs an output value for each frequency at each time interval for the ground state of an experimentee, and determines whether or not noise waves are combined with the electroencephalogram; and a display unit which displays the result of the electroencephalogram processing unit. 2. The apparatus of claim 1, wherein the electroencephalogram amplifying unit amplifies sensed electroencephalogram data and filter a predetermined frequency band of the data. 3. The apparatus of claim 1, wherein the electroencephalogram processing unit comprises: a Fourier transform unit which separates a sine wave from the electroencephalogram data amplified in the electroencephalogram amplifying unit; a data storage unit which stores the electroencephalogram of the ground state of the experimentee; and a control unit which calculates a relative output value for each frequency at each predetermined time interval of the ground state of the experimentee, by using the output value for each frequency of the electroencephalogram obtained by the Fourier transform unit, and the output value for each frequency obtained from the electroencephalogram data of the ground state of the experimentee. 4. The apparatus of claim 3, wherein the Fourier transform unit high-speed Fourier transforms the electroencephalogram data. 5. The apparatus of claim 3, wherein the data storage unit which stores the electroencephalogram of the ground state stores electroencephalogram data measured in an eye-closed and relaxed state, or in an interval before stimulus. 6. The apparatus of claim 3, wherein the control unit comprises: a first calculating unit which calculates an output value for each frequency of the electroencephalogram which is Fourier transformed; a second calculating unit which calculates the size of an output value for each frequency at each predetermined time interval using the output of the first calculating unit; a third calculating unit which calculates a relative output value for each frequency at each predetermined time interval of the ground state, by dividing the output of the second calculating unit by the output value for each frequency of the ground state of the experimentee; and a comparison unit which compares the output of the third calculating unit with a predetermined value. 7. The apparatus of claim 6, wherein in the second calculating unit which calculates an output value for each frequency at each predetermined time interval, the time interva... | Electronics and Telecommunications Research Institute | A61B005/04
US07087027 | B2 | 2006-08-08 | Device and method for monitoring respiration | 2002-04-22 | A system and method are provided to monitor a patient's respiration during and after medical/surgical procedures, as well as in clinical situations wherein the patient is at increased risk for developing central or obstructive respiratory depression and/or apnea. A sensor attached to the patient's face, which monitors nasal and oral air-flows. An electronic monitor analyzes patient respiratory patterns and initiates bedside and/or remote nursing station alarms in real time if and when a clinically significant respiratory event is detected. Respiration is measured by monitoring changes in the acoustic signature of the patient's breath passing over a corrugated or uneven surface, changes in the pressure of an inflatable closed volume of a sensor caused by the changes in temperature of the interior volume caused by the patient breathing over the sensor, or both. | 1. A respiration sensor apparatus comprising: a sensor capable of sensing relative high frequency acoustic signals with a frequency of at least 2 Khz, including an inflatable body having an outside surface and an interior volume, disposable between the mouth and nose of a patient; at least one tube having two ends, one end being in communication with said interior volume and extending from said sensor; and a monitor, connected to said second end of said tube, detecting respiration of the patient by detecting variations in at least one of a) pressure within said interior volume, and b) sound waves generated by the patient's respiration flowing over said sensor. | 1. A respiration sensor apparatus comprising: a sensor capable of sensing relative high frequency acoustic signals with a frequency of at least 2 Khz, including an inflatable body having an outside surface and an interior volume, disposable between the mouth and nose of a patient; at least one tube having two ends, one end being in communication with said interior volume and extending from said sensor; and a monitor, connected to said second end of said tube, detecting respiration of the patient by detecting variations in at least one of a) pressure within said interior volume, and b) sound waves generated by the patient's respiration flowing over said sensor. 2. A respiration sensor apparatus according to claim 1, further comprising corrugations disposed on said outside surface, wherein when the patient breathes over said sensor, said corrugations increase turbulence of the respiratory airflow over said outer surface to thereby increase the amount of sound generated by respiration. 3. A respiration sensor apparatus according to claim 1, said inflatable body further comprising: a bendable inelastic substrate; and a plastic inflatable film sealingly attached to said substrate to form said interior volume. 4. A respiration sensor apparatus according to claim 1, said inflatable body having a proximal surface disposable against the face of the patient, further comprising a hole formed in said proximal surface communicating with said interior volume, wherein when said sensor is disposed on the patient's face, said hole is covered and closed by the patient's face. 5. A respiration sensor apparatus according to claim 4, wherein if said sensor loses contact with the patient's face, said hole opens, and said interior volume becomes in communication with the ambient air, and said monitor detects the concomitant change in pressure of the interior volume. 6. A respiration sensor apparatus according to claim 4, further comprising an adhesive disposed on said proximal surface for adhering said sensor to the patient's face. 7. A respiration sensor apparatus according to claim 5, said monitor further comprising a pump in communication with said tube providing positive pressure into said tube and said interior volume. 8. A respiration sensor apparatus according to claim 1, said monitor further comprises a pressure transducer in communication with said tube for measuring the pressure in said tube and said interior volume. 9. A respiration sensor apparatus according to claim 8, wherein when the patient breathes over said sensor, the breathing causes a change in temperature of said interior volume which in turn causes a change in pressure of said interior volume. 10. A respiration sensor apparatus according to claim 9, said monitor further comprising a central processing unit (CPU), connected to said pressure transducer, analyzing pressure readings detected by said pressure transducer and determining a rate of respiration of the patient from the pressure... | - | A61B005/28
US07181268 | B2 | 2007-02-20 | Ischemia detection | 2002-11-25 | Techniques for detection and treatment of myocardial ischemia are described that monitor both the electrical and dynamic mechanical activity of the heart to detect and verify the occurrence of myocardial ischemia in a more reliable manner. The occurrence of myocardial ischemia can be detected by monitoring changes in an electrical signal such as an ECG or EGM, and changes in dynamic mechanical activity of the heart that are sensed by an accelerometer sensor. The heart acceleration signal can be obtained from an single- or multiple-axis accelerometer and/or a pressure sensor deployed within or near the heart. The techniques correlate contractility changes detected by an accelerometer or pressure sensor with changes in the ST electrogram segment detected by the electrodes to increase the reliability of ischemia detection. | 1. A method for detecting myocardial ischemia, the method comprising by: obtaining a first signal indicative of dynamic mechanical activity of a heart;obtaining a second signal indicative of electrical activity of the heart;detecting a myocardial isehemia condition based on both the first signal and the second signal;storing a peak endocardial acceleration value in a computer memory location;after a predetermined time interval, obtaining at least one new peak endocardial acceleration value;storing the at least one new peak endocardial acceleration signal; andcomparing the peak endocardial acceleration to the at least one new peak endocardial acceleration value;wherein obtaining the first signal includes obtaining a heart acceleration signal from an accelerometer deployed within the heart and wherein said heart acceleration signal is a peak endocardial acceleration value. | 1. A method for detecting myocardial ischemia, the method comprising by: obtaining a first signal indicative of dynamic mechanical activity of a heart;obtaining a second signal indicative of electrical activity of the heart;detecting a myocardial isehemia condition based on both the first signal and the second signal;storing a peak endocardial acceleration value in a computer memory location;after a predetermined time interval, obtaining at least one new peak endocardial acceleration value;storing the at least one new peak endocardial acceleration signal; andcomparing the peak endocardial acceleration to the at least one new peak endocardial acceleration value;wherein obtaining the first signal includes obtaining a heart acceleration signal from an accelerometer deployed within the heart and wherein said heart acceleration signal is a peak endocardial acceleration value. 2. A method according to claim 1, wherein the first signal and the second signal both satisfy criteria for indication of ischemia. 3. A method according to claim 2, wherein the criteria include a change in the first signal having a first predetermined relationship to a first threshold and a change in the second signal having second predetermined relationship to a second threshold. 4. A method according to claim 1, wherein said accelerometer is deployed spaced from an adjacent portion of endocardial tissue. 5. A method according to claim 4, wherein said adjacent portion of endocardial tissue is a one of the following: a portion of transseptal tissue, a portion of right ventricle tissue, a portion of ventricular outflow tract tissue, a portion of left ventricle tissue. 6. The method of claim 1, further comprising, when myocardial ischemia is detected, generating a signal for delivery of therapy to alleviate effects of the ischemia within the heart. 7. A method according to claim 1, wherein said accelerometer is a multiple axis accelerometer. 8. A computer readable medium for storing computer executable instructions for performing a method for detecting myocardial ischemia, comprising: executable instructions for obtaining a first signal indicative of dynamic mechanical activity of a heart from an accelerometer sensor;executable instructions for obtaining a second signal indicative of electrical activity of the heart;executable instructions for detecting a myocardial ischemia condition based on both the first signal and the second signal;executable instructions for storing the peak endocardial acceleration value in a computer memory location;executable instructions for, after a predetermined time interval, obtaining at least one new peak endocardial acceleration value;executable instructions for storing the at least one new peak endocardial acceleration signal; andexecutable instructions for comparing the peak endocardial acceleration to the at least one new peak endocardial acceleration valuewherein the executable instructions for obtaining the first signal further comprises... | Medtronic, Inc. | A61B005/0402
US07136492 | B2 | 2006-11-14 | Visual or audio playback of an audiogram | 2002-07-11 | A method for the visualization of the hearing capacity of a person. As a function of audiogram data, text is modified in any of several characteristics, including brightness, contrast, sharpness, and/or omission of letters. | 1. A method for visualization of a hearing capacity of a person, comprising the steps of: using audiogram data to modify a text sample in its brightness or its contrast against its background, in the sharpness of its contours, and/or through the total omission of individual letters, and displaying the modified text. | 1. A method for visualization of a hearing capacity of a person, comprising the steps of: using audiogram data to modify a text sample in its brightness or its contrast against its background, in the sharpness of its contours, and/or through the total omission of individual letters, and displaying the modified text. 2. The method as in claim 1, wherein a loudness perception is reflected as a function of the brightness or contrast selected, wherein for good loudness perception characters of the text are caused to appear dark and sharply contrasted while, in analogous fashion, for poor loudness perception the characters are caused to appear light to completely faded with correspondingly diminished contrast. 3. The method as in claim 2, wherein the clarity of perception is reflected in the sharpness of the contours of the text, wherein for clear perception the characters are caused to appear sharply contoured while for poor perception the characters appear poorly defined or fuzzy. 4. The method as in claim 1, further comprising the complete removal of individual letters of the text when on the basis of frequency-allocated auditory values in the audiogram and due to an existing hearing level such letters are partly or entirely imperceptible. 5. The method as in claim 1, wherein the audiogram comprises a predefined number of frequency ranges, it includes intensity levels from 1 to 120 dB, that a loudness perception subdivided and distributed over the frequencies adds up to 100%, that auditory or hearing-loss values entered in dB are multiplied by the corresponding percentages in the respective frequency range, and that the values are totaled with the resulting sum constituting the basis for determining the brightness or contrast of text characters on a scale from 0 dB=dark or black or maximum contrast to 120 dB=white or without contrast. 6. The method as in claim 1, wherein the audiogram comprises a predefined number of frequency ranges, that it includes loudness levels from 0 to 120 dB, that a speech intelligibility perception subdivided and distributed over the frequencies adds up to 100%, that individual auditory or hearing-loss values are multiplied by the corresponding frequencies, that the values taken together result in a composite value for the speech intelligibility perception, and that the contour sharpness of text characters is determined accordingly on a scale from 0 dB=sharply contoured characters to 120 dB=fuzzy to unrecognizable characters. 7. An audiogram module for an Internet-accessed website, characterized in that it permits a visualization of an audiogram by the method per claim 1, wherein a person who has accessed the website can either select predefined reference audiograms or enter individualized audiogram hearing-level data for the purpose of editing and/or visualizing a predefined text or an individually entered text by said method. 8. The module as in claim 7, further comprising predefined audiograms for levels ranging... | Phonak AG | H04R029/00 G10L019/00 A61B013/00 A61B005/00
US07065394 | B2 | 2006-06-20 | Guide catheter | 2001-12-12 | A guide catheter has properties that promote enhanced visibility of the guide catheter using fluoroscopic or ultrasonic imaging techniques. The tip of the guide catheter may incorporate a material that is both fluoro and echo visible. The first material also may be provided in various amounts along the length of the guide catheter. In addition, the guide catheter may incorporate a reinforcing braid that includes one or more strands of a second material that is radio-opaque. As a result, the result guide catheter achieves improved radio-opacity and echogenicity. | 1. A guide catheter comprising: an elongated sheath having proximal end, a distal tip, and an inner channel to accommodate travel of a medical component;a first material in the distal tip, wherein the first material is radio-opaque and echogenic; anda second material in a wall of the sheath, wherein the second material is a reinforcing braid formed within the sheath, wherein the braid includes a plurality of strands and at least one strand of the plurality of strands forming the reinforcing braid includes a radio-opaque material, and wherein jet-milled tungsten carbide particles are distributed within the polymeric material of the first material between approximately 70 to 75 percent by weight and have an average diameter approximately less than or equal to 500 nanometers. | 1. A guide catheter comprising: an elongated sheath having proximal end, a distal tip, and an inner channel to accommodate travel of a medical component;a first material in the distal tip, wherein the first material is radio-opaque and echogenic; anda second material in a wall of the sheath, wherein the second material is a reinforcing braid formed within the sheath, wherein the braid includes a plurality of strands and at least one strand of the plurality of strands forming the reinforcing braid includes a radio-opaque material, and wherein jet-milled tungsten carbide particles are distributed within the polymeric material of the first material between approximately 70 to 75 percent by weight and have an average diameter approximately less than or equal to 500 nanometers. 2. The guide catheter of claim 1, wherein the sheath includes a plurality of sheath segments extending along the length of the guide catheter, and each of the plurality of sheath segments is formed of a polymeric material containing jet-milled tungsten carbide particles. 3. The guide catheter of claim 1, wherein the radio-opaque material comprises a material selected from the group consisting of platinum iridium, gold, tantalum, platinum, and tungsten carbide. 4. A guide catheter comprising: an elongated sheath having proximal end, a distal tip, and an inner lumen sized to accommodate travel of medical components; anda first material in the distal tip, wherein the first material is radio-opaque and echogenic, and includes jet-milled tungsten carbide particles having an average diameter of less than 500 nanometers, and wherein the tungsten carbide articles are distributed within the polymeric material in the amount of approximately 70 to 75 percent by weight. 5. The guide catheter of claim 4, wherein the distal tip is formed of a polymeric material, and the first material comprises jet-milled tungsten carbide particles distributed within the polymeric material. 6. The guide catheter of claim 5, wherein the sheath includes a number of sheath segments extending along the length of the guide catheter, and each of the sheath segments is formed of a polymeric material containing jet-milled tungsten carbide particles. 7. The guide catheter of claim 6, wherein the polymeric material comprises a polyether block amide. 8. A guide catheter comprising: an elongated sheath having proximal end and a distal tip;a first material forming the distal tip, wherein the first material is formed of a polymeric material and jet-milled tungsten carbide particles; anda second material forming a wall of the sheath, wherein the second material is formed of a polymeric material and jet-milled tungsten carbide particles are distribute within a first portion of the polymeric material of the second material in an amount between approximately 40 to 75 percent by weight, wherein the jet-milled tungsten carbide particles are distributed within the polymeric material of the first material between approxima... | Medtronic, Inc | A61B005/05
US06999811 | B2 | 2006-02-14 | Method and device for the registration of two 3D image data sets | 2002-07-17 | A method and a device provide for the registration of two 3D image data sets of an object to be examined. The object to be examined is provided with a plurality of markers In order to enable a registration to be carried out, the positions of the markers in the 3D image data sets are first determined in a co-ordinate system associated with the relevant 3D image data set. The distances between the markers and/or the angles formed between lines which intersect in a marker and extend through two further markers are then determined. Finally, a transformation rule for the transformation of one of the 3D image data sets to the co-ordinate system of the other 3D image data set is determined. | 1. A method for the registration of two 3D image data sets of an object to be imaged which is provided with a plurality of markers which are contained in the 3D image data sets, the method comprising the steps of: a) positioning the plurality of markers on the object to be imaged;b) determining the positions of the markers in the 3D image data sets in a co-ordinate system associated with the relevant 3D image data set;c) determining at least one of a distance between two markers and an angle between two straight lines which intersect in one of the plurality of markers, each of the lines respectively defined by extending through another different marker of the plurality of markers;d) determining at least one of a corresponding distance and a corresponding angle in both of the 3D image data sets;e) determining the corresponding markers in each of the 3D image datasets on the basis of the at least one corresponding distance and angle; andf) determining a transformation rule for the transformation of one of the 3D image data sets to the co-ordinate system of the other 3D image data set;wherein the markers are arranged such that the distances between any two markers and the angles between any two straight lines which intersect in a marker and extend through two further markers are each unique. | 1. A method for the registration of two 3D image data sets of an object to be imaged which is provided with a plurality of markers which are contained in the 3D image data sets, the method comprising the steps of: a) positioning the plurality of markers on the object to be imaged;b) determining the positions of the markers in the 3D image data sets in a co-ordinate system associated with the relevant 3D image data set;c) determining at least one of a distance between two markers and an angle between two straight lines which intersect in one of the plurality of markers, each of the lines respectively defined by extending through another different marker of the plurality of markers;d) determining at least one of a corresponding distance and a corresponding angle in both of the 3D image data sets;e) determining the corresponding markers in each of the 3D image datasets on the basis of the at least one corresponding distance and angle; andf) determining a transformation rule for the transformation of one of the 3D image data sets to the co-ordinate system of the other 3D image data set;wherein the markers are arranged such that the distances between any two markers and the angles between any two straight lines which intersect in a marker and extend through two further markers are each unique. 2. The method of claim 1 wherein at least three markers are provided on the object to be imaged and are contained in the 3D image data set, and a combination of at least three of the distances and angles are determined. 3. The method of claim 1 wherein the corresponding markers are iteratively determined from the corresponding determined distances. 4. The method of claim 1 wherein the corresponding markers are obtained directly at the point of intersection of the straight lines from the determined corresponding angles. 5. The method of claim 1 wherein the markers are formed by at least one of artificial markers provided on the object to be imaged and anatomic features. 6. The method of claim 1 including the step of subtraction angiography in order to register and transform two 3D image data sets to a common co-ordinate system prior to the formation of a subtraction image. 7. The method of claim 1 including the steps of obtaining a first partial 3D image data set and a second partial 3D image dataset, and combining the incomplete first and second partial 3D image data sets so as to form a combined 3D image data set. 8. A device for the registration of two 3D image data sets of an object to be imaged which is provided with a plurality of markers which are contained in the 3D image data sets, the device comprising: a) means for determining the positions of the markers positioned on the object to be imaged in the 3D image data sets in a co-ordinate system associated with the relevant 3D image data set;b) means for determining at least one of a distance between two markers and an angle between two straight lines which intersect in one of the plurality of marke... | Koninklijke Philips Electronics N.V. | A61B005/05
US07266406 | B2 | 2007-09-04 | Magnetic resonance apparatus | 2002-03-20 | In a magnetic resonance apparatus, the transmission coil for generating the radiofrequency excitation field is stationarily arranged. The geometry of the transmission coil is permanently prescribed. For acquiring the magnetic resonance signal, the reception arrangement includes at least one first local coil. This is secured to a movable carrier, so that it can be placed against the person from a remote position. The carrier is secured in a carrier mount arranged stationarily within the examination volume. | 1. A magnetic resonance apparatus comprising: an apparatus housing;a basic magnetic field system which generates a static basic magnetic field in an imaging volume encompassed by said apparatus housing;a gradient field system which generates gradient fields in said examination volume;a radio frequency system having a transmission coil which emits a radio frequency excitation field in said examination volume, configured to trigger nuclear magnetic resonance signals in a subject in said examination volume, and a reception coil arrangement, comprising a local coil, configured to receive said nuclear magnetic resonance signals, said transmission coil having a non-variable geometry;a moveable carrier to which said local coil is attached, and a carrier mount, stationarily and permanently disposed within said examination volume, to which said moveable carrier is mounted, said moveable carrier being moveable from a remote position to a deployed position that causes said local coil to be placed against a subject in said examination volume;a restoring spring connected between said moveable carrier and said carrier mount that exerts a restoring force on said moveable carrier mount urging said carrier and said local coil towards said remote position, said restoring force exerted by said restoring spring being greater than a weight of said moveable carrier;a releasable holder that acts on said movable carrier to hold said local coil in said deployed position and, when released, allows said spring to automatically restore said local coil to said remote position; andan actuatable drive connected to said moveable carrier that, when actuated, applies a force to said movable carrier that is greater than said restoring force to move said carrier between said remote position and said deployed position against said subject. | 1. A magnetic resonance apparatus comprising: an apparatus housing;a basic magnetic field system which generates a static basic magnetic field in an imaging volume encompassed by said apparatus housing;a gradient field system which generates gradient fields in said examination volume;a radio frequency system having a transmission coil which emits a radio frequency excitation field in said examination volume, configured to trigger nuclear magnetic resonance signals in a subject in said examination volume, and a reception coil arrangement, comprising a local coil, configured to receive said nuclear magnetic resonance signals, said transmission coil having a non-variable geometry;a moveable carrier to which said local coil is attached, and a carrier mount, stationarily and permanently disposed within said examination volume, to which said moveable carrier is mounted, said moveable carrier being moveable from a remote position to a deployed position that causes said local coil to be placed against a subject in said examination volume;a restoring spring connected between said moveable carrier and said carrier mount that exerts a restoring force on said moveable carrier mount urging said carrier and said local coil towards said remote position, said restoring force exerted by said restoring spring being greater than a weight of said moveable carrier;a releasable holder that acts on said movable carrier to hold said local coil in said deployed position and, when released, allows said spring to automatically restore said local coil to said remote position; andan actuatable drive connected to said moveable carrier that, when actuated, applies a force to said movable carrier that is greater than said restoring force to move said carrier between said remote position and said deployed position against said subject. 2. A magnetic resonance imaging apparatus as claimed in claim 1 wherein said local coil has a non-variable geometry. 3. A magnetic resonance imaging apparatus as claimed in claim 1 further comprising a speed damper disposed to interact with said moveable carrier when said moveable carrier is moved to said remote position. 4. A magnetic resonance imaging apparatus as claimed in claim 1 further comprising a impact absorber disposed to interact with said moveable carrier when said moveable carrier is moved to said remote position. 5. A magnetic resonance imaging apparatus as claimed in claim 1 wherein said drive has a force limiter which limits a force with which said local coil can be placed against said subject in said deployed position. 6. A magnetic resonance imaging apparatus as claimed in claim 1 wherein said drive has a force regulator which regulates a force with which said local coil can be placed against said subject in said deployed position. 7. A magnetic resonance imaging apparatus as claimed in claim 1 further comprising a scissors mechanism which connects said moveable carrier to said carrier mount, allowing said carrier to move ... | Siemens Aktiengesellschaft | A61B005/05
US07151959 | B2 | 2006-12-19 | Systems and methods for assessing and modifying an individual's physiological condition | 2002-05-21 | Systems and methods for assessing an individual's physiological condition are provided. Cycle and shape parameters are derived from a recorded time trace containing heart rate data collected while an individual performs a cyclic exercise routine. Individually tailored exercise regimens that are based on these parameters are generated and modified as desired. | 1. A computer-readable medium containing instructions for controlling a computing device to determine a capacity index based on an electronic file containing a time trace of a heart rate of an individual obtained during an exercise test that includes a plurality of test cycles and at least one maximum effort cycle, said instructions comprising: instructions to determine at least one maximum effort cycle parameter; andinstructions to determine said capacity index based on said at least one maximum effort cycle parameter. | 1. A computer-readable medium containing instructions for controlling a computing device to determine a capacity index based on an electronic file containing a time trace of a heart rate of an individual obtained during an exercise test that includes a plurality of test cycles and at least one maximum effort cycle, said instructions comprising: instructions to determine at least one maximum effort cycle parameter; andinstructions to determine said capacity index based on said at least one maximum effort cycle parameter. 2. The computer-readable medium of claim 1 wherein said maximum effort cycle parameter is selected from a group consisting of a first peak heart rate corresponding to a maximum heart rate monitored during said maximum effort cycle. 3. A computer-readable medium containing instructions for controlling a computing device to determine a capacity index based on an electronic file containing a time trace of a heart rate of an individual obtained during an exercise test that includes a plurality of test cycles and at least one maximum effort cycle, said instructions comprising: instructions to determine at least one maximum effort cycle parameter; andinstructions to determine said capacity index based on said at least one maximum effort cycle parameter, wherein said instructions further comprise instructions to determine another parameter for use in said capacity index, said another parameter being selected from a group consisting of 1) a second peak heart rate monitored during another of said plurality of test cycles that is not said maximum effort cycle, 2) an upward slope of said heart rate monitored during a part of a stress portion of at least one of said plurality of test cycles, and 3) a downward slope of said heart rate monitored during a part of a relaxation portion of the at least one of said plurality of test cycles. | Lifewaves International, Inc. | A61B005/04
US06999812 | B2 | 2006-02-14 | Arteriosclerosis detection system | 2002-05-03 | An arteriosclerosis detection system including an electrocardiographic signal detection device, an eyeground image detection device for detecting an eyeground image in synchronization with an electrocardiographic signal detected by the detection device. The arteriosclerosis detection system further includes an eyeground vein constriction detection device for detecting the constriction of an eyeground vein in the vicinity of a site at which the eyeground vein and an eyeground artery cross each other. The constriction is detected based on the detected eyeground image in synchronization with the electrocardiographic signal. The arteriosclerosis detection system detects the eyeground image by executing an algorithm of software, which provides an eyeground image synchronized with an electrocardiographic signal by obtaining a stationary eyeground image synchronized with an arbitrary electrocardiographic signal from an animated eyeground image. | 1. An arteriosclerosis detection system comprising: electrocardiographic signal detection means for detecting an electrocardiographic signal;eyeground image detection means for continuously or discontinuously detecting an eyeground image in synchronization with the detected electrocardiographic signal; andeyeground vein constriction detection means for detecting constriction of an eyeground vein by using a ratio of an outer diameter of the eyeground vein in a vicinity of a site at which the eyeground vein and an eyeground artery cross each other to a portion of the eyeground vein distant from the crossing site,wherein the system is operable to detect the constriction on the basis of the eyeground image detected in synchronization with the electrocardiographic signal. | 1. An arteriosclerosis detection system comprising: electrocardiographic signal detection means for detecting an electrocardiographic signal;eyeground image detection means for continuously or discontinuously detecting an eyeground image in synchronization with the detected electrocardiographic signal; andeyeground vein constriction detection means for detecting constriction of an eyeground vein by using a ratio of an outer diameter of the eyeground vein in a vicinity of a site at which the eyeground vein and an eyeground artery cross each other to a portion of the eyeground vein distant from the crossing site,wherein the system is operable to detect the constriction on the basis of the eyeground image detected in synchronization with the electrocardiographic signal. 2. The arteriosclerosis detection system according to claim 1, wherein the electrocardiographic signal comprises at least one of an electrocardiographic R wave and T wave. 3. The arteriosclerosis detection system according to claim 1, wherein arteriosclerosis is detected by employment, as an arteriosclerosis index, of a degree of the constriction of an eye ground vein in a vicinity of a crossing site of an eyeground artery and the eyeground vein, the constriction degree being obtained using the eyeground image detection means. 4. The arteriosclerosis detection system according to claim 1, wherein the system is operable to detect a degree of aging of an artery of the subject by correlating normalized data obtained by normalizing, in accordance with the ages of a plurality of subjects, a degree of the constriction of an eyeground vein in a vicinity of a crossing site of an eyeground artery and the eyeground vein, the constriction degree being obtained using the arteriosclerosis detection system, with the degree of the constriction of each of the subjects. 5. An arteriosclerosis detection system comprising: electrocardiographic signal detection means for detecting an electrocardiographic signal;eyeground image detection means for continuously or discontinuously detecting an eyeground image in synchronization with an electrocardiographic signal detected by the detection means; andeyeground vein constriction detection means for detecting the constriction of an eyeground vein by using a ratio of an outer diameter of the eyeground vein in the vicinity of a site at which the eyeground vein and an eyeground artery cross each other to a portion of a eyeground vein distant from the crossing site,wherein the system is operable to detect the constriction of the eyeground vein based on the eyeground image detected in synchronization with the electrocardiographic signal, andwherein the system is operable to detect the eyeground image using software which can provide an eyeground image synchronized with an electrocardiographic signal by obtaining, by use of a computer, a stationary eyeground image synchronized with an arbitrary electrocardiographic signal from an animated eyeground image. 6. Th... | BML, Inc. Reiji Kawada | A61B005/02 A61B005/04 G06K009/00
US07062314 | B2 | 2006-06-13 | Cardiac rhythm management device with triggered diagnostic mode | 2001-12-28 | A device and method in which spectral analysis of heart rate variability is performed in order to assess autonomic balance. When a threshold level is reached, a diagnostic mode is triggered in which the device performs computationally intensive data analysis to assess the probability of a pathological event such as an arrhythmia occurring. | 1. A cardiac rhythm management device, comprising: a sensing channel for sensing ventricular electrogram signal and generating a ventricular sense (R wave) when the electrogram signal exceeds a specified threshold value; circuitry for measuring and collecting time intervals between successive R waves and storing the collected intervals as a discrete RR interval signal; bandpass filters for filtering the RR interval signal into defined high and low frequency bands; a power detector for determining the signal power of the RR interval signal in each of the low and high frequency bands, referred to LF and HF, respectively; and, circuitry for computing an LF/HF ratio and triggering a diagnostic mode if the LF/HF ratio exceeds a predetermined ratio threshold value, wherein upon entering the diagnostic mode, the device activates processing circuitry that performs data analysis to assess the probability of the occurrence of a pathological event. | 1. A cardiac rhythm management device, comprising: a sensing channel for sensing ventricular electrogram signal and generating a ventricular sense (R wave) when the electrogram signal exceeds a specified threshold value; circuitry for measuring and collecting time intervals between successive R waves and storing the collected intervals as a discrete RR interval signal; bandpass filters for filtering the RR interval signal into defined high and low frequency bands; a power detector for determining the signal power of the RR interval signal in each of the low and high frequency bands, referred to LF and HF, respectively; and, circuitry for computing an LF/HF ratio and triggering a diagnostic mode if the LF/HF ratio exceeds a predetermined ratio threshold value, wherein upon entering the diagnostic mode, the device activates processing circuitry that performs data analysis to assess the probability of the occurrence of a pathological event. 2. The device of claim 1 further comprising circuitry for resampling the RR interval signal to equalize the time intervals between values of the RR interval signal. 3. The device of claim 1 further comprising circuitry for detecting ectopic ventricular beats and filtering the RR intervals before and after such beats to derive the RR interval signal. 4. The device of claim 1 further comprising: circuitry for detecting ectopic ventricular beats and computing an ectopic beat density as the ratio of ectopic to non-ectopic beats over a specified time period; and, wherein the diagnostic mode is triggered only when the ectopic beat density also exceeds a predetermined density threshold value. 5. The device of claim 4 wherein the ectopic beat density threshold is dependent upon previous measurements. 6. The device of claim 5 wherein the ectopic beat density threshold is set to approximately 20% above the daily average of the previous day plus one standard deviation. 7. The device of claim 4 further comprising circuitry for delivering an appropriate therapy when the LF/HF ratio and ectopic beat density exceed their respective predetermined threshold values. 8. The device of claim 4 further comprising circuitry for transmitting an alert signal to an external device when the LF/HF ratio and ectopic beat density exceed their respective predetermined threshold values. 9. The device of claim 1 wherein the low frequency band approximately 0.04-0.15 Hz and the high frequency band is approximately 0.15-0.40 Hz. 10. The device of claim 1 further comprising circuitry for averaging the LF/HF ratio for a specified averaging period. 11. The device of claim 1 wherein the predetermined ratio threshold value is determined by the device based upon previous measurements. 12. The device of claim 11 wherein the ratio threshold is set to approximately 50% of the maximum computed LF/HF ratio value during a previous day. 13. A method for operating a cardiac rhythm management device, comprising: sensing ventricular electrogram sign... | Cardiac Pacemakers, Inc. | A61B005/0468
US07212851 | B2 | 2007-05-01 | Microstructured arrays for cortex interaction and related methods of manufacture and use | 2002-10-24 | A brain implant system consistent with embodiments of the present invention includes an electrode array having a plurality of electrodes for sensing neuron signals. A method for manufacturing the electrode array includes machining a piece of an electrically conductive substance to create a plurality of electrodes extending from a base member. Each electrode also has a corresponding base section. A nonconductive layer is provided around at least a portion of the base section of each electrode to support the plurality of electrodes. The base section of the electrodes are then cut to separate the base member from the plurality of electrodes supported by the nonconductive support layer. The present invention also includes a complete brain implant system using the above electrode array. | 1. A brain implant system comprising: an electrode configured to be inserted in a brain and for sensing electrical signals generated by brain neurons;a flexible wiring circuit connected to the electrode and adapted to receive the neuron electrical signals sensed by the electrode; anda processing unit for receiving the neuron electrical signals from the flexible wiring circuit, wherein the processing unit further includes: a detection module for detecting the occurrence of a neuron spike in the received neuron electrical signals; anda transmitter for transmitting data reflecting the occurrence of each detected neuron spike;wherein the electrode is a part of an electrode array, the electrode array comprising: a nonconductive layer of the flexible wiring circuit;an array of electrodes, each electrode having a base section and a tip section, wherein the base section of each electrode is inserted into the nonconductive layer, such that the electrodes are held together by the nonconductive layer; andan electrical connection located on the base section of each electrode to communicate with the respective electrode;wherein the nonconductive layer rests on a platform portion of each electrode after the base section of each electrode is inserted into the nonconductive layer. | 1. A brain implant system comprising: an electrode configured to be inserted in a brain and for sensing electrical signals generated by brain neurons;a flexible wiring circuit connected to the electrode and adapted to receive the neuron electrical signals sensed by the electrode; anda processing unit for receiving the neuron electrical signals from the flexible wiring circuit, wherein the processing unit further includes: a detection module for detecting the occurrence of a neuron spike in the received neuron electrical signals; anda transmitter for transmitting data reflecting the occurrence of each detected neuron spike;wherein the electrode is a part of an electrode array, the electrode array comprising: a nonconductive layer of the flexible wiring circuit;an array of electrodes, each electrode having a base section and a tip section, wherein the base section of each electrode is inserted into the nonconductive layer, such that the electrodes are held together by the nonconductive layer; andan electrical connection located on the base section of each electrode to communicate with the respective electrode;wherein the nonconductive layer rests on a platform portion of each electrode after the base section of each electrode is inserted into the nonconductive layer. 2. The system of claim 1, wherein a plurality of conductors are supported by the nonconductive layer for connection to respective ones of the plurality of electrodes, and wherein the nonconductive layer supports the plurality of electrodes. 3. The system of claim 1, wherein the nonconductive layer comprises an epoxy. 4. The system of claim 1, wherein the nonconductive layer comprises glass. 5. The system of claim 1, wherein the nonconductive layer comprises a flexible material. 6. The system of claim 5, wherein the flexible material comprises at least one of polyimide, parylene, and silicone. 7. The system of claim 1, wherein the electrodes are arranged in a two-dimensional matrix pattern. 8. The system of claim 1, wherein the electrodes are arranged in a honeycomb-like hexagonal pattern. 9. The system of claim 1, wherein the distances between neighboring electrodes varies. 10. The system of claim 1, wherein the electrodes increase in length from one side of the array to another side of the array. 11. The system of claim 1, wherein the plurality of electrodes have varying lengths. 12. The system of claim 11, wherein a first electrode has a length different that than of its immediately neighboring electrodes. 13. The system of claim 11, wherein the lengths of the plurality of electrodes are random. 14. The system of claim 1, wherein the plurality of electrodes have varying widths. 15. The system of claim 14, wherein a first electrode has a width different than that of each of its immediately neighboring electrodes. 16. The system of claim 1, wherein the width of the electrode is enlarged at or near the platform portion. 17. The system of claim 1, wherein the electrodes m... | Brown University Research Foundation | A61B005/04
US07160250 | B2 | 2007-01-09 | Method and equipment for analyzing biological signals representing intracranial and blood pressure fluctuations | 2002-07-23 | A method of analyzing biological pressure signals in a patient including obtaining arterial blood pressure, cerebral blood flow rate and intracranial pressure signals from the patient, and analyzing frequencies of the signals in relation to a range of target frequencies corresponding to at least one of type-B slow waves, infra-B waves and ultra-B waves; and an apparatus for analyzing biological pressure signals in a patient including an acquisition module, located adjacent to the patient and having input channels that receive signals from an arterial blood pressure (ABP) sensor, a cerebral blood flow rate (CBF) sensor, and an intracranial pressure (ICP) sensor; an analysis module that performs frequency analysis of the signals from a target frequency selected from the group consisting of type B slow waves corresponding to a frequency between 8×10−3 hertz and 50×10−3 hertz, infra-B waves corresponding to a frequency lower than 8×10−3 hertz, and ultra-B waves corresponding to a frequency including between 50×10−3 hertz and 200×10−3 hertz; and an exploitation and display module optionally located adjacent the patient in the form of an equipment unit or a local work station connected to the analysis module by an internal network or be constituted by a remote working station connected to the analysis module by a communication network. | 1. A method of analyzing biological pressure signals in a patient comprising: obtaining arterial blood pressure, cerebral blood flow rate and intracranial pressure signals from the patient, andanalyzing frequencies of said signals in relation to a range of target frequencies corresponding to at least one of type-B slow waves, infra-B waves and ultra-B waves (UB). | 1. A method of analyzing biological pressure signals in a patient comprising: obtaining arterial blood pressure, cerebral blood flow rate and intracranial pressure signals from the patient, andanalyzing frequencies of said signals in relation to a range of target frequencies corresponding to at least one of type-B slow waves, infra-B waves and ultra-B waves (UB). 2. The method according to claim 1, wherein the infra-B waves correspond to a frequency lower than 8×10−3 hertz. 3. The method according to claim 1, wherein the ultra-B waves (UB) correspond to a frequency between 50×10−3 hertz and 200×10−3 hertz. 4. The method according to claim 1, wherein frequency analysis comprises a Fast Fourier Transform (FFT) of the signals. 5. The method according to claim 1, wherein frequency analysis comprises a wavelet analysis. 6. The method according to claim 1, further comprising determining temporal shift between slow waves of the signals corresponding to two distinct inputs. 7. The method according to claim 1, wherein the type-B target frequency corresponds to frequencies between 8×10−3 hertz and 50×10−3 hertz. 8. The method according to claim 1, further comprising resampling at a sampling frequency lower than a first frequency to recording time-stamped signals in a memory. 9. The method according to claim 1, further comprising signal processing by filtering in a rereading module. 10. Apparatus for analyzing biological pressure signals in a patient comprising: an acquisition module adapted to be located adjacent to the patient and having input channels that receive signals from an arterial blood pressure (ABP) sensor, a cerebral blood flow rate (CBF) sensor, and an intracranial pressure (ICP) sensor;an analysis module that performs frequency analysis of the signals from a target frequency selected from the group consisting of: type B slow waves corresponding to a frequency comprised between 8×10−3 hertz and 50×10−3 hertz, infra-B waves corresponding to a frequency lower than 8×10−3 hertz, and ultra-B waves corresponding to a frequency comprised between 50×10−3 hertz and 200×10−3 hertz; andan exploitation and display module adapted to be located adjacent to the patient in the form of an equipment unit or a local work station connected to the analysis module by an internal network or in the form of a remote work station connected to the analysis module by a communication network. | Universite D'Auvergne | A61B005/00 A61B005/02
US07153262 | B2 | 2006-12-26 | Wearable human physiological data sensors and reporting system therefor | 2002-12-06 | A sensor array and computing apparatus are located on the human body while maintaining said sensors and apparatus within a proximity zone of the body such that the mobility and flexibility of the body are not deleteriously affected by the presence of the apparatus. The system permits the dynamic monitoring of human physiological status data without substantial interference in human motion and flexibility. A processor is mounted within a pod location with or adjacent to a sensor pod location, or said processor may be electrically connected to said sensor. | 1. An apparatus for detecting and reporting data relating to at least one of a human physiological and environmental condition, said apparatus adapted to be placed within at least one defined, preselected space proximate to a human body such that said apparatus avoids interference with the motion and flexibility of said human body, said apparatus comprising: a rigid pod having an inner surface adapted to accept an exterior surface of said human body and a convex outer surface, said inner surface of said pod having a convex portion adapted to be received by a generally concave portion of said exterior surface of said human body, said inner surface further having an outer edge, at least a portion of said pod being radiused at said outer edge of said inner surface;at least one sensor detecting at least one of a physiological conditional data point and an environmental conditional data point, said at least one sensor being mounted to said pod;processing means in electrical communication with said at least one sensor for converting said at least one data point from said at least one sensor into at least one of human physiological and environmental status output data; andoutput means in electrical communication with said processing means for communicating said output data to a user. | 1. An apparatus for detecting and reporting data relating to at least one of a human physiological and environmental condition, said apparatus adapted to be placed within at least one defined, preselected space proximate to a human body such that said apparatus avoids interference with the motion and flexibility of said human body, said apparatus comprising: a rigid pod having an inner surface adapted to accept an exterior surface of said human body and a convex outer surface, said inner surface of said pod having a convex portion adapted to be received by a generally concave portion of said exterior surface of said human body, said inner surface further having an outer edge, at least a portion of said pod being radiused at said outer edge of said inner surface;at least one sensor detecting at least one of a physiological conditional data point and an environmental conditional data point, said at least one sensor being mounted to said pod;processing means in electrical communication with said at least one sensor for converting said at least one data point from said at least one sensor into at least one of human physiological and environmental status output data; andoutput means in electrical communication with said processing means for communicating said output data to a user. 2. An apparatus according to claim 1, further comprising a flexible section, said rigid pod being supported by said flexible section. 3. An apparatus according to claim 2, said at least one sensor being supported by said flexible section. 4. An apparatus according to claim 2, said at least one sensor comprising two or more sensors, a first one of said two or more sensors being supported by said flexible section and a second one of said two or more sensors being supported by said rigid pod. 5. An apparatus according to claim 1, further comprising wrapping means coupled to said flexible section for maintaining contact between said pod and said human body. 6. An apparatus according to claim 1, further comprising a wireless transmitting device for transmitting at least one of said at least one data point and said output data to a receiving and processing device. 7. An apparatus according to claim 6, wherein said processing means and said transmitting means are mounted in said pod. 8. An apparatus according to claim 1, said pod having a width dimension as measured between a first lateral side and a second lateral side of said pod, at least a portion of said first lateral side and said second lateral side each having a taper such that said width dimension gradually decreases in a direction from said inner surface to said outer surface. | BodyMedia, Inc. | A61B005/00
US07197353 | B2 | 2007-03-27 | Sensitivity encoding MRI acquisition method | 2001-09-14 | A sensitivity encoding method (SENSE) is used to acquire an MR image having a reduced field of view. The number of aliased replicates caused by surrounding object boundaries is calculated for each image pixel location to obtain optimal aliasing artifact suppression without reducing image SNR. | 1. A method for producing a prescribed image of an object having a field of view D with a magnetic resonance imaging (MRI) system, the steps comprising: A) performing a pulse sequence with the MRI system to acquire calibration data with each of N local coils positioned near the object;B) performing a prescribed imaging pulse sequence with the MRI system to acquire image data with each of the N local coils;C) reconstructing calibration images with the calibration data;D) acquiring a calibration object image with the MRI system which depicts the object boundary disposed outside the field of view of the prescribed image;E) calculate a coil sensitivity image for each local coil using the calibration images;F) calculating an aliasing replicate number A for each pixel location in the prescribed image using the field of view D and the location of the object boundary;G) forming a sensitivity matrix S from the coil sensitivity images and the calculated aliasing replicate numbers A;H) reconstructing images I from the image data acquired in step b); andI) producing the prescribed image by calculating a proton distribution image M using the sensitivity matrix S and the images I. | 1. A method for producing a prescribed image of an object having a field of view D with a magnetic resonance imaging (MRI) system, the steps comprising: A) performing a pulse sequence with the MRI system to acquire calibration data with each of N local coils positioned near the object;B) performing a prescribed imaging pulse sequence with the MRI system to acquire image data with each of the N local coils;C) reconstructing calibration images with the calibration data;D) acquiring a calibration object image with the MRI system which depicts the object boundary disposed outside the field of view of the prescribed image;E) calculate a coil sensitivity image for each local coil using the calibration images;F) calculating an aliasing replicate number A for each pixel location in the prescribed image using the field of view D and the location of the object boundary;G) forming a sensitivity matrix S from the coil sensitivity images and the calculated aliasing replicate numbers A;H) reconstructing images I from the image data acquired in step b); andI) producing the prescribed image by calculating a proton distribution image M using the sensitivity matrix S and the images I. 2. The method as recited in claim 1 in which step I) is performed using the formula: M=[(S*S)−1S*]I,where S* is the complex conjugate of the sensitivity matrix S. 3. The method as recited in claim 1 in which the pulse sequence used in step A) is different than the prescribed imaging pulse sequence used in step B). 4. The method as recited in claim 3 in which the calibration data acquired in step A) and the image data acquired in step B) are acquired from substantially the same imaging volume in the object and the calibration object image is acquired from a substantially larger imaging volume. 5. The method as recited in claim 1 in which step F) includes: counting the number of times the aligned boundary location can be moved an increment D along the phase encoding axis before it passes the pixel location. 6. The method as recited in claim 1 in which step D) includes: i) producing a calibration object image which depicts the entire object disposed within a plane containing the slice defined by the prescribed image;ii) locating the object boundary in the calibration object image; andiii) aligning pixels in the calibration object image with pixels in the prescribed image. | General Electric Company | A61B005/05
US07251525 | B2 | 2007-07-31 | Treatment of arrhythmias via inhibition of a multifunctional calcium/calmodulin-dependent protein kinase | 2003-01-28 | This invention provides a method for treating or preventing arrhythmias in a human subject comprising the administration of an effective amount of a calcium/calmodulin-dependent protein kinase inhibitor. Also provided are pharmaceutical compositions comprising a calcium/calmodulin-dependent protein kinase inhibitor and a pharmaceutically acceptable carrier and methods for identifying agents useful for the treatment of arrhythmias. | 1. A method for identifying agents that inhibit calcium/calmodulin-dependent kinase II (CaM kinase II) and are useful for treating arrhythmias, the method comprising: (a) contactingCaM kinase II with an agent;(b) assaying the activity of the contacted CaM kinase II in vitro; and if CaM kinase II activity is reduced, then(c) administering said agent to an isolated heart cell, an isolated heart, or a test mammal; and then(d) detecting the presence or absence of arrhythmia;wherein a reduction of arrhythmia indicates that said agent is an inhibitor of CaM kinase II that is useful for treating arrhythmias. | 1. A method for identifying agents that inhibit calcium/calmodulin-dependent kinase II (CaM kinase II) and are useful for treating arrhythmias, the method comprising: (a) contactingCaM kinase II with an agent;(b) assaying the activity of the contacted CaM kinase II in vitro; and if CaM kinase II activity is reduced, then(c) administering said agent to an isolated heart cell, an isolated heart, or a test mammal; and then(d) detecting the presence or absence of arrhythmia;wherein a reduction of arrhythmia indicates that said agent is an inhibitor of CaM kinase II that is useful for treating arrhythmias. 2. The method of claim 1, wherein the assaying step comprises adding autocamtide-3 to the CaM kinase II in vitro and measuring phosphorylation of said autocamtide-3 by CaM kinase II. 3. The method of claim 1, wherein arrhythmia is detected in the test mammal using an electrocardiogram (ECG). 4. The method of claim 1, wherein the test mammal is a rabbit. 5. The method of claim 1, wherein arrhythmia is detected by measuring monophasic action potentials. | The Board of Trustees of the Leland Stanford Jr. University | A61B005/0452 G01N033/53 C12N009/12
US07048691 | B2 | 2006-05-23 | Method and apparatus for assessing hemodynamic parameters within the circulatory system of a living subject | 2001-03-22 | An improved method and apparatus for non-invasively assessing one or more hemodynamic parameters associated with the circulatory system of a living organism. In one aspect, the invention comprises a method of measuring a hemodynamic parameter by measuring a non-calibrated value of the parameter non-invasively, and inducing a stress of the circulatory system while measuring a second parameter. The response of the circulatory system to the stress is determined directly from the subject, and a calibration function is derived from the response and applied to the non-calibrated measured value to produce a calibrated measure of the actual value of the hemodynamic parameter. Methods of using backscattered acoustic energy for determination of hemodynamic markers are also disclosed. | 1. A method of assessing at least one hemodynamic parameter from a blood vessel, comprising: transmitting acoustic energy into the blood vessel;determining the power of reflections of said acoustic energy as a function of the depth within said blood vessel;identifying at least one first artifact within said power, said at least one artifact being related to the lumen associated with said blood vessel;identifying at least one wall of said blood vessel based at least in part on said first artifact and said act of determining the power; anddetermining said at least one hemodynamic parameter based at least in part on said at least one wall and said at least one artifact. | 1. A method of assessing at least one hemodynamic parameter from a blood vessel, comprising: transmitting acoustic energy into the blood vessel;determining the power of reflections of said acoustic energy as a function of the depth within said blood vessel;identifying at least one first artifact within said power, said at least one artifact being related to the lumen associated with said blood vessel;identifying at least one wall of said blood vessel based at least in part on said first artifact and said act of determining the power; anddetermining said at least one hemodynamic parameter based at least in part on said at least one wall and said at least one artifact. 2. The method of claim 1, wherein said act of identifying at least one first artifact comprises identifying an energy minimum within said power. 3. The method of claim 2, wherein said act of identifying at least one wall comprises identifying at least one region within said power wherein the energy level exceeds a mean lumen value. 4. The method of claim 1, wherein said at least one hemodynamic parameter comprises blood vessel diameter. 5. The method of claim 1, wherein said at least one hemodynamic parameter comprises blood flow kinetic energy. 6. An apparatus for measuring hemodynamic properties within the blood vessel of a living subject comprising: a first transducer adapted to measure at least a first hemodynamic parameter associated with said blood vessel;an applanation device adapted to compress said blood vessel while measuring said first hemodynamic parameter;a second transducer adapted to transmit ultrasonic energy into said blood vessel and receive echoes therefrom; anda signal processor operatively connected to said second transducer and configured to detect the lumen within said blood vessel based on said echoes, and estimate the diameter of said blood vessel based thereon, said detection of said lumen occurring prior to any wall detection associated with said estimate of the diameter, said estimate of diameter of said blood vessel being used at least in part to calibrate said measurement of said first hemodynamic parameter. 7. The apparatus of claim 6, wherein said blood vessel comprises the radial artery of a human being. 8. The apparatus of claim 7, wherein said first transducer comprises a pressure transducer disposed in proximity to said radial artery, and said second transducer comprises an ultrasonic transducer also disposed in proximity to said radial artery. 9. The apparatus of claim 8, wherein said processor is adapted to analyze A-mode signals derived from said echoes. 10. The apparatus of claim 6, wherein said applanation device comprises at least said first transducer, at least a portion of said transducer being disposed so as to exert a compressive force on the tissue of said living subject proximate to said blood vessel. 11. A method of assessing the arterial blood pressure of a living subject, comprising: applanating said blood vessel at least p... | Tensys Medical, Inc. | A61B005/02
US07203534 | B2 | 2007-04-10 | Method of assisting orientation in a vascular system | 2002-12-18 | The invention relates to a method of forming a very accurate vascular map or road map to be associated with the current X-ray image during an invasive vascular intervention. The vascular map is derived from a four-dimensional model of the vascular tree, that is, from a spatially three-dimensional, time-dependent model, the time dependency relating to natural motions of the body such as notably the heartbeat and/or the respiration. In particular the rhythm of the heartbeat or the respiration can be determined by means of a sensor so as to be used for synchronized selection and display of the corresponding three-dimensional vascular model. | 1. A method of assisting the orientation in the vascular system of a body volume, the method comprising: (a) acquisition of three-dimensional models of the vascular system in the body volume in different phases of a natural motion of the body based on changes due to one or more internal biological functions;(b) sensing a current phase of natural motion of the body;(c) selecting one of the three-dimensional models of the vascular system that best fits the current phase of natural motion;(d) forming a graphic representation on the basis of the selected three-dimensional model and a current image of the body volume; and(e) repeating the selecting step and forming step for each of the current images. | 1. A method of assisting the orientation in the vascular system of a body volume, the method comprising: (a) acquisition of three-dimensional models of the vascular system in the body volume in different phases of a natural motion of the body based on changes due to one or more internal biological functions;(b) sensing a current phase of natural motion of the body;(c) selecting one of the three-dimensional models of the vascular system that best fits the current phase of natural motion;(d) forming a graphic representation on the basis of the selected three-dimensional model and a current image of the body volume; and(e) repeating the selecting step and forming step for each of the current images. 2. The method of claim 1 wherein a two-dimensional projection of a three-dimensional model is formed in the step d). 3. The method of claim 2 wherein the current image is a fluoroscopic image and at least one of the projection direction and the projection plane of the three-dimensional model corresponds to at least one of the projection direction and the projection plane of the fluoroscopic image. 4. The method of claim 3 wherein at least one of the projection direction and the projection plane of the current fluoroscopic image is determined by comparing the fluoroscopic image with images of the body volume which were used for deriving the three-dimensional models of the vascular system. 5. The method of claim 1 wherein the position of an object within an imaging region is determined from the current fluoroscopic image and that the corresponding position is highlighted in the rendition of the models of the vascular system. 6. The method of claim 5 wherein the object within the imaging region comprises at least one of an instrument, a catheter and a guide wire. 7. The method of claim 1 wherein the natural motion of the body is caused by at least one of heartbeat and respiration. 8. An X-ray apparatus for monitoring an invasive vascular operation comprising: an X-ray source and a detector for forming X-ray images of a predetermined body volume;a memory for storing models of the vascular system in the body volume in different phases of a natural motion of the body based on changes due to one or more internal biological functions;a sensor for a current phase of natural motion of the body;means for selecting the model of the vascular system which best fits the current phase of the natural motion;means for updating the selection of the model as the current phase of natural motion changes; anda data processing unit which is coupled to the sensor and to the memory and is arranged in such a manner that it forms a graphic representation on the basis of the selected model of the vascular system and the current X-ray image of the body volume. 9. The X-ray apparatus of claim 8 comprising a device for determining at least one of a current projection direction and projection plane of the X-ray image. 10. The apparatus of claim 8, wherein the internal biologica... | Koninklijke Philips Electronics N.V. | A61B005/05
US07192403 | B2 | 2007-03-20 | Methods, apparatus and articles-of-manufacture for noninvasive measurement and monitoring of peripheral blood flow, perfusion, cardiac output biophysic stress and cardiovascular condition | 2002-08-16 | The invention relates to methods, apparatus, articles-of-manufacture, and coded data signals for measuring cardiac output, limb blood flow, perfusion, blood pressure, artery elasticity, and cardiovascular deterioration and disease, including performing these measurements on a continuous heart beat-by-beat basis, for humans and animals. Unlike empirical methods of other noninvasive blood pressure concepts, the invention is grounded on scientifically appropriate hemodynamic principles that studies have validated as accurate, and is practical for wide clinical use. Devices constructed in accordance with the invention can be comfortably employed for numerous applications, including hospital monitoring, physician's office cardiovascular disease management and drug therapy monitoring, home monitoring, and athletic applications.;The invention may be implemented in a variety of single or multi-sensor embodiments, such as: invasive pressure cannula sensor systems; non invasive pressure transducer arrays and piezo or other strain sensing materials that are placed against the skin above arteries; “upstream” pulsing-sensors (that apply single or multi-frequency vibrations that are measured “downstream” from the first placement location); other types of plethysmographic sensors; sonic/ultrasonic/Doppler sensors; MRI blood spin magnetizer/sensors; oxygen sensors; and electrocardiographic sensors, etc. | 1. A noninvasive method of estimating blood flow, comprising: noninvasively obtaining a pulsatile signal indicative of pressure pulsations in a patient's artery;using said pulsatile signal to compute a longitudinal pressure gradient waveform;computing plural harmonic components of said pressure gradient waveform; and,using said plural harmonic components of said pressure gradient waveform to estimate plural harmonic blood flow components. | 1. A noninvasive method of estimating blood flow, comprising: noninvasively obtaining a pulsatile signal indicative of pressure pulsations in a patient's artery;using said pulsatile signal to compute a longitudinal pressure gradient waveform;computing plural harmonic components of said pressure gradient waveform; and,using said plural harmonic components of said pressure gradient waveform to estimate plural harmonic blood flow components. 2. A method, as defined in claim 1, wherein computing plural harmonic components of said pressure gradient waveform involves an FFT process. 3. A method, as defined in claim 2, wherein said pressure gradient waveform is computed, at least in part, by differencing successive samples of said pulsatile signal. 4. A method, as defined in claim 3, wherein said estimated harmonic blood flow components are non-linearly dependent on harmonic frequency-related variables. 5. A method, as defined in claim 3, wherein said estimated harmonic blood flow components are non-linearly dependent on arterial radius-related variables. 6. A method, as defined in claim 3, wherein said estimated harmonic blood flow components are non-linearly dependent on elasticity-related variables. 7. A method, as defined in claim 3, wherein said pulsatile signal is obtained from a pressure cuff. 8. A method, as defined in claim 3, wherein plural pulsatile signals are obtained from plural respective pressure cuffs. 9. A method, as defined in claim 3, wherein estimating plural harmonic blood flow components involves estimating at least three harmonic blood flow components. 10. A method, as defined in claim 3, where said plural harmonic blood flow components are computed at each heartbeat. 11. A method, as defined in claim 1, wherein computing plural harmonic components of said pressure gradient waveform involves a DFT process. 12. A method of estimating blood flow, comprising: receiving a pulsatile signal related to pressure pulsations in a patient's artery;computing a longitudinal pressure gradient waveform, at least in part from said pulsatile signal;computing plural harmonic components of said pressure gradient waveform; and,using at least two of said plural harmonic components of said pressure gradient waveform to estimate blood flow. 13. A method, as defined in claim 12, wherein computing plural harmonic components of said pressure gradient waveform involves an FFT process. 14. A method, as defined in claim 12, wherein said pressure gradient waveform is computed, at least in part, by differencing successive samples of said pulsatile signal. 15. A method, as defined in claim 12, wherein said pulsatile signal is received from a pressure cuff. 16. A method, as defined in claim 12, wherein plural pulsatile signals are received from plural respective pressure cuffs. 17. A method, as defined in claim 12, wherein estimating blood flow involves estimating at least two harmonic blood flow components. 18. A method, as defined in claim 12,... | - | A61B005/02
US07041295 | B2 | 2006-05-09 | Compounds for treatment of infectious and immune system disorders and methods for their use | 2002-03-14 | The present invention provides polypeptides comprising an immunogenic epitope of a M. vaccae protein, polynucleotides encoding such polypeptides, and fusion proteins comprising at least one such polypeptide, together with genetic constructs comprising at least one inventive polynucleotide. Compositions comprising such polypeptides, polynucleotides, fusion proteins and/or genetic constructs may be employed in the treatment of infectious diseases and immune disorders. | 1. An isolated polypeptide comprising an amino acid sequence having at least 95% identity to SEQ ID NO: 71, wherein the polypeptide possesses at least one property selected from the group consisting of: (i) an ability to stimulate proliferation of T cells from individuals exposed to Mycobacterium tuberculosis, and(ii) an ability to stimulate interferon-gamma secretion from T cells from individuals exposed to Mycobacterium tuberculosis. | 1. An isolated polypeptide comprising an amino acid sequence having at least 95% identity to SEQ ID NO: 71, wherein the polypeptide possesses at least one property selected from the group consisting of: (i) an ability to stimulate proliferation of T cells from individuals exposed to Mycobacterium tuberculosis, and(ii) an ability to stimulate interferon-gamma secretion from T cells from individuals exposed to Mycobacterium tuberculosis. 2. A composition comprising at least one polypeptide according to claim 1, and at least one component selected from the group consisting of: physiologically acceptable carriers and immunostimulants. 3. A fusion protein comprising at least one polypeptide according to claim 1. 4. A composition comprising at least one fusion protein according to claim 3 and at least one component selected from the group consisting of: physiologically acceptable carriers and immunostimulants. 5. The composition of claim 4, wherein the composition further comprises a flt3 ligand. | Genesis Research & Development Corporation LTD | A01N037/18 A61K038/00 A61K039/00 C07K016/00 A61B005/055
US07232460 | B2 | 2007-06-19 | Nanodevices, microdevices and sensors on in-vivo structures and method for the same | 2002-04-25 | An in-vivo method and apparatus is disclosed that comprises at least one sensor for determining changes in a human's an animal's body and reporting said changes outside the body. The sensor may be embedded in a sheath. The apparatus may be used to monitor chemical or physical changes in the body fluids. Alternatively, the apparatus may be used to monitor and regulate chemical or physical levels in humans and animals. | 1. An apparatus comprising: a sheath having an inner and outer surface, wherein the outer surface is adapted for contacting a wall of a body passage, and the inner surface forms a lumen for enabling flow of body fluids therethrough;at least two sensors operatively attached to the sheath, for determining changes in a diameter of the lumen, wherein the at least two sensors further are arranged for determining changes due to collapse of the sheath; anda means for reporting the changes without an intervening wire. | 1. An apparatus comprising: a sheath having an inner and outer surface, wherein the outer surface is adapted for contacting a wall of a body passage, and the inner surface forms a lumen for enabling flow of body fluids therethrough;at least two sensors operatively attached to the sheath, for determining changes in a diameter of the lumen, wherein the at least two sensors further are arranged for determining changes due to collapse of the sheath; anda means for reporting the changes without an intervening wire. | Xillus, Inc. | A61F002/06 A61B005/00
US07288075 | B2 | 2007-10-30 | Methods and devices utilizing rheological materials | 2002-06-27 | Impression devices are provided for making an impression of a desired feature. The impression devices include: a flexible bladder having a cavity formed therein, the flexible bladder being capable of conforming to the details of the desired feature; a rheological material, such as a magnetorheological fluid, disposed in the cavity of the flexible bladder; and an activator for applying an external stimulus to the rheological material to change the state of the rheological material and thereby capture an impression of the desired feature. | 1. An impression device for making an impression of a desired feature, the impression device comprising: a flexible bladder bite bag having a cavity formed therein, the flexible bladder being conformable to a feature in a person's mouth as the desired feature;a magneto-rheological material disposed in the cavity of the flexible bladder; anda magnet for generating a magnetic field in at least a portion of the magneto-rheological fluid and capturing an impression of the desired feature thereby, wherein the magnet further comprises a u-shaped magnet having first and second ends adapted to be disposed between teeth in the person's mouth and an inner surface of a cheek. | 1. An impression device for making an impression of a desired feature, the impression device comprising: a flexible bladder bite bag having a cavity formed therein, the flexible bladder being conformable to a feature in a person's mouth as the desired feature;a magneto-rheological material disposed in the cavity of the flexible bladder; anda magnet for generating a magnetic field in at least a portion of the magneto-rheological fluid and capturing an impression of the desired feature thereby, wherein the magnet further comprises a u-shaped magnet having first and second ends adapted to be disposed between teeth in the person's mouth and an inner surface of a cheek. 2. The impression device of claim 1, wherein the flexible bladder further comprises a pocket corresponding to each of the first and second ends of the magnet for retaining the same therein. | Ethicon, Inc. | A61B005/103 A61B005/117 A61B005/00 A61C019/00 A61C009/00 A61C005/04 A61C019/04
US07054678 | B2 | 2006-05-30 | Systems and methods for assessing and modifying an individual's physiological condition | 2002-05-21 | Systems and methods for assessing an individual's physiological condition are provided. Cycle and shape parameters are derived from a recorded time trace containing heart rate data collected while an individual performs a cyclic exercise routine. Individually tailored exercise regimens that are based on these parameters are generated and modified as desired. | 1. A method of determining a capacity index while an individual undergoes an exercise test, said test comprising a plurality of test cycles, wherein said plurality of test cycles comprises a maximum effort cycle during which said individual attempts to exercise up to said individual's maximum physical capability, said method comprising: monitoring a heart rate of said individual during said test;determining a first peak heart rate corresponding to a maximum heart rate monitored during said maximum effort cycle; andassigning to said capacity index a value based on at least said first peak heart rate, wherein each of said plurality of test cycles comprises a stress portion followed by a relaxation portion, said method further comprising: determining a slope selected from a group consisting of (1) an upward slope of said heart rate monitored during a part of said stress portion of at least a first of test cycles and (2) a downward slope of said heart rate monitored during a part of said relaxation portion of said at least a first of test cycles; andassigning to said capacity index a value based on at least one of said peak heart rate, said upward slope, and said downward slope. | 1. A method of determining a capacity index while an individual undergoes an exercise test, said test comprising a plurality of test cycles, wherein said plurality of test cycles comprises a maximum effort cycle during which said individual attempts to exercise up to said individual's maximum physical capability, said method comprising: monitoring a heart rate of said individual during said test;determining a first peak heart rate corresponding to a maximum heart rate monitored during said maximum effort cycle; andassigning to said capacity index a value based on at least said first peak heart rate, wherein each of said plurality of test cycles comprises a stress portion followed by a relaxation portion, said method further comprising: determining a slope selected from a group consisting of (1) an upward slope of said heart rate monitored during a part of said stress portion of at least a first of test cycles and (2) a downward slope of said heart rate monitored during a part of said relaxation portion of said at least a first of test cycles; andassigning to said capacity index a value based on at least one of said peak heart rate, said upward slope, and said downward slope. 2. The method of claim 1 further comprising: determining a second peak heart rate corresponding to a maximum heart rate monitored during another of said plurality of test cycles that is not said maximum effort cycle; andassigning to said capacity index a value based on said first and second peak heart rates. 3. The method of claim 1 wherein said assigning comprises assigning to said capacity index a weighted sum of at least two of said upward slope, said downward slope, and said peak heart rate. 4. The method of claim 1 wherein said assigning comprises assigning to said capacity index an averaged sum of cycle parameters, said cycle parameters including at least two of said upward slope, said downward slope, and said peak heart rate. 5. The method of claim 1 wherein each of said test cycles comprises: said stress portion over a first duration; andsaid relaxation portion over a second duration. 6. The method of claim 5 wherein at least one of said test cycles has a first duration between about 30 seconds and about 90 seconds. 7. The method of claim 6 wherein said first duration is about one minute. 8. The method of claim 1 wherein said plurality of test cycles comprises: at least one warm-up cycle; anda plurality of intermediate cycles that follow said warm-up cycle, most of said test cycles having a target heart rate that is higher than its preceding cycle, and wherein said maximum effort cycle is after said plurality of intermediate cycles. 9. The method of claim 8 wherein at least one of said test cycles has a target heart rate that is higher than its immediately preceding cycle by up to about 15 beats per minute. 10. The method of claim 1 further comprising transmitting heart rate data from a test location to an analysis location through an electronic network, wher... | Lifewaves International, Inc. | A61B005/04
US07285117 | B2 | 2007-10-23 | Medical device control systems | 2002-03-15 | A medical device for use by an operator to perform a medical procedure in a body includes a handle to be held by the operator, a shaft attached to the handle and an actuating assembly to control functions associated with the medical procedure. The shaft includes an operative distal portion to perform a medical procedure in the body. The medical device may be part of an ablation catheter system that is programmed to correlate one or more functions with the actuation of the actuating assembly. Multiple actuating assemblies may be provided. Each actuating assembly may be a button. Other types of actuating assemblies may also be used, such as switches or a trackball. The actuating assembly may also be provided on a sleeve that may be selectively attached to the handle of the catheter or the physician operating the catheter, for example. | 1. A medical device, comprising: a shaft having a distal end and a proximal end;an operative element carried by said distal shaft end;a handle carried by said proximal shaft end; andan actuating element associated with said handle and configured for initiating a medical function; anda user programmable processing device configured for allowing a user to define an association between said medical function and a specific manner in which said actuating element is actuated. | 1. A medical device, comprising: a shaft having a distal end and a proximal end;an operative element carried by said distal shaft end;a handle carried by said proximal shaft end; andan actuating element associated with said handle and configured for initiating a medical function; anda user programmable processing device configured for allowing a user to define an association between said medical function and a specific manner in which said actuating element is actuated. 2. The medical device of claim 1, wherein said actuating element is configured for initiating a plurality of different medical functions, and wherein said processing device is configured for allowing a user to define an association between said plurality of medical functions and a specific manner in which said actuating element is actuated. 3. The medical device of claim 2, wherein said processing device is configured for allowing said user to associate a defined sequence of said plurality of medical functions with said actuating element, such that said plurality of medical functions is initiated in said sequence in response to said specific actuating manner. 4. The medical device of claim 3, wherein said specific actuating manner comprises repeatedly actuating said actuating element. 5. The medical device of claim 1, wherein said processing device is configured for allowing said user to define said medical function. 6. The medical device of claim 1, wherein said processing device is configured for allowing said user to define said specific actuating manner. 7. The medical device of claim 6, wherein said processing device is configured for allowing said user to define a number of actuations of said actuating element within a predetermined period of time. 8. The medical device of claim 6, wherein said processing device is configured for allowing said user to define a period of time in which said actuating element is actuated. 9. The medical device of claim 1, wherein said operative element is configured for being used in a performance of said medical function. 10. The medical device of claim 1, wherein said shaft is an intravascular catheter shaft. 11. The medical device of claim 1, wherein said processing device is coupled to said handle via an external cable. 12. A medical device, comprising: a shaft having a distal end and a proximal end;at least one operative element carried by said distal shaft end;a handle carried by said proximal shaft end; andan actuating element associated with said handle and configured for initiating a plurality of different medical functions in a predetermined sequence, wherein at least two of said plurality of medical functions are respectively initiated in response to at least two actuations of said actuating element. 13. The medical device of claim 12, wherein: a first of said plurality of medical functions is configured for being initiated when pressure is applied to said actuating element;said first function is configured for being term... | Boston Scientific Scimed, Inc. | A61B018/14 A61B005/04
US07206623 | B2 | 2007-04-17 | Optical sampling interface system for in vivo measurement of tissue | 2002-06-12 | An optical sampling interface system minimizes and compensates error resulting from sampling variations and measurement site state fluctuations. Components include: An optical probe placement guide having an aperture wherein the optical probe is received, facilitates repeatable placement accuracy on surface of a tissue measurement site with minimal, repeatable disturbance to surface tissue. The aperture creates a tissue meniscus that minimizes interference due to surface irregularities and controls variation in tissue volume sampled; an occlusive element placed over the tissue meniscus isolates the meniscus from environmental fluctuations, stabilizing hydration at the site and thus, surface tension; an optical coupling medium eliminates air gaps between skin surface and optical probe; a bias correction element applies a bias correction to spectral measurements, and associated analyte measurements. When the guide is replaced, a new bias correction is determined for measurements done with the new placement. Separate components of system can be individually deployed. | 1. An optical sampling interface system comprising: an optical probe placement guide for controlling variations in sample volume;means for controlling surface reflection of an optical signal;an occlusion plug for stabilizing surface hydration of a measurement site, said occlusion plug adapted to fit into an aperture in said placement guide between measurements, wherein said occlusion plug is removably attached at an exterior surface of said guide;wherein sampling error is minimized, eliminated, reduced or compensated. | 1. An optical sampling interface system comprising: an optical probe placement guide for controlling variations in sample volume;means for controlling surface reflection of an optical signal;an occlusion plug for stabilizing surface hydration of a measurement site, said occlusion plug adapted to fit into an aperture in said placement guide between measurements, wherein said occlusion plug is removably attached at an exterior surface of said guide;wherein sampling error is minimized, eliminated, reduced or compensated. 2. The system of claim 1, wherein said placement guide comprises: a mount having a contact surface, at least a portion of said contact surface being in contact with a surface proximate said measurement site during use; andan aperture, defined by said mount; adapted to receive said optical probe, wherein an area defined by said aperture comprises said measurement site;wherein said guide is removably attached at said measurement site so that said guide can repeatably couple said optical probe to the measurement site. 3. The system of claim 2, wherein means for attaching said guide comprises any of: an adhesive;at least one strap;at least one armband; andat least one suction element. 4. The system of claim 2, wherein size and shape of said aperture are matched to those of said optical probe, said probe fitting snugly Into said aperture, so that a mechanical registration in an x-y plane of said probe relative to said measurement site is provided, said guide further comprising at least one mechanical stop, said stop preventing over-penetration of said probe toward said measurement site, and providing registration along a z-axis of said probe relative to said measurement site. 5. The system of claim 2, said guide further comprising optical registration means wherein said probe is registered relative to said measurement site along any of x-, y-, and z-axes. 6. The system of claim 1, said means for controlling surface reflection comprising an optical coupling medium. 7. The system of claim 6, wherein said optical coupling medium comprises an optical coupling fluid. 8. The system of claim 7, wherein said optical coupling fluid comprises: one or more perfluoro compounds, wherein a quantity of said optical coupling fluid is placed at an interface of an optical probe and a measurement site, so that said probe and said measurement site are tightly optically coupled. 9. The system of claim 1, wherein shape and dimensions of a portion of said occlusion plug adapted to be received by said aperture mimics shape and dimensions of said aperture. 10. The system of claim 1, further comprising means for stabilizing surface temperature. 11. The system of claim 1, wherein said measurement site comprises a tissue measurement site and wherein surface tissue at said measurement site comprises skin, skin surface comprising stratum corneum. 12. The system of claim 11, further comprising means for correcting measurement bias resulting from one or both... | Sensys Medical, Inc. | A61B005/00
US07066891 | B2 | 2006-06-27 | Method and apparatus for gauging severity of myocardial ischemic episodes | 2002-12-20 | An implantable medical device (IMD) for gauging the severity of ischemia in a patient's heart includes a data collection module configured to receive data about the patient's heart, a data processing module configured to process the data to identify an episode of ischemia in the patient and to determine if the episode is stable or unstable, and a reporting module configured to provide an alarm to the patient if the episode is unstable. A method executable by an implantable medical device to detect ischemia in a human heart suitably includes the steps of receiving data about the heart at the implantable medical device, processing the data within the implantable medical device to determine the severity of ischemia, and providing a response from the implantable medical device to the patient if ischemia is indicated. | 1. A method executable by an implantable medical device (IMD) to detect ischemia in a patient's heart, the method comprising the steps of: receiving data about at least one physiologic parameter related to the severity of an ischemic episode of a patient's heart via an IMD;processing the data within the IMD to determine the severity of the ischemic episode, wherein said severity is deemed higher according to preselected criteria;in the event that the severity of the ischemic episode exceeds a present value of a variable threshold, providing an alert signal from the IMD to the patient, wherein the alert signal is determined as a function of the severity of the ischemic episode; andmodulating a palliative ischemia therapy based on the relative severity of the ischemic episodewherein the processing step comprises the steps of determining whether the ischemic episode can be classified as a stable ischemic episode or an unstable ischemic episode. | 1. A method executable by an implantable medical device (IMD) to detect ischemia in a patient's heart, the method comprising the steps of: receiving data about at least one physiologic parameter related to the severity of an ischemic episode of a patient's heart via an IMD;processing the data within the IMD to determine the severity of the ischemic episode, wherein said severity is deemed higher according to preselected criteria;in the event that the severity of the ischemic episode exceeds a present value of a variable threshold, providing an alert signal from the IMD to the patient, wherein the alert signal is determined as a function of the severity of the ischemic episode; andmodulating a palliative ischemia therapy based on the relative severity of the ischemic episodewherein the processing step comprises the steps of determining whether the ischemic episode can be classified as a stable ischemic episode or an unstable ischemic episode. 2. The method of claim 1 wherein the receiving step comprises receiving electrocardiogram (ECG) data. 3. The method of claim 2 wherein the processing step comprises interpreting an ST segment of the ECG data. 4. The method of claim 3 wherein the processing step further comprises interpreting accelerometer data. 5. The method of claim 1 further comprising: gathering feedback data about the heart during a temporal interval subsequent to the time the data was received and while the alert signal is being provided, andadjusting the alert signal as a function of the feedback data. 6. The method of claim 1 wherein modulating the palliative ischemia therapy comprises providing a dose of an anti-ischemia medication dispensed from one of the IMD and an implantable drug-delivery device. 7. The method of claim 1 wherein the palliative therapy comprises providing an electrical signal in electrical communication with excitable tissue of the heart. 8. The method of claim 1 wherein the palliative therapy comprises a neurological stimualtion therpy. 9. The method of claim 1 further comprising the step of storing results data in a memory if the ischemic episode is classified as a stable ischemic episode. 10. A method of gauging the severity of ischemia in a patient's heart, the method comprising the steps of: gathering electogram (EGM) data from a heart via an implantable medical device (IMD);processing the EGM data to isolate an ST segment for the patient;determining if the ST segment is one of elevated from a baseline and depressed from a baseline, and in the event that the ST segment is deemed elevated then classifying the event as an acutely severe episode of cardiac ischemia and in the event that the ST segment is deemed depressed then classifying the event as a relatively benign ischemic episode;in the event that the ST segment is deemed depressed, then evaluating accelerometer data to determine is the ischemic episode can be further classified as one of a stable episode and an unstable episode; andissuing a res... | Medtronic, Inc. | A61B005/02
US07020510 | B2 | 2006-03-28 | Optimal view map V.0.01 | 2002-07-25 | A method is disclosed which allows for optimally viewing a portion of a patient vascular system to facilitate at least one of diagnosis and treatment the vascular system. The method includes acquiring a model of said vascular system based on image data generated by an imaging device, identifying a portion of interest of the vascular system including determining a central vessel axis for a vessel of interest, generating a foreshortening map of the portion of interest based on viewing angle of the imaging device position, with respect to the patient, generating an overlap map to determine an amount of overlap present for particular viewing angles based on imaging device position, and generating a feasibility map to model dimensions of the patient based on patient characteristics, imaging device configuration, foreshortening and overlap. | 1. A method for optimally viewing a portion of a patient vascular system to facilitate at least one of diagnosis and treatment of said vascular system, the method comprising the steps of: acquiring a model of said vascular system based on image data generated by an imaging device;identifying a portion of interest of said vascular system including determining a central vessel axis for a vessel of interest;generating a foreshortening map of the portion of interest based on viewing angle and imaging device position, with respect to the patient;generating an overlap map to determine an amount of overlap for particular viewing angles based on imaging device position; andgenerating a feasibility map to model dimensions of the patient based on size of the patient, imaging device configuration, foreshortening and overlap. | 1. A method for optimally viewing a portion of a patient vascular system to facilitate at least one of diagnosis and treatment of said vascular system, the method comprising the steps of: acquiring a model of said vascular system based on image data generated by an imaging device;identifying a portion of interest of said vascular system including determining a central vessel axis for a vessel of interest;generating a foreshortening map of the portion of interest based on viewing angle and imaging device position, with respect to the patient;generating an overlap map to determine an amount of overlap for particular viewing angles based on imaging device position; andgenerating a feasibility map to model dimensions of the patient based on size of the patient, imaging device configuration, foreshortening and overlap. 2. A method for optimally viewing a portion of a patient vascular system to facilitate at least one of diagnosis and treatment of said vascular system, the method comprising the steps of: acquiring a model of said vascular system based on image data generated by an imaging device;identifying a portion of interest of said vascular system including determining a central vessel axis for a vessel of interest;generating a foreshortening map of the portion of interest based on viewing angle and imaging device position, with respect to the patient;generating an overlap map to determine an amount of overlap for particular viewing angles based on imaging device position;generating a feasibility map to model dimensions of the patient based on patient characteristics, imaging device configuration, foreshortening and overlap; andgenerating a composite view of said segment of interest in accordance with said foreshortening, feasibility and overlap maps. 3. A method as set forth in claim 2, wherein said step of generating a composite view includes generating an optimal view map in accordance with said composite view. 4. A method as set forth in claim 1, wherein said segment of interest may be identified with at least one marker for viewability. 5. A method as set forth in claim 4, wherein an optimal view of said optimal view map is generated automatically. 6. The method of claim 1, wherein the step of generating a feasibility map serves to eliminate, based on said size and said configuration, gantry position determinations for said viewing that would theoretically be possible if said size and said configuration were not taken into account. 7. The method of claim 1, wherein the step of generating a feasibility map takes into account that some gantry positions are theoretically possible but infeasible in practice. | Koninklijke Philips Electronics, N.V. | A61B005/05
US07282028 | B2 | 2007-10-16 | Method and apparatus for measuring animal's condition by acquiring and analyzing its biological signals | 2003-01-10 | A method and apparatus for measuring the biological condition of an animal by acquiring and analyzing its biological signals are provided. The biological signals from skin temperature, a photoplenthysmogram (PPG), an electrocardiogram (ECG), electrodermal activity (EDA), an electromyogram (EMG), and an electrogastrogram (EGG) are detected using a biological signal detection unit which is attached to the animal's skin. Feature vectors, including the mean heart rate of the photoplenthysmogram and its standard deviation, the very low frequency, low frequency, and high frequency components of heart rate variability, the frequency and mean amplitude of skin conductance responses, and the mean and maximum skin temperatures, are extracted from the detected biological signals. The biological condition, including needs and emotions, of the animal as to whether or not the animal feels hunger or fear, how much the animal is stressed, or whether or not the animal needs to have a bowel movement, is determined using a pattern classifier which has learned reference vectors, which reflect the behaviors, needs, and emotions of different kinds of animals for various biological conditions and are stored in a predetermined database. Therefore, the biological condition of the animal can be determined through instrumental communication, not through human languages, and the breeding of pets can be efficiently managed. | 1. A method for determining the biological condition of a non-human animal by acquiring and analyzing a biological signal, the method comprising: (a) acquiring a predetermined biological signal using a biological signal detection unit which is adapted to be attached to the non-human animal;(b) extracting a predetermined feature vector from the acquired predetermined biological signal, wherein the feature vectors include any one or any combination of the mean heart rate of the photoplethysmogram and its standard deviation, the very low frequency, low frequency, and high frequency components of heart rate variability of the non-human animal; and(c) analyzing and determining the biological condition of the non-human animal, including emotions and needs of the non-human animal, from the predetermined feature vector. | 1. A method for determining the biological condition of a non-human animal by acquiring and analyzing a biological signal, the method comprising: (a) acquiring a predetermined biological signal using a biological signal detection unit which is adapted to be attached to the non-human animal;(b) extracting a predetermined feature vector from the acquired predetermined biological signal, wherein the feature vectors include any one or any combination of the mean heart rate of the photoplethysmogram and its standard deviation, the very low frequency, low frequency, and high frequency components of heart rate variability of the non-human animal; and(c) analyzing and determining the biological condition of the non-human animal, including emotions and needs of the non-human animal, from the predetermined feature vector. 2. A method for determining the biological condition of an animal by acquiring and analyzing biological signals, the method comprising: (a) acquiring biological signals from skin temperature, a photoplethysmogram (PPG), an electrocardiogram (ECG), electrodermal activity (EDA), an electromyogram (EMG), and an electrogastrogram (EGG) using a biological signal detection unit which is adapted to be attached to the animal's skin; (b) extracting feature vectors from the acquired biological signals, the feature vectors including any one or any combination of the mean heart rate of the photoplethysmogram and its standard deviation, the very low frequency, low frequency, and high frequency components of heart rate variability, the frequency and mean amplitude of skin conductance responses which are obtainable from the electrodermal activity, and the mean and maximum skin temperatures; and (c) analyzing and determining the biological condition, including needs and emotions, of the animal as to whether or not the animal feels hunger or fear, how much the animal is stressed, or whether or not the animal needs to have a bowel movement, using a support vector machine which has learned reference vectors, which reflect the behaviors, needs, and emotions of different kinds of animals for various biological conditions and are stored in a predetermined database. 3. The method of claim 2, wherein extracting the mean heart rate as one of the feature vectors in step (b) from the photoplethysmogram, which is acquired as a biological signal in step (a), comprises: (b1) low-pass filtering the photoplethysmogram to remove high frequency noise; (b2) median-filtering an output signal in step (b1), from which the high frequency noise is removed, to estimate a baseline variation to a very low frequency; (b3) subtracting an output signal in step (b2) from the output signal in step (b1) to obtain a waveform from which the baseline variation and direct current and high frequency components are removed; and (b4) matching-filtering an output signal in step (b3) to obtain the location of peaks of the photoplethysmogram and measure the heart rate and heart rate variability... | Samsung Electronics Co., Ltd. | A61B005/00
US07191134 | B2 | 2007-03-13 | Audio psychological stress indicator alteration method and apparatus | 2002-03-25 | A method of changing psychological stress indices by evaluating manifestations of physiological change in the human voice wherein the utterances of a subject under examination are formatted as electrical signals and processed to alter selected characteristics which have been found to change with psycho-physiological state changes, such that the resultant output data signals are perceptually unchanged yet display none of the undesired physiological response characteristics. Apparatus for performing changes of this type includes a data input port, means for spectral alteration and a data output port. | 1. A method of changing physiological manifestations of psychological stress in a human source audio comprising the steps of: receiving an oral utterance by a subject in the form of electrical signals; and identification of portions of said electrical signals which represent variations in the psychological stress of the voice of the subject; and transformation of said identified portions of audio data into like audio data where the signal segments are transformed via selective variation of frequency domain characteristics such that the presence of stress indicative characteristics are removed from said signal; and providing the electrical signals for use as an electrical representation of audio utterances. | 1. A method of changing physiological manifestations of psychological stress in a human source audio comprising the steps of: receiving an oral utterance by a subject in the form of electrical signals; and identification of portions of said electrical signals which represent variations in the psychological stress of the voice of the subject; and transformation of said identified portions of audio data into like audio data where the signal segments are transformed via selective variation of frequency domain characteristics such that the presence of stress indicative characteristics are removed from said signal; and providing the electrical signals for use as an electrical representation of audio utterances. 2. An apparatus for producing electrical output signals approximating electrical input signals both representative of oral utterances, which removes indications of involuntary physiological manifestations of psychological stress, the apparatus comprising: means for receiving an input signal representing audio utterances; and processing means for reduction of stress indicators present in said input signal based on no less than one involuntary stress indicating metric of the identified segments of said electrical input signal; and means for providing resultant data as signals representing audio utterances. 3. An apparatus according to claim 2 wherein said processing includes a means for quantifying no fewer than one stress indicia selected from the group consisting of infrasonic frequency, micro-tremor, tremolo and counter homeostasis oscillation perturbation. 4. An apparatus according to claim 2 wherein said processing includes a means for counteracting no fewer than one stress indicia selected from the group consisting of infrasonic frequency, micro-tremor, tremolo and counter homeostasis oscillation perturbation. 5. An apparatus as in any of claims 2–4 wherein a said processing means is a digital signal processor. 6. An apparatus as in any of claims 2–4 wherein a said processing means is a microprocessor. 7. An apparatus as in any of claims 2–4 wherein a processing means is a digital filtering. 8. An apparatus as in any of claims 2–4 wherein a processing means is digital. 9. An apparatus as in any of claims 2–4 wherein a processing means is analog. 10. An apparatus for producing output electrical signals that are representative of electrical input signals where the difference between said input signals and said output signals is the absence in the output signals of indicators of involuntary physiological manifestations of psychological stress, the apparatus comprising the combination of: means for receiving electrical signals representative of oral utterances; and processing means for analyzing said electrical signals to identify involuntary stress indicating characteristics thereof; and additional processing means for altering said stress indicating characteristics of said electrical signals; and means for outputting the altered elec... | - | G10L021/00 A61B005/16
US07087214 | B2 | 2006-08-08 | Water-soluble porphyrin platinum compounds with high tumor selectivity and their use for the treatment of benign and malignant tumor diseases | 2003-01-27 | The invention relates to novel, water-soluble porphyrin platinum compounds of the tetraarylporphyrin platinum derivatives type or of the hematoporphyrin platinum derivatives type with high tumor selectivity and their use for the treatment of benign and malignant tumor diseases. In particular, the compounds are suitable for photodynamic anti-tumor therapy. | 1. A porphyrin platinum derivative of the (a) tetraarylporphyrin platinum derivative typeaccording to formula Ior according to formula IIor (b) hematoporphyrin platinum derivative typeaccording to formula IV | 1. A porphyrin platinum derivative of the (a) tetraarylporphyrin platinum derivative typeaccording to formula Ior according to formula IIor (b) hematoporphyrin platinum derivative typeaccording to formula IV 2. A compound according to claim 1, being in the form of a racemate, enantiomer or diastereoisomer thereof. 3. A method for photodynamic treatment of a cancer, comprising administering a compound according to claim 1 to a patient in need thereof. 4. A method for photodynamic treatment of a tumor, comprising administering a compound according to claim 1 to a patient in need thereof. 5. The method of claim 4, further comprising irradiating with electromagnetic radiation having a wavelength of 600 to 730 nm. 6. A composition comprising one or more compounds according to claim 1 and a physiologically tolerable carrier, diluent or auxiliary. 7. A method for photodynamic treatment of a cancer, comprising administering a compound according to claim 2 to a patient in need thereof. 8. A method for photodynamic treatment of a tumor, comprising administering a compound according to claim 2 to a patient in need thereof. 9. The method of claim 8, further comprising irradiating with electromagnetic radiation having a wavelength of 600 to 730 nm. 10. A composition comprising one or more compounds according to claim 2 and a physiologically tolerable carrier, diluent or auxiliary. 11. The compound according to claim 1, selected from the group consisting of: Diammine[2-(4-[10,15,20-tris[4-(1,4,7-trioxaocryl)phenyl]porphyrin-5-yl } phenoxy)malonato]platinum (II);(±)-trans-1,2-Diaminocyclohexane[2-(4{ 10,15,20-tris[4-(1,4,7,10 -tetraoxaundecyl)phenyl]porphyrin-5-yl} phenoxy) malonato[platinum(II);(±)-trans-1,2-Diaminocyclohexane{7,12-bis[1-( 1,4,7,10,13,16 -hexaoxaheptadecyl) ethyl]-3,8,13,17-tetramethylporphyrin-2,18-dipropionato} platinum(II); and2,2′ -Bipyridyl (7,12-bis[1-(1 ,4,7-trioxaocryl)ethyl]3,8,13,17-tetramethylporphyrin-2,18 -dipropionato } platinum(II). 12. A method for photodynamic treatment of a cancer, comprising administering a compound according to claim 11 to a patient in need thereof. 13. A method for photodynamic treatment of a tumor, comprising administering a compound according to claim 11 to a patient in need thereof. 14. The method of claim 13, further comprising irradiating with electromagnetic radiation having a wavelength of 600 to 730 nm. 15. A composition comprising one or more compounds according to claim 11 and a physiologically tolerable carrier, diluent or auxiliary. 16. A porphyrin platinum derivative of the hematoporphyrin platinum derivative type according to formula III 17. The compound of claim 16, wherein n=1-8. 18. The compound of claim 16, which is diammine{7,12-bis[1-(1,4,7-trioxaoctyl)ethyl]-3,8,13,17-tetramethylporphyrin-2,18-dipropionato } platinum(II). 19. A method for phorodynamic treatment of a cancer, comprising administering a compound according to claim 16 in a patient in need thereof. 20. A ... | Zentaris GmbH | A61B005/055 A61B010/00 C07B047/00 C07B005/10 A61K031/555
US07090645 | B2 | 2006-08-15 | Biased vascular temperature measuring device | 2002-07-02 | The present invention provides a catheter apparatus for temperature measurement of vascular tissue, comprising a flexible body, at least one resiliently biased projection depended from the body, a sensor carried by the projection and an electrical carrier for transmitting data from the sensor to a remote device, wherein the projection, in a deployed configuration, adopts an arcuate shape along at least part of its length, wherein the gradient of the arcuate portion of the projection, with respect to the longitudinal axis of the catheter, increases as a function of distance along the projection from the end attached to the catheter body. | 1. A catheter apparatus for temperature measurement of vascular tissue, comprising a flexible body, at least one resiliently biased projection depended from the body, a sensor carried by the projection and an electrical carrier for transmitting data from the sensor to a remote device, wherein the projection, in a deployed configuration, adopts an arcuate shape along of a major portion of its length under its own resilient bias, wherein the gradient of the arcuate portion of the projection, with respect to the longitudinal axis of the catheter, increases as a function of distance along the projection from the end attached to the catheter body. | 1. A catheter apparatus for temperature measurement of vascular tissue, comprising a flexible body, at least one resiliently biased projection depended from the body, a sensor carried by the projection and an electrical carrier for transmitting data from the sensor to a remote device, wherein the projection, in a deployed configuration, adopts an arcuate shape along of a major portion of its length under its own resilient bias, wherein the gradient of the arcuate portion of the projection, with respect to the longitudinal axis of the catheter, increases as a function of distance along the projection from the end attached to the catheter body. 2. An apparatus according to claim 1, wherein the maximum gradient of the projection, with respect to the longitudinal axis of the catheter body is less than 60°. 3. An apparatus according to claim 1, wherein the sensor is a thermistor. 4. An apparatus according to claim 1, wherein the projection comprises a superelastic material. 5. An apparatus according to claim 1, further comprising a sheath adapted to substantially encompass the sensor in a retracted configuration. 6. An apparatus according to claim 5, wherein the sheath is retractable, thus enabling the projections to take up a deployed configuration. 7. An apparatus according to claim 1, further comprising a heat shrink wrapping applied around at least a portion of the length of the projection. 8. An apparatus according to claim 1, wherein the catheter body has a longitudinal axis and at least part of the projection is extensible radially from the longitudinal axis of the catheter body. 9. An apparatus according to claim 8, wherein the projection lies substantially parallel to the longitudinal axis of the catheter body and against the bias in a retracted configuration and are extended radially from the body, with the bias, to contact the vascular wall in a deployed configuration. 10. An apparatus according to claim 8, wherein the sensor is positioned circumferentially about a longitudinal axis of the catheter body. 11. An apparatus according to claim 1, wherein the ratio of the area of the cross-sectional profile of the apparatus in the retracted and deployed configurations is in the range 1.75:1–1.25:1. 12. A method of maintaining a thermal contact with vascular tissue, comprising deploying at least one resiliently biased projection bearing a sensor from the body of a catheter to a vascular wall, wherein the projection adopts an arcuate shape along a major portion of its length under its own resilient bias, wherein the gradient of the arcuate portion of the projection, with respect to the longitudinal axis of the catheter, increases as a function of distance along the projection from the end attached to the catheter body, and withdrawing the catheter along a length of vascular tissue. | NV Thermocore Medical Systems SA | A61B005/00
US07263395 | B2 | 2007-08-28 | Venous pulse oximetry | 2003-01-31 | A method of non-invasively measuring venous oxygen saturation, comprising: applying a pressure transducer at a first site on a body, applying a drive signal to the external pressure transducer at a predetermined frequency, to cause a series of pulsations of a predetermined magnitude in the venous blood volume in the vicinity of said first site, applying an oximeter device at a second site on the body, measuring output signals received from said oximeter device, said output signals containing a component representative of the modulation of venous blood volume due to said pulsations, deriving a measure of venous oxygen saturation from the frequency response of said output signals. | 1. A method of non-invasively measuring venous oxygen saturation, comprising applying a pressure transducer at a first site on a body having venous and arterial systems,applying a drive signal to the pressure transducer at a predetermined frequency, to cause a series of pulsations of a predetermined magnitude in the venous blood volume in the vicinity of said first site, the magnitude of the pulsations being adjusted such that the arterial system is minimally disturbed,applying an oximeter device at a second site on the body,measuring output signals received from said oximeter device, said output signals containing a component representative of the modulation of venous blood volume due to said pulsations,deriving a measure of venous oxygen saturation from the frequency response of said output signals. | 1. A method of non-invasively measuring venous oxygen saturation, comprising applying a pressure transducer at a first site on a body having venous and arterial systems,applying a drive signal to the pressure transducer at a predetermined frequency, to cause a series of pulsations of a predetermined magnitude in the venous blood volume in the vicinity of said first site, the magnitude of the pulsations being adjusted such that the arterial system is minimally disturbed,applying an oximeter device at a second site on the body,measuring output signals received from said oximeter device, said output signals containing a component representative of the modulation of venous blood volume due to said pulsations,deriving a measure of venous oxygen saturation from the frequency response of said output signals. 2. A method according to claim 1 wherein the relationship between the distance between the first and second sites on the one hand, and the magnitude of the pulsations on the other hand, is such that a multiplicative term in the frequency spectrum of the measured signals, indicative of a disturbance to the arterial system, is minimised. 3. A method according to claim 2 wherein the magnitude of the pulsations is determined by progressively increasing the pulsation magnitude and observing a frequency response of the output signals, the appearance of a multiplicative term in the frequency spectrum being indicative of the maximum permissible magnitude having been reached. 4. A method according to claim 3, wherein, upon reaching the point where said multiplicative term appears, the pulsation magnitude is subsequently reduced to a point where the multiplicative term becomes insignificant. 5. A method according to claim 1 wherein the frequency of the drive signal is chosen such that the pulsations are distinguishable from the heart rate. 6. A method according to claim 5 wherein the frequency of the drive signal is determined iteratively on the basis of a real time measurement of the heart rate. 7. A method according to claim 1 wherein the relationship between the frequency ωm of the pulsations caused by said drive signal and the heart rate ωHR is chosen to comply with the following conditions: ωm≠n(ωHR±ΔωHR)andωm±(ωHR±ΔωHR)≠n(ωHR±ΔωHR)where n>1. 8. A method according to claim 1 wherein the magnitude of the pulsations is adjusted such as to cause a variation of less than 1% in the DC level of the output signals received from the oximeter device. 9. A method according to claim 8 wherein the magnitude of the pulsations is adjusted such as to cause a variation of approximately 0.1% in the DC level of the output signals received from the oximeter device. 10. A method according to claim 1 wherein the oximeter is placed on a digit and a further pressure transducer is placed away from a distal end of a limb, and arranged to occlude the supply of blood to and from said limb, and the measurement is performed during the period of occlusion. 11. A method acco... | Loughborough University Enterprises Ltd. | A61B005/00
US06993373 | B2 | 2006-01-31 | Invasive device provided with a segmented electrical connection conductor | 2001-11-21 | An invasive device 17, such as a catheter, is connected to a supply and control unit via a connection conductor 21 for the transfer of power supply energy and/or LF signals to and from an electrical circuit 20 in the distal part 18 of the device. During operation the connection conductor traverses the homogeneous magnetic field and the RF field of the MRI apparatus; this may give rise to heating of the connection conductor and/or to a disturbance of said fields. In order to mitigate these drawbacks, the connection conductor in accordance with the invention is subdivided into mutually separated segments 22-i, each of which is much shorter (for example, 1/20) than the wavelength of the RE field in the apparatus; the segments are separated from one another by separating elements 23-i (self-inductances) that provide a high inductance for said RF frequencies and a low inductance for comparatively low frequencies. | 1. An invasive device (17) that is intended to be introduced into an object (7) that is to be imaged by means of an MDI apparatus, said invasive device comprising a distal end (18) and is provided with a housing (19) that extends to the distal end, with a circuit (20) that is arranged at the area of the distal end, and also with an electrical connection conductor (21) that is connected to the circuit and extends through the housing, the connection conductor (21) comprises mutually separated segments (22-i), each of which is shorter than a predetermined value, and that the separation between the segments is realized by way of frequency-dependent separating elements (23-i) constructed as cores (25, 26) wound on a carrier (24) in such a manner that magnetic fields generated by current in the cores (25, 26) compensate one another, said elements (23-i) constitute a conductor for LF currents and an isolator for RF alternating current. | 1. An invasive device (17) that is intended to be introduced into an object (7) that is to be imaged by means of an MDI apparatus, said invasive device comprising a distal end (18) and is provided with a housing (19) that extends to the distal end, with a circuit (20) that is arranged at the area of the distal end, and also with an electrical connection conductor (21) that is connected to the circuit and extends through the housing, the connection conductor (21) comprises mutually separated segments (22-i), each of which is shorter than a predetermined value, and that the separation between the segments is realized by way of frequency-dependent separating elements (23-i) constructed as cores (25, 26) wound on a carrier (24) in such a manner that magnetic fields generated by current in the cores (25, 26) compensate one another, said elements (23-i) constitute a conductor for LF currents and an isolator for RF alternating current. 2. An invasive device as claimed in claim 1, wherein the predetermined value for the length of the segments (22-i) is less than 120 cm. 3. An invasive device as claimed in claim 2, wherein the predetermined value for the length of the segments (22-i) is less than 24 cm. 4. An invasive device as claimed in claim 1, wherein the separating elements are formed by self-inductances that do not contain a ferromagnetic material. 5. An invasive device as claimed in claim 4, wherein the self-inductances are formed in that the cores (25, 26) of the connection conductor (21) are wound so as to form bifilar coils. 6. An invasive device as claimed in claim 4, wherein the self-inductances have a value of at the most 1 μH. 7. An invasive device as claimed in claim 1, wherein the segments are formed by mutually twisting the cores (25, 26). | Koninklijke Philips Electronics, N.V. | A61B005/055
US07147603 | B2 | 2006-12-12 | Apparatus for measurement of living body | 2002-07-18 | Disclosed is an apparatus for measurement of a living body, comprising: a personal data input unit; a basal metabolism input unit; a standard value storage unit; a comparison unit; and a display unit. According to the invention the personal data input unit enters at least an age of a person to be measured, the basal metabolism input unit enters a basal metabolism of the person to be measured, and the standard value storage unit stores standard values of basal metabolism according to the ages. Furthermore, the comparison unit compares the basal metabolism entered by the input unit with the standard values of basal metabolism according to the ages stored in said storage unit, and calculates a basal metabolic age indicating what age the basal metabolism of the person to be measured corresponds to. Then, the display unit displays the information about the calculated basal metabolic age. | 1. An apparatus for measurement of a living body, comprising: a personal data input unit;a basal metabolism determining unit;a data storage unit;a comparison unit; anda display unit, whereinsaid personal data input unit enters at least an age of a person to be measured,said basal metabolism determining unit determines a basal metabolism of the person to be measured,said data storage unit stores the determined basal metabolism,said comparison unit compares the latest determined basal metabolism with an already stored basal metabolism, andsaid display unit displays the result of the comparison. | 1. An apparatus for measurement of a living body, comprising: a personal data input unit;a basal metabolism determining unit;a data storage unit;a comparison unit; anda display unit, whereinsaid personal data input unit enters at least an age of a person to be measured,said basal metabolism determining unit determines a basal metabolism of the person to be measured,said data storage unit stores the determined basal metabolism,said comparison unit compares the latest determined basal metabolism with an already stored basal metabolism, andsaid display unit displays the result of the comparison. 2. An apparatus for measurement of a living body according to claim 1 in which the basal metabolism is calculated on the basis of fat-free mass of the person to be measured. 3. An apparatus for measurement of a living body according to claim 2 in which said fat-free mass is calculated on the basis of bioelectrical impedance. 4. An apparatus for measurement of a living body according to claim 2 in which said fat-free mass is calculated on the basis of the thickness of subcutaneous fat. 5. An apparatus for measurement of a living body according to claim 1 in which said basal metabolism is calculated according to analysis of expiration. 6. An apparatus for measurement of a living body according to claim 1, wherein said comparison unit determines whether the basal metabolism has been increased or decreased, and said display unit displays the determined increase or decrease of the basal metabolism. 7. An apparatus for measurement of a living body according to claim 6 in which said fat-free mass is calculated on the basis of bioelectrical impedance. 8. An apparatus for measurement of a living body according to claim 6 in which said fat-free mass is calculated on the basis of the thickness of subcutaneous fat. 9. An apparatus for measurement of a living body, comprising: a personal data input unit;a basal metabolism determininci unit;a data storage unit;a comparison unit; anda display unit; whereinsaid personal data input unit enters at least an age of a person to be measured;said basal metabolism determining unit determines a basal metabolism of the person to be measured, said basal metabolism being calculated on the basis of fat-free mass of the person to be measured;said fat-free mass is calculated on the basis of bioelectrical impedance;said data storage unit stores the determined basal metabolism;said comparison unit compares the latest determined basal metabolism with an already stored basal metabolism to determine whether the basal metabolism has been increased or decreased; andsaid display unit displays the determined increase or decrease of the basal metabolism. | Tanita Corporation | A61B005/053 A61B005/083
US07303741 | B2 | 2007-12-04 | Systems and methods for high-resolution in vivo imaging of biochemical activity in a living organism | 2002-09-23 | This invention relates to bifunctional detection agents useful for providing high-resolution, in vivo imaging of biochemical activity in a living organism. Methods of using these bifunctional detection agents may comprise administering them into a living organism, and then estimating the localization of the detection agent using one modality (i.e., MRI), while concurrently estimating the level of biological activity using a second modality (i.e., optical imaging). One of the bifunctional detection agents comprises a magnetic resonance component and an optical imaging component. The magnetic resonance component comprises a contrast agent that is always activated or “on”. The optical imaging component comprises an activatable contrast agent or dye that is activated or turned “on” only in the presence of a particular event. For example, the optical imaging component may be activated by a certain wavelength of light and (1) by the presence of a particular biochemical marker, (2) by enzyme cleavage, or (3) by a change in the temperature or pH of the surrounding medium. These bifunctional detection agents allow both anatomical and functional/metabolic information to be obtained simultaneously. | 1. A bifunctional detection agent comprising: | 1. A bifunctional detection agent comprising: 2. A bifunctional detection agent comprising: 3. A bifunctional detection agent comprising: | General Electric Company | A61B005/055
US07051120 | B2 | 2006-05-23 | Healthcare personal area identification network method and system | 2001-12-28 | A system, method, apparatus, and computer program code for delivering a treatment includes generating a personal area network associated with a patient, said personal area network transmitting a patient identifier associated with said patient, retrieving treatment data associated with said patient identifier, and operating a treatment device pursuant to said treatment data. | 1. A method for delivering a treatment, comprising: a patient wearing a first device which is not implanted in the patient;bringing a second device in contact with the patient, said second device to selectively apply a proposed treatment to the patient;said first device transmitting to said second device a patient identifier via an electrical current signal transmitted via the patient's body, said patient identifier including information for identifying the patient;said second device using said transmitted patient identifier to confirm that said proposed treatment is intended for the patient; andsaid second device applying said proposed treatment to the patient after confirming that said proposed treatment is intended for the patient. | 1. A method for delivering a treatment, comprising: a patient wearing a first device which is not implanted in the patient;bringing a second device in contact with the patient, said second device to selectively apply a proposed treatment to the patient;said first device transmitting to said second device a patient identifier via an electrical current signal transmitted via the patient's body, said patient identifier including information for identifying the patient;said second device using said transmitted patient identifier to confirm that said proposed treatment is intended for the patient; andsaid second device applying said proposed treatment to the patient after confirming that said proposed treatment is intended for the patient. 2. The method of claim 1, wherein said second device is not implanted in the patient. 3. A method for delivering a treatment, comprising: a patient wearing a first device and a third device where the first device is not implanted in the patient;said first device transmitting to said third device a patient identifier via an electrical current signal transmitted via the patient's body, said patient identifier including information for identifying the patient;bringing a second device in contact with the patient, said second device to selectively apply a proposed treatment to the patient;said second device transmitting to said third device a message indicative of said proposed treatment, said message transmitted to said third device from said second device via an electrical current signal transmitted via the patient's body;said third device using said transmitted patient identifier and said transmitted message to confirm that said proposed treatment is intended for the patient;said third device transmitting a return message to said second device; andsaid second device responding to said return message by applying said proposed treatment to the patient. 4. The method of claim 3, wherein said second device is not implanted in the patient. 5. The method of claim 4, wherein said third device is not implanted in the patient. 6. An apparatus for delivering a treatment, comprising: a first device worn by a patient and not implanted in the patient;a second device in contact with the patient, said second device to selectively apply a proposed treatment to the patient;said first device operative to transmit to said second device a patient identifier via an electrical current signal transmitted via the patient's body, said patient identifier including information for identifying the patient;said second device operative to: use said transmitted patient identifier to confirm that said proposed treatment is intended for the patient; andapply said proposed treatment to the patient after confirming that said proposed treatment is intended for the patient. 7. The apparatus of claim 6, wherein said second device is not implanted in the patient. 8. An apparatus for delivering a treatment, comprising: a first device worn by a patien... | International Business Machines Corporation | G06F015/16 A61B005/00
US07104961 | B2 | 2006-09-12 | Method and system for evaluating cardiac ischemia with an exercise protocol | 2002-12-03 | A method of assessing cardiac ischemia in a subject to provide a measure of cardiac or cardiovascular health in that subject is described herein. The method comprises the steps of: (a) collecting a first RR-interval data set from said subject during a stage of gradually increasing heart rate up to a predetermined heart rate of at least 130 beats per minute; (b) collecting a second RR-interval data set from said subject during a stage of gradually decreasing heart rate; (c) comparing said first RR-interval data set to said second RR-interval data set to determine the difference between said data sets; and (d) generating from said comparison of step (c) a measure of cardiac ischemia during stimulation in said subject, wherein a greater difference between said first and second data sets indicates greater cardiac ischemia and lesser cardiac or cardiovascular health in said subject. | 1. A method of assessing cardiac ischemia in a subject to provide a measure of cardiac or cardiovascular health in that subject, said method comprising the steps of: (a) collecting a first RR-interval data set from said subject during a stage of gradually increasing heart rate up to a predetermined threshold of at least 130 beats per minute;(b) collecting a second RR-interval data set from said subject during a stage of gradually decreasing heart rate;(c) comparing said first RR-interval data set to said second RR-interval data set to determine the difference between said data sets; and(d) generating from said comparison of step (c) a measure of cardiac ischemia in said subject, wherein a greater difference between said first and second data sets indicates greater cardiac ischemia and lesser cardiac or cardiovascular health in said subject. | 1. A method of assessing cardiac ischemia in a subject to provide a measure of cardiac or cardiovascular health in that subject, said method comprising the steps of: (a) collecting a first RR-interval data set from said subject during a stage of gradually increasing heart rate up to a predetermined threshold of at least 130 beats per minute;(b) collecting a second RR-interval data set from said subject during a stage of gradually decreasing heart rate;(c) comparing said first RR-interval data set to said second RR-interval data set to determine the difference between said data sets; and(d) generating from said comparison of step (c) a measure of cardiac ischemia in said subject, wherein a greater difference between said first and second data sets indicates greater cardiac ischemia and lesser cardiac or cardiovascular health in said subject. 2. The method according to claim 1, wherein the predetermined threshold rate is defined by the formula PT=(220±V bpm)−age;where PT is the predetermined threshold heart rate, V is 20, bpm is beats per minute, and age is the age of said subject. 3. The method according to claim 2, wherein said first and second QT-interval data sets are collected under quasi-stationary conditions. 4. The method according to claim 1, further comprising: (e) concurrently with steps (a) and (b), collecting a first and second QT-interval data set from said subject;(f) comparing said first QT-interval data set to said second QT-interval data set to determine the difference between said data sets; and(g) generating from said comparison step (f) a measure of cardiac ischemia during stimulation in said subject, wherein a greater difference between said first and second data sets indicates greater cardiac ischemia and lesser cardiac or cardiovascular health in said subject. 5. The method according to claim 1, further comprising performing the method on a subject not afflicted with coronary artery disease. 6. The method according to claim 1, further comprising performing the method on a subject afflicted with coronary artery disease. 7. The method according to claim 1, further comprising performing the method on a subject afflicted with moderate coronary artery disease. 8. The method according to claim 1, wherein said first and second RR-interval data sets are collected under quasi-stationary conditions. 9. The method according to claim 1, wherein said stage of gradually increasing heart rate and said stage of gradually decreasing heart rate are each at least 3 minutes in duration. 10. The method according to claim 1, wherein said stage of gradually increasing heart rate and said stage of gradually decreasing heart rate are together carried out for a total time of from 6 minutes to 40 minutes. 11. The method according to claim 1, wherein: both said stage of gradually increasing heart rate and said stage of gradually decreasing heart rate are carried out between a peak rate and a minimum rate; andsaid peak rates of both said stage... | Mediwave Star Technology, Inc. | A61B005/0452
US07074198 | B2 | 2006-07-11 | Methods for treatment and prevention of disorders resulting from hypertension of neck and shoulder muscles | 2002-10-09 | Methods for treatment and prevention of disorders resulting from hypertension of neck and shoulder muscles, in particular hypertension of the scalenius muscles, comprise administering a training regime to a subject suffering from the disorders, whereby the training regime comprises manipulating an object while maintaining a pinch-grip force exerted by the fingers on the object below a given threshold value. The object preferably has a pressure sensor that generates a signal when the pinch-grip force exceeds the threshold value. The object may be a writing instrument whereby the training may comprise writing exercises. | 1. A method for treating a patient having or at risk of having hypertension of the scalenius muscles, said method comprising: providing an object held by pinch-grip force by the patient, said object generating a signal when the pinch-grip force exceeds a first threshold value in range of 100 grams–250 grams; andproviding instructions to the patient to manipulate the object with a pinch-grip force below the threshold value, wherein compliance is monitored based on the signal. | 1. A method for treating a patient having or at risk of having hypertension of the scalenius muscles, said method comprising: providing an object held by pinch-grip force by the patient, said object generating a signal when the pinch-grip force exceeds a first threshold value in range of 100 grams–250 grams; andproviding instructions to the patient to manipulate the object with a pinch-grip force below the threshold value, wherein compliance is monitored based on the signal. 2. A method according to claim 1, wherein the threshold value is in range from 150 grams–175 grams. 3. A method according to claim 1, wherein the object generates a second signal when the time measured from the moment when the first threshold value is exceeded, exceeds a second threshold value. 4. A method according to claim 3, wherein the first threshold value is in the range from 150 grams–175 grams. 5. A method according to claim 4, wherein the second threshold value for the measured time is 2.5 seconds. 6. A method according to claim 3, wherein only the second signal is generated. 7. A method according to claim 3, wherein the first and second signals are generated by feeding to a collective signaling device on the object. 8. A method according to claim 1, wherein the patient has hypertension and an exercise is continued until the hypertension of the scalenius muscles is reduced or released. 9. A method according to claim 8, whereby the release or reduction of hypertension of the scalenius muscles is measured by an increase in blood flow velocity in the artery radialis or artery subclavia. 10. A method according to claim 9, whereby the blood flow velocity is measured by ultrasound or optical Doppler velocimetry. 11. A method according to claim 9, whereby an exercise is continued until the mean blood flow velocity in the artery radialis or artery subclavia is increased by at least 10%. 12. A method according to claim 8, whereby the release or reduction of hypertension of the scalenius muscles is measured by an increase in the rotation angle of the Cervical Range of Motion as determined by the muscular resistance. 13. A method according to claim 12, whereby an exercise is continued until the rotation angle of the Cervical Range of Motion is increased by at least 10 degrees. 14. A method according to claim 8, whereby an exercise is continued for at least 1 minute. 15. A method according to claim 1, whereby an training regime comprises performing an exercise at least once per day. 16. A method according to claim 1, whereby the object is a writing instrument, a mouse or similar computer input device, a stroke-making element for use in sport, an element for playing a musical instrument, or a tool. 17. A method according to claim 16, whereby the writing instrument is a sensorpen that generates at least one signal when at least one of a first, second or third threshold value is exceeded, whereby at least one of the first, second or third threshold value of the equi... | Tensor, B.V. Smartsense, B.V. | A61B005/103 A61B005/117
US07138104 | B2 | 2006-11-21 | Simultaneous imaging of cardiac perfusion and a vitronectin receptor targeted imaging agent | 2002-08-07 | The present invention describes a method of concurrent imaging in a mammal comprising: a) administering to said mammal a vitronectin receptor targeted imaging agent and a perfusion imaging agent; andb) concurrently detecting the vitronectin receptor targeted imaging agent bound at the vitronectin receptor and the perfusion imaging agent; andc) forming an image from the detection of said vitronectin targeted imaging agent and said perfusion imaging agent. | 1. A method of concurrent imaging of organ blood flow and vitronectin receptors sites in a mammal comprising: I) administering to said mammal a vitronectin receptor targeted imaging agent and a perfusion imaging agent; andII) concurrently detecting the vitronectin receptor targeted imaging agent bound at the vitronectin receptor and the perfusion imaging agent; andIII) forming an image from the detection of said vitronectin targeted imaging agent and said perfusion imaging agent;wherein the vitronectin receptor targeted imaging agent is a diagnostic metallopharmaceutical comprising 111In, and a compound comprising: a) a chelator which is DOTA (1,4,7,10-tetraaza-4,7,10-tris(carboxymethyl)cyclododecane); andb) a targeting moiety which is bound to the chelator and is (2S)-2-[({2,6-dimethyl-4-(4-butoxyoyl)phenyl}sulfonyl)amino]-3-({7-[(imidazol-2-ylamino)methyl)]-1-methyl-4-oxo(3-hydroquinolyl)}-carbonylamino) propanoic acid; andc) a linking group between the targeting moiety and the chelator, which is 2-(2-aminoethylcarbonyl)-2-(2-amino-3-sulfo-propanoylamino)-ethanesulfonic acid;and wherein the perfusion imaging agent is 99mTc sestamibi. | 1. A method of concurrent imaging of organ blood flow and vitronectin receptors sites in a mammal comprising: I) administering to said mammal a vitronectin receptor targeted imaging agent and a perfusion imaging agent; andII) concurrently detecting the vitronectin receptor targeted imaging agent bound at the vitronectin receptor and the perfusion imaging agent; andIII) forming an image from the detection of said vitronectin targeted imaging agent and said perfusion imaging agent;wherein the vitronectin receptor targeted imaging agent is a diagnostic metallopharmaceutical comprising 111In, and a compound comprising: a) a chelator which is DOTA (1,4,7,10-tetraaza-4,7,10-tris(carboxymethyl)cyclododecane); andb) a targeting moiety which is bound to the chelator and is (2S)-2-[({2,6-dimethyl-4-(4-butoxyoyl)phenyl}sulfonyl)amino]-3-({7-[(imidazol-2-ylamino)methyl)]-1-methyl-4-oxo(3-hydroquinolyl)}-carbonylamino) propanoic acid; andc) a linking group between the targeting moiety and the chelator, which is 2-(2-aminoethylcarbonyl)-2-(2-amino-3-sulfo-propanoylamino)-ethanesulfonic acid;and wherein the perfusion imaging agent is 99mTc sestamibi. 2. The method of claim 1, wherein the images are displayed side-by-side to facilitate interpretation of the localization of the vitronectin targeted imaging in the body, relative to the distribution of the perfusion agent in the body. 3. The method of claim 1, wherein the images are overlayed to facilitate interpretation of the localization of the vitronectin targeted imaging in the body, relative to the distribution of the perfusion agent in the body. 4. The method of claim 1, for use in concurrent imaging sites of angiogenesis and organ perfusion. 5. The method of claim 1, for use in diagnosing and localizing sites of angiogenesis and perfusion abnormalities. 6. The method of claim 1, for use in concurrent detection and localization of sites of endothelial damage and perfusion abnormalities. 7. The method of claim 1, for use in the concurrent detection and localization of sites of vulnerable plaque and perfusion abnormalities. 8. The method of claim 1, wherein administering the vitronectin targeted imaging agent and a perfusion imaging agent is concurrent. 9. The method of claim 1, wherein administering the vitronectin targeted imaging agent and a perfusion imaging agent is sequential. 10. The method of claim 1, wherein the vitronectin targeted imaging agent and a perfusion imaging agent are administered in a synergistically effective amount. 11. The method of claim 1, wherein the gamma-emission energies of the vitronectin targeted imaging agent and the perfusion imaging agent are spectrally separable by pulse-height analysis. 12. The method of claim 1, wherein the difference in gamma emission spectral energies of the vitronectin antagonist diagnostic metallopharmaceutical and the perfusion imaging agent is >10 Kev. 13. The method of claim 1, wherein the concurrent detection of the vitronectin recepto... | Bristol-Myers Squibb Company | A61B005/055
US07024242 | B2 | 2006-04-04 | Implantable cardiac stimulation device which recommends ablation therapy and method | 2002-05-01 | An implantable cardiac stimulation device provides a stimulation therapy or an atrial ablation recommendation responsive to detection of atrial flutter of a heart. The device includes a sensing circuit that senses electrical activity of a heart and generates electrical signals representing electrical activity of the heart, an arrhythmia detector that detects atrial flutter of the heart, and a data processor that measures cardiac data responsive to the electrical signals. A therapy control responsive to the detection of atrial flutter recommends atrial ablation of the heart when the cardiac data satisfies a predetermined criteria and atrial flutter suppression when the cardiac data fails to satisfy the predetermined criteria. The device may further include a pulse generator that provides stimulation therapy to at least one atrium to terminate the atrial flutter. | 1. An implantable cardiac device comprising: a sensing circuit that senses electrical activity of a heart and generates electrical signals representing the electrical activity of the heart;an arrhythmia detector that receives the electrical signals from the sensing circuit and that detects atrial flutter of the heart;a data processor that provides cardiac data responsive to receipt of the electrical signals and measures interatrial conduction delays for the provided cardiac data; andan analyzer, connected to the data processor, that analyzes the interatrial conduction delays responsive to detection of atrial flutter and that indicates atrial ablation when the interatrial conduction delays are stable. | 1. An implantable cardiac device comprising: a sensing circuit that senses electrical activity of a heart and generates electrical signals representing the electrical activity of the heart;an arrhythmia detector that receives the electrical signals from the sensing circuit and that detects atrial flutter of the heart;a data processor that provides cardiac data responsive to receipt of the electrical signals and measures interatrial conduction delays for the provided cardiac data; andan analyzer, connected to the data processor, that analyzes the interatrial conduction delays responsive to detection of atrial flutter and that indicates atrial ablation when the interatrial conduction delays are stable. 2. The device of claim 1 wherein the arrhythmia detector further detects atrial fibrillation episodes of the heart, and wherein the data processor provides a recording of detected atrial fibrillation episodes and wherein the analyzer indicates atrial ablation if there has been a lack of detected atrial fibrillation episodes. 3. The device of claim 1 wherein the analyzer indicates atrial ablation if the interatrial conduction delays are normal. 4. The device of claim 3 wherein the data processor further measures cardiac intervals corresponding to the interatrial conduction delays and the analyzer determines normalcy of the interatrial conduction delays by determining if the interatrial conduction delays are within specified ranges in relation to the corresponding cardiac intervals. 5. The device of claim 1 wherein the analyzer indicates atrial ablation if the interatrial conduction delays are non-lengthening as a function of time, and wherein the analyzer determines if the interatrial conduction delays are lengthening. 6. An implantable cardiac device comprising: sensing means for sensing electrical activity of a heart and generating electrical signals representing the electrical activity of the heart;arrhythmia detecting means for detecting atrial flutter of the heart;data processing means for providing cardiac data responsive to the electrical signals and for measuring an interatrial conduction delay for the cardiac data; andanalyzing means for analyzing the the interatrial conduction delays responsive to detection of atrial flutter to determine if the interatrial conduction delays are stable and indicating atrial ablation when the interatrial conduction delays are stable. 7. The device of claim 6 wherein the arrhythmia detecting means includes atrial fibrillation detecting means for detecting atrial fibrillation episodes of the heart, and wherein the data processing means includes means for recording detected atrial fibrillation episodes and the analyzing means indicates atrial ablation if there has been a lack of detected atrial fibrillation episodes. 8. The device of claim 6 wherein analyzing means further analyzes the interatrial conduction delays to determine if the interatrial conduction delays are normal. 9. The device of claim 8 wher... | Pacesetter, Inc. | A61B018/18 A61B005/0464
US07146277 | B2 | 2006-12-05 | NMR assisted design of high affinity ligands for structurally uncharacterized proteins | 2001-06-13 | Methods of using NMR spectroscopy for identifying ligands that bind to structurally uncharacterized proteins, and improving the binding affinity of ligands for biological targets are disclosed. One aspect of the method includes preparing first NMR spectra of a first complex comprising the biological target and a paramagnetically labeled derivative of a first ligand, preparing a second NMR spectra of a second complex comprising the biological target and a second ligand, and analyzing the spectra to determine whether the second ligand binds to the biological target within the paramagnetic zone of the paramagnetically labeled derivative. The first and second steps may be performed simultaneously, consecutively, or in any order. | 1. A method for improving the binding affinity of a ligand for a biological target comprising: a) preparing first NMR spectra of a first complex comprising the biological target and a paramagnetically labeled derivative of a first ligand to define a paramagnetic zone;b) preparing second NMR spectra of a second complex comprising the biological target and a second ligand;c) analyzing the spectra to determine whether the second ligand binds to the biological target within the paramagnetic zone of the paramagnetically labeled derivative;d) deducing a relative three-dimensional orientation of the first and second ligands when bound to the biological target;e) deducing a distance of separation of the first and second ligands when bound to the biological target; andf) selecting or preparing a compound that contains the first and second ligands substantially in the relative orientation and distance. | 1. A method for improving the binding affinity of a ligand for a biological target comprising: a) preparing first NMR spectra of a first complex comprising the biological target and a paramagnetically labeled derivative of a first ligand to define a paramagnetic zone;b) preparing second NMR spectra of a second complex comprising the biological target and a second ligand;c) analyzing the spectra to determine whether the second ligand binds to the biological target within the paramagnetic zone of the paramagnetically labeled derivative;d) deducing a relative three-dimensional orientation of the first and second ligands when bound to the biological target;e) deducing a distance of separation of the first and second ligands when bound to the biological target; andf) selecting or preparing a compound that contains the first and second ligands substantially in the relative orientation and distance. 2. The method of claim 1 wherein step (c) is performed by: a) identifying peaks on the first NMR spectra that are perturbed by the paramagnetic label; andb) determining whether the second ligand perturbs peaks on the second NMR spectra that are also perturbed by the paramagnetic label. 3. The method of claim 1 wherein the first complex further comprises the second ligand, and step (c) is performed by determining whether the paramagnetically labeled derivative of the first ligand perturbs peaks associated with the second ligand. 4. The method of claim 3 further comprising, before step (c), preparing third NMR spectra of a mixture of the paramagnetically labeled derivative of the first ligand and the second ligand in the absence of the biological target. 5. The method of claim 1 wherein the distance of separation is determined as a function of the loss of intensity for NMR resonances from the second ligand. 6. The method of claim 1 wherein the three dimensional orientation is deduced by producing a field ordered state in a medium comprising the biological target, the first ligand, and the second ligand, and analyzing dipolar couplings within the first and second ligands. 7. The method of claim 6 wherein the field ordered state is produced by an aqueous dispersion of lipid bicelles having complementary charges to the biological target. 8. The method of claim 6 wherein the field ordered state is produced by an aqueous dispersion of bacteriophage having a domain of the biological target in the outer coat. 9. The method of claim 1 wherein the paramagnetic label is a nitroxide or metal chelate. 10. The method of claim 1 wherein the first and second NMR spectra are two dimensional heteronuclear single quantum coherence spectra. 11. The method of claim 1 wherein the biological target is isotopically labeled. 12. The method of claim 1 wherein the biological target is a protein, and NMR resonances from the protein are not assigned to a sequence of the protein. 13. The method of claim 1 wherein the biological target is a protein, and the three dimensional co... | James H. Prestegard | G06F019/00 G01F001/716 A61B005/055
US07020511 | B2 | 2006-03-28 | Method for three dimensional cine EBA/CTA imaging | 2002-08-14 | Certain embodiments of the present invention provide a method for three-dimensional cine EBA/CTA imaging. The method includes monitoring a cardiac cycle of a patient and selecting a trigger point along the cardiac cycle. When the cardiac cycle of the patient reaches the trigger point, a computed tomography (CT) scan of the patient is initiated. At least two CT scans of the patient are performed during a time period over two or more cardiac cycles. A cine angiography image is constructed from the at least two CT scans. | 1. A method for obtaining cine angiography images with a computed tomography (CT) scanner, comprising: monitoring a cardiac cycle of a patient;selecting at least one independently configurable trigger array, said at least one trigger array including a plurality of trigger points;initiating at least one CT scan of said patient based on said at least one trigger array;selecting at least one additional independently configurable trigger array, said at least one trigger array including an additional plurality of trigger points;initiating at least one additional CT scan of said patient based on said at least one additional trigger array;performing at least two CT scans of said patient during a time period over two or more cardiac cycles; andconstructing a cine angiography loop from said at least two CT scans. | 1. A method for obtaining cine angiography images with a computed tomography (CT) scanner, comprising: monitoring a cardiac cycle of a patient;selecting at least one independently configurable trigger array, said at least one trigger array including a plurality of trigger points;initiating at least one CT scan of said patient based on said at least one trigger array;selecting at least one additional independently configurable trigger array, said at least one trigger array including an additional plurality of trigger points;initiating at least one additional CT scan of said patient based on said at least one additional trigger array;performing at least two CT scans of said patient during a time period over two or more cardiac cycles; andconstructing a cine angiography loop from said at least two CT scans. 2. The method of claim 1, wherein said performing step obtains said at least two CT scans during a single cardiac cycle. 3. The method of claim 1, wherein said performing step obtains said at least two CT scans consecutively and beginning at different points within said time period. 4. The method of claim 1, wherein said performing step performs a complete CT scan in no more than 100 milliseconds. 5. The method of claim 1, further comprising sweeping an electron beam across a target ring to perform said at least two CT scans. 6. The method of claim 1, utilizing an x-ray fan beam to obtain said at least two CT scans. 7. The method of claim 1, further comprising combining a series of three dimensional images into a three dimensional cine loop based on said at least two CT scans. 8. The method of claim 1, further comprising displaying a series of moving three dimensional images based on said at least two CT scans. 9. The method of claim 1, wherein said initiating step includes prospective gating based on said cardiac cycle of the patient. 10. The method of claim 1, further comprising moving the patient with respect to the CT scanner between or during CT scans. 11. The method of claim 1, further comprising moving the patient with respect to the CT scanner during each of said at least two CT scans to obtain spiral scan. 12. The method of claim 1, wherein said performing step obtains multiple parallel CT slices from separate parallel rows of detectors in the CT scanner. 13. The method of claim 1, wherein said performing step obtains one image for each CT scan. 14. A method for obtaining cine loop images with a computed tomography (CT) scanner, comprising: monitoring a cardiac cycle of a patient;selecting at least one independently configurable trigger array, said at least one trigger array including a plurality of trigger points associated with two or more events with respect to said patient;initiating at least one CT scan of said patient based on said at least one trigger array;selecting at least one additional independently configurable trigger array, said at least one trigger array including an additional plurality of trigger points ass... | GE Medical Systems Global Technology Company, LLC | A61B005/05
US07167740 | B2 | 2007-01-23 | Measuring method in magnetic resonance imaging device and magnetic resonance imaging device | 2001-12-03 | In a measuring space of the object (1), an essential measurement region (41, 61, 71, 81, 111, 131) having a center region of the measuring space and a plurality of peripheral measurement regions (42, 62, 63, 72, 82, 112, 132) which do not have any region overlapped with the essential measurement region are set. Then, the essential measurement region is combined with a selected peripheral portion of the plurality of the peripheral measurement regions to measure in a preceding manner a nuclear magnetic resonance signal from the object as data of the measuring space. The essential measurement region is combined with the peripheral measurement region of the plurality of the peripheral measurement regions which has not been selected in the preceding measuring step to measure a nuclear magnetic resonance signal from the object as data of the measuring space. Sequentially, the essential measurement region is combined with the peripheral measurement region having not been measured in a preceding measuring step to measure a nuclear magnetic resonance signal from the object. The data of the essential measurement region and the data of the peripheral measurement region thus measured is supplemented with the data of the peripheral measurement region measured in the preceding measuring step to generate data of the measuring space. | 1. In a measuring method in a magnetic resonance imaging apparatus which repeatedly measures an object to be examined by utilizing a nuclear magnetic resonance phenomenon, the method comprising: a region setting step of setting, in a measuring space of the object (1), an essential measurement region (41, 61, 71, 81, 111, 131) and a divided measurement region (42, 62, 63, 72, 82, 112, 132) including plural divided portions which do not have any region overlapped with the essential measurement region;a preceding measuring step of combining the essential measurement region with a selected divided portion of the plural divided portions of the divided measurement region to measure a nuclear magnetic resonance signal from the object as data of the measuring space;a data generating step of combining the essential measurement region with the divided portion having not been selected in the preceding measuring step to measure a nuclear magnetic resonance signal from the object as data of the measuring space; sequentially combining the essential measurement region with the divided portion having not been measured in a preceding measuring step to measure a nuclear magnetic resonance signal from the object;supplementing the data of the essential measurement region and the data of the divided portion thus measured with the data of the divided portion measured in the preceding measuring step to generate data of the measuring space; anda step of generating an image of the object based on the measuring space generated in the data generating step. | 1. In a measuring method in a magnetic resonance imaging apparatus which repeatedly measures an object to be examined by utilizing a nuclear magnetic resonance phenomenon, the method comprising: a region setting step of setting, in a measuring space of the object (1), an essential measurement region (41, 61, 71, 81, 111, 131) and a divided measurement region (42, 62, 63, 72, 82, 112, 132) including plural divided portions which do not have any region overlapped with the essential measurement region;a preceding measuring step of combining the essential measurement region with a selected divided portion of the plural divided portions of the divided measurement region to measure a nuclear magnetic resonance signal from the object as data of the measuring space;a data generating step of combining the essential measurement region with the divided portion having not been selected in the preceding measuring step to measure a nuclear magnetic resonance signal from the object as data of the measuring space; sequentially combining the essential measurement region with the divided portion having not been measured in a preceding measuring step to measure a nuclear magnetic resonance signal from the object;supplementing the data of the essential measurement region and the data of the divided portion thus measured with the data of the divided portion measured in the preceding measuring step to generate data of the measuring space; anda step of generating an image of the object based on the measuring space generated in the data generating step. 2. In a measuring method in a magnetic resonance imaging apparatus which repeatedly measures an object to be examined by utilizing a nuclear magnetic resonance phenomenon, the method comprising: a region setting step of setting, in a measuring space of the object (1), an essential measurement region (41, 61, 71, 81, 111, 131) having a center region of the measuring space and a plurality of peripheral measurement regions (42, 62, 63, 72, 82, 112, 132) which do not have any region overlapped with the essential measurement region;a preceding measuring step of combining the essential measurement region with a selected peripheral portion of the plurality of the peripheral measurement regions to measure a nuclear magnetic resonance signal from the object as data of the measuring space;a data generating step of combining the essential measurement region with the peripheral measurement region of the plurality of the peripheral measurement regions which has not been selected in the preceding measuring step to measure a nuclear magnetic resonance signal from the object as data of the measuring space; sequentially combining the essential measurement region with the peripheral measurement region having not been measured in a preceding measuring step to measure a nuclear magnetic resonance signal from the object; and supplementing the data of the essential measurement region and the data of the peripheral measurement region thus mea... | Hitachi Medical Corporation | A61B005/05
US07033327 | B2 | 2006-04-25 | Method of determining the long axis of an object | 2002-09-13 | The long axis of an object that is complex in shape, such as a tooth, may be accurately determined using digital data representing the shape of the object. A polyhedron is defined such that the center of the polyhedron coincides with a point representing the centroid of the object. Axis lines are then determined by establishing reference lines through each face of the polyhedron and the point representing the centroid of the object. The moment of the object is calculated about each axis line, and the axis line that corresponds to the smallest calculated moment is selected. Additional iterations of the method are then carried out by recursively subdividing the corresponding face into smaller sections until a desired angular accuracy is achieved. The method is particularly useful for orthodontic diagnosis and treatment. | 1. A method of determining the long axis of an object comprising: defining a polyhedron that surrounds a point representing the centroid of an object;defining, for at least some of the faces of the polyhedron, an axis line that passes through the face and also through the point representing the centroid of the object;calculating the moment of the object about each axis line;selecting the axis line that corresponds to the smallest calculated moment;controlling the object within a three-dimensional (3D) modeling environment based upon the determined long axis; andconstructing an article in accordance with the selected axis lines. | 1. A method of determining the long axis of an object comprising: defining a polyhedron that surrounds a point representing the centroid of an object;defining, for at least some of the faces of the polyhedron, an axis line that passes through the face and also through the point representing the centroid of the object;calculating the moment of the object about each axis line;selecting the axis line that corresponds to the smallest calculated moment;controlling the object within a three-dimensional (3D) modeling environment based upon the determined long axis; andconstructing an article in accordance with the selected axis lines. 2. A method of determining the long axis of an object according to claim 1 wherein the act of defining a polyhedron includes the act of defining a polyhedron having a center that coincides wit the point representing the centroid of the object. 3. A method of determining the long axis of an object according to claim 2 wherein the act of defining a polyhedron tat surrounds a point representing the centroid of the object includes the act of establishing a polyhedron that surrounds the point and the act of translating the data representing the object relative to the polyhedron so that the center of the polyhedron coincides with the point. 4. A method of determining the long axis of an object according to claim 1 wherein the act of defining, for at least some of the faces of the polyhedron, an axis line that passes through the face includes the act of defining an axis line that passes through the centroid of the face. 5. A method of determining the long axis of an object according to claim 1 wherein the act of defining a polyhedron comprises the act of defining an icosahedron. 6. A method of determining the long axis of an object according to claim 1 and including the acts of: defining a reference line that passes through a vertex of the face corresponding to the selected axis line;determining the angle between the reference line and the selected axis line; andcomparing the determined angle to a pre-selected value. 7. A method of determining the long axis of an object according to claim 6 and including the acts of: subdividing into projected sections the face corresponding to the selected axis line if the determined angle is greater than the pre-selected value;defining, for at least some of the subdivided sections of the face, a subdivided section axis line that passes through the subdivided section and also passes through the point representing the centroid of the object;calculating the moment of the object about each subdivided section axis line; andselecting the subdivided section axis line that corresponds to the smallest calculated moment. 8. A method of determining the long axis of an object according to claim 7 wherein the act of defining, for at least some of the subdivided sections of the face, a subdivided section axis line includes the act of defining a subdivided section axis line that passes through the cent... | 3M Innovative Properties Company | A61B005/103 A61B005/117
US07029450 | B2 | 2006-04-18 | Dilation catheter assembly and related methods | 2001-12-14 | The invention, in one embodiment, is directed to a catheter assembly adapted to enable an operator to incise a human ureter from within, with reduced risk of damaging a crossing blood vessel. The dilation catheter includes an elongated body and a dilatable bladder. The dilatable bladder is adapted to dilate in a radially outward direction from the elongated body. A thermally responsive indicator incorporated with at least a portion of the dilation member exhibits a state change in response to a change in temperature. A temperature change indicates the existence of a nearby crossing vessel. In a surgical method, an operator repeatedly positions and inflates the dilatable bladder in a human ureter proximal to a desired incision location to map the location of any crossing vessels. The operator then determines an inner wall incision location that avoids the detected crossing vessels. | 1. A catheter assembly comprising, an elongated body having a first end portion and a second end portion,a dilatable bladder incorporated with said elongated body between the first end portion of the elongated body and the second end portion of the elongated body, the dilatable bladder being adapted to dilate in a radially outward direction from said elongated body,a thermally responsive indicator incorporated with at least a portion of said dilatable bladder, said thermally responsive indicator being adapted to visually change in response to a change in temperature, anda surgical cutter adapted to enable an operator to cut mammal flesh at a location other than the location of said thermally responsive indicator. | 1. A catheter assembly comprising, an elongated body having a first end portion and a second end portion,a dilatable bladder incorporated with said elongated body between the first end portion of the elongated body and the second end portion of the elongated body, the dilatable bladder being adapted to dilate in a radially outward direction from said elongated body,a thermally responsive indicator incorporated with at least a portion of said dilatable bladder, said thermally responsive indicator being adapted to visually change in response to a change in temperature, anda surgical cutter adapted to enable an operator to cut mammal flesh at a location other than the location of said thermally responsive indicator. 2. The catheter assembly of claim 1, wherein said elongated body defines an internal lumen extending from a first end of said elongated body to a second end of said elongated body, said dilatable bladder is in fluid communication with said internal lumen, and inflates in response to a positive fluid pressure in said internal lumen. 3. The catheter of claim 1, wherein said catheter is adapted for insertion into a body of a mammal and said thermally responsive indicator detects a change in temperature when said thermally responsive indicator is located in proximity of a blood vessel. 4. The catheter of claim 1, wherein said catheter is adapted for insertion into a human body and said thermally responsive indicator indicates a change in temperature when a portion of said dilatable bladder is disposed in a ureter and brought into proximity of a femoral artery or vein. 5. The catheter assembly of claim 1, wherein said dilatable bladder has an inner surface and said thermally responsive indicator is disposed on at least a portion of said inner surface. 6. The catheter assembly of claim 1, wherein said dilatable bladder has an outer surface and said thermally responsive indicator is disposed on at least a portion of said outer surface. 7. The catheter assembly of claim 1, wherein said thermally responsive indicator includes a thermochromatic material. 8. The catheter assembly of claim 7, wherein said dilatable bladder is formed from a first material and said thermochromatic material is disposed within said first material. 9. The catheter assembly of claim 1, further comprising a detector element adapted for detecting the visual change of said thermally responsive indicator. 10. The catheter assembly of claim 9, wherein said elongated body defines a detector lumen extending between the first end portion of said elongated body and the second end portion of said elongated body, and being adapted for receiving said detector element. 11. The catheter assembly of claim 9, wherein said detector element is a fiber optic camera adapted to enable an operator to view the visual change of said thermally responsive indicator. 12. The catheter assembly of claim 1, wherein said dilatable bladder extends around a first portion of a periphery of said... | Boston Scientific Scimed, Inc. | A61B005/00
US07131950 | B1 | 2006-11-07 | System and method for noise reduction in thermodilution for cardiac measurement | 2002-09-23 | A system and method for artifact reduction in thermodilution for cardiac measurement is provided by both respiration artifacts and noise artifacts are reduced. The system and method detects the beginning of a respiration cycle and estimates a respiration artifact which is subtracted from the thermosignal recorded. | 1. A method for measuring average flow using thermodilution, the method comprising the steps of: generating a pseudo-random heat pulse sequence;measuring a thermal signal downstream of the generated psuedo-random heat pulse sequence;generating a maximum length sequence noise estimate and a maximum length sequence average;subtracting the maximum length sequence noise estimate from the maximum length sequence average to generate an average thermal pulse; andestimating the average flow from the average thermal pulse. | 1. A method for measuring average flow using thermodilution, the method comprising the steps of: generating a pseudo-random heat pulse sequence;measuring a thermal signal downstream of the generated psuedo-random heat pulse sequence;generating a maximum length sequence noise estimate and a maximum length sequence average;subtracting the maximum length sequence noise estimate from the maximum length sequence average to generate an average thermal pulse; andestimating the average flow from the average thermal pulse. 2. The method of claim 1 wherein the flow measured is cardiac output. 3. The method of claim 1 wherein the maximum length sequence noise estimate is calculated by summing the thermal signal over a moving window equal in length to the maximum length sequence. 4. A method for measuring average flow using thermodilution, the method comprising the steps of: generating a pseudo-random heat pulse sequence;measuring a thermal signal downstream of the generated heat pulse sequence;subtracting an estimated periodic noise artifact from the measured thermal signal;generating a maximum length sequence noise estimate and a maximum length sequence average;subtracting the maximum length sequence noise estimate from the maximum length sequence average to generate an average thermal pulse; andestimating the average flow from the average thermal pulse. 5. The method of claim 4 wherein the estimated periodic noise artifact results from respiration. 6. The method of claim 4 wherein the step of subtracting the respiration artifact further comprising the steps of: detecting a respiration cycle;aligning and subtract an estimated average artifact from the thermal signal;updating the estimated average artifact by averaging successive artifacts aligned by the respiration cycle detection. 7. The method of claim 6 wherein a respiration cycle is detected from pressure fluctuations in a vein. 8. The method of claim 7 wherein the vein is selected from a group consisting of a superior vena cava and an inferior vena cava. 9. The method of claim 7 wherein the step of detecting a respiration cycle further comprises the steps of: detecting pressure fluctuations in a vein to generate a pressure fluctuation signal;filtering the pressure fluctuation signal to reduce cardiac related pressure fluctuations; anddetecting when the pressure fluctuation signal crosses a threshold based on previous signal maxima. 10. The method of claim 6 wherein the reference point of a respiration cycle is detected from monitoring nasal air flow. 11. The method of claim 6 wherein a respiration cycle is detected from monitoring chest expansion and contraction. 12. The method of claim 6 wherein a respiration cycle is determined from the output of a ventilator. 13. The method of claim 4 wherein the maximum length noise sequence estimate is calculated by summing the thermal signal over a moving window the equal in length to the maximum length sequence. 14. A method for measuring average f... | E.P. Limited | A61B005/02 A61B005/08
US07139600 | B2 | 2006-11-21 | Biophotometer | 2002-01-18 | The present invention provides an optical measurement device of biological tissue having a probe attachable to a moving part of a living body with a high curvature in biophotometry. A member having a high flexibility and a high friction coefficient is disposed on a surface for contact with the living body. In addition, a device for light irradiation and a device for collection of light are fixed with an elastic material for uniform pressure application. The resulting double structure or double layer has solved a problem. | 1. An optical measurement device of biological tissue for measuring a substance of metabolism within a subject to be measured by using a probe comprising: a device for light irradiation for irradiating the subject with light via an optical wave-guide; anda device for collection of light for collecting light irradiated from the device for light irradiation and propagated through an inside of the subject via another optical wave-guide,said probe having a plurality of guides, each of said guides fixing a respective one of said optical wave-guides, and having a portion for contact with said subject being composed of at least one member with a surface structure, said portion for contact with said subject being made of flexible material, said guides being fixed to said portion for contact with said subject so that the respective ends of said optical wave-guides contact with a surface of said subject when said device is worn, and having a holding portion which connects and holds each said optical wave-guide. | 1. An optical measurement device of biological tissue for measuring a substance of metabolism within a subject to be measured by using a probe comprising: a device for light irradiation for irradiating the subject with light via an optical wave-guide; anda device for collection of light for collecting light irradiated from the device for light irradiation and propagated through an inside of the subject via another optical wave-guide,said probe having a plurality of guides, each of said guides fixing a respective one of said optical wave-guides, and having a portion for contact with said subject being composed of at least one member with a surface structure, said portion for contact with said subject being made of flexible material, said guides being fixed to said portion for contact with said subject so that the respective ends of said optical wave-guides contact with a surface of said subject when said device is worn, and having a holding portion which connects and holds each said optical wave-guide. 2. The optical measurement device of biological tissue of claim 1, wherein respective end portions of the optical wave-guides of said devices for light irradiation and for collection of light are supported to have a distance between the optical wave-guides in a specified permissible range at a surface for contact with said subject. 3. The optical measurement device of biological tissue of claim 1, wherein said portion for contact with said subject is composed of a material having a high flexibility and a high friction coefficient. 4. The optical measurement device of biological tissue of claim 1, wherein said portion for contact with said subject is composed of a material which does not permeate or reflect a wavelength of said light. 5. The optical measurement device of biological tissue of claim 1, wherein a region of said member with the surface structure for contact with said subject which is surrounded by said optical wave-guides is provided with an aperture. 6. The optical measurement device of biological tissue of claim 1, wherein an end portion of said device for light irradiation is covered with a filter for light attenuation. 7. The optical measurement device of biological tissue of claim 1, wherein each of said devices for light irradiation and for collection of light has four or more optical wave-guides. 8. The optical measurement device of biological tissue of claim 1, wherein an end portion of each of said devices for light irradiation and for collection of light is vertically movable above a surface of the subject. 9. The optical measurement device of biological tissue of claim 1, wherein said devices for light irradiation and for collection of light are alternately positioned at points of a square lattice configuration or a rhombic lattice configuration. 10. The optical measurement device of claim 1, further comprising a plurality of members connected with each other so as to mount said probe on said subject. | Hitachi, Ltd. Hitachi Medical Corporation | A61B005/00
US07102531 | B2 | 2006-09-05 | Remote controller for television/video recorder having a function of measuring body fat and body fat meter | 2002-10-18 | A remote controller for a television/video recorder having a function of measuring a body fat includes a display section for displaying a measurement of the body fat and a message questioning whether or not a user to be examined carries a specific medical device such as a pacemaker, a reply entry section for entering a reply to the message, and a control section for controlling the action of a body fat meter. The control section directs the display section to display the message prior to the measurement of the body fat. Then, when receiving from the reply entry section a reply indicating that the user carries a pacemaker, the control section inhibits the measurement of the body fat and directs the display section to display a message of the inhibition. When receiving from the reply entry section a reply indicating that the user carries no pacemaker, the control section enables the measurement of the body fat of the user through registering the physical data including the weight and the height of the user. This safety feature can thus protect and hold the user who carries a specific medical device such as a pacemaker in the safety. | 1. A remote controller for a television/video recorder having a function of measuring a body fat which is capable of selecting a desired receiving/recording channel of the television broadcasting, comprising: an electrode section including a pair of current electrodes which feed a current across the body of a user to be examined for the body fat and a pair of measurement electrodes which detect an electric signal generated by the current being fed across the body of the user;an impedance measuring circuit which feeds the current between the current electrodes of the electrode section and measures an impedance across the body of the user to determine the body fat of the user from the electric signal detected by the measurement electrodes;a physical data setting section which is operated by the user for registering physical data including the weight, the height, the sex, and the age of the user;a body fat calculator which calculates the body fat of the user from the impedance across the body of the user measured by the impedance measuring circuit and the physical data of the user registered by the physical data setting section;a display which displays a resultant measurement of the body fat calculated by the body fat calculator and a message questioning whether or not the user carries a medical device including a pacemaker;a reply entry section operated by the user for entering a reply to the message; anda control section which directs the display to display the message prior to the measurement of the body fat of the user, inhibits the measurement of the body fat when a reply is received from the reply entry section indicating that the user carries a specific medical device including a pacemaker, and enables the measurement of the body fat when a reply is received from the reply entry section indicating that the user carries non of a specific medical device including a pacemaker. | 1. A remote controller for a television/video recorder having a function of measuring a body fat which is capable of selecting a desired receiving/recording channel of the television broadcasting, comprising: an electrode section including a pair of current electrodes which feed a current across the body of a user to be examined for the body fat and a pair of measurement electrodes which detect an electric signal generated by the current being fed across the body of the user;an impedance measuring circuit which feeds the current between the current electrodes of the electrode section and measures an impedance across the body of the user to determine the body fat of the user from the electric signal detected by the measurement electrodes;a physical data setting section which is operated by the user for registering physical data including the weight, the height, the sex, and the age of the user;a body fat calculator which calculates the body fat of the user from the impedance across the body of the user measured by the impedance measuring circuit and the physical data of the user registered by the physical data setting section;a display which displays a resultant measurement of the body fat calculated by the body fat calculator and a message questioning whether or not the user carries a medical device including a pacemaker;a reply entry section operated by the user for entering a reply to the message; anda control section which directs the display to display the message prior to the measurement of the body fat of the user, inhibits the measurement of the body fat when a reply is received from the reply entry section indicating that the user carries a specific medical device including a pacemaker, and enables the measurement of the body fat when a reply is received from the reply entry section indicating that the user carries non of a specific medical device including a pacemaker. 2. A remote controller for a television/video recorder having a function of measuring a body fat according to claim 1, wherein when receiving from the reply entry section the reply indicating that the user carries a specific medical device including a pacemaker, the control section directs the display to display a message indicating that the measurement of the body fat is inhibited. 3. A remote controller for a television/video recorder having a function of measuring a body fact according to claim 1, wherein when receiving from the reply entry section the reply indicating that the user carries non of a specific medical device including a pacemaker, the control section allows the registration of the physical data including the weight, the height, the sex, and the age of the user to be examined. 4. A remote controller for a television/video recorder having a function of measuring a body fat which is capable of selecting a desired receiving/recording channel of the television broadcasting, comprising: an electrode section including a pair of current electrodes which feed a... | Funai Electric Co., Ltd. | G08C019/00 H04N005/44 G01R027/02 G01R027/08 A61B005/05
US07047060 | B1 | 2006-05-16 | Multiple preparatory excitations and readouts distributed over the cardiac cycle | 2001-11-16 | A magnetic resonance cardiac imaging method for imaging during a cardiac cycle interval includes monitoring an electrocardiographic signal (90) associated with the imaged heart for a first trigger event (102). Responsive to the first trigger event, a data acquisition sequence (112, 120) is applied, including a first preparation sequence block (114), a first imaging sequence block (116) having at least one readout interval (228) that collects first data (118), a second preparation sequence block (122), and a second imaging sequence block (124) having at least one readout interval (228) that collects second data (126). The data acquisition sequence (112, 120) occupies an acquisition time interval which is less than the cardiac cycle interval of the imaged heart. | 1. A magnetic resonance cardiac imaging method for imaging a heart, the method comprising applying a data acquisition sequence including: a first preparation sequence block,a first imaging sequence block having at least one readout interval that collects first data, the first imaging sequence block effectuates data acquisition having a first image contrast type a second preparation sequence block, anda second imaging sequence block having at least one readout interval that collects second data, the second imaging sequence block effectuates data acquisition having a second image contrast type that is different from the first image contrast type,the data acquisition sequence occupying an acquisition time interval which is less than a cardiac cycle interval of the heart. | 1. A magnetic resonance cardiac imaging method for imaging a heart, the method comprising applying a data acquisition sequence including: a first preparation sequence block,a first imaging sequence block having at least one readout interval that collects first data, the first imaging sequence block effectuates data acquisition having a first image contrast type a second preparation sequence block, anda second imaging sequence block having at least one readout interval that collects second data, the second imaging sequence block effectuates data acquisition having a second image contrast type that is different from the first image contrast type,the data acquisition sequence occupying an acquisition time interval which is less than a cardiac cycle interval of the heart. 2. The method as set forth in claim 1, further including: monitoring an electrocardiographic signal associated with the heart for a first trigger event; andresponsive to the first trigger event, initiating the applying of the data acquisition sequence. 3. The method as set forth in claim 2, further including: providing a gating delay interval between the first trigger event and the start of the data acquisition sequence, the gating delay interval plus a time of the data acquisition sequence together being less than the cardiac cycle interval. 4. The method as set forth in claim 3 further including: providing a trigger window interval trailing the data acquisition sequence; andterminating the trigger window interval responsive to detection of a second trigger event. 5. The method as set forth in claim 1, wherein: the first preparation sequence block performs a first magnetization preparation affecting at least a portion of the heart; andthe second preparation sequence block performs a second magnetization preparation affecting at least a portion of the heart, wherein the second magnetization preparation is different from the first magnetization preparation. 6. The method as set forth in claim 1, wherein at least one of the first preparation sequence block and the second preparation sequence block performs at least one of spatial modulation of magnetization (SPAMM) and complementary spatial modulation of magnetization (CSPAMM) tagging of at least a portion of the cardiac muscle. 7. The method as set forth in claim 1, wherein one of: the first preparation sequence block combined with the first imaging sequence block, andthe second preparation sequence block combined with the second imaging sequence block,effectuates acquisition of imaging data with one of superimposed spatial modulation of magnetization (SPAMM) tagging and superimposed complementary spatial modulation of magnetization (CSPAMM) tagging. 8. The method as set forth in claim 1, wherein: the first preparation sequence block applies a first spatial modulation of magnetization tagging; andthe second preparation sequence block applies a second spatial modulation of magnetization tagging. 9. The method as set forth in c... | Koninklijke Philips Electronics N.V. | A61B005/05
US07264593 | B2 | 2007-09-04 | Method for determining, monitoring and updating correction data for liquid-filled transmission systems | 2001-09-25 | Method for determining, monitoring and updating correction data for correcting measured value distortions and for calibrating liquid-filled transmission systems, via which signals from the patient detected in the body of a patient are transmitted to an external measured value receiver, in particular for invasive pressure measurements in cardiology, intensive care medicine and anaesthesia. The correction data are determined in the frequency range for each frequency from statistics of the signals from the patient measured at this frequency, in particular from the transmission function obtained from the signal from the patient. In the determination of the correction data, empirically checked assumptions about the corresponding statistics of undistorted signals from the patient and general properties of the transmission function are used. | 1. Method for determining, monitoring and updating correction data for correcting measured value distortions caused by a liquid-filled transmission system and for calibrating the liquid-filled transmission system, as in invasive pressure measurements in cardiology, intensive care medicine and anaesthesia, the method comprising: measuring signals from a patient using the liquid-filled transmission system,generating spectral statistics of the measured signals from the patient, anddetermining correction data in a frequency range using logarithmic transformations for each frequency from the spectral statistics of the measured signals from the patient. | 1. Method for determining, monitoring and updating correction data for correcting measured value distortions caused by a liquid-filled transmission system and for calibrating the liquid-filled transmission system, as in invasive pressure measurements in cardiology, intensive care medicine and anaesthesia, the method comprising: measuring signals from a patient using the liquid-filled transmission system,generating spectral statistics of the measured signals from the patient, anddetermining correction data in a frequency range using logarithmic transformations for each frequency from the spectral statistics of the measured signals from the patient. 2. Method according to claim 1, wherein the correction data are determined from a transmission function obtained from the signals from the patient. 3. Method according to claim 2, wherein in determining the correction data, empirically checked assumptions about corresponding statistics of distorted signals from the patient and general properties of the transmission function are used. 4. Method according to claim 2, comprising: a) segmenting the measured signals from the patient,b) performing a complex Fourier transformation of the segments matched to segment length,c) determining segment by segment statistics of complex Fourier coefficients determined by the complex Fourier transformation, andd) determining the transmission function from a ratio of the segment by segment statistics of the complex Fourier coefficients determined to statistics of complex Fourier coefficients to be expected in an undistorted signal from the patient. 5. Method according to claim 4, wherein the complex Fourier transformation of the segments matched to the segment length comprises a fast Fourier transformation by means of zero padding. 6. Method according to claim 4, wherein the segmentation takes place in signal from the patient segments which are suitable with respect to their length for a Fourier transformation. 7. Method according to claim 6, wherein a length of the signal from the patient segment suitable for the Fourier transformation is determined in that from a presettable minimum length, a sign changed of a first derived function of the measured signal from the patient occurs according to time with a boundary condition of a minimum amount of a second derived function of the signal from the patient according to time. 8. Method according to claim 6, wherein a length of the signal segment from the patient suitable for the Fourier transformation is determined in that from a presettable minimum length, a minimum fault occurs in approximation of the signal from the patient by a Fourier series cut off in a higher frequency range. 9. Method according to claim 4, comprising a trend adjustment of the segmentation of the signal from the patient or pressure signal from the patient. 10. Method according to claim 4, comprising a multiplication of the segmented signal from the patient or pressure signal from the patient. ... | Deutsches Herzzentrum Berlin Medserv GmbH | A61B005/00
US07043292 | B2 | 2006-05-09 | Single or multi-mode cardiac activity data collection, processing and display obtained in a non-invasive manner | 2002-06-21 | The method of presenting concurrent information about the electrical and mechanical activity of the heart using non-invasively obtained electrical and mechanical cardiac activity data from the chest or thorax of a patient comprises the steps of: placing at least three active Laplacian ECG sensors at locations on the chest or thorax of the patient; where each sensor has at least one outer ring element and an inner solid circle element, placing at least one ultrasonic sensor on the thorax where there is no underlying bone structure, only tissue, and utilizing available ultrasound technology to produce two or three-dimensional displays of the moving surface of the heart and making direct measurements of the exact sites of the sensors on the chest surface to determine the position and distance from the center of each sensor to the heart along a line orthogonal to the plane of the sensor and create a virtual heart surface; updating the measurements at a rate to show the movement of the heart's surface; monitoring at each ultrasonic sensor site and each Laplacian ECG sensor site the position and movement of the heart and the passage of depolarization wave-fronts in the vicinity; treating those depolarization wave-fronts as moving dipoles at those sites to create images of their movement on the image of the beating heart's surface; and, displaying the heart's electrical activity on the dynamically changing image of the heart's surface with the goal to display an approximation of the activation sequence on the beating virtual surface of the heart | 1. A method of presenting concurrent information about the electrical and mechanical activity of the heart using non-invasively obtained electrical and mechanical cardiac activity data from the chest or thorax of a patient comprising the steps of: placing at least three active concentric ring Laplacian ECG sensors at locations on the chest or thorax of the patient, each sensor having at least one outer ring element and an inner solid circle element;placing at least one ultrasonic sensor on the thorax where there is no underlying bone structure, only tissue, and utilizing available ultrasound technology to produce two or three-dimensional displays of the moving surface of the heart and making direct measurements of the exact sites of the sensors on the chest surface to determine the position and distance from the center of each sensor to the heart along a line orthogonal to the plane of the sensor and create a virtual heart surface;updating the measurements at a rate to show the movement of the heart's surface;monitoring at each ultrasonic sensor site and each Laplacian ECG sensor site the position and movement of the heart and the passage of depolarization wave-fronts in the vicinity;treating those depolarization wave-fronts as moving dipoles at those sites to create images of their movement on the image of the beating heart's surface; and,displaying the heart's electrical activity on the dynamically changing image of the heart's surface with the goal to display an approximation of the activation sequence on the beating virtual surface of the heart. | 1. A method of presenting concurrent information about the electrical and mechanical activity of the heart using non-invasively obtained electrical and mechanical cardiac activity data from the chest or thorax of a patient comprising the steps of: placing at least three active concentric ring Laplacian ECG sensors at locations on the chest or thorax of the patient, each sensor having at least one outer ring element and an inner solid circle element;placing at least one ultrasonic sensor on the thorax where there is no underlying bone structure, only tissue, and utilizing available ultrasound technology to produce two or three-dimensional displays of the moving surface of the heart and making direct measurements of the exact sites of the sensors on the chest surface to determine the position and distance from the center of each sensor to the heart along a line orthogonal to the plane of the sensor and create a virtual heart surface;updating the measurements at a rate to show the movement of the heart's surface;monitoring at each ultrasonic sensor site and each Laplacian ECG sensor site the position and movement of the heart and the passage of depolarization wave-fronts in the vicinity;treating those depolarization wave-fronts as moving dipoles at those sites to create images of their movement on the image of the beating heart's surface; and,displaying the heart's electrical activity on the dynamically changing image of the heart's surface with the goal to display an approximation of the activation sequence on the beating virtual surface of the heart. 2. The method of claim 1 comprising the further steps of creating chest surface maps of isochrones and projecting those isochrones onto a virtual surface that represents the heart, which can indicate simultaneously depolarized zones for detecting zones of abnormal depolarization and contraction patterns. 3. The method of claim 1 comprising the further steps of creating Laplacian isopotential maps on the chest surface or by projecting the map onto a virtual and moving heart surface. 4. A method of measuring, analyzing and visualizing electrical and mechanical activities in a system comprising the steps of: a. collecting signals over a part of a surface of the system using a plurality of active concentric ring Laplacian sensors, where each sensor has at least one outer ring element and an inner solid circle element and a data acquisition unit,b. collecting ultrasonic information about the distance from each sensor to the surface of the moving heart below the each sensorc. determining the positions of the sensors,d. reducing measurement noise in the collected signals with digital filtering algorithms,e. computing and monitoring the equivalent moving dipoles which generate the detected signal at the Laplacian sensors and their positions from the ultrasonic image of the heart; and,f. displaying simultaneously still images of the processed signals in both the time and space domains, together with the vir... | - | A61B005/402
US07141025 | B2 | 2006-11-28 | Apparatus and method for qualitative assessment of pelvic floor muscular strength | 2002-11-29 | An apparatus and method for qualitative assessment of pelvic floor muscular strength includes inserting a spherical intracavitary device into vagina or anus above the level of the levator plate and measuring a force necessary to remove the inserted device from the vagina or anus. A diagnostic set for qualitative assessment of pelvic floor muscular strength, includes at least one spherical intracavitary device and a force gauge. The device and force gauge can be attached together after insertion of the device into the vagina or anus. | 1. A method for qualitative assessment of pelvic floor muscular strength comprises the steps of: inserting a spherical intracavitary device into the vagina above the level of the levator plate; andmeasuring a force necessary to remove said inserted device from said vagina with a force gauge. | 1. A method for qualitative assessment of pelvic floor muscular strength comprises the steps of: inserting a spherical intracavitary device into the vagina above the level of the levator plate; andmeasuring a force necessary to remove said inserted device from said vagina with a force gauge. 2. The method of claim 1, wherein said device is inserted into said vagina in the lithotomy position. 3. The method of claim 1, wherein said force is measured in the relaxation state of pelvic floor muscles. 4. The method of claim 1, wherein said force is measured during maximum contraction of pelvic floor muscles. 5. The method of claim 1, wherein said force is measured both in the relaxation of pelvic floor muscles and during contraction of pelvic floor muscles. 6. The method of claim 1, wherein said spherical intracavitary device is a sphere having a diameter of about 28 mm to 44 mm. 7. The method of claim 1, wherein said measuring is performed by means of a force gauge attachable to said spherical intracavitary device by a linker connected to said device, said linker after said insertion extending besides the vagina to allow attachment of said force gauge. 8. The method of claim 7, wherein said linker is a thread having means for attachment of said force gauge. | - | A61B005/103
US07149564 | B2 | 2006-12-12 | Automatic analysis in virtual endoscopy | 2002-03-28 | A computer system and a computer-implemented method are provided for interactively displaying a three-dimensional rendering of a structure having a lumen and for indicating regions of abnormal wall structure. A three-dimensional volume of data is formed from a series of two-dimensional images representing at least one physical property associated with the three-dimensional structure. An isosurface of a selected region of interest is created by a computer from the volume of data based on a selected value or values of a physical property representing the selected region of interest. A wireframe model of the isosurface is generated by the computer wherein the wireframe model includes a plurality of vertices. The vertices are then grouped into populations of contiguous vertices having a characteristic indicating abnormal wall structure by the computer. The wireframe model is then rendered by the computer in an interactive three-dimensional display to indicate the populations of abnormal wall structure. | 1. A method for interactively displaying three dimensional structures comprising the steps of: providing a three-dimensional volume of data representing at least one physical property associated with a three dimensional body;isolating a selected region of interest from the volume of data based on selected values of the physical property representing the selected region of interest;rendering the isolated region of interest as a surface rendered image in an interactive three-dimensional display; anddisplaying a volume-rendered image adjoining the surface-rendered image in an interactive three-dimensional display. | 1. A method for interactively displaying three dimensional structures comprising the steps of: providing a three-dimensional volume of data representing at least one physical property associated with a three dimensional body;isolating a selected region of interest from the volume of data based on selected values of the physical property representing the selected region of interest;rendering the isolated region of interest as a surface rendered image in an interactive three-dimensional display; anddisplaying a volume-rendered image adjoining the surface-rendered image in an interactive three-dimensional display. 2. The method as recited in claim 1 including displaying at least one planar image through the surface-rendered image of the interactive three-dimensional display. 3. A method for interactively displaying three dimensional structures comprising the steps of: providing a three-dimensional volume of data representing at least one physical property associated with a three dimensional body;isolating a selected region of interest from the volume of data based on selected values of the physical property representing the selected region of interest;determining an arbitrarily oriented cross-sectional area of the selected region of interest; andrendering the arbitrarily oriented cross-sectional area of the isolated region of interest in an interactive three-dimensional display. 4. An apparatus for automatic analysis in virtual endoscopy for interactively displaying a three-dimensional rendering of an anatomical structure having a lumen, the apparatus comprising: a computer-controlled device to acquire images of selected body organs, anda renderer to produce three-dimensional renderings of images of the anatomical structures, the renderer adapted to incorporate a surface rendering of an organ into a volume rendering of the surrounding anatomy using texture memory to achieve a combined surface/volume rendering,wherein the three-dimensional renderings are interactive to enable the display of movement about the rendered anatomical structure. 5. An apparatus for automatic analysis in virtual endoscopy for interactively displaying a three-dimensional rendering of an anatomical structure, the apparatus comprising: a computer-controlled device to acquire images of selected body organs, anda renderer to produce three-dimensional renderings of anatomical structures of the images, the renderer adapted to incorporate a surface rendering of an organ into a volume rendering of the surrounding anatomy using texture memory to achieve a hybrid rendering effect,wherein the surface rendered organ may be highlighted in the context of the volume-rendered surrounding anatomy, and wherein the three-dimensional renderings are interactive to enable the display of movement about the rendered anatomical structure. 6. An apparatus for automatic analysis in virtual endoscopy for interactively displaying a three-dimensional rendering of an anatomical structure, the apparat... | Wake Forest University Health Sciences | A61B005/00
US07207941 | B2 | 2007-04-24 | Birth monitoring system | 2002-06-05 | A medical transponder, including an ultrasonic sensor that detects impinging ultrasonic waves and generates electrical signals in response thereto; an electrical connection which receives said signals; and an electromagnetic RF transmitter coupled to said electrical connection and which generates an RF signal in response to said detected waves. | 1. A birth monitoring system, comprising: an ultrasonic transmitter comprising at least one transmission element;an electromagnetic RF receiver comprising at least one RF antenna;at least two transponders adapted to attach to a cervix and which generate RF signals responsive to excitation by an ultrasonic signal;a signal distinguisher which distinguishes between RF signals from different transponders; anda controller configured to analyzes said RF signals to produce an indication of a progress of a birth. | 1. A birth monitoring system, comprising: an ultrasonic transmitter comprising at least one transmission element;an electromagnetic RF receiver comprising at least one RF antenna;at least two transponders adapted to attach to a cervix and which generate RF signals responsive to excitation by an ultrasonic signal;a signal distinguisher which distinguishes between RF signals from different transponders; anda controller configured to analyzes said RF signals to produce an indication of a progress of a birth. 2. A system according to claim 1, wherein said indication comprises a dilation of said cervix. 3. A system according to claim 1, wherein said indication comprises a station of a fetal head. 4. A system according to claim 1, wherein said indication comprises an orientation of a fetal head. 5. A system according to claim 1, wherein said system is adapted to be worn by a mother being monitored. 6. A system according to claim 1, wherein said distinguisher is part of said controller. 7. A system according to claim 1, wherein at least one of the transponders comprises: an ultrasonic sensor that detects impinging ultrasonic waves and generates electrical signals in response thereto;an electrical connection which receives said signals; andan electromagnetic RF transmitter coupled to said electrical connection and which generates an RF signal in response to said detected waves. 8. A system according to claim 7, wherein said electrical connection comprises circuitry. 9. A system according to claim 8, wherein said circuitry comprises a driving circuitry. 10. A system according to claim 8, wherein said circuitry comprises a tuning circuitry. 11. A system according to claim 8, wherein said circuitry comprises a non-linear element which generates harmonics. 12. A system according to claim 8, wherein said circuitry modifies a frequency of said detected ultrasonic waves to generate a frequency for said transmitted RF waves. 13. A system according to claim 12, wherein said modifying comprises multiplying. 14. A system according to claim 12, wherein said modifying comprises a frequency shifting. 15. A system according to claim 12, wherein said circuitry comprises a variable element for generating different frequencies from a same base circuit. 16. A system according to claim 12, wherein said circuitry resonates with said impinging waves to generate said transmitted waves. 17. A system according to claim 12, wherein said circuitry comprises a modulation circuitry that uses said impinging waves to modulate said transmitted waves. 18. A system according to claim 12, comprising at least one additional sensor and wherein said circuitry modulates said transmitted wave using a signal from said sensor. 19. A system according to claim 7, wherein said electromagnetic RF transmitter has an output lower than said signal. 20. A system according to claim 7, wherein said electrical connection drives said RF transmitter with substantially no delay relative to ... | Barnev Ltd. | A61B008/00 A61B005/103
US07063674 | B2 | 2006-06-20 | Intravascular guide wire and method for manufacture thereof | 2002-10-09 | A guide wire, and a method for the manufacture thereof, having a core and a plastic jacket enclosing at least a portion of the core. In some embodiments, the plastic jacket comprises a proximal portion formed of a first plastic material and a distal jacket portion formed of a second plastic material. In some embodiments, the core member has a proximal portion and a distal portion, and a first plastic jacket is disposed about the proximal portion and a second jacket formed of a fluoroelastomer is disposed about the distal portion. | 1. A guidewire comprising an elongate core member having a proximal portion and a distal portion, a first plastic jacket disposed about the proximal portion and a second plastic jacket formed of a fluoroelastomer disposed about the distal portion, the first plastic jacket is in contact with the second plastic jacket. | 1. A guidewire comprising an elongate core member having a proximal portion and a distal portion, a first plastic jacket disposed about the proximal portion and a second plastic jacket formed of a fluoroelastomer disposed about the distal portion, the first plastic jacket is in contact with the second plastic jacket. 2. The guide wire of claim 1, wherein the core member comprises nitinol. 3. The guide wire of claim 1, further comprising a coil at the distal portion. 4. The guide wire of claim 3, wherein the second jacket is disposed about the coil. 5. The guide wire of claim 1, wherein a portion of the second jacket is disposed on a portion of the first jacket. 6. The guide wire of claim 1, further including a hydrophilic coating disposed on a portion of the first jacket. 7. The guide wire of claim 1, further including a hydrophilic coating disposed on a portion of the second jacket. 8. A guidewire comprising an elongate core member having a proximal portion and a distal portion, means for covering the proximal portion and means for covering the distal portion, wherein the means for covering the proximal portion comprises a plastic and means for covering the distal portion comprises a fluoroelastomer, the means for covering the proximal portion is in contact with the means for covering the distal portion. 9. The guidewire of claim 8, further including means for marking with fluoroscopy the distal portion. 10. The guide wire of claim 9, wherein the means for marking with fluoroscopy includes a coil disposed about the distal portion. 11. The guide wire of claim 10, wherein the means for covering the distal portion is disposed about the means for marking with fluoroscopy. | SciMed Life Systems, Inc. | A61B005/00 A61M025/00
US07058440 | B2 | 2006-06-06 | Dynamic computed tomography imaging using positional state modeling | 2002-06-25 | An apparatus for computed tomography (CT) imaging of a cyclically moving organ includes a positional state monitor (24, 40) that monitors a positional state of the cyclically moving organ such as the heart. A cone-beam CT scanner (10) acquires image data at least within a plurality of time windows. Each time window is centered about an occurrence of a selected positional state of the organ. A window analyzer (38) selects a data segment within each time window such that the data segments combine to form a complete data set covering a selected angular range. A reconstruction processor (44) reconstructs the selected data segments into an image representation. | 1. A method for acquiring diagnostic imaging data of a dynamically moving organ, the method comprising: monitoring positional state of the organ motion as a function of time by measuring a physiological parameter and applying a positional state relationship that relates positional state to the measured physiological parameter;identifying at least two occurrences of a selected positional state within each cycle of motion of the dynamically moving organ based on the measured physiological parameter and the state relationship;selecting time windows arranged about identified occurrences of the selected positional state including at least said identified at least two occurrences;acquiring imaging data of the dynamically moving organ at least during the time windows;combining portions of the imaging data acquired during the time windows corresponding to the selected positional state to form an image data set; andreconstructing an image representation of the moving organ in the selected positional state from the image data set. | 1. A method for acquiring diagnostic imaging data of a dynamically moving organ, the method comprising: monitoring positional state of the organ motion as a function of time by measuring a physiological parameter and applying a positional state relationship that relates positional state to the measured physiological parameter;identifying at least two occurrences of a selected positional state within each cycle of motion of the dynamically moving organ based on the measured physiological parameter and the state relationship;selecting time windows arranged about identified occurrences of the selected positional state including at least said identified at least two occurrences;acquiring imaging data of the dynamically moving organ at least during the time windows;combining portions of the imaging data acquired during the time windows corresponding to the selected positional state to form an image data set; andreconstructing an image representation of the moving organ in the selected positional state from the image data set. 2. The imaging method as set forth in claim 1, further including: defining a parametric functional model of the positional state relationship, the parametric functional model being used in relating the positional state to the measured physiological parameter. 3. The imaging method as set forth in claim 1, wherein the positional state includes a volume of a selected heart chamber, and the measuring of a physiological parameter includes: performing an electrocardiographic measurement. 4. The method as set forth in claim 1, wherein the combining step includes: selecting a data portion within each window such that the data portions provide a selected angular coverage. 5. The method as set forth in claim 1, wherein the positional state includes one of a spatial position of the organ and a spatial volume of the organ. 6. The method as set forth in claim 1, wherein the dynamically moving organ includes at least a portion of the heart, and the time windows include: a first time window arranged about a first occurrence of the selected positional state during the systolic cardiac cycle portion; anda second time window arranged about a second occurrence of the selected positional state during the diastolic cardiac cycle portion. 7. A computer medium programmed to implement the method of claim 1. 8. An imaging method for acquiring diagnostic imaging data of a dynamically moving organ, the method comprising: monitoring positional states of the moving organ as a function of time;selecting time windows arranged about occurrences of a selected positional state;acquiring imaging data of the dynamically moving organ at least during the time windows using an orbiting x-ray source;for each time window, selecting an angular segment of the orbiting x-ray source substantially centered within the time window;shifting each angular segment within its corresponding time window to provide a complete image data set for reconstruction while optimizing ... | Koninklijke Philips Electronics N.V. | A61B005/05
US07118539 | B2 | 2006-10-10 | Articulating guide wire for embolic protection and methods of use | 2002-02-26 | An intravascular guidewire having an articulating distal section for improved maneuverability beyond a lesion or other protrusion within a body lumen. An elongated core wire having an actuatable stop disposed thereon is adapted to have a reduced profile in a first position, and an enlarged profile in a second position. The actuatable stop can be adapted to expand by any number of means, including a spring coil, a polymeric tube or a mesh sleeve. In one particular embodiment, the actuatable stop comprises a distal section of a tubular member having a plurality of circumferentially disposed openings adapted to permit a plurality of struts disposed therebetween to expand in an outward direction when compressed axially. | 1. An articulating guidewire, comprising: an elongated core wire having a longitudinal axis, a proximal end and a distal end;an actuatable stop disposed at the distal end of said core wire, said actuatable stop moveable between a collapsed position and an expanded position the stop having a first diameter in the expanded position;an actuator positioned adjacent the stop, said actuator moveable between a first position and a second position to move the stop between the collapsed position and the expanded position; anda filter slidably disposed on the core wire proximally of said actuatable stop the filter having an expanded position with a second diameter, the second diameter being larger than the first diameter, allowing fluids that flow through the filter to flow around the stop. | 1. An articulating guidewire, comprising: an elongated core wire having a longitudinal axis, a proximal end and a distal end;an actuatable stop disposed at the distal end of said core wire, said actuatable stop moveable between a collapsed position and an expanded position the stop having a first diameter in the expanded position;an actuator positioned adjacent the stop, said actuator moveable between a first position and a second position to move the stop between the collapsed position and the expanded position; anda filter slidably disposed on the core wire proximally of said actuatable stop the filter having an expanded position with a second diameter, the second diameter being larger than the first diameter, allowing fluids that flow through the filter to flow around the stop. 2. The guidewire of claim 1, wherein said actuatable stop comprises a tubular member having a proximal end and a distal end. 3. The guidewire of claim 2, wherein the distal section of said tubular member includes a plurality of circumferentially disposed openings adapted to permit a plurality of struts disposed therebetween to expand in an outward direction. 4. The guidewire of claim 2, wherein the inner diameter of said tubular member is substantially similar to the outer diameter of the core wire. 5. The guidewire of claim 2, wherein the inner diameter of said tubular member is larger than the outer diameter of the core wire. 6. The guidewire of claim 1, further comprising a locking mechanism adapted to prevent relative motion between the actuatable stop and the core wire. 7. The guidewire of claim 6, wherein the locking mechanism comprises an enlarged outer diameter portion disposed on the core wire. 8. The guidewire of claim 6, wherein the locking mechanism comprises an enlarged outer diameter portion disposed on the core wire corresponding in size and shape to a reduced inner diameter portion disposed on the actuatable stop. 9. The guidewire of claim 6, wherein the locking mechanism comprises a locking hub disposed about a proximal portion of the core wire. 10. The guidewire of claim 1, wherein the actuator contacts the actuatable stop, causing the actuatable stop to move between contracted and expanded positions. 11. The guidewire of claim 10, wherein the actuator has proximal and distal ends, and the distal end of the actuator contacts the actuatable stop. 12. An articulating guidewire, comprising: an elongated core wire having a longitudinal axis, a proximal end and a distal end;an actuatable stop disposed at the distal end of said core wire, said actuatable stop moveable between a collapsed position and an expanded position;an actuator disposable about the core wire, said actuator having a proximal end and a distal end; anda filter slidably disposed on the core wire proximally of said actuatable stop. 13. The guidewire of claim 12, wherein said actuatable stop comprises a spring coil. 14. The guidewire of claim 12, wherein said actuatable stop comp... | Scimed Life Systems, Inc. | A61B005/00 A61M025/00 A61M029/00
US07074191 | B2 | 2006-07-11 | Method and device for non-intrusive measurement of blood pressure in a circuit of a dialysis machine | 2002-05-16 | A method for the non-intrusive measurement of blood pressure in a circuit of a dialysis machine, using a device provided with a sleeve forming a portion of the blood circuit and having at least one part which moves according to the pressure difference between the inside and the outside of the sleeve, and with an emitter and a receiver of electromagnetic waves. The method comprising the emission of a beam of electromagnetic waves and the receiving of a beam of electromagnetic waves reflected from the moving part, in order to measure the displacement of the moving part in such a way as to establish a relationship between the position of the moving part and the pressure of the blood circulating in the sleeve. | 1. A method for the non-intrusive measurement of blood pressure in a circuit of a dialysis machine, by means of a sleeve forming a portion of said circuit and comprising a part which moves according to a pressure difference between an inside and an outside of the sleeve, the method being characterized in that a beam of electromagnetic waves is directed toward said moving part and the beam reflected from said moving part is received in order to determine the displacement of the moving part with respect to a reference position of the moving part, said moving part being a rigid part of a wall of said sleeve and having an inner face which can be placed in contact with the blood circulating in said circuit. | 1. A method for the non-intrusive measurement of blood pressure in a circuit of a dialysis machine, by means of a sleeve forming a portion of said circuit and comprising a part which moves according to a pressure difference between an inside and an outside of the sleeve, the method being characterized in that a beam of electromagnetic waves is directed toward said moving part and the beam reflected from said moving part is received in order to determine the displacement of the moving part with respect to a reference position of the moving part, said moving part being a rigid part of a wall of said sleeve and having an inner face which can be placed in contact with the blood circulating in said circuit. 2. A method according to claim 1, wherein the electromagnetic waves are in the visible spectrum. 3. A method according to claim 1, wherein said beam of electromagnetic waves is emitted by an emitter and received by a receiver, said emitter and said receiver being located on the dialysis machine. 4. A method according to claim 1, wherein an emitted beam is directed in such a way that said emitted beam is inclined with respect to an outer face of said moving part. 5. A method according to claim 1, wherein said beam of electromagnetic waves is emitted by an emitter and received by a receiver, said method further comprising a zero setting stage, in which said reference position of the moving part with respect to the emitter is determined, said stage of zero setting comprising associating at least one sensitive cell of the receiver with a known value of pressure within the sleeve; said reference position being determined for an operating condition in which the pressure acting on the outer face is equal to the pressure acting on the inner face. 6. A method according to claim 5, wherein said pressure is the ambient pressure; said stage of zero setting being implemented before a start of a dialysis treatment. 7. A device for the non-intrusive measurement of blood pressure in a circuit of a dialysis machine, the device comprising a sleeve provided with a part which moves according to the pressure difference between the inside and the outside of the sleeve, the device being characterized in that it comprises an emitter of a beam of electromagnetic waves and a receiver of a beam of electromagnetic waves reflected from the moving part to measure the displacements of said moving part with respect to a reference position of the moving part, said moving part being a rigid cart of a wall of said sleeve. 8. A device according to claim 7, wherein the emitter and the receiver are fitted on the dialysis machine in a mounting for receiving said sleeve. 9. A device according to claim 7, wherein said moving part has an inner face which can be placed in contact with the blood flowing within the circuit and an outer face which can reflect the emitted beam. 10. A device according to claim 7, wherein the wall comprises a rigid fixed part, said moving part and a flex... | Gambro Lundia AB | A61B005/021 G01L009/00 G01L007/00 A61M001/14 B01D061/28
US07179450 | B2 | 2007-02-20 | Methods for in vivo evaluation of pulmonary physiology and/or function using NMR signals of polarized Xe | 2003-01-31 | In certain embodiments, methods of the present invention obtain dynamic data sets of an NMR spectroscopy signal of polarized 129Xe in a selected structure, environment, or system. The signal data can be used to evaluate: (a) the physiology of a membrane or tissue; (b) the operational condition or function of a body system or portion thereof (when at rest or under stimulation); and/or (c) the efficacy of a therapeutic treatment used to treat a diagnosed disorder, disease, or condition. Thus, the present invention provides methods for screening and/or diagnosing a respiratory, cardiopulmonary disorder or disease such as chronic heart failure, and/or methods for monitoring the efficacy of therapeutics administered to subject to treat the disorder or disease. | 1. An in vivo method for evaluating the physiological response of a subject to a therapeutic agent, comprising: delivering polarized 129Xe gas in vivo to a subject;obtaining a first NMR spectroscopic signal of the polarized gas in the subject at at least one chemical shift frequency to generate a first dynamic data set of the NMR spectroscopic signal values over time, with the polarization of the 129Xe being repetitively destroyed at predetermined delays in a compartment or compartments of interest, the dynamic data set being representative of the polarized gas in a physiologic structure, environment, or system of interest;administering a physiological active therapeutic agent to the subject;obtaining a second NMR spectroscopic signal of the polarized gas in the subject at the at least one chemical shift frequency to generate a second dynamic data set of the NMR spectroscopic signal values over time, with the polarization of the 129Xe being repetitively destroyed at predetermined delays in a compartment or compartments of interest; andcomparing the first and second dynamic data sets to evaluate the physiological response of the subject to the therapeutic agent. | 1. An in vivo method for evaluating the physiological response of a subject to a therapeutic agent, comprising: delivering polarized 129Xe gas in vivo to a subject;obtaining a first NMR spectroscopic signal of the polarized gas in the subject at at least one chemical shift frequency to generate a first dynamic data set of the NMR spectroscopic signal values over time, with the polarization of the 129Xe being repetitively destroyed at predetermined delays in a compartment or compartments of interest, the dynamic data set being representative of the polarized gas in a physiologic structure, environment, or system of interest;administering a physiological active therapeutic agent to the subject;obtaining a second NMR spectroscopic signal of the polarized gas in the subject at the at least one chemical shift frequency to generate a second dynamic data set of the NMR spectroscopic signal values over time, with the polarization of the 129Xe being repetitively destroyed at predetermined delays in a compartment or compartments of interest; andcomparing the first and second dynamic data sets to evaluate the physiological response of the subject to the therapeutic agent. 2. A method according to claim 1, wherein the compartment of interest comprises the alveolar-capillary membrane. 3. A method according to claim 1, further comprising the step of calculating the time constant associated with the time it takes the polarized 129Xe to diffuse across the alveolar-capillary membrane based on the data provided by the obtaining steps. 4. A method according to claim 1, further comprising the steps of generating a plurality of 129Xe uptake curves of signal strength versus time based on said first and second dynamic data sets and determining the thickness of the alveolar-capillary membrane based on data provided by said calculating and generating steps. 5. A method according to claim 1, wherein the evaluating step considers the response in the function of the cardiopulmonary system based on the comparing steps. 6. A method according to claim 1, wherein the evaluating step considers the response in the function or structure of the alveolar-capillary membrane. 7. A method according to claim 6, wherein the evaluating step comprises calculating and comparing the time constants associated with each data set to assess the function of the alveolar-capillary membrane. 8. A method according to claim 7, further comprising determining the thickness of the alveolar-capillary membrane before and after administration of the therapeutic agent based on the obtaining steps. 9. A method according to claim 1, wherein the step of administering 129Xe is carried out by the subject inhaling a quantity of the polarized 129Xe the 129Xe diffusing across the alveolar-capillary membrane into pulmonary blood, and wherein the first and second obtaining steps are carried out to include two frequencies, one associated with the 129Xe in tissue and one associated with 129Xe in the blood, and ... | Medi-Physics, Inc. | A61B005/055
US07187965 | B2 | 2007-03-06 | Cardiac rhythm monitoring device | 2002-05-29 | The present invention is a cardiac rhythm-monitoring device, which allows patients to perform a preliminary screening for supraventricular arrhythmia. The device detects beat-to-beat heart rhythms (i.e. the R—R interval between individual heart beats) and performs a screening test to determine if there are indications of arrhythmia. The test looks for variance in the R—R interval that is outside of the normal range, either using a pre-constructed chart based upon general population studies to determine the normal range of variance or using normal distribution analysis of the patient's own heart rhythm to determine the normal range of variance for determining irregular heartbeats, and if there are multiple irregularities within the sensed time frame, the patient is warned of potential supraventricular arrhythmia. By sensing both electrical impulses form the heart and mechanical responses to the heartbeat, the device may augment its analysis. | 1. A device comprising: a non-invasive sensor for detecting beat-to-beat heartbeat intervals, comprised of two conductive electrodes, wherein said sensor detects beat-to-beat heartbeat intervals by sensing the electrical impulses generated by a patient's heart;a means for storing the beat-to-beat heartbeat interval data sensed by said sensor;an arrhythmia detection algorithm which monitors the variance in the R—R interval between heartbeats and wherein said arrhythmia detection algorithm determines a normal range of variance if the variance in the R—R interval between heartbeats is less than or equal to approximately 125 milliseconds and determines an irregular variance if the variance in the R—R interval between heartbeats is greater than approximately 125 milliseconds;a microprocessor for executing said algorithm with the beat-to-beat heartbeat interval data sensed by said sensor, to generate a result;an output device for conveying the results of said microprocessor's application of said algorithm to the beat-to-beat heartbeat interval data sensed by said sensor; anda power source to provide power to said microprocessor, said output device, said sensor, and said means for storing. | 1. A device comprising: a non-invasive sensor for detecting beat-to-beat heartbeat intervals, comprised of two conductive electrodes, wherein said sensor detects beat-to-beat heartbeat intervals by sensing the electrical impulses generated by a patient's heart;a means for storing the beat-to-beat heartbeat interval data sensed by said sensor;an arrhythmia detection algorithm which monitors the variance in the R—R interval between heartbeats and wherein said arrhythmia detection algorithm determines a normal range of variance if the variance in the R—R interval between heartbeats is less than or equal to approximately 125 milliseconds and determines an irregular variance if the variance in the R—R interval between heartbeats is greater than approximately 125 milliseconds;a microprocessor for executing said algorithm with the beat-to-beat heartbeat interval data sensed by said sensor, to generate a result;an output device for conveying the results of said microprocessor's application of said algorithm to the beat-to-beat heartbeat interval data sensed by said sensor; anda power source to provide power to said microprocessor, said output device, said sensor, and said means for storing. 2. A device as in claim 1 wherein said microprocessor counts the number of irregular variances in the R—R interval between heartbeats and transmits a signal to said output device warning of a potential arrhythmia in the event that the number of irregular variances in the R—R interval between heartbeats exceeds a pre-set criteria. 3. A device comprising: a non-invasive sensor for detecting beat-to-beat heartbeat intervals, comprised of two conductive electrodes, wherein said sensor detects beat-to-beat heartbeat intervals by sensing the electrical impulses generated by a patient's heart;a means for storing the beat-to-beat heartbeat interval data sensed by said sensor;an arrhythmia detection algorithm which normalizes the sensed beat-to-beat heartbeat interval data by dividing each sensed beat-to-beat heartbeat interval by only the succeeding sensed beat-to-beat interval, and which employs statistical normal distribution analysis to classify said normalized heartbeat intervals as indicative of instability and which indicates an unstable heart rhythm possibly indicative of arrhythmia in the event that the number of irregular variances in the R—R interval between heartbeats exceeds a pre-set criteria;a microprocessor for executing said algorithm with the beat-to-beat heartbeat interval data sensed by said sensor, to generate a result;an output device for conveying the results of said microprocessor's application of said algorithm to the beat-to-beat heartbeat interval data sensed by said sensor; anda power source to provide power to said microprocessor, said output device, said sensor, and said means for storing. | - | A61B005/0452
US07076288 | B2 | 2006-07-11 | Method and system for detecting and/or predicting biological anomalies | 2003-01-29 | Biological anomalies are detected and/or predicted by analyzing input biological or physical data using a data processing routine. The data processing routine includes a set of application parameters associated with biological data correlating with the biological anomalies. The data processing routine uses an algorithm to produce a data series, e.g., a PD2i data series. The data series is used to detect or predict the onset of the biological anomalies. To reduce noise in the data series, the slope is set to a predetermined number if it is less than a predetermined value. To further reduce noise, a noise interval within the data series is determined and, if the noise interval is within a predetermined range, the data series is divided by another predetermined number, and new values are produced for the data series. | 1. A method of detecting or predicting biological anomalies, comprising the steps of: analyzing input biological or physical data using a data processing routine including a set of application parameters associated with biological data correlating with the biological anomalies to produce a data series,determining whether a slope of the data series is smaller than a predetermined value;if the slope is less than a predetermined value, setting the slope to a predetermined number; andusing the data series to detect or predict the onset of the biological anomalies,wherein the data processing routine uses the following algorithm to produce a data series PD2i: PD2i˜log C(n, r, nref*)/log rwhere ˜ means scales as, C is the correlation integral for PD2i in which n equals the data length, r equals the scaling range, and nref* equals a location of the reference vector for estimating the scaling region slope of log C/log r in a restricted small log-r range that is devoid of the effects of non-stationary data. | 1. A method of detecting or predicting biological anomalies, comprising the steps of: analyzing input biological or physical data using a data processing routine including a set of application parameters associated with biological data correlating with the biological anomalies to produce a data series,determining whether a slope of the data series is smaller than a predetermined value;if the slope is less than a predetermined value, setting the slope to a predetermined number; andusing the data series to detect or predict the onset of the biological anomalies,wherein the data processing routine uses the following algorithm to produce a data series PD2i: PD2i˜log C(n, r, nref*)/log rwhere ˜ means scales as, C is the correlation integral for PD2i in which n equals the data length, r equals the scaling range, and nref* equals a location of the reference vector for estimating the scaling region slope of log C/log r in a restricted small log-r range that is devoid of the effects of non-stationary data. 2. The method of claim 1, wherein the predetermined value is approximately 0.5. 3. The method of claim 1, wherein the predetermined number is zero. 4. The method of claim 1, wherein the input biological or physical data includes electrophysiological data. 5. The method of claim 4, wherein the input biological or physical data includes electrophysiological data and ECG data that is analyzed to detect or predict the onset of at least one of cardiac arrhythmias and cerebral epileptic seizure and/or to measure the severity of myocardial ischemia. 6. A method of detecting or predicting biological anomalies, comprising the steps of: analyzing input biological or physical data using a data processing routine including a set of application parameters associated with biological data correlating with the biological anomalies to produce a data series.determining whether a slope of the data series is smaller than a predetermined value;if the slope is less than a predetermined value, setting the slope to a predetermined number;using the data series to detect or predict the onset of the biological anomalies; determining a noise interval within the data series; andif the noise interval is within a predetermined range, dividing the data series by another predetermined number and repeating the step of analyzing to produce new values for the data series. 7. The method of claim 6, wherein the other predetermined number is two. 8. The method of claim 6, wherein the predetermined range is −x to +x, where x is any number. 9. The method of claim 8, wherein the predetermined range is −5 to +5. 10. A method of detecting or predicting biological anomalies, comprising the steps of: analyzing input biological or physical data using a data processing routine including a set of application parameters associated with biological data correlating with the biological anomalies to produce a data series;determining a noise interval within the data series; andif the noise interval... | Vicor Technologies, Inc. | A61B005/04
US07209886 | B2 | 2007-04-24 | System and method for implementing healthcare fraud countermeasures | 2003-01-22 | A method of discouraging healthcare fraud in conjunction with providing a health care service to a patient in which the patient provides a biometric signature. An event record is created where the event record creates an association between the biometric signature, a patient identifier, and location information indicating a location at which the biometric signature is obtained. The event record is stored in a central data store and can be later retrieved to verify the patient was present at the location. | 1. A method of discouraging phantom billing healthcare fraud comprising: providing a health care service;in conjunction with providing said health care service at a particular location to a patient, causing said patient to provide a biometric signature;capturing said biometric signature;creating an event record that records said patient was at said particular location at a particular time, said event record consisting of a patient identifier, said biometric signature, and location information indicating said particular location at which said biometric signature is obtained and at which said health care service is provided and said event record indicates the absence of phantom billing, wherein said event record is created using a computer system;transmitting said event record including said biometric signature to a computerized central data store that is accessible by a payer entity; andstoring said event record in said computerized central data store. | 1. A method of discouraging phantom billing healthcare fraud comprising: providing a health care service;in conjunction with providing said health care service at a particular location to a patient, causing said patient to provide a biometric signature;capturing said biometric signature;creating an event record that records said patient was at said particular location at a particular time, said event record consisting of a patient identifier, said biometric signature, and location information indicating said particular location at which said biometric signature is obtained and at which said health care service is provided and said event record indicates the absence of phantom billing, wherein said event record is created using a computer system;transmitting said event record including said biometric signature to a computerized central data store that is accessible by a payer entity; andstoring said event record in said computerized central data store. 2. The method of claim 1 wherein said act of storing said event record is performed before said biometric signature is used to identify said patient receiving said healthcare services. 3. The method of claim 1 wherein said act of storing said event record is performed before said biometric signature is used to authenticate said person receiving said healthcare service. 4. The method of claim 1 wherein said healthcare service comprises a service that is reimbursable by a private insurance plan and/or a public program. 5. The method of claim 1 further comprising obtaining a biometric signature at least two times during said provision of health care services. 6. The method of claim 1 further comprising: receiving a claim for reimbursement for said health care services;retrieving one or more event records corresponding to said health care services from said central data store;matching said claim for reimbursement to an event record for a particular patient to determine whether said one or more event records contain location information indicating that said patient was at a suitable location for providing said health care services; andauthorizing financial reimbursement of said health care service based in part upon said determination that said patient was present at a location suitable for providing said health care service. 7. The method of claim 1 further comprising: retrieving one or more records from said central data store; andperforming a fraud detection process using said retrieved records. 8. The method of claim 1 wherein said biometric signature comprises a fingerprint scan record. 9. The method of claim 1 further comprising: including within said event record an indication of consumable materials used during provision of said healthcare services. 10. The method of claim 1 further comprising: including within said event record and indication of equipment used during provision of said healthcare services. 11. The method of claim 1 further comprising: including within said event record ... | Biometric Technologies, Inc. | A61B005/00
US07187964 | B2 | 2007-03-06 | Cardiac catheter imaging system | 2002-09-26 | A device for measuring parameters of human tissue includes a multielectrode catheter for taking multiple measurements of the electrical characteristics of the human tissue, a concentric tube catheter located inside the multielectrode catheter, for providing structural support to the multi-electrode catheter and for serving as a conduit for advancing or withdrawing the multielectrode catheter over its surface; and an imaging catheter located inside the concentric tube catheter for taking multiple measurements of anatomical characteristics of the human tissue. | 1. A device for measuring electrical and geometric characteristics of body tissue from a blood-filled cavity within the tissue, comprising: a multielectrode lumen catheter, having multiple electrodes arranged in a fixed pattern on a continuous surface, wherein the electrodes are configured to take multiple simultaneous non-contact measurements of electrical potentials resulting from electrical activity from multiple locations in the tissue; andan anatomical imaging catheter having at least one imaging element for visualizing anatomical characteristics located inside the multielectrode lumen catheter, wherein the anatomical imaging catheter is configured to take multiple non-contact measurements of anatomical characteristics of the tissue, and for determining location and orientation of the multielectrode catheter with respect to the tissue;wherein the multielectrode catheter and the imaging catheter are configured to provide the measurements of electrical potentials and the measurements of anatomical characteristics to a data processing system for reconstruction of tissue surface electrograms. | 1. A device for measuring electrical and geometric characteristics of body tissue from a blood-filled cavity within the tissue, comprising: a multielectrode lumen catheter, having multiple electrodes arranged in a fixed pattern on a continuous surface, wherein the electrodes are configured to take multiple simultaneous non-contact measurements of electrical potentials resulting from electrical activity from multiple locations in the tissue; andan anatomical imaging catheter having at least one imaging element for visualizing anatomical characteristics located inside the multielectrode lumen catheter, wherein the anatomical imaging catheter is configured to take multiple non-contact measurements of anatomical characteristics of the tissue, and for determining location and orientation of the multielectrode catheter with respect to the tissue;wherein the multielectrode catheter and the imaging catheter are configured to provide the measurements of electrical potentials and the measurements of anatomical characteristics to a data processing system for reconstruction of tissue surface electrograms. 2. The device of claim 1, further comprising a coaxial tube catheter located inside the multielectrode lumen catheter, wherein the coaxial tube catheter provides structural support to the multielectrode lumen catheter, and serves as a conduit for advancing or withdrawing the multielectrode lumen catheter over the surface of the coaxial tube catheter, and for advancing or withdrawing the anatomical imaging catheter within a lumen of the coaxial tube catheter. 3. The device of claim 2, wherein the tissue comprises heart tissue and wherein the anatomical imaging catheter comprises an ultrasound transducer, which in operation continually provides tomographic sections of the heart cavity. 4. The device of claim 2, wherein the tissue comprises heart tissue and wherein the multielectrode lumen catheter is receptive to continuous cavitary electrical potentials, for producing signals for displaying graphical isopotential and isochronal maps of the electrical characteristics of the heart tissue. 5. The device of claim 2, wherein a roving electrode-catheter is inserted into the tissue cavity through the center of the multielectrode catheter, and the roving catheter emits and receives a location signal for determining the position and shape of the roving catheter. 6. The device of claim 2, wherein the data processing system comprises a data acquisition system, a data analysis system, and a data display system coupled to the device, wherein the data acquisition system is responsive to the multiple non-contact measurements of the electrical and anatomical characteristics to provide electrical and anatomical data to the data analysis system,wherein the data analysis system is responsive to the electrical and anatomical data to reconstruct the tissue surface electrog rams by solving Laplace's equation,wherein Laplace's equation is solved by employing the boundary elem... | Dirar S. Khoury | A61B005/04
US07049817 | B2 | 2006-05-23 | Thermostat-controlled containment cell for samples intended for NMR measurements and method for implementing it | 2002-04-05 | A thermostatically controlled containment cell for a sample, suitable notably for a measuring device such as a nuclear magnetic resonance (NMR) spectrometer, NMR type measuring device including the cell, and method for implementation is disclosed. Cell (8) comprises a body of a tubular sleeve (9) made of a material so selected as not to interfere with NMR measurements, (for example a non-magnetic and non electrically conducting material) ending at its opposite ends with terminal parts (14, 16), containing a sample (S), and means (19–22) for circulating within tubular sleeve (9) a controlled-temperature heat carrier for maintaining the sample at a substantially constant temperature during measurements which liquid is selected to not disturb the NMR measurements. The cell containing the sample is first heated by circulation of the heat carrier. An overpressure sufficient to prevent boiling is established in the cell. The cell can then be introduced into or removed from the NMR measuring device without stopping the circulation of the heating fluid. Applications notably are calibration of logging measurements on geologic samples or others. | 1. A removable thermostatically controlled containment cell configured for preconditioning at a predetermined temperature a sample prior to measurements in a nuclear magnetic resonance spectrometer comprising: magnets;coils placed in an air gap of the magnets allowing generation of a modulated magnetic field;a tubular sleeve made of a material which does not interfere with nuclear magnetic resonance measurements, ending at opposite ends thereof with terminal parts between which a sample is arranged; andmeans for circulating in the tubular sleeve a controlled temperature heat carrier fluid in order to bring the sample to a predetermined temperature, substantially constant during measurement; and whereinthe heat carrier fluid does not disturb the nuclear magnetic resonance measurements and the sleeve is introducible inside the coils without stopping a circulation of the heat carrier fluid. | 1. A removable thermostatically controlled containment cell configured for preconditioning at a predetermined temperature a sample prior to measurements in a nuclear magnetic resonance spectrometer comprising: magnets;coils placed in an air gap of the magnets allowing generation of a modulated magnetic field;a tubular sleeve made of a material which does not interfere with nuclear magnetic resonance measurements, ending at opposite ends thereof with terminal parts between which a sample is arranged; andmeans for circulating in the tubular sleeve a controlled temperature heat carrier fluid in order to bring the sample to a predetermined temperature, substantially constant during measurement; and whereinthe heat carrier fluid does not disturb the nuclear magnetic resonance measurements and the sleeve is introducible inside the coils without stopping a circulation of the heat carrier fluid. 2. A cell as claimed in claim 1, comprising means for establishing in the containment cell an overpressure sufficient to prevent a fluid saturated sample from boiling. 3. A cell as claimed in claim 1, wherein the body of the containment cell is made of a non-magnetic and non electrically conducting material. 4. A cell as claimed in claim 2, wherein the body of the containment cell is made of a non-magnetic and non electrically conducting material. 5. A method for conditioning at a substantially constant temperature a sample to be tested in a nuclear magnetic resonance device, comprising: first heating the sample to a required temperature in a thermostatically controlled containment cell comprising a body of a tubular sleeve made of a material which does not interfere with nuclear magnetic resonance measurements, ending at opposite ends thereof with terminal parts, containing the sample, and means for circulating in the tubular sleeve a controlled temperature heat carrier fluid for bringing or maintaining the sample at a predetermined temperature throughout the nuclear magnetic resonance measurements, the heat carrier fluid not disturbing the nuclear magnetic resonance measurements; andtransferring the containment cell into the nuclear magnetic resonance device without interrupting the heat carrier fluid circulation in the tubular sleeve. 6. A device configured for making nuclear magnetic resonance measurements on samples, comprising: a nuclear magnetic resonance apparatus including magnets, coils placed in an air gap of the magnets, an excitation signal generator and signal acquirer for generating the excitation signals creating an oscillating electric field and for acquiring a response of a sample to the excitation signals;a first tubular sleeve outside the coils and connected to a circuit in which a controlled temperature cooling fluid circulates;a removable thermostatically controlled containment cell for preconditioning at a predetermined temperature a sample prior to nuclear magnetic resonance measurements in the nuclear magnetic resonance apparatus;a t... | Institut Francais du Petrole | G01V003/00 A61B005/05
US07039458 | B2 | 2006-05-02 | Body fat measuring system for pregnant woman and health care system for pregnant woman | 2002-07-23 | Measurements by use of a conventional caliper process require skills and experiences. Further, since different pregnant women gain and keep fats in their bodies in different ways, results of measurements of their body fats vary. To overcome the inconvenience, there is provided a system for measuring a percent body fat or a total body fat by a bioelectrical impedance method, which comprises first input means, second input means and computation means, wherein personal data of a patient such as a height and a body weight is input into the first input means, a weight of a fetal part is input into the second input means, and the computation means computes a percent body fat or a total body fat by subtracting the body weight input into the second input means from the body weight input into the first input means. Further, to overcome the inconvenience, there is provided a system for measuring a percent body fat by a bioelectrical impedance method which comprises first input means, second input means and computation means, wherein personal data of a patient such as a height and a body weight is input into the first input means, a weight of a fetal part is input into the second input means, and the computation means computes a percent body fat by subtracting the body weight input into the second input means from the body weight input into the first input means. | 1. A system for measuring a percent body fat of a pregnant woman, the system comprising: means for measuring a bioelectrical impedance;first input means for receiving personal data of a patient,second input means for receiving a weight of a fetal part, andcomputation means for subtracting the body weight input into the second input means from a body weight input into the first input means to yield a subtraction result, and for computing a percent body fat based on the subtraction result and the measured bioelectrical impedance. | 1. A system for measuring a percent body fat of a pregnant woman, the system comprising: means for measuring a bioelectrical impedance;first input means for receiving personal data of a patient,second input means for receiving a weight of a fetal part, andcomputation means for subtracting the body weight input into the second input means from a body weight input into the first input means to yield a subtraction result, and for computing a percent body fat based on the subtraction result and the measured bioelectrical impedance. 2. The system of claim 1, wherein the weight of the fetal part is determined from the gestational weeks. 3. The system of claim 1, wherein the weight of the fetal part is determined from a length of uterine fundus. 4. The system of claim 1, wherein the weight of the fetal part is determined by ultrasonotomography. 5. The system of claim 1, wherein the weight of the fetal part is determined from the gestational weeks and a length of uterine fundus. 6. The system of claim 1, wherein the weight of the fetal part is determined from the gestational weeks and ultrasonotomography. 7. The system of claim 1, wherein the weight of the fetal part is determined from a length of uterine fundus and ultrasonotomography. 8. The system of claim 1, wherein the weight of the fetal part comprises a fetal body weight, a weight of an amniotic fluid and a weight of a placenta. 9. A system for measuring a total body fat of a pregnant woman, the system comprising: means for measuring a bioelectrical impedance;first input means for receiving personal data of a patient,second input means for receiving a weight of a fetal part, andcomputation means for subtracting the body weight input into the second input means from a body weight input into the first input means to yield a subtraction result, and for computing a total body fat based on the subtraction result and the measured bioelectrical impedance. 10. A system for measuring a percent body fat of a pregnant woman, the system comprising: means for measuring a bioelectrical impedance;first input means for receiving personal data of a patient,second input means for receiving a weight of a fetal part,computation means for subtracting the body weight input into the second input means from a body weight input into the first input means to yield a subtraction result, anddetermination means for determining the percent body fat based on gestational weeks, the subtraction result, and the measured bioelectrical impedance. 11. A system for measuring a total body fat of a pregnant woman, the system comprising: means for measuring a bioelectrical impedance;first input means for receiving personal data of a patient,second input means for receiving a weight of a fetal part,computation means for subtracting the body weight input into the second input means from a body weight input into the first input means to yield a subtraction result, anddetermination means for determining the total body fat based on ... | Tanita Corporation | A61B005/05
US07097825 | B2 | 2006-08-29 | Methods and reagents to acquire MRI signals and images | 2002-02-25 | Method for obtaining BOLD MRI images with allosteric effector compounds are disclosed. Methods for the determination of tumor oxygenation and for an optimal time for initiating radiation therapy are also disclosed. | 1. A method for measuring a blood oxygen level-dependent magnetic resonance imaging signal, compnsing a) administering an effective amount of an allosteric effector compound capable of decreasing hemoglobin binding affinity for oxygen; andb) performing a blood oxygen level-dependent magnetic resonance imaging scan, whereby said blood oxygen level-dependent magnetic resonance imaging signal is measured, wherein the allosteric effector compound is a compound selected from the group consisting of:a compound having the formula:where R1-5 may be hydrogen, halogen, or a substituted or unsubstituted C1-3 alkyl group and may be the same or different,R6-7 may each be hydrogen or methyl and may be the same or different, andR8 may be hydrogen, a substituted or unsubstituted C1-3 alkyl group, or a salt cation, and X and Z are CH2, NH, or O;a compound having the formula:where X and Z may each be CH2, CO, NH or O, and Y may be CO or NH, which the caveat that X, Y, and Z must all be different from each other, andR2-6 can be the hydrogen, halogen, substituted or unsubstituted C1-3 alkyl groups, and may be the same or different,R7-8 can be hydrogens, methyls, ethyls, or alkyl groups in a ring connecting the two, andR9 can be a hydrogen, lower alkyl, or salt cation;a compound having the formula:where R3-6 can be the hydrogen, halogen, substituted or unsubstituted C1-3 alkyl group, or a C1-3 ether or ester, and these moieties may be the same or different, or alkyl moieties of an aromatic or aliphatic ring incorporating two of the R3-6,R1 can be connected to any position on the phenyl ring, andsites R7-8 can be hydrogen, halogen, methyl, ethyl, and these moieties may be the same or different, or alkyl groups in a ring connecting the two, andR9 can be a hydrogen, halogen, C1-3 lower alkyl, or salt cation;a compound having the formula:where R1 can be connected to any position on the phenyl ring, andsites R7-8 can be hydrogen, halogen, methyl, ethyl, and these moieties may be the same or different, or alkyl groups in a ring connecting the two, andR2 is defined as a substituted or unsubstituted aromatic compound, a substituted or unsubstituted alkyl ring compound, or a substituted or unsubstituted phthalimide compound,X is a carboxyl,Y is a nitrogen,and R2 completes the phthalimide compound by being bonded to both X and Y; andwhere X, Y, and Z, may either be CH2, NH, O, or N, with the caveat that each are different from the other;a compound having the formula:where R2, R3, R4, R5and R6 may be hydrogen, halogen, or alkyl groups and may be the same or different,R7 and R8 may be hydrogen or methyl groups and may be the same or different, andwhere the R9 moiety is hydrogen or a salt cation;a compound having the formula:where R2 is a substituted or unsubstituted aromatic compound, or a substituted or unsubstituted alkyl ring compound, or a substituted or unsubstituted phthalimide compound that incorporates X and Y,X is a carbonyl,Y is a nitrogen, andR2 completes the phthalim... | 1. A method for measuring a blood oxygen level-dependent magnetic resonance imaging signal, compnsing a) administering an effective amount of an allosteric effector compound capable of decreasing hemoglobin binding affinity for oxygen; andb) performing a blood oxygen level-dependent magnetic resonance imaging scan, whereby said blood oxygen level-dependent magnetic resonance imaging signal is measured, wherein the allosteric effector compound is a compound selected from the group consisting of:a compound having the formula:where R1-5 may be hydrogen, halogen, or a substituted or unsubstituted C1-3 alkyl group and may be the same or different,R6-7 may each be hydrogen or methyl and may be the same or different, andR8 may be hydrogen, a substituted or unsubstituted C1-3 alkyl group, or a salt cation, and X and Z are CH2, NH, or O;a compound having the formula:where X and Z may each be CH2, CO, NH or O, and Y may be CO or NH, which the caveat that X, Y, and Z must all be different from each other, andR2-6 can be the hydrogen, halogen, substituted or unsubstituted C1-3 alkyl groups, and may be the same or different,R7-8 can be hydrogens, methyls, ethyls, or alkyl groups in a ring connecting the two, andR9 can be a hydrogen, lower alkyl, or salt cation;a compound having the formula:where R3-6 can be the hydrogen, halogen, substituted or unsubstituted C1-3 alkyl group, or a C1-3 ether or ester, and these moieties may be the same or different, or alkyl moieties of an aromatic or aliphatic ring incorporating two of the R3-6,R1 can be connected to any position on the phenyl ring, andsites R7-8 can be hydrogen, halogen, methyl, ethyl, and these moieties may be the same or different, or alkyl groups in a ring connecting the two, andR9 can be a hydrogen, halogen, C1-3 lower alkyl, or salt cation;a compound having the formula:where R1 can be connected to any position on the phenyl ring, andsites R7-8 can be hydrogen, halogen, methyl, ethyl, and these moieties may be the same or different, or alkyl groups in a ring connecting the two, andR2 is defined as a substituted or unsubstituted aromatic compound, a substituted or unsubstituted alkyl ring compound, or a substituted or unsubstituted phthalimide compound,X is a carboxyl,Y is a nitrogen,and R2 completes the phthalimide compound by being bonded to both X and Y; andwhere X, Y, and Z, may either be CH2, NH, O, or N, with the caveat that each are different from the other;a compound having the formula:where R2, R3, R4, R5and R6 may be hydrogen, halogen, or alkyl groups and may be the same or different,R7 and R8 may be hydrogen or methyl groups and may be the same or different, andwhere the R9 moiety is hydrogen or a salt cation;a compound having the formula:where R2 is a substituted or unsubstituted aromatic compound, or a substituted or unsubstituted alkyl ring compound, or a substituted or unsubstituted phthalimide compound that incorporates X and Y,X is a carbonyl,Y is a nitrogen, andR2 completes the phthalim... | Allos Therapeutics, Inc. | A61B005/055
US07077811 | B2 | 2006-07-18 | Guidewire tip construction | 2002-12-23 | Guidewire with an improved distal end. A guidewire with an improved distal end may include a core wire, a solder tip disposed at a distal end of the core wire, a radiopaque inner coil coupled to the distal end of the core wire, an outer coil disposed along the length of the core wire, and a distal hub coupled to the core wire and disposed proximate a distal end of the inner coil. | 1. A guidewire, comprising: an elongate shaft having a length, a proximal end and a distal end;a tip disposed at the distal end of the shaft;a radiopaque inner coil coupled to the distal end of the shaft, the inner coil having a proximal end and a distal end;an outer coil disposed along the length of the shaft, the outer coil having a proximal end and a distal end;a distal hub coupled to the shaft and disposed proximate the distal end of the inner coil; andfurther comprising a proximal hub coupled to the shaft. | 1. A guidewire, comprising: an elongate shaft having a length, a proximal end and a distal end;a tip disposed at the distal end of the shaft;a radiopaque inner coil coupled to the distal end of the shaft, the inner coil having a proximal end and a distal end;an outer coil disposed along the length of the shaft, the outer coil having a proximal end and a distal end;a distal hub coupled to the shaft and disposed proximate the distal end of the inner coil; andfurther comprising a proximal hub coupled to the shaft. 2. The guidewire in accordance with claim 1, wherein the shaft is comprised of stainless steel. 3. The guidewire in accordance with claim 1, wherein the shaft is comprised of nickel-titanium alloy. 4. The guidewire in accordance with claim 1, wherein the shaft is comprised of a composite of nickel-titanium alloy and stainless steel. 5. The guidewire in accordance with claim 1, wherein the inner coil is comprised of platinum. 6. The guidewire in accordance with claim 1, wherein the inner coil is comprised of gold. 7. The guidewire in accordance with claim 1, wherein the inner coil is comprised of tungsten alloy. 8. The guidewire in accordance with claim 1, wherein the outer coil is comprised of stainless steel. 9. The guidewire in accordance with claim 1, wherein the outer coil is comprised of nickel-titanium alloy. 10. The guidewire in accordance with claim 1, wherein the distal end of the outer coil is coupled to the tip. 11. The guidewire in accordance with claim 1, wherein the proximal end of the inner coil is coupled to the shaft. 12. The guidewire in accordance with claim 1, wherein the outer coil further comprises a proximal section that is coupled to the shaft. 13. The guidewire in accordance with claim 1, wherein the proximal hub is comprised of polytetrafluoroethylene. 14. The guidewire in accordance with claims 1, wherein the proximal hub prevents the inner coil from moving proximally relative to the shaft. 15. The guidewire in accordance with claim 1, further comprising a ribbon coupled to the shaft and extending between the shaft and the tip. 16. The guidewire in accordance with claim 1, wherein the outer coil is wound about the shaft in a first direction and the inner coil is wound about the shaft in a second direction, and wherein the first direction is substantially opposite of the second direction. 17. The guidewire in accordance with claim 1, wherein the inner coil is coupled to the outer coil. 18. A guidewire, comprising: an elongate shaft having a length, a proximal end and a distal end;a tip disposed at the distal end of the shaft;a radiopaque inner coil coupled to the distal end of the shaft, the inner coil having a proximal end and a distal end;an outer coil disposed along the length of the shaft, the outer coil having a proximal end and a distal end;a distal hub coupled to the shaft and disposed proximate the distal end of the inner coil; andwherein the distal hub is comprised of polytetrafluoroethyl... | Scimed Life Systems, Inc. | A61B005/00 A61M025/00
US07198776 | B2 | 2007-04-03 | Metal complex compounds | 2002-11-28 | The invention relates to metal complex compounds, contrast agents for MRI and in vivo NMR markers for NMR spectroscopy comprising said metal complex compounds and methods for in vivo determination of physiological parameters, e.g. enzyme activity or pH using said metal complex compounds. | 1. A metal complex compound comprising a paramagnetic chelate comprising a paramagnetic metal ion M and a chelating agent, said chelating agent is a chelating agent according to formula (I)wherein R1 represents hydrogen or C1–C15-alkyl which may optionally be substituted with one or more hydroxy groups,A and B are the same or different and represent CHR1R2, wherein R1 is of the definition as described above andR2 represents hydrogen, C1–C20-alkyl, C6–C20-aryl, C6–C20-aralkyl, said residues may optionally be substituted with one or more hydroxy groups, orA and B form together a bridge (CH2)m,Z represents NH2, NHR2, OH, O− or OR3, wherein R3 is a base equivalent or a metal ion equivalent,X represents a molecular moiety whose coordination distance to the paramagnetic metal ion chelated by the chelating agent of formula (I) changes upon influence of enzyme activity, being selected from the group of; alkyl-O—PO32− or aryl-O—PO32−which are substrates for phosphatates,alkyl-(NH)-(Glu)n which is a substrate for aminopeptidase A,4-alkyl-(C6R1R2R3R4)NH-aminoacid-NH2 wherein R1–R4 are hydrogen or F, Cl or NO2, which is a substrate for aminopeptidase,4-alkyl -(C6H4)—CO-aminoacid-CO2H which is a substrate for carboxypeptidase,4-alkyl-(C6H4)—CH2-NH3− which is a substrate for monoamine oxidase, and1-alkyl-β-O-glucoronic acid which is a substrate for β-glucoronidase,Y represents a fluorine atom or a hydrocarbon group comprising at least one fluorine atom,D represents a saturated or unsaturated straight or branched-chain hydrocarbon group containing 1 to 4 carbon atoms or a phenyl group,m represents an integer from 2 to 3 andn and o are the same or different and represent an integer from 1 to 3, wherein the coordination distance between X and M changes upon influence of enzymatic activity and thereby changing the chemical shift of the at least one fluorine atom. | 1. A metal complex compound comprising a paramagnetic chelate comprising a paramagnetic metal ion M and a chelating agent, said chelating agent is a chelating agent according to formula (I)wherein R1 represents hydrogen or C1–C15-alkyl which may optionally be substituted with one or more hydroxy groups,A and B are the same or different and represent CHR1R2, wherein R1 is of the definition as described above andR2 represents hydrogen, C1–C20-alkyl, C6–C20-aryl, C6–C20-aralkyl, said residues may optionally be substituted with one or more hydroxy groups, orA and B form together a bridge (CH2)m,Z represents NH2, NHR2, OH, O− or OR3, wherein R3 is a base equivalent or a metal ion equivalent,X represents a molecular moiety whose coordination distance to the paramagnetic metal ion chelated by the chelating agent of formula (I) changes upon influence of enzyme activity, being selected from the group of; alkyl-O—PO32− or aryl-O—PO32−which are substrates for phosphatates,alkyl-(NH)-(Glu)n which is a substrate for aminopeptidase A,4-alkyl-(C6R1R2R3R4)NH-aminoacid-NH2 wherein R1–R4 are hydrogen or F, Cl or NO2, which is a substrate for aminopeptidase,4-alkyl -(C6H4)—CO-aminoacid-CO2H which is a substrate for carboxypeptidase,4-alkyl-(C6H4)—CH2-NH3− which is a substrate for monoamine oxidase, and1-alkyl-β-O-glucoronic acid which is a substrate for β-glucoronidase,Y represents a fluorine atom or a hydrocarbon group comprising at least one fluorine atom,D represents a saturated or unsaturated straight or branched-chain hydrocarbon group containing 1 to 4 carbon atoms or a phenyl group,m represents an integer from 2 to 3 andn and o are the same or different and represent an integer from 1 to 3, wherein the coordination distance between X and M changes upon influence of enzymatic activity and thereby changing the chemical shift of the at least one fluorine atom. 2. A metal complex compound according to claim 1 wherein the paramagentic metal ion M is selected from the group consisting of divalent and trivalent ions of an element of atomic number 21 to 29, 42, 44 and 57 to 83. 3. A metal complex compound according to claim 1, wherein said chelating agent comprises at least one straight chain or branched chain alkyl group, aryl group or aralkyl group substituted by one or more fluorine atoms. 4. A metal complex compound according to claim 1, said metal complex compound further comprises a targeting vector. 5. Contrast agents or in vivo NMR markers comprising a metal complex compound according to claim 1. 6. Method of detecting abnormal physiological parameters in vivo comprising: a) administration of a contrast agent or an in vivo NMR marker comprising a metal complex compound according to claim 1 or a metal complex compound according claim 1 to a human or non-human animal body, andb) determination of abnormal physiological parameters by determining the change in fluorine chemical shift upon influence of said physiological parameters on said metal complex compo... | GE Healthcare AS | A61B005/055 C07F005/00
US07214194 | B2 | 2007-05-08 | Method for thermal diagnosis of pathology of a bioobject and device for carrying out said method | 2002-01-08 | The invention relates to the field of medical equipment and is designed for the temperature diagnostics of bioobject pathologies. The method comprises dynamical measuring a body temperature in a series of points. Based on the difference temperature field on the surface of a body, zones are determined inside the bioobject that are identified with pathology loci. A temperature is each point is estimated with taking into account a type of skin of the bioobject, a type of a temperature sensor and the form of the sensor temperature change curve in process of measurement. The device comprises temperature measurement units, temperature sensors being connected to said temperature measurement units, a RAM unit, a program realization unit, a comparator unit, a programming unit, and a visualization unit. The invention allows to increase speed and reliability of exposing pathology loci. | 1. A method for temperature diagnostics of bioobject pathologies, comprising: a) marking a bioobject surface or a body part with a plurality of regular points;b) measuring a series of temperatures by means of at least one body temperature sensor at said regular points with respect to time;c) obtaining a plurality of asymptotic temperature values at said plurality of regular points by plotting said series of temperature measurements on a curve with respect to time and obtaining from said curve an array of asymptotic temperature values;d) storing said array of asymptotic temperature values and obtaining a plurality of isotherm poles therefrom; e) from a magnitude value and sign for each of the plurality of isotherm poles, obtaining an averaged temperature value for the bioobject surface, Top, for a pathology locus of depth Z, for a nosology of said pathology locus;f) coordinating the array of asymptotic temperature values with a plurality of zones inside the bioobject body, the zones being identified with pathology loci responsible for the temperatures measured in (a) and (b), to getting coordinates of said pathology loci. | 1. A method for temperature diagnostics of bioobject pathologies, comprising: a) marking a bioobject surface or a body part with a plurality of regular points;b) measuring a series of temperatures by means of at least one body temperature sensor at said regular points with respect to time;c) obtaining a plurality of asymptotic temperature values at said plurality of regular points by plotting said series of temperature measurements on a curve with respect to time and obtaining from said curve an array of asymptotic temperature values;d) storing said array of asymptotic temperature values and obtaining a plurality of isotherm poles therefrom; e) from a magnitude value and sign for each of the plurality of isotherm poles, obtaining an averaged temperature value for the bioobject surface, Top, for a pathology locus of depth Z, for a nosology of said pathology locus;f) coordinating the array of asymptotic temperature values with a plurality of zones inside the bioobject body, the zones being identified with pathology loci responsible for the temperatures measured in (a) and (b), to getting coordinates of said pathology loci. 2. The method of claim 1, further comprising: g) iteratively repeating step (b) and obtaining a secondary array of surface temperature values with respect to time;h) obtaining an array of surface temperature difference values by subtracting the primary array from the secondary array and thereby obtaining a plurality of asymptotic temperature difference values at the plurality of regular points;i) storing said array of asymptotic temperature difference values and obtaining a plurality of isotherm poles therefrom:i) coordinating the array of asymptotic temperature difference values with a plurality of zones inside the bioobject body, each zones being identified with a pathology locus, with changes monitored in the pathology loci. 3. A method of claim 1 wherein, in case of a positive deviation values of the magnitude value and a positive sign of the temperature change took place in the isotherm pole as specified in step d) the depth of the pathology locus (Z) is determined by the criterion:Z∝r⁢2⁢⁢T2T02-T2-1,where r is the current radius counted from the isotherm pole, T0 is the temperature value at the isotherm pole, T is the current temperature value in polar coordinates with the center of the isotherm pole. 4. A method of claim 1 wherein, in case of a negative deviation values of the magnitude value and a negative sign of the temperature change took place in the isotherm pole as specified in step d) the depth of the pathology locus (Z) is determined by the criterion:Z∝r2(T0T)23-1,where r is the current radius counted from the isotherm pole. T0 is the temperature value at the isotherm pole, T is the current temperature value in polar coordinates with the center of the isotherm pole. 5. A method of claim 1, wherein, in case of the most dangerous pathology locus where temperature is higher than an average temperature Tcp, an aver... | - | A61B005/00
US06983185 | B2 | 2006-01-03 | Lead with terminal connector assembly | 2002-07-31 | A lead includes a lead body extending from a distal end to a proximal end, and at least one conductor disposed within the lead body and extending from the distal end to the proximal end of the lead body. The lead further includes an outer terminal ring, a terminal pin, and an insulative sleeve disposed between the outer terminal ring and the terminal pin, where the insulative sleeve is coupled with the outer terminal ring with a snap-fit connection. A pin latch is disposed on an outer periphery surface of the insulative sleeve, where the pin latch is rotatable about a hinge point. | 1. A lead comprising: a lead body extending from a distal end to a proximal end; at least one conductor disposed within the lead body and extending from the distal end to the proximal end of the lead body; an outer terminal ring coupled with the lead body; a terminal pin disposed at the proximal end of the lead body; and an insulative sleeve disposed between the outer terminal ring and the terminal pin, where the insulative sleeve is coupled with the outer terminal ring; and a pin latch disposed on an outer periphery surface of the insulative sleeve, the pin latch rotatable about a hinge point, the pin latch remains in a deflected position after the sleeve is coupled with the ring. | 1. A lead comprising: a lead body extending from a distal end to a proximal end; at least one conductor disposed within the lead body and extending from the distal end to the proximal end of the lead body; an outer terminal ring coupled with the lead body; a terminal pin disposed at the proximal end of the lead body; and an insulative sleeve disposed between the outer terminal ring and the terminal pin, where the insulative sleeve is coupled with the outer terminal ring; and a pin latch disposed on an outer periphery surface of the insulative sleeve, the pin latch rotatable about a hinge point, the pin latch remains in a deflected position after the sleeve is coupled with the ring. 2. The lead as recited in claim 1, wherein the sleeve and/or the ring include at least one anti-rotation feature. 3. The lead as recited in claim 2, wherein the anti-rotation features comprise axial grooves disposed on an inner surface of the outer terminal ring. 4. The lead as recited in claim 2, wherein the axial grooves comprise V-shaped grooves. 5. The lead as recited in claim 2, wherein the anti-rotation features comprise a flat formed on the terminal pin. 6. The lead as recited in claim 2, wherein the anti-rotation features comprise axial grooves disposed on an outer surface of the terminal pin. 7. A lead comprising: a lead body extending from a distal end to a proximal end; at least one conductor disposed within the lead body and extending from the distal end to the proximal end of the lead body; an outer terminal ring coupled with the lead body; a terminal pin disposed at the proximal end of the lead body; and an insulative sleeve disposed between the outer terminal ring and the terminal pin; and a pin latch disposed on an outer periphery surface of the insulative sleeve, the pin latch rotatable about a hinge point to deflect and couple the insulative sleeve with the terminal pin, the pin latch remains in a deflected position after the sleeve is coupled with the pin. 8. The lead as recited in claim 7, further comprising inner tubing disposed between the sleeve and the terminal pin. 9. The lead as recited in claim 8, wherein the insulative sleeve is coupled with the outer terminal ring with a snap-fit coupling. 10. The lead as recited in claim 7, wherein an inner surface of the sleeve is tapered from a first end to a second end. 11. The lead as recited in claim 7, wherein the insulative sleeve is formed of polyetheretherketone. 12. The lead as recited in claim 7, wherein the sleeve is longer than the terminal pin. 13. The lead as recited in claim 7, further comprising a means for preventing the insulative sleeve from rotating relative to the terminal pin and/or the outer terminal ring. 14. A lead comprising: a lead body extending from a distal end to a proximal end; at least one conductor disposed within the lead body and extending from the distal end to the proximal end of the lead body; an outer terminal ring coupled with the lead body; a termin... | Cardiac Pacemakers, Inc. | A61N001/05 A61B005/0408
US07288082 | B2 | 2007-10-30 | Method and apparatus for sterilely acquiring and separating a fluid | 2002-10-17 | A method and apparatus are provided by which a fluid may be drawn and packaged within a series of interconnected sterile units. Once filled with the fluid, the connectors joining the sterile units may be sealed and severed to produce a number of separate sterile units. The volume of the separate sterile units may correspond to a single dose of the fluid. The technique allows a fluid to be partitioned into sterile units without exposure to the air or other potential contaminants. | 1. A method for sterilely collecting fluid comprising: providing an interconnected sterile assembly comprising an upstream compartment, at least one intermediate compartment joined by a first connector to the upstream compartment, and a downstream compartment joined by a second connector to the intermediate compartment; andfilling the interconnected sterile assembly with a fluid, wherein the fluid passes through the upstream compartment and first connector and the at least one intermediate compartment and second connector to fill the downstream compartment, said filling the interconnected sterile assembly with a fluid accomplished without disconnecting any of the compartments from one another, wherein filling the interconnected sterile assembly with the fluid comprises operating a pressure regulator coupled to the downstream compartment such that pressure within the interconnected sterile assembly can be adjusted. | 1. A method for sterilely collecting fluid comprising: providing an interconnected sterile assembly comprising an upstream compartment, at least one intermediate compartment joined by a first connector to the upstream compartment, and a downstream compartment joined by a second connector to the intermediate compartment; andfilling the interconnected sterile assembly with a fluid, wherein the fluid passes through the upstream compartment and first connector and the at least one intermediate compartment and second connector to fill the downstream compartment, said filling the interconnected sterile assembly with a fluid accomplished without disconnecting any of the compartments from one another, wherein filling the interconnected sterile assembly with the fluid comprises operating a pressure regulator coupled to the downstream compartment such that pressure within the interconnected sterile assembly can be adjusted. 2. The method as recited in claim 1, further comprising storing the fluid within the interconnected sterile assembly. 3. The method as recited in claim 1, further comprising forming an airtight seal within the connectors connected to each compartment. 4. The method as recited in claim 3, wherein forming the airtight seal within the connectors connected to each compartment comprises at least one of heating and compressing the connectors to form the airtight seal. 5. The method as recited in claim 3, further comprising disconnecting each compartment at each seal such that three or more sterile storage compartments are formed from the disconnected sterile assembly. 6. The method as recited in claim 5, further comprising storing the fluid within the three or more sterile storage compartments. 7. The method as recited in claim 5, further comprising applying a dose of the fluid from one of the sterile storage compartments. 8. The method as recited in claim 1, further comprising drawing a patient's blood and wherein filling the interconnected sterile assembly with the fluid comprises filling the interconnected sterile assembly with the patient's blood. 9. The method as recited in claim 1, further comprising drawing a medication from a sterile stock supply and wherein filling the interconnected sterile assembly with the fluid comprises filling the interconnected sterile assembly with the medication. 10. The method as recited in claim 1, wherein filling the interconnected sterile assembly with the fluid comprises applying positive pressure to fill the interconnected sterile assembly with the fluid. 11. The method as recited in claim 1, wherein filling the interconnected sterile assembly with the fluid comprises applying negative pressure to fill the interconnected sterile assembly with the fluid. 12. An apparatus for storing a fluid comprising: an interconnected sterile assembly comprising: an upstream compartmentat least one intermediate compartment joined by a first connector to the upstream compartment;a downstream compartment join... | SeeFit Incorporated | A61B019/00 B65D081/00 A61B005/00
US07010338 | B2 | 2006-03-07 | Device for locating magnetic implant by source field | 2003-01-06 | An apparatus and method for locating a magnetic implant in a surgical application using the field of a source magnet for the implant guiding field. The source magnet is an electromagnet having a separate calibrated magnetic field component in addition to the guiding field, so that both the magnitude and orientation of the magnetic field as a function of position around the magnet are known. A magnetic implant is provided with a sensor, such as a three-axis Hall effect sensor, to provide an indication of the magnitude and orientation of an applied magnetic field when the implant is surgically implanted in a patient. After implantation, the source magnet is energized with a current having a modulated component. The modulated component is received and filtered from the signal received from the Hall effect sensor in the implant, and provided to a processor that computes the location of the implant. | 1. A magnetic surgical implant in combination with a device for providing position data for the magnetic surgical implant having a magnetic field sensor which is guided or propelled through a patient's body by an external electromagnet, said device comprising a current source for supplying a modulated electrical current to the external electromagnet to thereby generate an oscillating magnetic field component, a demodulator coupled to the magnetic field sensor and responsive to a magnetic field direction and magnitude of the oscillating magnetic field component to provide a signal indicative of the oscillating magnetic field component at the location of the magnetic field sensor, a memory containing a representation of a relationship of spatial locations to a magnetic field pattern produced by the external magnet, and a data processor coupled to the demodulator and the memory and configured to compute a spatial location of the magnetic field sensor as a function of the signal indicative of the oscillating magnetic field component. | 1. A magnetic surgical implant in combination with a device for providing position data for the magnetic surgical implant having a magnetic field sensor which is guided or propelled through a patient's body by an external electromagnet, said device comprising a current source for supplying a modulated electrical current to the external electromagnet to thereby generate an oscillating magnetic field component, a demodulator coupled to the magnetic field sensor and responsive to a magnetic field direction and magnitude of the oscillating magnetic field component to provide a signal indicative of the oscillating magnetic field component at the location of the magnetic field sensor, a memory containing a representation of a relationship of spatial locations to a magnetic field pattern produced by the external magnet, and a data processor coupled to the demodulator and the memory and configured to compute a spatial location of the magnetic field sensor as a function of the signal indicative of the oscillating magnetic field component. 2. The combination of claim 1, wherein the representation of the relationship of spatial locations to the magnetic field pattern produced by the electromagnet is a representation in a reference frame, the electromagnet is moveable, and the device further comprises a localizer configured to provide a signal representative of a position and orientation of the electromagnet to the processor, and the processor is configured to apply a correction based upon the position and orientation provided by the localizer. 3. The combination of claim 2 wherein the localizer comprises a sensor mounted on the electromagnet and a first plurality of fiducial markers adapted to be positioned on the patient and a second plurality of fiducial markers positioned on the electromagnet. 4. The combination of claim 1 and further comprising a display coupled to the processor, the processor further being configured to display on said display an image representing the spatial location of the magnetic field sensor and a preoperative image of a patient in which the magnetic field sensor is implanted, the location image and the preoperative image being superimposed on one another. 5. The combination of claim 1, in which the electromagnet is configured to generate a hemispherically symmetrical field. 6. The combination of claim 1, in which the electromagnet is configured to generate a hemispherically asymmetric field. 7. The combination of claim 1 in which the electromagnet comprises a plurality of separate coils. | Stereotaxis, Inc. | A61B005/00
US07117036 | B2 | 2006-10-03 | Using activity-based rest disturbance as a metric of sleep apnea | 2002-06-27 | An implantable cardiac device is programmed to monitor short term activity changes that occur while a patient is at rest to produce a sleep disturbance metric that is useful in analyzing and/or treating sleep apnea. After the implantable cardiac device confirms that a patient is at rest, the device monitors an instantaneous signal from an activity sensor to detect variances from normal rest mode activity. When the variances exceed a preset threshold for a short time period (e.g., less than 30–40 sec.), the patient is presumed to be experiencing a form of sleep disturbance as opposed to conscious or wakeful activity. These short term events are recorded as sleep disturbance events. The sleep disturbance metric are reported to a physician as a diagnostic to help ascertain the severity of sleep apnea or to evaluate the effectiveness of pacing therapies being applied to treat sleep apnea. | 1. A method implemented by an implantable cardiac device, comprising: detecting a heightened level of patient activity while the patient is at rest;timing a duration of the heightened level of patient activity;in an event that the duration is less than a predefined time limit, identifying the heightened level of patient activity as sleep disturbance;administering multiple different pacing therapies to treat sleep apnea; andevaluating the pacing therapies based on how the pacing therapies affect the sleep disturbances. | 1. A method implemented by an implantable cardiac device, comprising: detecting a heightened level of patient activity while the patient is at rest;timing a duration of the heightened level of patient activity;in an event that the duration is less than a predefined time limit, identifying the heightened level of patient activity as sleep disturbance;administering multiple different pacing therapies to treat sleep apnea; andevaluating the pacing therapies based on how the pacing therapies affect the sleep disturbances. 2. A method as recited in claim 1, wherein the detecting comprises determining when a patient is in a reclined position and monitoring the patient's activity while the patient is in the reclined position. 3. A method as recited in claim 1, wherein the detecting comprises monitoring an instantaneous signal from an activity sensor. 4. A method as recited in claim 1, wherein the time limit is approximately forty seconds. 5. A method as recited in claim 1, further comprising counting the sleep disturbances. 6. A method as recited in claim 1, further comprising storing at least one of a count or a frequency of the sleep disturbances. 7. A computer readable medium having computer-executable instructions that, when executed, direct an implantable cardiac device to perform the method as recited in claim 1. 8. A method comprising: determining that a patient is at rest;while the patient is at rest,monitoring an instantaneous activity signal indicative of a patient's activity;detecting when the instantaneous activity signal exceeds a threshold;timing a duration that the instantaneous activity signal exceeds the threshold;in an event that the instantaneous activity signal exceeds the threshold for less than a preset time, recording the patient's activity as a sleep disturbance;administering multiple different pacing therapies to treat sleep apnea;collecting sleep disturbance metrics for individual pacing therapies; andevaluating the pacing therapies based on how the pacing therapies affect the sleep disturbance metrics. 9. A method as recited in claim 8, further comprising determining a frequency of sleep disturbances that occur within a predetermined timeframe. 10. A method as recited in claim 8, further comprising counting the sleep disturbances. 11. A method as recited in claim 8, further comprising identifying the patient's activity as meaningful activity if the instantaneous activity signal exceeds the threshold for more than the preset time. 12. A computer readable medium having computer-executable instructions that, when executed, direct an implantable cardiac device to perform the method as recited in claim 8. 13. An implantable cardiac device comprising: sensing circuitry to sense a patient's activity;a sleep disturbance detector to detect when a patient, who is at rest, is experiencing sleep disturbance based on the patient's activity;a therapy module to prescribe a pacing therapy for treating sleep apnea from among multip... | Pacesetter, Inc. | A61N001/365 A61N001/37 A61B005/11
US07076309 | B2 | 2006-07-11 | Apparatus and method for fixedly engaging an electrical lead | 2002-04-30 | An electrical lead includes an elongate body having a proximal end and a fixation member joined to the elongate body and distally disposed from the proximal end of the elongate body, wherein the fixation member is capable of being extended from and retracted toward the elongate body. A method includes advancing an electrical lead into vasculature and extending a fixation member from the body of the electrical lead to fixedly engage the electrical lead within the vasculature. The method may further include retracting the fixation member toward the body of the electrical lead to unengage the electrical lead within the vasculature and retracting the electrical lead from the vasculature. | 1. A lead of an implantable medical device, comprising: an elongated body having a proximal end;a fixation member joined to the elongated body and distally disposed from the proximal end of the elongated body, wherein the fixation member is capable of being extended radially outward from and retracted inward toward the elongated body to provide fixation of a distal part of the elongated body within a portion of a great coronary vein of a subject:a locking mechanism for selectively restricting the fixation member from being extended radially outward from and retracted radially inward toward the elongated body, the locking mechanism comprising: a pin disposed in the fixation member;a plurality of bores disposed in the elongated body such that, as the fixation member is extended outward from and retracted inward toward the elongated body, the pin is capable of slidingly engaging one of the plurality of bores;a spring disposed in the fixation member; anda lever disposed between the spring and the pin such that, as the lever is moved in a first direction, the spring is compressed and the pin is retracted from the bore and, as the lever is moved in a second direction opposite to the first direction, the spring is uncompressed and the pin is extended into the bore. | 1. A lead of an implantable medical device, comprising: an elongated body having a proximal end;a fixation member joined to the elongated body and distally disposed from the proximal end of the elongated body, wherein the fixation member is capable of being extended radially outward from and retracted inward toward the elongated body to provide fixation of a distal part of the elongated body within a portion of a great coronary vein of a subject:a locking mechanism for selectively restricting the fixation member from being extended radially outward from and retracted radially inward toward the elongated body, the locking mechanism comprising: a pin disposed in the fixation member;a plurality of bores disposed in the elongated body such that, as the fixation member is extended outward from and retracted inward toward the elongated body, the pin is capable of slidingly engaging one of the plurality of bores;a spring disposed in the fixation member; anda lever disposed between the spring and the pin such that, as the lever is moved in a first direction, the spring is compressed and the pin is retracted from the bore and, as the lever is moved in a second direction opposite to the first direction, the spring is uncompressed and the pin is extended into the bore. 2. A lead, according to claim 1, further comprising: a stylet member;a fitting located within the fixation member capable of mating with the stylet; anda lumen, positioned in the elongated body, extending from the proximal end of the elongated body to the fittng of the fixation member, wherein the stylet is fed through the lumen and mated with the fitting of the fixation member. 3. A lead, according to claim 1, further comprising an electrical conductor extending from the proximal end to the fixation member, wherein the electrical conductor is capable of being used to extend the fixation member. 4. A lead, according to claim 1, wherein the elongated body includes a distal end and the fixation member is disposed proximate the distal end. 5. A lead, according to claim 1, further comprising a conductor extending within the elongated body and electrically coupled with the fixation member such that electrical stimuli may be transmitted through the conductor and emitted by the fixation member. 6. A lead, according to claim 1, further comprising a bladder positioned in the spring, the bladder containing silicone oil and having an air bubble disposed within the silicone oil. 7. A lead, according to claim 1, further comprising: a stylet insertable within a lumen of the lead;a fitting, positioned within the fixation member, capable of mating with the stylet; anda lumen extending from the proximal end of the elongated body to the fitting of the fixation member such that the stylet is fed through the lumen to move the lever and mate with the filling of the fixation member. 8. A lead, according to claim 1, further comprising a guide wire lumen positioned within the elongated body. 9. A lead, accor... | Medtronic, Inc. | A61N001/05 A61B005/042
US07155276 | B2 | 2006-12-26 | Method and apparatus to modify cranial electrical potentials to remediate psychiatric disorders and to enhance optimal brain functioning | 2002-07-17 | A method for modifying cranial electrical potentials utilizes an electrical lead having first and second electrodes. The electrical lead may also include either a resistive component for increasing the resistance of the electrical lead, or a diode for restricting the flow of electricity through the lead to a single direction. The electrical lead is attached between skin surface areas of a human head, such as between ears, or an ear and an area of the scalp or forehead. Preferably, two electrical leads are utilized, one lead extending from a left ear to a right portion of the head, and the other electrical lead extending from the right ear to the left portion of the head. Scalp electrical potentials are altered in a passive manner resulting in beneficial effects on a variety of psychiatric disorders and brain functioning. | 1. A method for modifying body electrical potentials, comprising the steps of: providing first and second electrical leads each having first and second electrodes;increasing the resistance of the leads using a resistive component;attaching the first electrode of the first lead to a skin surface area of a human head;attaching the second electrode of the first lead to another skin surface area on an opposite side of the human head;attaching the first electrode of the second lead to a skin surface area on the opposite side of the head as the first electrode of the first lead;attaching the second electrode of the second lead to another skin surface area on an opposite side of the human head as the first electrode of the second lead; andpassively dissipating electrical potentials of the head to treat psychiatric disorders and brain functioning. | 1. A method for modifying body electrical potentials, comprising the steps of: providing first and second electrical leads each having first and second electrodes;increasing the resistance of the leads using a resistive component;attaching the first electrode of the first lead to a skin surface area of a human head;attaching the second electrode of the first lead to another skin surface area on an opposite side of the human head;attaching the first electrode of the second lead to a skin surface area on the opposite side of the head as the first electrode of the first lead;attaching the second electrode of the second lead to another skin surface area on an opposite side of the human head as the first electrode of the second lead; andpassively dissipating electrical potentials of the head to treat psychiatric disorders and brain functioning. 2. The method of claim 1, wherein the first or second electrode comprises a conductive clip. 3. The method of claim 1, wherein the first or second electrode comprises a conductive flat or cup-shaped electrode. 4. The method of claim 1, including the step of restricting the flow of electricity through the lead to a single direction using a diode. 5. The method of claim 1, wherein the first electrical lead extends from a left ear to a portion of the scalp or forehead right of a midline of the head, and the second electrical lead extends from a right ear to a portion of the scalp or forehead left of the midline. 6. The method of claim 1, including a retentive member for securely holding the second electrodes in place on the scalp or forehead. 7. A method for modifying body electrical potentials, comprising the steps of: attaching a first electrode of a first electrical lead to a left ear of a human head, and attaching a second electrode of the first electrical lead to a skin surface area of the scalp or forehead right of a midline of the human head;attaching a first electrode of a second electrical lead to a right ear of a human head, and attaching a second electrode of the second electrical lead to a skin surface area of the scalp or forehead left of the midline of the human head; andpassively dissipating electrical potentials of the head to treat psychiatric disorders and brain functioning. 8. The method of claim 7, wherein the first electrode of the first and second leads comprises a conductive clip. 9. The method of claim 7, wherein the first or second electrode comprises a conductive flat or cup-shaped electrode. 10. The method of claim 7, including a retentive member for securely holding the second electrodes in place on the scalp or forehead. 11. The method of claim 7, including the step of increasing the resistance of the first and second electrical leads using a resistive component. | - | A61B005/0482
US07103407 | B2 | 2006-09-05 | Body fat monitoring system and method employing mobile terminal | 2002-06-28 | A system and method measures a parameter of body fat using body fat monitoring circuitry provided with a mobile communications device or device cover. A mobile terminal device includes a housing supporting a display, a processor, and a user interface for facilitating user interaction with the mobile terminal device. Body fat monitoring circuitry of the mobile terminal device includes a tactile interface supported by one or both of the housing and a cover of the housing. The monitoring circuitry communicates monitoring signals to a portion of a user's body via the tactile interface. The processor computes a resistance to the communication of the monitoring signals and computes the user's body fat parameter using the computed resistance. An acceleration sensor may be included to provide various pedometer functions. | 1. A mobile terminal device, comprising: a housing supporting a display;a cover detachably coupled to the housing;a processor;a user interface for facilitating user interaction with the mobile terminal device, the cover configured to cover at least a portion of the display and the user interface when coupled to the housing; andbody fat monitoring circuitry comprising a tactile interface, at least the tactile interface supported by the housing cover, the monitoring circuitry communicating monitoring signals to a portion of a user's body via the tactile interface, the processor computing a parameter associated with the user's body fat in response to the communication of the monitoring signals. | 1. A mobile terminal device, comprising: a housing supporting a display;a cover detachably coupled to the housing;a processor;a user interface for facilitating user interaction with the mobile terminal device, the cover configured to cover at least a portion of the display and the user interface when coupled to the housing; andbody fat monitoring circuitry comprising a tactile interface, at least the tactile interface supported by the housing cover, the monitoring circuitry communicating monitoring signals to a portion of a user's body via the tactile interface, the processor computing a parameter associated with the user's body fat in response to the communication of the monitoring signals. 2. The device of claim 1, wherein the processor computes a resistance to the communication of the monitoring signals and computes the user's body fat parameter using the computed resistance. 3. The device of claim 1, wherein the body fat parameter comprises percent body fat. 4. The device of claim 1, wherein the body fat parameter comprises one or both of fat mass and total body water. 5. The device of claim 1, wherein the tactile interface comprises a plurality of electrodes supported by the housing cover. 6. The device of claim 1, wherein the tactile interface comprises a plurality of electrodes supported by a common surface of the housing cover. 7. The device of claim 1, wherein the tactile interface comprises a plurality of electrodes supported by two or more surfaces of the housing cover. 8. The device of claim 1, wherein the monitoring signals produced by the monitoring circuitry comprise sinusoidal current or voltage signals. 9. The device of claim 1, wherein the monitoring signals produced by the monitoring circuitry comprise square wave current or voltage signals. 10. The device of claim 1, wherein the monitoring signals produced by the monitoring circuitry comprise a current or voltage signal having a frequency of about 50 kHz or greater. 11. The device of claim 1, wherein the monitoring signals produced by the monitoring circuitry comprise a current signal having an amperage less than about 1 mA. 12. The device of claim 1, wherein the monitoring circuitry maintains the monitoring signals at a substantially constant amperage or voltage. 13. The device of claim 1, wherein the user interface facilitates input of user information, the user information comprising one or more of gender, body weight, age, and height of the user, the processor using the user information to compute the parameter associated with the user's body fat. 14. The device of claim 1, wherein the tactile interface comprises a plurality of electrodes for communicating the monitoring signals to and from the portion of the user's body, and one or both of the housing and housing cover comprises an interface configured for communicating electrical signals between the processor and a device external to the mobile terminal device. 15. The device of claim 1, wherein the mobile ... | Nokia Corporation | A61B005/05
US07231243 | B2 | 2007-06-12 | Optical methods for tissue analysis | 2001-10-30 | The invention relates to methods and systems to optically analyze samples such as tissue based on speckle patterns of microscopic motion, such as Brownian motion. | 1. A method of analyzing tissue, the method comprising: illuminating a tissue with coherent or partially coherent light;receiving light reflected from the tissue at a detector and forming series of speckle patterns;analyzing changes in the speckle patterns at time intervals sufficient to measure changes caused by microscopic motion of objects within the tissue, wherein the tissue is at least one of in vivo or an internal tissue; andcompensating for macroscopic motion to isolate the microscopic motion. | 1. A method of analyzing tissue, the method comprising: illuminating a tissue with coherent or partially coherent light;receiving light reflected from the tissue at a detector and forming series of speckle patterns;analyzing changes in the speckle patterns at time intervals sufficient to measure changes caused by microscopic motion of objects within the tissue, wherein the tissue is at least one of in vivo or an internal tissue; andcompensating for macroscopic motion to isolate the microscopic motion. 2. The method of claim 1, wherein the microscopic motion is Brownian motion. 3. The method of claim 1, wherein the microscopic motion is motion of cells or cellular organelles. 4. The method of claim 1, wherein the illuminating step comprises providing an invasive device coupled to a light source, passing the device into a patient, placing the device in proximity to the tissue, and shining coherent or partially coherent light from the light source onto the tissue. 5. The method of claim 4, wherein the invasive device is selected from the group consisting of a catheter, an endoscope, and a laparoscope. 6. The method of claim 4, wherein the placing step includes placing the device in direct contact with the tissue. 7. The method of claim 1, wherein the coherent light comprises laser light. 8. The method of claim 1, wherein a detector is located farther than one wavelength of light from the tissue and detects far field speckle. 9. The method of claim 1, wherein a detector is located within one wavelength of light from the tissue and detects near field speckle. 10. The method of claim 1, wherein the analyzing step comprises comparing each of the series of speckle patterns to a series of reference speckle patterns, and quantifying the temporal correlation differences between the speckle patterns and the reference patterns. 11. The method of claim 10, wherein the analyzing step comprises digitizing each of the speckle patterns, and the quantifying step comprises evaluating a cross-correlation between the speckle patterns and the reference patterns. 12. The method of claim 11, wherein the analyzing step further comprises determining a decorrelation rate for the speckle patterns. 13. The method of claim 10, wherein the analyzing step comprises digitizing each of the speckle patterns, and the quantifying step comprises evaluating a maximum cross-correlation between the speckle patterns and the reference patterns. 14. The method of claim 1, wherein the analyzing step further comprises analyzing spatial characteristics of the speckle pattern to deduce structural characteristics of the tissue. 15. The method of claim 14, wherein the illuminating step comprises illuminating multiple locations of the tissue in succession, the receiving step comprises forming a separate series of speckle patterns for each respective section of the tissue, and the analyzing step comprises analyzing each separate series of speckle patterns and comparing the separate se... | The General Hospital Corporation | A61B005/00
US07228162 | B2 | 2007-06-05 | Analyte sensor | 2003-01-13 | A method of measuring an analyte concentration in body fluid in an animal body having skin and subcutaneous soft tissue that includes body fluid. The method comprises, introducing a portion of an analyte-measuring device into the subcutaneous soft tissue. The analyte-measuring device includes an analyte sensing element, having a sharpened distal end to facilitate introduction into the animal body. Additionally the sensing element has an indicating electrode that is adapted to be activated by electric power and to form a raw analyte measurement, when it is activated and in contact with the body fluid, in less than two minutes. The sensing device further has a reference electrode and an electric power and display device that is adapted to mate to the analyte-sensing element. The power and display device activates the analyte-sensing element by applying electric power to it. The analyte-sensing element is introduced into the animal body and is activated by the electric power and display device, thereby causing the analyte-sensing element to form a raw analyte measurement. The power and display device forms and displays a second analyte measurement that is related to the raw analyte measurement, within two minutes of the analyte sensing element being introduced into the body. | 1. A method of measuring an analyte concentration in body fluid in an animal body having skin and subcutaneous soft tissue that includes body fluid, said method comprising: (a) providing an analyte measuring device, including: (i) an analyte sensing element having an indicating electrode, said indicating electrode being adapted to be activated by electric power, and to form a raw analyte measurement when so activated and in contact with said body fluid; and(ii) an electric power, data processing and display device adapted to mate to said analyte sensing element and activate said analyte sensing element by applying electric power to it and adapted to receive said raw analyte measurement and to compute and display a refined analyte measurement, related to said raw analyte measurement;(b) introducing said analyte sensing element into said animal body subcutaneous soft tissue, thereby placing said indicating electrode into contact with said animal body subcutaneous soft tissue and said body fluid;(c) activating said analyte sensing element by applying electric power to it, thereby causing said analyte sensing element to form a raw analyte measurement;(d) removing said analyte sensing element from said soft tissue after the raw analyte measurement is formed; and(e) wherein said analyte sensing element retains body fluid in contact with said indicating electrode after said analyte sensing element is removed from said body, and wherein said analyte sensing element is removed from said body prior to a refined analyte measurement being formed. | 1. A method of measuring an analyte concentration in body fluid in an animal body having skin and subcutaneous soft tissue that includes body fluid, said method comprising: (a) providing an analyte measuring device, including: (i) an analyte sensing element having an indicating electrode, said indicating electrode being adapted to be activated by electric power, and to form a raw analyte measurement when so activated and in contact with said body fluid; and(ii) an electric power, data processing and display device adapted to mate to said analyte sensing element and activate said analyte sensing element by applying electric power to it and adapted to receive said raw analyte measurement and to compute and display a refined analyte measurement, related to said raw analyte measurement;(b) introducing said analyte sensing element into said animal body subcutaneous soft tissue, thereby placing said indicating electrode into contact with said animal body subcutaneous soft tissue and said body fluid;(c) activating said analyte sensing element by applying electric power to it, thereby causing said analyte sensing element to form a raw analyte measurement;(d) removing said analyte sensing element from said soft tissue after the raw analyte measurement is formed; and(e) wherein said analyte sensing element retains body fluid in contact with said indicating electrode after said analyte sensing element is removed from said body, and wherein said analyte sensing element is removed from said body prior to a refined analyte measurement being formed. 2. The method of claim 1, wherein said analyte measuring device further includes a housing and wherein said analyte sensing element is adapted to be withdrawn into said housing and placed therein in a state wherein reuse is impractical. 3. The method of claim 1, wherein said analyte sensing element further includes a sharpened distal and said indicating electrode and said sharpened distal end are physically coincident. 4. The method of claim 1, wherein said analyte sensing element is in the form of a sharpened tube and said indicating electrode is inside said tube. 5. The method of claim 4, wherein at least a portion of said sharpened tube serves as a reference electrode. 6. The method of claim 1, wherein said analyte sensing element includes a reference electrode that is adapted to abut the skin during the formation of an analyte measurement. 7. The method of claim 1, wherein said analyte measuring device includes a housing and wherein said analyte sensing element and said electric power, data processing and display device are operatively connected to said housing. 8. The method of claim 1, wherein said analyte measuring device includes a housing having a housing aperture and at least a first analyte sensing element and a second analyte sensing element and wherein before said analyte measuring device is ever used, both said first analyte sensing element and said second analyte sensing element are stored in a... | iSense Corporation | A61B005/05 A61B005/00
US07248906 | B2 | 2007-07-24 | Method and device for monitoring analyte concentration by optical detection | 2002-12-12 | A method and device are provided for the determination of the concentration of compounds in body tissue. The method utilises optical methods based on the interaction of light with compounds, whereby the concentration of the compound under analysis is determined. The method is especially suited for analysis for the concentration of glucose in blood or tissue of diabetic patients, a device being implanted underneath the skin of the patient and the method being carried out by using the implanted device. The device contains photo detectors at different levels connected by wires to an electronic circuit device. A differential analysis is performed on the signals from the detectors to reduce the effect of skin on the analysis. | 1. A device for implantation beneath the skin of a living creature, the device having outer surfaces of biocompatible material and comprising: detection means for detection of light transmitted through the skin, said detection means comprising: at least one substantially flat first area including a first photodetector, and at least one substantially flat second area including a second photodetector, the arrangement of the first and second areas being such that light transmitted through the skin to the device interacts with compounds and tissues over a different distance between the skin and the first area than between the skin and the second area; andtransmitting means for transmitting signals derived from said detection means to an external device. | 1. A device for implantation beneath the skin of a living creature, the device having outer surfaces of biocompatible material and comprising: detection means for detection of light transmitted through the skin, said detection means comprising: at least one substantially flat first area including a first photodetector, and at least one substantially flat second area including a second photodetector, the arrangement of the first and second areas being such that light transmitted through the skin to the device interacts with compounds and tissues over a different distance between the skin and the first area than between the skin and the second area; andtransmitting means for transmitting signals derived from said detection means to an external device. 2. The device according to claim 1, wherein the detection means includes first and second groups of corresponding first and second areas the first group of areas forms a common first level, the second group of areas forms a common second level, and the common first and second levels have a predetermined spacing from each other to provide the difference in distance over which light interacts with compounds and tissue between the skin and the first and second areas. 3. The device according to claim 2, wherein the first group of areas is formed at a base level of the device and the second group of areas is formed by projections from the said base level to a top level. 4. The device according to claim 3, wherein the detection means further includes at least one substantially flat third area including a third photodetector, forming a common level between the base level and the top level. 5. The device according to claim 1, further including a spacer, covering at least one second area so that compounds and tissue are unable to fill a volume defined between the spacer and the at least one second area, the volume between the spacer and the at least one second area providing the difference in distance over which light interacts with compounds and tissue. 6. The device according to claim 1, wherein at least a part of one of said first and second areas is covered by a permeable membrane. 7. The device according to claim 6, wherein the membrane is permeable to glucose. 8. A method for optically detecting the content of a compound in the body of a living creature, the method comprising: directing an external light source on different areas of an implanted device containing detection means, the distance, between the skin and said detection means, over which light from the light source interacts with compounds and tissue in respect of one area differing from the distance, between the skin and said detection means, over which light from the light source interacts with compounds and tissue in respect of another area;obtaining light-representing signals representing the light directed at the different areas of said detection means; andobtaining a value for the content of the compound by analysing the light-repre... | Danfoss A/S | A61B005/00
US07029449 | B2 | 2006-04-18 | Arteriosclerosis inspecting apparatus | 2003-02-03 | An arteriosclerosis inspecting apparatus including an information obtaining device which obtains information that is related to a velocity at which a pulse wave propagates in a living subject, a blood-pressure measuring device which measures a blood pressure of the subject, a heart-rate measuring device which measures a heart rate of the subject, a pre-ejection-period measuring device which measures a pre-ejection period from a time of starting of contraction of the heart of the subject to a time of starting of ejection of blood from the heart, an ejection-time measuring device which measures an ejection time from the time of starting of ejection of blood from the heart to a time of ending of ejection of blood from the heart, and an arteriosclerosis-inspection-related-augmentation-index determining means for determining an arteriosclerosis-inspection-related augmentation index of the subject, based on the obtained pulse-wave-velocity-related information, the measured blood pressure, the measured heart rate, the measured pre-ejection period, and the measured ejection time, according to a predetermined relationship between (A) (a1) pulse-wave-velocity-related information, (a2) blood pressure, (a3) heart rate, (a4) pre-ejection period and (a5) ejection time, and (B) arteriosclerosis-inspection-related augmentation index. | 1. An apparatus for inspecting arteriosclerosis of a living subject, comprising: a pulse-wave-velocity-related-information obtaining device which obtains pulse-wave-velocity-related information that is related to a velocity at which a pulse wave propagates in the subject;a blood-pressure measuring device which measures a blood pressure of the subject;a heart-rate measuring device which measures a heart rate of the subject;a pre-ejection-period measuring device which measures a pre-ejection period from a time of starting of contraction of the heart of the subject to a time of starting of ejection of blood from the heart;an ejection-time measuring device which measures an ejection time from the time of starting of ejection of blood from the heart to a time of ending of ejection of blood from the heart; andan arteriosclerosis-inspection-related-augmentation-index determining means for determining an arteriosclerosis-inspection-related augmentation index of the subject, based on the pulse-wave-velocity-related information obtained by the pulse-wave-velocity-related-information obtaining device, the blood pressure measured by the blood-pressure measuring device, the heart rate measured by the heart-rate measuring device, the pre-ejection period measured by the pre-ejection-period measuring device, and the ejection time measured by the ejection-time measuring device, according to a predetermined relationship between (A) (a1) pulse-wave-velocity-related information, (a2) blood pressure, (a3) heart rate, (a4) pre-ejection period and (a5) ejection time, and (B) arteriosclerosis-inspection-related augmentation index. | 1. An apparatus for inspecting arteriosclerosis of a living subject, comprising: a pulse-wave-velocity-related-information obtaining device which obtains pulse-wave-velocity-related information that is related to a velocity at which a pulse wave propagates in the subject;a blood-pressure measuring device which measures a blood pressure of the subject;a heart-rate measuring device which measures a heart rate of the subject;a pre-ejection-period measuring device which measures a pre-ejection period from a time of starting of contraction of the heart of the subject to a time of starting of ejection of blood from the heart;an ejection-time measuring device which measures an ejection time from the time of starting of ejection of blood from the heart to a time of ending of ejection of blood from the heart; andan arteriosclerosis-inspection-related-augmentation-index determining means for determining an arteriosclerosis-inspection-related augmentation index of the subject, based on the pulse-wave-velocity-related information obtained by the pulse-wave-velocity-related-information obtaining device, the blood pressure measured by the blood-pressure measuring device, the heart rate measured by the heart-rate measuring device, the pre-ejection period measured by the pre-ejection-period measuring device, and the ejection time measured by the ejection-time measuring device, according to a predetermined relationship between (A) (a1) pulse-wave-velocity-related information, (a2) blood pressure, (a3) heart rate, (a4) pre-ejection period and (a5) ejection time, and (B) arteriosclerosis-inspection-related augmentation index. 2. An apparatus according to claim 1, wherein the predetermined relationship is represented by a following expression: AIE=a×PWV+b×BP+c×HR+d×ET+e×PEP+fwhere PWV is pulse-wave-velocity-related information, BP is blood pressure,HR is heart rate,PEP is pre-ejection periodET is ejection time,AIE is arteriosclerosis-inspection-related augmentation index,a, b, c, d, and e are coefficients, andf is a constant. 3. An apparatus according to claim 1, further comprising a display device which displays the arteriosclerosis-inspection-related augmentation index of the subject determined by the arteriosclerosis-inspection-related-augmentation-index determining means. 4. An apparatus according to claim 2, wherein the predetermined relationship is predetermined for each living subject, and wherein the apparatus further comprises a display device which displays the coefficients a, b, c, d, and e and the constant f of the predetermined relationship. 5. An apparatus according to claim 1, further comprising a memory device which stores the predetermined relationship. 6. An apparatus for inspecting arteriosclerosis of a living subject, comprising: an augmentation-index measuring device which measures an augmentation index of the subject that is a proportion of a magnitude of a reflected-wave component of a pulse wave of the subject to a magnitude of an incide... | Colin Medical Technology Corporation | A61B005/00
US07254851 | B2 | 2007-08-14 | Patient support structure having a tabletop with a breast positioning aperture for a laser imaging apparatus | 2002-11-13 | A patient support structure for a laser imaging apparatus, comprises a tabletop to support a female patient in front-down, prone position. The tabletop includes an opening adapted to permit a breast of the patient to be vertically pendant below the tabletop. The opening is non-symmetric with respect to an axis of rotation of a scanning mechanism disposed below the tabletop. | 1. A patient support structure in combination with a laser imaging apparatus, including a laser scanning mechanism, comprising: a) a tabletop to support a female patient in front-down, prone position;b) said tabletop including an opening adapted to permit a breast of the patient to be vertically pendant below said tabletop;c) said opening being non-symmetric with respect to a vertical axis of rotation of said laser scanning mechanism disposed below said tabletop, said vertical axis projecting through said opening;d) said opening including a first portion defined by one side of a line across said opening through said axis and a first peripheral portion of said opening extending from said one side of said line toward the patient's feet;e) said opening including a second portion defined by an opposite side of said line and a second peripheral portion of said opening extending from said opposite side of said line toward the patient's head; andf) said second portion is elongated toward the head as compared to said first portion. | 1. A patient support structure in combination with a laser imaging apparatus, including a laser scanning mechanism, comprising: a) a tabletop to support a female patient in front-down, prone position;b) said tabletop including an opening adapted to permit a breast of the patient to be vertically pendant below said tabletop;c) said opening being non-symmetric with respect to a vertical axis of rotation of said laser scanning mechanism disposed below said tabletop, said vertical axis projecting through said opening;d) said opening including a first portion defined by one side of a line across said opening through said axis and a first peripheral portion of said opening extending from said one side of said line toward the patient's feet;e) said opening including a second portion defined by an opposite side of said line and a second peripheral portion of said opening extending from said opposite side of said line toward the patient's head; andf) said second portion is elongated toward the head as compared to said first portion. 2. A patient support structure as in claim 1, wherein: a) the peripheral edge portion of said opening is pliable. 3. A patient support structure as in claim 1, wherein: a) said first portion is semi-circular defined by a first radius having a center at said axis; andb) said second portion is partly defined by a circular arc having a second radius greater than said first radius and having a center at said axis. 4. A patient support structure as in claim 3, wherein: a) said semi-circle is joined to said circular arc with a pair of tangent lines. 5. A patient support structure as in claim 1, wherein: a) said opening is built into said tabletop. 6. A patient support structure as in claim 1, wherein: a) said opening is oval. 7. A patient support structure as in claim 1, wherein: a) said opening is elliptical. 8. A patient support structure as in claim 1, wherein: a) said opening is disposed in a removable disk; andb) said disk is received in a pocket having a larger opening on said tabletop. 9. A patient support structure as in claim 1, wherein: a) said laser scanning mechanism comprises a detector array including a laser beam directed toward the breast; andb) said detector array being disposed below said tabletop to image the internal structure of the breast, said detector array being rotatable about said vertical axis disposed through said opening. 10. A patient support structure as in claim 9, wherein said laser source is a semiconductor diode laser. 11. A patient support structure as in claim 9, wherein said laser source is a solid state laser. 12. A patient support structure as in claim 9, wherein said laser source is a near infrared light source. 13. A patient support structure as in claim 9, wherein said detector array includes a plurality of collimators associated with respective detectors, said collimators are directed toward said axis of rotation. 14. A patient support structure as in claim 13, wherein each o... | Imaging Diagnostic Systems, Inc. | A61B005/05
US07018337 | B2 | 2006-03-28 | Method and apparatus for automatic non-invasive blood pressure monitoring | 2002-11-11 | An automatic non-invasive blood pressure monitoring system includes a blood pressure monitor, a blood pressure cuff pneumatically connected to the blood pressure monitor, a deflation valve connected intermediate the blood pressure monitor and the blood pressure cuff, and a controller for automatically controlling the non-invasive blood pressure monitoring. | 1. An automatic non-invasive blood pressure monitoring system, comprising: a blood pressure monitor having an electronic dump valve;a blood pressure cuff;a deflation valve detachably and pneumatically connected with and intermediate said blood pressure monitor and said blood pressure cuff, said deflation valve being manually operable and configured to rapidly deflate said blood pressure cuff independent of said electronic dump valve; anda controller for automatically controlling said blood pressure monitor;wherein in response to a single-handed operation of said deflation valve by an operator, said blood pressure cuff is allowed to rapidly deflate absent a need for an operator to separate said blood pressure cuff from said blood pressure monitor. | 1. An automatic non-invasive blood pressure monitoring system, comprising: a blood pressure monitor having an electronic dump valve;a blood pressure cuff;a deflation valve detachably and pneumatically connected with and intermediate said blood pressure monitor and said blood pressure cuff, said deflation valve being manually operable and configured to rapidly deflate said blood pressure cuff independent of said electronic dump valve; anda controller for automatically controlling said blood pressure monitor;wherein in response to a single-handed operation of said deflation valve by an operator, said blood pressure cuff is allowed to rapidly deflate absent a need for an operator to separate said blood pressure cuff from said blood pressure monitor. 2. The automatic non-invasive blood pressure monitoring system of claim 1, further comprising: a monitor hose pneumatically connecting said blood pressure monitor to said deflation valve; anda cuff hose pneumatically connecting said deflation valve to said blood pressure cuff. 3. The automatic non-invasive blood pressure monitoring system of claim 2, wherein: said monitor hose and said cuff hose are detachably connected to said deflation valve. 4. The automatic non-invasive blood pressure monitoring system of claim 3, wherein said deflation valve further comprises: a valve body;a monitor port at a first end of said valve body for receiving said monitor hose;a cuff port at a second end of said valve body for receiving said cuff hose; andan air channel intermediate said monitor port and said cuff port, wherein said air channel comprises a sealable exhaust port for exhausting air from said air channel to ambient. 5. The automatic non-invasive blood pressure monitoring system of claim 4, wherein said deflation valve further comprises: an actuator assembly in operable communication with said sealable exhaust port;an exhaust port seal disposed proximate said sealable exhaust port, said exhaust port seal being responsive to said actuator assembly;a bias spring disposed within said valve body for biasing said actuator assembly in a first direction; andsaid sealable exhaust port being sealed when said actuator assembly is biased in said first direction and unsealed when said actuator assembly is biased in a second direction. 6. The automatic non-invasive blood pressure monitoring system of claim 5, wherein said actuator assembly further comprises: an actuator;a seal carrier for supporting said exhaust port seal; anda link disposed intermediate said actuator and said seal carrier for communicating a force between said actuator and said seal carrier. 7. The automatic non-invasive blood pressure monitoring system of claim 5, wherein said actuator assembly comprises a push-button actuator disposed within said valve body for one-hand operation of said deflation valve by an operator. 8. The automatic non-invasive blood pressure monitoring system of claim 6, wherein said bias spring is disposed proximate said sea... | GE Medical Systems Global Technology Company, LLC | A61B005/02
US07047062 | B2 | 2006-05-16 | Magnetic resonance imaging with nested gradient pulses | 2002-05-16 | Disclosed herein is a system and method for decreasing gradient field pulse sequence duration and reducing peripheral nerve stimulation with known gradient pulse areas for a magnetic resonance imaging system. The method comprising: receiving a first desired area corresponding to a first pulse; obtaining a second desired area corresponding to a second pulse; selecting the first pulse as a nested pulse if the first desired area is smaller than the second desired area, and establishing the second pulse as a nesting pulse, otherwise selecting the second pulse as the nested pulse and establishing the first pulse as the nesting pulse. The method also includes: determining an amplitude and pulse duration for the nested pulse and ascertaining an amplitude and pulse duration for the nesting pulse. Finally, the method includes arranging a plurality of gradient field pulse sequences to include the nested pulse and the nesting pulse. | 1. A method for decreasing gradient field pulse sequence duration and reducing peripheral nerve stimulation with known gradient pulse areas for a magnetic resonance imaging system, the method comprising: receiving a first desired area corresponding to a first pulse;obtaining a second desired area corresponding to a second pulse;selecting said first pulse as a nested pulse if said first desired area is smaller than said second desired area, and establishing said second pulse as a nesting pulse, otherwise selecting said second pulse as said nested pulse and establishing said first pulse as said nesting pulse, said nested pulse having portions coincident with said nesting pulse;determining an amplitude and pulse duration for said nested pulse;ascertaining an amplitude and pulse duration for said nesting pulse; andarranging a plurality of gradient field pulse sequences to include said nested pulse and said nesting pulse. | 1. A method for decreasing gradient field pulse sequence duration and reducing peripheral nerve stimulation with known gradient pulse areas for a magnetic resonance imaging system, the method comprising: receiving a first desired area corresponding to a first pulse;obtaining a second desired area corresponding to a second pulse;selecting said first pulse as a nested pulse if said first desired area is smaller than said second desired area, and establishing said second pulse as a nesting pulse, otherwise selecting said second pulse as said nested pulse and establishing said first pulse as said nesting pulse, said nested pulse having portions coincident with said nesting pulse;determining an amplitude and pulse duration for said nested pulse;ascertaining an amplitude and pulse duration for said nesting pulse; andarranging a plurality of gradient field pulse sequences to include said nested pulse and said nesting pulse. 2. The method of claim 1 wherein said determining further comprises: establishing if said nested pulse is triangular or trapezoidal and establishing a flat top time for said nested pulse if it is trapezoidal; andcalculating a ramp time for said nested pulse. 3. The method of claim 2 wherein said calculating comprises: setting a first duration corresponding to said nested pulse equivalent to about a second duration corresponding to a flattop portion of said nesting pulse; andcalculating a ramp time for said nesting pulse. 4. The method of claim 1 wherein said first pulse is a slice select rephaser pulse and said second pulse is a readout prephaser pulse. 5. The method of claim 1 wherein said first pulse is a readout prephaser pulse and said second pulse is a slice select rephaser pulse. 6. The method of claim 1 wherein said plurality of gradient field pulse sequences includes a slice select pulse sequence, a phase encoding pulse sequence and a readout pulse sequence. 7. The method of claim 6 wherein: said slice select pulse sequence comprises a slice select pulse, a slice select rephaser pulse and a slice select dephaser pulse; said phase encoding pulse sequence includes a phase encoding pulse and a phase encoding rewinder pulse; and said readout pulse sequence includes a readout prephaser pulse, a readout pulse and a readout dephaser pulse. 8. The method of claim 7 wherein said slice select dephaser pulse is configured to be about equivalent to said slice select rephaser pulse. 9. The method of claim 7 wherein said phase encoding rewinder pulse is configured to be about equivalent to said phase encoding pulse but of opposite sign. 10. The method of claim 7 wherein said readout dephaser pulse is configured to be about equivalent to said readout prephaser pulse. 11. The method of claim 1 further including determining an amplitude and duration for a slice select pulse and a readout pulse responsive to a desired area for said slice select pulse and another desired area for said readout pulse. 12. The method of claim 11 wherein... | GE Medical Systems Global Technology Company, LLC | A61B005/05
US07069067 | B2 | 2006-06-27 | Method for operating a magnetic resonance apparatus and magnetic resonance apparatus for implementing the method | 2002-05-09 | A method is provided for operating a magnetic resonance apparatus comprising a magnetic stimulation device, the method comprising bearing a region of a central nervous system of a living examination subject in an imaging volume of the magnetic resonance apparatus, whereby this region is to be imaged, operating the magnetic stimulation device for stimulating at least one prescribable location of the examination subject outside of the region to be imaged and generating a functional magnetic resonance image of the region to be imaged for imaging a neuronal activity of the central nervous system, whereby said neuronal activity is initiated by the stimulation. | 1. A method for operating a magnetic resonance apparatus comprising a magnetic stimulation device, the method comprising: bearing a region to be imaged of a central nervous system of a living examination subject in an imaging volume defined by the volume of a homogeneous basic magnetic field of the magnetic resonance apparatus;stimulating with the magnetic stimulation device at least one prescribable location of the examination subject outside of the imaging volume;initiating neuronal activity of the central nervous system by the stimulating; andgenerating a functional magnetic resonance image of the region to be imaged for imaging the neuronal activity. | 1. A method for operating a magnetic resonance apparatus comprising a magnetic stimulation device, the method comprising: bearing a region to be imaged of a central nervous system of a living examination subject in an imaging volume defined by the volume of a homogeneous basic magnetic field of the magnetic resonance apparatus;stimulating with the magnetic stimulation device at least one prescribable location of the examination subject outside of the imaging volume;initiating neuronal activity of the central nervous system by the stimulating; andgenerating a functional magnetic resonance image of the region to be imaged for imaging the neuronal activity. 2. The method according to claim 1, further comprising: simultaneously stimulating at least one further location of the examination subject. 3. The method according to claim 1, wherein the prescribable location is a deeply situated location in the examination subject. 4. The method according to claim 1, wherein the prescribable location is situated in an extremity or a torso of a patient. 5. The method according to claim 1, wherein the central nervous system comprises a brain of a patient. 6. The method according to claim 1, further comprising: providing an examination subject who has at least a partial moving incapability. 7. The method according to claim 1, further comprising: anesthetizing the examination subject; andmonitoring anesthesia of the examination subject based on the functional magnetic resonance image. 8. The method according to claim 1, further comprising: varying a signal controlling the magnetic stimulation device by at least one of amplitude, duration, and rate of increase. 9. A magnetic resonance device for implementing the method according to claim 1, comprising: a basic field magnetic system configured to generate an optimally homogeneous static basic magnetic field at least within an imaging volume defined by the volume of a homogeneous basic magnetic field;a gradient system configured to generate optimally linear gradient fields at least within the imaging volume and configured as a magnetic stimulation device for stimulating a prescribable location outside the imaging volume; anda high-frequency system configured to initiate and accept magnetic resonance signals of the area to be imaged; anda receiving device configured to receive a functional magnetic resonance image produced on the basis of the magnetic resonance signals. 10. A magnetic resonance apparatus, comprising: a basic field magnetic system configured to generate an optimally homogeneous static basic magnetic field at least within an imaging volume defined by the volume of a homogeneous basic magnetic field;a gradient system configured to generate optimally linear gradient fields at least within the imaging volume;a high-frequency system configured to initiate and accept magnetic resonance signals of the area to be imaged;a receiving device configured to receive a functional magnetic resonance image pro... | Siemens Aktiengesellschaft | A61B005/05
US07210480 | B2 | 2007-05-01 | Shock treatment systems and methods | 2002-04-08 | According to the invention, methods and devices for increasing cardiopulmonary circulation induced by chest compression and decompression when performing cardiopulmonary resuscitation are provided. According to one method, a pressure responsive inflow valve is coupled to a patient's airway. Chest compressions and chest decompressions are performed. During chest decompression the inflow valve prevents respiratory gases from entering the lungs until a certain negative intrathoracic pressure level is exceeded at which time the one inflow valve opens. In this way, the inflow valve assists in increasing the magnitude and duration of negative intrathoracic pressure during decompression to enhance the amount of blood flow into the heart and lungs. Further, the patient is supplied with a pressurized respiratory gas through the inflow valve when the inflow valve opens to ventilate the patient. | 1. A method for increasing the blood pressure in a spontaneously breathing person in need thereof, the method comprising the steps of: sensing at least one physiological parameter that is selected from a group consisting of the blood pressure, end tidal CO2 and oxygen saturation of the person;interfacing an inflow valve to the person's airway that is operable to vary resistance to respiratory gas inflow through the inflow valve to the person's lungs due to spontaneous inhalation;operating the inflow valve based on the sensed parameter to vary resistance to respiratory gas inflow to the person's lungs while the person is spontaneously inhaling to create a vacuum within the thorax of the person within the range from less than 0 cm H2O to about −50 cm H2O and increase blood flow back to the right heart of the person, thereby enhancing the person's blood pressure. | 1. A method for increasing the blood pressure in a spontaneously breathing person in need thereof, the method comprising the steps of: sensing at least one physiological parameter that is selected from a group consisting of the blood pressure, end tidal CO2 and oxygen saturation of the person;interfacing an inflow valve to the person's airway that is operable to vary resistance to respiratory gas inflow through the inflow valve to the person's lungs due to spontaneous inhalation;operating the inflow valve based on the sensed parameter to vary resistance to respiratory gas inflow to the person's lungs while the person is spontaneously inhaling to create a vacuum within the thorax of the person within the range from less than 0 cm H2O to about −50 cm H2O and increase blood flow back to the right heart of the person, thereby enhancing the person's blood pressure. 2. A method as in claim 1, wherein the person has low blood pressure due to conditions selected from a group consisting of blood loss, the administration of a drug, a high gravitational state, vasovagal syncope, cardiac tamponade, drowning, heat stroke, heart attack, right heart failure, return to earth after space flight, and sepsis. 3. A method as in claim 1, further comprising sending a signal to a controller that is representative of the sensed parameter and sending a signal from the controller to a mechanism to operate the valve based on the sensed parameter. 4. A method as in claim 1, wherein the inflow valve is operated to achieve a negative intrathoracic pressure within the range from about −5 cm H2O to about −15 cm H2O when the person breathes at flow rates in the range from about zero flow to about 70 liters per minute. 5. A method as in claim 1, wherein the inflow valve comprises an airway and an occlusion member, wherein the occlusion member is moved across the airway to vary the flow rate through the inflow valve. 6. A method as in claim 1, wherein the physiological parameter is repeatedly sensed, and wherein the valve is repeatedly operated based the sensed parameters. 7. A method as in claim 1, further comprising supplying supplemental ventilation or oxygen to the person. 8. A method as in claim 1, further comprising supplying a drug or a medicament to the person upstream or downstream of the inflow valve. 9. A method for increasing the blood pressure in a spontaneously breathing person, the method comprising the steps of: sensing at least one physiological parameter that is selected from a group consisting of the blood pressure, end tidal CO2 and oxygen saturation of the person;interfacing an inflow valve to the person's airway that is operable to increase or decrease respiratory gas flow to the person's lungs;operating the inflow valve based on the sensed parameter to increase or decrease respiratory gas flow to the person's lungs while the person is inhaling to create a vacuum within the thorax of the person and increase blood flow back to the right heart of the pe... | Advanced Circulatory Systems, Inc. | A62B009/02 A61M016/00 A61B005/00 A61B005/02 A61B005/08
US07172562 | B2 | 2007-02-06 | System, method and apparatus for locating, measuring and evaluating the enlargement of a foramen | 2002-11-22 | A system, apparatus, and method for locating, measuring and evaluating the enlargement of a foramen are provided. An instrument has a handle, an angled region, and a tip to be inserted into the foramen. An instrument can be pressed onto the foramen to determine its location and whether the empty space within the foramen is large enough for the maximum width of the tip to be inserted. If the tip cannot insert, the foramen can be cut to enlarge it enough for the tip to insert. A kit of several instruments or a single instrument with tips of different maximum widths are provided for determining the amount of empty space within the foramen through the course of the enlargement. | 1. An instrument for locating and measuring available open space for an exiting nerve in a neural foramen comprising: a handle;a substantially rigid tip having a pre-determined measurement dimension defined by the maximum width of the rigid tip; anda substantially rigid angled region coupled between the handle and the tip, said angled region making an angle of from 60 to 80 degrees from the tip distally to the axis of the handle,wherein the tip and the angled region are sized such that the maximum width of the tip is insertable into said neural foramen when said neural foramen is in at least one of an unenlarged and an enlarged state to obtain an accurate measurement of the size of said neural foramen. | 1. An instrument for locating and measuring available open space for an exiting nerve in a neural foramen comprising: a handle;a substantially rigid tip having a pre-determined measurement dimension defined by the maximum width of the rigid tip; anda substantially rigid angled region coupled between the handle and the tip, said angled region making an angle of from 60 to 80 degrees from the tip distally to the axis of the handle,wherein the tip and the angled region are sized such that the maximum width of the tip is insertable into said neural foramen when said neural foramen is in at least one of an unenlarged and an enlarged state to obtain an accurate measurement of the size of said neural foramen. 2. The instrument of claim 1, wherein the tip is substantially rounded. 3. The instrument of claim 1, wherein the tip is substantially spherical. 4. The instrument of claim 1, wherein the angled region forms an approximately 70 degree to 78 degree angle from the tip distally to the axis of the handle. 5. The instrument of claim 1, wherein the angled region forms an approximately 76 degree angle from the tip to the axis of the handle. 6. The instrument of claim 1, wherein the angled region is curved and having a diameter less than the measurement dimension. 7. The instrument of claim 1, wherein the maximum width of the tip is in the range of approximately 0.5 mm to 8 mm. 8. The instrument of claim 1, further comprising at least one substantially rigid substitute tip, wherein the tip and the substitute tip are removably coupleable to the angled region. 9. The instrument of claim 1, further comprising a second substantially rigid angled region having a second substantially rigid tip, said second angled region being coupled to the handle opposite the first angled region, wherein the second tip and the second angled region are sized such that the second tip is insertable into said neural foramen when said neural foramen is in at least one of an unenlarged and an enlarged state. 10. A system for finding and evaluating available open space for an exiting nerve in a neural foramen comprising: a plurality of instruments, each having a handle, a substantially rigid tip having a pre-determined measurement dimension defined by the maximum width, and a substantially rigid angled region between the handle and the tip, said angled region making an angle of from 60 to 80 degrees from the tip distally to the axis of the handle,wherein each tip and each angled region are sized such that the maximum width of each tip is insertable into said neural foramen when said neural foramen is in at least one of an unenlarged and an enlarged state, andwherein each tip of the plurality of instruments has a different maximum width than the tips of the others of the plurality of instruments. 11. The system of claim 10, wherein each tip is substantially rounded. 12. The system of claim 10, wherein each tip is substantially spherical. 13. The system of claim 10, wherein ... | - | A61B005/103 A61B005/107 G01B001/00
US07047067 | B2 | 2006-05-16 | Apparatus, methods, and computer program products for evaluating a risk of cardiac arrhythmias from restitution properties | 2002-05-31 | Methods, systems, and computer program products can be provided for evaluating a risk of the occurrence of cardiac arrhythmias in a heart. The system can include at least one sensing electrode and a controller. The sensing electrode can be configured to sense electrical signals from a heart over a range of heart rates. The controller can be configured to determine values of a at least one restitution property of the heart from the sensed electrical signals and to evaluate the risk of the occurrence of cardiac arrhythmias from the determine values of the at least one restitution property. | 1. A method for evaluating a risk of occurrence of cardiac arrhythmias in a heart, the method comprising: detecting electrical activity from a heart over a range of heart rates;determining values of at least one restitution property of the heart from the detected electrical activity over the range of heart rates; andevaluating the risk of occurrence of cardiac arrhythmias in the heart based on the determined values of the at least one restitution property by determining a rate of change of the determined restitution property values as a function of the heart rate; wherein the evaluating step comprises:determining an increased risk of the occurrence of cardiac arrhythmias when, in response to an abrupt change in the heart rate, the rate of change of the determined restitution property values is greater than a threshold value. | 1. A method for evaluating a risk of occurrence of cardiac arrhythmias in a heart, the method comprising: detecting electrical activity from a heart over a range of heart rates;determining values of at least one restitution property of the heart from the detected electrical activity over the range of heart rates; andevaluating the risk of occurrence of cardiac arrhythmias in the heart based on the determined values of the at least one restitution property by determining a rate of change of the determined restitution property values as a function of the heart rate; wherein the evaluating step comprises:determining an increased risk of the occurrence of cardiac arrhythmias when, in response to an abrupt change in the heart rate, the rate of change of the determined restitution property values is greater than a threshold value. 2. A method according to claim 1, wherein the detecting is carried out by detecting electrocardiogram signals over the range of heart rates. 3. A method according to claim 2, wherein the determining step comprises determining QT time interval values and TQ time interval values. 4. A method according to claim 3, wherein the evaluating step comprises: determining a rate of change of at least one of the QT time interval values and the TQ time interval values; anddetermining an increased risk of the occurrence of cardiac arrhythmias when the rate of change, in response to an abrupt change in the heart rate, is greater than a threshold value. 5. A method according to claim 1, wherein the detecting is carried out by detecting action potential signals and diastolic interval signals over the range of heart rates. 6. A method according to claim 5, wherein the evaluating step comprises: determining a rate of change of the action potential duration values; anddetermining an increased risk of the occurrence of cardiac arrhythmias when the rate of change, in response to an abrupt change in the heart rate, is greater than a threshold value. 7. A method according to claim 1, wherein: the detecting step is carried out by detecting T wave duration over the range of heart rates; andthe evaluating step is carried out by determining an increased risk of the occurrence of cardiac arrhythmias when the T wave duration, in response to a change in the heart rate, is greater than a threshold value. 8. A method for evaluating a risk of occurrence of cardiac arrhythmias in a heart, the method comprising: detecting electrical activity from a heart over a range of heart rates;determining values of at least one restitution property of the heart from the detected electrical activity over the range of heart rates; andevaluating the risk of occurrence of cardiac arrhythmias in the heart based on the determined values of the at least one restitution property by determining a rate of change of the determined restitution property values as a function of the heart rate, wherein the evaluating step comprises:identifying a restitution property of the heart a... | UAB Research Foundation | A61B005/452
US07027848 | B2 | 2006-04-11 | Apparatus and method for non-invasive spectroscopic measurement of analytes in tissue using a matched reference analyte | 2002-04-04 | A method is provided for building improved calibration models or for improving modifications to these models or validation of the models used in the non-invasive spectroscopic measurement of an analyte or attribute of tissue. The method uses a matched reference sample and measurement of that sample to ensure that the correct relationship between the spectra and analyte is made during the model building, modification or calibration process. A matched reference sample is one in which the analyte of interest or attribute of interest in the reference sample is representative of the analyte or attribute at the site being non-invasively sampled or the agreement between the reference concentration and the non-invasively sampled concentration is clinically significant. | 1. A method for performing modifications to calibration models in the non-invasive spectroscopic measurement of an analyte or attribute of tissue, the method comprising the steps of: providing a means for irradiating tissue with infrared energy;providing an output element, the output element operatively connected to a means for measuring a spectrum;irradiating tissue with the infrared energy so that there is absorption of at least a portion of the infrared energy in the tissue;collecting at least a portion of the infrared energy exiting the tissue through the output element and measuring the resulting spectrum;collecting at least one matched reference sample that is assayed to determine at least one reference value of the analyte or attribute of tissue; andusing the at least one reference value with the spectral measurement to perform a modification to the calibration model;wherein the matched reference sample is a fluid taken from a tissue compartment that is kinetically matched to the tissue which is irradiated. | 1. A method for performing modifications to calibration models in the non-invasive spectroscopic measurement of an analyte or attribute of tissue, the method comprising the steps of: providing a means for irradiating tissue with infrared energy;providing an output element, the output element operatively connected to a means for measuring a spectrum;irradiating tissue with the infrared energy so that there is absorption of at least a portion of the infrared energy in the tissue;collecting at least a portion of the infrared energy exiting the tissue through the output element and measuring the resulting spectrum;collecting at least one matched reference sample that is assayed to determine at least one reference value of the analyte or attribute of tissue; andusing the at least one reference value with the spectral measurement to perform a modification to the calibration model;wherein the matched reference sample is a fluid taken from a tissue compartment that is kinetically matched to the tissue which is irradiated. 2. The method of claim 1, wherein the matched reference sample is interstitial fluid. 3. The method of claim 1, wherein the matched reference sample is bulk fluid in the tissue. 4. The method of claim 1, wherein the matched reference sample is fluid obtained from a limb. 5. The method of claim 1, wherein the matched reference sample is lancet blood obtained from a forearm. 6. The method of claim 1, wherein the matched reference sample is a fluid taken from a tissue compartment that is spatially matched to the tissue which is irradiated. 7. The method of claim 1, wherein site-to-site variations are reduced through multiple collection of matched reference samples of the tissue over a sufficient area. 8. The method of claim 7, where the area optically sampled and the area used for procurement of multiple matched reference samples are nominally similar. 9. A method for performing subject specific modifications in the non-invasive spectroscopic measurement of an analyte or attribute of tissue, the method comprising the steps of: providing a means for irradiating tissue with infrared energy;providing an output element, the output element operatively connected to a means for measuring a spectrum;irradiating tissue with the infrared energy so that there is absorption of at least a portion of the infrared energy in the tissue;collecting at least a portion of the infrared energy exiting the tissue through the output element and measuring the resulting spectrum;collecting reference samples that are kinetically matched to the irradiated tissue and assaying the reference samples for the analyte or attribute of tissue; andusing the assayed reference samples with the spectral measurements to perform a subject-specific modification. 10. A method for building a calibration model for the non-invasive spectroscopic measurement of an analyte or attribute of tissue, the method comprising the steps of: providing a means for irradiating tissue with i... | InLight Solutions, Inc. University of New Mexico | A61B005/00
US07156813 | B2 | 2007-01-02 | Flow-independent parameter estimation based on tidal breathing exhalation profiles | 2003-02-05 | Parametric characterization of nitric oxide (NO) gas exchange using a two-compartment model of the lungs is a non-invasive technique to characterize inflammatory lung diseases. The technique applies the two-compartment model to parametric characterization of NO gas exchange from a tidal breathing pattern. The model is used to estimate up to six flow-independent parameters, and to study alternate breathing patterns. | 1. A method of characterizing nitric oxide exchange in lungs during tidal breathing, comprising: utilizing a two compartment model, wherein a first compartment represents an alveolar region of the lungs and a second compartment represents an airway to the lungs;selecting several flow-independent parameters; andestimating the values of the parameters by utilizing the two compartment model fitted to previously obtained experimental data gathered during tidal breathing. | 1. A method of characterizing nitric oxide exchange in lungs during tidal breathing, comprising: utilizing a two compartment model, wherein a first compartment represents an alveolar region of the lungs and a second compartment represents an airway to the lungs;selecting several flow-independent parameters; andestimating the values of the parameters by utilizing the two compartment model fitted to previously obtained experimental data gathered during tidal breathing. 2. The method of claim 1, wherein the flow independent parameters include one or more of J′awNO, DawNO, Calv,ss, {circumflex over (D)}alvNO, Vaw and Vds. 3. The method of claim 1, wherein the step of estimating further comprises characterizing at least Vaw and Vds of said several flow-independent parameters based on the two-compartment model. 4. The method of claim 1, wherein the step of estimating further comprises minimizing a difference between estimated values and experimentally obtained values. 5. The method of characterizing of claim 1, further comprising: representing estimated values by a data curve or an equation defining the curve;wherein said step of estimating comprises determining a first confidence interval after a first monitoring time; andwherein said step of estimating further comprises projecting at least one of said values having improved accuracy and an improved second confidence interval after a second predetermined monitoring time, wherein the projection is based on the data curve or the equation representing the curve. 6. The method of characterizing of claim 1, wherein the step of estimating further comprises improving the accuracy of the parameter values by reducing the frequency of tidal breathing during monitoring from an initial value of the frequency of tidal breathing. 7. The method of characterizing of claim 6, wherein the frequency of breathing is reduced by increasing the time for each tidal breath to in the range from 7 to 17 seconds. 8. The method of characterizing of claim 1, wherein the step of estimating further comprises improving the accuracy of the parameter values by increasing the flow rate ratio q of the inhalation flow rate QI, relative to the exhalation flow rate QE as compared to an initial magnitude of the ratio q. 9. The method of characterizing of claim 8, further comprising increasing the flow rate ratio q into the range from 6/5 to 12. 10. The method of characterizing of claim 1, further comprising the preliminary steps of: having a subject perform a tidal breathing maneuver;measuring and recording NO concentration and flow rate simultaneously during the maneuver; wherein the step of utilizing further comprises simulating the tidal breathing maneuver by the two-compartment model;the step of estimating further comprises the preliminary step of fitting recorded data to simulated data from the two-compartment model;wherein the method further comprises the additional preliminary steps of: comparing the recorded data and the sim... | The Regents of the University of California | A61B005/08
US07013174 | B2 | 2006-03-14 | Electromagnetic interference immune tissue invasive system | 2002-02-19 | An electromagnetic immune tissue invasive system includes a primary device housing. The primary device housing having a control circuit therein. A shielding is formed around the primary device housing to shield the primary device housing and any circuits therein from electromagnetic interference. A lead system transmits and receives signals between the primary device housing. The lead system is either a fiber optic system or an electrically shielded electrical lead system. | 1. An electromagnetic radiation immune tissue invasive sensing system, comprising: a photonic lead having a proximal end and a distal end;a light source, in the proximal end of said photonic lead, to produce a first light having a first wavelength;a wave-guide between the proximal end and distal end of said photonic lead;a bio-sensor, in the distal end of said photonic lead, to sense characteristics of a predetermined tissue region; anda distal sensor, in the distal end of said photonic lead, to convert the first light into electrical energy;said distal sensor, responsive to said bio-sensor, causing a second light, the second light having a second wavelength, to be propagated, through said wave-guide, to the proximal end of said photonic lead such that a characteristic of the second light is modulated to encode the sensed characteristics of the predetermined tissue region;said light source, in the proximal end of said photonic lead, producing said second light having said second wavelength;said distal sensor, in the distal end of said photonic lead, reflecting the second light back to the proximal end of said photonic lead such that a characteristic of the second light is modulated to encode the sensed characteristics of the predetermined tissue region. | 1. An electromagnetic radiation immune tissue invasive sensing system, comprising: a photonic lead having a proximal end and a distal end;a light source, in the proximal end of said photonic lead, to produce a first light having a first wavelength;a wave-guide between the proximal end and distal end of said photonic lead;a bio-sensor, in the distal end of said photonic lead, to sense characteristics of a predetermined tissue region; anda distal sensor, in the distal end of said photonic lead, to convert the first light into electrical energy;said distal sensor, responsive to said bio-sensor, causing a second light, the second light having a second wavelength, to be propagated, through said wave-guide, to the proximal end of said photonic lead such that a characteristic of the second light is modulated to encode the sensed characteristics of the predetermined tissue region;said light source, in the proximal end of said photonic lead, producing said second light having said second wavelength;said distal sensor, in the distal end of said photonic lead, reflecting the second light back to the proximal end of said photonic lead such that a characteristic of the second light is modulated to encode the sensed characteristics of the predetermined tissue region. 2. The electromagnetic radiation immune tissue invasive sensing system as claimed in claim 1, further comprising: a proximal sensor, in the proximal end of said photonic lead, to convert the modulated second light into electrical energy. 3. The electromagnetic radiation immune tissue invasive sensing system as claimed in claim 2, further comprising: a transmitter, in the proximal end of said photonic lead and operatively connected to said proximal sensor, to transmit, in response the electrical energy from the converted modulated second light, information representing the sensed characteristics of the predetermined tissue region. 4. The electromagnetic radiation immune tissue invasive sensing system as claimed in claim 1, wherein said light source includes a first emitter to emit the first light having the first wavelength and a second emitter to emit the second light having the second wavelength. 5. The electromagnetic radiation immune tissue invasive sensing system as claimed in claim 1, wherein said light source includes a first laser to produce the first light having the first wavelength and a second laser to produce the second light having the second wavelength. 6. The electromagnetic radiation immune tissue invasive sensing system as claimed in claim 1, wherein said distal sensor includes: an optical attenuator coupled to a mirror; andan optical-electrical conversion device to convert the first light into electrical energy;said optical attenuator attenuating the second light to encode the sensed characteristics of the predetermined tissue region. 7. The electromagnetic radiation immune tissue invasive sensing system as claimed in claim 6, wherein said optical attenuator attenuating the... | Biophan Technologies, Inc. | A61B005/04
US07295870 | B2 | 2007-11-13 | Method for the detection and automatic characterization of nodules in a tomographic image and a system of medical imaging by tomodensimetry | 2002-11-22 | For the automatic detection and characterization of nodules in a tomographic image of an anatomical zone of a patient, the image is segmented for identifying therein a region of interest, and the segmented image is processed for identifying the nodule.;During the processing step, an ellipsoid inscribed in the region of interest is modeled for deciding whether the image elements inscribed in this ellipsoid correspond to a nodule, and, for each zone of the region of interest extending beyond the ellipsoid, the image elements are identified which do not belong to the module according to mathematical morphologic criteria. | 1. Method of automatic detection and characterization of nodules (N) in a tomographic image of an anatomical zone of a patient, comprising a segmentation of the image for identifying therein a region of interest (N, V) by determining a contour in three dimensions using water separation lines of luminance gradient data of the image to determine a discontinuity of luminance in the data, and the processing of the segmented image for identifying the nodule (N), wherein during the step of processing, an ellipsoid inscribed within the region of interest is modeled, it is decided that the image elements inscribed within this ellipsoid correspond to a nodule and, for each zone of the region of interest extending outside the ellipsoid, the image elements are identified which do not belong to the nodule, according to criteria of mathematical morphology. | 1. Method of automatic detection and characterization of nodules (N) in a tomographic image of an anatomical zone of a patient, comprising a segmentation of the image for identifying therein a region of interest (N, V) by determining a contour in three dimensions using water separation lines of luminance gradient data of the image to determine a discontinuity of luminance in the data, and the processing of the segmented image for identifying the nodule (N), wherein during the step of processing, an ellipsoid inscribed within the region of interest is modeled, it is decided that the image elements inscribed within this ellipsoid correspond to a nodule and, for each zone of the region of interest extending outside the ellipsoid, the image elements are identified which do not belong to the nodule, according to criteria of mathematical morphology. 2. Method according to claim 1, wherein a Euclidean distance field is produced with respect to the perimeter of the segmented region of interest, the image elements locally having the greatest distance value are determined, and the center of the ellipsoid is positioned on an image element having the greatest distance value. 3. Method according to claim 2, wherein the said ellipsoid is constituted by a sphere delimited by the greatest spherical volume having the said center and inscribed within the element of interest. 4. Method according to claim 2, wherein a geodesic distance transformation is performed so as to create, starting from the said center, geodesic surfaces (S) respectively extending at increasing distances with respect thereto and each delimited by a set of image elements inscribed within the element of interest and equidistant from the center. 5. Method according to claim 4, wherein, during the step of identification of image elements not belonging to the nodule, a separate processing is performed of each geodesic surface (S) situated outside the said sphere so as to determine whether the image elements which constitute it correspond at least partially to a nodule. 6. Method according to claim 5, wherein, during the processing of each geodesic surface, a discriminant analysis is performed from at least one morphologic criterion chosen from among the distance separating it from the center of the sphere, the connectivity of the image elements of the geodesic surface with the preceding ones, and the maximum Euclidean distance on this image element. 7. Method according to claim 1, wherein, preceding the segmentation step, a preliminary filtering is performed during which a volume (V) is produced in which is inscribed a portion of the image likely to include a nodule, the said step of segmentation being implemented on the image elements of the said volume. 8. Method according to claim 7, wherein the step of filtering is implemented by a thresholding of numerical values associated with the image elements. 9. Method according to claim 7, wherein the said step of thresholding is implemented by... | General Electric Company | A61B005/05
US07141023 | B2 | 2006-11-28 | Micro perfusion device comprising a collecting container | 2002-08-21 | A micro perfusion device including a casing, a perfusate supply, an inner cannula forming an inner lumen between a front cannula opening which is distal when the device is subcutaneously positioned and a rear cannula opening, the inner lumen being connected or connectable to the perfusate supply via the rear cannula opening, a perfusion cannula projecting from a lower side of the casing and surrounding the inner cannula, an outer lumen between the inner cannula and the perfusion cannula, a collecting container for the perfusate, and a perfusate outlet connecting the outer lumen to the collecting container, wherein the collecting container has a variable volume and a partial vacuum with respect to a pressure prevailing in the perfusate supply may be generated in the collecting container. | 1. A micro perfusion device comprising a collecting container, for obtaining at least one constituent of a body fluid using a subcutaneously positioned perfusion cannula, said device comprising: a) a casing;b) a perfusate supply;c) an inner cannula, forming an inner lumen between a front cannula opening which is distal when the device is subcutaneously positioned and a rear cannula opening, said inner lumen being connected or connectable to said perfusate supply via said rear cannula opening, said inner cannula being accommodated by said casing so as to be slidable in the longitudinal direction from a front sliding position to a rear sliding position and said rear cannula opening being formed in a surface area of said inner cannula;d) a sealing ring surrounding said inner cannula in said casing under pressure and, in said rear sliding position of said inner cannula, pre-latching into said rear cannula opening;e) a perfusion cannula, projecting from a lower side of said casing and surrounding said inner cannula, an outer lumen remaining between said inner cannula and said perfusion cannula;f) a collecting container for said perfusate; andg) an outlet for said perfusate, connecting said outer lumen to said collecting container; wherein said collecting container is a variable-volume collecting container and a partial vacuum with respect to a pressure prevailing in said perfusate supply is generated in said collecting container. | 1. A micro perfusion device comprising a collecting container, for obtaining at least one constituent of a body fluid using a subcutaneously positioned perfusion cannula, said device comprising: a) a casing;b) a perfusate supply;c) an inner cannula, forming an inner lumen between a front cannula opening which is distal when the device is subcutaneously positioned and a rear cannula opening, said inner lumen being connected or connectable to said perfusate supply via said rear cannula opening, said inner cannula being accommodated by said casing so as to be slidable in the longitudinal direction from a front sliding position to a rear sliding position and said rear cannula opening being formed in a surface area of said inner cannula;d) a sealing ring surrounding said inner cannula in said casing under pressure and, in said rear sliding position of said inner cannula, pre-latching into said rear cannula opening;e) a perfusion cannula, projecting from a lower side of said casing and surrounding said inner cannula, an outer lumen remaining between said inner cannula and said perfusion cannula;f) a collecting container for said perfusate; andg) an outlet for said perfusate, connecting said outer lumen to said collecting container; wherein said collecting container is a variable-volume collecting container and a partial vacuum with respect to a pressure prevailing in said perfusate supply is generated in said collecting container. 2. A micro perfusion device comprising a collecting container, for obtaining at least one constituent of a body fluid using a subcutaneously positioned perfusion cannula, said device comprising: a) a casing;b) a perfusate supply;c) an inner cannula, forming an inner lumen between a front cannula opening which is distal when the device is subcutaneously positioned and a rear cannula opening, said inner lumen being connected or connectable to said perfusate supply via said rear cannula opening;d) the perfusion cannula, projecting from a lower side of said casing and surrounding said inner cannula, an outer lumen remaining between said inner cannula and said perfusion cannula;e) a collecting container for said perfusate;f) an outlet for said perfusate, connecting said outer lumen to said collecting container; wherein said collecting container is a variable-volume collecting container and a partial vacuum with respect to a pressure prevailing in said perfusate supply is generated in said collecting container; andg) a sensor of a measuring means, for measuring the concentration of the at least one constituent of said body fluid, the sensor being arranged in or on said casing, in said outlet. 3. The micro perfusion device as set forth in claim 2, wherein said sensor is arranged in an area of said outlet bordering said outer lumen. 4. The micro perfusion device as set forth in claim 2, wherein: the measuring means comprises a working electrode and a counter electrode; andsaid counter electrode is formed on a lower side of said ca... | Disetronic Licensing AG | A61B005/00 A61M003/00 A61M005/178 B65D081/00
US07209776 | B2 | 2007-04-24 | Method of determining the position of the articular point of a joint | 2002-12-03 | A non-invasive method for determining the articular point of a joint is disclosed. The method uses a surgical navigation system having a stereoscopic cameras which tracks the positions of infrared emitting or reflecting markers attached to appendages on either side of the joint. A movable marker is used to palpate known landmarks on the appendages to determine their positions. The appendages are moved to determine the trajectories of the landmarks relative to the joint. Positional information from the cameras is fed to a data processing system with resident software which uses the positional information and trajectories to mathematically determine the position of the joint articular point according to the laws of kinematics. | 1. A method of obtaining preoperative information determining the position of an articular point of a central joint for guiding orthopaedic surgery performed on said central joint, said central joint being between two substantially rigid bodies, the central joint being located between first and second outer joints located at the ends of said first and second rigid bodies distal to said central joint, respectively, said method comprising the steps of: identifying the position of a first point on said first rigid body located substantially at said central joint;determining the position of a first articular point of said first rigid body at said first outer joint;defining a first axis between said first articular point and said first point;identifying the respective positions of a second and a third point on said first rigid body on opposite sides of said central joint;defining a plane substantially perpendicular to said first axis and at substantially equal respective rectangular distances to said second and third points, the intersection of said plane and said first axis being an initial estimate of said articular point of said central joint;determining the position of a second articular point of said second rigid body at said second outer joint;defining a region in said plane having a predetermined size, said region including said initial estimate of said articular point of said central joint;moving said second articular point relatively to said first rigid body by rotating said second rigid body about said central joint;identifying a multiplicity of different positions of said second articular point during said rotation of said second rigid body;determining one point among a multiplicity of points within said region for which the position is substantially invariant for each of said positions of said second articular point, said one point being said articular point of said central joint; andusing the position of said articular point to guide an orthopaedic surgical procedure. | 1. A method of obtaining preoperative information determining the position of an articular point of a central joint for guiding orthopaedic surgery performed on said central joint, said central joint being between two substantially rigid bodies, the central joint being located between first and second outer joints located at the ends of said first and second rigid bodies distal to said central joint, respectively, said method comprising the steps of: identifying the position of a first point on said first rigid body located substantially at said central joint;determining the position of a first articular point of said first rigid body at said first outer joint;defining a first axis between said first articular point and said first point;identifying the respective positions of a second and a third point on said first rigid body on opposite sides of said central joint;defining a plane substantially perpendicular to said first axis and at substantially equal respective rectangular distances to said second and third points, the intersection of said plane and said first axis being an initial estimate of said articular point of said central joint;determining the position of a second articular point of said second rigid body at said second outer joint;defining a region in said plane having a predetermined size, said region including said initial estimate of said articular point of said central joint;moving said second articular point relatively to said first rigid body by rotating said second rigid body about said central joint;identifying a multiplicity of different positions of said second articular point during said rotation of said second rigid body;determining one point among a multiplicity of points within said region for which the position is substantially invariant for each of said positions of said second articular point, said one point being said articular point of said central joint; andusing the position of said articular point to guide an orthopaedic surgical procedure. 2. A method according to claim 1, wherein said step of determining the position of a first articular point comprises the steps of: moving said first point by rotating said first rigid body about said first outer joint;identifying a plurality of positions of said first point during the motion of said first point about said first outer joint; anddetermining mathematically a common point at said first outer joint having substantially the same distance to all of said positions of said first point, said common point being said first articular point. 3. A method according to claim 1, wherein said step of determining the position of said second articular point comprises the steps of: identifying the respective positions of a fourth and a fifth point on opposite sides of said second outer joint;defining a first line between said fourth and fifth points;identifying the position of a sixth point located in a plane substantially perpendicular to said first line and about equidistant ... | Aesculap AG & Co. KG | A61B005/00 A61B005/103 A61B019/00 A61F002/32
US07177671 | B2 | 2007-02-13 | RF coil, RF signal transmitter receiver, RF signal receiver, and magnetic resonance imaging system for the inferior abdomen | 2002-05-07 | For the purpose of providing an RF coil that enables satisfactory imaging of the prostate, the RF coil is composed of a first saddle coil and a second saddle coil. The first saddle coil has two loop portions opposed to each other so that they can sandwich the inferior abdomen of a human body antero-posteriorly. The second saddle coil has two loop portions opposed to each other so that they will permit insertion of the inferior limbs of a human body thereinto and can sandwich the inferior abdomen thereof laterally. | 1. An RF coil comprising: a first saddle coil having two loop portions opposed to each other, wherein said two loop portions of said first saddle coil configured to sandwich the inferior abdomen of a human body antero-posteriorly; anda second saddle coil having two loop portions opposed to each other, wherein said two loop portions of said second saddle coil configured to permit insertion of the inferior limbs of a human body thereinto and configured to sandwich the inferior abdomen thereof laterally, wherein one of said two loop portions of said second saddle coil comprises a single piece having a circumference configured to receive one of the inferior limbs. | 1. An RF coil comprising: a first saddle coil having two loop portions opposed to each other, wherein said two loop portions of said first saddle coil configured to sandwich the inferior abdomen of a human body antero-posteriorly; anda second saddle coil having two loop portions opposed to each other, wherein said two loop portions of said second saddle coil configured to permit insertion of the inferior limbs of a human body thereinto and configured to sandwich the inferior abdomen thereof laterally, wherein one of said two loop portions of said second saddle coil comprises a single piece having a circumference configured to receive one of the inferior limbs. 2. An RF coil of claim 1, further comprising a flexible circuit board that is shaped to be readily worn on the inferior abdomen of a human body and that has a circuit pattern relevant to said first saddle coil and a circuit pattern relevant to said second saddle coil printed thereon. 3. An RF coil comprising: a first coil pair having two loops opposed to each other, wherein said two loops of said first coil pair configured to sandwich the inferior abdomen of a human body antero-posteriorly; anda second coil pair having two loops opposed to each other, wherein said two loops of said second coil pair configured to permit insertion of the inferior limbs of a human body thereinto and configured to sandwich the inferior abdomen thereof laterally, wherein one of said two loops of said second coil pair comprises a single piece having a circumference configured to receive one of the inferior limbs. 4. An RF coil of claim 3, further comprising a flexible circuit board that is shaped to be readily worn on the inferior abdomen of a human body and that has a circuit pattern relevant to said first coil pair and a circuit pattern relevant to said second coil pair printed thereon. 5. An RF signal transmitter-receiver comprising: a first saddle coil having two loop portions opposed to each other, wherein said two loop portions of said first saddle coil configured to sandwich the inferior abdomen of a human body antero-posteriorly;a second saddle coil having two loop portions opposed to each other, wherein said two loop portions of said second saddle coil configured to permit insertion of the inferior limbs of a human body thereinto and configured to sandwich the inferior abdomen thereof laterally;a driving device for driving said first saddle coil and said second saddle coil using a quadrature technique; anda synthesizing device for synthesizing two RF signals received by said first saddle coil and said second saddle coil respectively using the quadrature technique. 6. An RF signal transmitter-receiver of claim 5, further comprising a flexible circuit board that is shaped to be readily worn on the inferior abdomen of a human body and that has a circuit pattern relevant to said first saddle coil and a circuit pattern relevant to said second saddle coil printed thereon. 7. An RF signal receiver compris... | GE Medical Systems Global Technology Company, LLC | A61B005/05
US07027856 | B2 | 2006-04-11 | Method for determining a metric of non-sustained arrhythmia occurrence for use in arrhythmia prediction and automatic adjustment of arrhythmia detection parameters | 2002-09-30 | An implantable medical device and associated method are provided for detecting non-sustained arrhythmias and determining a metric of non-sustained arrhythmias. The metric may be used for predicting the occurrence of a sustained arrhythmia or for automatically adjusting the parameters used for detecting a sustained arrhythmia. | 1. A method for determining a metric of the occurrence of non-sustained arrhythmias comprising; detecting non-sustained arrhythmia episodes as episodes;storing characteristic data related to the non-sustained arrhythmia; andcalculating a metric of non-sustained arrhythmias based on the stored data; andadjusting a parameter for detecting a sustained arrhythmia based on the metric of non-sustained arrhythmia occurrence, wherein stored characteristic data includes one of a number of non-sustained arrhythmia episodes within a specified period of time, a duration of a non-sustained arrhythmia episode, a cycle length during a non-sustained arrhythmia episode, and a cardiac signal morphology during a non-sustained arrhythmia episode. | 1. A method for determining a metric of the occurrence of non-sustained arrhythmias comprising; detecting non-sustained arrhythmia episodes as episodes;storing characteristic data related to the non-sustained arrhythmia; andcalculating a metric of non-sustained arrhythmias based on the stored data; andadjusting a parameter for detecting a sustained arrhythmia based on the metric of non-sustained arrhythmia occurrence, wherein stored characteristic data includes one of a number of non-sustained arrhythmia episodes within a specified period of time, a duration of a non-sustained arrhythmia episode, a cycle length during a non-sustained arrhythmia episode, and a cardiac signal morphology during a non-sustained arrhythmia episode. 2. The method according to claim 1, further including a method for predicting the occurrence of a sustained arrhythmia based on the metric of non-sustained arrhythmias comprising: calculating an arrhythmia risk score based on a metric of non-sustained arrhythmias;comparing the risk score to a predefined or automatically updated threshold criteria; andpredicting that a sustained arrhythmia is likely to occur with a high probability if the risk score meets or exceeds the threshold criteria. 3. The method according to claim 2, further including delivering a preventative therapy if a sustained arrhythmia is predicted to occur. 4. The method according to claim 2, further including adjusting the threshold criteria based on a detected sustained arrhythmia. 5. The method according to claim 1, wherein the parameter for detecting a sustained arrhythmia is adjusted by increasing the number of cardiac cycle intervals occurring in an arrhythmia detection zone required to detect a sustained arrhythmia. 6. The method according to claim 5, wherein the parameter is adjusted by varying the duration of non-sustained arrhythmia episodes. | Medtronic, Inc. | A61B005/0452
US07013178 | B2 | 2006-03-14 | Implantable medical device communication system | 2002-09-25 | An implantable medical device communication system communicates information between an implantable medical device and at least one slave device by way of a two-wire bus. Slave devices may include remote sensors and other implantable medical devices. The implantable medical device includes a communication unit to combine data and power for transmission over the two wire bus. The transmitted signal is selectively changeable between a first and second voltage. The slave device includes a recovery unit to recover data and power from the received signal. An extendable command set includes long commands to set up the system and shorter commands to conserve power. Selectively addressable multicast commands, and shortened quick trigger commands conserve power by lowering system current and increasing data throughput. | 1. An implantable medical device communication system, comprising: an implantable medical device comprising a power supply, a controller, and a communication unit, wherein the communication unit combines data from the controller and power from the power supply to produce a communication signal that is selectively changeable between at least a first and a second voltage, and the communication unit produces a reference signal;a bus to receive the communication signal on a first wire and to receive the reference signal on a second wire; anda slave device comprising a recovery unit to recover power from a difference in potential between the communication signal and the reference signal transmitted across the first and second wires of said bus, and a transceiver unit to decode data from the selectively changeable communication signal transmitted across the first wire,wherein one bit time of data to be communicated across the first wire is defined as a predetermined number of clock cycles, a logic “0” is communicated by holding the first wire at the first voltage for a first fraction of the predetermined number of clock cycles and a logic “1” is communicated by holding the first wire at the first voltage for a second fraction of the predetermined number of clock cycles. | 1. An implantable medical device communication system, comprising: an implantable medical device comprising a power supply, a controller, and a communication unit, wherein the communication unit combines data from the controller and power from the power supply to produce a communication signal that is selectively changeable between at least a first and a second voltage, and the communication unit produces a reference signal;a bus to receive the communication signal on a first wire and to receive the reference signal on a second wire; anda slave device comprising a recovery unit to recover power from a difference in potential between the communication signal and the reference signal transmitted across the first and second wires of said bus, and a transceiver unit to decode data from the selectively changeable communication signal transmitted across the first wire,wherein one bit time of data to be communicated across the first wire is defined as a predetermined number of clock cycles, a logic “0” is communicated by holding the first wire at the first voltage for a first fraction of the predetermined number of clock cycles and a logic “1” is communicated by holding the first wire at the first voltage for a second fraction of the predetermined number of clock cycles. 2. The implantable medical device communication system according to claim 1, wherein the predetermined number of clock cycles is 6, the first fraction is ⅓ and the second fraction is ⅔, such that a logic “0” is communicated by holding the first wire at the first voltage for 2 clock cycles and a logic “1” is communicated by holding the first wire at the first voltage for 4 clock cycles. 3. The implantable medical device communication system according to claim 1, wherein the communication unit of said implantable medical device transmits logic information by selectively driving the first wire at the first voltage and selectively pulling the first wire to the second voltage at the first or second fraction of the predetermined number of clock cycles. 4. The implantable medical device communication system according to claim 3, wherein the communication unit transmits a logic “0” by selectively driving the first wire at the first voltage for 2 clock cycles and transmits a logic “1” by selectively driving the first wire at the first voltage for 4 clock cycles. 5. The implantable medical device communication system according to claim 3, wherein the transceiver unit of said slave device decodes the logic information by sensing the first or second fraction of the predetermined number of clock cycles. 6. The implantable medical device communication system according to claim 5, wherein the transceiver unit of said slave device decodes a logic “0” by sensing driving of the first wire at the first voltage for 2 clock cycles and decodes a logic “1” by sensing driving of the first wire at the first voltage for 4 clock cycles. 7. The implantable medical device communication system according to cla... | Medtronic, Inc. | A61N001/37 A61B005/04
US07282027 | B2 | 2007-10-16 | Service center system and method as a component of a population diagnostic for sleep disorders | 2002-08-07 | A method for evaluating sleep disorders in mammals and distributing control information related to the sleep disorders through electronic communication networks. The method includes acquiring analog information related to physiological signals through a mobile device from a mammal user. The mammal user is at a geographic location outside of a sleep laboratory staffed by trained technicians. The mobile device is coupled to a sensing device, which is coupled to monitoring region(s) of the mammal. The method includes converting the analog information into a digital format using a signal processing device and transmitting information related to the physiological signals in the digital format to one or more service computers through an electronic communications network. Each of the service computers is operably coupled to a health information service facility. The method determines if the mammal user has related control information at one of the health information service facilities. The related control information is received through the communication network. The method processes the information related to the physiological signals at one or more of the service computers to provide a report associated with the information related to the physiological signals and the mammal user. The method communicates report information in the report associated with the information related to the physiological signals upon indication of receipt of at least control account information of the mammal user based upon the determining the related account information. | 1. A method for evaluating sleep disorders in mammals and distributing control information related to the sleep disorders through electronic communication networks, the method comprising: acquiring analog information related to physiological signals through a mobile device from a mammal user, the mammal user being at a geographic location outside of a sleep laboratory staffed by trained technicians, the mobile device coupled to one or more sensing devices, the sensing device(s) being coupled to one or more monitoring regions of the mammal, the sensing device(s) comprising at least an external microphone device for acquiring the analog information related to the physiological signals;converting the analog information into a digital format using a signal processing device;transmitting information related to the physiological signals in the digital format to one or more service computers through an electronic communications network, one or more of the service computers being operably coupled to a health information service facility;determining if the mammal user has related control information at one of the health information service facilities, the related control information being received through the communication network;processing the information related to the physiological signals at one or more of the service computers to provide a report associated with the information related to the physiological signals and the mammal user; andcommunicating report information in the report associated with the information related to the physiological signals directed to evaluating a sleep disorder upon indication of receipt of at least control account information of the mammal user based upon the determining the related account information. | 1. A method for evaluating sleep disorders in mammals and distributing control information related to the sleep disorders through electronic communication networks, the method comprising: acquiring analog information related to physiological signals through a mobile device from a mammal user, the mammal user being at a geographic location outside of a sleep laboratory staffed by trained technicians, the mobile device coupled to one or more sensing devices, the sensing device(s) being coupled to one or more monitoring regions of the mammal, the sensing device(s) comprising at least an external microphone device for acquiring the analog information related to the physiological signals;converting the analog information into a digital format using a signal processing device;transmitting information related to the physiological signals in the digital format to one or more service computers through an electronic communications network, one or more of the service computers being operably coupled to a health information service facility;determining if the mammal user has related control information at one of the health information service facilities, the related control information being received through the communication network;processing the information related to the physiological signals at one or more of the service computers to provide a report associated with the information related to the physiological signals and the mammal user; andcommunicating report information in the report associated with the information related to the physiological signals directed to evaluating a sleep disorder upon indication of receipt of at least control account information of the mammal user based upon the determining the related account information. 2. The method of claim 1 wherein the geographic location is the home or residence of the mammal user. 3. The method of claim 1 wherein the geographic location is a nursing home or skilled nursing facility. 4. The method of claim 1 wherein the geographic location is a hospital room. 5. The method of claim 1 wherein the processing includes identifying a sleep disorder, the sleep disorder being selected from at least one selected from sleep apnea, the upper airway resistance syndrome, periodic leg movements of sleep, or Cheyne-Stokes respiration. 6. The method of claim 1 wherein the processing includes identifying a sleep disorder, the sleep disorder being snoring. 7. The method of claim 1 wherein the determining is provided before processing. 8. The method of claim 1 wherein the determining is provided after or concurrently with processing. 9. The method of claim 1 wherein the one or more of the monitoring regions includes a neck portion. 10. The method of claim 1 wherein one or more of the monitoring regions includes a suprasternal notch. 11. The method of claim 1 wherein one or more of the monitoring regions is selected from regions of the thorax including the pericardium, the sternal area, the anterior chest,... | Apneos Corporation | A61B005/00
US07264627 | B2 | 2007-09-04 | Wicking methods and structures for use in sampling bodily fluids | 2002-08-29 | A sampling member defines a passageway including an inlet opening for reception of a bodily fluid therein, and a wicking member having a first portion coupled with the sampling member adjacent to the inlet opening. The wicking member further includes a second portion displaced from the first portion for initially contacting the bodily fluid. In use the sampling member is placed adjacent to an incision site with the second portion of the wicking member positioned to have the bodily fluid first contact the wicking member and thereby draw the fluid from the incision site to the inlet opening. In related embodiments there is provided a sampling system and method which further includes a lancing device to create the incision and optionally expression members to urge the bodily fluid from the incision, preferably in an integrated unit. | 1. A device, comprising: an incision forming device configured to form an incision;an analysis area configured to analyze bodily fluid from the incision;a sampling device defining a capillary passageway that extends from the analysis area for drawing the bodily fluid via capillary action to the analysis area;a wicking member extending from the sampling device proximal the capillary passageway to draw the bodily fluid from the incision to the sampling passageway, wherein the wicking member is an elongate member that is hydrophilic to draw the bodily fluid; andwherein the wicking member includes a folded section configured to fold against the sampling device until deployment. | 1. A device, comprising: an incision forming device configured to form an incision;an analysis area configured to analyze bodily fluid from the incision;a sampling device defining a capillary passageway that extends from the analysis area for drawing the bodily fluid via capillary action to the analysis area;a wicking member extending from the sampling device proximal the capillary passageway to draw the bodily fluid from the incision to the sampling passageway, wherein the wicking member is an elongate member that is hydrophilic to draw the bodily fluid; andwherein the wicking member includes a folded section configured to fold against the sampling device until deployment. 2. The device of claim 1, wherein the wicking member has a uniform size and shape. 3. The sampling device of claim 1, wherein: the sampling device has a top surface, a bottom surface and an aperture extending from the top surface to the bottom surface, the aperture being configured to surround the incision;the capillary passageway opens into the aperture; andthe wicking member extends within the aperture. 4. The sampling device of claim 3, further comprising a sealing member surrounding the aperture at the bottom surface of the sampling device to minimize fluid leakage. 5. The sampling device of claim 4, wherein the sealing member includes a hydrophobic surface to resist passage of the bodily fluid. 6. The sampling device of claim 4, further comprising a constricting system extending from the bottom surface to create a bulge of skin around the incision. 7. The sampling device of claim 3, further comprising a constricting system provided around the aperture to create a bulge of skin around the incision. 8. The sampling device of claim 7, wherein the constricting system includes at least two deformable elements positioned on opposite sides of the aperture for engaging and constricting the skin. 9. The sampling device of claim 3, wherein the sampling device has a recessed surface between the capillary passageway and the bottom surface to promote fluid flow into the capillary passageway. 10. The sampling device of claim 1, wherein the sampling device has a constricting system extending therefrom to create a bulge of skin around the incision. 11. The sampling device of claim 1, wherein the analysis area includes a test reagent covering absorbent material. 12. The sampling device of claim 1, wherein: the sampling device has a top surface, a bottom surface and an end edge; andthe capillary passageway communicates with the end edge at a location displaced from the bottom surface. 13. The sampling device of claim 12, wherein the sampling device has a recessed surface connecting between the capillary passageway and the bottom surface. 14. The sampling device of claim 12, further comprising a sealing member located on the bottom surface to block passage of the bodily fluid. 15. The sampling device of claim 1, further comprising: wherein the sampling device defines an apertu... | Roche Diagnostics Operations, Inc. | A61B005/00
US07083577 | B2 | 2006-08-01 | Guide wire | 2002-10-24 | A guide wire having an elongated core wire provided with a tabular distal end portion, and a coiled wire provided on an outer circumference of the core wire, wherein the distal end portion of the core wire is provided with a cutout section or a rased section so that a diametrical cross-sectional area of the distal end portion of the core wire gradually decreases toward a distal end thereof. | 1. A guide wire comprising an elongated core wire having a distal end and a proximal end, and a coil wire provided on an outer circumference of the distal end of the core wire, said distal end comprising a tabular shaped section at a distal most end thereof, characterized in that said tabular shaped section is provided with a cutout section extending in an axial direction of the core wire so that a diametrical cross-sectional area of said tabular shaped section gradually decreases toward the distal most end of the core wire, said cutout section being formed by cutting out of said tabular shaped section a triangular shape having a base at a distal end of the cutout section and an apex at a proximal end of the cutout section. | 1. A guide wire comprising an elongated core wire having a distal end and a proximal end, and a coil wire provided on an outer circumference of the distal end of the core wire, said distal end comprising a tabular shaped section at a distal most end thereof, characterized in that said tabular shaped section is provided with a cutout section extending in an axial direction of the core wire so that a diametrical cross-sectional area of said tabular shaped section gradually decreases toward the distal most end of the core wire, said cutout section being formed by cutting out of said tabular shaped section a triangular shape having a base at a distal end of the cutout section and an apex at a proximal end of the cutout section. 2. The guide wire according to claim 1, wherein said tabular shaped section is further provided with at least two rased sections which are symmetrical with respect to a longitudinal axis of said tabular shaped section. 3. The guide wire according to claim 1, wherein said cutout section is formed so as to extend axially 10 to 70% the length of said tabular shaped section from said distal end toward the proximal end. 4. The guide wire according to claim 1, wherein the diametrical cross-sectional area of said tabular shaped section phases down toward said distal end thereof. 5. The guide wire according to claim 1, wherein said tabular shaped section of the core wire is composed of a material different from that of the core wire. | Nipro Corporation | A61B005/00 A61M025/00
US07228168 | B2 | 2007-06-05 | Systems and methods for assessing and modifying an individual's physiological condition | 2002-05-21 | Systems and methods for assessing an individual's physiological condition are provided. Cycle and shape parameters are derived from a recorded time trace containing heart rate data collected while an individual performs a cyclic exercise routine. Individually tailored exercise regimens that are based on these parameters are generated and modified as desired. | 1. A computer-readable medium containing instructions for prescribing an exercise regimen for an individual, said instructions comprising: instructions to receive a metric indicative of said individual's physiological condition;instructions to generate said exercise regimen using an algorithm that relates said metric to said regimen;instructions to export said exercise regimen to said individual using an output device; andinstructions for generating a guide based on said regimen, said guide directing said individual in real-time on how to exercise, wherein said metric is based on heart rate information of an individual obtained during an exercise test. | 1. A computer-readable medium containing instructions for prescribing an exercise regimen for an individual, said instructions comprising: instructions to receive a metric indicative of said individual's physiological condition;instructions to generate said exercise regimen using an algorithm that relates said metric to said regimen;instructions to export said exercise regimen to said individual using an output device; andinstructions for generating a guide based on said regimen, said guide directing said individual in real-time on how to exercise, wherein said metric is based on heart rate information of an individual obtained during an exercise test. 2. The computer-readable medium of claim 1 wherein said exercise regimen comprises a plurality of exercise sessions to be performed by said individual, wherein each of said plurality of exercise sessions comprising: at least one warm up cycle; andat least one critical cycle that has a target heart rate, wherein said target heart rate is based on said metric. 3. The computer-readable medium of claim 1 further comprising instructions to determine a capacity index based on an electronic file containing a time trace of a heart rate of said individual obtained during said exercise test that includes a plurality of test cycles and at least one maximum effort cycle, wherein said instructions to determine a capacity index comprises: instructions to determine at least one maximum effort cycle parameter; andinstructions to determine said capacity index based on said at least one maximum effort cycle parameter. 4. The computer-readable medium of claim 3, wherein said maximum effort cycle parameter is selected from a group consisting of a first peak heart rate corresponding to a maximum heart rate monitored during said maximum effort cycle. 5. The computer-readable medium of claim 3, wherein said instructions to determine a capacity index further comprises instructions to determine another parameter for use in said capacity index, said another parameter being selected from a group consisting of a second peak heart rate monitored during another of said plurality of test cycles that is not said maximum effort cycle, an upward slope of said heart rate monitored during a part of a stress portion of at least one of said plurality of test cycles, and a downward slope of said heart rate monitored during a part of a relaxation portion of the at least one of said plurality of test cycles. 6. A computer-readable medium containing instructions for prescribing an exercise regimen for an individual, said instructions comprising: instructions to receive a metric indicative of said individual's physiological condition;instructions to generate said exercise regimen using an algorithm that relates said metric to said regimen;instructions to export said exercise regimen to said individual using an output device;instructions for generating a guide based on said regimen, said guide directing said individual in real-time on h... | Lifewaves International, Inc. | A61B005/04
US07162293 | B2 | 2007-01-09 | MR arrangement for localizing a medical instrument | 2002-10-10 | The invention relates to an MR arrangement which includes a medical instrument (3), an RF arrangement (17) which is provided thereon and in which at least one microcoil (1) and a capacitor (2) are connected so as to form a resonant circuit, and also an optical signal lead (4) via which a light signal is applied to the RF arrangement. In order to achieve simple and fast determination of the position of the medical instrument (3) within the examination zone of an MR apparatus, the invention proposes to provide a modulator (20) for modulating the light signal applied to the RF arrangement and also an optoelectrical converter (5) which converts the modulated light signal into an electrical signal and is coupled to the resonant circuit in such a manner that the resonant circuit is triggered by the electrical signal from the optoelectrical converter (5) so as to emit an RF field. | 1. An MR apparatus comprising: a medical instrument;an RF arrangement disposed on the medical instrument, the RF arrangement comprising: at least one microcoil; anda capacitor connected to the microcoil so as to form a resonant circuit;an optical signal lead via which a light signal is applied to the RF arrangement;a modulator for modulating the light signal applied to the RF arrangement; andan optoelectrical converter for converting the modulated light signal into an electrical signal, the optoelectrical converter being coupled to the resonant circuit in such a manner that the resonant circuit is triggered by the electrical signal from the optoelectrical converter so as to emit an RF field. | 1. An MR apparatus comprising: a medical instrument;an RF arrangement disposed on the medical instrument, the RF arrangement comprising: at least one microcoil; anda capacitor connected to the microcoil so as to form a resonant circuit;an optical signal lead via which a light signal is applied to the RF arrangement;a modulator for modulating the light signal applied to the RF arrangement; andan optoelectrical converter for converting the modulated light signal into an electrical signal, the optoelectrical converter being coupled to the resonant circuit in such a manner that the resonant circuit is triggered by the electrical signal from the optoelectrical converter so as to emit an RF field. 2. An MR arrangement as claimed in claim 1, wherein the optoelectrical converter is a photodiode. 3. An MR arrangement as claimed in claim 2, wherein the microcoil, the capacitor and the photodiode are connected in parallel in the resonant circuit. 4. A method of localizing a medical instrument in an examination volume of an MR apparatus having an RF frequency, the method comprising: generating a light signal;coupling the light signal along an optical signal lead to a optoclectrical converter;converting the light signal to a photocurrent using the optoelectrical converter; andemitting an RF field in the vicinity of the medical instrument in response to the photocurrent using a resonant circuit disposed on the medical instrument wherein the light signal is modulated with the resonance frequency of the MR apparatus into the optical signal lead. 5. The method according to claim 4 wherein the resonant circuit comprises at least one coil and at least one capacitor. 6. The method according to claim 4 wherein the RF field in the vicinity of the medical instrument is at the resonance frequency of the MR apparatus. 7. A method as claimed in claim 4 wherein the modulated light signal is coupled into the optical signal lead in a pulsed fashion, the MR signal generated in the vicinity of the microcoil being detected each time subsequent to a light pulse in order to be evaluated for the determination of the position. 8. A method as claimed in claim 4 wherein the light signal coupled into the optical signal lead is modulated with the resonance frequency of the MR apparatus in a pulsed fashion, the MR signal generated in the vicinity of the microcoil being detected each time in the time intervals in which the light signal is not modulated. 9. A medical instrument comprising: an RF arrangement disposed at the distal end of the medical instrument, the RF arrangement comprising at least one microcoil and a capacitor connected so as to form a resonant circuit; andan optical signal lead via which a modulated light signal is applied to the RF arrangement;wherein the resonant circuit is coupled to a photodiode, in which the modulated light signal generates a photocurrent, in such a manner that the resonant circuit is triggered by the photocurrent so as to emit an RF field.... | - | A61B005/055
US07117026 | B2 | 2006-10-03 | Physiological model based non-rigid image registration | 2002-06-12 | A method for non-rigid registration and fusion of images with physiological modeled organ motions resulting from respiratory motion and cardiac motion that are mathematically modeled with physiological constraints. A method of combining images comprises the steps of obtaining a first image dataset (24) of a region of interest of a subject and obtaining a second image dataset (34) of the region of interest of the subject. Next, a general model of physiological motion for the region of interest is provided (142). The general model of physiological motion is adapted with data derived from the first image data set (140) to provide a subject specific physiological model (154). The subject specific physiological model is applied (172) to the second image dataset (150) to provide a combined image (122). | 1. A method of combining images, the method comprising the steps of: obtaining a first image dataset from a subject; obtaining a second image dataset from the subject; segmenting the first image data set; segmenting the second image data set; providing a physiological model; morphing the physiological model to the segmented first image data set to produce a subject specific physiological phantom; registering the subject specific physiological phantom with the segmented second image data set to generate a transform; and applying the transform to one of the first and second image data sets to provide a fused image. | 1. A method of combining images, the method comprising the steps of: obtaining a first image dataset from a subject; obtaining a second image dataset from the subject; segmenting the first image data set; segmenting the second image data set; providing a physiological model; morphing the physiological model to the segmented first image data set to produce a subject specific physiological phantom; registering the subject specific physiological phantom with the segmented second image data set to generate a transform; and applying the transform to one of the first and second image data sets to provide a fused image. 2. A method of combining images, the method comprising the steps of: obtaining a first image dataset of a region of interest of a subject; obtaining a second image dataset of the region of interest of the subject; providing a general model of physiological motion for the region of interest; adapting the general model of physiological motion with data derived from the first image data set to provide a subject specific physiological model; and applying the subject specific physiological model to the second image dataset to provide a combined image. 3. The method of combining images of claim 2 wherein the first image dataset is obtained from a first imaging modality and the second image dataset is obtained from a second imaging modality different than the first imaging modality. 4. The method of combining images of claim 3 wherein one of the first and second image datasets is primarily functional image data and the other image dataset is primarily anatomical image data. 5. The method of combining images of claim 2 wherein the first image dataset is obtained from a discrete imaging and the second image dataset is obtained from a discrete imaging system. 6. A method of combining images, the method comprising the steps of: obtaining a computed tomography image dataset from a subject; obtaining a nuclear medicine image dataset from the subject; segmenting the computed tomography image data set; segmenting the nuclear medicine image data set; providing a general physiological model; morphing the physiological model to the segmented computed tomography image data set to produce a subject specific physiological phantom; registering the subject specific physiological phantom with the segmented nuclear medicine image data set to generate a transform; and applying the transform to the computed tomography image data set to provide a fused image. 7. A method of combining images, the method comprising the steps of: obtaining a computed tomography image dataset from a subject; obtaining a nuclear medicine image dataset from the subject; segmenting the computed tomography image data set; segmenting the nuclear medicine image data set; providing a general physiological model; morphing the physiological model to the segmented computed tomography image data set to produce a subject specific physiological phantom; registering the subject specific phy... | Koninklijke Philips Electronics N.V. | A61B005/05
US07006860 | B2 | 2006-02-28 | Magnetic resonance imaging coil having vest support structure | 2002-04-12 | The vest provides a simple and intuitive coil support assuring proper alignment of the coils with the patient and providing support for coil elements encircling the opposed arms.;A coil for receiving and/or transmitting radio-frequency signals for magnetic resonance imaging of the chest, and particularly for ventilation studies of the lungs, is supported in a flexible vest having arm holes locating the coil elements about the patient. | 1. A vest coil comprising; a vest structure wearable by a patient and including opposed arm holes for receiving the patient's arms therethrough; anda coil structure supported by the vest to receive magnetic resonance image signals from the region of the patient's lungs wherein the coil structure includes at least one electrically continuous coil encircling an arm hole. | 1. A vest coil comprising; a vest structure wearable by a patient and including opposed arm holes for receiving the patient's arms therethrough; anda coil structure supported by the vest to receive magnetic resonance image signals from the region of the patient's lungs wherein the coil structure includes at least one electrically continuous coil encircling an arm hole. 2. The vest coil of claim 1 wherein the coil structure includes two coils, one passing around each of the opposed arm holes. 3. The vest coil of claim 2 wherein the vest structure includes side panels having the armholes and a front and back panel extending between the side panels and wherein the front and back panels include coils. 4. The vest coil of claim 3 wherein the coil elements encircling each of the opposed arm holes provides a first Helmholtz coil pair and wherein the front and back panels support a second Helmholtz coil pair perpendicular to the first. 5. The vest coil of claim 3 wherein the coil element encircling each of the opposed armholes and the coil elements in the front and back panels provide independent signals for phased array studies. 6. The vest coil of claim 1 wherein the vest structure is constructed of a flexible material to wrap around the patient. 7. The vest coil of claim 1 wherein the vest structure includes a separable portion disposed between the opposed armholes along an opening of the vest structure so that the vest structure may be put on by a patient in part by wrapping it about the patient. 8. The vest coil of claim 7 wherein the separable portion includes a closure means holding the separable portions in varying spatial relationships when the vest structure is wrapped about the patient to accommodate patients of different chest sizes. 9. The vest coil of claim 1 wherein the closure means includes engaging strips of hook and loop fasteners. 10. The vest coil of claim 1 wherein the flexible vest structure includes an outer fabric shell supporting therein a set of layered, independently slidable flexible leaves including at least one flexible printed circuit board. 11. The vest coil of claim 10 wherein at least one of the leaves is an elastomeric plastic foam. 12. The vest coil of claim 11 wherein the elastomeric plastic foam includes a sliding layer on at least one surface. 13. The vest coil of claim 1 including further a back tab extending downward from the remainder of the vest structure at the rear of the vest holding signal processing electronics. 14. A method of performing a lung ventilation study comprising the steps of: (a) introducing a hyperpolarized gas into a patient's lungs; and(b) measuring a magnetic resonance signal from the hyperpolarized gas using a vest structure wearable by a patient and including opposed armholes for receiving the patient's arms therethrough and a coil structure supported by the vest to receive magnetic resonance image signals from the region of the patient's lungs wherein the coil structure incl... | Invivo Corporation | A61B005/05
US07103930 | B1 | 2006-09-12 | Device for the calibration and standardization of hip rotation | 2002-03-21 | The present invention concerns a device for the calibration and standardization of the rotation and movement of a patient's hip. More particularly the present invention relates to a device for the calibration and standardization of the rotation and movement of the hip for use in connection with improving the reproducibility and accuracy of a diagnostic procedure such as, for example, bone density measurement using dual x-ray absorptiometry (DXA). | 1. An apparatus for the calibration and standardization of the rotation and movement of a patient's hip, said device comprising: a table having a horizontal upper surface for receiving a patient,a base member positioned on said table upper surface and having an open upper portion; anda saddle member having sides; an upper curved surface that is capable of engaging a knee of the patient; a lower rigid curved surface that rotates within the base member open upper portion to position the patient's knee; and opposing side surfaces between the upper and lower curved surfaces,wherein each side surface of the saddle member has an arcuate groove for slidingly engaging the open upper portion of the base member and the upper portion of the base member has protrusions for slidingly engaging the grooves of the saddle member. | 1. An apparatus for the calibration and standardization of the rotation and movement of a patient's hip, said device comprising: a table having a horizontal upper surface for receiving a patient,a base member positioned on said table upper surface and having an open upper portion; anda saddle member having sides; an upper curved surface that is capable of engaging a knee of the patient; a lower rigid curved surface that rotates within the base member open upper portion to position the patient's knee; and opposing side surfaces between the upper and lower curved surfaces,wherein each side surface of the saddle member has an arcuate groove for slidingly engaging the open upper portion of the base member and the upper portion of the base member has protrusions for slidingly engaging the grooves of the saddle member. 2. The apparatus of claim 1, wherein the saddle member is provided with a handle on its lower rigid curved surface. 3. The apparatus of claim 1, wherein the saddle member is provided with a series of holes and a mating pin to retain the saddle member in a desired position. 4. The apparatus of claim 1, wherein the inner surface of the saddle member is provided with a firm foam retaining means for securing the patient's knee. 5. The apparatus of claim 1, wherein positioning of the saddle member is controlled manually. 6. The apparatus of claim 1, wherein positioning of the saddle member is controlled via a motor and the location of the saddle member with respect to the base is synchronized with a digital output which is displayed on a monitor. 7. The apparatus of claim 1 which also comprises means for performing a medical orthopedic or diagnostic procedure. 8. The apparatus of claim 7, wherein the means is capable of measuring bone density. 9. The apparatus of claim 8, wherein the means is capable of measuring bone density using dual x-ray absorptiometry. 10. The apparatus of claim 7, wherein the means is capable of performing an MRI or CAT scan. 11. In a medical procedure that concerns the calibration and standardization of the rotation and movement of a patient's hip to measure bone density using dual x-ray absorptiometry, the improvement wherein a knee of the patient is positioned in a device comprising a base member having an open upper portion and a saddle member having sides; an upper curved surface that is capable of engaging the knee of the patient; a lower rigid curved surface that rotates within the base member open upper portion to position said knee; and opposing side surfaces between the upper and lower curved surfaces, wherein each side surface of the saddle member has an arcuate groove for slidingly engaging the open upper portion of the base member and the upper portion of the base member has protrusions for slidingly engaging the grooves of the saddle member. 12. A device for the calibration and standardization of the rotation and movement of a patient's hip, said device comprising: a base member having a open u... | The Trustees of Columbia University | A61B005/055
US07291106 | B2 | 2007-11-06 | Diagnostic system and portable telephone device | 2002-07-31 | A patient registers his audible sound when he is in good health. When he is not feeling well, and even when he is in good health, he may contact regularly with a medical institution. Once he pronounces words to his mobile phone, the audible sound will be stored in a storing portion for storing audible sound information of the patient and will be processed by a computer that gives diagnosis to the patient and the patient will be informed of the diagnosis through a public line. | 1. A diagnosing system diagnosing a patient to communicate with a computer of a medical institution through a public line, comprising: a stethoscope having a sound pick-up portion picking up a sound information and an image pick-up device picking up image information of the patient which is disposed inside the sound pick-up portion;a signal transformer processing the sound information and the image information digitally, being connected with the computer, and transmitting the processed sound and image information digitally through the public line;wherein the computer has:receiving means for receiving the transmitted sound information and image information transmitted by the signal transformer; andfirst transmitting means for transmitting diagnosis information obtained from the medical institution based on the received sound information and image information by the receiving means. | 1. A diagnosing system diagnosing a patient to communicate with a computer of a medical institution through a public line, comprising: a stethoscope having a sound pick-up portion picking up a sound information and an image pick-up device picking up image information of the patient which is disposed inside the sound pick-up portion;a signal transformer processing the sound information and the image information digitally, being connected with the computer, and transmitting the processed sound and image information digitally through the public line;wherein the computer has:receiving means for receiving the transmitted sound information and image information transmitted by the signal transformer; andfirst transmitting means for transmitting diagnosis information obtained from the medical institution based on the received sound information and image information by the receiving means. 2. The diagnosing system as set forth in claim 1, further comprising: audible sound information transmitting means for transmitting audible sound information actually pronounced by the patient to the computer;wherein the computer has:storing means for storing the audible sound information of the patient obtained when the patient is in good health, as first audible sound information;audible sound information receiving means for receiving the transmitted audible sound information;analyzing means for comparing second audible sound information of the patient with the first audible sound information stored in the storing means and obtaining an audible sound-diagnosis information; andsecond transmitting means for transmitting the obtained audible sound-diagnosis information to the patient. 3. The diagnosing system as set forth in claim 2, wherein the audible sound information transmitting means is a mobile phone carried by the patient. 4. The diagnosing system as set forth in claim 1, wherein the image pick-up device is a charge coupled device (CCD) or a complementary metal oxide semiconductor (CMOS) device. 5. The diagnosing system as set forth in claim 1, further comprising: an USB interface capable of outputting at least one of the sound information picked up by the sound pick-up portion and the image information picked-up by the image pick-up device. | - | A61B005/00
US07027852 | B2 | 2006-04-11 | Lead with distal tip surface electrodes connected in parallel | 2002-05-21 | An body implantable stimulation lead is provided including tip electrode patterns and configurations producing low noise, clean sensed signals devoid of far field components, such sensed signals being generated irrespective of the direction of the incident depolarization wavefront. The invention also provides high pacing impedances and advantageous anode-to-cathode surface area ratios. Further, implantable leads utilizing the features of the present invention are particularly suitable for left side stimulation therapies. | 1. A body implantable lead for transmitting electrical signals between an electrical connector at a proximal end of the lead and selected body tissue, the electrical connector being adapted to be received by a receptacle in an implantable medical device, the lead comprising: a lead body having a distal end;a distal tip surface at the distal end of the lead body, the distal tip surface carrying at least two parallel-connected anode sensing electrodes, and further carrying at least one cathode sensing electrode, the respective electrodes being adapted to electrically communicate with the selected body tissue;a first conductor electrically coupling the at least two anode sensing electrodes with the connector assembly; anda second conductor electrically coupling the at least one cathode sensing electrode with the connector assembly;wherein each of the first and second conductors comprises a pair of redundant conductors. | 1. A body implantable lead for transmitting electrical signals between an electrical connector at a proximal end of the lead and selected body tissue, the electrical connector being adapted to be received by a receptacle in an implantable medical device, the lead comprising: a lead body having a distal end;a distal tip surface at the distal end of the lead body, the distal tip surface carrying at least two parallel-connected anode sensing electrodes, and further carrying at least one cathode sensing electrode, the respective electrodes being adapted to electrically communicate with the selected body tissue;a first conductor electrically coupling the at least two anode sensing electrodes with the connector assembly; anda second conductor electrically coupling the at least one cathode sensing electrode with the connector assembly;wherein each of the first and second conductors comprises a pair of redundant conductors. 2. The lead of claim 1 in which: the respective electrodes are configured and positioned to generate a net voltage output in response to an incident depolarization wavefront, the net voltage output being sufficient to be acknowledged by the implantable medical device. 3. The lead of claim 2 in which: the respective electrodes are arranged on the distal tip surface so as to generate The net voltage output irrespective of the direction of the incident depolarization wavefront. 4. The lead of claim 1 in which: the at least two anode sensing electrodes have a total active surface area greater than the total active surface area of the at least one cathode sensing electrode. 5. The lead of claim 4 in which: the total active surface area of the at least two anode sensing electrodes is at least about twice that of the total active surface area of the at least one cathode sensing electrode. 6. The lead of claim 5 in which: the total active surface area of the at least two parallel-connected electrodes is at least about 1.0 mm2 and the total active surface area of the at least one additional electrode is about 0.3 mm2. 7. The lead of claim 4 in which: the active surface area of each of the at least two anode sensing electrodes and the at least one cathode sensing electrode are substantially equal. 8. The lead of claim 1 in which: all of the electrodes are arranged in a closely spaced cluster so as to sense local electrical events and not far field signals. 9. The lead of claim 1 in which: the electrodes have active surface areas providing an impedance of at least about 1,500 ohms. 10. The lead of claim 1 in which: each of the electrodes has a small active surface area, the electrodes being arranged in a closely spaced cluster, the electrodes being substantially non-responsive to far field signals. 11. The lead of claim 1 in which: the distal end of the lead carries an extendable/retractable screw-in helix for anchoring the distal end in tissue within the right side of the heart. 12. The lead of claim 1 in which: each of the electrode... | Pacesetter, Inc. | A61B005/042 A61N001/05
US07207945 | B2 | 2007-04-24 | System and method for providing diagnosis and monitoring of respiratory insufficiency for use in automated patient care | 2002-05-20 | An automated system and method for diagnosing and monitoring respiratory insufficiency and outcomes thereof is described. A plurality of monitoring sets is retrieved from a database. Each of the monitoring sets include recorded measures relating to patient information recorded on a substantially continuous basis. A patient status change is determined by comparing at least one recorded measure from each of the monitoring sets to at least one other recorded measure. Both recorded measures relate to the same type of patient information. Each patient status change is tested against an indicator threshold corresponding to the same type of patient information as the recorded measures which were compared. The indicator threshold corresponds to a quantifiable physiological measure of a pathophysiology indicative of respiratory insufficiency. | 1. A system for providing diagnosis and monitoring of respiratory insufficiency for use in automated patient care, comprising: a comparison module comparing at least one recorded physiological measure to at least one other recorded physiological measure on a substantially regular basis to quantify a change in patient pathophysiological status for equivalent patient information; andan analysis module evaluating an absence, an onset, a progression, a regression, and a status quo of respiratory insufficiency dependent upon the change in patient pathophysiological status. | 1. A system for providing diagnosis and monitoring of respiratory insufficiency for use in automated patient care, comprising: a comparison module comparing at least one recorded physiological measure to at least one other recorded physiological measure on a substantially regular basis to quantify a change in patient pathophysiological status for equivalent patient information; andan analysis module evaluating an absence, an onset, a progression, a regression, and a status quo of respiratory insufficiency dependent upon the change in patient pathophysiological status. 2. A system according to claim 1, further comprising: a diagnostic module comparing the change in patient pathophysiological status to an indicator threshold corresponding to a quantifiable physiological measure indicative of respiratory insufficiency. 3. A system according to claim 1, further comprising: a management module managing the respiratory insufficiency through administration of at least one of antibiotic and antiviral therapies, bronchodilator therapies, oxygen therapies, antiinflammation therapies, electrical therapies, and mechanical therapies. 4. A system according to claim 1, further comprising: a database module retrieving the at least one recorded physiological measure and the at least one other recorded physiological measure from monitoring sets stored in a database. 5. A system according to claim 4, further comprising: a server system collecting the at least one recorded physiological measure and the at least one other recorded physiological measure into each monitoring set recorded on a substantially continuous basis or derived therefrom. 6. A system according to claim 5, further comprising: at least one of an implantable medical device and an external medical device recording the at least one recorded physiological measure and the at least one other recorded physiological measure. 7. A system according to claim 1, further comprising: the analysis module evaluating an absence, an onset, a progression, a regression, and a status quo of diseases other than respiratory insufficiency dependent upon the change in patient pathophysiological status. 8. A system according to claim 1, further comprising: a diagnostic module comparing at least one recorded quality of life measure to at least one other recorded quality of life measure on a substantially regular basis to qualify a change in patient pathophysiological status. 9. A system according to claim 1, further comprising: a stored stickiness Indicator defined for at least one physiological measure corresponding to a temporal boundary on one of patient diagnosis and treatment;a diagnostic module timing each change in patient pathophysiological status for the equivalent patient information and determining one of a revised patient diagnosis and treatment responsive to each change in patient pathophysiological status with a tuning exceeding the stickiness indicator. 10. A system according to claim 1, further compr... | Cardiac Intelligence Corporation | A61B005/08 A61B005/02 A61B005/00
US07103400 | B2 | 2006-09-05 | Artifact elimination in time-gated anatomical imaging | 2002-11-08 | A method for anatomical imaging includes acquiring image data representative of three-dimensional volume segments of an image volume of interest in a subject. The image data are acquired in synchronism with corresponding physiological cycles of the subject. Each volume segment contains image data distributed in three dimensions. Acquiring image data includes selecting a sequence of scan lines for each respective volume segment configured to minimize an occurrence of motion artifacts throughout the image volume. The image data representative of the volume segments is combined to produce a representation of a three-dimensional anatomical image of the image volume. | 1. A computer implemented method for anatomical imaging, comprising: a) acquiring image data representative of three-dimensional, individually-continuous volume segments of an image volume of interest in a subject, in synchronism with corresponding physiological cycles of the subject, each of the volume segments containing image data distributed in three dimensions, wherein, for a pair of spatially consecutive ones of said volume segments, respective sequences of scan lines traversing said ones, said acquiring pair-wise mirror-images said sequences to minimize an occurrence of motion artifacts throughout the image volume; andb) combining the image data representative of the volume segments to produce a representation of a three-dimensional anatomical image of the image volume. | 1. A computer implemented method for anatomical imaging, comprising: a) acquiring image data representative of three-dimensional, individually-continuous volume segments of an image volume of interest in a subject, in synchronism with corresponding physiological cycles of the subject, each of the volume segments containing image data distributed in three dimensions, wherein, for a pair of spatially consecutive ones of said volume segments, respective sequences of scan lines traversing said ones, said acquiring pair-wise mirror-images said sequences to minimize an occurrence of motion artifacts throughout the image volume; andb) combining the image data representative of the volume segments to produce a representation of a three-dimensional anatomical image of the image volume. 2. The method of claim 1, wherein acquiring image data further includes selecting the sequence of scan lines within each respective volume segment to minimize the occurrence of motion artifacts across adjacent volume segment boundaries. 3. The method of claim 1, wherein the motion artifacts include one selected from the group consisting of color flow and angio artifacts. 4. The method of claim 1, wherein the anatomical imaging includes one selected from the group consisting of ultrasound, magnetic resonance (MR), and CT imaging. 5. The method of claim 4, wherein the anatomical imaging includes ultrasound imaging, said method further comprising: completing the ultrasound imaging of a given volume size in less than ten (10) seconds. 6. The method of claim 1, further comprising: repeating said acquiring and said combining to generate a sequence of images; andforming one of the following selected from the group consisting of a time history of images and a motion picture of images. 7. The method of claim 1, wherein acquiring image data further includes configuring the sequence of scan lines within each respective volume segment to minimize time discontinuities within each volume segment, in addition to minimizing time discontinuities across adjacent volume segment boundaries. 8. The method of claim 1, wherein acquiring image data further includes configuring individual scan lines of a sequence to reduce a local time gradient throughout the volume segment to less than ten percent (10%) of a total time for acquiring image data for each volume segment. 9. The method of claim 1, wherein boundary sides of said ones of the volume segments are acquired at approximately the same exact phase of the physiological cycle when the anatomical feature is in the same position during a single physiological cycle of the subject. 10. The method of claim 1, wherein acquiring image data comprises synchronizing acquisition of the image data to a selected phase of the subject's physiological cycle, and synchronizing acquisition of the image data includes one selected from the group consisting of cardiac gating and respiratory gating. 11. A system for anatomical imaging, comprising: an acquis... | Koninklijke Philips Electronics, N.V. | A61B005/05
US07069071 | B2 | 2006-06-27 | Distributed multi-user replication-based system for real time data access during cardiology procedures | 2000-12-29 | A distributed multi-user system for real time data access during cardiology procedures. The system includes an interactive computer network which can be used to simultaneously display data from a cardiology procedure on a plurality of devices and at a plurality of locations. The system is implemented through various replication topologies which allow the simultaneous access a cardiology study during the procedure. The study can be displayed at any of the plurality of locations, which may be local or remote, during the procedure. | 1. A method of accessing a study record taken during a cardiac catheterization procedure, the procedure being conducted in a cardiac catheterization lab, comprising the acts of (a) inserting at least one catheter into a patient comprising a heart, the catheters terminating in a position proximate to the heart and comprising one or more sensors configured to sense data from the heart;(b) stimulating the heart with electrical signals from the catheter;(c) sensing data from the heart;(d) transmitting the data from the sensors to a data collection device;(e) transmitting the data from the data collection device to a central publisher;(f) replicating the data;(g) transmitting the replicated data from the central publisher to a plurality of client workstations; and(h) simultaneously displaying the data on the plurality of client workstations during the cardiac catheterization procedure. | 1. A method of accessing a study record taken during a cardiac catheterization procedure, the procedure being conducted in a cardiac catheterization lab, comprising the acts of (a) inserting at least one catheter into a patient comprising a heart, the catheters terminating in a position proximate to the heart and comprising one or more sensors configured to sense data from the heart;(b) stimulating the heart with electrical signals from the catheter;(c) sensing data from the heart;(d) transmitting the data from the sensors to a data collection device;(e) transmitting the data from the data collection device to a central publisher;(f) replicating the data;(g) transmitting the replicated data from the central publisher to a plurality of client workstations; and(h) simultaneously displaying the data on the plurality of client workstations during the cardiac catheterization procedure. 2. The method, as set forth in claim 1, wherein the catheterization procedure comprises a electrophysiology procedure. 3. The method, as set forth in claim 1, wherein the catheterization procedure comprises an ablation procedure. 4. The method, as set forth in claim 1, wherein the central publisher comprises a server. 5. The method, as set forth in claim 4, wherein the server is located locally with respect to the cardiac catheterization lab. 6. The method, as set forth in claim 1, wherein the central publisher is located within the cardiac catheterization lab. 7. The method, as set forth in claim 1, wherein each of the plurality of client workstations comprises a display monitor. 8. The method, as set forth in claim 1, wherein act (c) comprises the act of sensing electrical data from the heart. 9. The method, as set forth in claim 1, wherein act (c) comprises the act of sensing hemodynamic data from the heart. 10. The method, as set forth in claim 1, wherein act (h) comprises simultaneously displaying the data on a plurality of client systems during the procedure, wherein at least one of the plurality of client workstations is located remotely with respect to the cardiac catheterization lab. 11. The method, as set forth in claim 1, wherein act (h) occurs in real time. 12. A method of accessing a study record taken during a cardiac catheterization procedure, the procedure being conducted in a cardiac catheterization lab, comprising the acts of: (a) transmitting data from a catheter to a data collection device;(b) transmitting the data from the data collection device to a central publisher;(c) replicating the data;(d) transmitting the replicated data from the central publisher to a plurality of client workstations; and(e) simultaneously displaying the data on the plurality of client workstations during the cardiac catheterization procedure. 13. The method, as set forth in claim 12, wherein the catheterization procedure comprises a electrophysiology procedure. 14. The method, as set forth in claim 12, wherein the catheterization procedure comprises an ablatio... | GE Medical Systems, Inc. | A61B005/402
US07166297 | B2 | 2007-01-23 | Lipid-based drug delivery systems against parasitic infections | 2001-04-11 | The present invention is directed to a lipid-based delivery system for administration of an active drug substance selected from lysolipid derivatives which are particularly useful in the treatment or detection of parasitic infections, especially parasitic infections which cause an elevated PLA2 level in the infected mammal. Preferred parasitic infections are infections wherein the life cycle of the parasite involves the liver and/or spleen of the infected organism. | 1. A method for the systemic treatment of parasitic infections of a mammal in need of treatment thereof, said parasitic infection being characterized by increasing the PLA2 level in said mammal by administering to the mammal an efficient amount of a lipid-based drug delivery system comprising (A) a first active drug substance selected from lysolipid derivatives, wherein the active drug substance is present in the lipid-based system in the form of a prodrug, said prodrug being a lipid derivative having (a) an aliphatic group of a length of at least 7 carbon atoms and an organic radical having at least 7 carbon atoms, and (b) a hydrophilic moiety, said prodrug furthermore being a substrate for extracellular phospholipase A2 to the extent that the organic radical can be hydrolytically cleaved off, whereas the aliphatic group remains substantially unaffected, whereby the first active drug substance is liberated in the form of a lysolipid derivative which is not a substrate for lysophospholipase, and (B) a second drug substance which is a therapeutically and/or prophylactically active substance selected from the group consisting (i) anti-parasitic agents, (ii) antibiotics and anti-fungals, and (iii) anti-inflammatory agents, wherein the lipid-based drug delivery system is in the form of liposomes. | 1. A method for the systemic treatment of parasitic infections of a mammal in need of treatment thereof, said parasitic infection being characterized by increasing the PLA2 level in said mammal by administering to the mammal an efficient amount of a lipid-based drug delivery system comprising (A) a first active drug substance selected from lysolipid derivatives, wherein the active drug substance is present in the lipid-based system in the form of a prodrug, said prodrug being a lipid derivative having (a) an aliphatic group of a length of at least 7 carbon atoms and an organic radical having at least 7 carbon atoms, and (b) a hydrophilic moiety, said prodrug furthermore being a substrate for extracellular phospholipase A2 to the extent that the organic radical can be hydrolytically cleaved off, whereas the aliphatic group remains substantially unaffected, whereby the first active drug substance is liberated in the form of a lysolipid derivative which is not a substrate for lysophospholipase, and (B) a second drug substance which is a therapeutically and/or prophylactically active substance selected from the group consisting (i) anti-parasitic agents, (ii) antibiotics and anti-fungals, and (iii) anti-inflammatory agents, wherein the lipid-based drug delivery system is in the form of liposomes. 2. A method according to claim 1, wherein the increased level of PLA2 is localized to a specific tissue and/or organ of the mammal, said tissue and/or organ being infected by the parasite. 3. A method according to claim 1, wherein the parasitic infection involves the liver and/or the spleen of the mammal. 4. A method according to claim 1, wherein the organic radical which can be hydrolytically cleaved off, is an auxiliary drug substance or an efficiency modifier for the active drug substance. 5. A method according to claim 1, wherein the prodrug is a lipid derivative of the following formula:wherein X and Z independently are selected from O, CH2, NH, NMe, S, S(O), and S(O)2;Y is —OC(O)—, Y then being connected to R2 via either the oxygen or carbonyl carbon atom;R1 is an aliphatic group of the formula Y1Y2;R2 is an organic radical having at least 7 carbon atoms;where Y1 is —(CH2)n1—(CH═CH)n2—(CH2)n3—(CH═CH)n4—(CH2)n5—(CH═CH)n6—(CH2)n7—(CH═CH)n8r—(CH2)n9, and the sum of n1+2n2+n3+2n4+n5+2n6+n7+2n8+n9 is an integer of from 9 to 29; n1 is zero or an integer of from 1 to 29, n3 is zero or an integer of from 1 to 20, n5 is zero or an integer of from 1 to 17, n7 is zero or an integer of from 1 to 14, and n9 is zero or an integer of from 1 to 11; and each of n2, n4, n6 and n8 is independently zero or 1; and Y2 is CH3 or CO2H; where each Y1–Y2 independently may be substituted with halogen or C1-4-alkyl,R3 is selected from phosphatidic acid (PO2—OH), derivatives of phosphatidic acid and bioisosters to phosphatic acid and derivatives thereof. 6. A method according to claim 5, wherein R2 is an aliphatic group of a length of at least 7 carbon atoms. 7. A method acc... | LiPlasome Pharma A/S | A61K009/127 A61K051/00 A61B005/055 A61B008/00
US07272429 | B2 | 2007-09-18 | Methods and apparatus for facilitating a reduction in artifacts | 2002-11-27 | A method for facilitating a reduction in artifacts includes receiving data regarding a first energy spectrum of a scan of an object, and receiving data regarding a second energy spectrum of a scan of the object, wherein the second energy spectrum is different than the first energy spectrum. The method further includes reconstructing at least one original first energy image using the first energy spectrum data, reconstructing at least one original second energy image using the second energy spectrum data, transforming at least one original second energy image into at least one transformed first energy image, and combining at least one original first energy image with at least one transformed first energy image to generate a combined first energy image. | 1. A method for facilitating a reduction in artifacts, said method comprising: receiving data regarding a scan of an object, the data corresponding to a plurality of sequential projections comprising a first set of projections corresponding to a first energy spectrum and a second set of projections corresponding to a second energy spectrum, the first set of projections interleaved with the second set of projections;reconstructing at least one original first energy image using the data corresponding to the first set of projections;reconstructing at least one original second energy image using the data corresponding to the second set of projections;transforming at least one original first energy image into at least one transformed second energy image;transforming at least one original second energy image into at least one transformed first energy image;combining at least one original first energy image with at least one transformed first energy image to generate at least one combined first energy image; andcombining at least one original second energy image with at least one transformed second energy image to generate at least one combined second energy image. | 1. A method for facilitating a reduction in artifacts, said method comprising: receiving data regarding a scan of an object, the data corresponding to a plurality of sequential projections comprising a first set of projections corresponding to a first energy spectrum and a second set of projections corresponding to a second energy spectrum, the first set of projections interleaved with the second set of projections;reconstructing at least one original first energy image using the data corresponding to the first set of projections;reconstructing at least one original second energy image using the data corresponding to the second set of projections;transforming at least one original first energy image into at least one transformed second energy image;transforming at least one original second energy image into at least one transformed first energy image;combining at least one original first energy image with at least one transformed first energy image to generate at least one combined first energy image; andcombining at least one original second energy image with at least one transformed second energy image to generate at least one combined second energy image. 2. A method in accordance with claim 1 further comprising decomposing the combined first and second energy images to generate at least one quantitative material characterization image. 3. A method in accordance with claim 2 wherein said decomposing comprises at least one of decomposing by a Compton and photoelectric decomposition, decomposing by a basis material decomposition (BMD), and decomposing by a logarithm subtraction decomposition (LSD). 4. A method in accordance with claim 1 wherein said receiving data comprises receiving data comprising fan-beam projection data, said method further comprising rebinning the fan-beam projection data into equivalent parallel-beam projection data. 5. A method for facilitating a reduction in artifacts, said method comprising: receiving data from a scan of an object, the data corresponding to a plurality of sequential projections comprising a first set of projections corresponding to a first energy spectrum and a second set of projections corresponding to a second energy spectrum, the first set of projections interleaved with the second set of projections;reconstructing at least one original first energy image using the data corresponding to the first set of projections;reconstructing at least one original second energy image using the data corresponding to the second set of projections;transforming at least one original first energy image into at least one transformed second energy image;transforming at least one original second energy image into at least one transformed first energy image;reprojecting at least one transformed first energy image to generate reprojected first energy spectrum data;reprojecting at least one transformed second energy image to generate reprojected second energy spectrum data;combining the first energy spectrum data with t... | GE Medical Systems Global Technology Company, LLC | A61B005/05
US07197358 | B2 | 2007-03-27 | Identifying infants at risk for sudden infant death syndrome | 2002-06-18 | Identifying infants at risk for SIDS includes applying electrodes to an infant, receiving electrical signals from the electrodes, analyzing the received electrical signals to measure alternans of a heart of the infant, and identifying whether the infant is at risk for SIDS. A system for identifying infants at risk for SIDS includes an input unit configured to receive electrical signals from electrodes applied to an infant, a processor connected to the input unit and configured to process the received electrical signals to measure alternans of a heart of the infant, and a comparator configured to compare the measured alternans with alternans in a population of infants. | 1. A method for identifying infants at risk for SIDS, comprising: applying electrodes to an infant;receiving electrical signals from the electrodes;analyzing the received electrical signals to measure alternans of a heart of the infant; andidentifying whether the infant is at risk for SIDS using the measured alternans,wherein alternans comprise a beat-to-beat ABABAB change in an amplitude of an ECG waveform of the infant. | 1. A method for identifying infants at risk for SIDS, comprising: applying electrodes to an infant;receiving electrical signals from the electrodes;analyzing the received electrical signals to measure alternans of a heart of the infant; andidentifying whether the infant is at risk for SIDS using the measured alternans,wherein alternans comprise a beat-to-beat ABABAB change in an amplitude of an ECG waveform of the infant. 2. The method of claim 1 wherein: analyzing the received electrical signals comprises measuring T-wave alternans; andidentifying whether the infant is at risk for SIDS comprises comparing the measured T-wave alternans with T-wave alternans in a population of infants. 3. The method of claim 1 wherein identifying whether the infant is at risk for SIDS comprises analyzing and classifying the measured alternans of the heart. 4. The method of claim 1 further comprising elevating a heart rate of the infant. 5. The method of claim 4 wherein elevating the heart rate of the infant includes stressing the infant. 6. The method of claim 1 further comprising analyzing the received electrical signals to measure a QT interval of the heart of the infant. 7. The method of claim 6 further comprising comparing the measured QT interval with QT intervals in a population of infants. 8. The method of claim 1 wherein the received electrical signals comprise electrocardiogram signals. 9. The method of claim 1 wherein analyzing the received electrical signals includes sampling the received electrical signals at a frequency less than or equal to twice a frequency corresponding to alternans. 10. The method of claim 1 wherein analyzing the received electrical signals includes using a spectral approach to measure alternans. 11. The method of claim 1 wherein analyzing the received electrical signals includes using an analytical approach to measure alternans. 12. The method of claim 11 wherein using the analytical approach includes: processing the received electrical signals to create a processed signal having an asymmetric spectrum; andprocessing the processed signal to measure alternans in the received electrical signals. 13. The method of claim 12 wherein processing the received electrical signals to create a processed signal comprises creating the processed signal as an analytical signal. 14. The method of claim 13 wherein creating the processed signal as an analytical signal comprises generating a frequency domain representation of the received electrical signals, modifying the frequency domain representation to remove components corresponding to negative frequencies, and generating the analytical signal as a time domain representation of the modified frequency domain representation. 15. The method of claim 12 wherein processing the processed signal includes processing samples of the processed signal spaced by intervals greater than or equal to half of a period of alternans. 16. The method of claim 12 wherein processing the received electri... | Cambridge Heart, Inc. | A61B005/0444
US07139609 | B1 | 2006-11-21 | System and method for monitoring cardiac function via cardiac sounds using an implantable cardiac stimulation device | 2003-01-17 | Techniques are provided for performing internal measurement of heart sounds to estimate patient cardiac function in terms of stroke volume, cardiac output, or a maximum rate of change of aortic pressure with time (max dP/dt). Control parameters of the medical device are then automatically adjusted so as to optimize overall cardiac function or to provide for ventricular resynchronization therapy. In one example, heart sound signals are derived from acceleration signals received from an accelerometer. The heart sound signals are analyzed to identify S1 and S2 heart sounds as well as ejection period and isovolumic interval (ISOV). Proxies for max dP/dt, stroke volume and cardiac output are then derived from the S1 and S2 heart sounds, the ejection period and the ISOV. Alternative techniques, not requiring detection of ISOV, are employed for use if the patient has heart value regurgitation. | 1. In an implantable cardiac stimulation device for implant within a patient, a method comprising: detecting heart sound signals;determining whether systolic regurgitant murmurs are occurring during an ejection period; andprocessing the heart sound signals to derive one or more parameters representative of cardiac function of the patient, wherein the one or more parameters comprise stroke volume, cardiac output, and maximum rate of change of aortic pressure over time. | 1. In an implantable cardiac stimulation device for implant within a patient, a method comprising: detecting heart sound signals;determining whether systolic regurgitant murmurs are occurring during an ejection period; andprocessing the heart sound signals to derive one or more parameters representative of cardiac function of the patient, wherein the one or more parameters comprise stroke volume, cardiac output, and maximum rate of change of aortic pressure over time. 2. The method of claim 1 wherein determining one or more parameters representative of cardiac function of the patient comprises: determining two or more parameters representative of cardiac function of the patient. 3. The method of claim 1 wherein detecting heart sound signals comprises: detecting acceleration signals representative of the acceleration of tissues within the patient; andconverting the acceleration signals into heart sound signals. 4. The method of claim 3 wherein determining one or more parameters representative of cardiac function comprises: identifying S1 and S2 heart sounds within the heart sound signals. 5. The method of claim 3 further comprising detecting electrocardiographic signals representative of electrical characteristics of the heart. 6. The method of claim 5 wherein identifying S1 and S2 heart sounds comprises: detecting ventricular depolarization signals and ventricular repolarization signals within the electrocardiographic signals;defining S1 and S2 search windows within the heart sound signals subsequent to the ventricular depolarization and repolarization signals, respectively; anddetecting the S1 and S2 heart sounds within, respectively, the S1 and S2 search windows. 7. The method of claim 6 wherein detecting the S1 and S2 heart sounds comprises rectifying the heart sound signals;determining a maximum of the rectified heart sound signal within each of the S1 and S2 windows;setting S1 and S2 detection thresholds based on the maximum of the rectified heart sound signals; andidentifying the S1 and S2 heart sounds based on a comparison of the rectified heart sound signal within the detection thresholds within the respective S1 and S2 search windows. 8. The method of claim 6 further comprising detecting the duration of an ejection period based on the S1 and S2 heart sounds. 9. The method of claim 8 wherein detecting the duration of the ejection period is performed to determine the time interval from the end of the S1 sound to the beginning of the S2 sound. 10. The method claim 9 further comprising detecting heart rate based on the electrocardiographic signals. 11. The method of claim 10 further comprising detecting a maximum of the S1 heart sound (max_S1). 12. The method of claim 11 wherein determining whether systolic regurgitant murmurs are occurring during the ejection period comprises determining a peak to peak maximum of the heart sound signal during a murmur detection interval and comparing the maximum heart sound in the murmur detectio... | Pacesetter, Inc. | A61B005/02 A61N001/365
US07251513 | B2 | 2007-07-31 | Method of measuring biological information using light and apparatus of measuring biological information using light | 2003-01-23 | A compact apparatus of measuring biological information using light capable of measuring biological information with high reproducibility and accuracy is provided. The apparatus of measuring biological information using light comprises a light source part irradiating an organism, a light receiving part receiving light propagating from the light source part through the inside of the organism and outgoing from the surface of the organism, a forming part forming the surface of the organism into a predetermined shape by applying a pressure thereto, and a calculation part calculating information of the relation between the amount of received light and the biological information of the organism previously determined based on the amount of light received in the light receiving part. | 1. A method of measuring biological information utilizing an apparatus of measuring said biological information using light comprising: a light source part adapted to irradiate an organism; a light receiving part receiving light propagating from said light source part through an inside of said organism and outgoing from a surface of said organism; a forming part adapted to form said surface of said organism into a predetermined shape by applying a pressure thereto; a protrusion part being provided on a face of said forming part contacting said surface of said organism; a pressure detecting part detecting said pressure applied to said surface of said organism by said forming part, said pressure detecting part is connected to said forming part; and a calculation part calculating said biological information of said organism using information of a relation between an amount of said received light and said biological information of said organism previously determined, based on said amount of said light received in said light receiving part, at least one of said light source part and said light receiving part is provided on said protrusion part, said calculation part calculates said biological information of said organism based on said amount of said received light when it is detected that said pressure reaches said level equal to or greater than said predefined value, said method comprising: a first step of utilizing said forming part to form said surface of said organism into said predetermined shape by applying said pressure thereto;a second step of utilizing said light source part to irradiate said organism with said light;a third step of utilizing said light receiving part to receive said light propagating through said inside of said organism and outgoing from said surface of said organism; anda fourth step of utilizing said calculation part to calculate said biological information of said organism using said information of said relation between said amount of said received light and said biological information of said organism previously determined, based on said amount of said light received in said third step,wherein in said fourth step, said biological information of said organism is calculated using said information of said relation between said amount of said received light, which has been acquired after said pressure reaches a level equal to or greater than a predefined value, and said biological information of said organism previously determined, based on said amount of said received light acquired after said pressure reaches said level equal to or greater than said predefined value. | 1. A method of measuring biological information utilizing an apparatus of measuring said biological information using light comprising: a light source part adapted to irradiate an organism; a light receiving part receiving light propagating from said light source part through an inside of said organism and outgoing from a surface of said organism; a forming part adapted to form said surface of said organism into a predetermined shape by applying a pressure thereto; a protrusion part being provided on a face of said forming part contacting said surface of said organism; a pressure detecting part detecting said pressure applied to said surface of said organism by said forming part, said pressure detecting part is connected to said forming part; and a calculation part calculating said biological information of said organism using information of a relation between an amount of said received light and said biological information of said organism previously determined, based on said amount of said light received in said light receiving part, at least one of said light source part and said light receiving part is provided on said protrusion part, said calculation part calculates said biological information of said organism based on said amount of said received light when it is detected that said pressure reaches said level equal to or greater than said predefined value, said method comprising: a first step of utilizing said forming part to form said surface of said organism into said predetermined shape by applying said pressure thereto;a second step of utilizing said light source part to irradiate said organism with said light;a third step of utilizing said light receiving part to receive said light propagating through said inside of said organism and outgoing from said surface of said organism; anda fourth step of utilizing said calculation part to calculate said biological information of said organism using said information of said relation between said amount of said received light and said biological information of said organism previously determined, based on said amount of said light received in said third step,wherein in said fourth step, said biological information of said organism is calculated using said information of said relation between said amount of said received light, which has been acquired after said pressure reaches a level equal to or greater than a predefined value, and said biological information of said organism previously determined, based on said amount of said received light acquired after said pressure reaches said level equal to or greater than said predefined value. 2. The method of measuring biological information according to claim 1, wherein said predefined value of said pressure is about 7 kPa or greater. 3. A method of measuring biological information utilizing an apparatus of measuring said biological information using light comprising: a light source part adapted to irradiate an organism; a light receiving part r... | Matsushita Electric Industrial Co., Ltd. | A61B005/00
US07052470 | B2 | 2006-05-30 | Breathing detection/confirmation device | 2003-02-07 | A breath detection/confirmation device includes a sensor for determining if an infant or other person is breathing, and a means for generating a human perceptible signal in response thereto. In use, the sensor is placed adjacent the infants or person's mouth/nose, and a light or some other signal indicates to the user if the infant or person is breathing. | 1. A breath detection/confirmation device comprising: an elongate body having a proximal, handle portion, and a distal, end portion;the distal end portion comprising:a sensor detecting each breath of a person;a means for generating a human perceptible signal with each breath to confirm that the person is breathing; anda light for illuminating at least a part of a person's face allowing placement of the distal end of the device adjacent to the person's mouth or nose. | 1. A breath detection/confirmation device comprising: an elongate body having a proximal, handle portion, and a distal, end portion;the distal end portion comprising:a sensor detecting each breath of a person;a means for generating a human perceptible signal with each breath to confirm that the person is breathing; anda light for illuminating at least a part of a person's face allowing placement of the distal end of the device adjacent to the person's mouth or nose. 2. The breath detection/confirmation device according to claim 1, wherein the device further comprises a switch for selectively turning on the light. 3. A method for confirming the breathing of a person, the method comprising: selecting a body having a proximal, handle end configured for holding in a human hand and a distal end, the distal end having a sensor disposed therein for confirming breathing of a person; andholding the device in the hand and positioning the device so that the distal end is disposed adjacent to the mouth or nose of a sleeping or unconscious person; andgenerating a human perceptible signal when the person breathes that the person is breathing; andwherein the method comprises illuminating at least a portion of the distal end of the device to facilitate placement of the distal end adjacent to the mouth or nose of the person. 4. A method according to claim 3, wherein the method comprises holding the device in the user's hand while confirming breathing. 5. A method according to claim 3, wherein the method comprises sensing breathing by use of an infrared sensor. 6. A method according to claim 3, wherein the method comprises sensing breathing by detection of condensation. 7. A method according to claim 3, wherein the method comprises sensing breathing by detection of humidity. 8. A method according to claim 3, wherein the method comprises sensing breathing by detection of heat. 9. A method according to claim 3, wherein the mehtod comprises sensing breathing by detection of fluid flow. 10. A method for confirming the breathing of a person, the method comprising: selecting a body having a proximal, handle end configured for holding in a human hand and a distal end, the distal end having a sensor disposed therein for confirming breathing of a person; andholding the device in the hand and positioning the device so that the distal end is disposed adjacent to the mouth or nose of a sleeping or unconscious person; andgenerating a human perceptible signal when the person breathes that the person is breathing; andwherein the human perceptible signal comprises illuminating a light in response to each breath. 11. A method for confirming the breathing of a person, the method comprising: selecting a body having a proximal, handle end configured for holding in a human hand and a distal end, the distal end having a sensor disposed therein for confirming breathing of a person; andholding the device in the hand and positioning the device so that the distal end is disposed ... | - | A61B005/08
US07184820 | B2 | 2007-02-27 | Tissue monitoring system for intravascular infusion | 2002-08-23 | An infusion system has a capability to monitor infusion complications such as extravasation, tissue necrosis, infiltration, phlebitis, and blood clots. The infusion system has at least partially transparent flexible film barrier dressing in a flexible membrane that incorporates a plurality of sensors capable of detecting tissue condition and a control unit capable of coupling to the film barrier dressing that monitors signals from the sensors. A device is capable of executing non-invasive physiological measurements to characterize physiologic information from cross-sectional surface and subcutaneous tissue to detect the presence or absence of tissue conditions such as infiltration or extravasation during intravascular infusion. In some examples, the device utilizes depth-selective methods to sense, detect, quantify, monitor, and generate an alert notification of tissue parameters. | 1. A tissue monitoring apparatus comprising: a film barrier dressing comprising a flexible membrane integrating one or more sensor elements adapted to interrogate tissue at multiple electrical signal frequencies and an adhesive layer adapted to couple the film barrier dressing to a patient's skin;a control unit adapted to couple to the one or more sensor elements in the film barrier dressing, the control unit adapted to control the sensor elements to interrogate the tissue at selected multiple frequencies, sense one or more electrical signal parameters indicative of tissue condition in three-dimensional space, and execute a complex number three-dimensional pattern recognition operation on the one or more sensed parameters to determine a subcutaneous mapping indicative of a tissue condition;a bio-impedance sensor including a plurality of sensor elements configured to acquire impedance signals in three-dimensional space; andan electrical impedance tomography processor executable in the control unit and adapted to execute an object recognition operation on the impedance signals. | 1. A tissue monitoring apparatus comprising: a film barrier dressing comprising a flexible membrane integrating one or more sensor elements adapted to interrogate tissue at multiple electrical signal frequencies and an adhesive layer adapted to couple the film barrier dressing to a patient's skin;a control unit adapted to couple to the one or more sensor elements in the film barrier dressing, the control unit adapted to control the sensor elements to interrogate the tissue at selected multiple frequencies, sense one or more electrical signal parameters indicative of tissue condition in three-dimensional space, and execute a complex number three-dimensional pattern recognition operation on the one or more sensed parameters to determine a subcutaneous mapping indicative of a tissue condition;a bio-impedance sensor including a plurality of sensor elements configured to acquire impedance signals in three-dimensional space; andan electrical impedance tomography processor executable in the control unit and adapted to execute an object recognition operation on the impedance signals. 2. A tissue monitoring apparatus according to claim 1 wherein: the flexible membrane is at least partially transparent so that a patient's tissue underlying the film barrier dressing can be visually inspected. 3. A tissue monitoring apparatus according to claim 1 wherein: the film barrier dressing is a breathable barrier film that protects against infection and functions as a structural member that is capable of securing an intravenous catheter. 4. A tissue monitoring apparatus according to claim 1 further comprising: a bio-impedance sensor, the control unit adapted to couple to the bio-impedance sensor and execute a pattern recognition operation in one or more dimensions on the impedance signals to generate one or more sensor pathways using depth-selective sensing of tissue bio-impedance in a selected frequency range, generate cross-sectional surface measurements and subcutaneous measurements at one or more selected tissue depths by controlling the field extension of the sensor pathway, and determine the tissue condition. 5. A tissue monitoring apparatus according to claim 1 further comprising: a bio-impedance sensor and an optical sensor, the control unit adapted to couple to the bio-impedance sensor and to the optical sensor and adapted to execute a pattern recognition operation in one or more dimensions on the impedance signals and the optical signals to determine the tissue condition. 6. A tissue monitoring apparatus according to claim 1 further comprising: a bio-impedance sensor including a plurality of sensor elements configured to acquire impedance signals in three-dimensional space; andan electrical impedance tomography processor executable in the control unit and adapted to execute a three-dimensional tomography operation on the impedance signals and generate one or more sensor pathways using depth-selective sensing of tissue bio-impedance in a selected freque... | Subqiview, Inc. | A61B005/05
US07209781 | B2 | 2007-04-24 | Method for localizing at least one focal lesion in a biological tissue section | 2003-01-03 | In a method for localizing at least one focal lesion in a biological tissue section, the lesion exhibiting an electrical property different from the tissue section and the electrical property in the tissue section being essentially constant, a sequence of electrical excitation signals with different frequency is applied to the tissue section, electrical response signals at a number of measurement locations at a surface of the tissue section are measured that occur thereat due to the excitation signals, frequency-independent signal parts in the response signals are determined and further-processed input values of a localization step. Modeling of the tissue section and a determination of a set of lead fields are based on the input values. The lead fields transformed. The frequency-independent signal parts are compared to the transformed lead fields in the localization step, and the location of the transformed lead fields that best reproduce the frequency-independent signal parts is output as location of the lesion. | 1. A method for localizing at least one lesion in a biological tissue section, said lesion exhibiting an electrical property different from said tissue section, and said electrical property in said tissue section being substantially constant, comprising the steps of: applying a sequence of electrical excitation signals with respectively different frequencies to said tissue section;measuring electrical response signals, arising due to said excitation signals, at a plurality of measurement locations at a surface of said tissue section;determining frequency-independent signal parts in said response signals and further-processing said frequency-independent signal parts to produce input values for a localization procedure; andin said localization procedure, modeling said tissue section dependent on said input values to produce a tissue section model and from said tissue section model determining a set of lead fields, and transforming said lead fields to obtain transformed lead fields, and comparing said frequency-independent signal parts to said transformed lead fields to identify a location of said transformed lead fields that best reproduces said frequency-independent signal parts, and providing said location as an output as the location of said lesion. | 1. A method for localizing at least one lesion in a biological tissue section, said lesion exhibiting an electrical property different from said tissue section, and said electrical property in said tissue section being substantially constant, comprising the steps of: applying a sequence of electrical excitation signals with respectively different frequencies to said tissue section;measuring electrical response signals, arising due to said excitation signals, at a plurality of measurement locations at a surface of said tissue section;determining frequency-independent signal parts in said response signals and further-processing said frequency-independent signal parts to produce input values for a localization procedure; andin said localization procedure, modeling said tissue section dependent on said input values to produce a tissue section model and from said tissue section model determining a set of lead fields, and transforming said lead fields to obtain transformed lead fields, and comparing said frequency-independent signal parts to said transformed lead fields to identify a location of said transformed lead fields that best reproduces said frequency-independent signal parts, and providing said location as an output as the location of said lesion. 2. A method as claimed in claim 1 comprising transforming said lead fields by normalizing said lead fields in a first transformation, thereby obtaining normalized lead fields. 3. A method as claimed in claim 2 comprising transforming said lead fields by transforming said normalized lead fields, in a second transformation, into orthogonal lead fields. 4. A method as claimed in claim 3 comprising determining said orthogonal lead fields from said transformed lead field by a singular value resolution. 5. A method as claimed in claim 1 comprising modeling said tissue section as an electrical volume conductor. 6. A method as claimed in claim 1 comprising determining electrical admittance data from said response signals dependent on a location at said surface of said tissue section at which at least one maximum of said electrical admittance data occurs, and identifying at least one depth position in said tissue section of said lesion beneath said position on said surface of said maximum, dependent on at least one position of a peak selected from the group consisting of maximum peaks and minimum peaks. 7. A method as claimed in claim 1 comprising determining a frequency-dependency of a bioelectrical signal source from said response signals and classifying said lesion dependent on said frequency-dependency. 8. A method as claimed in claim 7 comprising representing said bioelectrical signal source by a multi-pole development. 9. A method as claimed in claim 1 comprising generating a matrix containing said response signals, and determining said frequency-independent signal parts in said response signals by singular value resolution of said matrix. | Siemens Aktiengesellschaft | A61B005/05
US07004914 | B2 | 2006-02-28 | Crimp and cut tool for sealing and unsealing guide wires and tubular instruments | 2002-10-03 | An apparatus for modulating the pressure of a fluid such as a gas within the expandable portion of a guide wire catheter. In one embodiment, the pressure and volume of fluid in the tubular guide wire is maintained by utilizing an apparatus to crimp the end of the tubular guide wire, effectively sealing the inside bore of the tubular guide wire. Here, a deflation tool might be employed to sever the crimp from the rest of the tubular guide wire, unsealing the bore of the tubular guide wire. Using this apparatus, the tubular guide wire can be re-sealed and re-opened as necessary. | 1. An apparatus for modulating the pressure of a fluid within an expandable member on a guide wire or tubular instrument, the apparatus comprising an inflation tool which comprises: means for controllably gripping an open, proximal end of a guide wire or tubular instrument;means for introducing a fluid into said expandable member on said guide wire or tubular instrument through said open end, thereby inflating said expandable member on said guide wire or tubular instrument to a desired extent; andmeans for effecting the sealing of said open proximal end of said guide wire or tubular instrument to maintain said fluid in said guide wire or tubular instrument, while maintaining said expandable member inflated to a desired extent, the apparatus further comprising a means for re-opening said proximal end of said guide wire or tubular instrument, said re-opening means comprising a cutting tool arranged to sever said guide wire or tubular instrument distally of said sealed proximal end, said cutting tool further being arranged to entrap said severed proximal end. | 1. An apparatus for modulating the pressure of a fluid within an expandable member on a guide wire or tubular instrument, the apparatus comprising an inflation tool which comprises: means for controllably gripping an open, proximal end of a guide wire or tubular instrument;means for introducing a fluid into said expandable member on said guide wire or tubular instrument through said open end, thereby inflating said expandable member on said guide wire or tubular instrument to a desired extent; andmeans for effecting the sealing of said open proximal end of said guide wire or tubular instrument to maintain said fluid in said guide wire or tubular instrument, while maintaining said expandable member inflated to a desired extent, the apparatus further comprising a means for re-opening said proximal end of said guide wire or tubular instrument, said re-opening means comprising a cutting tool arranged to sever said guide wire or tubular instrument distally of said sealed proximal end, said cutting tool further being arranged to entrap said severed proximal end. 2. The apparatus of claim 1, wherein said sealing means comprises a crimping tool. 3. The apparatus of claim 2, wherein said crimping tool comprises a plurality of dies arranged to crimp said guide wire or tubular instrument. 4. In combination with the apparatus of claim 3, a guide wire or tubular instrument and a deformable material to facilitate sealing, wherein said guide-wire or tubular instrument comprises an inside surface that defines a bore, with said deformable material coating said inside surface. 5. In combination with the apparatus of claim 3, a guide wire or tubular instrument and a deformable material to facilitate sealing, wherein said guide-wire or tubular instrument comprises an inside surface that defines a bore, with said deformable material contacting at least a portion of said inside surface. 6. A valve mechanism for modulating pressure in a balloon or expandable member on or about a catheter or guide wire, the valve mechanism comprising; a) an inflation tool, the inflation tool comprising a gripping means, the gripping means comprising a deformable member and a deforming mechanism, the deforming mechanism acting upon the deformable member to cause gripping of the catheter or guide wire;an inflating means, the inflating means comprising a charge of fluid, and a means for modulating the pressure of the charge of fluid, and a pierceable barrier to prevent premature leakage of the charge of fluid; anda crimping means; the crimping means comprising a tool suitable for sealing and maintaining pressure in the catheter or guide wire, by applying radial pressure by a plurality of dies upon said catheter or guide wire, thereby creating a crimp and sealing said catheter or guide wire; andb) a deflation tool, the deflation tool comprising cutting means arranged to sever the crimp of the catheter or guide wire from the remainder of the catheter or guide wire, thereby unsealing sai... | Kensey Nash Corporation | A61B005/00 A61M025/00
US07172624 | B2 | 2007-02-06 | Medical device with magnetic resonance visibility enhancing structure | 2003-02-06 | A medical device that inhibits distortion of medical resonance images taken of the device. In particular, various structures are utilized to allow visibility proximate, and inside of, a tubular member, such as a stent. In one embodiment, the stent does not contain electrically conductive loops. In an alternative embodiment, rings in the stent are arranged such that current in one ring portion is opposed by current in another ring portion. | 1. A stent comprising; a generally tubular structure having a plurality of non conductive coil-shaped structural members;a radiopaque layer covering the coil shaped structural members, wherein each coil-shaped structural member defines a gap in the radiopaque layer; anda coating of insulating material applied to the radiopaque layer, wherein the coating of the insulating material is disposed in the gaps of the radiopaque layer to prevent the formation of an electrical loop across the gaps. | 1. A stent comprising; a generally tubular structure having a plurality of non conductive coil-shaped structural members;a radiopaque layer covering the coil shaped structural members, wherein each coil-shaped structural member defines a gap in the radiopaque layer; anda coating of insulating material applied to the radiopaque layer, wherein the coating of the insulating material is disposed in the gaps of the radiopaque layer to prevent the formation of an electrical loop across the gaps. 2. The stent of claim 1, wherein the insulating material is a polymeric material. 3. The stent of claim 1 wherein the structural members are selected from the group comprising ceramics and polymers. 4. A stent comprising: a generally tubular structure having a plurality of non conductive coil shaped structural members;a radiopaque layer applied to the coil shaped structural members, wherein each coil-shaped structural member defines a gap in the radiopaque layer; anda coating of insulating material applied to the radiopaque layer to prevent the formation of an electrical loop across the gaps,said radiopaque layer and said insulating material being adapted to allow magnetic changes in a region immediately proximate the generally tubular structure, induced by a magnetic resonance imaging process, to be substantially unobstructed by the generally tubular structure. 5. The stent of claim 4 wherein the generally tubular structure comprises a plurality of generally coaxially arranged ring connected to one another such that current induced in adjacent ring flows in opposite directions in the adjacent ring. 6. The stent of claim 5 wherein the ring structures comprise: a first ring portion including a first end and a second end spaced apart from the first end;a second ring portion including a first end and a second end spaced apart from the first end of the second ring portion;a first connector connecting the first end of the first ring portion and the first end of the second ring portion; anda second connector connecting the second end of the first ring portion and the second end of the second ring portion, wherein a current flows in a first direction in the first ring portion and in a second direction, opposite the first direction, in the second ring portion. 7. The stent of claim 6 wherein the first ring portion and the second ring portion are positioned substantially parallel to each other. 8. The stent of claim 6 wherein the first and second ring portions are made of a substantially non-ferromagnetic material. 9. The stent of claim 6 wherein the first ring portion and the second ring portion define a gap in an axial direction along an outer periphery of the generally tubular structure. 10. The stent of claim 5, wherein the generally tubular structure comprises: a plurality of rings, each ring compromising: a first ring portion disposed about a central axis and including a first end and a second end spaced apart from the first end;a second ring portion disp... | Boston Scientific Scimed, Inc. | A61F002/06 A61B005/05
US07094206 | B2 | 2006-08-22 | System for measuring respiratory function | 2002-09-09 | The present invention relates to a system for measuring respiratory function of living organisms. More particularly, signals indicative of the change in lung volume, defined as active and passive work, required to breathe and the airflow through the respiratory system of the living organism are obtained and processed as waveforms to provide a signal indicative of the respiratory restriction. The methods of the present invention measure clinical forms of airway obstruction, airway reactivity and lung volume and may be used to continuously or intermittently monitor patients with compromised respiratory function. In a preferred embodiment, a head-out, breath in respiratory plethysmograph system provides the signals indicative of change in lung volume as related to pressure changes in a chamber. Further, flowmetric variables are generated that provide a characterization of airway obstructions. | 1. The method for diagnosing sleep apnea conditions of a living organism comprising the steps of: obtaining a first input signal from a first sensor indicative of a change in lung volume during breathing by said living organism;obtaining a second input signal from a second sensor indicative of airflow through the respiratory system of said living organism during the change in lung volume; andprocessing said first input signal and said second input signal by analyzing phase and magnitude differences between the first input signal and the second input signal, and to calculate a third signal indicative of a respiration restriction condition that adversely effects sleep of said living organism. | 1. The method for diagnosing sleep apnea conditions of a living organism comprising the steps of: obtaining a first input signal from a first sensor indicative of a change in lung volume during breathing by said living organism;obtaining a second input signal from a second sensor indicative of airflow through the respiratory system of said living organism during the change in lung volume; andprocessing said first input signal and said second input signal by analyzing phase and magnitude differences between the first input signal and the second input signal, and to calculate a third signal indicative of a respiration restriction condition that adversely effects sleep of said living organism. 2. The method according to claim 1, wherein the processing step comprises the step of comparing said first input signal to said second input signal. 3. The method according to claim 2, further comprising of obtaining the first input signal through use of at least one of respiratory induction plethysmography, a head out, breath-in plethysmograph, an impedance plethysmography, an optoelectric plethysmograph and a fiber optic respiration plethysmograph. 4. The method according to claim 1, further comprising determining the location of a respiratory restriction by analyzing whether phase and magnitude differences occur in one of an inspiratory portion of said first and second signals and an expiratory portion of said first and second signals. 5. The method according to claim 1, further comprising of obtaining the second input signal through use of one of a nasal cannula and a pneumotachographic measurement device. 6. The method according to claim 1, further comprising of obtaining the second input signal through use of an ultrasonic device. 7. The method according to claim 1, further comprising of obtaining the second input signal through use of one of a thermistor, and an acoustic sensor. 8. The method according to claim 1, further comprising of obtaining the second input signal through use of a breath-sound intensity device. 9. The method according to claim 1, further comprising the step of calibrating first input signal. 10. The method according to claim 1, wherein the first input signal is processed to provide a measure of functional residual capacity and specific airway resistance. | The Trustees of Tufts College | A61B005/00
US07004910 | B2 | 2006-02-28 | System and method for monitoring body temperature | 2002-12-12 | The system comprises one or more sensor devices, where each sensor device is capable of measuring a temperature at a known location of a body of a human or other animal. Each sensor transmits a temperature measurement value and sensor identification to a monitor device. The monitor device receives the temperature measurement value and sensor identification, and computes an adjusted temperature value based upon the position of the measuring sensor, body age, and time of day. The adjusted temperature value is tested for alarm conditions and displayed for use by a caretaker. | 1. A method for monitoring body temperature, comprising the steps of: measuring one or more temperature values by sensors at known locations on one or more bodies;transmitting the measured temperature values and corresponding sensor identifications from the sensors to a monitoring device;receiving the measured temperature values and corresponding sensor identifications by the monitoring device;comparing the received sensor identifications with sensor identifications in a sensor registration table for determining a unique sensor manufacturer registration number, a sensor position, a body age, and alarm limits associated with each sensor;computing adjusted temperature values from the measured temperature values based on the known body locations of the corresponding sensors;storing the received measured temperature values, the adjusted temperature values, the sensor identification numbers, a sensor status, a time and date indication, an end indicator, and a new entry indicator as a sensor entry in a multiple sensor data table; anddisplaying the adjusted temperature values for use by a caretaker. | 1. A method for monitoring body temperature, comprising the steps of: measuring one or more temperature values by sensors at known locations on one or more bodies;transmitting the measured temperature values and corresponding sensor identifications from the sensors to a monitoring device;receiving the measured temperature values and corresponding sensor identifications by the monitoring device;comparing the received sensor identifications with sensor identifications in a sensor registration table for determining a unique sensor manufacturer registration number, a sensor position, a body age, and alarm limits associated with each sensor;computing adjusted temperature values from the measured temperature values based on the known body locations of the corresponding sensors;storing the received measured temperature values, the adjusted temperature values, the sensor identification numbers, a sensor status, a time and date indication, an end indicator, and a new entry indicator as a sensor entry in a multiple sensor data table; anddisplaying the adjusted temperature values for use by a caretaker. 2. The method of claim 1, wherein: the step of measuring further comprises measuring at predetermined time intervals; andthe step of transmitting further comprises transmitting at the predetermined time intervals. 3. The method of claim 1, wherein: the step of measuring further comprises measuring in response to requests received from the monitor; andthe step of transmitting further comprises transmitting in response to the requests received from the monitor. 4. The method of claim 1, further comprising the step of calculating estimated remaining battery life of each sensor and transmitting the estimated battery life from each sensor to the monitoring device. 5. The method of claim 1, further comprising the step of detecting sensor errors/status, setting sensor error/status flags, and transmitting the sensor error/status flags from each sensor to the monitoring device. 6. The method of claim 1, further comprising the step of activating a sensor status indicator. 7. The method of claim 1, further comprising the step of switching the sensor to a low power mode when not measuring and transmitting. 8. The method of claim 1, wherein: the step of transmitting the measured temperature values comprises transmitting the measured temperature values via a wireless communication link; andthe step of receiving the measured temperature values comprises receiving the measured temperature values via a wireless communication link. 9. The method of claim 1, wherein: the step of transmitting the measured temperature values comprises transmitting the measured temperature values via a wireless area network; andthe step of receiving the measured temperature values comprises receiving the measured temperature values via a wireless area network. 10. The method of claim 1, wherein the step of displaying the adjusted temperature values comprises displaying the adjusted tempe... | Alert Care, Inc | A61B005/00
US07123953 | B2 | 2006-10-17 | Method and system for evaluating arrhythmia risk with QT-RR interval data sets | 2002-12-03 | A method of assessing the cardiac arrhythmia risk in a subject to provide a measure of cardiac or cardiovascular health in that subject is described herein. In one embodiment, the method comprises the steps of: (a) collecting at least one QT and RR interval data set from the subject; (b) separating fluctuations from slow trends in said at least one QT and RR interval data set; (c) comparing said QT and RR fluctuations to one another and (d) generating from the comparison of step (c) partial measures of risk of cardiac arrhythmia in said subject. A greater difference between QT and RR fluctuations indicates greater risk of cardiac arrhythmia in said subject. The data sets are collected in such a manner that they reflect almost exclusively the conduction in the heart muscle and minimize the effect on the data sets of rapid transients due to autonomic nervous system and hormonal control. | 1. A method of screening for risk of cardiac arrhythmia in a subject, the method comprising the steps of: (a) collecting at least one QT interval data set and at least one RR interval data set from the subject during (i) a stage of gradually increasing heart rate, (ii) a stage of gradually decreasing heart rate, or (iii) both a stage of gradually increasing heart rate and gradually decreasing heart rate;(b) separating fluctuations from slow trends in said at least one QT interval data set and in said at least one RR interval data set to provide QT fluctuations and RR fluctuations;(c) comparing said QT fluctuations and said RR fluctuations to one another to determine the difference therebetween; and(d) generating from the comparison of step (c) a measure of risk of cardiac arrhythmia in said subject, wherein a greater difference between QT fluctuations and RR fluctuations indicates greater risk of cardiac arrhythmia in said subject. | 1. A method of screening for risk of cardiac arrhythmia in a subject, the method comprising the steps of: (a) collecting at least one QT interval data set and at least one RR interval data set from the subject during (i) a stage of gradually increasing heart rate, (ii) a stage of gradually decreasing heart rate, or (iii) both a stage of gradually increasing heart rate and gradually decreasing heart rate;(b) separating fluctuations from slow trends in said at least one QT interval data set and in said at least one RR interval data set to provide QT fluctuations and RR fluctuations;(c) comparing said QT fluctuations and said RR fluctuations to one another to determine the difference therebetween; and(d) generating from the comparison of step (c) a measure of risk of cardiac arrhythmia in said subject, wherein a greater difference between QT fluctuations and RR fluctuations indicates greater risk of cardiac arrhythmia in said subject. 2. The method according to claim 1, wherein said QT and RR interval data sets are collected under quasi-stationary conditions. 3. The method according to claim 1, wherein said stage of gradually increasing heart rate and said stage of gradually decreasing heart rate are each at least 3 minutes in duration. 4. The method according to claim 1, wherein said stage of gradually increasing heart rate and said stage of gradually decreasing heart rate are together carried out for a total time of from 6 minutes to 40 minutes. 5. The method according to claim 1, wherein: both said stage of gradually increasing heart rate and said stage of gradually decreasing heart rate are carried out between a peak rate and a minimum rate; andsaid peak rates of both said stage of gradually increasing heart rate and said stage of gradually decreasing heart rate are the same. 6. The method according to claim 1, wherein said stage of gradually decreasing heart rate is carried out at at least three different heart-rate stimulation levels. 7. The method according to claim 6, wherein said stage of gradually increasing heart rate is carried out at at least three different heart-rate stimulation levels. 8. The method according to claim 1, wherein said stage of gradually increasing heart rate and said stage of gradually decreasing heart rate are carried out sequentially in time. 9. The method according to claim 1, wherein said stage of gradually increasing heart rate and said stage of gradually decreasing heart rate are carried out separately in time. 10. The method according to claim 1, wherein said heart rate during said stage of gradually increasing heart rate does not exceed more than 120 beats per minute. 11. The method according to claim 1, wherein said heart rate during said stage of gradually increasing heart rate exceeds 120 beats per minute. 12. The method according to claim 1, further comprising the step of: (e) comparing said measure of risk of cardiac arrhythmia risk to at least one reference value; and then(f) generating from s... | Mediwave Star Technology Inc. | A61B005/04
US07207947 | B2 | 2007-04-24 | System and method for detecting circadian states using an implantable medical device | 2003-01-10 | Techniques are provided for detecting the circadian state of a patient using an implantable medical device based on selected blood carbon dioxide (CO2) parameters. In one example, the implantable device tracks changes in end tidal CO2 (etCO2) levels and changes in maximum variations of pCO2 levels per breathing cycle (ΔcycleCO2) over the course of the day and determines the circadian state based thereon. It has been found that average etCO2 levels are generally highest and average ΔcycleCO2 levels are generally lowest while a patient is asleep and opposite while a patient is awake. Hence, by tracking changes in average etCO2 and ΔcycleCO2 levels over the course of the day, circadian states can be detected. Minute ventilation and activity levels are used to assist in the determination of the circadian state. Additional techniques are directed to detecting the stage of sleep. | 1. In an implantable medical device for implant within a patient, a method for detecting a circadian state of the patient, the method comprising: detecting changes in one or more blood carbon dioxide (CO2) parameters;determining the circadian state of the patient based upon the changes in the one or more blood CO2 parameters;detecting minute ventilation levels of the patient and wherein the step of determining the circadian state of the patient additionally takes into account the detected minute ventilation levels;determining a ratio of minute ventilation to blood CO2 levels of the patient and wherein the step of determining the circadian state of the patient additionally takes into account the determined ratio; anddetecting activity levels of the patient;wherein detecting changes in selected blood CO2 parameters is performed to detect changes in etCO2 levels and changes in ΔcycleCO2; andwherein the step of determining the circadian state of the patient takes into account a combination of ΔcycleCO2 levels, etCO2 levels, activity levels and minute ventilation levels. | 1. In an implantable medical device for implant within a patient, a method for detecting a circadian state of the patient, the method comprising: detecting changes in one or more blood carbon dioxide (CO2) parameters;determining the circadian state of the patient based upon the changes in the one or more blood CO2 parameters;detecting minute ventilation levels of the patient and wherein the step of determining the circadian state of the patient additionally takes into account the detected minute ventilation levels;determining a ratio of minute ventilation to blood CO2 levels of the patient and wherein the step of determining the circadian state of the patient additionally takes into account the determined ratio; anddetecting activity levels of the patient;wherein detecting changes in selected blood CO2 parameters is performed to detect changes in etCO2 levels and changes in ΔcycleCO2; andwherein the step of determining the circadian state of the patient takes into account a combination of ΔcycleCO2 levels, etCO2 levels, activity levels and minute ventilation levels. 2. The method of claim 1 wherein detecting changes in selected blood CO2 parameters comprises detecting changes in maximum variation of pCO2 level per breathing cycle (ΔcycleCO2). 3. The method of claim 1 wherein detecting changes in selected blood CO2 parameters comprises detecting changes in maximum variation of blood pH level per breathing cycle (ΔcyclepH). 4. The method of claim 1 wherein determining the circadian state of the patient based upon the changes in the selected blood CO2 parameters comprises identifying whether the patient is asleep or awake. 5. The method of claim 1 further comprising determining a sleep stage of the patient based on the ratio of minute ventilation to blood CO2 levels of the patient. 6. The method of claim 1 further comprising controlling device functions based on the detected circadian state of the patient. 7. In an implantable medical device for implant within a patient, a method for detecting a circadian state of the patient, the method comprising: detecting changes in one or more blood carbon dioxide (CO2) parameters to detect changes in etCO2 levels and changes in ΔcycleCO2 levels;determining minute ventilation levels of the patient;detecting activity levels of the patient; anddetermining the circadian state of the patient comprising:generating histograms representative of a range of values of the etCO2 levels, ΔcycleCO2 levels, activity levels, and minute ventilation levels detected over a period of time; andidentifying the circadian state of the patient based upon the shape of the histograms. 8. In an implantable medical device for implant within a patient, a control system comprising: a blood CO2 sensor detecting changes in etCO2 levels and changes in ΔcycleCO2;an activity sensor detecting activity levels of the patient;a minute ventilation sensor detecting minute ventilation levels of the patient; anda controller operative to control d... | Pacesetter, Inc. | A61B005/08
US07235056 | B2 | 2007-06-26 | Body fluid sampling device and methods of use | 2002-04-23 | A device for sampling body fluid includes a housing having a sleeve at a forward end thereof which is displaceable in response to being pressed against a user's skin to trigger the firing of a lancet. After the lancet is removed from the incision, the sleeve is repeatedly pressed against the skin to depress a ring of body tissue in surrounding relationship to the incision to express body fluid outwardly through the incision. A pusher member is then actuated to push a capillary tube through a front end of the housing for drawing-in body fluid. The lancet is a disposable lancet which includes a body supporting a skin-lancing member and the capillary tube. The disposable lancet passes through an upper end of a lancet carrier when being installed or removed. The device cannot be armed until the disposable lancet is installed in the housing, because the capillary tube functions to push a safety device to a non-safety position. | 1. A method of collecting a bodily fluid sample, comprising: forming an incision in skin with an integrated bodily fluid sampling device that has a sampling tube; projecting an end of the sampling tube from the integrated bodily fluid sampling device after said forming; and a collecting the bodily fluid sample from the incision with the sampling tube. | 1. A method of collecting a bodily fluid sample, comprising: forming an incision in skin with an integrated bodily fluid sampling device that has a sampling tube; projecting an end of the sampling tube from the integrated bodily fluid sampling device after said forming; and a collecting the bodily fluid sample from the incision with the sampling tube. 2. The method of claim 1, further comprising: providing the integrated bodily fluid sampling device including an incising device and the sampling tube, the sampling tube having an inlet opening, the sampling tube being separate from the incising device, the sampling tube having a retracted position and a sampling position relative to the integrated bodily fluid sampling device; placing the sampling device against the skin in a use position with the sampling tube in the retracted position having the inlet opening in a first location displaced from the skin; wherein said projecting the end of the sampling tube includes moving the sampling tube from the retracted position to the sampling position at which the inlet opening is positioned adjacent the incision in the skin; and maintaining the sampling device in the use position during said forming the incision and said projecting the end of the sampling tube. 3. The method of claim 1, wherein: the device includes a lancet member; and said forming the incision includes cutting the skin with the lancet member. 4. The method of claim 3, wherein: the sampling tube is slidably coupled to the lancet member; and said projecting includes sliding the end of the tube outwardly from the lancet member. 5. The method of claim 3, further comprising retracting the lancet member away from the incision after said cutting. 6. The method of claim 1, further comprising drawing out the sample from the sampling tube onto a test strip. 7. The method of claim 1, wherein: the device includes a pusher member positioned proximal the tube; and said projecting includes pushing the tube with the pusher member. 8. The method of claim 1, further comprising pressing the device against the skin to draw up the bodily fluid sample from the incision. 9. The method of claim 8, wherein said pressing includes repeatedly reciprocating the device against the skin. 10. A lancet member, comprising: a body; an incision-forming member extending from said body, said incision-forming member being adapted to form an incision in skin; and a tube coupled to said body, said tube having an end adapted to move outwardly from said body, said tube having an extended position at which said end of said tube extends past said incision-forming member to collect a bodily fluid sample from the incision. 11. The lancet member of claim 10, wherein said incision-forming member includes a needle. 12. The lancet member of claim 10, wherein: said incision-forming member includes a needle; said body includes at least one lower boss to align said body in a lancet carrier and at least a pair of upper bosses... | Amira Medical | A61B005/00
US06947565 | B2 | 2005-09-20 | Physiological condition monitors utilizing very low frequency acoustic signals | 2002-07-01 | Physiological condition monitors utilizing very low frequency acoustic signals and signals indicative of body orientation are disclosed. The physiological condition monitors comprise a sensor that is capable of detecting low frequency acoustic signals in the frequency range of one tenth Hertz to thirty Hertz. The sensor comprises a chamber having portions that form a cavity and a low frequency microphone placed within the cavity. An alternate embodiment of the invention comprises a chamber having portions that form a resonant cavity, a microphone mounted in the resonant cavity, and a membrane that covers the resonant cavity. Low frequency acoustic signals that are incident on the membrane cause the membrane to move and amplify the acoustic signals within the resonant cavity. The sensor provides information concerning physiological conditions, such as respiration and cardiac activity. The sensor in a physiological condition monitor does not need to be directly coupled to the skin of the person being monitored. The physiological condition monitor simultaneously provides information concerning cardiac activity, and respiration activity, and the movement and position orientation of the monitored person's body. | 1. An apparatus for detecting very low frequency acoustic signals and signals indicative of body orientation comprising: a low frequency sensor capable of being acoustically coupled to a source of low frequency acoustic signals, said low frequency sensor capable of receiving low frequency acoustic signals in the frequency range of one tenth Hertz to thirty Hertz and generating electronic signals indicative of the intensity of said low frequency acoustic signals; and an apparatus capable of evaluating movement of a body relative to an environment of the type comprising: a sensor, associable with said body, that senses dynamic and static accelerative phenomena of said body; and a processor, associated with said sensor, that processes said sensed dynamic and static accelerative phenomena as a function of at least one accelerative event characteristic to thereby determine whether said evaluated body movement is within environmental tolerance. | 1. An apparatus for detecting very low frequency acoustic signals and signals indicative of body orientation comprising: a low frequency sensor capable of being acoustically coupled to a source of low frequency acoustic signals, said low frequency sensor capable of receiving low frequency acoustic signals in the frequency range of one tenth Hertz to thirty Hertz and generating electronic signals indicative of the intensity of said low frequency acoustic signals; and an apparatus capable of evaluating movement of a body relative to an environment of the type comprising: a sensor, associable with said body, that senses dynamic and static accelerative phenomena of said body; and a processor, associated with said sensor, that processes said sensed dynamic and static accelerative phenomena as a function of at least one accelerative event characteristic to thereby determine whether said evaluated body movement is within environmental tolerance. 2. An apparatus as claimed in claim 1 further comprising signal processing circuitry coupled to said low frequency sensor capable of processing said electronic signals from said low frequency sensor to obtain electronic signals indicative of the intensity of said low frequency acoustic signals in the frequency range of one tenth Hertz to two Hertz. 3. An apparatus as claimed in claim 1 further comprising signal processing circuitry coupled to said low frequency sensor capable of processing said electronic signals from said low frequency sensor to obtain electronic signals indicative of the intensity of said low frequency acoustic signals in the frequency range of ten Hertz to thirty Hertz. 4. The apparatus as claimed in claim 1 wherein said low frequency sensor comprises: a chamber capable of being acoustically coupled to a source of acoustic signals, said chamber having portions that define a cavity within said chamber; and a microphone capable of receiving low frequency acoustic signals within said cavity of said chamber in the frequency range of one tenth Hertz to thirty Hertz and capable of generating electronic signals indicative of the intensity of said low frequency acoustic signals. 5. The apparatus as claimed in claim 4 wherein said chamber comprises one of: a closed chamber filled with fluid and an open chamber filled with air. 6. The apparatus as claimed in claim 4 wherein said chamber is formed having nonrigid wall are capable of moving inwardly and outwardly with respect to the interior of said cavity in response to the presence of low frequency acoustic energy. 7. A physiological condition monitor for detecting very low frequency acoustic signals of a person and signals indicative of a person's body orientation comprising: a low frequency sensor capable of being acoustically coupled to the body of the person being monitored and capable of receiving low frequency acoustic signals in the range of one tenth Hertz to thirty Hertz and capable of generating electronic signals indicative of the in... | iLife Solutions, Inc. | A61B007/04 A61B005/02
US06866642 | B2 | 2005-03-15 | Enhanced method for joining two core wires | 2002-11-25 | An intravascular guide wire having at least two core materials joined together. There is a wire core having a proximal core section with a proximal end and a distal end and a distal core section with a proximal end and a distal end. The distal end of the proximal core section and the proximal end of the distal core section are formed into complementary shapes, and then placed into a flexible sleeve in opposing directions. Inside the flexible sleeve, the complementary shaped ends are joined together through bonding, welding, brazing, cementing, or soldering. The flexible sleeve can be either a stretched coil or a polyimide sleeve, each with an outer diameter similar to the outer diameter of the core wire, therefore the guide wire does not require additional grinding to reduce the outer diameter of the joined section. | 1. An intravascular guide wire, comprising: a core having a proximal core section with proximal and distal ends and a distal core section with proximal and distal ends; and a flexible sleeve covering the distal end of the proximal core section and the proximal end of the distal core section, the sleeve having proximal and distal ends, each end of the sleeve circumferentially surrounding the core, the sleeve defining a plurality of openings such that each end of the sleeve has a plurality of circumferential discontinuities; wherein the distal end of the proximal core section and the proximal end of the distal core section overlap one another. | 1. An intravascular guide wire, comprising: a core having a proximal core section with proximal and distal ends and a distal core section with proximal and distal ends; and a flexible sleeve covering the distal end of the proximal core section and the proximal end of the distal core section, the sleeve having proximal and distal ends, each end of the sleeve circumferentially surrounding the core, the sleeve defining a plurality of openings such that each end of the sleeve has a plurality of circumferential discontinuities; wherein the distal end of the proximal core section and the proximal end of the distal core section overlap one another. 2. The guide wire of claim 1, further comprising a mass of hardened material at a joint connecting the distal end of the proximal core section to the proximal end of the distal core section. 3. The guide wire of claim 2, wherein the mass of hardened material is disposed between the distal end of the proximal core section and the proximal end of the distal core section and encases the ends within the flexible sleeve. 4. The guide wire of claim 2, wherein the mass of hardened material includes a bonding material selected from the group consisting of solders, brazes, adhesives, epoxies, glues, laser welds, and spot welds. 5. The guide wire of claim 1, wherein the flexible sleeve has an outer diameter that is generally equal to the outer diameter of the core. 6. The guide wire of claim 1, wherein at least one of the distal end of the proximal core section and the proximal end of the distal core section includes a taper. 7. The guide wire of claim 1, wherein at least one of the distal end of the proximal core section and the proximal end of the distal core section has a reduced diameter. 8. The guide wire of claim 1, wherein the flexible sleeve is a polymeric sleeve. 9. The guide wire of claim 8, wherein the polymeric sleeve is a polyimide sleeve. 10. The guide wire of claim 8, wherein the polymeric sleeve has an outer diameter generally equal to the outer diameter of the core. 11. The guide wire of claim 8, further comprising a mass of hardened material connecting the distal end of the proximal core section to the proximal end of the distal core section inside the polymeric sleeve. 12. A method for joining two intravascular guide wire core materials, comprising: providing a proximal core section with a proximal and distal end and a distal core section with a proximal and distal end; forming a complementary shape into the distal end of the proximal core section and the proximal end of the distal core section; providing a flexible sleeve having proximal and distal ends; creating a plurality of openings in the sleeve so that each end of the sleeve has a plurality of circumferential discontinuities; and inserting the proximal and distal core sections into the flexible sleeve in opposing directions with the complementary shaped ends overlapping one another. 13. The method of claim 12, wherein forming the... | Advanced Cardiovascular Systems, Inc. | A61B005/00
US06939313 | B2 | 2005-09-06 | Device for sensing parameters of a hollow body organ | 2002-05-14 | A device for sensing the profile of a hollow body organ includes a hollow guidewire or probe carrying a sensor. The hollow probe includes a distal coiled end tube with a tip bent at an angle between about 10° to 90° relative to the probe longitudinal axis and having a torque response of no more than about 0.4 ozf in. The flexural rigidity of the distal end tube relative to the mid and proximal end tubes forming the probe is generally in the range of about 1:(4 to 4,400):(110 to 13,000), respectively. An actuator is detachably secured to the probe and includes one or several motors which allow the sensor to travel in a rotational and axial path through the body organ. | 1. A diagnostic probe suitable for insertion into a blood vessel comprising: an elongate stem including a proximal end portion and a distal end portion terminating in a coil and having at least a segment thereof biased away from the rotational axis of the probe and having a torque response of no more than about 0.4 ozf in.; a transducer mounted to the coil; a connector at the proximal end portion of the stem; and at least one conductor connecting the transducer to the connector. | 1. A diagnostic probe suitable for insertion into a blood vessel comprising: an elongate stem including a proximal end portion and a distal end portion terminating in a coil and having at least a segment thereof biased away from the rotational axis of the probe and having a torque response of no more than about 0.4 ozf in.; a transducer mounted to the coil; a connector at the proximal end portion of the stem; and at least one conductor connecting the transducer to the connector. 2. The diagnostic probe of claim 1 wherein an elongate shaping mandrel extends through the coil. 3. The diagnostic probe of claim 1 wherein the coil is rotatable relative to the proximal end portion of the stem and the probe further comprises a rotatable actuation shaft extending through the proximal end portion of the stem and operably associated with the coil. 4. The diagnostic probe of claim 3 further comprising an actuator adapted for coupling to the proximal end portion of the stem and including a motor operably coupled to the actuation shaft. 5. The diagnostic probe of claim 3 wherein the proximal end portion of the stem defines a sheath slidable on the actuation shaft and operably coupled to the coil for allowing the axial reciprocal movement of the coil. 6. The diagnostic probe of claim 1 wherein the coil is biased at an angle between about 10° to 90° relative to the longitudinal axis of the stem. 7. The diagnostic probe of claim 6 wherein the coil is biased at an angle between about 40° to 60° relative to the longitudinal axis of the stem. 8. The diagnostic probe of claim 6 wherein the coil is biased away from the longitudinal axis of the stem. 9. The diagnostic probe of claim 6 wherein the coil is biased towards the longitudinal axis of the stem. 10. The diagnostic probe of claim 1 wherein the stem includes a mid portion between the proximal end portion and the coil and the flexural rigidity ratio between the proximal end portion and the mid portion and the coil is about 1 to 4,400 to 13,000, respectively. 11. The diagnostic probe of claim 1 further comprising a sensor mounted to the coil adjacent to the transducer for measuring an impedance against the blood vessel. 12. The diagnostic probe of claim 1 wherein the transducer is adapted to have a measurement response time between about 10 to 100 msec. 13. The diagnostic probe of claim 1 wherein the probe is adapted to be withdrawn longitudinally through the blood vessel at a rate of 0.3 to 10 mm/sec. 14. The diagnostic probe of claim 1 wherein the probe is adapted to rotate about a longitudinal axis defined by the elongate stem at a rate of between 1 to 100 rpm. 15. A device for sensing the profile of a hollow body organ, comprising: a hollow guidewire having a relaxed bent configuration defining an angle between about 10° to 90° relative to a longitudinal axis of the guidewire, and a contracted configuration; and a sensor connected to the guidewire and moveable therewith, the sensor being displace... | - | A61B005/103 A61B005/117
US06931273 | B2 | 2005-08-16 | Database of body surface ECG P wave integral maps for localization of left-sided atrial arrhythmias | 2001-04-11 | A system and method are provided for developing a database of body surface ECG P-wave maps for classification and localization of left-sided atrial arrhythmias. The system and method include generating and receiving P-wave data in a subject by left atrial pacing or receiving P-wave data in a subject during spontaneously occurring or induced left atrial arrhythmias; computing (e.g. potential or integral) maps of the P-wave data; classifying the maps specific to a left atrial ectopic origin; verifying the classification procedure; averaging the classified maps into mean maps; and storing and accessing the mean maps in the database. The mean maps of the P-wave data in the database can be used to automatically classify and localize P-wave data from a patient obtained during a left atrial arrhythmia such as atrial tachycardia, focal atrial fibrillation, or orthodromic atrioventricular reentrant tachycardia. | 1. A system for developing a database of body surface ECG P wave data for classification and localization of left atrial arrhythmias wherein said database is derived from P wave data of two or more subjects comprising: (a) means for receiving said P wave data from a subject; (b) means for classifying said P wave data and thereby creating classified P wave data; (c) means for averaging said classified P wave data and thereby creating mean P wave data; and (d) means for storing and accessing said mean P wave data in said database. | 1. A system for developing a database of body surface ECG P wave data for classification and localization of left atrial arrhythmias wherein said database is derived from P wave data of two or more subjects comprising: (a) means for receiving said P wave data from a subject; (b) means for classifying said P wave data and thereby creating classified P wave data; (c) means for averaging said classified P wave data and thereby creating mean P wave data; and (d) means for storing and accessing said mean P wave data in said database. 2. The system as set forth in claim 1, wherein one of said left atrial arrhythmias is an atrial tachycardia, a focal atrial fibrillation, or an orthodromic AV reentrant tachycardia involving an accessory pathway. 3. The system as set forth in claim 1, wherein said P wave data is in the form of maps. 4. The system as set forth in claim 3, wherein said maps are in the form of integral maps. 5. The system as set forth in claim 3, wherein said maps are in the form of potential maps. 6. The system as set forth in claim 1, further comprises means for obtaining said P wave data by electrically stimulating a left atrium of said subject using a probe. 7. The system as set forth in claim 6, wherein said P wave data is obtained using a transseptal procedure. 8. The system as set forth in claim 6, wherein said P wave data is obtained using a retrograde aortic procedure. 9. The system as set forth in claim 1, further comprises means for obtaining said P wave data from spontaneously occurring left atrial arrhythmias. 10. The system as set forth in claim 1, further comprises means for obtaining said P wave data by inducing left atrial arrhythmias. 11. The system as set forth in claim 1, wherein said receiving means further comprises means for sensing heart cycle signals while said subject is spontaneously producing a left atrial arrhythmia, while a left atrial arrhythmia is induced in said subject, or while a left atrial pacing is being performed in said subject. 12. The system as set forth in claim 11, wherein said receiving means further comprises means for detecting an electrical heart signal with a plurality of sensors proximate a subject's torso. 13. The system as set forth in claim 12, wherein said receiving means further comprises means for separating from said electrical heart signal an atrial signal obscured by a ventricular signal. 14. The system as set forth in claim 12, wherein said receiving means further comprises means for selecting at least one reference cycle from among a plurality of heart cycles to determine said P wave data. 15. The system as set forth in claim 14, wherein said receiving means further comprises means for selecting a time interval of said reference cycle and comparing signals from said plurality of sensors during said selected time interval. 16. The system as set forth in claim 15, wherein said receiving means further comprises means for generating a data matrix by integrating said sig... | University of California San Francisco | A61B005/04
US06884214 | B2 | 2005-04-26 | Daily-life disability related physical information determining apparatus | 2002-10-21 | An apparatus firstly acquires a value involving in a physical condition, such as a value relating to a body constitution, a value relating to a basic physical strength or a value relating to a motive ability, by acquisition means 101, secondly, based on said value, the apparatus specifies in specification means 103 the physical information involving in the physical strength related phase of the daily-life disability, such as a value relating to the basic strength, a value relating to the motive ability, or a value relating to a daily-life disability, all of which are configured for an aged person placed as a subject and stored in storage means 102, and finally the apparatus in output means 104 outputs thus specified information. Thereby, the determination on the physical information relating to the physical strength related phase of the daily-life disability may become possible automatically. | 1. A daily-life disability related physical information determining apparatus, comprising acquisition means, storage means, specification means and output means, wherein said acquisition means acquires at least two physical condition values selected from a group consisting of an entire body muscle mass, a percent entire body muscle, a leg muscle mass, a percent leg muscle, an arm muscle mass, a percent body fat, a body fat mass, a lean body mass, a percent lean body, a leg lean body mass, a percent leg lean body, an arm lean body mass, a gravity center sway distance, a gravity center sweep area, an extracellular fluid volume, an intracellular fluid volume, and a pulse differential count between sitting position and standing position; said storage means stores an active motion disabled level, which have been associated with the at least two values acquired by said acquisition means; said specification means specifies the active motion disabled level stored in said storage means based on the at least two values acquired by said acquisition means; and said output means outputs the active motion disabled level specified by said specification means. | 1. A daily-life disability related physical information determining apparatus, comprising acquisition means, storage means, specification means and output means, wherein said acquisition means acquires at least two physical condition values selected from a group consisting of an entire body muscle mass, a percent entire body muscle, a leg muscle mass, a percent leg muscle, an arm muscle mass, a percent body fat, a body fat mass, a lean body mass, a percent lean body, a leg lean body mass, a percent leg lean body, an arm lean body mass, a gravity center sway distance, a gravity center sweep area, an extracellular fluid volume, an intracellular fluid volume, and a pulse differential count between sitting position and standing position; said storage means stores an active motion disabled level, which have been associated with the at least two values acquired by said acquisition means; said specification means specifies the active motion disabled level stored in said storage means based on the at least two values acquired by said acquisition means; and said output means outputs the active motion disabled level specified by said specification means. 2. The apparatus according to claim 1, wherein said storage means further stores a first advice information and a second advice information, the first advice information being designated as text information presented as an advice to a person to be measured, including a warning, an evocation, a reminder and a caution with respect to a result of the active motion disabled level specified by said specification means, and the second advice information being designated as text information presented as an advice to the person to be measured, including a warning, an evocation, a reminder and a caution with respect to a result for each of the at least two physical condition values acquired by said acquisition means, said specification means further specifies the first advice information stored in said storage means based on the previously specified active motion disabled level and the second advice information stored in said storage means based on each of the at least two values acquired by said acquisition means, and said output means further output, at the same time, the first information and the second information specified by said specification means. 3. The apparatus according to claim 1, wherein said output means is display means for indicating a result of the active motion disabled level specified by said specification means, the display means indicating the result in a color associated with the active motion disabled level to provide a graphical representation of the active motion disabled level, by associating the result with a graph having a graph axis which is associated with each of the at least two physical condition values acquired by said acquisition means. 4. The apparatus according to anyone of claim 1 or 3, wherein said acquisition means acquires anyone at least two physical condition values... | Tanita Corporation | A61B005/00
US06934581 | B2 | 2005-08-23 | Non-invasive method for disease diagnosis | 2002-08-01 | A diagnostic method comprises selecting 24 biologically active points (BAPs), measuring the skin resistance at the points, plotting the average resistance for the points as an isoelectric line and defining a normal corridor. A stimulation is applied to each of the points and thereafter the skin resistance is once again measured. The results of the two sets of measurements are compared and diagnostic conclusions are reached from their comparison. | 1. Diagnostic method, comprising the steps of selecting 24 biologically active points (BAPs), measuring the skin resistance at said points, plotting the average resistance for said points as an isoelectric line and defining a normal corridor, further comprising the steps of: A-a stimulation is applied to each of the said points, B-the skin resistance is measured once again at said points after applying the stimulation, C-the results of the two sets of measurements are compared, and D-diagnostic conclusions are reached from their comparison. | 1. Diagnostic method, comprising the steps of selecting 24 biologically active points (BAPs), measuring the skin resistance at said points, plotting the average resistance for said points as an isoelectric line and defining a normal corridor, further comprising the steps of: A-a stimulation is applied to each of the said points, B-the skin resistance is measured once again at said points after applying the stimulation, C-the results of the two sets of measurements are compared, and D-diagnostic conclusions are reached from their comparison. 2. Method according to claim 1, wherein the diagnostic conclusions comprise the following: I-If a result obtained from the first set of measurements falls outside the normal corridor, this is considered a potential indicator of disease activity, II-if the corresponding result obtained from the second set of measurements (after stimulation) also lies outside the normal corridor, this is considered as a true indication of the presence of a disease, III-If, however, the corresponding result from the second set of measurements falls within the normal corridor, then the first measurement is considered not to be an indication of a true disease state and is therefore disregarded, and IV-If there are deviations in the meridian (outside the normal corridor) only after the stimulation, this is considered an indication of the presence of functional problems only in the corresponding organ. 3. Method according to claim 1, further comprising the steps of, if an indication of a disease state in a given organ or organ system is found: a-measuring the electrical resistance of the relevant announcement point; b-measuring the electrical resistance of the abdominal zero point; c-calculating the ratio between the said two results; and d-drawing diagnostic conclusions from said calculated ratio. 4. Method according to claim 1, comprising the following steps: I-measuring the electric resistance of the skin at the 24 BAPs in the feet and hands of the patient; II-applying a stimulation to the BAPs by means of a skin electropunctural probe; III-repeating the measurement of the electric resistance of the skin at the 24 BAPs; IV-taking electric resistance measurements of the 10 announcement points on the abdomen; V-measuring the parameters of the 12 sympathetic points in the paravertebral zone; VI-measuring the zero energy for announcement points and for sympathetic points; and VII-mathematically processing the collected information. 5. Method according to claim 1, further comprising previously treating the BAPs with 70% ethyl alcohol solution. 6. Method according to claim 1, wherein the stimulation is carried out repeatedly on each BAP for one minute each time. 7. Method according to claim 1, further comprising the steps of adding together the 48 data received from 24 BAPs before and after the stimulation, dividing the results by 24, plotting the figures thus obtained as the ordinate on a diagram on which the BAPs are marked ... | Medex Screen Ltd. | A61B005/05
US06980852 | B2 | 2005-12-27 | Film barrier dressing for intravascular tissue monitoring system | 2002-08-23 | A medical apparatus comprises a dressing, one or more sensor elements integrated into the dressing, a flexible membrane coupled to the dressing, and an adhesive layer coupled to the dressing and coupled to the flexible membrane, the adhesive layer capable of attaching the apparatus to a patient's skin. | 1. An apparatus comprising: a dressing;one or more sensor elements integrated into the dressing further comprising: patterned conductive inks integrated into the dressing;dielectric insulators selectively patterned in combination with the patterned conductive inks to form the one or more sensor elements; anda foam layer coupled to the patterned conductive inks and the dielectric insulators;a flexible membrane coupled to the dressing; andan adhesive layer coupled to the dressing and coupled to the flexible membrane, the adhesive layer capable of attaching the apparatus to a patient's skin. | 1. An apparatus comprising: a dressing;one or more sensor elements integrated into the dressing further comprising: patterned conductive inks integrated into the dressing;dielectric insulators selectively patterned in combination with the patterned conductive inks to form the one or more sensor elements; anda foam layer coupled to the patterned conductive inks and the dielectric insulators;a flexible membrane coupled to the dressing; andan adhesive layer coupled to the dressing and coupled to the flexible membrane, the adhesive layer capable of attaching the apparatus to a patient's skin. 2. An apparatus according to claim 1 wherein: the flexible membrane is at least partially transparent so that a patient's tissue underlying the apparatus can be visually inspected. 3. An apparatus according to claim 1 wherein the flexible membrane is a breathable barrier film that protects against and functions as a structural member that can secure an intravenous catheter. 4. An apparatus according to claim 1 further comprising: a bio-impedance sensor integrated into the one or more sensor elements. 5. An apparatus according to claim 1 further comprising: a bio-impedance sensor and an optical sensor integrated into the one or more sensor elements. 6. An apparatus according to claim 1 further comprising: a bio-impedance sensor including a plurality of sensor elements configured to acquire impedance signals in three-dimensional space. 7. An apparatus according to claim 1 further comprising: a sensor electrode array including equally-spaced electrodes in a geometrical configuration. 8. A film barrier dressing comprising: a flexible frame with an interior void;a flexible protective film coupled to the flexible frame and extending over the interior void, the flexible protective film being at least partly transparent to permit viewing through the flexible frame interior void;one or more sensor elements integrated into one or more portions of the flexible frame and the flexible protective film comprising: patterned conductive inks integrated into the film barrier dressing;dielectric insulators selectively patterned in combination with the patterned conductive inks to form the one or more sensor elements; anda foam layer coupled to the patterned conductive inks and the dielectric insulators. 9. A film hairier dressing according to claim 8 further comprising: an adhesive layer coupled to the film barrier dressing and coupled to the flexible membrane, the adhesive layer capable of attaching the apparatus to a patient's skin. 10. A film baffier dressing according to claim 8 wherein: the flexible protective film further comprises a microporous breathable barrier film capable of passing water vapor while preventing bacteria and viral transmission. 11. A film barrier dressing according to claim 8 wherein: the flexible protective film further comprises a dense monolithic medium breathable barrier film capable of passing water vapor while preventing bacteria and vira... | SubQIVIEW Inc. | A61B005/05
US06950694 | B2 | 2005-09-27 | Automatic input impedance balancing for electrocardiogram (ECG) sensing applications | 2002-12-17 | A voltage sensing system includes input impedance balancing for electrocardiogram (ECG) sensing or other applications, providing immunity to common-mode noise signals while capable of use with two electrodes. Signals are received at first and second electrodes having associated impedances. An impedance circuit includes a feedback controller that adjusts an effective impedance associated with the second electrode based on a difference signal, a common mode signal, a phase-shifted (e.g., quadrature common mode) signal, and an impedance associated with the first electrode. As a result, signals associated with each electrode undergo a similar degree of gain/attenuation and/or phase-shift. This reduces common mode noise and enhances the signal-to-noise characteristics of a desired ECG or other output signal, without requiring the use of more than two electrodes. | 1. An apparatus for sensing a difference between first and second input voltages, the apparatus comprising: a reference circuit, where the reference circuit includes an input receiving an amplified signal based on first input voltage, and an output providing a reference voltage; a compensation circuit, where the compensation circuit includes an input receiving an amplified signal based on the second input voltage, and an output providing a compensated voltage based on reference voltage; a first amplification circuit, where the first amplification circuit includes a first input receiving a signal based on the reference voltage, a second input receiving a signal based on the compensated voltage, and an output providing a differential output signal based on a difference between the signals of the first and second inputs of the first amplification circuit; an averager circuit, where the averager circuit includes a first input receiving a signal based on the reference voltage, a second input receiving a signal based on the compensated voltage, and an output providing a common mode (CM) output signal based on the signals at the first and second inputs of the averager circuit; and an impedance circuit means for adjusting an impedance of the compensation circuit, based on the signals received from the outputs of the first amplification circuit and the averager circuit for reducing an impact of the CM output signal on the output of the first amplification circuit. | 1. An apparatus for sensing a difference between first and second input voltages, the apparatus comprising: a reference circuit, where the reference circuit includes an input receiving an amplified signal based on first input voltage, and an output providing a reference voltage; a compensation circuit, where the compensation circuit includes an input receiving an amplified signal based on the second input voltage, and an output providing a compensated voltage based on reference voltage; a first amplification circuit, where the first amplification circuit includes a first input receiving a signal based on the reference voltage, a second input receiving a signal based on the compensated voltage, and an output providing a differential output signal based on a difference between the signals of the first and second inputs of the first amplification circuit; an averager circuit, where the averager circuit includes a first input receiving a signal based on the reference voltage, a second input receiving a signal based on the compensated voltage, and an output providing a common mode (CM) output signal based on the signals at the first and second inputs of the averager circuit; and an impedance circuit means for adjusting an impedance of the compensation circuit, based on the signals received from the outputs of the first amplification circuit and the averager circuit for reducing an impact of the CM output signal on the output of the first amplification circuit. 2. The apparatus of claim 1, further includes a phase-shifter circuit means for providing a phase-shifted CM output signal from an output of the averager circuit. 3. The apparatus of claim 2, in which the impedance circuit means includes a feedback controller circuit, where the feedback controller circuit includes: first mixer means, coupled for receiving the differential output signal from the first amplification circuit and the CM signal from the averager circuit, for providing a first mixer output based on the differential and CM output signals; and second mixer means, coupled for receiving the differential output signal from the first amplification circuit and the phase-shifted CM output signal from the phase-shifter circuit means, and for providing a second mixer output based on the differential and the phase-shifted CM output signals. 4. The apparatus of claim 3, in which the impedance circuit means further includes: first filter means for coupling the differential output signal from the output of the first amplification circuit to the second mixer means; second filter means for coupling the CM output signal from the output of the phase-shifter circuit means to the first mixer means; and third filter means for coupling the phase-shifted CM output signal from the output of the phase-shifter circuit means to the second mixer means. 5. The apparatus of claim 1, further comprising a first input circuit coupled between a first electrode and the input of the reference circuit and in which... | Cardiac Pacemakers, Inc. | A61B005/04
US06912414 | B2 | 2005-06-28 | Electrode systems and methods for reducing motion artifact | 2003-01-28 | An electrode system for reducing noise from an electronic signal, the system including an electrode that provides the electronic signal, and a sensor that senses motion and provides a motion signal. The electrode system includes a controller that determines a noise value based on an analysis of the motion signal, and subtracts the noise value from the electronic signal. The electrode system can reduce or eliminate motion artifact from an electronic signal that can result in misdiagnosis, prolong procedural duration and inappropriate treatment of a patient. | 1. An electrode system for reducing motion artifact from an electronic signal, the system comprising: an electrode that provides an electronic signal; a sensor that senses electrode motion and provides a motion signal; and a controller that determines a noise value based on an analysis of the electronic signal and the motion signal, and subtracts the noise value from the electronic signal. | 1. An electrode system for reducing motion artifact from an electronic signal, the system comprising: an electrode that provides an electronic signal; a sensor that senses electrode motion and provides a motion signal; and a controller that determines a noise value based on an analysis of the electronic signal and the motion signal, and subtracts the noise value from the electronic signal. 2. The electrode system of claim 1, further comprising the electrode and the sensor coupled together and connected to a patch that is placed on the skin of a patient. 3. The electrode system of claim 1, wherein the sensor senses the motion in at least one dimension. 4. The electrode system of claim 1, wherein the controller is part of a patient monitoring system. 5. The electrode system of claim 4, the controller uses a noise cancellation circuit or routine to subtract the noise value from the electronic signal. 6. The electrode system of claim 1, further comprising a plurality of electrodes that produce a plurality of electronic signals that are differential physiological signals of a human being. 7. A method for reducing motion artifact from an electronic signal, comprising: measuring an electronic signal through an electrode; sensing electrode motion through a sensor; generating a motion signal based on the sensed electrode motion; determining a noise value based on an analysis of the electronic signal and the motion signal; and subtracting the noise value from the electronic signal. 8. The method of claim 7, further comprising coupling the electrode and the sensor together and connecting them to a patch that is placed on the skin of a patient. 9. The method of claim 7, further comprising sensing the motion in at least one dimension. 10. The method of claim 7, wherein a patient monitoring system determines the noise value based on the analysis of the electronic signal and motion signal, and subtracts the noise value from the electronic signal. 11. The method of claim 10, the patient monitoring system uses a noise cancellation circuit or routine to subtract the noise value from the electronic signal. 12. The method of claim 7, further comprising producing a plurality of electronic signals that are differential physiological signals of a human being. 13. A method for reducing motion artifact from differential electronic signals, comprising: measuring differential electronic signals produced by at least two recording electrodes; sensing electrode motion through a sensor associated with each of the at least two recording electrodes; generating motion signals based on the sensed electrode motion of each of the at least two recording electrodes; analyzing the motion signals; determining noise values based on the analysis of the motion signals; and subtracting the noise values from the differential electronic signals. 14. The method of claim 13, further comprising determining the motion signals based on at least one dimension. 15. The method of cla... | Southwest Research Institute | A61B005/0408
US06926675 | B2 | 2005-08-09 | Catheter heating circuit for continuous cardiac output measurement | 2002-01-29 | The present invention concerns a catheter heating circuit for continuous cardiac output measurement. The circuit comprises a DC power source for supplying an operating voltage for the circuit, wherein the heating power from the heating circuit is fed to the catheter via an isolating transformer. Isolation is known per se and it is not described in detail. According to the invention the heating circuit further comprises means for generating a bipolar square wave with variable duty cycle, which square wave is supplied to said isolating transformer and further to the catheter filament. | 1. A catheter heating circuit for heating a catheter for continuous cardiac output measurement, the circuit comprising an isolation transformer couplable to the catheter; a single supply DC power source for supplying an operating voltage for the circuit, wherein heating power from the heating circuit is adapted to be fed to the catheter via the isolating transformer; and a subcircuit connected between the DC source and the isolation transformer for generating a bipolar square wave with variable duty cycle. | 1. A catheter heating circuit for heating a catheter for continuous cardiac output measurement, the circuit comprising an isolation transformer couplable to the catheter; a single supply DC power source for supplying an operating voltage for the circuit, wherein heating power from the heating circuit is adapted to be fed to the catheter via the isolating transformer; and a subcircuit connected between the DC source and the isolation transformer for generating a bipolar square wave with variable duty cycle. 2. Catheter heating circuit according to claim 1, wherein the subcircuit for generating a bipolar square wave with variable duty cycle comprises a bridge circuit, comprising four on/off controllable switches elements, and a bridge controller, which is connected to said bridge circuit for controlling the state of said switching elements according to a predetermined controlling scheme. 3. Catheter heating circuit according to claim 2, wherein said catheter heating circuit further comprises a filter circuit, which is connected either in a non-floating side of the heating circuit between said bridge circuit and said isolating transformer or in a floating side of the heating circuit between said isolating transformer and the catheter. 4. Catheter heating circuit according to claim 2, wherein said circuit further comprises a measurement circuit for measuring power delivered to the catheter; and a feedback loop for supplying to said bridge controller an indication of the power delivered to the catheter. 5. Catheter heating circuit according to claim 2, wherein said bridge controller is connected to said DC power source for measuring and output voltage of said DC power source. 6. Catheter heating circuit according to claim 4, wherein said feedback loop comprises an isolator circuit for galvanically isolating said power indication delivered to the catheter. 7. Catheter heating circuit according to claim 6, wherein said isolator circuit is an opto-isolator. 8. Catheter heating circuit according to claim 2, wherein said bridge controller comprises a microcontroller and supporting circuitry. 9. Catheter heating circuit according to claim 2, wherein the that the two halves of the said bridge circuit are adapted to be controlled by identical square waves with an adjustable phase shift between them. 10. Catheter heating circuit according to claim 1, wherein said circuit further comprises a measurement circuit for measuring the power delivered to the catheter. 11. Catheter heating circuit according to claim 1, wherein said heating circuit further comprises a safety switch for disconnecting the heating power if needed. | GE Healthcare Finland Oy | A61B005/028
US06849045 | B2 | 2005-02-01 | Computerized medical diagnostic and treatment advice system including network access | 2002-11-19 | A system and method for providing computerized, knowledge-based medical diagnostic and treatment advice. The medical advice is provided to the general public over networks, such as a telephone network or a computer network. The invention also includes a stand-alone embodiment that may utilize occasional connectivity to a central computer by use of a network, such as the Internet. New authoring languages, interactive voice response and speech recognition are used to enable expert and general practitioner knowledge to be encoded for access by the public. “Meta” functions for time-density analysis of a number of factors regarding the number of medical complaints per unit of time are an integral part of the system. A re-enter feature monitors the user's changing condition over time. A symptom severity analysis helps to respond to the changing conditions. System sensitivity factors may be changed at a global level or other levels to adjust the system advice as necessary. | 1. A medical diagnostic system comprising: a computing environment arranged to receive patient information, comprising: a communications network; a server computer connected to the network; a client processor connected to the network and in data communication with the server computer via the network; an input device, connected to the client processor, to receive patient information from a user who is the patient or a proxy or assistant thereto; a script database; an output device, connected to the client processor, to provide information indicative of a diagnosed medical condition information; the environment executing programs for diagnosing by scoring a medical condition based on said patient information and communicating said diagnosed condition via said output device; said programs comprising: a plurality of medical complaint algorithms selectively executed based on at least a portion of the received information, wherein any one of the medical complaint algorithms scores at least a portion of the received information and diagnoses a medical condition associated with the executed medical complaint algorithm if the score reaches or passes a threshold; and wherein the program is arranged to read data from, and write data to, a medical history database stored in association with the client processor or said server computer, wherein any such scored information transferred between the client processor and the server computer is securely transmitted; and wherein each medical complaint algorithm is executed by said client processor, by a script engine program which is arranged: to interpret a medical complaint script file comprising lists of disease data, symptom data and question data; to output questions to said user; to input answers from said user via said input device; and to score said patient data and diagnose a said medical condition and update said medical history database using said scoring; and wherein said server computer is arranged to selectively download a said medical complaint script file to said client processor. | 1. A medical diagnostic system comprising: a computing environment arranged to receive patient information, comprising: a communications network; a server computer connected to the network; a client processor connected to the network and in data communication with the server computer via the network; an input device, connected to the client processor, to receive patient information from a user who is the patient or a proxy or assistant thereto; a script database; an output device, connected to the client processor, to provide information indicative of a diagnosed medical condition information; the environment executing programs for diagnosing by scoring a medical condition based on said patient information and communicating said diagnosed condition via said output device; said programs comprising: a plurality of medical complaint algorithms selectively executed based on at least a portion of the received information, wherein any one of the medical complaint algorithms scores at least a portion of the received information and diagnoses a medical condition associated with the executed medical complaint algorithm if the score reaches or passes a threshold; and wherein the program is arranged to read data from, and write data to, a medical history database stored in association with the client processor or said server computer, wherein any such scored information transferred between the client processor and the server computer is securely transmitted; and wherein each medical complaint algorithm is executed by said client processor, by a script engine program which is arranged: to interpret a medical complaint script file comprising lists of disease data, symptom data and question data; to output questions to said user; to input answers from said user via said input device; and to score said patient data and diagnose a said medical condition and update said medical history database using said scoring; and wherein said server computer is arranged to selectively download a said medical complaint script file to said client processor. 2. The system of claim 1, wherein the medical condition comprises a disease. 3. The system of claim 1, wherein the communications network comprises the Internet. 4. The system of claim 1, wherein the client processor utilizes a web browser. 5. The system of claim 1, additionally comprising a treatment table accessed according to the diagnosed medical condition. 6. The system of claim 1, wherein each of the plurality of medical complaint algorithms is associated with one or more medical conditions. 7. The system of claim 1, wherein the diagnosed medical condition is communicated to a patient. 8. The system of claim 1, wherein the diagnosed medical condition is communicated to a physician. 9. The system of claim 8, wherein the score associated with the diagnosed medical condition is communicated to the physician. 10. The system of claim 1, wherein a plurality of potentially diagnosed medical conditions are accumulat... | First Opinion Corporation | A61B005/00
US06947790 | B2 | 2005-09-20 | Neurocognitive function EEG measurement method and system | 2002-04-15 | An efficient, objective testing method and system for evaluating changes in mental function is described. The method and system are based on measuring an individual's behavioral responses and brain function during a brief cognitive test battery and passive control conditions. The method and system is designed to assess an individual's fundamental cognitive functions, and whether those functions have been significantly affected by a variety of factors such as progressive disease processes, medication, stress, fatigue, training, or the passage of time. The method and system can be used to determine whether drugs being evaluated to treat diseases or conditions affecting cognitive brain function have a significant positive effect on delaying or improving the symptoms of such a disease or condition, especially during clinical trials for drug approval and subsequent marketing. The method and system may also be employed as part of the successful diagnosis or ongoing treatment of neurological diseases or conditions that directly or indirectly affect human neurocognitive performance. The method and system may also be used to determine transitory changes in overall cognitive function due to emotional stress or fatigue, and more long lasting changes in overall cognitive function following training and educational programs. | 1. The method of measuring changes in a human subject's fundamental cognitive brain functions, the changes, for example, being due to disease, injury, remedial treatment, the utilization of medicines and normal variation within and between days and over a period of time, including the steps of: (a) presenting an attention-damanding task to the subject, which engages one or more of the subject's fundamental cognitive functions, and, simultaneously; (b) at least once, measuring the subject's behavioral responses to the task, and neuroelectric activity at the subject's scalp using a set of electroencephalograph (EEG) electrodes and amplifier and analog/digital converters, to provide a set of baseline digital data representing the subject's baseline state behavioral responses and neuroelectric activity response to the task; (c) at least once again, measuring the subject's behavioral responses to the task, and neuroelectric activity at the subject's scalp using a set of electroencephalograph (EEG) electrodes and amplifier and analog/digital converters, to provide a set of possibly altered state digital data representing the subject's possibly altered state behavioral responses and neuroelectric activity in response to the task; (d) in a computer system, comparing the subject's baseline state and possibly altered state digital data using a mathematical function derived from behavioral responses and EEG derived neuroelectric activity responses of a reference group of subjects performing the same task recorded in their baseline and altered state conditions, the comparison using a multivariate statistical method combining measures of task performance with brain function measures in a single comparison, and (e) deriving one or more scores for the subject based on the comparison described in (d) and determining the significance of those scores. | 1. The method of measuring changes in a human subject's fundamental cognitive brain functions, the changes, for example, being due to disease, injury, remedial treatment, the utilization of medicines and normal variation within and between days and over a period of time, including the steps of: (a) presenting an attention-damanding task to the subject, which engages one or more of the subject's fundamental cognitive functions, and, simultaneously; (b) at least once, measuring the subject's behavioral responses to the task, and neuroelectric activity at the subject's scalp using a set of electroencephalograph (EEG) electrodes and amplifier and analog/digital converters, to provide a set of baseline digital data representing the subject's baseline state behavioral responses and neuroelectric activity response to the task; (c) at least once again, measuring the subject's behavioral responses to the task, and neuroelectric activity at the subject's scalp using a set of electroencephalograph (EEG) electrodes and amplifier and analog/digital converters, to provide a set of possibly altered state digital data representing the subject's possibly altered state behavioral responses and neuroelectric activity in response to the task; (d) in a computer system, comparing the subject's baseline state and possibly altered state digital data using a mathematical function derived from behavioral responses and EEG derived neuroelectric activity responses of a reference group of subjects performing the same task recorded in their baseline and altered state conditions, the comparison using a multivariate statistical method combining measures of task performance with brain function measures in a single comparison, and (e) deriving one or more scores for the subject based on the comparison described in (d) and determining the significance of those scores. 2. The method of measuring changes in a human subject's fundamental cognitive brain functions, the changes, for example, being due to disease, injury, remedial treatment, the testing of medicines and normal variation within and between days and over a period of time, including the steps of: (a) presenting an attention-demanding task to the subject, which engages the subject's fundamental cognitive functions, and, simultaneously; (b) at least once, measuring the subject's behavioral responses to the task, and associated brain function using functional magnetic resonance imaging (fMRI), alone or in combination with EEG, to provide a set of baseline digital data representing the subject's behavioral responses and brain activity in response to the task; (c) at least once again, measuring the subject's behavioral responses to the task, and associated brain function using functional magnetic imaging (fMRI), alone or in combination with EEG, to provide a set of possibly altered state digital data representing the subject's behavioral responses and brain activity in response to the task; (d) in a computer system, comparing... | SAM Technology, Inc. | A61B005/04 A61B005/05
US06955649 | B2 | 2005-10-18 | Arteriosclerosis evaluating apparatus | 2002-07-15 | An apparatus for evaluating arteriosclerosis of a living subject, including a blood-pressure measuring device which measures at least one blood-pressure value of the subject, a pulse-wave detecting device which detects a pulse wave from a portion of the subject, a relationship determining device for determining a relationship between pulse-wave magnitude and blood pressure, based on at least one magnitude of the pulse wave detected by the pulse-wave detecting device and said at least one blood-pressure value measured by the blood-pressure measuring device, an incident-wave-peak determining device for determining a magnitude of a peak point of an incident-wave component which is contained in the pulse wave detected by the pulse-wave detecting device, a reflected-wave-occurrence-time determining device for determining a time of occurrence of a peak point of a reflected-wave component which is contained in the pulse wave, a pressure-difference determining device for determining, according to the relationship between pulse-wave magnitude and blood pressure, a pressure difference corresponding to a magnitude difference between a magnitude of the pulse wave at the time of occurrence of the peak point of the reflected-wave component, and the magnitude of the peak point of the incident-wave component, and a display device which displays the pressure difference determined by the pressure-difference determining device. | 1. An apparatus for evaluating arteriosclerosis of a living subject, comprising: a blood-pressure measuring device which measures at least one blood-pressure value of the subject;a pulse-wave detecting device which detects a pulse wave from a portion of the subject;a relationship determining means for determining a relationship between pulse-wave magnitude and blood pressure, based on at least one magnitude of the pulse wave detected by the pulse-wave detecting device and said at least one blood-pressure value measured by the blood-pressure measuring device;an incident-wave-peak determining means for determining a magnitude of a peak point of an incident-wave component which is contained in the pulse wave detected by the pulse-wave detecting device;a reflected-wave-occurrence-time determining means for determining a time of occurrence of a peak point of a reflected-wave component which is contained in the pulse wave detected by the pulse-wave detecting device;a pressure-difference determining means for determining, according to the relationship between pulse-wave magnitude and blood pressure determined by the relationship determining means, a pressure difference corresponding to a magnitude difference between a magnitude of the pulse wave at the time of occurrence of the peak point of the reflected-wave component determined by the reflected-wave-occurrence-time determining means, and the magnitude of the peak point of the incident-wave component determined by the incident-wave-peak determining means; anda display device which displays the pressure difference determined by the pressure-difference determining means. | 1. An apparatus for evaluating arteriosclerosis of a living subject, comprising: a blood-pressure measuring device which measures at least one blood-pressure value of the subject;a pulse-wave detecting device which detects a pulse wave from a portion of the subject;a relationship determining means for determining a relationship between pulse-wave magnitude and blood pressure, based on at least one magnitude of the pulse wave detected by the pulse-wave detecting device and said at least one blood-pressure value measured by the blood-pressure measuring device;an incident-wave-peak determining means for determining a magnitude of a peak point of an incident-wave component which is contained in the pulse wave detected by the pulse-wave detecting device;a reflected-wave-occurrence-time determining means for determining a time of occurrence of a peak point of a reflected-wave component which is contained in the pulse wave detected by the pulse-wave detecting device;a pressure-difference determining means for determining, according to the relationship between pulse-wave magnitude and blood pressure determined by the relationship determining means, a pressure difference corresponding to a magnitude difference between a magnitude of the pulse wave at the time of occurrence of the peak point of the reflected-wave component determined by the reflected-wave-occurrence-time determining means, and the magnitude of the peak point of the incident-wave component determined by the incident-wave-peak determining means; anda display device which displays the pressure difference determined by the pressure-difference determining means. 2. An apparatus according to claim 1, further comprising: a blood-pressure-related-information obtaining device which obtains blood-pressure-related information which changes in relation with blood pressure of the subject;a memory device which stores a plurality of sets of display information each of which includes a pressure difference determined by the pressure-difference determining means and a piece of blood-pressure-related information obtained by the blood-pressure-related-information obtaining device and; anda graph displaying means for operating the display device to display a two-dimensional graph defined by a first axis indicative of pressure difference and a second axis indicative of blood-pressure-related information, and additionally display, in the two-dimensional graph, a plurality of symbols at a plurality of positions, respectively, which correspond to said plurality of sets of display information, respectively, that are stored by the memory device. 3. An apparatus according to claim 2, further comprising an input device which is operable by an operator to input identification information identifying the subject. 4. An apparatus according to claim 3, further comprising a subject identifying means for identifying the subject based on the identification information inputted through the input device, wherein the memory ... | Colin Medical Technology Corporation | A61B005/02
US06968226 | B2 | 2005-11-22 | Method and system for terminating an atrial arrhythmia | 2002-01-30 | A method and system is provided for responding, from internally within a patient, to an atrial arrhythmia in a heart including measuring from within the patient at least one electrocardiogram characteristic indicative of the atrial arrhythmia, and controlling from within the patient drug therapy delivery to the patient responsive to measuring the at least one electrocardiogram characteristic. Drug therapy is initiated to the patient responsive to measuring the at least one electrocardiogram characteristic. According to one aspect of the present invention, the drug therapy is staged within the patient prior to measuring the at least one electrocardiogram characteristic. According to another example embodiment, the heart is paced from within the patient at a predefined rate responsive to measuring the at least one electrocardiogram characteristic, pacing occurring alone, or in combination with drug therapy. | 1. A method of responding, from internally within a patient, to an atrial arrhythmia in a heart, the method comprising: monitoring electrocardiograms indicative of cardiac activity within the heart;measuring from within the patient at least one electrocardiogram characteristic indicative of the atrial arrhythmia;initiating a drug delivery therapy responsive to the indicated atrial arrhythmia; andcontrolling from within the patient the drug therapy delivery to the patient responsive to indications of ventricular proarrhythmia from the monitored electrocardiogram. | 1. A method of responding, from internally within a patient, to an atrial arrhythmia in a heart, the method comprising: monitoring electrocardiograms indicative of cardiac activity within the heart;measuring from within the patient at least one electrocardiogram characteristic indicative of the atrial arrhythmia;initiating a drug delivery therapy responsive to the indicated atrial arrhythmia; andcontrolling from within the patient the drug therapy delivery to the patient responsive to indications of ventricular proarrhythmia from the monitored electrocardiogram. 2. The method of claim 1, further comprising: pacing the heart from within the patient at a predefined rate responsive to measuring at least one electrocardiogram characteristic after initiating drug therapy delivery. 3. The method of claim 1, further comprising initiating drug therapy delivery to the patient responsive to measuring the at least one electrocardiogram characteristic, the drug therapy being staged within the patient prior to measuring the at least one electrocardiogram characteristic. 4. The method of claim 3, wherein controlling the drug delivery therapy further comprises preventing the initiation of the drug therapy delivery to the patient that was responsive to measuring the at least one electrocardiogram characteristic, the drug therapy being staged internal to the patient prior to measuring the at least one electrocardiogram characteristic. 5. The method of claim 4, wherein drug therapy includes an anti-arrhythmic drug. 6. The method of claim 5, wherein the anti-arrhythmic drug includes Ibutilide. 7. The method of claim 4, wherein the at indication of ventricular proarrhythmia includes a Q–T interval. 8. The method of claim 7, wherein the indication of ventricular proarrhythmia is that the Q–T interval is greater than 480 mS. 9. The method of claim 1, further comprising: pacing the heart from within the patient at a predefined rate responsive to measuring the at least one electrocardiogram characteristic and prior to initiating the drug delivery therapy. 10. The method of claim 9, wherein controlling includes signaling the patient to initiate drug therapy delivery responsive to measuring the at least one electrocardiogram characteristic. 11. The method of claim 10, wherein the at least one electrocardiogram characteristic includes a time duration difference between a monitored Q–T interval and a selected Q–T interval duration range. 12. The method of claim 9, further comprising initiating drug therapy delivery to the patient responsive to the at least one electrocardiogram characteristic, the drug therapy being staged internal to the patient prior to the atrial arrhythmia. 13. The method of claim 12, wherein drug therapy delivery includes delivering an anti-arrhythmic drug. 14. The method of claim 13, wherein the anti-arrhythmic drug includes Ibutilide. 15. The method of claim 12, wherein the at least one electrocardiogram characteristic includes a time d... | Medtronic, Inc. | A61B005/04
US06978169 | B1 | 2005-12-20 | Personal physiograph | 2002-04-04 | Disclosed are systems and methods which provide a relatively small, preferably easily transportable, interpretive biotelemetric monitor, such as may utilize platforms developed for use as personal digital assistants, adapted for use by non-health care personnel. A preferred embodiment biotelemetric monitor comprises a self-contained EKG monitor providing ICU cardiac monitoring with accurate real-time diagnostic heart rhythm analysis directly to the wearer. Preferred embodiment biotelemetric monitors implement multiple interactive screens to guide a user through a series of screens to help understand the monitored conditions. Moreover, the biotelemetric monitor may operate to analyze and interpret the monitored data to thereby present information with respect to the probable implications of the monitored condition and/or make recommendations which may be useful to the patient at that time. Preferably, communications functionality is included to facilitate communication between the biotelemetric monitor and a host system. | 1. A portable electrocardiograph monitor system comprising: a display providing for display of both textual and graphical information;a screen hierarchy for presenting information upon said display;a probe interface for receiving 3 leads of electrocardiograph data from a plurality of probes;a processor operable to accept data received via said probe interface and present at least a portion thereof upon said display, wherein a first screen of said screen hierarchy provides said at least a portion of said data presented upon said display; andan alarm for providing audible instructions to a patient relating to interpreted electrocardiograph data; andwherein said first screen further provides alarm information associated with said at least a portion of said electrocardiograph data displayed upon said display, wherein a second screen of said screen hierarchy provides alarm specific details associated with said alarm information, and wherein said second screen is accessible from said first screen by operation of an alarm button by a user. | 1. A portable electrocardiograph monitor system comprising: a display providing for display of both textual and graphical information;a screen hierarchy for presenting information upon said display;a probe interface for receiving 3 leads of electrocardiograph data from a plurality of probes;a processor operable to accept data received via said probe interface and present at least a portion thereof upon said display, wherein a first screen of said screen hierarchy provides said at least a portion of said data presented upon said display; andan alarm for providing audible instructions to a patient relating to interpreted electrocardiograph data; andwherein said first screen further provides alarm information associated with said at least a portion of said electrocardiograph data displayed upon said display, wherein a second screen of said screen hierarchy provides alarm specific details associated with said alarm information, and wherein said second screen is accessible from said first screen by operation of an alarm button by a user. 2. The system of claim 1, wherein said display comprises a touch screen element for accepting user input. 3. A portable electrocardiograph monitor system comprising: a display providing for display of both textual and graphical information;a screen hierarchy for presenting information upon said display;a probe interface for receiving 3 leads of electrocardiograph data from a plurality of probes;a processor operable to accept data received via said probe interface and present at least a portion thereof upon said display, wherein a first screen of said screen hierarchy provides said at least a portion of said data presented upon said display; andan alarm for providing audible instructions to a patient relating to interpreted electrocardiograph data; andan event recorder operable to record data received via said probe interface for later access wherein said event recorder is controlled to record an event upon input by a user, wherein said input by a user is provided by operation of an event recorder button available to said user throughout a plurality of screens of said screen hierarchy. 4. The system of claim 3, wherein said event recorder is controlled to record an event upon interpretation of said data received via said probe interface by said processor. 5. The system of claim 4, wherein said interpretation provides a determination that said data received via said probe interface indicates an irregularity. 6. The system of claim 4, wherein said interpretation provides a determination that said data received via said probe interface is abnormal as compared to a user's normal data as monitored by said system. 7. A portable electrocardiograph monitor system comprising: a display providing for display of both textual and graphical information;a screen hierarchy for presenting information upon said display;a probe interface for receiving 3 leads of electrocardiograph data from a plurality of probes;a processor operab... | - | A61B005/044
US06922580 | B2 | 2005-07-26 | Blood flow gated MRI | 2002-06-04 | A magnetic resonance imaging method and apparatus includes a navigator region defined within the subject by selective excitation. Blood flow is measured within the selected region using the principles of phase contrast MR angiography. A cardiac cycle plot is constructed from Fourier transformed data that represents measured velocity of blood flow through the navigator region as a function of time. On the basis of the cardiac cycle plot and the navigator measurements, data acquisition is synchronized or gated to portions of the cardiac cycle. | 1. A method of magnetic resonance comprising: disposing a region of interest and a region with flowing blood in a main magnetic field; applying at least one phase contrast navigator sequence to the flowing blood region to generate navigator echoes; determining a measurement of blood movement from the navigator echoes; gating an imaging sequence in accordance with the blood movement measurement. | 1. A method of magnetic resonance comprising: disposing a region of interest and a region with flowing blood in a main magnetic field; applying at least one phase contrast navigator sequence to the flowing blood region to generate navigator echoes; determining a measurement of blood movement from the navigator echoes; gating an imaging sequence in accordance with the blood movement measurement. 2. The method as set forth in claim 1, wherein the navigator sequence generates navigator echoes that indicate blood flow velocity. 3. The method as set forth in claim 1, wherein the selected region also includes non-blood tissue and further including: acquiring a navigator base line measurement indicative of a contribution of non-blood tissue to the navigator echo. 4. A method of magnetic resonance comprising: disposing a region of interest and a region with flowing blood in a main magnetic field; repeatedly applying phase contrast navigators over at least one cardiac cycle and determining a plurality of blood movement measurements over the cardiac cycle; determining a measurement of blood movement from the navigator echoes; generating a blood movement versus time plot indicative of the cardiac cycle from the blood flow measurements; gating an imaging sequence in accordance with the blood movement measurement. 5. The method as set forth in claim 4, further including: setting a blood flow measurement threshold relative to the cardiac cycle plot to select a portion of the cardiac cycle to be imaged. 6. The method as set forth in claim 4, further including: correlating the cardiac cycle plot and electra cardio-graph signals. 7. The method as set forth in claim 4, further including: obtaining first derivatives of the cardiac cycle plot. 8. The method as set forth in claim 7, wherein the applied navigator sequences generate navigator echoes that are velocity and acceleration encoded and further including: comparing the blood flow velocity and acceleration with amplitudes and derivatives of the cardiac cycle plot to determine points in the cardiac cycle. 9. The method as set forth in claim 8, further including: triggering the imaging sequence data acquisition in response to the determined points in the cardiac cycle. 10. A method of magnetic resonance comprising: disposing a region of interest and a region with flowing blood in a main magnetic field; applying at least one phase contrast navigator sequence to the flowing blood region to generate navigator echoes; determining a measurement of blood movement from the navigator echoes; applying at least one extrapolation algorithm to the cardiac cycle plot based on the first derivative of the cardiac cycle plot to determine a triggering time; and, triggering a data acquisition of the region of interest at the triggering time. 11. The method as set forth in claim 1 further including: adjusting the navigator sequence to adjust alignment of the flowing blood region with a blood vessel. 12. A magne... | Koninklijke Philips Electronics N.V. | A61B005/05
US06936015 | B2 | 2005-08-30 | Active guide wire and method of making the same | 2002-06-06 | An active guide wire (1) that can be used in a narrow blood vessel in a complicate vascular system such as intracerebral is disclosed as well as a method of making the same. The active wire comprises a bias coil spring (3); an outer tube (4) having the bias coil spring (3) sheathed therewith and providing an external wall for the active wire guide (1); and a shape memory alloy (SMA) actuator (5) in the form of a flat sheet anchored in and securely connected to the outer tube (4), wherein the flat sheet SMA actuator (5) is shape memorized so as to bend in a direction of the thickness of the flat sheet and has a shape running zigzag longitudinally thereof. The actuator (5) is energized with an electric current and thereby heated to have its leading end bent. | 1. An active guide wire comprising: a bias coil spring; an outer tube having the bias coil spring sheathed therewith and providing an external wall for the active guide wire; and a shape memory alloy (SMA) actuator in the form of a flat sheet anchored in and securely connected to said outer tube, wherein said flat sheet SMA actuator is shape memorized so as to bend in a direction of the thickness of said flat sheet and has a shape running zigzag longitudinally thereof. | 1. An active guide wire comprising: a bias coil spring; an outer tube having the bias coil spring sheathed therewith and providing an external wall for the active guide wire; and a shape memory alloy (SMA) actuator in the form of a flat sheet anchored in and securely connected to said outer tube, wherein said flat sheet SMA actuator is shape memorized so as to bend in a direction of the thickness of said flat sheet and has a shape running zigzag longitudinally thereof. 2. An active guide wire as set forth in claim 1 wherein said bias coil spring is composed of an electrically conductive elastic material. 3. An active guide wire as set forth in claim 1 wherein said outer tube is composed of a polymeric material. 4. An active guide wire as set forth in claim 1, further comprising an electrically conductive spacer whereby an end portion of said actuator and an end portion of said bias coil spring are brought together, a plurality of electrically nonconductive spacers whereby a plurality of intermediate regions of said actuator are fastened to a plurality of intermediate regions of said bias coil spring, respectively; a further electrically nonconductive spacer whereby a base portion of said actuator is fastened to a base portion of said bias coil spring; and a lead conductor connected to the base portion of said actuator. | Masayoshi Esashi Yoichi Haga | A61B005/00 A61M025/00
US06961605 | B2 | 2005-11-01 | Method for processing biomagnetic field data, magnetic field contour mapping, forming their waveforms and a biomagnetic instrument using the same | 2001-10-30 | A method for processing biomagnetic fields generated by biocurrents resulting from activities of human brain or myocardia and its mapping apparatus are provided, which features biomagnetic measurement and its analysis, magnetic field mapping and its imaging and their waveform generation by a simple operation. Biomagnetic fields emitted from the patient are measured at a plurality of measurement positions, and a contour map of magnetic field obtained as a result of processing of these measured biomagnetic fields is imaged in the magnetic contour map display apparatus, which display apparatus comprises: the process function display area indicating process function items including measurement; and the analysis data display area which displays the waveform together with a designated measurement time, the waveform being generated during measurement at least based on the measured biomagnetic fields and during the designated measurement period of time. | 1. A method of processing measured data of a magnetic field generated by a patient and displaying a magnetic field diagram or map by adjusting a bias current and an offset voltage Voff of each of a plurality of SQUID channels arranged in rows and columns, comprising the steps of: selecting a value of the bias current at which amplitude ΔV of a φ-v characteristic curve becomes maximum when the bias current is changed for each SQUID channel where V is a voltage corresponding to an external magnetic field when the external magnetic field is applied to each SQUID channel and φ is magnetic flux which expresses the external magnetic field; obtaining a maximum value Vmax and a minimum value Vmin of the voltage by changing the external magnetic field applied to each SQUID channel and finding a value of the offset voltage by calculating the equation Voff=(Vmax+Vmin)/2 for each SQUID channel; and displaying φ-V characteristic curves for all of the SQUID channels. | 1. A method of processing measured data of a magnetic field generated by a patient and displaying a magnetic field diagram or map by adjusting a bias current and an offset voltage Voff of each of a plurality of SQUID channels arranged in rows and columns, comprising the steps of: selecting a value of the bias current at which amplitude ΔV of a φ-v characteristic curve becomes maximum when the bias current is changed for each SQUID channel where V is a voltage corresponding to an external magnetic field when the external magnetic field is applied to each SQUID channel and φ is magnetic flux which expresses the external magnetic field; obtaining a maximum value Vmax and a minimum value Vmin of the voltage by changing the external magnetic field applied to each SQUID channel and finding a value of the offset voltage by calculating the equation Voff=(Vmax+Vmin)/2 for each SQUID channel; and displaying φ-V characteristic curves for all of the SQUID channels. 2. A method of processing measured data of a magnetic field generated by a patient and displaying a magnetic field diagram or map by adjusting a bias current and an offset voltage Voff of each of a plurality of SQUID channels arranged in rows and columns, comprising the steps of: selecting a value of the bias current at which amplitude ΔV of a φ-V characteristic curve becomes maximum when the bias current is changed for each SQUID channel where V is a voltage corresponding to an external magnetic field when the external magnetic field is applied to each SQUID channel and φis a magnetic flux which expresses the external magnetic field; obtaining a maximum value vmax and a minimum value Vmin of the voltage by changing the external magnetic field applied to each SQUID channel and finding a value of the offset voltage by calculating the equation Voff=(Vmax+Vmin)/2 for SQUID each channel to detect a SQUID channel which is in a state unsuitable for measurement, and thereafter, displaying single-waveforms for all the SQUID channels. 3. A method of measuring a magnetic field generated by a heart of a patient at a plurality of positions, comprising the steps of: displaying the plurality of measurement positions on a display screen; selecting a measurement position to be displayed out of the plurality of measurement positions; distinctively displaying in different colors individual pieces of information regarding strengths of the measured magnetic field at the plurality of measurement positions according to the strengths of the measured magnetic field so as to be capable of identifying the individual pieces of information displayed in different colors as magnetic field strengths in corresponding relation to color-coded magnetic field strengths in a magnetic field strength index frame; and displaying the plurality of measurement positions along with the presence or absence of selection of the measurement position to be displayed, with the individual pieces of information displayed. 4. A method of mea... | Hitachi, Ltd. | A61B005/05
US06878121 | B2 | 2005-04-12 | Sleep scoring apparatus and method | 2002-11-01 | A method and an apparatus are disclosed for scoring the quality of sleep experienced by a subject wearing the apparatus for a specified period. This sleep monitoring apparatus includes a motion sensor that quantifies the temporal variation in the subject's motion, an analog to digital converter to sample the sensor data, a microprocessor with embedded programmable memory to store control and processing firmware, a memory device for data storage, a display means, and a switch for staring and resetting the device. The firmware for this apparatus: (i) directs the microprocessor to sample sensor output at a prescribed time interval and to temporarily store this data in the memory device, (ii) directs the microprocessor to analyze the temporarily stored data to compute a sleep score based on the magnitude of movements recorded during the monitoring period, and (iii) controls the operation of the display means. Upon termination of the sleep period, the apparatus' display means will display a sleep score based on a scale of 0 to 100. A high score indicates restful sleep (very low movements), while a low score indicates excessive restlessness during the monitored sleep period. | 1. A sleep monitoring apparatus for scoring the quality of sleep experienced by a subject wearing said apparatus for a specified period, said sleep monitoring apparatus comprising: a means for sensing the temporal variation in the motion at the location where said apparatus is being worn, an analog to digital converter connected to said sensing means for sampling the data sensed by said motion sensing means, a microprocessor connected to said converter, said microprocessor having embedded programmable memory, a means connected to said microprocessor for storing data, a display means connected to said microprocessor, and firmware that controls the operation of said apparatus by directing the microprocessor to sample the output of said sensing means at a prescribed time interval and to temporarily store said sampled data in said data storage means, by directing the microprocessor to analyze the temporarily stored data to compute a sleep score based on the magnitude of movements recorded during said monitoring period, and by controlling the operation of said display means to display said sleep score. | 1. A sleep monitoring apparatus for scoring the quality of sleep experienced by a subject wearing said apparatus for a specified period, said sleep monitoring apparatus comprising: a means for sensing the temporal variation in the motion at the location where said apparatus is being worn, an analog to digital converter connected to said sensing means for sampling the data sensed by said motion sensing means, a microprocessor connected to said converter, said microprocessor having embedded programmable memory, a means connected to said microprocessor for storing data, a display means connected to said microprocessor, and firmware that controls the operation of said apparatus by directing the microprocessor to sample the output of said sensing means at a prescribed time interval and to temporarily store said sampled data in said data storage means, by directing the microprocessor to analyze the temporarily stored data to compute a sleep score based on the magnitude of movements recorded during said monitoring period, and by controlling the operation of said display means to display said sleep score. 2. A sleep monitoring apparatus as recited in claim 1, further comprising a means connected to said microprocessor for resetting and starting the operation of said apparatus. 3. A sleep monitoring apparatus as recited in claim 1, wherein said apparatus is adapted to be worn on a subject's arm. 4. A sleep monitoring apparatus as recited in claim 2, wherein said apparatus is adapted to be worn on a subject's arm. 5. A sleep monitoring apparatus as recited in claim 1, wherein said firmware controls sampling of said output of said motion sensing means so as to sample at a sampling frequency that encompasses the frequency range of movements expected at the location where said apparatus is worn. 6. A sleep monitoring apparatus as recited in claim 4, wherein said firmware controls sampling of said output of said motion sensing means so as to sample at a sampling frequency that encompasses the frequency range of movements expected at the location where said apparatus is worn. 7. A sleep monitoring apparatus as recited in claim 1, wherein said firmware controls computation of said sleep score by utilizing an algorithm that relates movement counts per specified time period to one of a specified number of levels that are defined as comprising said sleep score. 8. A sleep monitoring apparatus as recited in claim 6, wherein said firmware controls computation of said sleep score by utilizing an algorithm that relates movement counts per specified time period to one of a specified number of levels that are defined as comprising said sleep score. 9. A sleep monitoring apparatus as recited in claim 1, wherein said motion sensing means is a piezoelectric motion sensor having signal conditioning and filtering circuitry is to yield specified levels of signal fidelity. 10. A sleep monitoring apparatus as recited in claim 8, wherein said motion sensing means is a p... | - | A61B005/103 A61B005/117
US06969507 | B2 | 2005-11-29 | Imaging sexual response | 2002-11-08 | The present invention relates to methods for contrast-enhanced imaging of sexual response. In particular, the present invention relates to methods of magnetic resonance imaging (MRI), computerized X-ray tomography (CT), ultrasound, and optical imaging using contrast agents to measure blood flow during sexual response, especially female sexual response. The invention relates to methods of analyzing normal sexual function and sexual dysfunction and provides a means for evaluating and screening potential therapeutic compounds for enhancing sexual function or treating sexual dysfunction. | 1. A method for assessing the effect of a compound on female sexual arousal disorder or female orgasmic disorder comprising: (a) administering a magnetic resonance imaging contrast agent to a female subject; (b) subjecting the subject to magnetic resonance imaging; (c) measuring an imaging signal characteristic of the contrast agent to establish a baseline; (d) subjecting the female subject to a stimulus designed to cause sexual arousal; (e) monitoring an imaging signal characteristic of the contrast agent to assess the female sexual arousal disorder or female orgasmic disorder after the stimulus; (f) administering a compound hypothesized to produce an effect on the female sexual arousal disorder of female orgasmic disorder to the female subject; and (g) determining if the female sexual arousal disorder or female orgasmic disorder has been affected as a result of administration of the compound. | 1. A method for assessing the effect of a compound on female sexual arousal disorder or female orgasmic disorder comprising: (a) administering a magnetic resonance imaging contrast agent to a female subject; (b) subjecting the subject to magnetic resonance imaging; (c) measuring an imaging signal characteristic of the contrast agent to establish a baseline; (d) subjecting the female subject to a stimulus designed to cause sexual arousal; (e) monitoring an imaging signal characteristic of the contrast agent to assess the female sexual arousal disorder or female orgasmic disorder after the stimulus; (f) administering a compound hypothesized to produce an effect on the female sexual arousal disorder of female orgasmic disorder to the female subject; and (g) determining if the female sexual arousal disorder or female orgasmic disorder has been affected as a result of administration of the compound. 2. The method according to claim 1, wherein the magnetic resonance imaging contrast agent is MS-325. 3. The method according to claim 1, wherein the compound hypothesized to produce an effect on the female sexual arousal disorder or female orgasmic disorder is selected from the group consisting of sildenafil citrate, tadalafil, and vardenafil. 4. The method according to claim 1 wherein the subject is a mammal. 5. The method according to claim 4 wherein the subject is a human. 6. The method according to claim 1, wherein said determining step comprises: (h) subjecting the female subject after said compound administration to a stimulus designed to cause sexual arousal; (i) monitoring an imaging signal characteristic of the contrast agent to assess the female sexual arousal disorder or female orgasmic disorder after the stimulus; and (j) comparing the imaging signal characteristics obtained in step (e) and step (i) to assess the effect of the compound on the female sexual arousal disorder or female orgasmic disorder. 7. The method of claim 1 or claim 6, wherein the imaging signal characteristic of the contrast agent is monitored in the pelvis of said female subject. 8. The method of claim 1 or claim 6, wherein the imaging signal characteristic of the contrast agent that is monitored includes monitoring the number of pixels in a given region of interest. 9. The method of claim 1 or claim 6, wherein the imaging signal characteristic of the contrast agent that is monitored includes monitoring the blood volume in a given region of interest. 10. The method of claim 1 or claim 6, wherein the stimulus designed to cause sexual arousal is selected from somatosensory, audiovisual, physical, and pharmacological stimuli. 11. The method of claim 1, wherein said female subject is pre-menopausal. 12. The method of claim 1, wherein said female subject is post-menopausal. 13. The method of claim 1, wherein said compound is a PDE5 inhibitor. 14. The method of claim 8, wherein said region of interest is the genitalia. 15. The method of claim 9, wherein said regio... | Epix Pharmaceuticals, Inc. | A61K049/00 A61B005/055
US06865407 | B2 | 2005-03-08 | Calibration technique for non-invasive medical devices | 2002-07-11 | There is a need for a non-invasive method of calibrating medical devices at the point of care, where the calibration is performed without the removal of blood or bodily fluids. The invention is directed to an approach for calibrating a first non-invasive sensor in which the tissue being measured is modulated in some way so as to after the value of the parameter being measured by the first optical sensor. A second sensor detects another parameter that also changes with the modulation. The second sensor is absolutely calibrated. Where there is a known relationship between the first and second parameters, a calibration may be derived for the first sensor. Such a technique is applicable to calibrating non-invasive sensors for monitoring a wide variety of physiologic parameters including, inter alia, glucose, blood gases, blood electrolytes and blood pH. | 1. A method of calibrating a non-invasive sensor for a first physiologic parameter in tissue, comprising: (a) non-invasively measuring first values of the first physiologic parameter and of a second physiologic parameter in the tissue, a known relationship existing between the first and second physiologic parameters; (b) acting on the tissue so as to change the first and second physiologic parameters; (c) non-invasively measuring second values of the first physiologic parameter and of the second physiologic parameter in the tissue, the measurements of the first and second values of the second physiologic parameter being absolutely calibrated; and (d) determining a calibrated value for at least one of the first and second values of the first physiologic parameter from the first and second values of both the first and second physiologic parameters; wherein non-invasively measuring the first values of the first physiologic parameter includes exciting the tissue with excitation light and detecting a resulting fluorescent signal. | 1. A method of calibrating a non-invasive sensor for a first physiologic parameter in tissue, comprising: (a) non-invasively measuring first values of the first physiologic parameter and of a second physiologic parameter in the tissue, a known relationship existing between the first and second physiologic parameters; (b) acting on the tissue so as to change the first and second physiologic parameters; (c) non-invasively measuring second values of the first physiologic parameter and of the second physiologic parameter in the tissue, the measurements of the first and second values of the second physiologic parameter being absolutely calibrated; and (d) determining a calibrated value for at least one of the first and second values of the first physiologic parameter from the first and second values of both the first and second physiologic parameters; wherein non-invasively measuring the first values of the first physiologic parameter includes exciting the tissue with excitation light and detecting a resulting fluorescent signal. 2. A method as recited in claim 1, wherein non-invasively measuring the first and second values of the first parameter include measuring light received from the tissue. 3. A method as recited in claim 1, wherein non-invasively measuring the first and second values of the second parameter include measuring light received from the tissue. 4. A method as recited in claim 1, wherein the values for the first and second physiologic parameters are measured in the same region of the tissue. 5. A method as recited in claim 1, wherein acting on the tissue includes applying a systemic stimulus to the tissue. 6. A method as recited in claim 5, wherein applying a systemic stimulus to the tissue includes changing a metabolic rate of the tissue. 7. A method as recited in claim 1, wherein acting on the tissue includes applying a stimulus to the tissue localized to a region of the tissue where the first and second physiologic parameters are non-invasively measured. 8. A method as recited in claim 1, wherein acting on the tissue includes applying pressure to a region of the tissue where the first and second physiologic parameters are measured. 9. A method as recited in claim 1, wherein acting on the body includes changing temperature of a region of the body where the first and second physiologic parameters are measured. 10. The method of claim 1 wherein non-invasively measuring the first values of the first physiologic parameter further includes measuring tight absorbed by the tissue. 11. The method of claim 1 wherein non-invasively measuring the first values of the first physiologic parameter further includes measuring light reflected by the tissue. 12. A method of calibrating a non-invasive sensor for a first physiologic parameter in tissue of a patient, comprising: (a) non-invasively measuring first values of the first physiologic parameter and of a second physiologic parameter in the tissue, a known relationship existing betwee... | Optical Sensors, Inc. | A61B005/00
US06954663 | B2 | 2005-10-11 | Continuous wave optical imaging assuming a scatter-law | 2003-01-07 | There is provided a method and a system for determining the concentration of chromophores and reconstructing images in turbid media, such as animal tissues, using a continuous wave optical approach. In particular the approach is based on measurements of attenuation signals and the calculation of concentrations of chromophores using a predetermined scatter law. The system comprises a continuous wave photon migration model calculator coupled to an optical source and detector for estimating concentrations of chromophores and scatter parameters used in image reconstruction. | 1. A method for determining concentration of one or more chromophores in a turbid medium the method comprising: i) measuring an attenuation signal using continuous wave (CW) illumination at a predetermined number of wavelengths; ii) providing a scatter law in which scatter is a function of wavelength; iii) estimating scatter parameters of said provided scatter law and a concentration for the one or more chromophores, said chromophores having a predetermined relationship between chromophore absorption and concentration as a function of wavelength; iv) calculating attenuation values at said predetermined wavelengths using the estimated concentrations and scatter parameters; v) adjusting the concentration for the one or more chromophores and the scatter parameters until the calculated attenuation values and the measured attenuation differ by less than a predetermined value thereby determining said concentration within said medium; and wherein said predetermined number of wavelengths is sufficient to provide a desired degree of accuracy in said determination of said concentration. | 1. A method for determining concentration of one or more chromophores in a turbid medium the method comprising: i) measuring an attenuation signal using continuous wave (CW) illumination at a predetermined number of wavelengths; ii) providing a scatter law in which scatter is a function of wavelength; iii) estimating scatter parameters of said provided scatter law and a concentration for the one or more chromophores, said chromophores having a predetermined relationship between chromophore absorption and concentration as a function of wavelength; iv) calculating attenuation values at said predetermined wavelengths using the estimated concentrations and scatter parameters; v) adjusting the concentration for the one or more chromophores and the scatter parameters until the calculated attenuation values and the measured attenuation differ by less than a predetermined value thereby determining said concentration within said medium; and wherein said predetermined number of wavelengths is sufficient to provide a desired degree of accuracy in said determination of said concentration. 2. The method as claimed in claim 1 wherein said scatter parameters are maintained constant during said step of adjusting said concentration of said one or more chromophores. 3. The method as claimed in claim 1 wherein the general form of the scatter law is given by Aλ−B. 4. The method as claimed in claim 1 wherein said estimations of said concentration of one or more chromophores and said scatter parameters are based on known concentrations and scatter parameters in similar turbid media. 5. The method as claimed in claim 1, wherein the estimation of the concentrations and the scatter parameters is based on a comparison of the measured attenuation function with previously calculated attenuations functions for turbid media substantially identical to said turbid medium. 6. The method as claimed in claim 1 wherein the concentration of the one or more chromophores is inhomogeneous within said turbid medium and wherein the attenuation values are calculated using a photon diffusion equation. 7. A method for optical imaging of a region of interest in a turbid medium, said method comprising: i) injecting light at a predetermined number of wavelengths into said tissue at one or more injection ports; ii) detecting said light at one or more detection ports to obtain a measured attenuation function; iii) providing a scatter law; iv) estimating a concentration for one or more chromophores and scatter parameters at a plurality of voxels in said region of interest; v) calculating attenuation values at said predetermined number of wavelengths using the estimated concentration and scatter parameters and a photon diffusion equation to generate a calculated attenuation function; vi) adjusting said concentration and said scatter parameters until said calculated attenuation function and said measured attenuation function differ by less than a predetermined value; vii) generating an image... | ART Advanced Research Technologies Inc. | A61B005/00
US06912416 | B2 | 2005-06-28 | Detecting or preventing tissue damage | 2002-07-18 | Systems and methods are described for, among other things, detecting tissue injury and for evaluating the integrity of an anatomic structure. A system may include a first sensor disposable through a first space on a first side of the anatomic structure, a second sensor disposed in a second space on a second side of the anatomic structure, and a processor coupled to the first sensor and to the second sensor, the processor including a meter, the meter responsive to a communication between the first sensor and the second sensor, wherein an invasion of the anatomic structure by the first sensor alters the communication between the first sensor and the second sensor. | 1. A system for evaluating the integrity of an anatomic structure, comprising: a first sensor component comprising a surgical tool including a needle, the tool being attached to a length of supporting material, the first sensor component being disposable though a first space on a first side of the anatomic structure; a second sensor component disposable in a second space on a second side of the anatomic structure; and a processor coupled to the first sensor component and to the second sensor component the processor including a meter, the meter responsive to a communication between the first sensor component and the second sensor component; wherein an invasion of the anatomic structure by the first sensor component alters the communication between the first sensor component and the second sensor component. | 1. A system for evaluating the integrity of an anatomic structure, comprising: a first sensor component comprising a surgical tool including a needle, the tool being attached to a length of supporting material, the first sensor component being disposable though a first space on a first side of the anatomic structure; a second sensor component disposable in a second space on a second side of the anatomic structure; and a processor coupled to the first sensor component and to the second sensor component the processor including a meter, the meter responsive to a communication between the first sensor component and the second sensor component; wherein an invasion of the anatomic structure by the first sensor component alters the communication between the first sensor component and the second sensor component. 2. The system of claim 1, further comprising a conductive medium adapted to be instilled into the second space and coupled to the second sensor component. 3. The system of claim 2, wherein the conductive medium is adapted to coat the second side of the anatomic structure. 4. The system of claim 1, wherein the processor provides an output signal indicative of the integrity of the anatomic structure. 5. The system of claim 1, wherein the second sensor component comprises a catheter conductor. 6. The system of claim 5, wherein the catheter conductor comprises at least one of a Foley catheter and a catheter guide. 7. The system of claim 1, wherein the first sensor component and the second sensor component are coupled electrically to the processor, and the communication between the first sensor component and the second sensor component comprises an electrical conduction. 8. The system of claim 7, wherein the meter comprises at least one of an ammeter, a voltmeter, a resistometer, and an electrical conductometer. 9. The system of claim 1, wherein the atomic structure comprises a wall of a urinary bladder, the first space comprises the Space of Retzius, and the second space comprises a lumen of the urinary bladder. 10. The system of claim 1, further comprising a power source providing power to at least one of the processor, the first sensor component and the second sensor component. 11. The system of claim 10, wherein the power source comprises a battery. 12. The system of claim 1, wherein at least one part of the system is disposable. 13. The system of claim 1, wherein at least one part of the system is reusable. 14. An electrical tissue injury detection system, comprising: a first conductor comprising a surgical tool including a needle, the tool being attached to a length of of supporting material, the first conductor being disposable through a first space on a first side of an anatomic structure; a second conductor disposable in a second space on a second side of the anatomic structure; a conductive medium adapted to be instilled into the second space and in communication with the second conductor; and a processor coupled to the first ... | - | A61B005/05
US06926668 | B2 | 2005-08-09 | System and method for analyzing normalized patient voice feedback in an automated collection and analysis patient care system | 2002-09-20 | A system and method for analyzing normalized patient voice feedback in an automated collection and analysis patient care system is described. Device measures providing physiological measures collected by an implantable medical device on a substantially continuous basis are received for storage into a patient care record. Voice feedback spoken by an individual patient substantially contemporaneous to the collection of at least one set of the device measures is received. The voice feedback is processed into normalized quality of life measures for storage into the patient care record. The physiological measures and the quality of life measures stored in the patient care record are analyzed relative to at least one of other physiological measures and other quality of life measures to determine a patient status indicator. | 1. An analysis system for providing patient physiological and voice feedback data for use in automated patient care, comprising: a database to store a set of collected device measures and voice feedback, the collected device measures set providing quantitative health care data indicators regularly recorded by a medical device for a patient wider automated patient care, the voice feedback spoken by the patient substantially contemporaneous to the collection of the collected device measures; and a processor to retrieve the collected device measures set and the voice feedback, to analyze the collected device measures set through derived measure determination and statistical value calculation into indicators of patient wellness, and to process the voice feedback into normalized quality of life measures for inclusion with the indicator, of patient wellness. | 1. An analysis system for providing patient physiological and voice feedback data for use in automated patient care, comprising: a database to store a set of collected device measures and voice feedback, the collected device measures set providing quantitative health care data indicators regularly recorded by a medical device for a patient wider automated patient care, the voice feedback spoken by the patient substantially contemporaneous to the collection of the collected device measures; and a processor to retrieve the collected device measures set and the voice feedback, to analyze the collected device measures set through derived measure determination and statistical value calculation into indicators of patient wellness, and to process the voice feedback into normalized quality of life measures for inclusion with the indicator, of patient wellness. 2. An analysis system according to claim 1, wherein at least one of the collected device measures and the voice feedback is respectively compared to at least one of a collected device measures and voice feedback stored for the patient, a patient peer group, and a patient population. 3. An analysis system according to claim 1, wherein the spoken voice feedback is requested through pre-determined prompts, each such prompt being associated with a quality of life measure and parsing the spoken voice feedback in accordance with a voice grammar and speech vocabulary. | Cardiac Intelligence Corporation | A61B005/04
US06925324 | B2 | 2005-08-02 | System and device for multi-scale analysis and representation of physiological data | 2002-04-17 | System comprised of a medical device and method for analyzing physiological and health data and representing the most significant parameters at different levels of detail which are understandable to a lay person and a medical professional. Low, intermediate and high-resolution scales can exchange information between each other for improving the analyses; the scales can be defined according to the corresponding software and hardware resources. A low-resolution Scale I represents a small number of primary elements such as intervals between the heart beats, duration of electrocardiographic PQ, QRS, and QT-intervals, amplitudes of P-, Q-, R-, S-, and T-waves. This real-time analysis is implemented in a portable device that requires minimum computational resources. The set of primary elements and their search criteria can be adjusted using intermediate or high-resolution levels. At the intermediate-resolution Scale II, serial changes in each of the said elements can be determined using a mathematical decomposition into series of basis functions and their coefficients. This scale can be implemented using a specialized processor or a computer organizer. At the high-resolution Scale III, combined serial changes in all primary elements can be determined to provide complete information about the dynamics of the signal. This scale can be implemented using a powerful processor, a network of computers or the Internet. The system can be used for personal or group self-evaluation, emergency or routine ECG analysis, or continuous event, stress-test or bed-side monitoring. | 1. A method for analysis of physiological or health data in at least two levels of detail, said method comprising: analyzing a plurality of primary elements in said data in first scale, low level resolution to detect one-time changes in such primary elements and thereby identify abnormal or unstable primary elements; comparing said primary elements with reference values for such primary elements, and representing deviation of primary elements from the reference values in a form understandable to a lay person; analyzing at least some of the same said primary elements in said data in second scale, higher level resolution using mathematical decomposition to provide detailed characterization of serial changes in said abnormal or unstable primary elements; and exchanging information between said analyzing in said first and second levels of resolution to improve at least one of said first and second analyses. | 1. A method for analysis of physiological or health data in at least two levels of detail, said method comprising: analyzing a plurality of primary elements in said data in first scale, low level resolution to detect one-time changes in such primary elements and thereby identify abnormal or unstable primary elements; comparing said primary elements with reference values for such primary elements, and representing deviation of primary elements from the reference values in a form understandable to a lay person; analyzing at least some of the same said primary elements in said data in second scale, higher level resolution using mathematical decomposition to provide detailed characterization of serial changes in said abnormal or unstable primary elements; and exchanging information between said analyzing in said first and second levels of resolution to improve at least one of said first and second analyses. 2. A method as set forth in claim 1 in which said analyzing in a second scale, higher level resolution uses orthogonal mathematical decomposition. 3. A method as set forth in claim 2 in which said analyzing in said second scale, higher resolution further includes: combining a plurality of said serial changes in the abnormal or unstable primary elements and analyzing the combined changes using orthogonal mathematical decomposition; and selecting a combination of parameters for tracking the changes in said abnormal or unstable elements and sending information respecting the selected combination to said low-resolution analysis for adjustment of monitoring parameters. 4. A method as set forth in claim 2 in which said analyzing in said second scale, higher resolution comprises: forming a time series from said serial changes; characterizing said time series using Principal Component Analysis (PCA) and generating PCA-coefficients indicative of both linear and nonlinear changes in the individual pattern; and determining the magnitude of said linear and nonlinear changes by using time varying mean and variance of said PCA-coefficients and determining the complexity of said linear and nonlinear changes by calculating the number of PCA-coefficients that exhibit substantially simultaneous changes. 5. A method as set forth in claim 2 in which said analyzing in said second scale, higher resolution comprises: forming time series from said serial changes; characterizing said time series using Principal Component Analysis (PCA) and generating PCA-coefficients indicative of both linear and nonlinear changes in the individual pattern; presenting said PCA-coefficients as a vector in n-dimensional space wherein n is the number of variables analyzed; and determining the magnitude and direction of the combined changes in the characteristic variables by the changes in the magnitude and direction of said n-dimensional vector. 6. A method as set forth in claim 2 which includes adjusting said analyzing in low resolution based on said detailed characterization of se... | - | A61B005/044
US06950699 | B1 | 2005-09-27 | Water content probe | 2001-12-12 | A method and system to determine brain stiffness is disclosed. A probe to measure tissue water content is inserted through an aperture (burr hole) in the cranium into brain tissue. The probe has two electrically separated plate conductors with a dielectric which forms a capacitor plane. One conductor has a surface mount resistor to allow exact impedance matching to the core of a coaxial cable. The other conductor attaches electrically to the shield of the coaxial cable. The probe is stabilized in the brain tissue through a plastic ventriculostomy bolt which has been secured by screw tapping into the cranium. The coaxial cable connects to a spectrum analyzer. Brain water content and blood congestion alter the resonant frequency of the probe, allowing a realtime readout of apparent tissue water content. By monitoring the momentary shift in center resonant frequency or, alternatively, the standing wave ratio slightly off resonant frequency, a beat-to-beat pulsatile waveform is derived relating to the perfusion of the brain. A strain gauge intracranial pressure sensor (ICP) is separately affixed through the bolt and adjacent to the water content probe. By comparing the phase angle or lag time difference between the pressure tracing and the perfusion tracing, a realtime measurement of organ stiffness or compliance is derived. | 1. A probe for measuring tissue water content in a region of interest in the brain, the probe comprising: an implantable tissue water content sensor having two plates with a proximal and distal end, the two plates being separated by a dielectric material and the distal end being implantable in brain tissue;an impedance matching circuit coupled to the proximal end of one of the plates;a first output terminal coupled to the matching circuit and a second output terminal coupled to one of the plates;a remotely positioned frequency spectrum analyzer receiving an output signal from the first and second output terminals; anda digital computer having a display, the digital computer having an input coupled to the output signal from the water content probe and the spectrum analyzer, the computer programmed to display the resonant frequency of the sensor indicative of water content in the brain tissue. | 1. A probe for measuring tissue water content in a region of interest in the brain, the probe comprising: an implantable tissue water content sensor having two plates with a proximal and distal end, the two plates being separated by a dielectric material and the distal end being implantable in brain tissue;an impedance matching circuit coupled to the proximal end of one of the plates;a first output terminal coupled to the matching circuit and a second output terminal coupled to one of the plates;a remotely positioned frequency spectrum analyzer receiving an output signal from the first and second output terminals; anda digital computer having a display, the digital computer having an input coupled to the output signal from the water content probe and the spectrum analyzer, the computer programmed to display the resonant frequency of the sensor indicative of water content in the brain tissue. 2. The probe of claim 1 wherein the two plates are coated with insulation material sufficient to provide DC isolation. 3. The probe of claim 1 wherein the impedance matching circuit includes a resistor. 4. The probe of claim 1 further comprising a coaxial cable having a core conductor coupled to the impedance matching circuit and a circumferential conductor coupled to the proximal end of the other plate, the coaxial cable being coupled to the spectrum analyzer. 5. The probe of claim 1 wherein the plates and the dielectric material have a series of transverse holes. 6. The probe of claim 1 further comprising an intracranial pressure sensor located in substantially parallel orientation with the water content sensor and reading the pressure of the region of interest. 7. The probe of claim 6 further comprising: an analog to digital converter having an output and an input coupled to the intracranial pressure sensor; andwherein the computer is coupled to the output of the analog to digital converter and is programmed to display simultaneous tracings of apparent water content pulsatility due to tissue perfusion and compression based on the signal from the spectrum analyzer and the intracranial pressure waveform. 8. The probe of claim 7 wherein the pressure sensor is a tissue-implanted strain gauge. 9. The probe of claim 7 wherein the pressure sensor is a fiberoptic sensor. 10. The probe of claim 7 further comprising: a wireless transmitter coupled to the intracranial sensor and the water content sensor; anda wireless receiver coupled to the digital computer, the receiver tuned to signals from the transmitter. 11. The probe of claim 10 wherein the impedance matching and transmitter circuit components are an implantable component integrated circuit of the sensor probe. 12. The probe of claim 7 wherein the digital computer determines apparent water content pulsatility due to tissue perfusion and compression by plotting the change in standing wave ratio to the side of the return loss curve on the spectrum analyzer and determines where the standing wave ratio ... | Brain Child Foundation | A61B005/05
US06859663 | B2 | 2005-02-22 | Myoelectric-pattern classification method and apparatus | 2002-01-29 | A myoelectric pattern that is an action potential generated by a coordinated action of a plurality of muscles is measured by a group of surface electrodes on a skin surface, and a feature pattern is extracted from the measured myoelectric pattern. Redundant coding is used to encode the extracted feature pattern into a bit pattern, the coded bit pattern is subjected to pattern classification, and output control signal is generated. | 1. A myoelectric-pattern classification method comprising using a group of surface electrodes on a skin surface to measure a myoelectric pattern that is an action potential generated by a coordinated action of a plurality of muscles, extracting a feature pattern from the measured myoelectric pattern, using redundant coding to encode the extracted feature pattern into a bit pattern, performing pattern classification on the encoded bit pattern and generating an output control signal. | 1. A myoelectric-pattern classification method comprising using a group of surface electrodes on a skin surface to measure a myoelectric pattern that is an action potential generated by a coordinated action of a plurality of muscles, extracting a feature pattern from the measured myoelectric pattern, using redundant coding to encode the extracted feature pattern into a bit pattern, performing pattern classification on the encoded bit pattern and generating an output control signal. 2. A myoelectric-pattern classification apparatus comprising a group of surface electrodes that measures on a C. skin surface a myoelectric pattern that is an action potential generated by a coordinated action of a plurality of muscles, a feature-pattern extraction apparatus that extracts a feature pattern from the measured myoelectric pattern, an encoder that uses redundant coding to encode the extracted feature pattern into a bit pattern, and a pattern classifier that classifies the encoded bit pattern and generates an output control signal. | National Institute of Advanced Industrial Science and Technology Isamu Kajitani Tetsuya Higuchi | A61B005/04
US06942621 | B2 | 2005-09-13 | Method and apparatus for detecting weak physiological signals | 2002-07-11 | A method and apparatus are provided for measuring a weak physiological signal in the presence of stronger signals. The method and apparatus utilize motion sensing devices placed on or about the patient to obtain a motion data set from each motion sensing device describing the aggregate detected motion. The data sets are synchronized and then used to solve a set of simultaneous equations which allows for the separation of the various motion components which may include such components of interest as fetal heart motion. | 1. A method for detecting a physiological signal of interest comprising: disposing two or more motion sensing devices on or about a patient; calibrating the two or more motion sensors such that a synchronization adjustment is determined for each motion sensor; collecting a data set from each motion sensor; synchronizing the data sets using the synchronization adjustment; and using the synchronized data sets to solve a set of simultaneous equations in which one variable represents the physiological signal of interest. | 1. A method for detecting a physiological signal of interest comprising: disposing two or more motion sensing devices on or about a patient; calibrating the two or more motion sensors such that a synchronization adjustment is determined for each motion sensor; collecting a data set from each motion sensor; synchronizing the data sets using the synchronization adjustment; and using the synchronized data sets to solve a set of simultaneous equations in which one variable represents the physiological signal of interest. 2. The method of claim 1, wherein the physiological signal of interest is fetal heart motion. 3. The method of claim 2, wherein disposing a plurality of motion sensing devices comprises placing a pad to which motion sensors have been associated upon the patient. 4. The method of claim 2 further comprising determining at least one of a fetal heart rate and a fetal heart waveform. 5. The method of claim 1, wherein the plurality of motion sensing devices is at least two motion sensing devices. 6. The method of claim 1, wherein solving the set of simultaneous equations comprises performing a tensor analysis upon the plurality of data sets. 7. The method of claim 1, further comprising checking the validity of the solution to the simultaneous equations. 8. The method of claim 7, wherein checking the validity comprises analyzing a non-periodic component of the plurality of data sets. 9. A method for monitoring fetal heart motion comprising: collecting a motion data set from each of two or more motion sensing devices disposed on or about a patient; synchronizing the motion data sets; and solving a set of simultaneous equations to determine two or more motion components of the synchronized motion data sets related to fetal heart motion. 10. The method of claim 9, wherein one of the determined motion components is attributable to a fetal heartbeat. 11. The method of claim 10, further comprising determining a fetal heart rate from the fetal heartbeat component. 12. The method of claim 10, further comprising determining a fetal heart waveform from the fetal heartbeat component. 13. The method of claim 9, wherein the number of simultaneous equations is equal to the number of synchronized motion data sets and is greater than or equal to the number of motion components. 14. The method of claim 9, further comprising disposing the two or more motion sensing devices upon a pad. 15. The method of claim 9, wherein solving the set of simultaneous equations comprises performing a tensor analysis upon the synchronized motion data sets. 16. The method of claim 9, further comprising checking the validity of the solution to the simultaneous equations. 17. The method of claim 16, wherein checking the validity comprises analyzing a non-periodic component of the synchronized motion data sets. 18. A physiological monitoring system for monitoring a physiological signal of interest comprising: a sensor array comprising two or more motion sensitiv... | GE Medical Systems Information Technologies, Inc. | A61B005/02 A61B005/11
US06928317 | B2 | 2005-08-09 | Heart rate transmitter | 2002-12-12 | A heart rate transmitter comprising a strap and a body. The strap is in long belt shape, a hole is in the center of the strap, a conducting rubber plate each is on both ends of the hole, a contacting tip stretches out from the conducting rubber plate to the inner side of the hole. The body is installed inside the hole of the strap; two contacting points are on both ends of the body. When the body is inserted into the hole of the strap, two contacting points are pressed to contact the contacting tips of the conducting rubber plate inside the hole of the strap; the conducting rubber plate contacts human body to measure the heartbeat rate. | 1. A heart rate transmitter comprising: a strap shaped as a single long belt including a hole in the center of said strap, said strap further comprising at least one conducting rubber plate extending along a length of said strap from said hole, and a contacting tip of said conducting rubber plate that stretches out from said conducting rubber plate adjacent to the inner side of said hole; a body installed inside said hole of said strap, said body comprising at least one electrical contacting point at a periphery of said body to connect said contacting tip of said conducting rubber plate when said body is inserted into said hole of said strap; and a transmitter device comprised of said body. | 1. A heart rate transmitter comprising: a strap shaped as a single long belt including a hole in the center of said strap, said strap further comprising at least one conducting rubber plate extending along a length of said strap from said hole, and a contacting tip of said conducting rubber plate that stretches out from said conducting rubber plate adjacent to the inner side of said hole; a body installed inside said hole of said strap, said body comprising at least one electrical contacting point at a periphery of said body to connect said contacting tip of said conducting rubber plate when said body is inserted into said hole of said strap; and a transmitter device comprised of said body. 2. The heart rate transmitter according to claim 1, wherein said contacting tip of said conducting rubber plate extends generally perpendicular to a length of said conducting rubber plate. 3. The heart rate transmitter according to claim 2, wherein said contacting tip of said conducting rubber plate is inset from an extreme end of said conducting rubber plate. 4. The heart rate transmitter according to claim 2, wherein said contacting tip of said conducting rubber plate is not within said hole. 5. The heart rate transmitter according to claim 1, wherein said contacting tip of said conducting rubber plate extends generally perpendicular to a length of said strap. 6. The heart rate transmitter according to claim 1, wherein said strap is integral. 7. The heart rate transmitter according to claim 1, comprising two straps. | - | A61B005/0402
US06925322 | B2 | 2005-08-02 | Optical MRI catheter system | 2002-07-25 | An assembly for delivering optical signals that has a nuclear magnetic resonance system, an optical interface assembly, and a catheter assembly with a fiber optic cable assembly and an optical interface and an antenna. The catheter assembly converts electromagnetic signals received by the antenna into optical signals. The optical interface assembly converts optical signals into electrical signals and electrical signals into optical signals. | 1. An assembly for delivering optical signals, wherein said assembly is comprised of a nuclear magnetic resonance system, a optical interface assembly electrically connected to said nuclear magnetic resonance system, and a catheter assembly connected to said optical interface assembly, wherein: (a) said optical interface assembly converts optical signals into electrical signals; and (b) said catheter assembly is comprised of a fiber optic cable assembly, and an optical connection assembly at the proximal end of said catheter assembly, and an antenna disposed at the distal end of said catheter assembly, wherein said catheter assembly comprises means for converting electromagnetic signals received by said antenna into optical signals, and wherein said means for converting electromagnetic signals received by said antenna is disposed in said distal end of said catheter assembly. | 1. An assembly for delivering optical signals, wherein said assembly is comprised of a nuclear magnetic resonance system, a optical interface assembly electrically connected to said nuclear magnetic resonance system, and a catheter assembly connected to said optical interface assembly, wherein: (a) said optical interface assembly converts optical signals into electrical signals; and (b) said catheter assembly is comprised of a fiber optic cable assembly, and an optical connection assembly at the proximal end of said catheter assembly, and an antenna disposed at the distal end of said catheter assembly, wherein said catheter assembly comprises means for converting electromagnetic signals received by said antenna into optical signals, and wherein said means for converting electromagnetic signals received by said antenna is disposed in said distal end of said catheter assembly. 2. The assembly as recited in claim 1, wherein said nuclear magnetic resonance system is a nuclear magnetic resonance imaging system. 3. The assembly as recited in claim 1, wherein said nuclear magnetic resonance system is a nuclear magnetic resonance spectrometer. 4. The assembly as recited in claim 1, wherein said catheter assembly is detachably connected to said optic interface assembly. 5. The assembly as recited claim 1 wherein said nuclear magnetic resonance system is comprised of a signal input channel. 6. The assembly as recited in claim 5, wherein said optical interface assembly is connected to said signal input channel. 7. The assembly as recited in claim 1, wherein said nuclear magnetic resonance system is comprised of a gating output channel. 8. The assembly as recited in claim 7, wherein said optical interface assembly is connected to said gating output channel. 9. The assembly as recited in claim 1, further comprising means for generating a calibration reference optical signal disposed in the distal end of said catheter assembly. 10. The assembly as recited in claim 1, wherein said fiber optic cable assembly is comprised of one strand of optical fiber. 11. The assembly as recited in claim 1, wherein said fiber optic cable assembly is comprised of a multiplicity of strands of optical fiber. 12. The assembly as recited in claim 1, wherein said fiber optic cable assembly is comprised of a hollow lumen. 13. The assembly as recited in claim 12, wherein said hollow lumen is a nonmagnetic, nonconductive hollow lumen. 14. The assembly as recited in claim 1, wherein said optical interface assembly is comprised of means for converting analog optical signals to electrical signals. 15. The assembly as recited in claim 1, wherein said optical interface assembly is comprised of means for converting digital optical signals to electrical signals. 16. The assembly as recited in claim 1, wherein said catheter assembly is comprised of means for converting electrical signals into analog optical signals. 17. The assembly as recited in claim 1, wherein said catheter ... | Biophan Technologies, Inc. | A61B005/05 A61B006/00
US06855118 | B2 | 2005-02-15 | Method and apparatus for measuring leg length in an upright position | 2002-08-28 | An apparatus for measuring body symmetry is disclosed. The apparatus includes a pair of supports, a level, and means for determining support displacement. The supports position the apparatus on left and right body points that are generally symmetrical about a patient's spinal column. The level indicates whether the first support is higher or lower than the second support when placed upon the left and right body points thereby identifying that the body points are not actually symmetrical. The means for determining support displacement determine the displacement of one support relative to the other in a vertical direction when one body point is higher or lower than the other body point. The apparatus is utilized to determine hip or shoulder symmetry and displacement. | 1. An apparatus for measuring body symmetry, comprising: a pair of supports, for positioning the apparatus on left and right body points that are generally symmetrical about a patient's spinal column; a first level indicator, mounted to one of the pair of supports; a second level indicator, mounted to the other of the pair of supports, configured to register a common value with the first level indicator when the pair of supports are placed in the same horizontal plane and vertically parallel to each other; a tube, shared by the first and second level indicators, filled with fluid to indicate the relative displacement of the supports; and means for determining the displacement of one support relative to the other in a vertical direction when one body point is higher or lower than the other body point. | 1. An apparatus for measuring body symmetry, comprising: a pair of supports, for positioning the apparatus on left and right body points that are generally symmetrical about a patient's spinal column; a first level indicator, mounted to one of the pair of supports; a second level indicator, mounted to the other of the pair of supports, configured to register a common value with the first level indicator when the pair of supports are placed in the same horizontal plane and vertically parallel to each other; a tube, shared by the first and second level indicators, filled with fluid to indicate the relative displacement of the supports; and means for determining the displacement of one support relative to the other in a vertical direction when one body point is higher or lower than the other body point. 2. The apparatus as in claim 1 further comprising a horizontal level indicator placed on a top portion of the apparatus generally between the first and second supports. 3. The apparatus as in claim 1 wherein the supports are horizontally adjustable between body points for accurate placement on different patients or different body points of the same patient. 4. An apparatus for measuring inequality of limb length or hip symmetry of a patient, comprising: a pair of supports for positioning the apparatus on either iliac crest of the patient; a first level indicator, mounted to one of the pair of supports; a second level indicator, mounted to the other of the pair of supports, configured to register a common value with the first level indicator when the pair of supports are placed in the same horizontal plane and vertically parallel to each other; a tube, shared by the first and second level indicators, filled with fluid to indicate the relative displacement of the supports; and means for determining the displacement of one support relative to the other in a vertical direction and measuring the amount of that displacement to determine the limb length inequality. | - | A61B005/103
US06916461 | B2 | 2005-07-12 | Metal complexes that contain perfluoroalkyl, process for their production and their use in NMR diagnosis | 2002-07-18 | The invention relates to new, monomeric, perfluoroalkyl-substituted metal complexes, processes for their production, and their use in diagnosis and therapy.;The compounds according to the invention are suitable especially as in vivo contrast media in nuclear spin tomography (MRT). They can preferably be used as blood-pool agents and as lymphographic agents. | 1. A method of imaging lymph nodes, comprising conducting said imaging in conjunction with administration of a contrast agent comprising a perfluoroalkyl-containing compound of formula I RF—L—A I in which RF is a perfluorinated, straight-chain or branched carbon chain with formula —CnF2nX, in which X represents a terminal fluorine, chlorine, bromine, iodine or hydrogen atom, and n stand for numbers 4-30, L means a direct bond, a methylene group, an —NHCO group, a group whereby p means numbers 0 to 10, q and u, independently of one another, mean numbers 0 or 1 and R1 is a hydrogen atom, a methyl group, a —CH2—OH group, a —CH2—CO2H group or a C2-C15 chain, which optionally is interrupted by 1 to 3 oxygen atoms, 1 to 2>C groups or an optionally substituted aryl group and/or is substituted with 1 to 4 hydroxy groups, 1 to 2 C1-C4 alkoxy groups, 1 to 2 carboxy groups, a group —SO3H, or L is a straight-chain, branched, saturated or unsaturated C2-C30 carbon chain, which optionally contains 1 to 10 oxygen atoms, 1 to 3 —NR1 groups, 1 to 2 sulfur atoms, a piperazine, a —CONR1 group, an —SO2 group, an —NR1—CO2 group, 1 to 2 —CO groups, a group —CO—N—T —N(R1)—SO2—RF or 1 to 2 optionally substituted aryls and/or is interrupted by these groups and/or is optionally substituted with 1 to 3 —OR1 groups, 1 to 2 oxo groups, 1 to 2 —NH—COR1 groups, 1 to 2 —CONHR1 groups, 1 to 2 —(CH2)p—CO2H groups, 1 to 2 groups of —(CH2)p—(O)q—CH2CH2—RF, whereby R1, RF and p and q have the above-indicated meanings, and T means a C2-C10 chain, which optionally is interrupted by 1 to 2 oxygen atoms or 1 to 2 —NHCO groups, A stands for a complexing agent or metal complex or their salts of organic and/or inorganic bases or amino acids or amino acid amides, specifically for a complexing agent or complex of general formula II in which R3, Z1 and Y are independent of one another, and R3 has the meaning of R1 or means —(CH2)m—L—RF, whereby m is 0, 1 or 2 and L and RF have the above-mentioned meaning, Z1 independently of one another, mean a hydrogen atom or a metal ion equivalent of atomic numbers 21-29, 39, 42, 44 or 57-83, Y means —OZ1 or whereby Z1, L, RF and R3 have the above-mentioned meanings, or A stands for a complexing agent or complex of general formula III in which R3 and Z1 have the above-mentioned meanings and R2 has the meaning of R1, or A stands for a complexing agent of complex of general formula IV in which Z1 has the above-mentioned meaning, or A stands for a complexing agent or complex of general formula V in which Z1 has the above-mentioned meaning, and o and q stand for numbers 0 or 1 and the sum o+q=1 results, or A stands for a complexing agent or complex of general formula VI in which Z1 has the above-mentioned meaning or A stands for a complexing agent or complex of general formula VII in which Z1 and Y have the above-mentioned meanings or A stands for a complexing agent or complex of general formula VIII in which R3 and Z1 have the above-mentioned mean... | 1. A method of imaging lymph nodes, comprising conducting said imaging in conjunction with administration of a contrast agent comprising a perfluoroalkyl-containing compound of formula I RF—L—A I in which RF is a perfluorinated, straight-chain or branched carbon chain with formula —CnF2nX, in which X represents a terminal fluorine, chlorine, bromine, iodine or hydrogen atom, and n stand for numbers 4-30, L means a direct bond, a methylene group, an —NHCO group, a group whereby p means numbers 0 to 10, q and u, independently of one another, mean numbers 0 or 1 and R1 is a hydrogen atom, a methyl group, a —CH2—OH group, a —CH2—CO2H group or a C2-C15 chain, which optionally is interrupted by 1 to 3 oxygen atoms, 1 to 2>C groups or an optionally substituted aryl group and/or is substituted with 1 to 4 hydroxy groups, 1 to 2 C1-C4 alkoxy groups, 1 to 2 carboxy groups, a group —SO3H, or L is a straight-chain, branched, saturated or unsaturated C2-C30 carbon chain, which optionally contains 1 to 10 oxygen atoms, 1 to 3 —NR1 groups, 1 to 2 sulfur atoms, a piperazine, a —CONR1 group, an —SO2 group, an —NR1—CO2 group, 1 to 2 —CO groups, a group —CO—N—T —N(R1)—SO2—RF or 1 to 2 optionally substituted aryls and/or is interrupted by these groups and/or is optionally substituted with 1 to 3 —OR1 groups, 1 to 2 oxo groups, 1 to 2 —NH—COR1 groups, 1 to 2 —CONHR1 groups, 1 to 2 —(CH2)p—CO2H groups, 1 to 2 groups of —(CH2)p—(O)q—CH2CH2—RF, whereby R1, RF and p and q have the above-indicated meanings, and T means a C2-C10 chain, which optionally is interrupted by 1 to 2 oxygen atoms or 1 to 2 —NHCO groups, A stands for a complexing agent or metal complex or their salts of organic and/or inorganic bases or amino acids or amino acid amides, specifically for a complexing agent or complex of general formula II in which R3, Z1 and Y are independent of one another, and R3 has the meaning of R1 or means —(CH2)m—L—RF, whereby m is 0, 1 or 2 and L and RF have the above-mentioned meaning, Z1 independently of one another, mean a hydrogen atom or a metal ion equivalent of atomic numbers 21-29, 39, 42, 44 or 57-83, Y means —OZ1 or whereby Z1, L, RF and R3 have the above-mentioned meanings, or A stands for a complexing agent or complex of general formula III in which R3 and Z1 have the above-mentioned meanings and R2 has the meaning of R1, or A stands for a complexing agent of complex of general formula IV in which Z1 has the above-mentioned meaning, or A stands for a complexing agent or complex of general formula V in which Z1 has the above-mentioned meaning, and o and q stand for numbers 0 or 1 and the sum o+q=1 results, or A stands for a complexing agent or complex of general formula VI in which Z1 has the above-mentioned meaning or A stands for a complexing agent or complex of general formula VII in which Z1 and Y have the above-mentioned meanings or A stands for a complexing agent or complex of general formula VIII in which R3 and Z1 have the above-mentioned mean... | Schering AG | A61B005/055 A61K051/00
US06975896 | B2 | 2005-12-13 | Fiducial markers for MRI | 2002-05-23 | Fiducials (50) are disposed in an imaging region (14) along with a subject. The fiducials are mounted either to the subject itself, a surgical tool (52), an imaging probe or receive coil (80), or the like. The fiducials are filled with a liquid or gel of a fluorine 19 (Fl19) compound which has a resonance frequency that is only 6% off from the resonance frequency of protons. This enables a common transmitter (22), transmit and receive coils, and receiver (34) to be utilized for both proton and fluorine 19 imaging. The proton and fluorine resonance signals are separately reconstructed into corresponding image memories (42, 64). From the positions of the fiducials, a position calculator (66) determines the position of a fiducial carrying surgical tool or probe relative to the proton image. A depiction of the tool or probe from a look-up table (70) is appropriately positioned and rotated (68) and superimposed on the proton image by a video processor (44). | 1. A method of magnetic resonance comprising: disposing a plurality of fiducial markers on at least one of a subject and a medical accessory, the fiducial markers including dipoles having a first resonance frequency different from a proton resonance frequency;disposing a subject in an imaging region of a main magnetic field;exciting magnetic resonance in both dipoles of the fiducial markers and dipoles in at least a portion of the subject concurrently with common radio frequency pulses, the dipoles of the subject having a second resonance frequency that is the proton resonance frequency;spatially encoding the magnetic resonance with magnetic field gradients which spread the first and second resonance frequencies over spatially dependent first and second frequency spectra, the first and second spectra being separated from each other;concurrently receiving magnetic resonance signals from resonating dipoles of the subject and the fiducial markers;separating the magnetic resonance signals based on frequency spectra;separately reconstructing: an image representation of the fiducial markers from the first spectra resonance signals; and,an image representation of the at least a portion of the subject from the second spectra resonance signals. | 1. A method of magnetic resonance comprising: disposing a plurality of fiducial markers on at least one of a subject and a medical accessory, the fiducial markers including dipoles having a first resonance frequency different from a proton resonance frequency;disposing a subject in an imaging region of a main magnetic field;exciting magnetic resonance in both dipoles of the fiducial markers and dipoles in at least a portion of the subject concurrently with common radio frequency pulses, the dipoles of the subject having a second resonance frequency that is the proton resonance frequency;spatially encoding the magnetic resonance with magnetic field gradients which spread the first and second resonance frequencies over spatially dependent first and second frequency spectra, the first and second spectra being separated from each other;concurrently receiving magnetic resonance signals from resonating dipoles of the subject and the fiducial markers;separating the magnetic resonance signals based on frequency spectra;separately reconstructing: an image representation of the fiducial markers from the first spectra resonance signals; and,an image representation of the at least a portion of the subject from the second spectra resonance signals. 2. A method of magnetic resonance comprising: disposing a plurality of fiducial markers on at least one of a subject and a medical accessory, the fiducial markers including dipoles having a first resonance frequency;disposing a subject in an imaging region of a main magnetic field, dipoles of the subject having a second resonance frequency;concurrently exciting magnetic resonance at the first and second resonance frequencies using the same radio frequency transmitter and RF coils;spatially encoding the magnetic resonance with magnetic field gradients which spread the first and second resonance frequencies over spatially dependent first and second frequency spectra, the first and second spectra being separated from each other;concurrently receiving the first and second resonance signals using the same RF coil and receiver;separating the magnetic resonance signals based on frequency spectra;reconstructing an image representation of the fiducial markers from the first spectra resonance signals; and,reconstructing an image representation of the at least a portion of the subject from the second spectra resonance signals. 3. The method as set forth in claim 2, further including: calculating positions of the plurality of fiducial markers; and,calculating an orientation of a tool on which the plurality of fiducial markers are mounted. 4. The method as set forth in claim 2, wherein: the fiducials include fluorine 19 and the first resonance frequency is a resonance frequency of fluorine 19; andthe second resonance frequency is a resonance frequency of protons. 5. The method as set forth in claim 4 wherein the main magnetic field is 0.23 T, the first resonance frequency is about 9.2 MHZ, the second resonance frequency is ... | Koninklijke Philips Electronics N.V. | A61B005/05
US06980847 | B2 | 2005-12-27 | Apparatus for dynamic angular position tracking | 2002-03-25 | A magnetic resonance imaging (“MRI”) system and method of correcting for the motion of a subject during an MRI scan. The motion of a subject in an MRI scan is corrected by acquiring nuclear magnetic resonance (“NMR”) spectra of an ordered material attached to a subject, wherein there is a dependence on the NMR spectra of the ordered material's orientation in a magnetic field, and determining an orientation of the subject. Correction information is then supplied to a primary spectrometer channel which is used to execute an MRI scan, and an imaging pulse sequence of the MRI scan is adjusted to correct the subject's orientation. | 1. A method of correcting for the motion of a subject during a magnetic resonance imaging (“MRI”) scan, comprising: acquiring a nuclear magnetic resonance (“NMR”) spectra of an ordered material attached to the subject, wherein there is a dependence on the NMR spectra of the ordered material's orientation in a magnetic field; determining an orientation of the subject based on the acquired spectra, and correcting the MRI scan. | 1. A method of correcting for the motion of a subject during a magnetic resonance imaging (“MRI”) scan, comprising: acquiring a nuclear magnetic resonance (“NMR”) spectra of an ordered material attached to the subject, wherein there is a dependence on the NMR spectra of the ordered material's orientation in a magnetic field; determining an orientation of the subject based on the acquired spectra, and correcting the MRI scan. 2. The method according to claim 1, further comprising: supplying correction information to a primary spectrometer channel, which is used to execute an MRI scan, based on the determined orientation of the subject to adjust at least one MRI scan parameter. 3. The method according to claim 2, further comprising: adjusting an imaging pulse sequence of the MRI scan, based on the correction information, to correct for the subject's orientation. 4. The method according to claim 3, wherein the imaging pulse sequence of the MRI scan is adjusted based on a request. 5. The method according to claim 1, wherein the ordered material is a single crystal, oriented polymer, or liquid crystal. 6. The method according to claim 1, further comprising: attaching a radio frequency (“RF”) transmit/receive circuit to the subject. 7. The method according to claim 6, wherein the ordered material is embedded in a solenoid coil in the RF transmit/receive circuit. 8. The method according to claim 1, wherein the spectra of the ordered material are acquired with an independent spectrometer channel of the MRI system that operates at a separate frequency from the primary spectrometer channel executing the MRI scan. 9. The method according to claim 1, wherein the orientation of the subject is determined by correlating changes in the peak splittings of the NMR spectra of the ordered material with the changes in the subject's orientation. 10. The method according to claim 9, wherein the subject's orientation is determined by accessing calibration tables pertaining to the ordered material. 11. The method according to claim 1, wherein the orientation of the subject is determined by correlating changes in the peak shiftings of the NMR spectra of the ordered material with the changes in the subject's orientation. 12. The method according to claim 11, wherein the subject's orientation is determined by accessing calibration tables pertaining to the ordered material. 13. The method according to claim 1, further comprising: recording an image after the imaging pulse sequence is adjusted. 14. The method according to claim 1, wherein two ordered materials are employed. 15. The method according to claim 14, wherein two ordered materials with different resonance amplitudes are employed. 16. The method according to claim 14, wherein at least one ordered material is oriented to detect changes in pitch. 17. The method according to claim 14, wherein at least one ordered material is oriented to detect changes in yaw. 18. A method of correcting for the moti... | New York State Department of Mental Hygiene Office of Mental Health | A61B005/055
US06939303 | B2 | 2005-09-06 | Method and system of determining cardiac contractility | 2002-08-09 | A method and system is provided for determining the cardiac contractility of a cardiovascular system. The method includes utilizing values representing stroke volume, preload, and afterload to calculate the cardiac contractility. The calculation may include a mathematical relationship in which the cardiac contractility is represented by a first edge of a tetrahedron, the preload is represented by a second edge of the tetrahedron, the afterload is represented by a third edge of the tetrahedron, and the stroke volume is represented by the volume of the tetrahedron. A display apparatus is also provided. The display apparatus includes a schematic representation of a tetrahedron and indicia indicating that the cardiac contractility, the preload, the afterload, and the stroke volume are represented by mutually distinct parameters of the tetrahedron. | 1. A method of determining the cardiac contractility of a cardiovascular system, said method comprising: determining a value for the stroke volume of the cardiovascular system; determining a value for the preload of the cardiovascular system; determining a value for the afterload of the cardiovascular system; utilizing the value of said stroke volume, the value of said preload, and the value of said afterload to determine said cardiac contractility; and wherein said preload comprises a parameter selected from the group consisting of end-diastolic volume, end-diastolic volume index, end-diastolic pressure, pulmonary capillary pressure, end-diastolic stress, and left ventricle end-diastolic dimension. | 1. A method of determining the cardiac contractility of a cardiovascular system, said method comprising: determining a value for the stroke volume of the cardiovascular system; determining a value for the preload of the cardiovascular system; determining a value for the afterload of the cardiovascular system; utilizing the value of said stroke volume, the value of said preload, and the value of said afterload to determine said cardiac contractility; and wherein said preload comprises a parameter selected from the group consisting of end-diastolic volume, end-diastolic volume index, end-diastolic pressure, pulmonary capillary pressure, end-diastolic stress, and left ventricle end-diastolic dimension. 2. The method of claim 1 wherein said preload comprises said left ventricle end-diastolic dimension. 3. The method of claim 2 wherein said preload is determined using an echocardiogram. 4. The method of claim 2 wherein said stroke volume comprises said left ventricle end-diastolic dimension minus left ventricle end-systolic dimension. 5. The method of claim 1 wherein determining said cardiac contractility utilizes a relationship in which said cardiac contractility is represented by a first edge of a tetrahedron, said preload is represented by a second edge of said tetrahedron, said afterload is represented by a third edge of said tetrahedron, and said stroke volume is represented by the volume of said third edge of said tetrahedron, and said stroke volume is represented by the volume of said tetrahedron. 6. A method of determining the cardiac contractility of a cardiovascular system, said method comprising: determining a value for the stroke volume of the cardiovascular system; determining a value for the preload of the cardiovascular system; determining a value for the afterload of the cardiovascular system; utilizing the value of said stroke volume, the value of said preload, and the value of said afterload to determine said cardiac contractility; and wherein said afterload comprises a parameter selected from the group consisting of arterial pressure, left ventricle systolic pressure, arterial impedance, vascular resistance, systolic pressure, left ventricle meridional stress, left ventricle circumferential stress, and systolic stress. 7. The method of claim 6 wherein said afterload comprises said left ventricle systolic pressure. 8. The method of claim 7 wherein said left ventricle systolic pressure is determined by adding the systolic gradient of said cardiovascular system to said arterial pressure. 9. The method of claim 6 wherein said afterload comprises a member of the group consisting of systolic pressure, left ventricle meridional stress, left ventricle circumferential stress, and systolic stress. 10. The method of claim 9 wherein said systolic pressure, said systolic stress, said left ventricle meridional stress, and said left ventricle circumferential stress may be utilized interchangeably for said afterload. 11. The method of c... | Provecto Tecnico R-2000, Ca | A61B005/02
US06878120 | B2 | 2005-04-12 | Methods and apparatus for suctioning and pumping body fluid from an incision | 2002-04-08 | A sampling device for sampling body fluid includes a casing which is pressed against a skin surface to produce a seal therewith. A cocking mechanism is pulled rearwardly to place a lancing device in a cocked state. By releasing a trigger, the lancing device is driven forward to produce an incision in the skin, and then is retracted out of the incision. A plunger mounted on the cocking mechanism is then retracted to generate a negative pressure at the front end of the casing to draw body fluid from the incision. | 1. A sampling device for sampling body fluid, comprising: a casing defining a longitudinal axis; a skin-lancing mechanism for extending a lancet through a longitudinal front end of the casing and against a skin surface to form an incision therein, and then retracting the lancet back into the casing; a plunger being longitudinally rearwardly movable relative to the skin lancing mechanism, the plunger being constructed and arranged with the sampling device such that a negative pressure is generated at the front end of the casing by movement of the plunger away from the front end of the casing, the negative pressure acting to draw body fluid from the incision; means for driving the plunger rearwardly relative to the skin-lancing mechanism; and a latch constructed and arranged to release the plunger such that the plunger is rearwardly driven relative to the skin-lancing mechanism. | 1. A sampling device for sampling body fluid, comprising: a casing defining a longitudinal axis; a skin-lancing mechanism for extending a lancet through a longitudinal front end of the casing and against a skin surface to form an incision therein, and then retracting the lancet back into the casing; a plunger being longitudinally rearwardly movable relative to the skin lancing mechanism, the plunger being constructed and arranged with the sampling device such that a negative pressure is generated at the front end of the casing by movement of the plunger away from the front end of the casing, the negative pressure acting to draw body fluid from the incision; means for driving the plunger rearwardly relative to the skin-lancing mechanism; and a latch constructed and arranged to release the plunger such that the plunger is rearwardly driven relative to the skin-lancing mechanism. 2. The device of claim 1, wherein the means for driving the plunger includes a spring. 3. The device of claim 2, wherein the spring is coupled to the plunger. 4. The device of claim 1, wherein: the plunger defines a notch; and the latch releasably engages the notch in the plunger to retain the plunger in a longitudinally forward state. 5. The device of claim 1, further comprising: a cocking mechanism coupled to the skin-lancing mechanism for positioning the skin-lancing mechanism in a cocked state; and wherein the plunger is longitudinally rearwardly movable relative to the cocking mechanism. 6. The device of claim 5, wherein the cocking mechanism includes a handle to manually position the skin-lancing mechanism in the cocked state. 7. The device of claim 5, wherein the cocking mechanism includes the latch adapted to be releasably latched through insertion of the lancet into the casing. 8. The device of claim 1, further comprising a trigger to releasably hold the skin-lancing mechanism in a cocked state. 9. The device of claim 8, further comprising an actuator to release the skin-lancing mechanism from the trigger. 10. A bodily fluid sampling device, comprising: a casing defining an opening; a skin-lancing mechanism disposed in the casing to extend a lancet from the opening of the casing to form an incision in skin; a plunger disposed in the casing, the plunger being independently moveable with respect to the skin-lancing mechanism to reduce pressure at the opening to draw bodily fluid from the incision; and a hollow member mounted around the opening in the casing, the hollow member having a skin engaging surface to establish a seal with the skin; and an inner member positioned inside the hollow member, the inner member being movable relative to the hollow member to break the seal with the skin. 11. The device of claim 10, further comprising means for moving the plunger relative to the skin-lancing mechanism. 12. The device of claim 11, wherein the means for moving the plunger includes a spring. 13. The device of claim 11, further comprising a latch construc... | Roche Diagnostics Operations, Inc. | A61B005/00
US06904306 | B1 | 2005-06-07 | Method and apparatus for evaluation of contrast agent uptake based on derived parametric images | 2002-01-23 | A magnetic resonance imaging (MRI) apparatus (10) acquires a plurality of parametric images (60) with at least one varying imaging parameter. A parametric map (62) is constructed from the plurality of parametric images (60). At least one pilot parameter (64) is identified from at least the parametric map (62). The at least one identified pilot parameter (64) includes at least a volume of interest for a diagnostic image. A contrast agent (54) is administered to the patient (18). The identified volume of interest is imaged during influx of the administered contrast agent (54) into the identified volume of interest. The imaging uses the at least one identified pilot parameter (64). | 1. A method for imaging tissue viability or vascular function in a patient using a magnetic resonance imaging (MRI) apparatus, the method comprising: acquiring a parametric map having blood oxygenation level dependent contrast; estimating an improved slice orientation based on the parametric map; optimizing the slice orientation using at least one iteration of the acquiring and estimating; subsequent to the acquiring and optimizing, administering a magnetic contrast agent to the patient; and imaging during a transient distribution of the contrast agent in the patient at the optimized slice orientation. | 1. A method for imaging tissue viability or vascular function in a patient using a magnetic resonance imaging (MRI) apparatus, the method comprising: acquiring a parametric map having blood oxygenation level dependent contrast; estimating an improved slice orientation based on the parametric map; optimizing the slice orientation using at least one iteration of the acquiring and estimating; subsequent to the acquiring and optimizing, administering a magnetic contrast agent to the patient; and imaging during a transient distribution of the contrast agent in the patient at the optimized slice orientation. 2. The method as set forth in claim 1, further including: administering a stress-causing agent to the patient to alter patient metabolism during the acquiring of the parametric map. 3. The method as set forth in claim 2, wherein the stress-causing agent includes at least one of dobutamine, dipyridamole, and adenosine. 4. The method as set forth in claim 1, wherein the step of imaging a transient distribution of the contrast agent in the patient includes: acquiring first pass image data during a time when the contrast agent is taken up into an organ of interest to be imaged. 5. The method as set forth in claim 1, wherein the step of imaging a transient distribution of the contrast agent in the patient includes: acquiring late enhancement image data during a time when the contrast agent is removed from an organ of interest. 6. The method as set forth in claim 1, wherein the step of imaging a transient distribution of the contrast agent in the patient further includes: delaying imaging for a selected time; subsequent to the imaging delay, acquiring a plurality of delayed parametric images; and deriving a parametric map having a late enhancement contrast from the delayed parametric images. 7. The method as set forth in claim 1, wherein the acquiring of a parametric map includes: acquiring a plurality of images having a varying T2* or T2 weighting; and constructing a T2* or T2 map from the plurality of images. 8. The method as set forth in claim 1, wherein the acquiring of a parametric map includes: acquiring a plurality of images during a multiple-echo readout; and estimating a pixel intensity decay time constant or rate for each pixel based on the plurality of images. 9. The method as set forth in claim 1, wherein the acquiring of a parametric map includes: administering a stress agent to the patient; prior to the administering of the stress agent, acquiring a first set of images; subsequent to the administering of the stress agent, acquiring a second set of images; and combining the first and second sets of images to construct a parametric map identifying stressed tissues. 10. The method as set forth in claim 9, wherein the combining of the first and second sets of images includes: calculating at least one unstressed statistical quantity for each pixel of the first set of images; calculating at least one stressed statistical quantity... | Koninklijke Philips Electronics N.V. | A61B005/055
US06939308 | B2 | 2005-09-06 | Acoustic sensor array for non-invasive detection of coronary artery disease | 2002-11-08 | Methodology for determining the bounds of a patient's acoustic window is described. Medical application acoustic array designs with apertures accommodated by patient acoustic windows and merged acoustic windows are exemplified. | 1. A method for visualizing the acoustic window of a person, which comprises: (i) causing a person to assume a supine position; (ii) drawing a series of indicia on the chest of the person along the left sternal border at the beginning of each intercostal space for each of a selected plurality of intercostal spaces 1 to 6 of the person; (iii) placing a probe at the left sternal border of each of the selected intercostal spaces and moving said probe along each of the intercostal spaces until lung tissue is encountered; (iv) placing an indicator on the chest of the person to identify as to where at each intercostal space lung tissue is encountered; and (v) drawing a series of lines connecting said lung indicator indicia to the left sternal border on the chest of the person, wherein said lines indicate the bounds of the acoustic window perimeter of one corresponding intercostal space. | 1. A method for visualizing the acoustic window of a person, which comprises: (i) causing a person to assume a supine position; (ii) drawing a series of indicia on the chest of the person along the left sternal border at the beginning of each intercostal space for each of a selected plurality of intercostal spaces 1 to 6 of the person; (iii) placing a probe at the left sternal border of each of the selected intercostal spaces and moving said probe along each of the intercostal spaces until lung tissue is encountered; (iv) placing an indicator on the chest of the person to identify as to where at each intercostal space lung tissue is encountered; and (v) drawing a series of lines connecting said lung indicator indicia to the left sternal border on the chest of the person, wherein said lines indicate the bounds of the acoustic window perimeter of one corresponding intercostal space. 2. The method of claim 1, further comprising visualizing a boundary line with a perimeter geometry that encloses all of said individual intercostals space lines to define total acoustic window space of the person. 3. The method of claim 2, further comprising determining an acoustic window centroid of the total acoustic window space. 4. The method of claim 1, wherein the plurality of selected intercostals spaces includes at least spaces 4 and 5. 5. The method of claim 4, wherein the plurality of selected intercostals spaces includes all spaces 1-6. 6. The method of claim 1, wherein the probe is an ultrasound probe. 7. The method of claim 1, further comprising providing a sensor array with a plurality of sensor elements and arranging the sensor array on the patient so that the sensor elements occupy a major portion of and extends substantially within the total visualized acoustic window space. 8. The method of claim 7, wherein the sensor array includes about 32 sensor elements. 9. The method of claim 8, further comprising non-invasively determining if the patient has coronary artery disease. 10. The method of claim 7, wherein the sensor elements are configured to non-invasively detect heart and/or blood flow sounds. 11. The method of claim 10, wherein the sensor elements are acoustic sensors. 12. The method of claim 7, further comprising providing a sensor array and arranging the sensor array on the patient so that it's sensor elements extend across a major portion of the lines of the selected intercostals spaces but substantially within the total visualized acoustic window space. 13. The method of claim 7, wherein the sensor array is a polyvinyldifluoride (PVDF) sensor array. 14. The method of claim 7, wherein the sensor elements are piezoelectric film sensors. 15. The method of claim 7, wherein the sensor array is a plurality of spaced-apart linear sensor arrays. 16. The method of claim 15, wherein each of the linear sensor arrays are positioned on the person above selected respective intercostals spaces. 17. The method of claim 7, wherein the sensor ar... | MedAcoustics, Inc. | A61B005/00
US06860851 | B2 | 2005-03-01 | Vulnerable plaque diagnosis and treatment | 2002-11-27 | Vulnerable plaque is identified through a diagnostic method by identifying pathologic markers such as C-reactive proteins or pH change generated by the pathologic site. First and second detectors are advanced through a blood vessel with each of the detectors selected to detect the pathological marker. The detectors are spaced apart while seeking the marker. Differential concentration of the pathological marker as measured by the first and second detectors indicates a presence of the pathological site in proximity to at least one of the detectors. | 1. A diagnostic method for identifying a pathologic site in a blood vessel, of a patient's vasculature wherein said pathologic site is characterized by a generation of a pathologic marker, said method comprising: advancing a first detector and a second detector into said vasculature with each of said first and second detectors selected to detect an intensity of said marker, said detectors being arranged to avoid contact with a wall of the vasculature; spacing said first and second detectors apart while said detectors are seeking detection of said marker; advancing at least one of said detectors through said vessel with each of said detectors measuring an intensity of said marker at a location in said vasculature of each of said detectors; and comparing an intensity of said marker as measured by each of said detectors and noting an increase in intensity of said marker as measured by said at least one of said detectors relative to an intensity of said marker as measured by the other of said detectors; wherein a differential of intensity of said marker between said first and second detectors indicates a presence of said pathologic site in proximity to said at least one of the detectors and with said differential compensating for a systemic presence of said marker. | 1. A diagnostic method for identifying a pathologic site in a blood vessel, of a patient's vasculature wherein said pathologic site is characterized by a generation of a pathologic marker, said method comprising: advancing a first detector and a second detector into said vasculature with each of said first and second detectors selected to detect an intensity of said marker, said detectors being arranged to avoid contact with a wall of the vasculature; spacing said first and second detectors apart while said detectors are seeking detection of said marker; advancing at least one of said detectors through said vessel with each of said detectors measuring an intensity of said marker at a location in said vasculature of each of said detectors; and comparing an intensity of said marker as measured by each of said detectors and noting an increase in intensity of said marker as measured by said at least one of said detectors relative to an intensity of said marker as measured by the other of said detectors; wherein a differential of intensity of said marker between said first and second detectors indicates a presence of said pathologic site in proximity to said at least one of the detectors and with said differential compensating for a systemic presence of said marker. 2. A method according to claim 1 comprising narrowing a spacing between said first and second detectors after said indication of said presence of said pathologic site to more narrowly focus on said pathologic site. 3. A method according to claim 1 wherein said pathologic site is a vulnerable plaque. 4. A method according to claim 1 wherein said pathologic marker is a C-reactive protein. 5. A method according to claim 1 wherein said pathologic marker is a hydrogen ion concentration (pH). 6. A method according to claim 1 wherein said pathologic marker is an increase in temperature. 7. A method according to claim 1 wherein said pathologic marker is a localized inflammation marker. 8. A method according to claim 1 wherein each of said first and second detectors is placed in said vessel. | EnteroMedics Inc. | A61B005/00
US06850789 | B2 | 2005-02-01 | Combination SPO2/temperature measuring apparatus | 2002-07-29 | A medical diagnostic instrument includes at least one blood oxygen saturation sensor and at least one temperature sensor. The sensors are being capable of measuring blood oxygen saturation and temperature simultaneously after insertion of a probe portion of the instrument relative to a defined body site, such as the axilla, rectum, or sublingual pocket of a patient. At least a portion of the probe portion of the instrument is disposable. | 1. A medical diagnostic instrument comprising: a hollow elongate probe portion which is insertable into a body site; means for measuring the blood oxygen saturation of a patient; means for measuring the temperature of the patient, each of said measuring means being arranged within said elongate probe portion and capable of measuring simultaneously with the other of said measuring means, said blood oxygen saturation measuring means including at least one light transmitter and at least one light detector and said temperature measuring means including at least one temperature sensing element; and a transparent sheath sized for covering said probe portion. | 1. A medical diagnostic instrument comprising: a hollow elongate probe portion which is insertable into a body site; means for measuring the blood oxygen saturation of a patient; means for measuring the temperature of the patient, each of said measuring means being arranged within said elongate probe portion and capable of measuring simultaneously with the other of said measuring means, said blood oxygen saturation measuring means including at least one light transmitter and at least one light detector and said temperature measuring means including at least one temperature sensing element; and a transparent sheath sized for covering said probe portion. 2. The instrument of claim 1, wherein said body site is the axilla of a patient into which said elongate probe portion can be inserted. 3. The instrument of claim 1, further including means for measuring the pulse rate of a patient disposed within said elongate probe portion of said instrument. 4. The instrument of claim 1, including display means for displaying at least one of said measured parameters. 5. The instrument of claim 1, wherein said body site is the sublingual pocket of a patient into which said elongate probe portion of said instrument can be inserted. 6. The instrument of claim 1, wherein said body site is the rectum into which said elongate probe portion of said instrument can be inserted. 7. The instrument of claim 1, wherein said temperature sensing element is disposed at a distal end thereof and in which said at least one light transmitter and at least one light detector are arranged adjacent to said temperature sensing element. 8. The instrument of claim 7, wherein said at least one light transmitter and light detector includes at least one optical fiber disposed in the distal end of said probe portion, said at least one optical fiber being capable of at least one of transmitting and receiving light relative to a said body site. 9. The instrument of claim 1, including at least one light transmitter and at least one light sensor disposed in relation to at least one optical fiber disposed in the distal end of said hollow elongate probe portion, said at least one optical fiber being capable of transmitting and receiving light relative to said body site, said at least one light transmitter and said at least one light sensor being disposed remotely from the distal end of said probe portion. 10. The instrument of claim 1, wherein said temperature measuring means includes at least one thermistor and resistance measuring means connected electrically to said at least one thermistor, wherein said at least one thermistor is disposed in a distal end of said probe portion. 11. The instrument of claim 1, wherein said temperature measuring means includes at least one thermocouple. 12. The instrument of claim 1, wherein said probe portion includes a substrate having said blood oxygen saturation means and said temperature measuring means disposed thereupon. 13. The instrument of cla... | Welch Allyn, Inc. | A61B005/00
US06875180 | B2 | 2005-04-05 | Electromagnetic interference immune tissue invasive system | 2002-02-19 | An electromagnetic immune tissue invasive system includes a primary device housing. The primary device housing having a control circuit therein. A shielding is formed around the primary device housing to shield the primary device housing and any circuits therein from electromagnetic interference. A lead system transmits and receives signals between the primary device housing. The lead system is either a fiber optic system or an electrically shielded electrical lead system. | 1. An electromagnetic radiation immune sensing system, comprising: a photonic lead having a proximal end and a distal end; a light source, in the proximal end of said photonic lead, to produce a first light having a first wavelength and a second light having a second wavelength; a wave-guide between the proximal end and distal end of said photonic lead; a bio-sensor, in the distal end of said photonic lead, to measure changes in an electric field located outside a body, the electric field being generated by the shifting voltages on a body's skin surface; and a distal sensor, in the distal end of said photonic lead, to convert the first light into electrical energy and, responsive to said bio-sensor, to reflect the second light back the proximal end of said photonic lead such that a characteristic of the second light is modulated to encode the measured changes in the electric field. | 1. An electromagnetic radiation immune sensing system, comprising: a photonic lead having a proximal end and a distal end; a light source, in the proximal end of said photonic lead, to produce a first light having a first wavelength and a second light having a second wavelength; a wave-guide between the proximal end and distal end of said photonic lead; a bio-sensor, in the distal end of said photonic lead, to measure changes in an electric field located outside a body, the electric field being generated by the shifting voltages on a body's skin surface; and a distal sensor, in the distal end of said photonic lead, to convert the first light into electrical energy and, responsive to said bio-sensor, to reflect the second light back the proximal end of said photonic lead such that a characteristic of the second light is modulated to encode the measured changes in the electric field. 2. The electromagnetic radiation immune sensing system as claimed in claim 1, further comprising: a proximal sensor, in the proximal end of said photonic lead, to convert the modulated second light into electrical energy. 3. The electromagnetic radiation immune sensing system as claimed in claim 2, further comprising: a transmitter, in the proximal end of said photonic lead and operatively connected to said proximal sensor, to transmit, in response the electrical energy from the converted modulated second light, information representing the measured changes in the electric field. 4. The electromagnetic radiation immune sensing system as claimed in claim 1, wherein said light source includes a first emitter to emit the first light having the first wavelength and a second emitter to emit the second light having the second wavelength. 5. The electromagnetic radiation immune sensing system as claimed in claim 1, wherein said light source includes a first laser to produce the first light having the first wavelength and a second laser to produce the second light having the second wavelength. 6. The electromagnetic radiation immune sensing system as claimed in claim 1, wherein said distal sensor includes: an optical attenuator coupled to a mirror; and an optical-electrical conversion device to convert the first light into electrical energy; said optical attenuator attenuating the second light to encode the measured changes in the electric field. 7. The electromagnetic radiation immune sensing system as claimed in claim 6, wherein said optical attenuator attenuating the second light to create pulses of light having equal intensity and periods of no light, the periods of no light differing in time in response to the measured changes in the electric field. 8. The electromagnetic radiation immune sensing system as claimed in claim 6, wherein said optical attenuator attenuating the second light to create light having differing intensities over a period of time. 9. The electromagnetic radiation immune sensing system as claimed in claim 6, further comprising: a beam spli... | Biophan Technologies, Inc. | A61B005/00
US06908441 | B1 | 2005-06-21 | Urine collection device | 2002-09-27 | A urine collection device for facilitating collection of urine samples from a user. The urine collection device includes a funnel member having a bowl portion and a stem portion. The bowl portion has perimeter wall defining a cavity. The cavity of the bowl portion has an open top whereby the open top of the cavity is designed for permitting urine to enter the cavity of the bowl portion. The stem portion has a bore. The bore is in fluid communication with the cavity of the bowl portion whereby the bore of the stem portion is designed for permitting urine collected in the bowl portion to pass through the stem portion. A collection member is selectively coupled to the stem portion of the funnel member opposite the bowl portion for collecting and storing urine when the urine has passed through the bore of the stem portion of the funnel member. | 1. A urine collection device for collecting urine from a user for medical analysis, the urine collection device comprising: a funnel member having a bowl portion and a stem portion, said stem portion being integrally coupled to said bowl portion to inhibit leaking between said stem portion and said bowl portion, said bowl portion having a perimeter wall defining a cavity, said cavity of said bowl portion having an open top such that said open top of said cavity is adapted for permitting urine to enter said cavity of said bowl portion, said stem portion having a bore, said bore being in fluid communication with said cavity of said bowl portion such that said bore of said stem portion is adapted for permitting urine collected in said bowl portion to pass through said stem portion;a collection member being selectively coupled to said stem portion of said funnel member opposite said bowl portion, said collection member having a peripheral wall defining an interior space, said interior space of said collection member being in fluid communication with said bore of said stem portion of said funnel member, said interior space of said collection member being adapted for collecting and storing urine when the urine has passed through said bore of said stem portion of said funnel member;said stem portion of said funnel member having a first end portion and a medial portion, said first portion end portion of said stem portion being coupled to said bowl portion of said funnel member, said first end portion being pivotally coupled to said medial portion of said stem portion, said bowl portion being substantially pivotally positionable with respect to said medial portion such that said bowl portion is adapted for being positioned for comfort when said bowl portion is positioned between the legs of the user;said stem portion of said funnel member having a second end portion, said second end of said stem portion being rotatably coupled to said medial portion opposite said first end portion, said second end of said of said stem portion being selectively coupled to said collection member, said second end portion of said stem portion being for permitting said collection member to be rotatably positioned around an axis of said medial portion of said stem portion. | 1. A urine collection device for collecting urine from a user for medical analysis, the urine collection device comprising: a funnel member having a bowl portion and a stem portion, said stem portion being integrally coupled to said bowl portion to inhibit leaking between said stem portion and said bowl portion, said bowl portion having a perimeter wall defining a cavity, said cavity of said bowl portion having an open top such that said open top of said cavity is adapted for permitting urine to enter said cavity of said bowl portion, said stem portion having a bore, said bore being in fluid communication with said cavity of said bowl portion such that said bore of said stem portion is adapted for permitting urine collected in said bowl portion to pass through said stem portion;a collection member being selectively coupled to said stem portion of said funnel member opposite said bowl portion, said collection member having a peripheral wall defining an interior space, said interior space of said collection member being in fluid communication with said bore of said stem portion of said funnel member, said interior space of said collection member being adapted for collecting and storing urine when the urine has passed through said bore of said stem portion of said funnel member;said stem portion of said funnel member having a first end portion and a medial portion, said first portion end portion of said stem portion being coupled to said bowl portion of said funnel member, said first end portion being pivotally coupled to said medial portion of said stem portion, said bowl portion being substantially pivotally positionable with respect to said medial portion such that said bowl portion is adapted for being positioned for comfort when said bowl portion is positioned between the legs of the user;said stem portion of said funnel member having a second end portion, said second end of said stem portion being rotatably coupled to said medial portion opposite said first end portion, said second end of said of said stem portion being selectively coupled to said collection member, said second end portion of said stem portion being for permitting said collection member to be rotatably positioned around an axis of said medial portion of said stem portion. 2. The urine collection device as set forth in claim 1, further comprising: said collection member having indicia, said indicia being positioned on an exterior surface of said collection member, said indicia being adapted for indicating information needed for identifying the urine sample stored in said collection member. 3. The urine collection device as set forth in claim 1, further comprising: said bowl portion having an upper edge forming an upper crest and a lower crest, said bowl portion being elongated such that said bowl portion is adapted for being positioned between the legs of the user, said upper crest being adapted for being positioned along a front of a groin of the user for inhibiting urine fr... | - | A61B005/00 B65D081/00
US06959211 | B2 | 2005-10-25 | Device for capturing thermal spectra from tissue | 2002-08-06 | A device and method are provided for use with a noninvasive optical measurement system, such as a thermal gradient spectrometer, for improved determination of analyte concentrations within living tissue. In one embodiment, a wearable window is secured to a patient's forearm thereby isolating a measurement site on the patient's skin for determination of blood glucose levels. The wearable window effectively replaces a window of the spectrometer, and thus forms an interface between the patient's skin and a thermal mass window of the spectrometer. When the spectrometer must be temporarily removed from the patient's skin, such as to allow the patient mobility, the wearable window is left secured to the forearm so as to maintain a consistent measurement site on the skin. When the spectrometer is later reattached to the patient, the wearable window will again form an interface between the spectrometer and the same location of skin as before. | 1. A device for use with a noninvasive optical measurement system for capturing thermal spectra from living tissue, said device comprising: a window holder; anda window connected to said window holder, said window comprising a material of high thermal conductivity so as to permit thermal spectra to pass through said window;said device defining a skin contact surface configured for placement in intimate thermal contact with the skin of a patient;said device further defining a system contact surface opposite said skin contact surface; said system contact surface being configured for removable placement against said noninvasive optical measurement system;further comprising a heating element disposed upon said window;wherein said heating element comprises a grid structure. | 1. A device for use with a noninvasive optical measurement system for capturing thermal spectra from living tissue, said device comprising: a window holder; anda window connected to said window holder, said window comprising a material of high thermal conductivity so as to permit thermal spectra to pass through said window;said device defining a skin contact surface configured for placement in intimate thermal contact with the skin of a patient;said device further defining a system contact surface opposite said skin contact surface; said system contact surface being configured for removable placement against said noninvasive optical measurement system;further comprising a heating element disposed upon said window;wherein said heating element comprises a grid structure. 2. The device of claim 1, wherein said window holder has an aperture allowing unimpeded transmission of thermal spectra through said window holder. 3. The device of claim 1, wherein said window holder is made of injection-molded plastic. 4. The device of claim 1, wherein said material is polycrystalline float zone silicon. 5. The device of claim 1, wherein said material has a thickness of about 0.25 millimeters. 6. The device of claim 1, wherein said grid structure comprises a plurality of bridging sub-busses spaced apart from one another and a plurality of heating wires spaced apart from one another, and at least two busses disposed on opposite sides of said heating grid. 7. The device of claim 6, wherein said sub-busses each has a width of about 50 micrometers. 8. The device of claim 6, wherein said sub-busses are spaced apart from one another by about 1.0 millimeters on center. 9. The device of claim 6, wherein said heating wires each has a width of about 20 micrometers. 10. The device of claim 6, wherein said busses comprise an electrical connection whereby electrical communication is established between said heating element and a power supply. 11. The device of claim 10, wherein said power supply is in operative communication with a timed switching device which intermittently supplies electrical power to said heating element via said electrical connection. 12. A device for use with a noninvasive optical measurement system for capturing thermal spectra from living tissue, said device comprising: a window holder; anda window connected to said window holder, said window comprising a material of high thermal conductivity so as to permit thermal spectra to pass through said window;said device defining a skin contact surface configured for placement in intimate thermal contact with the skin of a patient;said device further defining a system contact surface opposite said skin contact surface; said system contact surface being configured for removable placement against said noninvasive optical measurement system;further comprising a heating element disposed upon said window, said heating element comprising a grid structure which in turn comprises a plurality of bridging sub... | OptiScan Biomedical Corp. | A61B005/00
US06862542 | B2 | 2005-03-01 | Erythema measuring device | 2002-01-17 | The present invention provides an erythema meter comprising a light guide that carries light of two specific wavelengths (probing and reference) at two distinct frequencies that are generated and modulated by a either a single or multiple source(s), a photodetector mounted in the tip of the guide that receives light reflected from the surface being examined, and circuitry electrically coupled to the guide for processing the light data, and determining the level of erythema present on the examined surface. The probing and reference wavelengths are delivered in sinusoidal or amplitude modulated fashion, thereby permitting electronic filtering of the received data. A calculating circuit determines the quotient of the two wavelengths after having been reflected off of a surface, such as mucosal or dermal surfaces, which is representative of the severity of erythema present in the surface. | 1. A system for measuring erythema, comprising: a. a light guide having first and second opposite ends; b. a waveform generator for generating at least first and second signals, and modulating said first and second signals at first and second predetermined frequencies, respectively; c. a light source electrically interconnected to said waveform generator for transmitting light to a surface at said first and second frequencies; d. a photo transistor mounted in said light guide at said second end for receiving reflected light and generating a second signal; e. a calculator circuit for determining a level of erythema based on said second signal; and f. a display unit electrically interconnected to said calculator circuit for displaying said level of erythema. | 1. A system for measuring erythema, comprising: a. a light guide having first and second opposite ends; b. a waveform generator for generating at least first and second signals, and modulating said first and second signals at first and second predetermined frequencies, respectively; c. a light source electrically interconnected to said waveform generator for transmitting light to a surface at said first and second frequencies; d. a photo transistor mounted in said light guide at said second end for receiving reflected light and generating a second signal; e. a calculator circuit for determining a level of erythema based on said second signal; and f. a display unit electrically interconnected to said calculator circuit for displaying said level of erythema. 2. An erythema measuring device, comprising: a. a waveform generator for generating first and second waveforms; b. a light guide interconnected to said waveform generator for guiding and transmitting light to a surface at first and second frequencies corresponding to said first and second waveforms, respectively; c. a photo transistor connected to said light guide for receiving light reflected from said surface; d. an electronic highpass filter for filtering said reflected light received by said photo transistor; e. an analog to digital converter electrically coupled to said filter; and f. a display unit electrically coupled to said analog to digital converter. 3. A method for measuring erythema present on a surface, comprising the steps of: a. transmitting first and second wavelengths of light towards said surface; b. receiving said first and second wavelengths of light after they have been reflected off of said surface; c. electronically filtering said first and second wavelengths of reflected light; d. calculating a level of erythema based on said filtered first and second wavelengths of light; and e. displaying said level of erythema. 4. An erythema measuring device, comprising: a. means for generating first and second waveforms; b. means for transmitting light to a surface at first and second frequencies corresponding to said first and second waveforms, respectively; c. means for receiving light reflected from said surface; d. means for filtering said reflected light received by said light receiving means; e. an analog to digital converter electrically coupled to said filtering means; and f. a display unit electrically coupled to said analog to digital converter. 5. An erythema measuring device, comprising: a. means for generating first and second waveforms and modulating said first and second waveforms at first and second predetermined frequencies, respectively; b. means for transmitting light to a surface at said first and second frequencies; c. means for receiving light reflected from said surface; d. means for calculating a level of erythema based on said light reflected from said surface; and e. a display unit for displaying said level of erythema. 6. A method for measuring ... | Charlotte-Mecklenburg Hospital | G06F019/00 A61B005/00
US06881190 | B2 | 2005-04-19 | Standard pulse-wave-propagation-velocity-related-value determining apparatus and pulse-wave-propagation-velocity-related-value obtaining apparatus | 2003-02-03 | An apparatus for determining a standard pulse-wave-propagation-velocity-related value for a living subject, comprising: a blood-pressure measuring device for measuring a blood pressure of the subject; an age determining device for determining an age of the subject; and a standard-pulse-wave-propagation-velocity-related-value determining device for determining the standard pulse-wave-propagation-velocity-related value for the subject, based on the blood pressure of the subject measured by the blood-pressure measuring device and the age of the subject determined by the age determining device, according to a predetermined relationship between (A) (a1) blood pressure and (a2) age and (B) standard pulse-wave-propagation-velocity-related value. | 1. An apparatus for determining a standard pulse-wave-propagation-velocity-related value for a living subject, comprising: a blood-pressure measuring device for measuring a blood pressure of the subject; an age determining device for determining an age of the subject; a standard-pulse-wave-propagation-velocity-related-value determining device for determining the standard pulse-wave-propagation-velocity-related value for the subject, based on the blood pressure of the subject measured by the blood-pressure measuring device and the age of the subject determined by the age determining device, according to a predetermined relationship between (A) (a1) blood pressure and (a2) age and (B) standard pulse-wave-propagation-velocity-related value; and a gender determining device for determining a gender of the subject, wherein the predetermined relationship between (A)(a1) blood pressure and (a2) age and (B) standard pulse-wave-propagation-velocity-related value comprises two predetermined relationships between (A)(a1) blood pressure and (a2) age and (B) standard pulse-wave-propagation-velocity-related value that correspond to male and female, respectively, and wherein the standard-pulse-wave-propagation-velocity-related-value determining device determines the standard pulse-wave-propagation-velocity-related value for the subject, according to one of the two predetermined relationships between (A)(a1) blood pressure and (a2) age and (B) standard pulse-wave-propagation-velocity-related value that corresponds to the gender of the subject determined by the sender determining device. | 1. An apparatus for determining a standard pulse-wave-propagation-velocity-related value for a living subject, comprising: a blood-pressure measuring device for measuring a blood pressure of the subject; an age determining device for determining an age of the subject; a standard-pulse-wave-propagation-velocity-related-value determining device for determining the standard pulse-wave-propagation-velocity-related value for the subject, based on the blood pressure of the subject measured by the blood-pressure measuring device and the age of the subject determined by the age determining device, according to a predetermined relationship between (A) (a1) blood pressure and (a2) age and (B) standard pulse-wave-propagation-velocity-related value; and a gender determining device for determining a gender of the subject, wherein the predetermined relationship between (A)(a1) blood pressure and (a2) age and (B) standard pulse-wave-propagation-velocity-related value comprises two predetermined relationships between (A)(a1) blood pressure and (a2) age and (B) standard pulse-wave-propagation-velocity-related value that correspond to male and female, respectively, and wherein the standard-pulse-wave-propagation-velocity-related-value determining device determines the standard pulse-wave-propagation-velocity-related value for the subject, according to one of the two predetermined relationships between (A)(a1) blood pressure and (a2) age and (B) standard pulse-wave-propagation-velocity-related value that corresponds to the gender of the subject determined by the sender determining device. 2. An apparatus according to claim 1, further comprising: an input device operable for inputting the gender of the subject, wherein the gender determining device determines, as the gender of the subject, the gender inputted through the input device. 3. An apparatus for obtaining a pulse-wave-propagation-velocity-related value for a living subject, comprising: a standard-pulse-wave-propagation-velocity-related-value determining device according to claim 1; an actual-pulse-wave-propagation-velocity-related-value measuring device for measuring an actual pulse-wave-propagation-velocity-related value that is related to a velocity at which a pulse wave propagates in the subject; and a display device which displays the standard pulse-wave-propagation-velocity-related value determined by the standard-pulse-wave-propagation-velocity-related-value determining device and the actual pulse-wave-propagation-velocity-related value measured by the actual-pulse-wave-propagation-velocity-related-value measuring device. 4. An apparatus according to claim 3, wherein the actual-pulse-wave-propagation-velocity-related-value measuring device comprises: an upper-arm-pulse-wave detecting device which includes an upper-arm cuff adapted to be wound around an upper arm of the subject and detects an upper-arm pulse wave of the subject; an ankle-pulse-wave detecting device which includes an ankle cuff ad... | Colin Medical Technology Corporation | A61B005/02
US06970736 | B2 | 2005-11-29 | Analysis system of matter adhered to inside wall of vessel | 2001-10-29 | An analysis system comprises an endoscope 10 insertable into the vessel and a Raman analysis apparatus 40. An insert cable 54 of the Raman analysis apparatus 40 is inserted into a channel 10a of the endoscope 10. An excitation optical fiber 60 and a bundle 70A of a plural number of light receivingoptical fibers 70 are received in the insert cable 54. A transparent small piece 63 having a film-like excitation optical filter 64 is abutted against the distal end of the fiber 60. A transparent plate 71 having a film-like light reciving optical filter 72 is abutted against the distal end of fibers bundle 70A. The plate 71 has a center hole 71a and the piece 63 is fitted into the hole 71a. | 1. An analysis system of matter adhered to an inside wall of a vessel comprising: a guiding apparatus including a main body portion,a flexible inset portion extending from said main body portion and being able to be inserted said vessel and having a window formed in a distal end thereof, anda channel extending through said menu body portion and insert portion and reaching said window, anda Raman analysis apparatus including a light source emitting an excitation light,a spectroscope spectrally analyzing said matter adhered to the inside wall of said vessel,a flexible insert cable to be inserted into said channel and whose distal end is faced with said window,an excitation optical fiber being received in said insert cable and whose basal end is connected to said light source for transmitting said excitation light and whose distal end is arranged at a central area of the distal end portion of said insert cable,a plural number of light receiving optical fibers being received in said insert cable and whose basal end are connected to said spectroscope and whose distal end are arranged in such a manner as to surround said excitation optical fiber at the distal end portion of said insert cable,a transparent excitation small piece having a film-like excitation optical filter adhered to a surface thereof and the surface being abutted against a distal end of said excitation optical fiber, said film-like excitation optical filter cutting all light only excepting a light having a predetermined wavelength, anda transparent light receiving plate having a film-like light receiving optical filter adhered to a surface thereof and the surface being abutted against distal ends of said plural number of light receiving optical fibers, said film-like light receiving optical filter cutting only a light having said predetermined wavelength,wherein said excitation light along said excitation optical fiber is projected through said window and Raman scattered by impinging on said matter adhered to the inside wall of said vessel, said scattered light is made incident to said window, and said incident light is transmitted along said light receiving optical fibers to said spectroscope for analyzing, andwherein said transparent light receiving plate has a center hole and said transparent excitation small piece is fitted into said center hole. | 1. An analysis system of matter adhered to an inside wall of a vessel comprising: a guiding apparatus including a main body portion,a flexible inset portion extending from said main body portion and being able to be inserted said vessel and having a window formed in a distal end thereof, anda channel extending through said menu body portion and insert portion and reaching said window, anda Raman analysis apparatus including a light source emitting an excitation light,a spectroscope spectrally analyzing said matter adhered to the inside wall of said vessel,a flexible insert cable to be inserted into said channel and whose distal end is faced with said window,an excitation optical fiber being received in said insert cable and whose basal end is connected to said light source for transmitting said excitation light and whose distal end is arranged at a central area of the distal end portion of said insert cable,a plural number of light receiving optical fibers being received in said insert cable and whose basal end are connected to said spectroscope and whose distal end are arranged in such a manner as to surround said excitation optical fiber at the distal end portion of said insert cable,a transparent excitation small piece having a film-like excitation optical filter adhered to a surface thereof and the surface being abutted against a distal end of said excitation optical fiber, said film-like excitation optical filter cutting all light only excepting a light having a predetermined wavelength, anda transparent light receiving plate having a film-like light receiving optical filter adhered to a surface thereof and the surface being abutted against distal ends of said plural number of light receiving optical fibers, said film-like light receiving optical filter cutting only a light having said predetermined wavelength,wherein said excitation light along said excitation optical fiber is projected through said window and Raman scattered by impinging on said matter adhered to the inside wall of said vessel, said scattered light is made incident to said window, and said incident light is transmitted along said light receiving optical fibers to said spectroscope for analyzing, andwherein said transparent light receiving plate has a center hole and said transparent excitation small piece is fitted into said center hole. 2. An analysis system of matter adhered to an inside wall of a vessel according to claim 1, wherein another excitation optical filter for cutting all light only excepting a light having said predetermined wavelength is disposed between said light source and said excitation optical fiber. 3. An analysis system of matter adhered to an inside wall of a vessel according to claim 1, wherein said guiding apparatus is an endoscope through which an interior of said vessel can be observed. | Machida Endoscope Co., Ltd. | A61B005/02 A61B006/00 G02B006/04
US06957098 | B1 | 2005-10-18 | Markers for interventional devices in magnetic resonant image (MRI) systems | 2002-06-27 | Methods and apparatuses for locating a catheter, or other medical international devices are described herein. In one aspect of the invention, an exemplary method includes providing a pressure sensitive device attached to a body of a catheter, providing a resonant circuit coupled to the pressure sensitive device in the catheter, the resonant circuit having an oscillating frequency, generating a magnetic field, the magnetic field inducing and causing the resonant circuit to oscillate, applying a pressure on the pressure sensitive device, the pressure applied on the pressure sensitive device changing the frequency of the resonant circuit, detecting the oscillation of the resonant circuit, and determining location of the catheter based on the oscillation of the resonant circuit. Other methods and apparatuses are also described. | 1. A method for locating a catheter, the method comprising: providing a pressure sensitive device attached to a body of a catheter;providing a circuit coupled to the pressure sensitive device in the catheter,generating a magnetic field, the magnetic field inducing and causing the circuit to oscillate;applying a pressure on the pressure sensitive device, the pressure applied on the pressure sensitive device changing an oscillating frequency of the circuit;detecting oscillation of the circuit; anddetermining location of the catheter based on the oscillation of the circuit. | 1. A method for locating a catheter, the method comprising: providing a pressure sensitive device attached to a body of a catheter;providing a circuit coupled to the pressure sensitive device in the catheter,generating a magnetic field, the magnetic field inducing and causing the circuit to oscillate;applying a pressure on the pressure sensitive device, the pressure applied on the pressure sensitive device changing an oscillating frequency of the circuit;detecting oscillation of the circuit; anddetermining location of the catheter based on the oscillation of the circuit. 2. The method of claim 1, wherein the pressure sensitive device is attached to an outer wall of the catheter. 3. The method of claim 1, wherein the pressure sensitive device is incorporated into a wall of the catheter. 4. The method of claim 1, wherein the circuit is located at a distal tip region of the catheter. 5. The method of claim 1, wherein the circuit comprises an LC circuit incorporated into a tip structure of the catheter. 6. The method of claim 5, wherein the LC circuit comprises: a capacitor;an inductor coil coupled to the capacitor in parallel or in series; andwherein the inductor coil is wound around the catheter in a way such that the circuit contains a maximum Q factor. 7. The method of claim 6, wherein the LC circuit comprises multiple inductor coils. 8. The method of claim 7, wherein the multiple inductor coils are separated by capacitors coupled in series or in parallel. 9. The method of claim 1, wherein the magnetic field is generated through a magnetic resonant image (MRI) system. 10. The method of claim 1, wherein the pressure applied on the pressure sensitive device is received from an internal part of a patient's body. 11. The method of claim 10, wherein the pressure applied on the pressure sensitive device is received from a blood pressure of the patient's body. 12. The method of claim 1, wherein the pressure applied on the pressure sensitive device is received from an external pressure source through a pressure controller. 13. The method of claim 12, further comprising adjusting the pressure applied on the pressure sensitive device through the pressure controller, such that the circuit is operating in and out of a resonant state. 14. The method of claim 1, wherein the pressure sensitive device is a pressure sensitive resistor, wherein the resistance of the pressure sensitive resistor changes based on the pressure applied on the pressure sensitive resistor. 15. The method of claim 1, further comprising: generating a radio frequency (RF) signal through the oscillation of the circuit;detecting the oscillation of the circuit through sensing the intensity of the RF signal; anddetermining the location of the catheter distal tip based on signals produced through the RF signal. | Advanced Cardiovascular Systems, Inc. | A61B005/05 A61M001/00
US06923769 | B2 | 2005-08-02 | Method and apparatus for monitoring biological abnormality and blood pressure | 2002-05-21 | In a clinical and physiological abnormality monitoring apparatus, and blood pressure monitoring apparatus detects a blood pressure abnormality and the like of a body by employing a pulse wave signal. A frequency analysis is carried out with respect to a pulse wave signal, while this pulse wave signal corresponds to time sequential data of pulse waves. As a result, both a C-frequency component indicative of a fluctuation component of a base line of the pulse wave signal, and also an A-frequency component representative of the respective pulse waves are acquired. A ratio C/A of power of a peak contained in the C-frequency component with respect to power of a peak contained in the A-frequency component is calculated to determine abnormality of the blood pressure. | 1. A clinical and physiological abnormality monitoring apparatus comprising: means for detectiong vasomotion of a body in a percutaneous manner; and means for determining a clinical and physiological abnormality upon the detected vasomotion, wherein the detecting means detects a change contained in a blood flow condition of the blood vessels in the percutaneous manner so as to acquire a fluctuation of the blood flow condition caused by the vasomotion, the acquired fluctuation being used to determine the clinical and physiological abnormality. | 1. A clinical and physiological abnormality monitoring apparatus comprising: means for detectiong vasomotion of a body in a percutaneous manner; and means for determining a clinical and physiological abnormality upon the detected vasomotion, wherein the detecting means detects a change contained in a blood flow condition of the blood vessels in the percutaneous manner so as to acquire a fluctuation of the blood flow condition caused by the vasomotion, the acquired fluctuation being used to determine the clinical and physiological abnormality. 2. A clinical and physiological abnormality monitoring apparatus as in claim 1, wherein: the detecting means detects frequency component corresponding to the fluctuation of the blood flow condition, which is lower than a frequency component corresponding to a pulse interval. 3. A clinical and physiological abnormality monitoring apparatus as in claim 1, wherein: the detecting means detects both frequency components corresponding to the pulse interval and corresponding to the fluctuation of the blood flow condition, which is lower than the frequency component corresponding to the pulse interval, the detected frequency components being used to determine the clinical and physiological abnormality. 4. A clinical and physiological abnormality monitoring apparatus as in claim 1, wherein: the detecting means detects a change contained in fluctuations of the blood flow condition after the measuring operation is commenced, the detected change contained in the fluctuations being used to determine the clinical and physiological abnormality. 5. A clinical and physiological abnormality monitoring apparatus as in claim 1, wherein: the detecting means detects a change contained in fluctuations of the blood flow condition before a predetermined period is detected, the detected change being used to determine the clinical and physiological abnormality. 6. A clinical and physiological abnormality monitoring apparatus as in claim 1, wherein: the detecting means includes a blood flow condition detecting apparatus unit for acquiring the fluctuation of the blood flow condition. 7. A clinical and physiological abnormality monitoring apparatus as in claim 1, wherein: the determining means determines the clinical and physiological abnormality based upon both the fluctuation of the blood flow condition and body motion. 8. A clinical and physiological abnormality monitoring apparatus as in claim 7, wherein: the detecting means includes a blood flow condition detecting apparatus unit for detecting the fluctuation of the blood flow condition; and a body motion detecting apparatus unit for detecting the body motion. 9. A clinical and physiological abnormality monitoring apparatus as in claim 1, wherein: the detecting means acquires a C-frequency component corresponding to the fluctuation of the blood flow condition from a measured pulse wave signal, the acquired C-frequency component being used to determine the clinical an... | Denso Corporation Junichiro Hayano | A61B005/00
US06892085 | B2 | 2005-05-10 | Glucose sensor package system | 2002-11-26 | A glucose sensor package system that includes a glucose sensor and a protective package that indicates exposure to temperature changes to indicate proper temperature control. Also covered are methods of transporting and sterilizing the package. In addition, glucose sensors directed to various sizing and positioning of the electrodes on the glucose sensor are covered. | 1. A glucose sensor comprising: a substrate; a working electrode that includes at least one enzyme and that is coupled to the substrate; a counter electrode coupled to the substrate; and a reference electrode coupled to the substrate, wherein the counter electrode is formed larger than the working electrode, and the working electrode is formed larger than the reference electrode. | 1. A glucose sensor comprising: a substrate; a working electrode that includes at least one enzyme and that is coupled to the substrate; a counter electrode coupled to the substrate; and a reference electrode coupled to the substrate, wherein the counter electrode is formed larger than the working electrode, and the working electrode is formed larger than the reference electrode. 2. The glucose sensor according to claim 1, wherein the working electrode is placed between the counter electrode and the reference electrode. 3. A glucose sensor according to claim 1, wherein the glucose sensor is adapted to provide continuous glucose sensor measurements. 4. A glucose sensor according to claim 1, wherein the glucose sensor is adapted to provide near continuous glucose sensor measurements. 5. A glucose sensor according to claim 1, wherein the glucose sensor is adapted to provide intermittent glucose sensor measurements. 6. A glucose sensor according to claim 1, wherein the glucose sensor is adapted to fit a needle equal to or smaller than 18 gauge. 7. A glucose sensor according to claim 1, wherein the glucose sensor is adapted to fit a 22 gauge needle. 8. A glucose sensor according to claim 1, wherein the glucose sensor is adapted to fit in a needle size range of 18 to 30 gauge. 9. A glucose sensor according to claim 1, wherein the glucose sensor is adapted for placement in subcutaneous tissue to provide glucose sensor measurements. 10. A glucose sensor according to claim 1, wherein the glucose sensor is adapted for placement in tissue to provide glucose sensor measurements. 11. A glucose sensor according to claim 1, wherein the glucose sensor is sized to minimize pain on insertion. 12. A glucose sensor according to claim 1, wherein the glucose sensor is a surface mounted glucose sensor that is adapted to utilize interstitial fluid harvested from the skin. 13. A glucose sensor according to claim 1, wherein the working electrode, the counter electrode and the reference electrode are formed in a non-circular pattern. 14. A glucose sensor according to claim 1, wherein the working electrode, the counter electrode and the reference electrode are formed in a linear pattern. 15. A glucose sensor according to claim 1, wherein the working electrode, the counter electrode and the reference electrode end in a generally polygonal shape. 16. An analyte sensor comprising: a substrate; a working electrode that includes at least one enzyme and that is coupled to the substrate; a counter electrode coupled to the substrate; and a reference electrode coupled to the substrate, wherein the counter electrode is formed larger than the working electrode, and the working electrode is formed larger than the reference electrode. 17. An analyte sensor according to claim 16, wherein the working electrode is placed between the counter electrode and the reference electrode. 18. An analyte sensor according to claim 16, wherein the analyte glucose sensor is adapted to... | Medtronic MiniMed, Inc. | A61B005/05 A61B005/00 A61B005/04
US06970733 | B2 | 2005-11-29 | System and method for electrode localization using ultrasound | 2002-10-21 | The present invention is a device localization system that uses one or more ultrasound reference catheters to establish a fixed three-dimensional coordinate system within a patient's heart using principles of triangulation. The coordinate system is represented graphically in three-dimensions on a video monitor and aids the clinician in guiding other medical devices, which are provided with ultrasound transducers, through the body to locations at which they are needed to perform clinical procedures. In one embodiment of a system according to the present invention, the system is used in the heart to help the physician guide mapping catheters for measuring electrical activity, and ablation catheters for ablating selected regions of cardiac tissue, to desired locations within the heart. | 1. A medical system, comprising: one or more reference catheters having a plurality of ultrasound reference transducers; a medical device having one or more ultrasound target transducers; and localization processing circuitry coupled to the ultrasound reference and target transducers, the processing circuitry configured for causing the ultrasound reference transducers to transmit and receive ultrasound signals between each other in order to establish a three-dimensional coordinate system, and for causing the plurality of ultrasound reference transducers and the one or more ultrasound target transducers to transmit and receive ultrasound signals between each other in order to determine a position of the one or more ultrasound target transducers within the three-dimensional coordinate system. | 1. A medical system, comprising: one or more reference catheters having a plurality of ultrasound reference transducers; a medical device having one or more ultrasound target transducers; and localization processing circuitry coupled to the ultrasound reference and target transducers, the processing circuitry configured for causing the ultrasound reference transducers to transmit and receive ultrasound signals between each other in order to establish a three-dimensional coordinate system, and for causing the plurality of ultrasound reference transducers and the one or more ultrasound target transducers to transmit and receive ultrasound signals between each other in order to determine a position of the one or more ultrasound target transducers within the three-dimensional coordinate system. 2. The system of claim 1, wherein the processing circuitry is configured for causing the one or more ultrasound target transducers to receive ultrasound signals from the plurality of reference transducers in order to determine the position of the one or more ultrasound target transducers within the three-dimensional coordinate system. 3. The system of claim 1, wherein the processing circuitry is further configured for measuring elapsed time between transmission of ultrasound signals by transmitting transducers and receipt of the signals by receiving transducers, and for calculating distances between the transducers using the measured elapsed times. 4. The system of claim 1, further comprising a graphical display, wherein the processing circuitry is further configured for generating a three-dimensional image of at least a portion of the medical device for display on the graphical display. 5. The system of claim 1, wherein the medical device comprises a catheter. 6. The system of claim 5, wherein the catheter is a mapping catheter. 7. The system of claim 6, wherein the mapping catheter has one or more mapping electrodes, the system further comprising electrophysiology processing circuitry coupled to the one or more mapping electrodes for receiving mapping signals corresponding to electrical activity measured by the one or more mapping electrodes. 8. The system of claim 7, further comprising a graphical display, wherein the localization processing circuitry is further configured for generating a three-dimensional image of at least a portion of the mapping catheter for display on the graphical display, and for generating data representing the received mapping signals for display on the graphical display. 9. The system of claim 5, wherein the catheter is an ablation catheter. 10. The system of claim 9, wherein the ablation catheter comprises one or more ablation electrodes, the system further comprising a radio frequency generator coupled to the one or more ablation electrodes for creating a lesion on tissue. 11. The system of claim 10, further comprising a graphical display, wherein the localization processing circuitry is further configured for generat... | SciMed Life Systems, Inc. | A61B005/04
US06923770 | B2 | 2005-08-02 | Pulse-wave-characteristic-point determining apparatus, and pulse-wave-propagation-velocity-related-information obtaining apparatus employing the pulse-wave-characteristic-point determining apparatus | 2002-04-15 | An apparatus for determining a characteristic point of a pulse wave detected from a living subject, the apparatus including a pulse-wave detecting device which detects a first pulse wave from a prescribed portion of the subject, a pulse-wave modifying device for modifying a waveform of the first pulse wave detected by the pulse-wave detecting device, according to a predetermined relationship between first pulse wave, and second pulse wave whose characteristic point is clear, and a characteristic-point determining device for determining a characteristic point of the modified waveform of the first pulse wave. | 1. An apparatus for determining a characteristic point of a pulse wave detected from a living subject, the apparatus comprising: a pulse-wave detecting device which detects a first pulse wave from a prescribed portion of the subject; a pulse-wave modifying means for modifying a waveform of the first pulse wave detected by the pulse-wave detecting device according to a predetermined relationship between first pulse wave and a second pulse wave, whose characteristic point is clear; and a characteristic-point determining means for determining a characteristic point of the modified waveform of the first pulse wave. | 1. An apparatus for determining a characteristic point of a pulse wave detected from a living subject, the apparatus comprising: a pulse-wave detecting device which detects a first pulse wave from a prescribed portion of the subject; a pulse-wave modifying means for modifying a waveform of the first pulse wave detected by the pulse-wave detecting device according to a predetermined relationship between first pulse wave and a second pulse wave, whose characteristic point is clear; and a characteristic-point determining means for determining a characteristic point of the modified waveform of the first pulse wave. 2. An apparatus for obtaining pulse-wave-propagation-velocity-related information that is related to a velocity at which a pulse wave propagates in a living subject, the apparatus comprising: a pulse-wave detecting device which detects, as a first heartbeat-synchronous signal, a first pulse wave from a first prescribed portion of the subject; a pulse-wave modifying means for modifying a waveform of the first pulse wave detected by the pulse-wave detecting device according to a predetermined relationship between first pulse wave and a second pulse wave, whose characteristic point is clear; a characteristic-point determining means for determining a characteristic point of the modified waveform of the first pulse wave; a heartbeat-synchronous-signal detecting device which detects, from a second prescribed portion of the subject that is different from the first prescribed portion, a second heartbeat-synchronous signal; and a pulse-wave-propagation-velocity-related-information obtaining means for obtaining the pulse-wave-propagation-velocity-related information, based on a time of occurrence of the determined characteristic point of the modified waveform of the first pulse wave and a time of occurrence of a prescribed point of the second heartbeat-synchronous signal detected by the heartbeat-synchronous-signal detecting device. 3. An apparatus according to claim 2, further comprising a display device which displays the pulse-wave-propagation-velocity-related information obtained by the pulse-wave-propagation-velocity-related-information obtaining means. 4. An apparatus according to claim 2, wherein the heartbeat-synchronous-signal detecting device comprises a heart-sound microphone which detects a heart sound produced from the heart of the subject. 5. An apparatus according to claim 4, wherein the pulse-wave detecting device comprises an inflatable cuff which detects a brachial-artery pulse wave produced from an upper arm of the subject. 6. An apparatus according to claim 5, wherein the pulse-wave-propagation-velocity-related-information obtaining means determines, as the pulse-wave-propagation-velocity-related information, a pulse-wave propagation time equal to a time difference between the time of occurrence of the prescribed point of the heart sound as the second heartbeat-synchronous signal detected by the heart-sound microphone as t... | Colin Medical Technology Corporation | A61B005/02
US06913578 | B2 | 2005-07-05 | Method for customizing audio systems for hearing impaired | 2002-04-24 | Improved approaches to assist those having hearing loss are disclosed. One approach pertains to providing customization of personal audio systems for a hearing impaired individual. Another approach pertains to designing and producing an audio product that includes components or software that can be customized for hearing impaired individuals. The customization provided by either approach can be performed on-line or off-line. | 1. A method for on-line customization of an audio system for a particular subject with hearing loss, said method comprising the acts of: performing a first on-line hearing test for a reference subject to produce reference parameters; performimg a second on-line hearing test for the particular subject to produce participant parameters; determining a hearing loss for the particular subject based on the participant parameters and the reference parameters; and determining customization parameters for the audio system based on the hearing loss, wherein the audio system is a personal audio system, and wherein the personal audio system is a media player. | 1. A method for on-line customization of an audio system for a particular subject with hearing loss, said method comprising the acts of: performing a first on-line hearing test for a reference subject to produce reference parameters; performimg a second on-line hearing test for the particular subject to produce participant parameters; determining a hearing loss for the particular subject based on the participant parameters and the reference parameters; and determining customization parameters for the audio system based on the hearing loss, wherein the audio system is a personal audio system, and wherein the personal audio system is a media player. 2. A method as recited in claim 1, wherein said method further comprises: modifying the audio system in accordance with the customization parameters, thereby customizing the audio system for compensation for the hearing loss of the participating subject. 3. A method as recited claim 2, wherein the audio system is a standard telephone. 4. A method as recited in claim 2, wherein the audio system is a mobile telephone. 5. A method for on-line customization of an audio system for a particular subject with hearing loss, said method comprising the acts of: performing a first on-line hearing test for a reference subject to produce reference parameters; performing a second on-line hearing test for the particular subject to produce participant parameters; determining a hearing loss for the particular subject based on the participant parameters and the reference parameters; and determining customization parameters for the audio system based on the hearing loss, wherein the audio system is a personal audio system, and wherein the personal audio system is a mobile telephone. 6. A method as recited in claim 5, wherein said method further comprises: modifying the audio system in accordance with the customization parameters, thereby customizing the audio system for compensation for the hearing loss of the participating subject. 7. A method for on-line customization of an audio system for a particular subject with hearing loss, said method comprising the acts of: performing a first on-line hearing test for a reference subject to produce reference parameters; performing a second on-line hearing test for the particular subject to produce participant parameters; determining a hearing loss for the particular subject based on the participant parameters and the reference parameters; and determining customization parameters for the audio system based on the hearing loss, wherein the audio system is a telephone. 8. A method for on-line customization of an audio system for a particular subject with hearing loss, said method comprising the acts of: performing a first on-line hearing test for a reference subject to produce reference parameters; performing a second on-line hearing test for the particular subject to produce participant parameters; determining a hearing loss for the particular subject based on the partic... | Apherma Corporation | A61B005/00 H04M011/00
US06976963 | B2 | 2005-12-20 | Apparatus and method for precision vital signs determination | 2002-09-30 | Apparatus and techniques are provided for precise measuring and monitoring of certain vital signs of patients that have been difficult to measure especially in emergency situations. Sensor may be used for detecting contraction and expansion in the vascular bed of the lining tissue of the external auditory canal during a cardiac cycle to obtain a better indication of certain physiological parameters. Such a signal generally may be superimposed with an additional signal that is primarily due to breathing activity of the patient. Based on various scenarios different sensors may be used to determine the signal due to breathing activity and thus the physiological parameter of interest may be derived. The signal corresponding to the physiological parameter of interest, for example, blood pressure may then be used to monitor the vital signs or control other medical equipment. | 1. A method comprising: a. sensing a first signal corresponding to a blood pressure superimposed with a signal component corresponding to a respiratory cycle by detecting contraction and expansion in the vascular bed of the lining tissue of the external auditory canal during a cardiac cycle;b. sensing a second signal representing the signal component corresponding to the respiratory cycle; andc. removing the second signal from the first signal to derive a third signal corresponding to the blood pressure. | 1. A method comprising: a. sensing a first signal corresponding to a blood pressure superimposed with a signal component corresponding to a respiratory cycle by detecting contraction and expansion in the vascular bed of the lining tissue of the external auditory canal during a cardiac cycle;b. sensing a second signal representing the signal component corresponding to the respiratory cycle; andc. removing the second signal from the first signal to derive a third signal corresponding to the blood pressure. 2. The method as in claim 1, wherein the blood pressure comprises of a diastolic blood pressure and a systolic blood pressure. 3. The method as in claim 2, further comprising of determining a blood pressure paradox. 4. The method as in claim 3, wherein the blood pressure paradox comprises a time variation of the systolic blood pressure with changes in a respiratory cycle. 5. The method as in claim 3, wherein the changes in the respiratory cycle comprise changes during breathing in and changes during breathing out. 6. The method as in claim 3, wherein the blood pressure paradox is used as an indicator of a disease state. | - | A61B005/02 A61B005/08
US06876886 | B1 | 2005-04-05 | Magnetically shielded conductor | 2002-08-26 | A magnetically shielded conductor assembly containing a conductor disposed within an insulating matrix, and nanomagentic material and nanoelectrical material disposed around the conductor. The conductor has a resistivity at 20 degrees Centigrade of from about 1 to about 100 microohm-centimeters. The insulating matrix is composed of nano-sized particles having a maximum dimension of from about 10 to about 100 nanometers. The insulating matrix has a resistivity of from about 1,000,000,000 to about 10,000,000,000,000 ohm-centimeter. The nanomagnetic material has an average particle size of less than about 100 nanometers. The nanomagnetic material has a saturation magnetization of from about 200 to about 26,000 Gauss. The magnetically shielded conductor assembly is flexible, having a bend radius of less than 2 centimeters. | 1. A magnetically shielded conductor assembly comprised of a conductor disposed within an insulating matrix, nanomagnetic material disposed around said conductor, and nanoelectrical material disposed around said conductor, provided that such nanomagnetic material is not contiguous with said conductor, and further provided that such nanoelectrical material is not contiguous with said conductor, wherein: (a) said conductor has a resistivity of from about 1 to about 100 micro ohm-centimeters, (b) said insulating matrix is comprised of nano-sized particles, wherein at least about 90 weight percent of said particles have a maximum dimension of from about 10 to about 100 nanometers; (c) said insulating matrix has a resistivity of from about 1,000,000,000 to about 10,000,000,000,000 ohm-centimeter; (d) said nanomagnetic material has an average particle size of less than about 100 nanometers; (e) said nanomagnetic material has a saturation magnetization of from about 200 to about 26,000 Gauss and a thickness of less than about 2 microns; (f) said nanoelectrical material has a resisitivity of from about 1.6 to about 100 microohm-centimeters; and (g) said magnetically shielded conductor assembly is flexible, having a bend radius of less than 2 centimeters. | 1. A magnetically shielded conductor assembly comprised of a conductor disposed within an insulating matrix, nanomagnetic material disposed around said conductor, and nanoelectrical material disposed around said conductor, provided that such nanomagnetic material is not contiguous with said conductor, and further provided that such nanoelectrical material is not contiguous with said conductor, wherein: (a) said conductor has a resistivity of from about 1 to about 100 micro ohm-centimeters, (b) said insulating matrix is comprised of nano-sized particles, wherein at least about 90 weight percent of said particles have a maximum dimension of from about 10 to about 100 nanometers; (c) said insulating matrix has a resistivity of from about 1,000,000,000 to about 10,000,000,000,000 ohm-centimeter; (d) said nanomagnetic material has an average particle size of less than about 100 nanometers; (e) said nanomagnetic material has a saturation magnetization of from about 200 to about 26,000 Gauss and a thickness of less than about 2 microns; (f) said nanoelectrical material has a resisitivity of from about 1.6 to about 100 microohm-centimeters; and (g) said magnetically shielded conductor assembly is flexible, having a bend radius of less than 2 centimeters. 2. The magnetically shielded conductor assembly as recited in claim 1, wherein said magnetically shielded conductor assembly further comprises a layer comprised of said nanomagnetic material disposed around said conductor, and a layer of said nanoelectrical material disposed around said conductor. 3. The magnetically shielded conductor assembly as recited in claim 1, wherein disposed around said conductor is a layer comprised of said nanomagnetic material and said nanoelectrical material. 4. The magnetically shielded conductor assembly as recited in claim 1, wherein said nanoelectrical material is comprised of aluminum and iron. 5. The magnetically shielded conductor assembly as recited in claim 4, wherein said nanoelectrical material consists essentially of aluminum and iron. 6. The magnetically shielded conductor assembly as recited in claim 5, wherein at least 9 moles of aluminum are present for each mole of iron in said nanoelectrical material. 7. The magnetically shielded conductor assembly as recited in claim 1, wherein said conductor and said second conductor consists essentially of copper. 8. The magnetically shielded conductor assembly as recited in claim 1, wherein said nanomagnetic material is comprised of atoms of iron. 9. The magnetically shielded conductor assembly as recited in claim 1, wherein said nanomagnetic material is comprised of atoms of cobalt. 10. The magnetically shielded conductor assembly as recited in claim 1, wherein said nanomagnetic material is comprised of atoms of nickel. 11. The magnetically shielded conductor assembly as recited in claim 1, further comprising a power supply operatively connected to said conductor. 12. The magnetically shielded conductor ass... | Nandset, LLC | A61B005/0408 A61N001/05 H01B011/06
US06918881 | B2 | 2005-07-19 | Biopsy needle with integrated guide pin | 2002-05-31 | A biopsy instrument for use in a biopsy gun includes a hollow outer needle defining a lumen therethrough and an inner needle slidably engaged within the lumen. The inner needle defines a cavity adjacent its distal end for receiving a tissue sample. A first hub is attached to the proximal end of the inner needle, and a second hub is attached to the proximal end of the outer needle. An elongated pin having a first end and a second end is fixedly attached to the first hub at the first end of the pin. The second end is in contact with the second hub so that the hub and pin are slideable relative to each other. A stop engagement between the second end of the pin and the second hub is configure to prevent disengagement of the needles and to maintain a specific physical relationship between the hubs. | 1. A biopsy instrument for use in a biopsy gun, comprising: a hollow outer needle having a proximal end and a distal end and defining a lumen therethrough; an inner needle slidably disposed within said lumen, said inner needle having a proximal end and a distal end and defining a cavity adjacent said distal end for receiving a tissue sample; said inner and outer needles slideable relative to one another between a first position wherein said distal end of said inner needle is retracted within said lumen so that said outer needle covers said cavity and a second position wherein said distal end of said inner needle is extended away from said outer needle to expose said cavity; a first hub attached to said proximal end of said inner needle; a second hub attached to said proximal end of said outer needle, said second hub having a track running parallel to said outer needle and a stop member projecting from a surface of said second hub adjacent said track, wherein said track is defined by said surface, a wall projecting from said surface and a dog projecting from said wall; and an elongated pin having a first end, a second end and an arm projecting from said second end, said first end fixedly attached to said first hub with said shaft parallel to said inner needle and said pin slidingly engaged to said second hub within said track, said second hub slideable along said pin between a location near said first end and said second end where said arm meets said stop member. | 1. A biopsy instrument for use in a biopsy gun, comprising: a hollow outer needle having a proximal end and a distal end and defining a lumen therethrough; an inner needle slidably disposed within said lumen, said inner needle having a proximal end and a distal end and defining a cavity adjacent said distal end for receiving a tissue sample; said inner and outer needles slideable relative to one another between a first position wherein said distal end of said inner needle is retracted within said lumen so that said outer needle covers said cavity and a second position wherein said distal end of said inner needle is extended away from said outer needle to expose said cavity; a first hub attached to said proximal end of said inner needle; a second hub attached to said proximal end of said outer needle, said second hub having a track running parallel to said outer needle and a stop member projecting from a surface of said second hub adjacent said track, wherein said track is defined by said surface, a wall projecting from said surface and a dog projecting from said wall; and an elongated pin having a first end, a second end and an arm projecting from said second end, said first end fixedly attached to said first hub with said shaft parallel to said inner needle and said pin slidingly engaged to said second hub within said track, said second hub slideable along said pin between a location near said first end and said second end where said arm meets said stop member. 2. The biopsy instrument of claim 1 wherein said second hub is slideable along said pin a distance sufficient to enable operation of the biopsy instrument in a biopsy gun. 3. The biopsy instrument of claim 1 wherein said second hub is slideable along said pin a predetermined distance to enable insertion of said biopsy needle in a biopsy gun. 4. The biopsy instrument of claim 1 further comprising a ramp projecting from said surface at a location distal to said stop member, said ramp inclined towards said stop member. 5. The biopsy instrument of claim 4 wherein said ramp includes a ramp head facing said stop member, said ramp head and said stop member defining a space, said space having a length that is at least as large a cross sectional dimension of said arm. 6. The biopsy instrument of claim 5 wherein said ramp has a height sufficient to prevent said arm from backing over said ramp as the biopsy instrument is inserted into the biopsy gun when said arm is positioned in said space. 7. The biopsy instrument of claim 6 wherein said ramp has a height of at least about one third of a height of said arm. 8. The biopsy instrument of claim 6 wherein said ramp has a height sufficient to provide a tactile indication when said arm slides over ramp and into said space. 9. The biopsy instrument of claim 6 wherein said track is defined by said surface, a wall projecting from said surface and a dog projecting from said wall. 10. A biopsy instrument for use in a biopsy gun, comprising: a hollow... | Promex Technologies, LLC | A61B005/00 B65D081/00
US06934575 | B2 | 2005-08-23 | Position tracking and imaging system for use in medical applications | 2002-09-03 | A method of monitoring a position of a medical instrument with respect to a patient's body, including positioning a patient reference unit on a portion of the patient's body; positioning a remote unit with respect to the medical instrument; associating at least one field sensor with at least one of the patient reference unit and the remote unit; generating an electromagnetic position characteristic field; producing position data responsive to a sensor output signal that is responsive to the presence of the electromagnetic position characteristic field such that the sensor output signal is representative of a position of the remote unit with respect to the patient reference unit; determining image data representing a plurality of images of the patient's body; and displaying an image of the patient's body based on the image data and responsive to position image data that varies for different position data. | 1. A method of monitoring a position of a medical instrument with respect to a patient's body and for displaying images of the patient's body responsive to the position of the medical instrument, said method comprising: securing a patient reference unit from movement with respect to the patient's body while permitting the patient's body to be moved within an operating environment; securing a remote unit from movement with respect to the medical instrument; positioning an operating environment reference unit within the operating environment; associating at least one field generator with at least one of the patient reference unit, the remote unit and the operating environment reference unit; associating at least one field sensor with at least one of the patient reference unit, the remote unit and the operating environment reference unit; generating a position characteristic field with the at least one field generator; producing position data responsive to a sensor output signal produced by the at least one field sensor that is responsive to the presence of the position characteristic field such that the sensor output signal is representative of a position of the remote unit with respect to the patient reference unit; obtaining image data representing a plurality of images of the patient's body; and displaying an image of the patient's body based on the image data and responsive to position image data that varies for different position data. | 1. A method of monitoring a position of a medical instrument with respect to a patient's body and for displaying images of the patient's body responsive to the position of the medical instrument, said method comprising: securing a patient reference unit from movement with respect to the patient's body while permitting the patient's body to be moved within an operating environment; securing a remote unit from movement with respect to the medical instrument; positioning an operating environment reference unit within the operating environment; associating at least one field generator with at least one of the patient reference unit, the remote unit and the operating environment reference unit; associating at least one field sensor with at least one of the patient reference unit, the remote unit and the operating environment reference unit; generating a position characteristic field with the at least one field generator; producing position data responsive to a sensor output signal produced by the at least one field sensor that is responsive to the presence of the position characteristic field such that the sensor output signal is representative of a position of the remote unit with respect to the patient reference unit; obtaining image data representing a plurality of images of the patient's body; and displaying an image of the patient's body based on the image data and responsive to position image data that varies for different position data. 2. The method of claim 1, wherein said associating at least one field generator step comprises attaching at least one field generator to the patient reference unit; and wherein said associating at least one field sensor step comprises attaching at least one field sensor to the remote unit. 3. The method of claim 1, wherein said associating at least one field generator step comprises attaching at least one field generator to the operating environment reference unit; and wherein said associating at least one field sensor step comprises attaching at least one field sensor to the remote unit. 4. The method of claim 1, further comprising positioning a reference sensor a fixed distance from at least one of the at least one field generator and the at least one field sensor. 5. The method of claim 1, further comprising generating a reference output signal for use in monitoring the position characteristic field. 6. The method of claim 1, wherein said generating a position characteristic field comprises generating a position characteristic field that is a non-linear field that emanates from substantially a singular point. 7. The method of claim 1, wherein said generating a position characteristic field comprises generating a position characteristic filed that is at least two time multiplexed fields that emanate from substantially the same point. 8. The method of claim 1, further comprising registering a spatial orientation of prerecorded scan images with a spatial orientation of the position characteristic field. ... | GE Medical Systems Global Technology Company, LLC | A61B005/00
US06968225 | B2 | 2005-11-22 | Real-time localization, monitoring, triggering and acquisition of 3D MRI | 2001-08-24 | A technique embodied in a method and apparatus is disclosed for acquiring MR images in 2D and 3D using a common pulse sequence switchable in real-time. The technique includes identifying a 3D imaging volume and applying a 3D pulse sequence having the slice encoding and rewinder gradients disabled in one dimension. The technique then allows localizing and monitoring by acquiring 2D MR data. When it is desirable to enter a 3D imaging mode, the slice encoding and rewinder gradients are re-enabled in the common pulse sequence for acquisition of 3D MR data. | 1. A method of MR data acquisition comprising: prescribing a 3D imaging volume;applying a pulse sequence that is applicable as a 3D pulse sequence with slice encoding and rewinder gradients disabled in one dimension;acquiring 2D MR data to localize the 3D imaging volume;enabling the disabled encoding and rewinder gradients in the pulse sequence;applying the pulse sequence in three dimensions; andacquiring 3D MR data of the 3D imaging volume. | 1. A method of MR data acquisition comprising: prescribing a 3D imaging volume;applying a pulse sequence that is applicable as a 3D pulse sequence with slice encoding and rewinder gradients disabled in one dimension;acquiring 2D MR data to localize the 3D imaging volume;enabling the disabled encoding and rewinder gradients in the pulse sequence;applying the pulse sequence in three dimensions; andacquiring 3D MR data of the 3D imaging volume. 2. The method of claim 1 further comprising the step of modifying the pulse sequence between a 2D pulse sequence and a 3D pulse sequence in real-time and on-the-fly. 3. The method of claim 1 further comprising the step of allowing adjustment of at least one of an FOV, a slice thickness, flip angle, matrix size, sampling bandwidth, and spatial saturation between real-time data acquisitions. 4. The method of claim 3 further comprising the step of acquiring full k-space data for one MR data acquisition after an adjustment, and then acquiring partial k-space data thereafter until a subsequent adjustment. 5. The method of claim 1 wherein the steps of acquiring 2D MR data is sped up by first acquiring one set of complete k-space data, and then acquiring a subset of k-space data thereafter. 6. The method of claim 1 further comprising the step of detecting object movement during real-time data acquisition and if object movement is detected, acquiring full k-space data for at least one MR data acquisition and acquiring partial k-space data thereafter. 7. The method of claim 6 wherein the step of detecting is performed by one of an automated detection using a navigator echo technique and manual observation of real-time generated images. 8. The method of claim 1 further comprising the step of switching from 2D MR data acquisition to 3D data acquisition within one repetition time (TR). 9. The method of claim 1 as used in an MRA exam and further comprising: injecting a contrast agent;continuously applying the pulse sequence, acquiring 2D MR data, and displaying images in real-time until an arrival of the contrast agent in a monitor station; and then,switching the pulse sequence to 3D acquisition and acquiring 3D MR data for the prescribed 3D imaging volume. 10. The method of claim 9 further comprising the step of adaptively switching between a 2D monitor mode and a 3D acquisition mode for each of a number of prescribed 3D imaging volumes. 11. An MRI apparatus to acquire MR images and switch between 2D and 3D image acquisition in real-time comprising: a magnetic resonance imaging (MRI) system having a plurality of gradient coils positioned about a bore of a magnet to impress a polarizing magnetic field and an RF transceiver system and an RF switch controlled by a pulse module to transmit RF signals to an RF coil assembly to acquire MR images; anda computer programmed to: modify a selected pulse sequence upon demand between a 2D pulse sequence and a 3D pulse sequence;apply the selected pulse sequence and acquire M... | General Electric Company | A61B005/055
US06901278 | B1 | 2005-05-31 | Methods for reducing the risk of breast cancer in and improving the health of women | 2003-01-29 | The subject invention provides methods for evaluating women in designing appropriate menopausal and postmenopausal therapies to match the particular needs of the individual. In one embodiment, the subject invention provides methods for evaluating particular factors in menopausal and postmenopausal women that are then taken into account in identifying a therapy program to promote good health and reduce the risk of any adverse health effects as a result of the therapy, specifically the role of hormone induced/associated breast cancer. More specifically, the subject invention includes methods for evaluating menopausal and postmenopausal women to determine whether hormone replacement therapy is appropriate in this regard, and if so, the specifics of how that therapy is to be administered. The subject invention, by incorporating various factors into a comprehensive, function breast hormone profile, allows for the identification of untreated women with estrogen-related metabolic abnormalities that may increase the risk of breast cancer. The subject invention determines specific therapies to correct and/or inhibit the aberrant biosynthetic pathways, and thus, the risk of cancer. | 1. A method for providing an individualized therapy for a menopausal or postmenopausal patient, wherein said method comprises conducting a diagnostic mammogram; genotyping at least one enzyme system that regulates hormone biosynthesis and metabolism; evaluating breast endocrine metabolism; and based on the results of these tests, providing the individualized therapy to the patient. | 1. A method for providing an individualized therapy for a menopausal or postmenopausal patient, wherein said method comprises conducting a diagnostic mammogram; genotyping at least one enzyme system that regulates hormone biosynthesis and metabolism; evaluating breast endocrine metabolism; and based on the results of these tests, providing the individualized therapy to the patient. 2. The method according to claim 1, wherein the individualized therapy is a hormonal replacement therapy that treats conditions associated with menopause and postmenopause. 3. The method according to claim 1, wherein the individualized therapy is a hormonal replacement therapy that reduces the likelihood of developing future adverse health effects in the patient. 4. The method according to claim 1, wherein the step of evaluating breast endocrine metabolism comprises conducting a test for at least one bio-active hormone. 5. The method according to claim 4, wherein the bio-active hormone is estradiol. 6. The method according to claim 5, further comprising correlating the results of the mammogram with the results of the test for estradiol. 7. The method according to claim 6, wherein when the mammogram results are normal and the estradiol serum level is less than 5 pg/mL, the individualized therapy for the patient comprises low-dose, oral or transdermal administration of hormone replacement therapy. 8. The method according to claim 6, wherein when the results of the mammogram are highly dense tissue and the estradiol serum level is less than 5 pg/mL, the individualized therapy for the patient comprises administering selective estrogen receptor modulators. 9. The method according to claim 8, further comprising the step of administering low-dose, oral or transdermal hormone replacement therapy in combination with selective estrogen receptor modulators. 10. The method according to claim 8, further comprising the step of conducting a breast metabolism biopsy. 11. The method according to claim 4, further comprising conducting a test for at least one pro-hormone when the bio-active hormone level is too high. 12. The method according to claim 11, wherein the bio-active hormone is estradiol and the pro-hormone is estrone. 13. The method according to claim 12, wherein when the estradiol serum level is greater than or equal to 5 pg/ml and the estrone level is normal, the individualized therapy for the patient comprises administering selective estrogen receptor modulators to the patient. 14. The method according to claim 13, further comprising the step of conducting a breast metabolism biopsy. 15. The method according to claim 12, wherein when the estradiol serum level is greater than or equal to 5 pg/ml and the estrone level is higher than normal, an estrone sulfate level is assessed. 16. The method according to claim 15, wherein when the estrone sulfate level is above normal levels, the individualized therapy for the patient comprises administering sulfatase inhibit... | - | A61B005/00
US06944489 | B2 | 2005-09-13 | Method and apparatus for shunting induced currents in an electrical lead | 2002-01-29 | An electrical lead includes an elongate body having a proximal end portion and a distal end portion, a first electrode disposed adjacent and joined to the distal end portion of the elongate body. Current flow within the first electrode is limited when a predetermined condition occurs, such as the generation of an electromagnetic field having a predetermined frequency range. The medical electrical lead may further comprise one or more second electrodes disposed adjacent the first electrode and joined to the elongate body to shunt current to body tissue when the predetermined condition occurs. | 1. An electrical lead, comprising: an elongate body having a proximal end portion and a distal end portion; a first electrode disposed near the distal end portion of the elongate body; a first conductor electrically coupled to the first electrode; a second electrode disposed adjacent the elongate body; a capacitive device electrically coupled to the first conductor and the second electrode; and a component coupled to the first conductor to selectively limit current through the first electrode. | 1. An electrical lead, comprising: an elongate body having a proximal end portion and a distal end portion; a first electrode disposed near the distal end portion of the elongate body; a first conductor electrically coupled to the first electrode; a second electrode disposed adjacent the elongate body; a capacitive device electrically coupled to the first conductor and the second electrode; and a component coupled to the first conductor to selectively limit current through the first electrode. 2. The lead of claim 1, wherein the component is an inductor. 3. The lead of claim 1, wherein the component is a switch. 4. The lead of claim 3, wherein the component is coupled to a sensor capable of detecting an electromagnetic field. 5. The lead of claim 4, wherein the sensor is capable of detecting an electromagnetic field in a predetermined frequency range. 6. The lead of claim 5, wherein the switch includes active circuit components. 7. The lead of claim 5, wherein the switch includes Field Effect Transistors (FETs). 8. The lead of claim 5, wherein the switch includes at least one Micro-Electrical-Mechanical System (MEMs) switch. 9. The lead according to claim 1, further comprising a second conductor extending between the proximal end portion of the elongate body and the second electrode end being electrically coupled to the second electrode. 10. The lead according to claim 1, further comprising: a third electrode disposed adjacent the second electrode and joined to the elongate body; and a second conductor extending between the proximal end portion of the elongate body and the third electrode and being electrically coupled to the third electrode. 11. The lead according to claim 1, wherein the capacitive device further comprises a discoidal capacitor having a central contact and a peripheral contact; the first conductor is electrically coupled with the central contact; and the second electrode is electrically coupled with the peripheral contact. 12. The lead according to claim 1, wherein: the first electrode and the second electrode are capable of being electrically coupled with body tissue; the first conductor is capable of transmitting an electrical current; and the capacitive device has an electrical impedance substantially less than electrical impedance of the body tissue at signal frequencies greater than those used to sense biological signals or deliver electrical stimulation to the body tissue. 13. The lead according to claim 1, wherein: the first electrode and the second electrode are capable of being electrically coupled with body tissue; the first conductor is capable of transmitting an electrical current; and the capacitive device has an electrical impedance substantially greater than an electrical impedance of the body tissue at signal frequencies used to sense biological signals or deliver electrical stimulation to the body tissue. 14. The lead according to claim 1, wherein an area of the second electrode is at least thr... | Medtronic, Inc. | A61B005/04
US06905472 | B2 | 2005-06-14 | Incompetent cervix aide | 2002-11-13 | Two embodiments of an annular cervical ring have an internal diameter 14 for surrounding the cervix 12. Each embodiment also includes a tightening device 16 for decreasing the internal diameter 14 to tighten the annular ring down around the cervix 12 to hold the cervix 12 closed. | 1. A method for holding the cervix closed comprising the steps of: placing an internal diameter of an annular cervical ring around the cervix in the vagina of a woman, decreasing the internal diameter of the cervical ring by telescoping a first portion of the cervical ring down into a second portion of the cervical ring to decrease the internal diameter and tighten the cervical ring around the cervix, and holding the cervix closed for an extended period of time. | 1. A method for holding the cervix closed comprising the steps of: placing an internal diameter of an annular cervical ring around the cervix in the vagina of a woman, decreasing the internal diameter of the cervical ring by telescoping a first portion of the cervical ring down into a second portion of the cervical ring to decrease the internal diameter and tighten the cervical ring around the cervix, and holding the cervix closed for an extended period of time. 2. A method as set forth in claim 1 including ratcheting said first portion into said second portion and mechanically holding said portions together. | - | A61B005/103 A61B005/117
US06881192 | B1 | 2005-04-19 | Measurement of sleep apnea duration and evaluation of response therapies using duration metrics | 2002-06-12 | An implantable cardiac device is programmed to detect an episode of sleep apnea, measure the duration of the episode, and store this information in memory. When multiple episodes are recorded, the device computes statistics on the apnea durations, such as average apnea duration and total apnea duration for a preselected time period (e.g., 8-hour rest period, 24-hour day, etc.). The implantable cardiac device may further be used to administer pacing therapy in response to detecting sleep apnea. Under the control of a physician, the implantable device may be programmed to administer different types of responsive therapies to evaluate whether certain therapies are more effective at treating apnea than others. For instance, a pacing therapy that results in lowering the average apnea duration or total apnea duration may be preferred over other pacing therapies that do not achieve such results. | 1. An implantable cardiac device comprising: sensing circuitry to sense whether a patient is at rest; a sleep apnea detector to detect when a patient, who is at rest, is experiencing an episode of sleep apnea and to measure a duration of the episode of sleep apnea; a therapy module to prescribe a pacing therapy for treating sleep apnea from among multiple different pacing therapies; a pulse generator to generate pacing pulses according to the prescribed pacing therapy for an evaluation timeframe; and the sleep apnea detector measuring durations of episodes experienced during the evaluation timeframe of the prescribed pacing therapy. | 1. An implantable cardiac device comprising: sensing circuitry to sense whether a patient is at rest; a sleep apnea detector to detect when a patient, who is at rest, is experiencing an episode of sleep apnea and to measure a duration of the episode of sleep apnea; a therapy module to prescribe a pacing therapy for treating sleep apnea from among multiple different pacing therapies; a pulse generator to generate pacing pulses according to the prescribed pacing therapy for an evaluation timeframe; and the sleep apnea detector measuring durations of episodes experienced during the evaluation timeframe of the prescribed pacing therapy. 2. An implantable cardiac device as recited in claim 1, wherein the sensing circuitry comprises an activity sensor to sense a patient's activity to determine whether the patient is at rest. 3. An implantable cardiac device as recited in claim 1, wherein the sensing circuitry comprises a position sensor to sense a patient's position as a way to determine whether the patient is at rest. 4. An implantable cardiac device as recited in claim 1, wherein the sensing circuitry comprises an accelerometer that is used to determine whether the patient is at rest. 5. An implantable cardiac device as recited in claim 1, wherein the sensing circuitry is configured to sense a respiration-related signal and the sleep apnea detector detects the episode of sleep apnea based upon the respiration-related signal. 6. An implantable cardiac device as recited in claim 5, wherein the sensing circuitry comprises an impedance sensor to detect the respiration-related signal. 7. An implantable cardiac device as recited in claim 5, wherein the respiration-related signal is selected from a group of signals comprising a signal indicative of tidal volume, a signal indicative of respiration rate, and a signal indicative of minute ventilation, and O2 saturation. 8. An implantable cardiac device as recited in claim 5, wherein the sleep apnea detector is configured to detect the episode of sleep apnea by determining when the respiration-related signal exceeds a first threshold indicative of hyperventilation and then falls below a lower second threshold for a predetermined time period. 9. An implantable cardiac device as recited in claim 1, wherein the sensing circuitry is configured to sense an oxygen-related signal and the sleep apnea detector detects the episode of sleep apnea based upon the oxygen-related signal. 10. An implantable cardiac device as recited in claim 1, further comprising a memory to store apnea data pertaining to the episode of sleep apnea and the duration of the episode. 11. An implantable cardiac device as recited in claim 1, further comprising a statistical module to derive statistics on the duration of the sleep apnea episode. 12. An implantable cardiac device as recited in claim 1, further comprising a pulse generator to generate pacing pulses applied in response to the sleep apnea detector detecting the episode of sl... | Pacesetter, Inc. | A61B005/08
US06965794 | B2 | 2005-11-15 | Apparatus for routing electromyography signals | 2002-04-10 | An EMG monitoring system includes a matrix of detection electrode arrays that is positioned on the patient along his or her spine. The arrays are electrically connected through a switching mechanism to EMG amplifiers that are included in conventional monitoring instrumentation. A switch controller operates the switches to provide signals to the EMG amplifiers from selected sets of detection electrode arrays. The controller controls the switches to provide signals simultaneously to each EMG amplifier. If the monitoring instrumentation includes four EMG amplifiers, the matrix of electrode arrays and the switching mechanism may be used to take simultaneous measurements at two different spinal levels. If additional EMG amplifiers are included, simultaneous measurements at additional spinal levels may also be made. The matrix may also include redundant arrays and/or electrodes, such that the clinician can select the sets of electrodes that conform to the size of the patient. | 1. A system for routing electromyography signals to one or more electromyography amplifiers, the system including: A. a matrix for placement along a spine, the matrix including a plurality of electrodes arranged in sets of parallel arrays with the arrays arranged on either side of the spine when the matrix is in place along the spine the arrays being spaced vertically to correspond to the size of a relatively small patient;B. a switching mechanism that electrically connects selected electrode arrays to the amplifiers; andC. a switch controller that operates the switching mechanism to control the selection of the arrays, the switch controller directing the switching mechanism to connect to the amplifiers selected arrays that are spaced vertically to correspond to the size of the patient on which the matrix is arranged. | 1. A system for routing electromyography signals to one or more electromyography amplifiers, the system including: A. a matrix for placement along a spine, the matrix including a plurality of electrodes arranged in sets of parallel arrays with the arrays arranged on either side of the spine when the matrix is in place along the spine the arrays being spaced vertically to correspond to the size of a relatively small patient;B. a switching mechanism that electrically connects selected electrode arrays to the amplifiers; andC. a switch controller that operates the switching mechanism to control the selection of the arrays, the switch controller directing the switching mechanism to connect to the amplifiers selected arrays that are spaced vertically to correspond to the size of the patient on which the matrix is arranged. 2. The system of claim 1 wherein the matrix is arranged in two adhesive strips that are placed on opposite sides of the spine when the system is in use. 3. The system of claim 1 wherein the matrix is housed in a re-usable pad. 4. The system of claim 3 further including a two-sided adhesive sheet that holds the pad in place along the spine when the system is in use. 5. The system of claim 4 wherein the adhesive sheet includes holes that correspond to the locations of the electrodes in the matrix. 6. The system of claim 3 wherein the pad further houses the switching mechanism. 7. A system for routing electromyography signals to one or more electromyography amplifiers, the system including: A. a matrix for placement along a spine, the matrix including a plurality of electrodes arranged in arrays with certain or all of the respective arrays further including one or more redundant electrodes that are spaced horizontally to correspond to various sizes of patients;B. a switching mechanism that electrically connects selected electrode arrays to the amplifiers; andC. a switch controller that operates the switching mechanism to control the selection of the arrays, the switch controller directing the switching mechanism to connect to the amplifiers, from selected respective arrays, selected electrodes and redundant electrodes that are spaced vertically to correspond to the size of the patient on which the matrix is arranged. 8. A system for routing electromyography signals to one or more electromyography amplifiers, the system including: A. a matrix for placement along a spine, the matrix including a plurality of electrodes arranged in arrays;B. a switching mechanism that electrically connects selected electrode arrays to the amplifiers; andC. a switch controller that operates the switching mechanism to control the selection of the arrays the switch controller being programmed to select arrays in accordance with various protocols and further programmed to select arrays in accordance with the size of the patient. 9. A system for routing electromyography signals to one or more electromyography amplifiers, the system including: A. a matri... | Fasstech, Inc. | A61B005/04
US06961607 | B2 | 2005-11-01 | Method for assessing myocardial angiogenesis | 2002-07-31 | Polymeric contrast agents are used to monitor myocardial neovascularization. The agent is injected and images of the myocardium are obtained at various time intervals after injection. Localized image enhancement indicates an area of increased angiogenesis. | 1. A method for monitoring myocardial angiogenesis, the method comprising: intravenously administering a polymeric contrast agent to a subject, the polymeric contrast agent having a conformation with a length that is 5 to 500 times greater than the average diameter of the polymeric contrast agent;obtaining at least two images of the myocardium of the subject; andcomparing the images to assess the growth of new blood vessels, wherein a localized image enhancement indicates an area of increased myocardial angiogenesis. | 1. A method for monitoring myocardial angiogenesis, the method comprising: intravenously administering a polymeric contrast agent to a subject, the polymeric contrast agent having a conformation with a length that is 5 to 500 times greater than the average diameter of the polymeric contrast agent;obtaining at least two images of the myocardium of the subject; andcomparing the images to assess the growth of new blood vessels, wherein a localized image enhancement indicates an area of increased myocardial angiogenesis. 2. A method as in claim 1 wherein the step of intravenously administering a polymeric contrast agent comprises administering a contrast agent having a polypeptide backbone. 3. A method as in claim 1 wherein the step of intravenously administering a polymeric contrast agent comprises administering a contrast agent having a backbone formed from a polypeptide selected fro the group consisting of polylysine, polyglutamic acid, polyaspartic acid and copolymers of lysine and either glutamic acid or aspartic acid. 4. A method as in claim 1 wherein the step of intravenously administering a polymeric contrast agent comprises administering a contrast agent having a polymer backbone having covalently bound thereto at least one member from the group consisting of diethylenetriaminepentaacetic acid (DTPA), 1,4,7,10-Tetraazacyclododecane-1,4,7,10-tetraacetic acid (DOTA), 1,4,7,10-Tetraazacyclododecane-1,4,7,10-tetrakis(2-propionic acid) (DOTMA), 1,4,8,11-tetraazacyclotetradecane-1,4,8,11 -tetraacetic acid (TETA), 1,4,7,10-tetraazacyclododecane-1,4,7,10-tetrakis(3-(4-carboxyl)-butanoic acid), 1,4,7,10-tetraazacyclododecane-1,4,7,10-tetrakis(acetic acid-methyl amide), 1,4,7,10-tetraazacyclododecane-1,4,7,10-tetrakis(methylene phosphonic acid), and p-isothiocyanatobenzyl-1,4,7,10-tetraazacyclododecane-1,4,7,10-tetraacetic acid (p-SCN-Bz-DOTA), bis(thiosemicarbazone), derivatives of bis(thiosemicarbazone), porphyrins, derivatives of porphyrins, 2,3-bis(2-thioacetamido)propionates, derivatives of 2,3-bis(2-thioacetamido)propionates, N,N′-bis(mercaptoacetyl)-2,3-diaminopropanoate, bis(aminoethanethiol) and derivatives of bis(aminoethanethiol). 5. A method as in claim 1 wherein the step of intravenously administering a polymeric contrast agent comprises administering a contrast agent at a dose in the range of about 0.01 mmoles Gd/Kg to about 0.1 mmoles Gd/Kg. 6. A method as in claim 1 wherein the step of intravenously administering a polymeric contrast agent comprises administering a contrast agent having a polypeptide backbone having a length of 80 to 1500 amino acid residues. 7. A method as in claim 1 wherein the step of intravenously administering a polymeric contrast agent comprises administering a contrast agent having a diameter in the range of 20 Angstroms to 50 Angstroms. 8. A method as in claim 1 wherein the step of intravenously administering a polymeric contrast agent comprises administering a contrast agent having one or more paramagne... | - | A61B005/05 A61B005/055
US06928313 | B2 | 2005-08-09 | System and method for accessing the coronary sinus to facilitate insertion of pacing leads | 2003-01-27 | A guiding catheter system employs an outer jacket with one or more inflation lumens provided along the length of the jacket. The lumens are in fluid communication with one or more fluted balloons mounted at a distal end of the outer jacket. A torquable stylet is movably disposed within an open lumen of the outer jacket. A guidewire is movably disposed in an open lumen of the torquable stylet. The outer jacket can be provided with break-away features allowing the jacket to be peeled off of a pacing lead. | 1. A guiding catheter system for accessing a patient's heart, comprising: an outer jacket, comprising: an open guide lumen extending from a proximal tip of the outer jacket to a distal tip of the outer jacket; at least one inflation lumen provided along an exterior surface of the outer jacket; and at least one segment of a distal balloon fixably mounted at a distal portion of the outer jacket, the distal balloon in fluid communication with the at least one inflation lumen; and a stylet disposed within the guide lumen of the outer jacket and comprising a distal pre-formed bend, the stylet rotatable within the outer jacket along a longitudinal axis of the outer jacket, the distal pre-formed bend of the stylet dimensioned to deflect the outer jacket laterally of the longitudinal axis of the outer jacket upon rotation of the stylet. | 1. A guiding catheter system for accessing a patient's heart, comprising: an outer jacket, comprising: an open guide lumen extending from a proximal tip of the outer jacket to a distal tip of the outer jacket; at least one inflation lumen provided along an exterior surface of the outer jacket; and at least one segment of a distal balloon fixably mounted at a distal portion of the outer jacket, the distal balloon in fluid communication with the at least one inflation lumen; and a stylet disposed within the guide lumen of the outer jacket and comprising a distal pre-formed bend, the stylet rotatable within the outer jacket along a longitudinal axis of the outer jacket, the distal pre-formed bend of the stylet dimensioned to deflect the outer jacket laterally of the longitudinal axis of the outer jacket upon rotation of the stylet. 2. The guiding catheter system of claim 1, wherein the outer jacket further comprises a peel away feature provided along a longitudinal length of the outer jacket. 3. The guiding catheter system of claim 2, wherein the peel away feature comprises an undercut between the at least one inflation lumen and the exterior surface of the outer jacket. 4. The guiding catheter system of claim 2, wherein the peel away feature comprises a thickened section at an interface between the at least one inflation lumen and the exterior surface of the outer jacket. 5. The guiding catheter system of claim 1, wherein the stylet comprises an open lumen, the guiding catheter system further comprising a guidewire movably disposed within the open lumen of the stylet. 6. The guiding catheter system of claim 5, wherein the guidewire comprises a loop at a distal tip of the guidewire. 7. The guiding catheter system of claim 5, wherein the guidewire is formed from a material such that a distal tip of the guidewire is substantially straight at ambient temperature and the distal tip assumes a shape of a loop at body temperature. 8. The guiding catheter system of claim 5, wherein the guidewire comprises a sensor at a distal tip of the guidewire. 9. The guiding catheter system of claim 8, wherein the sensor comprises an electrode. 10. The guiding catheter system of claim 1, wherein the open guide lumen of the outer jacket is dimensioned to receive a medical electrical lead. 11. The guiding catheter system of claim 1, further comprising one or more secondary balloons integrated into the inflation lumen, the secondary balloons proximal to the distal balloon segment. 12. A method of cannulating a destination blood vessel, comprising: introducing an outer jacket and a stylet of a catheter system into an access vessel, the outer jacket comprising a guide lumen extending from a proximal tip of the outer jacket to a distal tip of the outer jacket; rotating the stylet within the guide lumen of the outer jacket so that a pre-formed bend on a distal end of the stylet deflects the outer jacket to steer a distal end of the outer jacket; advancing a guidew... | Cardiac Pacemakers, Inc. | A61B005/04
US06904313 | B1 | 2005-06-07 | Methods and systems for monitoring heart rate recovery | 2002-05-23 | Methods and systems for monitoring heart rate recovery are described. In one embodiment, an implantable monitoring device monitors a patient's heart activity. The monitor looks for periods of increased activity, such as those that are experienced during some level of exercise, and periods of lesser activity (such as rest) directly following. The patient's heart rate recovery is monitored as the heart recovers from the active period to the rest period. Data associated with the patient's heart rate recovery can then be diagnostically used to ascertain certain conditions associated with the patient. | 1. An implantable monitoring device comprising: one or more processors; one or more sensors operably coupled with the one or more processors and configured to sense a patient's heart beat; and memory coupled to the one or more processors, wherein the memory stores instructions which, when executed by the one or more processors, cause the one or more processors to: determine a patients heart rate from heart beats sensed by the one or more sensors; detect an active period associated with a patient's physical exertion and characterized by an elevated heart rate; detect termination of the active period as characterized by a heart rate that is slower than the elevated heart rate; monitor the patient's heart rate recovery as the patient's heart recovers from a generally elevated rate associated with the active period, to a slower rate that is not associated with the active period; and ascertain whether the monitored heart rate recovery is normal for that patient. | 1. An implantable monitoring device comprising: one or more processors; one or more sensors operably coupled with the one or more processors and configured to sense a patient's heart beat; and memory coupled to the one or more processors, wherein the memory stores instructions which, when executed by the one or more processors, cause the one or more processors to: determine a patients heart rate from heart beats sensed by the one or more sensors; detect an active period associated with a patient's physical exertion and characterized by an elevated heart rate; detect termination of the active period as characterized by a heart rate that is slower than the elevated heart rate; monitor the patient's heart rate recovery as the patient's heart recovers from a generally elevated rate associated with the active period, to a slower rate that is not associated with the active period; and ascertain whether the monitored heart rate recovery is normal for that patient. 2. The implantable monitoring device of claim 1, wherein the instructions cause the one or more processors to generate a notification in an event that the monitored heart rate recovery is not normal. 3. The implantable monitoring device of claim 2, wherein the instructions cause the one or more processors to effect transmission of the notification to an external device. 4. The implantable monitoring device of claim 1, wherein the instructions cause the one or more processors to save data associated with the monitored heart rate recovery on the one or more computer-readable media. 5. An implantable device comprising: one or more processors; one or more sensors operably coupled with the one or more processors and configured to sense a patients heart beat; one or more activity sensors operably coupled with the one or more processors and configured to sense when a patient is active and inactive; memory coupled to the one or more processors, wherein the memory stores instructions which, when executed by the one or more processors, cause the one or more processors to: determine a patient's heart rate from heart beats sensed by the one or more sensors; determine when a patient is active and inactive using the one or more activity sensors; define a normal heart rate recovery for a patient that pertains to how the patient's heart recovers from an elevated rate associated with an active period when the patient is active, to a lower rate associated with an inactive period when the patient is less active; monitor the patient's heart rate recovery from an active period to a less active period; ascertain whether the monitored heart rate recovery is trending away from the normal heart rate recovery; and if the monitored heart rate recovery is trending away from the normal heart rate recovery, take a remedial action associated with remedying the monitored heart rate recovery. 6. The implantable device of claim 5, wherein the instructions cause the one or more processors to take a remedial action that d... | Pacesetter, Inc. | A61B005/04
US06837857 | B2 | 2005-01-04 | Method for the recording of acoustic parameters for the customization of hearing aids | 2002-07-29 | Acoustic parameters of an ear are recorded by measuring the impedance of the ear canal. A first pressure microphone measures the sonic pressure. A second pressure microphone measures the sonic pressure through a calibrated acoustic resistance. The sound energy flux is determined by the difference between the measured sonic pressures. The impedance is calculated from these results. | 1. Method of measuring an impedance of an ear canal comprising the steps of: measuring a first sonic pressure with a first pressure microphone; determining a sound energy flux by the difference of the measured first sonic pressure in relation to a second sonic pressure measured with a second pressure microphone through a calibrated acoustic resistance; and calculating the impedance from said first and second sonic pressures. | 1. Method of measuring an impedance of an ear canal comprising the steps of: measuring a first sonic pressure with a first pressure microphone; determining a sound energy flux by the difference of the measured first sonic pressure in relation to a second sonic pressure measured with a second pressure microphone through a calibrated acoustic resistance; and calculating the impedance from said first and second sonic pressures. 2. Method according to claim 1 wherein the calibrated acoustic resistance comprises porous sinter bronze arranged concentrically around a channel of the first pressure microphone. 3. Method according to claim 1, wherein the step of measuring is performed by pressure microphones arranged in a hearing aid shell, and wherein the hearing aid shell is introduced into the ear canal before starting the step of measuring. 4. Method according to claim 3 wherein the step of calculating comprises evaluation of signals from said first and second pressure microphones by electronic components arranged in the hearing aid shell. 5. Method according to claim 1 further comprising the step of implementing curve fit proceedings onto the measuring results, to calculate or estimate the following parameters: impedance, and/or effective distance of a hearing aid outlet to an eardrum, and/or compliance of a rest volume of the ear canal, and/or overall compliance at deep frequencies, and/or individual transitional function of an inner ear, and/or effect of leaks. 6. Method according to claim 1 wherein the measuring is made under variable, adjustable static pressure in the ear canal. 7. Method according claim 6 wherein the static pressure is adjusted over an additional static pressure line for the ear canal. 8. A measuring device for performing a method of measuring an impedance of an ear canal comprising: a first pressure microphone measuring first sonic pressure; a second pressure microphone for measuring a second sonic pressure through a calibrated acoustic resistance; means for determining a sound energy flux by the difference of the measured first sonic pressure in relation to the measured second sonic pressure; means for calculating the impedance from said first and second sonic pressures; and a probe adapted to be at least partly introduced into the ear canal, wherein said first and second microphones are arranged within the probe, and wherein an acoustic resistance is pre-connected to at least one of said first and second microphones. 9. Measuring device according to claim 8 wherein the calibrated acoustic resistance comprises a porous sinter bronze. 10. Measuring device according to claim 8 arranged within a shell of a hearing aid adapted to a shape of the ear canal, and comprising evaluation electronics arranged within the shell and connected to said first and second microphones of the measuring device. | Phonak AG | A61B005/00
US06904304 | B2 | 2005-06-07 | CT detector reflector useful in detector scintillator array | 2002-05-01 | A method for making a detector array comprising a photosensor and scintillators, including placing a thermoplastically encased reflective film between scintillators of a scintillator array and optically coupling the scintillators with the photosensor. | 1. A method for fabricating a photodetector array including a photosensor array and scintillators, said method comprising the steps of: placing a thermoplastically encased reflective film between and in front of the scintillators of a scintillator array; and optically coupling the scintillators with the photosensor array. | 1. A method for fabricating a photodetector array including a photosensor array and scintillators, said method comprising the steps of: placing a thermoplastically encased reflective film between and in front of the scintillators of a scintillator array; and optically coupling the scintillators with the photosensor array. 2. A method in accordance with claim 1 wherein the encased reflective film is prepared by positioning a reflective film within a thermoplastic film and affixing said thermoplastic film to the reflective film. 3. A method in accordance with claim 2 wherein said step of affixing comprises laminating the thermoplastic material to the reflective film. 4. A method in accordance with claim 2 wherein the reflective film is encased in a process comprising depositing the reflective film directly on a thermoplastic sheet and laminating it with another layer of thermoplastic material. 5. A method in accordance with claim 2 wherein said reflective film is a low x-ray damageable material. 6. A photodetector array comprising: an array of scintillators optically coupled to a photosensor array; and a thermoplastically encased reflective film positioned between and in front of said scintillators of said array of scintillators. 7. The photodetector array of claim 6 wherein said array is employed in a computed tomography imaging system. 8. The photodetector array of claim 7 wherein said reflective film is thermoplastically encased in a process comprising depositing the reflective film directly on a thermoplastic sheet and laminating it with another layer of thermoplastic material. 9. The photodetector array of claim 8 wherein said reflective film is a low x-ray damageable material. 10. A computed tomography imaging system comprising a rotating gantry, a radiation source, and a detector array on the rotating gantry configured to detect attenuated radiation from the radiation source passing through an object between the radiation source and the detector array, the detector array comprising a photosensor array, an array of scintillators, and a thermoplastically encased reflective film within said scintillator array, wherein said photosensor array is optically coupled to said scintiliator array, and said thermoplastically encased reflective film is positioned between and in front of scintillators of said scintillator array. 11. The imaging system of claim 10 wherein the object is a patient. 12. The imaging system of claim 10 wherein said reflective film is encased in a process comprising depositing the reflective film directly on a thermoplastic sheet and laminating it with another layer of thermoplastic material. 13. The imaging system of claim 10 wherein said reflective film is a low x-ray damageable material. | GE Medical Systems Global Technology Company, LLC | A61B005/05 G01F023/00
US06898455 | B2 | 2005-05-24 | Method for providing optimal drug dosage | 2002-06-03 | A method for treating attention deficit/hyperactivity disorder (ADHD) includes performing an objective ADHD behavioral assay for a patient, administering a drug to the patient at varying doses, and performing an fMRI on the patient for the varying drug doses to determine an optimal dosage for the patient. The method can further include performing a further objective ADHD behavioral assay after the drug administration. | 1. A method for treating attention deficit/hyperactivity disorder (ADHD), comprising: performing an objective ADHD behavioral assay for a patient; administering a drug to the patient at varying doses; and performing an fMRI on the patient for the varying drug doses to determine an optimal dosage for the patient. | 1. A method for treating attention deficit/hyperactivity disorder (ADHD), comprising: performing an objective ADHD behavioral assay for a patient; administering a drug to the patient at varying doses; and performing an fMRI on the patient for the varying drug doses to determine an optimal dosage for the patient. 2. The method according to claim 1, further including performing a further objective ADHD behavioral assay after the drug administration. 3. The method according to claim 1, further including performing the objective ADHD behavioral assay using monotonous computerized task. 4. The method according to claim 1, further including performing the objective ADHD behavioral assay using an Optax test. 5. The method according to claim 1, further including administering methylphenidate to the patient. 6. The method according to claim 1, further including performing the fMRI to obtain T2 RT information to assess blood volume changes in a cerebellum of the patient in response to the drug doses. 7. The method according to claim 6, further including selecting a first one of the varying drug doses based upon blood flow levels to the cerebellar vermis of the patient in response to the drug doses. 8. The method according to claim 1, further including performing the fMRI at predetermined times after the administration of the drug. 9. The method according to claim 1, further including screening the patient by evaluating the patient for a predetermined number of ADHD symptoms. 10. The method according to claim 1, further including obtaining a baseline fMRI for the patient. 11. The method according to claim 1, further including administering each of the varying drug doses for about a week. 12. The method according to claim 1, further including plotting a T2 RT change for each of the drug doses from a baseline T2 RT. 13. The method according to claim 12, further including plotting the T2 RT changes versus a baseline objective ADHD test of the patient. 14. The method according to claim 12, further including selecting the optimal one of the varying drug doses corresponding to a largest positive percent change in the T2 RT from baseline. 15. The method according to claim 1, further including identifying a region of interest (ROI) in the fMRI; computing median pixel intensity values in the ROI; and determining a T2 relaxation time from the median pixel intensity values. 16. A method for selecting an optimal drug dosage to treat ADHD, comprising: performing a first objective ADHD test on a patient; performing an fMRI on the patient to obtain a baseline T2 RT in the cerebellum; administering a drug to the patient at a first dosage for a first period of time; performing an fMRI on the patient to obtain T2 RT information including blood flow changes in the cerebellum in response to the first dosage; performing a second objective ADHD test on the patient a known time after the administration of the first dose; administering the drug to the patient at a... | The McLean Hospital Corporation | A61B005/05
US06887208 | B2 | 2005-05-03 | Method and system for analyzing respiratory tract sounds | 2003-01-09 | A method and system for analyzing respiratory tract sounds in an individual. A plurality of N transducers are fixed over the thorax. The ith transducer is fixed at a location xi, and generates a signal P(xi,t) indicative of pressure waves at the location xi, for i=1 to N. A processor receives the signals P(xi,t) and determines an average acoustic energy {tilde over (P)}(x,t1,t2) at at least one position x over a time interval where {tilde over (P)}is determined in an algorithm involving at least one of the signals P(xi,t). | 1. A system for analyzing sounds in at least a portion of an individual's respiratory tract comprising: (a) a plurality of N transducers, each transducer configured to be fixed on a surface of the individual over the thorax, the ith transducer being fixed at a location xi and generating a signal P(xi, t) indicative of pressure waves at the location xi; for i=1 to N; and (b) a processor configured to receive the signals P(xi,t) and determine an average acoustic energy {tilde over (P)}(x, t1, t2) at at least one position x over a time interval from a first time t1 to a second time t2, {tilde over (P)} being determined in an algorithm involving at least one of the signals P(xi,t). | 1. A system for analyzing sounds in at least a portion of an individual's respiratory tract comprising: (a) a plurality of N transducers, each transducer configured to be fixed on a surface of the individual over the thorax, the ith transducer being fixed at a location xi and generating a signal P(xi, t) indicative of pressure waves at the location xi; for i=1 to N; and (b) a processor configured to receive the signals P(xi,t) and determine an average acoustic energy {tilde over (P)}(x, t1, t2) at at least one position x over a time interval from a first time t1 to a second time t2, {tilde over (P)} being determined in an algorithm involving at least one of the signals P(xi,t). 2. The system according to claim 1 further comprising a two-dimensional display device. 3. The system according to claim 2 wherein the processor is further configured to display a representation of the function {tilde over (P)} on the display device. 4. The system according to claim 1 wherein the processor is further configured to compare the average acoustic energy {tilde over (P)} to one or more predetermined functions {tilde over (F)} and determine a function {tilde over (F)}0 from among the functions {tilde over (F)} most similar to {tilde over (P)}. 5. The system according to claim 4 wherein the processor is further configured to make a diagnosis based upon the determined function. 6. The system according to claim 1 wherein the average acoustic energy {tilde over (P)} over a time interval from t1 to t2 is determined at a location xi of a transducer using the algebraic expression: P~⁢(xi,t1,t2)=∫t1t2⁢P2⁢(xi,t)⁢ ⁢ⅆt. 7. The system according to claim 1 wherein the function {tilde over (P)} is determined at one or more locations x in an algorithm comprising: (One) determining an average acoustic energy {tilde over (P)}(xi,t1,t2) over a time interval from t1 to t2 at a plurality of locations xi of transducers; and (Two) determining an average acoustic energy {tilde over (P)}(x,t1,t2) at at least one location x by interpolation of the determined {tilde over (P)}(x,t1,t2). 8. The system according to claim 7 wherein an average acoustic energy {tilde over (P)}(xi,t1, t2) is determined over a time interval from t1 to t2 at a plurality of locations xi of transducers using the algebraic expression: P~⁢(xi,t1,t2)=∫t2t2⁢P2⁢(xi,t)⁢ ⁢ⅆt. 9. The system according to claim 7 wherein an average acoustic energy is determined at at least one location x by interpolation of the determined {tilde over (P)}(xi,t1,t2) using the algebraic expression: P~⁡(x,t1,t2)=∑i=1N⁢P~⁡(xi,t1,t2)⁢g⁡(x,xi,σ)(2)where g(x,xi,σ) is a kernel satisfying ∇2⁢g=∂g∂σ(3)∑i=1N⁢g⁡(x,xi,σ)⁢ ⁢is⁢ ⁢approximately⁢ ⁢equal⁢ ⁢to⁢ ⁢1.(4) 10. The system according to claim 9 wherein g(x,xiσ)is the kernel g(x,xi,σ)= Exp-((x1-xi1⁢σ)22⁢σ)·Exp-((x2-xi2⁢σ)22⁢σ).(5) 11. The system according to claim 1 wherein the processor is configured to determine an average acoustic energy over a plurality of successive time intervals,... | Deepbreeze Ltd. | A61B005/08
US06904309 | B2 | 2005-06-07 | Microdialysis probes and methods of use | 2002-06-21 | The subject invention pertains to novel microdialysis probes having highly advantageous structural features, particularly for clinical applications. A microdialysis probe of the subject invention can be introduced into a biological tissue using insertion means such that the insertion means can the be completely dissociated from the biological tissue and the microdialysis probe. The microdialysis probes of the subject invention can further include guiding means composed of imageable material that may be used to guide the microdialysis probes to a target tissue using the appropriate sensing equipment. A further aspect of the subject invention relates to methods of using novel microdialysis probes to assess the presence or absence of a substance within biological tissue. The subject invention further relates to methods of using novel microdialysis probes to deliver a substance within biological tissue. | 1. A microdialysis probe comprising: a semi-premeable membrane portion; an outer tube, wherein said outer tube has a proximal end, a distal end, an inner surface, and an outer surface, wherein said membrane portion is disposed on said distal end of said outer tube; an inner tube, wherein said inner tube has a proximal end, a distal end, an inner surface, and an outer surface, and wherein said inner tube is arranged within said outer tube; and a guiding means located on said probe in a known spatial relationship with said membrane portion such that imaging said guiding means provides information as to the location of said membrane portion when said probe is inserted into a medium. | 1. A microdialysis probe comprising: a semi-premeable membrane portion; an outer tube, wherein said outer tube has a proximal end, a distal end, an inner surface, and an outer surface, wherein said membrane portion is disposed on said distal end of said outer tube; an inner tube, wherein said inner tube has a proximal end, a distal end, an inner surface, and an outer surface, and wherein said inner tube is arranged within said outer tube; and a guiding means located on said probe in a known spatial relationship with said membrane portion such that imaging said guiding means provides information as to the location of said membrane portion when said probe is inserted into a medium. 2. The microdialysis probe of claim 1, wherein said guiding means comprises an imageable material. 3. The microdialysis probe of claim 1, wherein said guiding means comprises an imageable material having an image signature that is discernable from that of biological tissue. 4. The microdialysis probe of claim 1, wherein said guiding means comprises an echogenic material. 5. The microdialysis probe of claim 1, wherein said guiding means comprises an echogenic material having an acoustic impedance different from that of the surrounding medium into which said guiding means is introduced. 6. The microdialysis probe of claim 1, wherein said guiding means comprises an imageable material selected from the group consisting of aluminum, plastic, sand, and metal particles. 7. The microdialysis probe of claim 1, wherein said guiding means comprises an imageable material that is radio-opaque. 8. The microdialysis probe of claim 1, wherein said guiding means comprises an imageable material that can be imaged with magnetic resonance imaging or magnetic resonance spectroscopy. 9. The microdialysis probe of claim 1, wherein said guiding means comprises an imageable material that can be imaged with single photon emission tomography or positron emission tomography. 10. The microdialysis probe of claim 1, wherein said guiding means comprises an imageable material that is formulated as a coating on said microdialysis probe. 11. The microdialysis probe of claim 1, wherein said guiding means also provides information as to the spatial orientation of said membrane portion when said membrane portion is inserted into a medium and when said guiding means is imaged. 12. The microdialysis probe of claim 1, wherein said guiding means provides information as to at least one geometric parameter of said membrane portion when said guiding means is imaged. 13. The microdialysis probe of claim 1, wherein said membrane portion is defined by boundaries, and wherein said guiding means occupies an area on said microdialysis probe such that said guiding means delineates one or more of said boundaries when said guiding means is imaged. 14. The microdialysis probe of claim 13, wherein said guiding means is coextensive with one or more of said boundaries. 15. The microdialysis probe of claim 1, whe... | University of Florida | A61B005/05
US06920348 | B2 | 2005-07-19 | Method and apparatus for determining metabolic factors from an electrocardiogram | 2003-01-31 | Methods and devices are provided for determining metabolic factors using electrocardiogram measurements from a person's Wilson points. A first derivative of an electrocardiogram measurement is calculated. A ratio is calculated of the absolute value of the positive spike of the first derivative to the sum of the absolute values of the positive and negative spikes. In some embodiments, the ratio is multiplied by a constant to determine a metabolic factors. Further operations may be performed on the ratio to determine other metabolic factors. In some embodiments, a garment is provided for easily locating Wilson points. Methods and devices are provided for taking and processing EGC measurements to determine metabolic factors and for using the metabolic factors to optimize an exercise program. | 1. A non-invasive method for determining a metabolic factor, comprising the steps of: obtaining a first derivative of an electrocardiogram measurement; determining Rx, an absolute value of a positive spike of the first derivative; determining RSx, a sum of absolute values of the positive and negative spikes of the first derivative; and dividing Rx by RSx to determine Vx, a number proportional to the metabolic factor. | 1. A non-invasive method for determining a metabolic factor, comprising the steps of: obtaining a first derivative of an electrocardiogram measurement; determining Rx, an absolute value of a positive spike of the first derivative; determining RSx, a sum of absolute values of the positive and negative spikes of the first derivative; and dividing Rx by RSx to determine Vx, a number proportional to the metabolic factor. 2. The method of claim 1, further comprising the step of multiplying Vx by a constant to determine the metabolic factor. 3. The method of claim 2, wherein: the electrocardiogram measurement is taken from the V6 Wilson point; the metabolic factor is aerobic capacity; and the constant is approximately 120. 4. The method of claim 2, wherein: the electrocardiogram measurement is taken from the V2 Wilson point; the metabolic factor is lactacidemia; and the constant is approximately 30. 5. The method of claim 2, wherein: the electrocardiogram measurement is taken from the V3R Wilson point; the metabolic factor is phosphocreatine capacity; and the constant is approximately 115. 6. An apparatus for determining metabolic factors from electrocardiogram measurements, comprising: means for obtaining a first derivative of an electrocardiogram measurement; means for determining Rx, an absolute value of a positive spike of the first derivative; means for determining RSx, a sum of absolute values of the positive and negative spikes of the first derivative; and means for dividing Rx by RSx to determine Vx, a number proportional to the metabolic factor. 7. The apparatus of claim 6, further comprising means for multiplying Vx by a constant to determine the metabolic factor. 8. The apparatus of claim 6, wherein the means for obtaining the first derivative comprises a processor. 9. The apparatus of claim 6, wherein the means for obtaining the first derivative comprises a circuit to which an electrocardiogram measurement is input and which outputs the first derivative of the electrocardiogram measurement. 10. The apparatus of claim 6, further comprising means for converting analog signals from the electrocardiogram unit into digital form. 11. The apparatus of claim 6, further comprising display means. 12. The apparatus of claim 6, further comprising communication means for communicating with one or more devices. 13. The apparatus of claim 7, wherein: the electrocardiogram measurement is taken from the V6 Wilson point; the metabolic factor is aerobic capacity; and the constant is approximately 120. 14. The apparatus of claim 7, wherein: the electrocardiogram measurement is taken from the V2 Wilson point; the metabolic factor is lactacidemia; and the constant is approximately 30. 15. The apparatus of claim 7, wherein: the electrocardiogram measurement is taken from the V3R Wilson point; the metabolic factor is phosphocreatine capacity; and the constant is approximately 115. 16. An apparatus for determining metabolic factors from elect... | Eugene V. Vasin | A61B005/0402
US06948501 | B2 | 2005-09-27 | Patient support for external counterpulsation cardiac assist device | 2002-04-29 | A patient support for an external counterpulsation cardiac assist procedure. The patient support includes: a base; and at least one support connected to the base for maintaining at least a portion of the patient above a tabletop during a positive pressure portion of the external counterpulsation cardiac assist procedure. In one embodiment, the base is a tabletop and the at least one support is movably supported on the table top in lateral and/or vertical directions. In another embodiment, the base is higher than the tabletop and the at least one support is a sling. | 1. A patient support for an external counterpulsation cardiac assist procedure, the patient support comprising: a base, wherein the base is the tabletop; at least one support connected to the tabletop for maintaining at least a portion of the patient above the tabletop during a positive pressure portion of the external counterpulsation cardiac assist procedure; and moving means for movably fixing the at least one support to the tabletop such that the at least one support can be movably positioned under the portion of the patient to be maintained above the tabletop, wherein the moving means comprises the at least one support having a slide portion which is slidably arranged in a corresponding slot on a top surface of the tabletop; wherein the at least one support further having a locking means for locking the at least one support at a predetermined position corresponding to the portion of the patient to be maintained above the tabletop; wherein the locking means comprises a locking shaft having a keyed portion on one end and a threaded portion on another end, the keyed portion being slidingly disposed in a corresponding keyway in the tabletop, the locking means further comprising a threaded knob threadingly engaged with the threaded portion, wherein tightening of the knob on the shaft causes the keyed portion to engage a lower surface of the tabletop to thereby lock the at least one support thereon. | 1. A patient support for an external counterpulsation cardiac assist procedure, the patient support comprising: a base, wherein the base is the tabletop; at least one support connected to the tabletop for maintaining at least a portion of the patient above the tabletop during a positive pressure portion of the external counterpulsation cardiac assist procedure; and moving means for movably fixing the at least one support to the tabletop such that the at least one support can be movably positioned under the portion of the patient to be maintained above the tabletop, wherein the moving means comprises the at least one support having a slide portion which is slidably arranged in a corresponding slot on a top surface of the tabletop; wherein the at least one support further having a locking means for locking the at least one support at a predetermined position corresponding to the portion of the patient to be maintained above the tabletop; wherein the locking means comprises a locking shaft having a keyed portion on one end and a threaded portion on another end, the keyed portion being slidingly disposed in a corresponding keyway in the tabletop, the locking means further comprising a threaded knob threadingly engaged with the threaded portion, wherein tightening of the knob on the shaft causes the keyed portion to engage a lower surface of the tabletop to thereby lock the at least one support thereon. 2. The patient support of claim 1, wherein the at least one support comprises a padded portion fabricated of a soft material. 3. The patient support of claim 2, wherein the at least one support further comprises a fixing means for fixing the at least a portion of the patient thereto. 4. The patient support of claim 3, wherein the fixing means comprises a strap having a first end fixed to the at least one support and a corresponding buckle fixed to the at least one support. 5. The patient support of claim 1, wherein the base is the tabletop. 6. The patient support of claim 5, further comprising height adjustment means for movably fixing the at least one support to the tabletop such that a height of the at least one support above the tabletop can be adjusted. 7. The patient support of claim 6, wherein the height adjustment means comprises at least one rod connected to the at least one support, the at least one rod being slidingly disposed in a corresponding bore in the tabletop. 8. The patient support of claim 7, further comprising a locking means for locking the rod at a predetermined height, wherein the locking means comprises the at least one rod having at least one groove and a detent housed in the tabletop for operatively engaging with the at least one groove. 9. The patient support of claim 1, further comprising biasing means for biasing the at least one support above the tabletop. 10. The patient support of claim 9, wherein the biasing means comprises a constant tension spring corresponding to each of the at least one supports, the const... | P.P.T.T. L.L.P. | A61G015/00 A61F013/00 A61B019/00 A61B005/37
US06898460 | B2 | 2005-05-24 | Method and apparatus for uterine contraction monitoring using linear predictive modeling of abdominal surface EMG signals | 2002-01-15 | A system and a method are provided for processing electromyogram (EMG) input signals from an abdominal surface to detect uterine contractions. The system comprises a sensor configured to detect a uterine EMG signal and generate an EMG input signal. The system further includes a signal processor coupled to the sensor and configured to generate an EMG prediction error signal. The signal processor performs signal-processing operations on the EMG input signal to generate the EMG prediction error signal, where the EMG prediction error signal corresponds to the magnitude of at least one contraction event and the periodicity of multiple contraction events. | 1. A system for processing electromyogram (EMG) input signals from an external abdominal surface to detect uterine contractions, said system comprising: a sensor configured to detect an EMG signal and to generate a corresponding EMG input signal; and a signal processor coupled to said sensor and configured to generate a respective EMG prediction error signal, which represents a magnitude of at least one contraction event and periodicity of a set of multiple contraction events. | 1. A system for processing electromyogram (EMG) input signals from an external abdominal surface to detect uterine contractions, said system comprising: a sensor configured to detect an EMG signal and to generate a corresponding EMG input signal; and a signal processor coupled to said sensor and configured to generate a respective EMG prediction error signal, which represents a magnitude of at least one contraction event and periodicity of a set of multiple contraction events. 2. The system of claim 1 further comprising a display device coupled to said signal processor. 3. The system of claim 2, wherein said display device is selected from the group consisting of a computer monitor, a instrument display monitor, a bedside display monitor, a printer, and a strip chart recorder. 4. The system of claim 1, wherein the sensor comprises two EKG electrodes configured to be placed in contact with said exterior abdominal surface adjacent to a uterus. 5. The system of claim 1, wherein said signal processor further comprises: an amplifier coupled to the sensor and adapted to amplify said EMG input signal and provide an amplified representation of said EMG input signal; a low pass filter coupled to said amplifier and configured to filter said amplified representation of said EMG input signal to generate a low-pass filtered representation of said EMG input signal; an analog to digital convener coupled to said low pass filter and configured to process said low-pass filtered representation of said EMG input signal to generate a digitized representation of said EMG input signal; a noise filter coupled to said analog to digital converter and configured to remove a power line structure from said digitized representation of said EMG input signal to generate a noise-filtered version of said EMG input signal; and a digital processor coupled to said noise filter and configured to process said noise-filtered version of the EMG input signal to compute said EMG prediction error signal. 6. The system of claim 5, wherein said digital processor comprises an adaptation algorithm, wherein said adaptation algorithm comprises a parameter estimating portion that is configured to be coupled to a prediction error filter, wherein said parameter estimating portion is configured to compute said at least one prediction coefficient to optimize an EMG prediction error signal performance index. 7. The system of claim 6, wherein said adaptation algorithm is further configured to filter said digitized representation of said EMG input signal to compute said EMG prediction signal. 8. The system of claim 6, wherein said adaptation algorithm is selected from the group consisting of a Least Square adaptation algorithm and a Burg adaptation algorithm. 9. The system of claim 6, wherein said prediction error filter is configured to have model order in a range from 2 to 10. 10. The system of claim 6, wherein said adaptation algorithm is a Least Square adaptation algorithm that has a time... | General Electric Company | A61B005/04
US06878111 | B2 | 2005-04-12 | System for measuring subjective well being | 2002-07-09 | A system for measuring subjective well-being utilizes an interactive communications module, an electronic transmission module and an intervention module. The interactive communications module receives data or input that reflects a person's subjective well-being and makes a temporal record of the input. The electronic transmission module transmits that data or input to a communications device. The intervention module is coupled to the interactive communications module to communicate a trend of subjective well-being. Trends may be created using a correlation module adapted to correlate and compare subjective and/or objective data. The correlation module may also include an inquiry module to allow a person to inquire about a person's subjective well-being. | 1. A system for subjectively measuring a person's physical and/or emotional well-being, the system comprising: a. an implanted interactive communications module adapted to receive data that reflects a person's subjective well-being; b. an interactive electronic transmission module coupled to the interactive communications module to transmit the data to a communications device; and c. an intervention module coupled to the interactive communications module to communicate a trend of subjective well-being. | 1. A system for subjectively measuring a person's physical and/or emotional well-being, the system comprising: a. an implanted interactive communications module adapted to receive data that reflects a person's subjective well-being; b. an interactive electronic transmission module coupled to the interactive communications module to transmit the data to a communications device; and c. an intervention module coupled to the interactive communications module to communicate a trend of subjective well-being. 2. The primary input system of claim 1, wherein the system further comprises a correlation module that correlates the person's input to create a trend of subjective well-being. 3. The primary input system of claim 1, wherein the system further comprises a correlation module that correlates the person's input to create a trend of subjective well-being and further correlates the input with objective data on the person's well-being to create a comparative trend of subjective well-being data to objective well-being data. 4. The correlation module of claim 2 wherein the correlation module further comprises an inquiry module. 5. The correlation module of claim 3 wherein the correlation module further comprises an inquiry module. 6. The interactive communications module of claim 1, wherein the interactive communications module makes a temporal record of the person's input. 7. The interactive communications module of claim 1, wherein the interactive communications module automatically prompts the person to enter subjective well-being data. 8. The interactive communications module of claim 7, wherein the prompt is an electrically induced. 9. The interactive communications module of claim 8, wherein the electrically induced prompt induces a physical response in the person. 10. The interactive communications module of claim 9, wherein the physical response in the person comprises a muscle twitch. 11. The interactive communications module of claim 7, wherein the prompt is audibly induced. 12. The interactive communications of module of claim 1, wherein the interactive communications module comprises an accelerometer. 13. The interactive communications of module of claim 1, wherein the representation of the person's subjective well-being comprises a verbal response. 14. The interactive communications of module of claim 1, wherein the representation of the person's subjective well-being comprises an electronic response. 15. The electronic response of claim 14, wherein the response is entered through a device in electronic communication with the interactive communications module. 16. The interactive communications of module of claim 1, wherein the interactive communications module comprises a global computer network. 17. The interactive electronic transmission module of claim 1, wherein the representation of the person's subjective well-being comprises a plurality of icons. 18. The icons of claim 17, wherein the icons comprise a happy face, a st... | Cardiac Pacemakers, Inc. | A61B005/00
US06852086 | B2 | 2005-02-08 | Detection of signs of attempted deception and other emotional stresses by detecting changes in weight distribution of a standing or sitting person | 2002-06-18 | Non-obtrusive monitoring apparatus for detecting the emotional stress of attempted deception by a standing or seated subject. The apparatus has sensors for measuring parameters of posture, sway, weight distribution, cardiac function, and respiratory function, and outputs computer-readable data for processing and reporting. Also disclosed are: systems for synchronizing output data with information about questions asked of the subject as well as other interrogation stimuli administered to the subject; and systems for analyzing the output data and interrogation stimuli to identify patterns indicating heightened anxiety. | 1. A method for detecting emotional stress of a human subject in response to cognitive stimuli by measuring variations in the pressure distribution across the subject's footprint, the method comprising: (a) causing the subject to apply force on a flat surface, said flat surface having embedded therein a plurality of pressure sensitive sensors; (b) connecting said pressure sensitive sensors to a signal analysis processor; (c) presenting the subject with cognitive stimuli, the subject's disposition to which is of interest in the context of a given interview; (d) monitoring variations in pressure distribution between said sensors, and deriving variation signals therefrom; (e) filtering said variation signals to derive filtered variation signals; and (d) matching said filtered variation signals to a pattern that represents emotional stress. | 1. A method for detecting emotional stress of a human subject in response to cognitive stimuli by measuring variations in the pressure distribution across the subject's footprint, the method comprising: (a) causing the subject to apply force on a flat surface, said flat surface having embedded therein a plurality of pressure sensitive sensors; (b) connecting said pressure sensitive sensors to a signal analysis processor; (c) presenting the subject with cognitive stimuli, the subject's disposition to which is of interest in the context of a given interview; (d) monitoring variations in pressure distribution between said sensors, and deriving variation signals therefrom; (e) filtering said variation signals to derive filtered variation signals; and (d) matching said filtered variation signals to a pattern that represents emotional stress. 2. The method of claim 1, wherein said subject sits in a chair, and at least one leg of said chair is positioned on said flat surface. 3. The method of claim 1, wherein said flat surface is a mat. 4. The method of claim 3, wherein said mat is embedded in a chair and said subject sits in said chair. 5. The method of claim 4, wherein a plurality of said mats is used and the subject touches different mats of said plurality of mats with different parts of his or her body. 6. The method of claim 5, wherein at least one leg of said chair is positioned upon at least one mat of said plurality of mats. 7. A system for detecting emotional stress of a human subject, in response to cognitive stimuli by measuring variations in pressure distribution across the subject's footprint, the system comprising: (a) a flat surface having embedded therein a plurality of pressure sensitive sensors; (b) a signal conditioning processing means coupled to said pressure sensitive sensors for detecting shifts in the subject's weight distribution according to signals from said pressure sensitive sensors; (c) a signal analysis processing means coupled to said signal conditioning processing means, said signal analysis processing means for analyzing said shifts in weight distribution to detect patterns indicative of emotional stress; and (d) a display coupled to said signal analysis processing means. 8. The system of claim 7, wherein at least one of said sensors is a high-capacitance piezo electric pressure sensor. 9. The system of claim 7, wherein at least one of said sensors is a fiber-optic pressure sensor. 10. The system of claim 7, wherein at least one of said sensors is a strain gauge. 11. The system of claim 7, wherein said signal analysis processing means includes software operative to identify pressure change correlates of flight-or-fight responses. 12. The system of claim 11, wherein said software is expert system software. | - | A61B005/103
US06922584 | B2 | 2005-07-26 | Method and apparatus for discrimination atrial fibrillation using ventricular rate detection | 2001-11-28 | A method and apparatus for method for discriminating heart rhythms includes computing a first predetermined number of RR intervals from received QRS intervals, and computing a median RR interval corresponding to a predetermined number of the first predetermined number of RR intervals. A predetermined beat-to-beat variation and a corresponding predetermined count are determined based on the computed median RR interval. Beat-to-beat variation differences between the first predetermined number of RR intervals are computed and a determination is made as to whether the computed beat-to-beat variation differences are greater than the predetermined beat-to-beat variation, and as to whether a number of the computed beat-to-beat variation differences that are greater than the predetermined beat-to-beat variation is greater than the predetermined count. The heart rhythm is identified as an irregular rhythm in response to the number being greater than or equal to the predetermined count. | 1. A method for discriminating heart rhythms in an implantable medical device, comprising the steps of: receiving a QRS interval corresponding to the heart rhythm and computing a first predetermined number of RR intervals from the received QRS intervals; computing a median RR interval corresponding to a predetermined number of the first predetermined number of RR intervals; determining a rate zone in response to the computed median RR interval; determining a predetermined beat-to-beat variation corresponding to the rate zone; varying a count in response to the rate zone and the predetermined beat-to-beat variation; computing beat-to-beat variation differences between the first predetermined number of RR intervals; comparing the beat-to-beat variation differences to the predetermined beat-to-beat variation to determine whether the number of the computed beat-to-beat variation differences greater than the predetermined beat-to-beat variation is greater than the count; and identifying the heart rhythm in response to the number of the computed beat-to-beat variation differences relative to the count. | 1. A method for discriminating heart rhythms in an implantable medical device, comprising the steps of: receiving a QRS interval corresponding to the heart rhythm and computing a first predetermined number of RR intervals from the received QRS intervals; computing a median RR interval corresponding to a predetermined number of the first predetermined number of RR intervals; determining a rate zone in response to the computed median RR interval; determining a predetermined beat-to-beat variation corresponding to the rate zone; varying a count in response to the rate zone and the predetermined beat-to-beat variation; computing beat-to-beat variation differences between the first predetermined number of RR intervals; comparing the beat-to-beat variation differences to the predetermined beat-to-beat variation to determine whether the number of the computed beat-to-beat variation differences greater than the predetermined beat-to-beat variation is greater than the count; and identifying the heart rhythm in response to the number of the computed beat-to-beat variation differences relative to the count. 2. The method of claim 1, wherein the step of computing beat-to-beat variation differences comprises calculating the difference of the absolute value of RR(n)-RR(n-1), wherein RR(n) and RR(n-1) are consecutive RR intervals of the first predetermined number of RR intervals. 3. The method of claim 1, wherein the step of identifying comprises the steps of: identifying the heart rhythm as an irregular rhythm in response to the number of the computed beat-to-beat variation differences greater than the predetermined beat-to-beat variation being greater than or equal to the count; and identifying the heart rhythm as a regular rhythm in response to the number of the computed beat-to-beat variation differences greater than the predetermined beat-to-beat variation being less than the count. 4. The method of claim 3, wherein the irregular rhythm corresponds to atrial fibrillation. 5. The method of claim 1, wherein the predetermined beat-to-beat variation is set equal to 50 ms in response to the rate zone being greater than 500 ms, to 25 ms in response to the rate zone being less than or equal to 500 ms and greater than 400 ms, and to 15 ms in response to the rate zone being less than or equal to 400 ms. 6. The method of claim 5, wherein the count is set equal to 8 in response to the rate zone being greater than 400 ms and to 5 in response to the rate zone being less than or equal to 400 ms. 7. A method for discriminating between heart rhythms in an implantable medical device, comprising the steps of: computing a first predetermined number of RR intervals from received QRS intervals; computing a median RR interval corresponding to a predetermined number of the first predetermined number of RR intervals; determining a rate zone in response to the computed median RR interval; determining a predetermined beat-to-beat variation and varying a predetermined count ... | Medtronic, Inc. | A61B005/0452
US06976982 | B2 | 2005-12-20 | Flexible microchip devices for ophthalmic and other applications | 2002-01-09 | Microchip device arrays that can conform to a curved surface are provided for the controlled release or exposure of reservoir contents. The arrays comprise two or more microchip device elements, each of which includes a plurality of reservoirs that contain molecules for controlled release or components for selective exposure, and a means for flexibly connecting the device elements. The reservoirs can contain one or more drugs and/or one or more secondary devices, such as a sensor or a component thereof. Preferably, the microchip devices contain and controllably release therapeutic, prophylactic, and diagnostic molecules to and into the eye of a patient in need thereof. | 1. A microchip device array for the controlled release or exposure of reservoir contents comprising: two or more microchip device elements, each of which includes a plurality of reservoirs which contain molecules for controlled release or components for selective exposure; anda means for flexibly connecting said device elements, so as to form a flexible array which can conform to a curved surface. | 1. A microchip device array for the controlled release or exposure of reservoir contents comprising: two or more microchip device elements, each of which includes a plurality of reservoirs which contain molecules for controlled release or components for selective exposure; anda means for flexibly connecting said device elements, so as to form a flexible array which can conform to a curved surface. 2. The microchip device array of claim 1, wherein the means for flexibly connecting comprises a flexible supporting layer attached to a surface of the device elements. 3. The microchip device array of claim 2, wherein the flexible supporting layer comprises a polymer. 4. The microchip device array of claim 3, wherein the polymer is selected from the group consisting of polyimides, polyesters, parylenes, and hydrogels. 5. The microchip device array of claim 2, wherein the flexible supporting layer is porous or permeable to molecules releasable from the reservoirs or provided with one or more apertures through said flexible supporting layer. 6. The microchip device array of claim 1, wherein the means for flexibly connecting comprises one or more hinges or flexible tethers connecting two or more of the device elements. 7. The microchip device array of claim 1 for implantation onto or into a patient, wherein the array can conform to the curvature of a tissue surface. 8. The microchip device array of claim 7 for implantation into or onto the eye of the patient, wherein the tissue surface comprises ophthalmic tissue. 9. The microchip device array of claim 7, wherein the tissue surface is selected from the group consisting of the stratum comeum, mucosal membranes, blood vessels, bone, brain, and bladder. 10. The microchip device array of claim 1, wherein the microchip device elements further comprise a plurality of discrete reservoir caps over the molecules in the reservoirs, wherein each reservoir cap controls release of the molecules from one of the reservoirs. 11. The microchip device array of claim 1 further comprising a means for wirelessly communicating with the microchip device elements. 12. The microchip device array of claim 11, wherein the communicating means comprises a photocell to receive incident light energy. 13. The microchip device array of claim 1 further comprising an energy storage means. 14. The microchip device array of claim 13, wherein the energy storage means comprises a capacitor, a battery, or both. 15. The microchip device array of claim 1 further comprising electrical connections between two or more of the microchip device elements, such that the microchip device elements can be powered or controlled by a common energy source or control source, respectively. 16. The microchip device array of claim 1, wherein the reservoirs contain drug molecules. 17. The microchip device array of claim 1, wherein the reservoirs contains one or more secondary devices for exposure. 18. The microchip device array of claim 17, wherein ... | MicroCHIPS, Inc. | A61K009/22 A61M031/00 A61B005/05
US06937889 | B2 | 2005-08-30 | Method and apparatus for determining pregnancy possibility | 2001-10-02 | Disclosed are method and apparatus for making a decision of possibility of pregnancy in terms of how the values of bioelectrical impedance vary after expiration of a certain length of time since the ovulation day. | 1. A pregnancy possibility determining apparatus comprising: a plurality of pairs of electrodes; a bioelectrical impedance meter; a bioelectrical impedance storing device; an ovulation decision-making unit or an ovulation presuming unit; a pregnancy decision-making unit; and a warning unit, wherein said pairs of electrodes can be applied to selected points of the outer layer of the skin of a woman's body; said bioelectrical impedance meter measures the value of the bioelectrical impedance appearing between one of said pairs of electrodes; said boielectrical impedance storing device stores the so measured values of bioelectrical impedance; said ovulation decision-making unit makes a decision as to when the woman has an ovulation day on the basis of a time-sequence transition or historical record of the values of bioelectrical impedance stored in the bioelectrical impedance storing device; or said ovulation presuming unit counts a predetermined number of days from the beginning of a next menstration period presumable from the time-sequence transition or historical record of the values of bioelectrical impedance stored in the bioelectrical impedance storing device, and said ovulation presuming unit presumes that the day so counted backward from the presumable beginning day of the next menstration period falls on the ovulation day; said pregnancy decision-making unit makes a decision of the possibility of pregnancy in terms of whether or not a predetermined number of days have passed from the ovulation day actually determined or presumed with the values of bioelectrical impedance remaining low; and said warning unit informs the woman of the result decision-making of the possibility of pregnancy. | 1. A pregnancy possibility determining apparatus comprising: a plurality of pairs of electrodes; a bioelectrical impedance meter; a bioelectrical impedance storing device; an ovulation decision-making unit or an ovulation presuming unit; a pregnancy decision-making unit; and a warning unit, wherein said pairs of electrodes can be applied to selected points of the outer layer of the skin of a woman's body; said bioelectrical impedance meter measures the value of the bioelectrical impedance appearing between one of said pairs of electrodes; said boielectrical impedance storing device stores the so measured values of bioelectrical impedance; said ovulation decision-making unit makes a decision as to when the woman has an ovulation day on the basis of a time-sequence transition or historical record of the values of bioelectrical impedance stored in the bioelectrical impedance storing device; or said ovulation presuming unit counts a predetermined number of days from the beginning of a next menstration period presumable from the time-sequence transition or historical record of the values of bioelectrical impedance stored in the bioelectrical impedance storing device, and said ovulation presuming unit presumes that the day so counted backward from the presumable beginning day of the next menstration period falls on the ovulation day; said pregnancy decision-making unit makes a decision of the possibility of pregnancy in terms of whether or not a predetermined number of days have passed from the ovulation day actually determined or presumed with the values of bioelectrical impedance remaining low; and said warning unit informs the woman of the result decision-making of the possibility of pregnancy. 2. A pregnancy possibility determining apparatus according to claim 1 wherein it further comprises a physical condition decision-making unit, which can make a decision as to which physical phase of the monthly body condition the woman has presently on the basis of the time-sequence transition or historical record of the values of bioelectrical impedance stored in the bioelectrical impedance storing device, whereby the warning unit may inform the woman of which physical phase of the monthly body condition the woman has presently. 3. A pregnancy possibility determining apparatus according to claim 1 or 2 wherein the possibility of pregnancy is presumed when 21 days have passed from the ovulation day actually determined or from the presumed ovulation day with the values of bioelectrical impedance remaining low. 4. A pregnancy possibility determining apparatus according to claim 1 or 2 wherein pieces of information provided by the warning unit are changed depending on how many days have passed from the ovulation day actually determined or from the presumed ovulation day with the values of bioelectrical impedance remaining low. 5. A pregnancy possibility determining apparatus according to claim 1 or 2 wherein the possibility of pregnancy is given by numerical... | Tanita Corporation | A61B005/05
US06869591 | B2 | 2005-03-22 | Paramagnetic particles that provide improved relaxivity | 2002-05-21 | An improved contrast agent for magnetic resonance imaging comprises particles to each of which is coupled a multiplicity of chelating agents containing paramagnetic ions. In the improved agent, the position of the ion is offset from the surface of the particle so as to improve the relaxivity imparted by the contrast agent. | 1. A contrast agent for magnetic resonance imaging (MRI), which agent comprises particles, wherein said particles are coupled to a chelator containing a paramagnetic ion which ion is positioned offset from the surface of the particles, so as to provide said ion substantial access to hydrogen nuclei in a surrounding liquid, whereby the relaxivity of said nuclei is enhanced, wherein said particles are microparticles or nanoparticles comprised of an inert core comprising fluorocarbon liquid surrounded by a lipid/surfactant coating, and wherein said offset is such that the particle provides a ρ1 of at least about 0.5×106 (s*mM)−1 or a ρ2 of at least about 1×106 (s*mM)−1 at a field strenth of 1.5 T on a per particle basis; and wherein the particles are emulsified in a liquid emulsion. | 1. A contrast agent for magnetic resonance imaging (MRI), which agent comprises particles, wherein said particles are coupled to a chelator containing a paramagnetic ion which ion is positioned offset from the surface of the particles, so as to provide said ion substantial access to hydrogen nuclei in a surrounding liquid, whereby the relaxivity of said nuclei is enhanced, wherein said particles are microparticles or nanoparticles comprised of an inert core comprising fluorocarbon liquid surrounded by a lipid/surfactant coating, and wherein said offset is such that the particle provides a ρ1 of at least about 0.5×106 (s*mM)−1 or a ρ2 of at least about 1×106 (s*mM)−1 at a field strenth of 1.5 T on a per particle basis; and wherein the particles are emulsified in a liquid emulsion. 2. The agent of claim 1, wherein said offset is such that the particle provides a ρ1 of at least about 10 (s*mM)−1 or a ρ2 of at least about 20 (s*mM)−1 at a field strength of 1.5 T on a per ion basis. 3. The agent of claim 1, wherein said offset is such that ρ1 is increased at least about 1.5-fold or ρ2 is increased at least about 1.5-fold at a field strength of 1.5 T on a per particle basis as compared to ρ1 or ρ2 of particles wherein the paramagnetic ion resides at less than 5 Å from the surface. 4. The agent of claim 1, wherein said inert core comprises a perfluorocarbon compound. 5. The agent of claim 4, wherein the inert core comprises a mixture of fluorocarbons and oils. 6. The agent of claim 1, wherein the chelator is selected from the group consisting of a porphyrin, ethylenediaminetetraacetic acid (EDTA), diethylenetriamine-N,N,N′,N″,N″-pentaacetate (DTPA), 1,4,10,13-tetraoxa-7,16-diazacyclooctadecane-7,16-diacetate, N-2-(azol-1(2)-yl)ethyliminodiacetic acid, 1,4,7,10-tetraazacyclododecane-N,N′,N″,N′″-tetraacetic acid, 1,7,13-triaza-4,10,16-trioxacyclo-octadecane-N,N′,N″-triacetate, tetraethylene glycol,1,5,9-triazacyclododecane-N,N′,N″,-tris(methylene)phosphonic acid, and N,N′,N″-trimethylammonium chloride. 7. The agent of claim 6, wherein the chelator is DTPA. 8. The agent claim 1, wherein the paramagnetic ion is selected from the group consisting of scandium, titanium, vanadium, chromium, manganese, iron, cobalt, nickel, copper, molybdenum, ruthenium, cerium, praseodymium, neodymium, promethium, samarium, europium, gadolinium, terbium, dysprosium, holmium, erbium, thulium, and ytterbium. 9. The agent of claim 8, wherein the paramagnetic ion is gadolinium. 10. The agent of claim 1, wherein the lipid/surfactant coating comprises at least one compound selected from the group consisting of a natural phospholipid, a synthetic phospholipid, a fatty acid, a cholesterol, a lysolipid, a sphingomyelin, a tocopherol, a glucolipid, a stearylamine, a cardiolipin, a lipid with an ether-linker fatty acid, a lipid with an ester linked fatty acid, a polymerized lipid, and a polyethylene glycol-conjugated lipid. 11. The agent of claim 1 wherein said particles are c... | Barnes-Jewish Hospital | A61B005/055
US06850795 | B2 | 2005-02-01 | Method of noninvasively determining a patient's susceptibility to arrhythmia | 2003-01-09 | A system and method for detecting a patient's susceptibility to arrhythmias and cardiac tissue abnormality is disclosed. The method consists of using a computer 27, a display 23, software loaded onto the computer that generates graphical user interfaces (GUIs), an electronic interface 18, and a plurality of electrodes. The electronics interface 18 is in electronic communication with the computer 27, and further in electronic communication with the electrodes that are placed by self-adhesion at predetermined locations on a test subject 35. According to one aspect of the invention, the method enables a user, typically a medical professional, to initiate, with minimal input, certain diagnostic tests involving observing and analyzing a series of QRS complexes 130, some of which are biased with a subpacing current, and others of which are unbiased. The signals are then compared, and the differences are analyzed to detect a patient's susceptibility to arrhythmias and cardiac tissue abnormality. | 1. A method for detecting patient susceptibility to arrhythmias comprising the steps of: providing means for injecting low level electromagnetic energy stimulations into a patient's body suitable to alter at least one cardiac signal to allow a comparison of acquired distinguishable signal differences with at least one other cardiac signal which has not been affected by said injected energy; connecting a lead system for both sensing the cardiac signals and delivering the injected energy to said means for injecting low level electromagnetic energy into the patient's body, said lead system, comprising a plurality of energy sensing and energy delivery leads and a plurality of energy sensing and energy delivery electrodes; utilizing sensing improvement means for improving the quality of the sensed electromagnetic energy during the stimulation and acquisition process; and evaluating the sensed signal for one or more arrhythmia predictive factors by evaluating variability over multiple QRST complexes by evaluating the relative timing of any combination of P-wave, Q-wave, R-wave, S-wave, or T-wave in a single QRST complex. | 1. A method for detecting patient susceptibility to arrhythmias comprising the steps of: providing means for injecting low level electromagnetic energy stimulations into a patient's body suitable to alter at least one cardiac signal to allow a comparison of acquired distinguishable signal differences with at least one other cardiac signal which has not been affected by said injected energy; connecting a lead system for both sensing the cardiac signals and delivering the injected energy to said means for injecting low level electromagnetic energy into the patient's body, said lead system, comprising a plurality of energy sensing and energy delivery leads and a plurality of energy sensing and energy delivery electrodes; utilizing sensing improvement means for improving the quality of the sensed electromagnetic energy during the stimulation and acquisition process; and evaluating the sensed signal for one or more arrhythmia predictive factors by evaluating variability over multiple QRST complexes by evaluating the relative timing of any combination of P-wave, Q-wave, R-wave, S-wave, or T-wave in a single QRST complex. | Harbinger Medical, Inc. | A61B005/0452
US06960198 | B2 | 2005-11-01 | Tube and syringe holding device for blood collection and for producing vacuum into the tube | 2002-05-15 | The invention consists of a tube syringe holding device for blood collection and for producing vacuum into the tube which target is to enable the collection of blood, that can be manufactured in a light material where two bodies, one of them cylindrical and the other semi-cylindrical, both hollow and of the same length, form one structure. Its bottom ends are opened. The top end of the cylindrical body is continued by a conduit with a male luer adapter, and the top end of the semi-cylindrical one is continued by two conduits with threads. | 1. A tube syringe holding device for blood collection, the holding device comprising: a cylindrical body having a bottom end, a top end, a diameter and a length; and a semi-cylindrical body having a bottom end, a top end, a diameter and a length; wherein the cylindrical body and semi-cylindrical body form one structure; and wherein the cylindrical body and semi-cylindrical body are adapted to produce a vacuum into the tube; wherein the cylindrical body further comprises a male luer adapter at the top end; wherein the semi-cylindrical body comprises a first threaded female luer adapter and a second male luer adapter at the top end. | 1. A tube syringe holding device for blood collection, the holding device comprising: a cylindrical body having a bottom end, a top end, a diameter and a length; and a semi-cylindrical body having a bottom end, a top end, a diameter and a length; wherein the cylindrical body and semi-cylindrical body form one structure; and wherein the cylindrical body and semi-cylindrical body are adapted to produce a vacuum into the tube; wherein the cylindrical body further comprises a male luer adapter at the top end; wherein the semi-cylindrical body comprises a first threaded female luer adapter and a second male luer adapter at the top end. 2. A tube syringe holding device, according to claim 1, further comprising an elongated tube having a first and second female luer adapter at each end, wherein the first female luer of the elongated tube is connected to the male luer adapter of the cylindrical body. 3. A tube syringe holding device, according to claim 2, further comprising a first and second needle having a threaded male luer adapter, wherein each needle is connected to the threaded female luer adapter of the semi-cylindrical body. 4. A tube syringe holding device, according to claim 3, wherein the second female luer of the elongated tube is connected to the first needle having a threaded male luer adapter. 5. A tube syringe holding device, according to claim 4, wherein a needle having a female luer adapter is connected to the second needle having a threaded male luer adapter. 6. A tube syringe holding device, according to claim 5, further comprising a plunger adapted to be introduced into the bottom end of the cylindrical body and a blood sample tube adapted to be introduced into the bottom end of the semi-cylindrical body; wherein the blood sample tube contains air. 7. A tube syringe holding device, according to claim 6, wherein the semi-cylindrical body is adapted to receive blood samples tubes having different diameters. 8. A tube syringe holding device, according to claim 7, wherein the blood is drawn into the tube while it is being collected from the patient. 9. A tube syringe holding device, according to claim 6, wherein the holding device produces a vacuum by aspirating air from blood test tubes at the same time the blood is drawn. 10. A tube syringe holding device, according to claim 1, wherein the diameter and length of the cylindrical body are similar to those of the holding device. 11. A tube syringe holding device, according to claim 1, wherein the semi-cylindrical body further includes two walls that are opened and partially lengthened to form one structure with the cylindrical body. 12. A tube syringe holding device, according to claim 1, wherein the length of the semi-cylindrical body is similar to the length of the cylindrical body. 13. A tube syringe holding device, according to claim 1, wherein the holding device is reusable or disposable. 14. A tube syringe holding device for blood collection, the holding device compris... | - | A61B019/00 A61B005/00 A61M005/32 B65D081/00
TWI236891 |  | 2005-08-01 | 血壓計之校正裝置及其方法 | 2004-02-13 | 一種血壓計校正裝置及其方法，用以校正一血壓計。標準壓力源用以產生標準壓力。校正裝置接收標準壓力並將之轉送給血壓計，並輸出對應於標準壓力之標準壓力值至血壓計。血壓計係依據接收之標準壓力及標準壓力值而進行校正。 | 一種血壓計的校正方法，該方法包括：輸出一標準壓力；產生對應於該標準壓力之一標準壓力值；利用該血壓計量測該標準壓力並據以產生一量測信號；產生一顯示壓力值，其係依據該量測信號而得；以及校正該血壓計，其係依據該標準壓力值與該顯示壓力值而得。 | 一種血壓計的校正方法，該方法包括：輸出一標準壓力；產生對應於該標準壓力之一標準壓力值；利用該血壓計量測該標準壓力並據以產生一量測信號；產生一顯示壓力值，其係依據該量測信號而得；以及校正該血壓計，其係依據該標準壓力值與該顯示壓力值而得。 如申請專利範圍第1項所述之校正方法，其中在產生該顯示壓力值之步驟中，係依據一第一轉換關係而依據該量測信號以決定該顯示壓力值。 如申請專利範圍第2項所述之校正方法，其中校正該血壓計之步驟係依據該標準壓力值與該顯示壓力值的比較結果而建立一第二轉換關係。 如申請專利範圍第3項所述之校正方法，其中若該標準壓力值與該顯示壓力值有誤差，則以該量測信號與該標準壓力值做為該第二轉換關係之一參考點。 一種血壓計校正系統，用以校正一血壓計，該校正系統包括：一標準壓力源，用以產生一標準壓力；以及一校正裝置，與該標準壓力源耦接，接收該標準壓力並將之轉送給該血壓計，並輸出對應於該標準壓力之一標準壓力值至該血壓計；其中，該血壓計係依據接收之該標準壓力及該標準壓力值而進行校正。 一種血壓計，該血壓計具有一校正模式及一量測模式，當該血壓計於該校正模式時，係接收一標準壓力及對應於該標準壓力之一標準壓力值，該血壓計包括：一壓力感測器，用以接收該標準壓力並據以輸出一量測信號；以及一微控制器，依據該量測信號而產生一顯示壓力值；其中，當該血壓計於該校正模式時，該微控制器係依據該顯示壓力值及該標準壓力值而校正該血壓計。 如申請專利範圍第6項所述之血壓計，其中該血壓計更包括一類比/數位轉換器，用以將該壓力感測器輸出之該量測信號數位化後輸出至該微控制器。 如申請專利範圍第6項所述之血壓計，其中該微控制器係依據一第一轉換關係而由該量測信號決定該顯示壓力值。 如申請專利範圍第8項所述之血壓計，其中當該血壓計於該校正模式時，該微控制器係依據該標準壓力值與該顯示壓力值的比較結果而建立一第二轉換關係。 如申請專利範圍第9項所述之血壓計，其中若該標準壓力值與該顯示壓力值有誤差，則以該量測信號與該標準壓力值做為該第二轉換關係之一參考點。 |  | A61B005/02 G01L027/00
TWI235647 |  | 2005-07-11 | 醫療裝置 | 2003-04-11 | 提供一種醫療裝置，其可以兼顧操作性和對患者的安全性。該醫療裝置係由將患者的血液在體外進行迴圈的同時，並進行淨化的透析裝置所構成，該透析裝置具備有：操作輸入部(12)，係操作者用以進行操作輸入；動作部，其根據該操作輸入部(12)的輸入，進行用於醫療所規定的動作其中該醫療裝置具備有：人體檢測機構(18)，可以檢測出操作者在上述操作輸入部(12)附近；控制部(19)，控制成只有在人體檢測機構檢測出操作者時，允許用以指示上述動作部動作的上述操作輸入部(12)來進行輸入操作，或者上述動作部會根據上述操作輸入部(12)的輸入而動作。 | 一種醫療裝置，係具備有：操作者用以進行輸入操作的操作輸入部、和根據該操作輸入部的輸入以進行用於醫療所規定的動作之動作部，該醫療裝置之特徵為具備有：人體檢測機構，可以檢測出操作者在該操作輸入部附近；控制部，控制成只是在人體檢測機構檢測出操作者時，允許指示該動作部動作的該操作輸入部進行輸入操作，或者該動作部根據該操作輸入部的輸入而動作。 | 一種醫療裝置，係具備有：操作者用以進行輸入操作的操作輸入部、和根據該操作輸入部的輸入以進行用於醫療所規定的動作之動作部，該醫療裝置之特徵為具備有：人體檢測機構，可以檢測出操作者在該操作輸入部附近；控制部，控制成只是在人體檢測機構檢測出操作者時，允許指示該動作部動作的該操作輸入部進行輸入操作，或者該動作部根據該操作輸入部的輸入而動作。 如申請專利範圍第1項之醫療裝置，其係可解除該控制部的控制。 如申請專利範圍第1項之醫療裝置，其中具備有顯示機構，可進行該動作部的動作狀況等規定的顯示，同時該顯示機構在該人體檢測機構檢測到操作者時，則進行規定的顯示。 如申請專利範圍第2項之醫療裝置，其中具備有顯示機構，可進行該動作部的動作狀況等規定的顯示，同時該顯示機構在該人體檢測機構檢測到操作者時，則進行規定的顯示。 如申請專利範圍第3項之醫療裝置，其中該顯示機構係由具有背燈的液晶顯示幕構成，若用該人體檢測機構檢測出操作者，可提高該背燈的發光或亮度，進行規定的顯示。 如申請專利範圍第4項之醫療裝置，其中該顯示機構係由具有背燈的液晶顯示幕構成，若用該人體檢測機構檢測出操作者，可提高該背燈的發光或亮度，進行規定的顯示。 如申請專利範圍第1項至第6項中任一項之醫療裝置，其中具備有聲音輸出機構，根據該動作部的動作狀況而輸出規定的聲音，同時該聲音輸出機構的輸出是根據該人體檢測機構之操作者的檢測而被控制。 如申請專利範圍第1項至第6項中任一項之醫療裝置，其中具備有外部顯示燈，根據該動作部的動作狀況發光，同時可根據該人體檢測機構之操作者的檢測以控制該外部顯示燈的發光量。 如申請專利範圍第7項之醫療裝置，其中具備有外部顯示燈，根據該動作部的動作狀況發光，同時可根據該人體檢測機構之操作者的檢測以控制該外部顯示燈的發光量。 如申請專利範圍第1項之醫療裝置，其中該人體檢測機構具有照射紅外線的發光機構和接受其反射的受光機構，藉由受光機構接受反射以檢測操作者位在該操作輸入部附近。 |  | A61B005/02
TWI245622 |  | 2005-12-21 | 自律神經之分析方法及其裝置 | 2003-12-26 | 一種自律神經之分析方法及其裝置，此分析裝置包括感測元件、電腦設備與輸出設備。其感測元件之電極貼片係黏貼於受測者之手臂皮膚表面，以偵測與輸出受測者之心跳訊號。電腦設備則藉由訊號收集線自電極貼片接收心跳訊號，並對心跳訊號予以放大、濾波、數位化與轉換計算後得到多數個心率變異性參數，且針對這些心率變異性參數作計算、比較與分析後至資料庫中之內建表尋找相對應之分析資料。最後，則由輸出設備輸出分析資料這些心率變異性參數檢查資料。 | 一種自律神經之分析方法，係以非侵體方式對受測者之自律神經系統作診斷，該分析方法包括：輸入受測者之一基本資料，並量測受測者之一心跳訊號；對該心跳訊號作轉換，並獲得多數個心率變異性參數；對該些心率變異性參數至少其中之一作自然對數運算，並得到運算後之該些心率變異性參數；以一資料庫之多數個參考值對運算後之該些心率變異性參數作計算與最佳化，並輸出所得到之多數個標準差；根據該基本資料與該些標準差至一內建表中尋找相符之一分析資料；以及輸出整合該些心率變異性參數、該分析資料、該基本資料與該些標準差之一檢查資料。 | 一種自律神經之分析方法，係以非侵體方式對受測者之自律神經系統作診斷，該分析方法包括：輸入受測者之一基本資料，並量測受測者之一心跳訊號；對該心跳訊號作轉換，並獲得多數個心率變異性參數；對該些心率變異性參數至少其中之一作自然對數運算，並得到運算後之該些心率變異性參數；以一資料庫之多數個參考值對運算後之該些心率變異性參數作計算與最佳化，並輸出所得到之多數個標準差；根據該基本資料與該些標準差至一內建表中尋找相符之一分析資料；以及輸出整合該些心率變異性參數、該分析資料、該基本資料與該些標準差之一檢查資料。 如申請專利範圍第1項所述之自律神經之分析方法，其中該些心率變異性參數包括一波峰間距、一低頻變異性參數、一高頻變異性參數與一低頻/高頻變異性參數比值。 如申請專利範圍第2項所述之自律神經之分析方法，其中根據該基本資料與該些標準差至一內建表中尋找相符之該分析資料之步驟包括：以該些標準差中代表該波峰間距之值與該內建表中之多數個內建值作比較，並得到該波峰間距之功能狀態；以該些標準差中代表該低頻變異性參數之值與該內建 表中之該些內建值作比較，並得到該低頻變異性參數之功能狀態；以該些標準差中代表該高頻變異性參數之值與該內建表中之該些內建值作比較，並得到該高頻變異性參數之功能狀態；以該些標準差中代表該低頻/高頻變異性參數比值之值與該內建表中之該些內建值作比較，並得到該低頻/高頻變異性參數比值之功能狀態；以及根據得到之該些心率變異性參數之功能狀態輸出相對應之該分析資料。 如申請專利範圍第1項所述之自律神經之分析方法，其中對該心跳訊號作轉換，並獲得該些心率變異性參數之步驟包括：對該心跳訊號作數位轉換，並偵測該數位心跳訊號之多數個波峰；統計並確認每一該些波峰；計算與得到該些波峰之多數個波峰間距，並統計與確認每一該些波峰間距；以及對該些波峰間距作計算，並得到該些心率變異性參數之頻率域。 如申請專利範圍第4項所述之自律神經之分析方法，其中對該些波峰間距計算係使用快速傅立葉轉換。 如申請專利範圍第1項所述之自律神經之分析方法，其中該分析資料包括整體自律神經功能活性、生理年 齡曲線與心跳。 如申請專利範圍第1項所述之自律神經之分析方法，其中該檢查資料更包括極低頻變異性參數、功率密度頻譜(PSD)與總功率。 一種自律神經之分析裝置，係以非侵體方式對受測者之自律神經作診斷，該分析裝置包括：一感測元件，具有多數個電極貼片與多數個訊號收集線，該些電極貼片係黏貼於受測者之手臂皮膚表面，以偵測與輸出受測者之一心跳訊號；一電腦設備，具有一資料庫，該電腦設備藉由該些訊號收集線接收該心跳訊號，並對該心跳訊號予以放大、濾波、數位化與一轉換計算後得到多數個心率變異性參數，且於針對該些心率變異性參數作計算、比較與分析後至該資料庫中之一內建表尋找相對應之一分析資料；以及一輸出設備，耦接至該電腦設備，以接收及輸出整合該分析資料與該些心率變異性參數之一檢查資料。 如申請專利範圍第8項所述之自律神經之分析裝置，其中該些心率變異性參數包括一波峰間距、一低頻變異性參數、一高頻變異性參數與一低頻/高頻變異性參數比值。 如申請專利範圍第8項所述之自律神經之分析裝置，其中該轉換算為快速傅立葉轉換。 如申請專利範圍第8項所述之自律神經之分析裝置，其中該分析資料包括身體狀況指標、受測者之體質、 整體自律神經功能活性、生理年齡曲線與心跳。 如申請專利範圍第8項所述之自律神經之分析裝置，其中該檢查資料更包括極低頻變異性參數、功率密度頻譜(PSD)與總功率。 如申請專利範圍第8項所述之自律神經之分析裝置，其中該輸出設備為一顯示器，用以顯示該檢查資料。 如申請專利範圍第8項所述之自律神經之分析裝置，其中該輸出設備為一印表機，用以列印該檢查資料。 如申請專利範圍第8項所述之自律神經之分析裝置，其中該輸出設備為一燒錄機，用以將該檢查資料燒錄於光碟片上。 如申請專利範圍第8項所述之自律神經之分析裝置，其中該輸出設備為一網路系統，用以將該檢查資料傳送至遠端終端機。 如申請專利範圍第8項所述之自律神經之分析裝置，其中該電腦設備包括至少一放大器、一濾波器與一類比/數位轉換器。 如申請專利範圍第8項所述之自律神經之分析裝置，其中該電腦設備為一具數位訊號處理(DSP)能力之電腦，用以作頻域分析、時域分析與非線性分析。 |  | A61B005/024
TWI227665 |  | 2005-02-11 | 非侵襲性血管機械特性檢測技術 | 2003-06-18 | 本發明提供一種安全且方便之非侵襲性方法，動態檢測人體週邊動脈血管機械性質，包括血管硬度(彈性係數)以及黏彈特性。本創作利用超音波儀器與壓力感測器同步監視血壓與血管管徑，藉由適當的超音波固定裝置與特殊的影像處理技術，取得活體週邊動脈血管機械性質，量測結果有很好的重複性且不受人為因素影響，可以應用於血液循環相關疾病之檢測與藥物控制的療效的佐證。同時，在血管上游動態施加物理性改變，可以動態觀察週邊動脈血管的機械性質變化與反應。本創作所提出之血管動態黏彈特性(能量耗損率)，在週邊循環不良的動脈血管有很明顯的改變，可以反應管壁材質變化，是加速血液循環疾病發展的主要機制。 | 一種非侵襲性血管機械黏彈特性的檢測方法，至少包含下列步驟：◎提供一夾具將超音波探頭及一壓力計組合，以固定間隔隔開；並以適當的鬆緊帶固定於人體血管之體表；◎量測該部位血管的收縮壓與輸張壓以校正血壓波型；◎將觸診壓力計訊號以適當的放大電路處理，連接超音波系統，將血管的二維影像、血管直徑脈動(M-mode)與相對應的血壓波型同步顯示於螢幕；◎將螢幕中含數個心跳週期血管變化的影像記錄傳輸至電腦平台；◎以影像分析方法擷取同步之血管脈動波型、管壁內中層組織之厚度變形量與血壓波型；◎由影像掃描深度的刻度訂定影像比例尺度；◎調整血管等效管徑值與等效脈動值；◎自動化影像處理，將同步同位置取得的血壓與管徑變形資料，計算血管管徑與管壁內中層組織(intima-media thickness; IMT)之應力-應變關係曲線、分析其動態楊氏係數及能量消散率(EDR)。 | 一種非侵襲性血管機械黏彈特性的檢測方法，至少包含下列步驟：◎提供一夾具將超音波探頭及一壓力計組合，以固定間隔隔開；並以適當的鬆緊帶固定於人體血管之體表；◎量測該部位血管的收縮壓與輸張壓以校正血壓波型；◎將觸診壓力計訊號以適當的放大電路處理，連接超音波系統，將血管的二維影像、血管直徑脈動(M-mode)與相對應的血壓波型同步顯示於螢幕；◎將螢幕中含數個心跳週期血管變化的影像記錄傳輸至電腦平台；◎以影像分析方法擷取同步之血管脈動波型、管壁內中層組織之厚度變形量與血壓波型；◎由影像掃描深度的刻度訂定影像比例尺度；◎調整血管等效管徑值與等效脈動值；◎自動化影像處理，將同步同位置取得的血壓與管徑變形資料，計算血管管徑與管壁內中層組織(intima-media thickness; IMT)之應力-應變關係曲線、分析其動態楊氏係數及能量消散率(EDR)。 如申請專利範圍第1項所述之非侵襲性血管機械黏彈特性的檢測方法，以彈性繃帶將夾具(超音波探頭與壓力計)固定於受測者血管上。 如申請專利範圍第1項所述之非侵襲性血管機械黏彈特性的檢測方法，以自動化影像處理方式將血管管徑脈動與觸診血壓變化同步分析。 如申請專利範圍第3項所述之非侵襲性血管機械黏彈特性的檢測方法，自動的血管邊緣影像搜尋，採最大灰階梯度所對應之最佳灰階閥值，以固定灰階閥值之影像內插法，追蹤血管壁邊緣隨時間的移動，並以灰階等位圖提供自動化的管徑脈動波型。 如申請專利範圍第1項所述之非侵襲性血管機械黏彈特性的檢測方法，以血管管壁之應力-應變關係(負載曲線)所衍生的機械性質參數，如硬度、動態楊氏係數及能量消散率(EDR)，以及各數值沿著管壁各層之變化。 如申請專利範圍第1項所述之非侵襲性血管機械黏彈特性的檢測方法，以管壁內中層組織(IMT)之應力-應變關係(負載曲線)所衍生的機械性質參數，如硬度(intima-media stiffness)、動態楊氏係數及能量消散率(intima-media EDR)等。 |  | A61B005/021 A61B008/00
US06871084 | B1 | 2005-03-22 | High-impedance optical electrode | 2001-07-03 | High-impedance optical electrodes modulate light in response to a life-form bio-potential and then converts the modulated light to an electrical signal that provides traditional EEG and EEC type output. Light splitters are used to provide multiple electrodes and an electronic reference source. A pilot tone is used to achieve high sensitivity and synchronize multiple units while an optical phase-shift modulator is used to reduce optical noise. | 1. A high-impedance optical electrode used for measuring bio-potentials comprising: a) a light source; b) an electro-optic modulator: (1) receiving light from said light source; (2) modulating said light in response to a bio-potential; and (3) providing a modulated light output proportional to said bio-potential; c) a photodetector for receiving and converting said modulated light output from said electro-optic modulator to an electrical signal; d) electronic circuitry for providing an electronic output signal; and e) a pilot tone generated by said electronic circuitry and superimposed on said bio-potential. | 1. A high-impedance optical electrode used for measuring bio-potentials comprising: a) a light source; b) an electro-optic modulator: (1) receiving light from said light source; (2) modulating said light in response to a bio-potential; and (3) providing a modulated light output proportional to said bio-potential; c) a photodetector for receiving and converting said modulated light output from said electro-optic modulator to an electrical signal; d) electronic circuitry for providing an electronic output signal; and e) a pilot tone generated by said electronic circuitry and superimposed on said bio-potential. 2. A high-impedance optical electrode used for measuring bio-potentials comprising: a) a light source; b) an electro-optic modulator: (1) receiving light from said light source; (2) modulating said light in response to a bio-potential; and (3) providing a modulated light output proportional to said bio-potential; c) an optical splitter for splitting said light from said light source into at least a second light portion wherein said second light portion is used as an optical reference signal. 3. The high-impedance optical electrode used for measuring bio-potentials according to claim 2 wherein said second a third light portion is received by a second electro-optical modulator. 4. The high-impedance optical electrode used for measuring bio-potentials according to claim 2 further comprising an optical phase-shift modulator. 5. A high impedance optical electrode for measuring bio-potentials comprising: a) a light source; b) a bio-potential; c) an electro-optic modulator; (1) receiving light from said light source; (2) modulating said light in response to a bio-potential; and (3) providing a modulated light output; d) a photodetector for receiving and converting said modulated light output from said electro-optic modulator into an electrical output; and e) wherein at least one end of said electro-optic modulator connections to at least one member of a group of members consisting of: an optical fiber, said light source, and said photodetector, is formed at an angle to vertical. 6. The high impedance optical electrode according to claim 5 wherein said electrical output is a voltage. 7. The high impedance optical electrode according to claim 5 wherein said light source is a laser diode. 8. The high impedance optical electrode according to claim 7 wherein said laser diode is a highly coherent laser diode. 9. The high impedance optical electrode according to claim 7 wherein said laser diode is a low coherent laser diode. 10. The high impedance optical electrode for measuring bio-potentials according to claim 5 wherein said light source is a distributed feedback laser. 11. The high impedance optical electrode for measuring bio-potentials according to claim 5 wherein said light source is a Fabry-Perot laser. 12. The high impedance optical electrode for measuring bio-potentials according to claim 5 wherein said light source is a vertical ... | Srico, Inc. | A61B005/04
US06884223 | B2 | 2005-04-26 | Method for detecting α-oxoaldehydes in the whole blood, blood plasma and/or serum of a patient | 2001-05-29 | A method for detecting α-oxoaldehydes in the whole blood, blood plasma and/or serum of a patient. A respiratory air sample is analyzed for the presence of at least one α-oxoaldehyde, and the presence of α-oxoaldehydes in the whole blood, blood plasma and/or serum of the patient is determined from the results of the analysis. | 1. A method for detecting α-oxoaldehydes in at least one of whole blood, blood plasma and serum of a patient, the method comprising analyzing a breath air sample for the presence of at least one α-oxoaldehyde and deducing the presence of α-oxoaldehydes in the at least one of whole blood, blood plasma and serum of the patient from the analysis result. | 1. A method for detecting α-oxoaldehydes in at least one of whole blood, blood plasma and serum of a patient, the method comprising analyzing a breath air sample for the presence of at least one α-oxoaldehyde and deducing the presence of α-oxoaldehydes in the at least one of whole blood, blood plasma and serum of the patient from the analysis result. 2. The method according to claim 1 wherein analyzing a breath air sample for the presence of at least one α-oxoaldehyde includes derivatizing the at least one α-oxoaldehyde in the breath air sample by means of a derivatization medium and determining the resulting derivatives. 3. The method according to claim 2 wherein determining the resulting derivatives includes determining the resulting derivatives by one of the following methods fluorometrically; photometrically; and, chromatographically. 4. The method according to claim 1 wherein analyzing a breath air sample for the presence of at least one α-oxoaldehyde includes at least one of enclosing the breath air sample in a sample bag, enclosing the breath air sample in another suitable sample container and directing the breath air sample through an impinger. 5. The method according to claim 4 wherein analyzing a breath air sample for the presence of at least one α-oxoaldehyde includes derivatizing the at least one α-oxoaldehyde in the breath air sample by means of a derivatization medium and determining the resulting derivatives. 6. The method according to claim 1 wherein analyzing a breath air sample for the presence of at least one α-oxoaldehyde and deducing the presence of α-oxoaldehydes in the at least one of whole blood, blood plasma and serum of the patient from the analysis result includes converting methylglyoxal which is contained in the breath air sample into lactate with glyoxalase and subsequently detecting the lactate with a sensor. 7. The method according to claim 1 wherein analyzing a breath air sample for the presence of at least one α-oxoaldehyde includes detecting the α-oxoaldehydes by at least one of the following methods: electrochemically; spectrometrically; enzymatically; and, immunochemically. 8. The method according to claim 1 wherein analyzing a breath air sample for the presence of at least one α-oxoaldehyde includes determining the α-oxoaldehydes in the breath air sample mass-spectrometrically. 9. The method according to claim 1 wherein analyzing a breath air sample for the presence of at least one α-oxoaldehyde includes determining the concentration of α-oxoladehydes in the breath air sample, and deducing the presence of α-oxoaldehydes in the at least one of whold blood, blood plasma and serum of the patient includes deducing the concentration of α-oxoaldehydes in the at least one of whole blood, blood plasma and serum of the patient therefrom. 10. The method according to claim 9 wherein analyzing a breath air sample for the presence of at least one α-oxoaldehyde includes analyzing a breath air sample for the presen... | Institut fur Chemo-und Biosensorik Munster E.V. | A61B005/08
US06882874 | B2 | 2005-04-19 | Compensation of human variability in pulse oximetry | 2002-02-15 | The invention relates to the calibration of a pulse oximeter intended for non-invasively determining the amount of at least two light-absorbing substances in the blood of a subject. In order to bring about a solution by means of which the effects caused by the tissue of the subject can be taken into account in connection with the calibration of a pulse oximeter, initial characterization measurements are carried out for a pulse oximeter calibrated under nominal conditions. Based on the characterization measurements, nominal characteristics are established describing the conditions under which nominal calibration has been defined, and reference data indicating the nominal characteristics are stored. In-vivo measurements are then performed on living tissue and based on the in-vivo measurements and the reference data stored, tissue-induced changes in the nominal characteristics are determined. Subject-specific variation in the in-vivo measurements is compensated for by correcting the nominal calibration on the basis of the tissue-induced changes. | 1. A method for compensating for subject-specific variability in an apparatus intended for non-invasively determining the amount of at least two light-absorbing substances in the blood of a subject and provided with emitter means for emitting radiation at a minimum of two different wavelengths and with detector means for receiving the radiation emitted, the method comprising the steps of calibrating the apparatus using a nominal calibration, carrying out initial characterization measurements, said measurements to include measuring radiation emitted by said emitter means and received by the detector means without transmission through tissue, based on the characterization measurements, establishing nominal characteristics describing conditions under which the nominal calibration is used, storing reference data indicating the nominal characteristics established, performing in-vivo measurements on living tissue, wherein radiation emitted by the emitter means through the living tissue and received by the detector means is measured, based on the in-vivo measurements and the reference data stored, determining tissue-induced changes in the nominal characteristics, and compensating for subject-specific variation in the in-vivo measurements by correcting the nominal calibration on the basis of the tissue-induced changes. | 1. A method for compensating for subject-specific variability in an apparatus intended for non-invasively determining the amount of at least two light-absorbing substances in the blood of a subject and provided with emitter means for emitting radiation at a minimum of two different wavelengths and with detector means for receiving the radiation emitted, the method comprising the steps of calibrating the apparatus using a nominal calibration, carrying out initial characterization measurements, said measurements to include measuring radiation emitted by said emitter means and received by the detector means without transmission through tissue, based on the characterization measurements, establishing nominal characteristics describing conditions under which the nominal calibration is used, storing reference data indicating the nominal characteristics established, performing in-vivo measurements on living tissue, wherein radiation emitted by the emitter means through the living tissue and received by the detector means is measured, based on the in-vivo measurements and the reference data stored, determining tissue-induced changes in the nominal characteristics, and compensating for subject-specific variation in the in-vivo measurements by correcting the nominal calibration on the basis of the tissue-induced changes. 2. A method according to claim 1, including compensation for effects causing wavelength shift. 3. A method according to claim 2, wherein the compensation for effects causing wavelength shift includes defining subject-specific extinction coefficients for the substances. 4. A method according to claim 3, wherein said establishing step includes determining DC transmission characteristics of the emitter and detector means, spectral characteristics of the emitter and detector means and nominal transmission characteristics for the tissue. 5. A method according to claim 4, wherein said establishing step further includes determining a temperature in which the nominal calibration is used. 6. A method according to claim 5, wherein the extinction coefficients εij are determined according to the following formula: ɛijeffective=1W⁢∫Δ⁢ ⁢λ⁢ɛj⁡(λ)*LEDi⁡(λ⁡(T))*DET⁡(λ)*tissue⁡(λ)⁢∂λ,where εijeffective is the effective extinction for the emitter means as defined by the integration over an emission spectrum LEDi(λ) of the emitter means, DET(λ) represents a spectral sensitivity of the detector means, tissue(λ) is a spectral transmission of radiation through the tissue, εj is the extinction coefficient of the substance, T is a temperature, λ is a wavelength of the radiation, W is a normalization factor, and i and j are matrix indices. 7. A method according to claim 6, wherein the step of establishing nominal characteristics includes defining a nominal extinction matrix with a nominal extinction coefficient for each substance/wavelength pair, and the step of determining tissue-induced changes includes updating the nominal extinction matrix, whereby the ... | Datex-Ohmeda, Inc. | A61B005/00
US06853854 | B1 | 2005-02-08 | Noninvasive measurement system | 2000-01-21 | The noninvasive measurement system provides a technique for manipulating wave data. In particular, wave data reflected from a biological entity is received, and the reflected wave data is correlated to a substance in the biological entity. The wave data may comprise light waves, and the biological entity may comprise a human being or blood. Additionally, a substance may comprise, for example, a molecule or ionic substance. The molecule may be, for example, a glucose molecule.;Furthermore, the wave data is used to form a matrix of pixels with the received wave data. The matrix of pixels may be modified by techniques of masking, stretching, or removing hot spots.;Then, the pixels may be integrated to obtain an integration value that is correlated to a glucose level. The correlation process may use a lookup table, which may be calibrated to a particular biological entity. Moreover, an amplitude and phase angle may be calculated for the reflected wave data and used to identify a glucose level in the biological entity.;The glucose level may be displayed on a monitor attached to the computer. The computer may be a portable, self-contained unit that comprises a data processing system and a wave reflection capture system. On the other hand, the computer may be attached to a network of other computers, wherein the reflected wave data is received by the computer and forwarded to another computer in the network for processing. | 1. A method of manipulating wave data, comprising: receiving wave data reflected from a biological entity; and, correlating the reflected wave data to a substance in the biological entity, wherein the correlating step comprises the steps of: calculating an amplitude for the reflected wave data; calculating a phase angle for the reflected wave data; and, identifying a glucose level in the biological entity by using the amplitude and the phase angle. | 1. A method of manipulating wave data, comprising: receiving wave data reflected from a biological entity; and, correlating the reflected wave data to a substance in the biological entity, wherein the correlating step comprises the steps of: calculating an amplitude for the reflected wave data; calculating a phase angle for the reflected wave data; and, identifying a glucose level in the biological entity by using the amplitude and the phase angle. 2. The method of claim 1, wherein the wave data comprises light waves. 3. The method of claim 2, wherein the biological entity is an eye, and the method further comprising the step of forming a plurality of pixels based on the reflected wave data. 4. The method of claim 3, further comprising: integrating the pixels to form an integrated value; and correlating the integrated value to a glucose level. 5. The method of claim 3, further comprising identifying a center of the eye. 6. The method of claim 3, wherein the eye comprises an iris and a pupil. 7. The method of claim 6, further comprising calculating an average brightness around the pupil of the eye. 8. The method of claim 7, further comprising equalizing the iris using the brightness around the pupil as a baseline. 9. The method of claim 6, further comprising masking the pupil of the eye. 10. The method of claim 2, wherein the received light waves are reflected from an eye and the method further comprises forming a matrix comprised of pixels based on the received light waves. 11. The method of claim 10, wherein the amplitude is calculated by summing all of the pixels. 12. The method of claim 10 wherein calculating the phase angle comprises: summing rows of pixels in the matrix to obtain an xGRU value; summing columns of pixels in the matrix to obtain a yGRU value; and calculating a ratio of the xGRU value and the yGRU value. 13. The method of claim 12, wherein the ratio is calculated by dividing a larger of the xGRU value and the yGRU value by a smaller of the xGRU value and the yGRU value. 14. The method of claim 10, further comprising masking a portion of the matrix. 15. The method of claim 10, further comprising calculating a true amplitude. 16. The method of claim 15, wherein calculating the true amplitude comprises subtracting a value based on the phase angle from a summation of pixels formed by the light waves. 17. The method of claim 2, further comprising applying a filter to the reflected light waves. 18. The method of claim 1, wherein the biological entity comprises a human being. 19. The method of claim 1, wherein the biological entity comprises blood. 20. The method of claim 19, wherein the substance comprises glucose molecules in the blood. 21. The method of claim 1, further comprising forming a matrix of pixels with the received wave data. 22. The method of claim 21, further comprising masking a portion of the matrix. 23. The method of claim 21, wherein calculating an amplitude comprises integrating the pixels... | Q Step Technologies, LLC | A61B005/00
US06885196 | B2 | 2005-04-26 | Biosensor | 2001-07-17 | A highly sensitive biosensor that uses a smaller amount of sample for measurement is described. The biosensor includes a first insulating base plate having a working electrode, a second insulating base plate having a counter electrode opposed to the working electrode, and a reagent layer containing at least an oxidoreductase. The biosensor further includes a sample supply pathway formed between the first and second insulating base plates, where the working electrode, counter electrode and reagent layer are exposed to an inside of the sample supply pathway, and the distance between the working electrode and the counter electrode is 150 μm or less. | 1. A biosensor comprising a first insulating base plate having a working electrode, a second insulating base plate having a counter electrode opposed to said working electrode, a reagent layer comprising at least an oxidoreductase, and a sample supply pathway formed between the first and second insulating base plates, wherein said working electrode, counter electrode and reagent layer are exposed to an inside of said sample supply pathway, and the distance between said working electrode and said counter electrode is 150 μm or less, and the area of said counter electrode exposed to the sample supply pathway is smaller than the area of said working electrode exposed to the sample supply pathway. | 1. A biosensor comprising a first insulating base plate having a working electrode, a second insulating base plate having a counter electrode opposed to said working electrode, a reagent layer comprising at least an oxidoreductase, and a sample supply pathway formed between the first and second insulating base plates, wherein said working electrode, counter electrode and reagent layer are exposed to an inside of said sample supply pathway, and the distance between said working electrode and said counter electrode is 150 μm or less, and the area of said counter electrode exposed to the sample supply pathway is smaller than the area of said working electrode exposed to the sample supply pathway. 2. The biosensor in accordance with claim 1, wherein the area of said working electrode exposed to the sample supply pathway is 0.1 to 2.0 mm2, and the area of said counter electrode exposed to the sample supply pathway is in a range of not less than 0.05 mm2 and less than 2.0 mm2. 3. The biosensor in accordance with claim 1, wherein the distance between said working electrode and said counter electrode is 40 μm or more. 4. The biosensor in accordance with claim 1, wherein a spacer member having a slit is interposed between the first and second insulating base plates such that said sample supply pathway is formed in said slit. 5. A biosensor comprising a first insulating base plate having a working electrode, a second insulating base plate having a counter electrode opposed to said working electrode, a reagent layer comprising at least an oxidoreductase, and a sample supply pathway formed between the first and second insulating base plates, wherein said working electrode, counter electrode and reagent layer are exposed to said sample supply pathway, and the area of said counter electrode exposed to the sample supply pathway is less than the area of said working electrode exposed to the sample supply pathway. 6. The biosensor in accordance with claim 5, wherein the distance between said working electrode and said counter electrode is 150 μm or less. | Matsushita Electric Industrial Co., Ltd. | A61B005/05
US06895268 | B1 | 2005-05-17 | Medical workstation, imaging system, and method for mixing two images | 2000-06-28 | In a system, method and workstation, images of a first subject are acquired with an image signal acquisition unit, the position of the image signal acquisition unit is determined, the position of a second subject is determined and the position of the second subject relative to the image signal acquisition unit is also determined and an image of the second subject is mixed into an image of the first subject acquired with the image signal acquisition unit. | 1. A system comprising: an image signal acquisition unit for acquiring 2D image signals of an examination subject in real time, and an imaging unit for producing a 2D image of the examination subject in an image plane from said 2D image signals; an instrument having a tip adapted for interaction with the examination subject; a navigation system including a position acquisition unit for determining a position of said image signal acquisition unit and for determining a position of said tip relative to said position acquisition unit; and a mixing unit connected to said imaging unit for mixing a representation of said tip into said 2D image and, if said tip is not located in said image plane, for determining a distance of said tip from said image plane and for mixing a designation of said distance into said 2D image, said designation being alterable and indicating a magnitude of said distance. | 1. A system comprising: an image signal acquisition unit for acquiring 2D image signals of an examination subject in real time, and an imaging unit for producing a 2D image of the examination subject in an image plane from said 2D image signals; an instrument having a tip adapted for interaction with the examination subject; a navigation system including a position acquisition unit for determining a position of said image signal acquisition unit and for determining a position of said tip relative to said position acquisition unit; and a mixing unit connected to said imaging unit for mixing a representation of said tip into said 2D image and, if said tip is not located in said image plane, for determining a distance of said tip from said image plane and for mixing a designation of said distance into said 2D image, said designation being alterable and indicating a magnitude of said distance. 2. A system as claimed in claim 1 wherein said navigation system includes identifiers, selected from the group consisting of detectable marks and position sensors, which are respectively attachable to said image signal acquisition unit and to said second subject and which are identifiable as to position by said position acquisition unit. 3. A system as claimed in claim 1 wherein said image signal acquisition unit comprises an ultrasound laparoscope. 4. A system as claimed in claim 1 wherein said image signal acquisition unit comprises an X-ray source and an X-ray receiver. 5. A system as claimed in claim 1 wherein said imaging unit produces a 3D image of said first subject from said image signals. 6. A system as claimed in claim 1 wherein said position acquisition unit simultaneously identifies the position of said image signal acquisition unit and the position of said second subject. 7. A system as claimed in claim 1 further comprising an acceptance device for said first subject and wherein said position acquisition device identifies a position of said acceptance device simultaneously with identifying the position of said image signal acquisition unit and the position of said second subject. 8. A system as claimed in claim 1 wherein, if said tip is not located in said image plane, said mixing unit mixes a projection of said tip in said image plane into said 2D image. 9. A system as claimed in claim 1 wherein said mixing unit mixes a circle, as said designation, into said 2D image, said circle having a diameter which is alterable dependent on said magnitude of said distance. 10. A system as claimed in claim 1 wherein said mixing unit mixes said designation into said 2D image with an alterable color intensity indicating said magnitude of said distance. 11. A system comprising: a C-arm X-ray image data acquisition unit for acquiring 3D image data of a patient, and an image unit for producing a 3D image of the patient from said image data; a support mechanism for support the first subject; a position acquisition system for determining a position of said... | Siemens Aktiengesellschaft | A61B005/05
US06866639 | B2 | 2005-03-15 | Handheld low voltage testing device | 2002-09-23 | A portable hand-held electrical testing device including a processor housed within an enclosure. The processor is configured to operate on commands by a user to process sub-microvolt electrical signals received through an input/output interface. The input/output interface includes a capacitive coupled amplifier with adjustable gain settings. Onboard memory linked to the processor stores processing data and instructions. A display device is mounted to said enclosure and is operatively connected to the processor to display processing results in real time. | 1. A sub-microvolt electrical signal testing device, comprising: a portable hand-held enclosure; a processor housed by said enclosure; a memory operatively coupled to said processor, said memory storing one or more operating instructions; a display device mounted to said enclosure, said display device being operatively connected to said processor; a multi-channel input-output interface operatively coupled to said processor with a bi-directional communications channel, said multi-channel input-output interface configured with at least one input channel to receive external electrical signals and communicate said external electrical signals to said processor; a plurality of electrical shock protection components coupled between said at least one input channel and said bi-directional communications channel; a switching network coupled between each of said plurality of electrical shock protection components and said bi-directional communications channel, said switching network configured to selectively couple one of a plurality of electrodes to said bi-directional communications channel; a differential operational amplifier coupled between said switching network and said bi-directional communications channel, said differential operational amplifier having an adjustable gain setting; and wherein said processor is configured to utilize said one or more operating instructions to perform one or more tests on signals received from said multi-channel input-output interface, and to display results of said one or more tests on said display device. | 1. A sub-microvolt electrical signal testing device, comprising: a portable hand-held enclosure; a processor housed by said enclosure; a memory operatively coupled to said processor, said memory storing one or more operating instructions; a display device mounted to said enclosure, said display device being operatively connected to said processor; a multi-channel input-output interface operatively coupled to said processor with a bi-directional communications channel, said multi-channel input-output interface configured with at least one input channel to receive external electrical signals and communicate said external electrical signals to said processor; a plurality of electrical shock protection components coupled between said at least one input channel and said bi-directional communications channel; a switching network coupled between each of said plurality of electrical shock protection components and said bi-directional communications channel, said switching network configured to selectively couple one of a plurality of electrodes to said bi-directional communications channel; a differential operational amplifier coupled between said switching network and said bi-directional communications channel, said differential operational amplifier having an adjustable gain setting; and wherein said processor is configured to utilize said one or more operating instructions to perform one or more tests on signals received from said multi-channel input-output interface, and to display results of said one or more tests on said display device. 2. The sub-microvolt electrical signal testing device of claim 1 wherein said bi-directional communications channel has a 20-bit signal resolution. 3. The sub-microvolt electrical signal testing device of claim 1 wherein said differential operational amplifier is coupled between said switching network and said bi-directional communications channel by one or more capacitors. 4. The sub-microvolt electrical signal testing device of claim 1 wherein said multi-channel input-output interface includes a common mode cancellation amplifier configured to reduce signal noise levels in said electrical signals. 5. The sub-microvolt electrical signal testing device of claim 1 wherein said multi-channel input-output interface includes an otoacoustic emission testing interface and an auditory brain stem testing interface; and wherein said processor is configured with said one or more stored operating instructions for auditory tests selected from the group comprising otoacoustic emission test procedures, auditory brainstem response test procedures, and combinations thereof for a test subject. 6. The sub-microvolt electrical signal testing device of claim 1 further including said processor having an internal memory, said processor configured to utilize said internal memory to execute said one or more operating instructions, whereby signal noise levels from external bus access are reduced. 7. The sub-microvolt electrical sign... | Everest Biomedical Instruments | A61B005/00 A61B005/04
US06889070 | B2 | 2005-05-03 | Magnetic resonance imaging apparatus and assembly process thereof | 2001-11-14 | Magnetic resonance imaging apparatus having a pair of a first magnet device and a second magnet device for generating a magnetostatic field, the pair of magnet devices being installed opposite each other across an examination space for accommodating an examinee; a gradient-field generating device; a high-frequency field generating device; and a yoke for combining the first and second magnetic devices to guide magnetic fluxes generated by the first and second magnetic devices to thereby form a closed magnetic circuit, wherein the yoke includes a first plate member fixed to a first magnet device, a second plate member fixed to a second magnet device and one or more support-post members interconnecting the first plate member and the second plate member, each of the first and second plate members and support-post members includes a plurality of segments formed in a predetermined shape. | 1. Magnetic resonance imaging apparatus comprising: a pair of a first magnet device and a second magnet device for generating a magnetostatic field, said pair of magnet devices being installed in face-to-face relation with each other across an examination space for accommodating an examinee; a gradient-field generating device; a high-frequency field generating device; and a yoke for combining said first and second magnetic devices to guide magnetic fluxes generated by said first and second magnetic devices to thereby form a closed magnetic circuit, wherein said yoke includes a first plate member fixed to a first magnet device, a second plate member fixed to a second magnet device and one or more support-post members interconnecting said first plate member and said second plate member, each of said first and second plate members and support-post members includes a plurality of segments formed in such a shape as to minimize leakage field strength from said first magnet device and said second magnet device. | 1. Magnetic resonance imaging apparatus comprising: a pair of a first magnet device and a second magnet device for generating a magnetostatic field, said pair of magnet devices being installed in face-to-face relation with each other across an examination space for accommodating an examinee; a gradient-field generating device; a high-frequency field generating device; and a yoke for combining said first and second magnetic devices to guide magnetic fluxes generated by said first and second magnetic devices to thereby form a closed magnetic circuit, wherein said yoke includes a first plate member fixed to a first magnet device, a second plate member fixed to a second magnet device and one or more support-post members interconnecting said first plate member and said second plate member, each of said first and second plate members and support-post members includes a plurality of segments formed in such a shape as to minimize leakage field strength from said first magnet device and said second magnet device. 2. Magnetic resonance imaging apparatus according to claim 1, wherein each of said first and second plate members and said support-post members includes a plurality of segments of different shapes and said segments are combined in a pattern to match lines of magnetic induction generated by said first and second magnetic devices. 3. Magnetic resonance imaging apparatus according to claim 2, wherein said first and second plate members have different numbers of segments at different positions based on a calculated magnetic flux distribution in said yoke, and said first and second plate members have segments of shapes to match a calculated magnetic flux distribution. 4. Magnetic resonance imaging apparatus according to claim 1, wherein said first and second magnet devices are superconducting magnet devices, and said superconducting magnet devices each include a cryostat unit containing a superconducting coil to cool said superconducting coil, said cryostat units are mounted between said plurality of segments. 5. Magnetic resonance imaging apparatus according to claim 1, wherein said support-post members each include two support posts for connecting said first and second plate members at separate positions, and said two support posts are each formed by a combination of a plurality of segments. 6. Magnetic resonance imaging apparatus according to claim 5, wherein said cryostat unit has mounted therein a part for achieving a desired field uniformity in said examination space. 7. Magnetic resonance imaging apparatus according to claim 1, wherein said shape to minimize said leakage field strength from said first and second magnet devices is a shape to match lines of magnetic induction generated by said first and second magnetic devices. 8. Magnetic resonance imaging apparatus according to claim 1, wherein each of said first and second plate members and said support-post members includes a plurality of segments of different shapes and said segments... | Hitachi Medical Corporation | A61B005/055
US06882872 | B2 | 2005-04-19 | Biological information detecting probe, biological information measuring apparatus, fabrication method for biological information detecting probe, and method of measuring biological information | 2001-02-06 | Provided are a biological information detecting probe and a biological information measuring apparatus that are easy to handle and can easily carry out highly accurate measurements by improving contact between a surface of a living body and a light sensor while minimizing damage to the body tissue, and a biological information measurement method for implementing the same. The invention also achieves a biological information measuring apparatus that can easily measure biological information concerning deep portions of a living body by improving contact between the surface of the living body and the light sensor, and a method for implementing the same. The biological information detecting probe comprises a raised portion 5 forming a recessed portion which is pressed against a living tissue in intimate contact relationship, a substrate 4, an exit end face 9 from which detection light exits through one part of the recessed portion, and a light receiving means 11 which is provided in another part of the recessed portion and into which the detection light is introduced, wherein with the raised portion 5 and substrate 4 held pressed against the living tissue, the detection light is passed through the living tissue and introduced into the receiving means 11. | 1. A biological information detecting probe by comprising: pressing means which is pressed against a living tissue; detection light emitting means of emitting detection light through light emission surface of said pressing means; and a detection light entrance which said detection light is introduced through a light detection surface of said pressing means; wherein: said light emission surface is substantially parallel to a first plane; said light detection surface is substantially parallel to a second plane; the first plane and the second plane form an angle of less than 180°; at least one of said detection light emitting means or said detection light entrance is formed from a material that has a higher refractive index than said living tissue such that; said detection light is capable of being passed rectilinearly through said living tissue placed in contact with said pressing means and is capable of being introduced into said detection light entrance; and said detection light is refracted at least at one of said light emission surface or said light detection surface with said pressing means held pressed against said living tissue. | 1. A biological information detecting probe by comprising: pressing means which is pressed against a living tissue; detection light emitting means of emitting detection light through light emission surface of said pressing means; and a detection light entrance which said detection light is introduced through a light detection surface of said pressing means; wherein: said light emission surface is substantially parallel to a first plane; said light detection surface is substantially parallel to a second plane; the first plane and the second plane form an angle of less than 180°; at least one of said detection light emitting means or said detection light entrance is formed from a material that has a higher refractive index than said living tissue such that; said detection light is capable of being passed rectilinearly through said living tissue placed in contact with said pressing means and is capable of being introduced into said detection light entrance; and said detection light is refracted at least at one of said light emission surface or said light detection surface with said pressing means held pressed against said living tissue. 2. The biological information detecting probe as set forth in claim 1, further comprising: a living tissue pressing part which is pressed against said living tissue and thereby deforms a portion of said living tissue; and a base part which contacts a portion of said living tissue other than the portion thereof against which said living tissue pressing part is pressed; wherein said pressing means is formed extending over said living tissue pressing part and said base part. 3. The biological information detecting probe as set forth in claim 2, wherein said living tissue pressing part and/or said base part include secretion removing means of removing secretion released from said living tissue, said secretion removing means being located in a portion contacting said living tissue. 4. The biological information detecting probe as set forth in claim 3, wherein said secretion removing means is a hole or a groove. 5. The biological information detecting probe as set forth in claim 2, wherein: said detection light emitting means is provided in said living tissue pressing part, and said detection light entrance is provided in said base part. 6. The biological information detecting probe as set forth in claim in 5, wherein said detection light emitting means and/or said detection light entrance means include an optical waveguide. 7. The biological information detecting probe as set forth in claim 6, wherein said optical waveguide has a Y-branch shape or a plate-like shape. 8. The biological information detecting probe as set forth in claim 6, wherein: said optical waveguide in said detection light emitting means is for receiving external input light, and an end face of said optical waveguide, from which said detection light is not emitted, is formed so as to guide said input light to an end face from which said detec... | Matsushita Electric Industrial Co., Ltd. | A61B005/00
US06843801 | B2 | 2005-01-18 | Method and system for identifying and displaying groups of cardiac arrhythmic episodes | 2002-04-08 | A medical device system that receives cardiac data representing a plurality of stored arrhythmic episodes, and analyzing the cardiac data to identify and display a subset of stored arrhythmic episodes as a function of user-specified episode criteria. The medical device system presents a query window on an interactive display in order to receive user-specified episode criteria via one or more input fields. The medical device displays only those episodes matching the episode criteria such as arrhythmia type, zone of detection, date of occurrence and average heart rate in beats per minute (BPM). | 1. A medical device system comprising: a first device including: first electronic circuitry adapted to receive a cardiac signal and sense a plurality of arrhythmic episodes from the cardiac signal; and a memory circuit adapted to store cardiac data representative of the plurality of arrhythmic episodes; and a second device, communicatively coupled to the first device, the second device including: second electronic circuitry adapted to receive the cardiac data from the first device and analyze the cardiac data to identify a subset of the plurality of arrhythmic episodes based on user-specified episode criteria; and an interactive user input and presentation device, coupled to the second electronic circuitry, the interactive user input and presentation device adapted to receive the user-specified episode criteria and present the identified subset of the plurality of arrhythmic episodes. | 1. A medical device system comprising: a first device including: first electronic circuitry adapted to receive a cardiac signal and sense a plurality of arrhythmic episodes from the cardiac signal; and a memory circuit adapted to store cardiac data representative of the plurality of arrhythmic episodes; and a second device, communicatively coupled to the first device, the second device including: second electronic circuitry adapted to receive the cardiac data from the first device and analyze the cardiac data to identify a subset of the plurality of arrhythmic episodes based on user-specified episode criteria; and an interactive user input and presentation device, coupled to the second electronic circuitry, the interactive user input and presentation device adapted to receive the user-specified episode criteria and present the identified subset of the plurality of arrhythmic episodes. 2. The system of claim 1, wherein the first device is a cardiac defibrillator. 3. The system of claim 1, wherein the first device is coupled to the second device via RF telemetry. 4. The system of claim 1, wherein the interactive user input and presentation device includes an interactive display screen adapted to display a query window having a plurality of input fields to receive the user-specified episode criteria and display the identified subset of the plurality of arrhythmic episodes. 5. The system of claim 4, wherein the interactive display screen is adapted to display the identified subset of the plurality of arrhythmic episodes in a tabular format. 6. The system of claim 1, wherein the cardiac data representative of the plurality of arrhythmic episodes include at least one of: an episode number; an episode time stamp indicating date and time of an occurrence; an episode type; an arrhythmia zone of detection; an episode heart rate; a therapy delivered prior to an episode detection; and episode detection enhancement criteria. 7. The system of claim 6, wherein the user-specified episode criteria include at least one of: an arrhythmia type; an arrhythmia zone of detection; a time of occurrence; an average heart rate; and a heart rate range. 8. The system of claim 7, wherein the second device further includes a printer, coupled to the second electronic circuitry, the printer adapted to print the identified subset of the plurality of arrhythmic episodes. 9. The system of claim 7, wherein the second device includes a removable storage medium, coupled to the second electronic circuitry, the removable storage medium adapted to store the identified subset of the plurality of arrhythmic episodes. 10. A system comprising: communication circuitry adapted to receive cardiac data representative of a plurality of arrhythmic episodes; processing circuitry, coupled to the communication circuitry, the processing circuitry adapted to analyze the cardiac data to identify a subset of the plurality of arrhythmic episodes based on user-specified episode criteria; and a... | Cardiac Pacemakers, Inc. | A61B005/0472 A61B005/044
US06843775 | B2 | 2005-01-18 | Blood drawing system | 2001-05-29 | A blood drawing system including a closeable first conduit through which drawn blood can flow; a first pressurizeable collection vessel for receiving blood therein; a hollow-shaft member; a second conduit leading from the first conduit to the first collection vessel; and a third conduit leading from the first conduit to the hollow shaft member connectable to a negative-pressure second collection vessel. Closing the first conduit allows blood transfer from the first vessel into the second vessel without exposure to or inclusion of any external devices to accomplish such transfer. Also included is an intravenous system with a conduit through which medicament liquid can be delivered and from which medicament liquid-free blood can be withdrawn. A first pressurizeable vessel first removes and conserves for reintroduction residual medicament liquid from the conduit, while thereafter a second negative-pressure vessel draws subsequent medicament liquid-free blood from the conduit. | 1. A blood drawing system for drawing blood from a blood vessel of a living being, the system comprising: a) a first conduit through which blood drawn from a blood vessel can flow, said first conduit having a closure member for selectively closing the first conduit at a closure site thereof; b) a first collection vessel for receiving blood and having an exteriorly-operable pressure controller for application of negative or positive pressure within said first collection vessel; c) a hollow-shaft member; d) a second conduit having a first end in fluid communication with the first conduit and a second end in fluid communication with the first collection vessel; and e) a third conduit having a first end in communication with the first conduit and a second end in communication with the hollow shaft member; wherein when said first conduit is closed by said closure member and positive pressure is applied within the first collection vessel by the pressure controller, liquid is transferred from the first collection vessel through the hollow shaft member. | 1. A blood drawing system for drawing blood from a blood vessel of a living being, the system comprising: a) a first conduit through which blood drawn from a blood vessel can flow, said first conduit having a closure member for selectively closing the first conduit at a closure site thereof; b) a first collection vessel for receiving blood and having an exteriorly-operable pressure controller for application of negative or positive pressure within said first collection vessel; c) a hollow-shaft member; d) a second conduit having a first end in fluid communication with the first conduit and a second end in fluid communication with the first collection vessel; and e) a third conduit having a first end in communication with the first conduit and a second end in communication with the hollow shaft member; wherein when said first conduit is closed by said closure member and positive pressure is applied within the first collection vessel by the pressure controller, liquid is transferred from the first collection vessel through the hollow shaft member. 2. A blood drawing system as claimed in claim 1 wherein the closure member of the first conduit is a clamp. 3. A blood drawing system as claimed in claim 1 wherein the first collection vessel is a syringe. 4. A blood drawing system as claimed in claim 3 wherein the pressure controller is a plunger within said syringe. 5. A blood drawing system as claimed in claim 1 wherein the hollow shaft member is a needle. 6. A blood drawing system for drawing blood from a blood vessel of a living being, the system comprising: a) a first conduit through which blood drawn from a blood vessel can flow, said first conduit having a clamp provided thereto for selectively closing the first conduit; b) a syringe for receiving blood and having an exteriorly-operable plunger for application of negative or positive pressure within the syringe; c) a hollow needle; d) a second conduit having a first end in fluid communication with the first conduit and a second end in fluid communication with the syringe; and e) a third conduit having a first end in communication with the first conduit and a second end in communication with the hollow needle; wherein when said first conduit is closed and positive pressure is applied within the syringe by the plunger, liquid is transferred from the syringe through the hollow needle. 7. A method of drawing blood from a blood vessel of a living being, the method comprising: a) connecting a blood vessel conduit leading from the blood vessel to a blood drawing system, the system comprising: i) a first conduit through which blood from the blood vessel conduit can flow, said first conduit having a closure member for selectively closing the first conduit at a closure site thereof; ii) a first collection vessel for receiving blood and having an exteriorly-operable pressure controller for application of negative or positive pressure within said first collection vessel; iii) a hollow-shaft member; i... | Smiths Medical ASD, Inc. | A61B005/00
US07070569 | B2 | 2006-07-04 | Non-invasive determination of conditions in the circulatory system of a subject | 2002-03-04 | A method for non-invasively determining functional cardiac output (FCO) and/or venous blood CO2 partial pressure (PvCO2). The amount of CO2 (VCO2N) released from the blood and end capillary blood CO2 content (CcCO2N) are determined from measurements from exhaled breathing gases. The CO2 content of the breathing gases inhaled by the subject is increased and values for VCO2R and CcCO2R are obtained. A regression analysis is performed using the obtained VCO2N, VO2R, CcCO2N, and CcCO2R values. The regression line is extrapolated to obtain a value for CcCO2 when (VCO2) is zero so that CvCO2 becomes known. The CvCO2 thus determined can be inserted in a non-differential form in the Fick equation, along with VCO2 and CcCO2 values from normal breathing, to determine FCO. To determine PvCO2, CvCO2 is altered in accordance with the amount of oxygen in the venous blood, to correctly indicate PvCO2. The continuing validity of the FCO measurement can be examined on a breath-by-breath basis by noting changes in an indicator variable, such as VCO2 or end tidal CO2 amounts. | 1. A method for non-invasively determining a condition of the circulatory system of a subject, the subject inhaling and exhaling breathing gases during breathing, said method comprising the steps of: (a) measuring the amount of CO2 in the breathing gases exhaled by the subject and the CO2 concentration of the breathing gases exhaled by the subject for a first (1) breathing condition of the subject;(b) determining at least one value of the amount of CO2 released from the circulatory system of the subject (VCO21) using the amount of CO2 in the breathing gases exhaled when the subject is in the first condition;(c) determining at least one value for a quantity indicative of the end capillary blood CO2 content of the subject using the CO2 concentration of the breathing gases exhaled when the subject in the first condition;(d) altering the CO2 concentration in the lungs of the subject;(e) measuring the amount of CO2 in the breathing gases exhaled by the subject and the CO2 concentration of the breathing gases exhaled by the subject for at least one breath of the subject under a second (2) breathing condition of altered CO2 in the lungs of the subject;(f) determining at least one value for the amount of CO2 released from the circulatory system of the subject (VCO22), the determination of the value being carried out in a time period less than that required for blood leaving the lungs of the subject to pass through the circulatory system of the subject and return to the lungs, the determination of the value using the amount of CO2 in the exhaled breathing gases for the second breathing condition;(g) determining at least one value for a quantity indicative of the end capillary blood CO2 content of the subject, the determination of the value being carried out in a time period less than that required for blood leaving the lungs of the subject to pass through the circulatory system of the subject and return to the lungs, the determination of the value using the CO2 concentration of the breathing gases exhaled for the second breathing condition;(h) performing a regression analysis using the determined VCO21, VCO22, and end capillary blood CO2 quantity values to establish a regression line;(i) extrapolating the regression line to obtain a value for the end capillary blood CO2 quantity when the amount of CO2 released from the circulatory system of the subject (VCO2) is zero;determining a further value for the quantity indicative of the end capillary blood CO2 content for breathing of the subject in the first breathing condition;forming a relationship between the value for the quantity indicative of the end capillary blood CO2 content for breathing in the first breathing condition used in the regression analysis and the value obtained by extrapolating the regression line in step (i); andapplying the relationship to the further determined value for a quantity indicative of the end capillary blood CO2 content to provide a new value for the value which was obtained ... | 1. A method for non-invasively determining a condition of the circulatory system of a subject, the subject inhaling and exhaling breathing gases during breathing, said method comprising the steps of: (a) measuring the amount of CO2 in the breathing gases exhaled by the subject and the CO2 concentration of the breathing gases exhaled by the subject for a first (1) breathing condition of the subject;(b) determining at least one value of the amount of CO2 released from the circulatory system of the subject (VCO21) using the amount of CO2 in the breathing gases exhaled when the subject is in the first condition;(c) determining at least one value for a quantity indicative of the end capillary blood CO2 content of the subject using the CO2 concentration of the breathing gases exhaled when the subject in the first condition;(d) altering the CO2 concentration in the lungs of the subject;(e) measuring the amount of CO2 in the breathing gases exhaled by the subject and the CO2 concentration of the breathing gases exhaled by the subject for at least one breath of the subject under a second (2) breathing condition of altered CO2 in the lungs of the subject;(f) determining at least one value for the amount of CO2 released from the circulatory system of the subject (VCO22), the determination of the value being carried out in a time period less than that required for blood leaving the lungs of the subject to pass through the circulatory system of the subject and return to the lungs, the determination of the value using the amount of CO2 in the exhaled breathing gases for the second breathing condition;(g) determining at least one value for a quantity indicative of the end capillary blood CO2 content of the subject, the determination of the value being carried out in a time period less than that required for blood leaving the lungs of the subject to pass through the circulatory system of the subject and return to the lungs, the determination of the value using the CO2 concentration of the breathing gases exhaled for the second breathing condition;(h) performing a regression analysis using the determined VCO21, VCO22, and end capillary blood CO2 quantity values to establish a regression line;(i) extrapolating the regression line to obtain a value for the end capillary blood CO2 quantity when the amount of CO2 released from the circulatory system of the subject (VCO2) is zero;determining a further value for the quantity indicative of the end capillary blood CO2 content for breathing of the subject in the first breathing condition;forming a relationship between the value for the quantity indicative of the end capillary blood CO2 content for breathing in the first breathing condition used in the regression analysis and the value obtained by extrapolating the regression line in step (i); andapplying the relationship to the further determined value for a quantity indicative of the end capillary blood CO2 content to provide a new value for the value which was obtained ... | Instrumentarium Corp. | A61B005/08 A61B005/02
US07229447 | B1 | 2007-06-12 | Photodynamic therapy utilizing a solution of photosensitizing compound and surfactant | 2000-02-26 | The invention relates to a method of photoeradication of cells at a tissue site, such as an infection or sterilization site or cancer cell activity site, including applying a solution such as methylene blue, toluidene blue, polymyxin B, SDS, or other surfactants and combinations thereof to the tissue site and exposing the tissue site with a light having a light wavelength and light dosage and a light dosage rate. The solution may have a concentration of methylene blue between 5 μg/ml to about 100 μg/ml. The wavelength may range from about 610 nm to about 670 nm. The light dosage may range from about 0 J/cm2 to about 200 J/cm2. The light dosage rate may range from about 0 mw/cm2 to about 150 mw/cm2. Treatable organisms include candida albicans, escherechia coli, pseudomonas aeruginosa, staphylococcus aureus, streptococcus pneumoniae, and clostridia. | 1. A method of photodynamic disruption of cells comprising the steps of: identifying an area of cell activity;applying a concentration including a combination of a surfactant and a photosensitizing agent to the area of cell activity, said surfactant disorienting a cell membrane so that said membrane no longer functions as an effective osmotic barrier, and so that said photosensitizing agent is able to pass through the disoriented cell membrane; andexposing the area of cell activity to light having a light wavelength, light dosage and a light dosage rate to cause photodynamic cellular disruption, wherein the surfactant is SDS provided in a solution having an SDS concentration range of between 0.003% to 0.01%. | 1. A method of photodynamic disruption of cells comprising the steps of: identifying an area of cell activity;applying a concentration including a combination of a surfactant and a photosensitizing agent to the area of cell activity, said surfactant disorienting a cell membrane so that said membrane no longer functions as an effective osmotic barrier, and so that said photosensitizing agent is able to pass through the disoriented cell membrane; andexposing the area of cell activity to light having a light wavelength, light dosage and a light dosage rate to cause photodynamic cellular disruption, wherein the surfactant is SDS provided in a solution having an SDS concentration range of between 0.003% to 0.01%. 2. The method of photodynamic disruption of cells of claim 1 wherein the step of identifying an area of cell activity includes an examination of a portion of a living body. 3. The method of photodynamic disruption of cells of claim 1 wherein the light wavelength ranges from about 400 nm to about 800 nm, the light dosage ranges from about 10 J/cm2 to about 100 J/cm2 and the light dosage rate ranges from about 50 mw/cm2 to about 200 mw/cm2. 4. The method of photodynamic disruption of cells of claim 1 wherein the wavelength ranges from about 600 nm to about 700 nm. 5. A method of photodynamic disruption of acellular organisms comprising the steps of: identifying an area of acellular organism activity;applying a concentration including a combination of a surfactant and a photosensitizing agent to the area of acellular organism activity, said surfactant disorienting an acellular organism membrane so that said membrane no longer functions as an effective osmotic barrier, and so that said photosensitizing agent is able to pass through the disoriented acellular organism membrane; andexposing the area of acellular organism activity to light having a light wavelength, light dosage and a light dosage rate, wherein the surfactant is SDS provided in a solution having an SDS concentration range of between 0.003% to 0.01%. 6. The method of photodynamic disruption of acellular organisms of claim 5 wherein the step of identifying an area of acellular activity includes the step of identifying an area of virus activity. 7. The method of photodynamic disruption of acellular organisms of claim 5, wherein the step of identifying an area of acellular organism activity includes an examination of a portion of a living body. 8. The method of photodynamic disruption of acellular organisms of claim 5, wherein the light wavelength ranges from about 400 nm to about 800 nm, the light dosage ranges from about 10 J/cm2 to about 100 J/cm2 and the light dosage rate ranges from about 50 mw/cm2 to about 200 mw/cm2. 9. The method of photodynamic disruption of acellular organisms of claim 5 wherein the wavelength ranges from about 600 nm to about 700 nm. | Advanced Photodynamics Technologies, Inc. | A61B005/02
US07087068 | B1 | 2006-08-08 | Lancet | 2001-08-21 | A skin pricker has a barrel (1,11) in which a hammer (2,29) released by a trigger mechanism (32,34) can be shot forwards by a compressed spring (3,38) to impact on a lancet (24) and momentarily project its tip. The rear end of the spring acts against a barrier (4,39) adjustable axially of the barrel, so that the spring can be more or less compressed before release, causing the hammer to act with greater or lesser force on the lancet. The barrier (4,39) may have projections (6,8,42) that engage in skew slots (7,9,43) in the barrel, so that rotation of the barrier axially adjusts it. A sleeve (44) over the rear part of the barrel instrumental in priming the device may co-operate with the projections and be rotatable to set the desired spring force. | 1. A medical skin piercing device comprising a barrel, means for carrying a lancet in the forward part of the barrel to allow the tip of the lancet to advance from a retracted to a projecting position, a hammer arranged when released from a rearward position to impact on the rear of the lancet to cause such advance, a barrier to the rear of the hammer and user-adjustable axially within the barrel by cam action as a result of rotation of the barrier, a spring acting between the barrier and the hammer, and a trigger mechanism for holding the hammer in and releasing the hammer from said rearward position with the spring compressed to a degree determined by the axial adjustment of the barrier, wherein the user-set degree of compression of the spring determines the force propelling the hammer after release of the hammer,wherein the barrel has slots skew to the axis of the barrel in which projections on the barrier engage,wherein the rear portion of the barrel is encased by a captive sleeve spring-urged forwardly, the sleeve having a lost motion connection through the rear end of the barrel and through the barrier to the hammer, whereby pulling back the sleeve retracts the hammer to said rearward position, and release of the sleeve allows the sleeve to revert to its forward position disconnected from the hammer, andwherein the sleeve when pulled back reveals the slots for adjustment of the projections in the slots. | 1. A medical skin piercing device comprising a barrel, means for carrying a lancet in the forward part of the barrel to allow the tip of the lancet to advance from a retracted to a projecting position, a hammer arranged when released from a rearward position to impact on the rear of the lancet to cause such advance, a barrier to the rear of the hammer and user-adjustable axially within the barrel by cam action as a result of rotation of the barrier, a spring acting between the barrier and the hammer, and a trigger mechanism for holding the hammer in and releasing the hammer from said rearward position with the spring compressed to a degree determined by the axial adjustment of the barrier, wherein the user-set degree of compression of the spring determines the force propelling the hammer after release of the hammer,wherein the barrel has slots skew to the axis of the barrel in which projections on the barrier engage,wherein the rear portion of the barrel is encased by a captive sleeve spring-urged forwardly, the sleeve having a lost motion connection through the rear end of the barrel and through the barrier to the hammer, whereby pulling back the sleeve retracts the hammer to said rearward position, and release of the sleeve allows the sleeve to revert to its forward position disconnected from the hammer, andwherein the sleeve when pulled back reveals the slots for adjustment of the projections in the slots. 2. A device as claimed in claim 1, wherein the slots have short portions non-skew to said axis to locate the projections in set positions. 3. A device as claimed in claim 1, wherein the projections are on resilient formations integral with the barrier, allowing the projections to be moved radially inwards for insertion of the barrier into the barrel, the projections springing outwardly when they register with the slots. 4. A device as claimed in claim 1, wherein a nose section of the barrel is removable to expose the lancet carrying means and allow lancets to be removed and replaced. 5. A medical skin piercing device, comprising a barrel, means for carrying a lancet in the forward part of the barrel to allow the tip of the lancet to advance from a retracted to a projecting position, a hammer arranged when released from a rearward position to impact on the rear of the lancet to cause such advance, a barrier to the rear of the hammer and user-adjustable axially within the barrel by cam action as a result of rotation of the barrier, a spring acting between the barrier and the hammer, and a trigger mechanism for holding the hammer in and releasing the hammer from said rearward position with the spring compressed to a degree determined by the axial adjustment of the barrier, wherein the user-set degree of compression of the spring determines the force propelling the hammer after release of the hammer,wherein the barrel has slots skew to the axis of the barrel in which projections on the barrier engage,wherein the rear portion of the barrel is... | Owen Mumford Limited | A61B017/34 A61B005/151
US07087025 | B2 | 2006-08-08 | Blood pressure determination based on delay times between points on a heartbeat pulse | 2003-01-29 | Blood pressure determination with resolution sufficient to resolve Pulsus Paradoxus is based on measurement of delay time between points on a heartbeat pulse. | 1. A method of monitoring blood pressure, which comprises: determining delay time between predetermined points on a heartbeat pulse, and using the delay time to obtain a measurement of blood pressure, wherein the predetermined points on the heartbeat pulse are on an onset slope of the pulse, and wherein plural delay times are measured between predetermined points on the onset slope that are located on a portion of the onset slope that extends between the peak of the pulse and 50% of the peak of the pulse. | 1. A method of monitoring blood pressure, which comprises: determining delay time between predetermined points on a heartbeat pulse, and using the delay time to obtain a measurement of blood pressure, wherein the predetermined points on the heartbeat pulse are on an onset slope of the pulse, and wherein plural delay times are measured between predetermined points on the onset slope that are located on a portion of the onset slope that extends between the peak of the pulse and 50% of the peak of the pulse. 2. A method according to claim 1, wherein the plural delay times are averaged. 3. Apparatus for monitoring blood pressure, comprising: a sensor that detects a heartbeat pulse at a site on a body through which blood flows, and a unit that determines delay time between predetermined points on the front side or between predetermined points on the back side of the detected heartbeat pulse and that uses the delay time to obtain a measurement of blood pressure, wherein the unit measures plural delay times between the predetermined points on the heartbeat pulse. 4. Apparatus according to claim 3, wherein the plural delay times are averaged. 5. A method of monitoring blood pressure, which comprises: determining plural delay times between predetermined points on the front side or between predetermined points on the back side of a heartbeat pulse, and using the plural delay times to obtain a measurement of blood pressure. 6. A method according to claim 5, wherein the plural delay times are averaged. 7. A process of determining changes in blood pressure, which comprises: repeating the method of claim 5 for each of a series of heartbeat pulses over a period of time; and producing an output of blood pressure measurement over the period of time. 8. A process of determining changes in blood pressure, which comprises: repeating the method of claim 6 for each of a series of heartbeat pulses over a period of time; and producing an output of blood pressure measurement over the period of time. 9. A process according to claim 7, wherein variations of the blood pressure output are used to determined Pulsus Paradoxus. 10. A process according to claim 9, wherein the output is filtered to isolate Pulsus Paradoxus modulations. | Empirical Technologies Corporation | A61B005/00
US07268155 | B2 | 2007-09-11 | Photosensitizer conjugates for pathogen targeting | 2002-05-09 | Conjugate molecules which include photosensitizer compositions conjugated to non-antibody non-affinity pair targeting moieties and methods of making and using such conjugates are described. | 1. A method of treating a subject, for a disorder characterized by the presence of an unwanted organism, comprising: administering to the subject a conjugate comprising a polylysine backbone to which is coupled a targeting moiety and a porphyrin photosensitizer, wherein the targeting moiety is a cationic antimicrobial peptide; irradiating the subject with energy of wavelength appropriated to produce a cytotoxic effect by the photosensitizer; thereby treating the subject, for the disorder characterized by the presence of an unwanted organism, wherein the unwanted organism is a bacterium and is located in the oral cavity including throat and tonsil, in the sinus, in the ear, in the nose, in the peritoneal cavity, or on the epidermis. | 1. A method of treating a subject, for a disorder characterized by the presence of an unwanted organism, comprising: administering to the subject a conjugate comprising a polylysine backbone to which is coupled a targeting moiety and a porphyrin photosensitizer, wherein the targeting moiety is a cationic antimicrobial peptide; irradiating the subject with energy of wavelength appropriated to produce a cytotoxic effect by the photosensitizer; thereby treating the subject, for the disorder characterized by the presence of an unwanted organism, wherein the unwanted organism is a bacterium and is located in the oral cavity including throat and tonsil, in the sinus, in the ear, in the nose, in the peritoneal cavity, or on the epidermis. 2. The method of claim 1, wherein the source of said irradiation is a laser. 3. The method of claim 1, wherein said bacterium is located in a wound. 4. The method of claim 1, wherein in addition to said disorder, said subject is suffering from an acquired immune disorder. 5. The method of claim 1, wherein said disorder is a periodontal disorder. 6. The method of claim 1, wherein the method includes topically administering the conjugate to an area of the subject which is infected with an unwanted organism. 7. The method of claim 1, wherein the conjugate is administered to periodontal tissue. 8. The method of claim 1, wherein said bacterium is an oral bacterial species. 9. The method of claim 1, wherein said bacterium is Porphyromonas (Bacteroides) gingivalis. 10. The method of claim 5, wherein the periodontal disorder is a periodontal infection. 11. The method of claim 1, wherein the disorder is a tooth surface infection. 12. The method of claim 1, wherein the disorder is a soft tissue infection. 13. The method of claim 12, wherein the soft tissue infection is an abscess. 14. The method of claim 3, wherein the wound is infected. 15. The method of claim 3, wherein the wound is a burn. 16. The method of claim 1, wherein the bacterium is an anaerobic bacterium. 17. The method of claim 1, wherein the bacterium is selected from the group consisting of a Porphyromonas, Bacteroides, Peptostreptococcus, Prevotella, Streptococcus, Actinobacillus, Fusobacterium, Pseudomonas, Staphylococcus, Helicobacter, Escherichia, Klebsiella, Enterococcus, Capnocytophage, Haemophilus, Salmonella, Chlamydia, Vibrio, Borrelia, Campylobacter, Bordetella, Neisseria, Legionella, Shigella, Erwinia, Enterobacter, Actinobacillus, Listeria, Yersinia, Propionibacterium and Corynebacterium. 18. The method of claim 17, wherein the Bacteroides is Bacteroides forsythus. 19. The method of claim 17, wherein the Actinobacillus is Actinobacillus actinomycetemcomitans. 20. The method of claim 17, wherein the Fusobacterium is Fusobacterium nucleatum. 21. The method of claim 17, wherein the Pseudomonas is Pseudomonas aeruginosa. 22. The method of claim 17, wherein the Staphylococcus is Staphylococcus aureus. 23. The method of claim 17,... | The General Hospital Corporation | A61K031/40 A61K038/00 A61B005/00
US07097826 | B2 | 2006-08-29 | Chlorin and bacteriochlorin-based difunctional aminophenyl DTPA and N | 2002-06-21 | A compound comprising a chemical combination of a photodynamic tetra-pyrrolic compound with a plurality of radionuclide element atoms such that the compound may be used to enhance MR imaging and also be used as a photodynamic compound for use in photodynamic therapy to treat hyperproliferative tissue. The preferred compounds have the structural formula:; where R1, R2, R2a R3, R3a R4, R5, R5a R6, R7, R7a, and R8 cumulatively contain at least two functional groups that will complex or combine with an MR imaging enhancing element or ion. The compound is intended to include such complexes and combinations and includes the use of such compounds for MR imaging and photodynamic therapy treatment of tumors and other hyperproliferative tissue. | 1. A composition comprising a chemical combination of a photodynamic tetra-pyrrolic compound having a porphyrin or chlorin structure having four pendant groups attached to at least one pyrolle ring, said compound being provided with a plurality of radionuclide element atoms connected to a plurality of functional groups that will complex or combine with an MR imaging enhancing element or ion selected from the group consisting of —(CH2)2CONHphenyleneCH2DTPA; —CH2CH2CON(CONHphenylene-CH2DTPA)2such that the compound may be used to enhance MR imaging wherein —CH2CH2CON(CONHphenylene-CH2DTPA)2 counts as two such functional groups. | 1. A composition comprising a chemical combination of a photodynamic tetra-pyrrolic compound having a porphyrin or chlorin structure having four pendant groups attached to at least one pyrolle ring, said compound being provided with a plurality of radionuclide element atoms connected to a plurality of functional groups that will complex or combine with an MR imaging enhancing element or ion selected from the group consisting of —(CH2)2CONHphenyleneCH2DTPA; —CH2CH2CON(CONHphenylene-CH2DTPA)2such that the compound may be used to enhance MR imaging wherein —CH2CH2CON(CONHphenylene-CH2DTPA)2 counts as two such functional groups. 2. The composition of claim 1 wherein the radionuclide element is selected from the group consisting of Technetium99, Indium111, radioactive iodine and a Lanthanum series element including Gadolinium. 3. The composition of claim 1 wherein the radionuclide element forms cations and the compound is a chelate of the radionuclide element with the porphyrin or chlorin structure. 4. The composition of claim 1 wherein the radionuclide element forms anions and the compound is a direct chemical combination of the radionuclide element with a porphyrin or chlorin structure. 5. A compound having the structural formula:where R1, R2, R2a R3, R3a R4, R5, R5a R6, R7, R7a, and R8 cumulatively contain at least two functional groups that complex or combine with an MR imaging enhancing element or ion; R1 is —CH═CH2, —CHO, COOH,where R9=—OR10 where R10 is lower alkyl of 1 through 8 carbon atoms, —((CH2)2O)nCH3 where n is 0 through 3, or —(CH2)2CONHphenyleneCH2DTPA; R2, R2a, R3, R3a, R4, R5, R5a, R7, and R7a are independently hydrogen, lower alkyl or substituted lower alkyl so as to contain a functional group that will combine with a radionuclide element selected from the group consisting of Technetium99, Indium111, radioactive iodine and a Lanthanum series element including Gadolinium or two R2, R2a, R3, R3a, R5, R5a, R7, and R7a groups on adjacent carbon atoms may be taken together to form a covalent bond or two R2, R2a, R3, R3a, R5, R5a, R7, and R7a groups on the same carbon atom may form a double bond to a divalent pendant group; R2 and R3 may together form a 5 or 6 membered heterocyclic ring containing oxygen, nitrogen or sulfur; R6 is —CH2—, —NR11— or a covalent bond; R8 is —(CH2)2CO2CH3, —(CH2)2CONHphenyleneCH2DTPA, —CH2CH2CON(CONHphenyleneCH2DTPA)2, —CH2R11 orwhere R11 isand polynuclide complexes thereof; wherein the functional groups that will complex or combine with an MR imaging enhancing element or ion are selected from the group consisting of —(CH2)2CONHphenyleneCH2DTPA; —CH2CH2CON(CONHphenyleneCH2DTPA)2;wherein —CH2CH2CON(CONHphenylene-CH2DTPA)2 counts as two such functional groups. 6. The compound of claim 5 wherein the compound is a dinuclide complex of gadolinium III. 7. The compound of claim 6 where R8 is —CH2CH2CON(CONHphenyleneCH2DTPA)2. 8. The compound of claim 7 wherein R2, R2a, R3, R3a, R4, R5, R5a, R7, and R7a are in... | Health Research, Inc. | A61B005/055
US07171252 | B1 | 2007-01-30 | Methods, computer program products, and devices for calibrating chronically tissue implanted sensors using chronically tissue | 2000-03-29 | Calibration of in vivo oxygen and pH sensor systems can be performed by generating a constituent element of an environment proximate to an in vivo sensor electrode via an in vivo generating electrode and determining a level of the constituent element in the tissue via the in vivo sensor electrode. Accordingly, accurate monitoring of tissue can be achieved while reducing the need to calibrate the in vivo sensor systems using invasive procedures. Related electrode assemblies are also discussed. | 1. A method for calibrating an in vivo sensor system, the method comprising the steps of: generating a constituent element of an environment proximate to an in vivo sensor electrode via an in vivo generating electrode; anddetermining a level of the constituent element in the tissue via the in vivo sensor electrode, wherein the step of generating comprises the steps of:generating a first constituent element at a first time at the in vivo generating electrode; andgenerating a second constituent element at a second time at the in vivo generating electrode. | 1. A method for calibrating an in vivo sensor system, the method comprising the steps of: generating a constituent element of an environment proximate to an in vivo sensor electrode via an in vivo generating electrode; anddetermining a level of the constituent element in the tissue via the in vivo sensor electrode, wherein the step of generating comprises the steps of:generating a first constituent element at a first time at the in vivo generating electrode; andgenerating a second constituent element at a second time at the in vivo generating electrode. 2. A method according to claim 1 further comprising the step of: calibrating the in vivo sensor circuit based on the level of the constituent element using a chronically tissue-implanted device. 3. A method according to claim 1, wherein the step of determining comprises the steps of: determining a first level of the constituent element between the first time and the second time; anddetermining a second level of the constituent element after the second time. 4. A method according to claim 1, wherein the step of generating the first constituent element comprises the step of generating oxygen, and wherein the step of generating the second constituent element comprises the step of generating hydrogen. 5. A method according to claim 1, wherein the step of generating the first constituent element comprises the step of generating an oxygen saturated environment proximate to the in vivo sensor electrode, and wherein the step of generating the second constituent element comprises the step of generating a hydrogen saturated environment proximate to the in vivo sensor electrode. 6. A method according to claim 1, wherein the step of generating comprises the step of generating the constituent element using electrolysis. 7. A method according to claim 1, wherein the step of generating the second constituent element is preceded by the step of waiting a time interval to allow the generated first constituent element to dissipate. 8. A method according to claim 7, wherein the time interval comprises about 1 minute. 9. A method according to claim 1, wherein the step of generating the first constituent element comprises the step of conducting a current from a first in vivo generating electrode to a second in vivo generating electrode through the tissue. 10. A method according to claim 1, wherein the step of calibrating comprises the step of determining a relationship between a current conducted by the in vivo sensor electrode and a third level of the constituent element using the first and second levels of the constituent element. 11. A method according to claim 10, wherein the step of calibrating comprises the steps of: determining a maximum rate of change of the voltage; andassociating the maximum rate of change of the voltage with a voltage level that is associated with a neutral pH level. 12. A method according to claim 1, wherein the step of generating a constituent element of an environment proximate... | Sicel Technologies, Inc. North Carolina State University | A61B005/05 A61B005/07 A61B006/00
US06996487 | B2 | 2006-02-07 | Automatic calibration system for computer-aided surgical instruments | 2001-03-15 | A system for automatic calibration of instruments (S) having varying cross-sectional dimensions within a predetermined range and having detectable elements (110, 112, 113, 114) thereon for computer-aided surgery, comprising a calibration base (C) having detectable elements (43, 44, 46, 48) secured thereto for detecting a position and an orientation thereof in space by sensors (204) connected to a position calculator (202). The calibration base is adapted to receive and to releasably secure a working shaft (100) of any of the instruments (S) and provides an abutting surface (14) for a tip (102) thereof in such a way that a position and orientation of the tip (102) of the instrument (S) secured therein is calculable when working shaft cross-section dimensions thereof are known. The position calculator (202) receives instrument data (214) and calibration data (218) from an operator through a user interface (206) and stores the instrument data (214) and calibration data (218) for subsequent calibrations. | 1. A calibration base to automatically calibrate any of a plurality of instruments having varying cross-sectional dimensions within a predetermined range for computer-aided surgery, each of the instrument, having a detectable portion, a working shaft and a tip at an end of the working shaft, said calibration base comprising; first detectable means secured thereto for detecting a position and an orientation thereof in space by sensor means connected to a position calculator; and said calibration base adapted to receive and to releasably secure the working shaft of any of the instruments and providing a first abutting surface for the tip thereof in such a way that a position and orientation of the tip of any of the instruments secured therein is calculable when working shaft cross-section dimensions thereof are known; whereby any of the instruments is calibrated when the position and orientation of the tip of the working shaft thereof is calculated. | 1. A calibration base to automatically calibrate any of a plurality of instruments having varying cross-sectional dimensions within a predetermined range for computer-aided surgery, each of the instrument, having a detectable portion, a working shaft and a tip at an end of the working shaft, said calibration base comprising; first detectable means secured thereto for detecting a position and an orientation thereof in space by sensor means connected to a position calculator; and said calibration base adapted to receive and to releasably secure the working shaft of any of the instruments and providing a first abutting surface for the tip thereof in such a way that a position and orientation of the tip of any of the instruments secured therein is calculable when working shaft cross-section dimensions thereof are known; whereby any of the instruments is calibrated when the position and orientation of the tip of the working shaft thereof is calculated. 2. The calibration base as defined in claim 1, wherein the first abutting surface extends in a direction generally opposed to a longitudinal axis of the working shaft of any one of the instruments. 3. The calibration base as defined in claim 2, wherein the working shaft of any of the instruments is abutted against a second and a third abutting surface. 4. The calibration base as defined in claim 3, wherein said second, and said third abutting surfaces define a V-shaped channel. 5. The calibration base as defined in claim 4, wherein said first, second end third abutting surfaces are each in a perpendicular relation with respect one to another. 6. The calibration base as defined claim 1, wherein the working shaft of any one of the instruments is releasably secured against abutting surfaces of said calibration bass by a biasing member. 7. The calibration base as defined in claim 6, wherein said biasing member comprises a spring-loaded lever. 8. The calibration base as defined in claim 7, wherein said lever has a generally rounded end surface for engaging contact with the working shaft. 9. The calibration base as defined in claim 1, wherein said first detectable means comprises three optically detectable spheres. 10. The calibration base as defined in claim 9, wherein each said detectable sphere is retro-reflective. 11. The calibration base as defined in claim 9, wherein said detectable spheres are detachable from said calibration base. 12. The calibration base as defined in claim 1, wherein said first detectable means are disposed in a concavity of said calibration base for impact protection thereof. 13. The calibration base as defined in claim 12, wherein said concavity is comprised of at least a pair of protrusions. 14. A method for calibrating any of a plurality of instruments having varying cross-sectional dimensions within a predetermined range with a position calculator connected to sensor means for computer-aided surgery, any one of the instruments having first detectable means being s... | Orthosoft Inc. | A61B005/05
US07189208 | B1 | 2007-03-13 | Method for measuring heart electrophysiology | 2000-04-12 | A mapping catheter is positioned in a heart chamber, and active electrode sites are activated to impose an electric field within the chamber. The blood volume and wall motion modulates the electric field, which is detected by passive electrode sites on the preferred catheter. Electrophysiology measurements, as well as geometry measurements, are taken from the passive electrodes and used to display a map of intrinsic heart activity. | 1. A method of presenting medical diagnostic information comprising: (a) measuring the anatomy of an organ to create an organ data file representing the anatomy of the organ at a specific time, said anatomy having a surface representing a wall surface of the organ;(b) measuring an electrophysiologic variable having a spatial relationship with the organ;(c) generating an electrophysiologic data file representing the dynamic behavior of said electrophysiologic variable; and(d) combining said organ data file and said electrophysiologic data file by displaying the dynamic behavior of said electrophysiologic variable on an image representing said wall surface of the organ created from said static organ data file. | 1. A method of presenting medical diagnostic information comprising: (a) measuring the anatomy of an organ to create an organ data file representing the anatomy of the organ at a specific time, said anatomy having a surface representing a wall surface of the organ;(b) measuring an electrophysiologic variable having a spatial relationship with the organ;(c) generating an electrophysiologic data file representing the dynamic behavior of said electrophysiologic variable; and(d) combining said organ data file and said electrophysiologic data file by displaying the dynamic behavior of said electrophysiologic variable on an image representing said wall surface of the organ created from said static organ data file. 2. A method of presenting cardiac medical diagnostic information comprising: (a) measuring the anatomy of an organ by collecting data throughout a heart beat cycle creating a data set representing the anatomy of the organ through the heart beat cycle;(b) reducing the data set to provide a geometric representation of the anatomy at a specific time in the heart beat cycle to create an organ data file, the anatomy having a surface representing a wall surface of the organ;(c) measuring an electrophysiologic variable having a spatial relationship with the organ;(d) generating an electrophysiologic data file representing dynamic behavior of said electrophysiologic variable; and(e) combining said organ data file and said electrophysiologic data file by displaying the dynamic behavior of said electrophysiologic variable on an image of said wall surface of the organ created from said organ data file. 3. A method of presenting cardiac medical diagnostic information wherein comprising: (a) measuring the anatomy of an organ by collecting data throughout a heart beat cycle to create a data set representing the anatomy of the organ through the heart beat cycle;(b) reducing said data set to provide a geometric representation of the anatomy at a specific time in the heart beat cycle to create an organ data file, the anatomy having a surface representing a surface of the organ;(c) measuring an electrophysiologic variable having a spatial relationship with the organ;(d) generating an electrophysiologic data file representing dynamic behavior of said electrophysiologic variable; and(e) combining said organ data file and said electrophysiologic data file by displaying the dynamic behavior of said electrophysiologic variable on an image created from said organ data file, whereinstep (a) comprises collecting data throughout a heart beat cycle with a catheter electrode in contact with a moving wall of the organ to create a data set representing the anatomy of the organ throughout the heart beat cycle and then reducing the data set to provide a geometric representation of an anatomic surface of the organ at a specific time in the heart beat cycle. 4. The method of presenting cardiac medical diagnostic information of claim 2 wherein: step (c) comprises generating ... | Endocardial Solutions, Inc. | A61B005/103 A61B005/02
US07215988 | B1 | 2007-05-08 | Sensor for measuring tissue perfusion | 1999-10-04 | The invention relates to a method and a sensor for measurement of tissue perfusion. The sensor is provided with a reservoir (4) for a fluid or gaseous tracer and a tracer-permeable barrier (3), a sub-surface of which is in contact with the surrounding tissue and another sub-surface of which is in contact with a detection cavity (5) which is connected to a suitable apparatus for the measurement of tracer concentration in the detection cavity. The concentration of the tracer in the detection cavity is a measure of perfusion in the surrounding tissue. According to another embodiment of the invention it is also possible to carry out measurements of perfusion in the surface layers of the skin or of an organ. | 1. Sensor for the measurement of tissue perfusion where a fluid or gaseous tracer is being supplied from a tracer source via a reservoir (4) to the tissue, the perfusion and of which is to be measured, and detected by a detection device via a detection cavity (5) comprising: said reservoir (4) having a reservoir wall with a tracer-permeable reservoir wall portion that permits tracer to be supplied from the reservoir to surrounding tissue, said detection cavity having a detection cavity wall having a tracer-permeable detection cavity wall portion, said tracer-permeable reservoir wall portion and the tracer-permeable detection wall portion respectively, communicating with the surroundings, such that a part of the tracer leaving said reservoir (4) can arrive at said detection cavity (5), the reservoir and the detection cavity are mutually interspaced, elongated cavities and the tracer-permeable reservoir wall portion and the tracer permeable detection cavity wall portion are elongated side wall portions. | 1. Sensor for the measurement of tissue perfusion where a fluid or gaseous tracer is being supplied from a tracer source via a reservoir (4) to the tissue, the perfusion and of which is to be measured, and detected by a detection device via a detection cavity (5) comprising: said reservoir (4) having a reservoir wall with a tracer-permeable reservoir wall portion that permits tracer to be supplied from the reservoir to surrounding tissue, said detection cavity having a detection cavity wall having a tracer-permeable detection cavity wall portion, said tracer-permeable reservoir wall portion and the tracer-permeable detection wall portion respectively, communicating with the surroundings, such that a part of the tracer leaving said reservoir (4) can arrive at said detection cavity (5), the reservoir and the detection cavity are mutually interspaced, elongated cavities and the tracer-permeable reservoir wall portion and the tracer permeable detection cavity wall portion are elongated side wall portions. 2. Sensor according to claim 1, wherein said reservoir (4) communicates partly with said surrounding tissue through a spatially extended tracer-permeable barrier (3), having a first surface (18) which forms a first area, and partly with said detection cavity (5) through the same spatially extended tracer permeable barrier (3), having a second surface (13) which forms a second area, said tracer-permeable barrier is a common permeable wall for said reservoir and said detection cavity, said detection cavity wall is distinct from said common permeable wall, and said reservoir wall is distinct from said common permeable wall. 3. Sensor according to claim 1, wherein said reservoir (4) communicates with said surrounding tissue through a spatially extended tracer-permeable barrier (14), a first surface (14′) of which forms a first area, and partly with said detection cavity (5) via said tissue and through another spatially extended tracer-permeable barrier (15), a second surface (15′) of which forms a second area. 4. Sensor according to claim 1, wherein said reservoir (4) and said detection cavity (5) are separated by a barrier (3, 19), and the reservoir (4), barrier (3, 19) and cavity (5) are built together to form a longitudinal sensor. 5. Sensor according to claim 1, wherein said reservoir (4), said detection cavity (5) are separated by spatially extended tracer-permeable barriers (3, 14, 15) are adapted to be located between a large surface of a tracer-impermeable panel or disc (17) and the surface (20) of the skin or organ of a patient, the perfusion of the surface layers of which skin or organ is to be measured, and a longitudinal axis 11 extending substantially parallel with said large surface of the panel or disc (17), such that said spatially extended tracer-permeable barriers (3, 14, 15) are partly in contact with the surface of the skin or organ, and such that tracer can move from said reservoir (4) into said skin or organ and either from her... | - | A61B005/00 A61B005/05
US07011815 | B2 | 2006-03-14 | Diagnostic imaging contrast agents with extended blood retention | 2003-01-30 | The present invention relates to contrast agents for diagnostic imaging with prolonged blood retention. In particular, this invention relates to novel compounds that are characterized by an image enhancing moiety (IEM); a protein plasma binding moiety (PPBM); and a blood half-life extending moiety (BHEM). This invention also relates to pharmaceutical compositions comprising these compounds and to methods of using the compounds and compositions for blood half-life extension and contrast enhancement of diagnostic imaging. | 1. A method of MR imaging of a mammal, said method comprising the steps of: (a) administering to said mammal a contrast agent or a pharmaceutically acceptable salt thereof, said contrast agent having the formula: IEM-[Lm-{BHEMs-PPBMo}p]q,wherein m, s, o, p, and q are one;wherein said IEM is an image enhancing moiety selected from the group consisting of: wherein one of R1–R11 is -[Lm-{BHEMs-PPBMo}p] and the R1–R11 groups that are not -[Lm-{BHEMs-PPBMo}p] are hydrogen;wherein R12, R13, and R14 can be the same or different and are selected from the group consisting of O− and NH2;wherein R15 is H, CH2CH(OH)CH3, hydroxy alkyl, or CH2COR12;wherein said M is a paramagnetic metal ion selected from the group consisting of Gd(III), Fe(III), Mn(II), Mn(III), Cr(III), Cu(II), Dy(III), Tb(III), Ho(III), Er(III), and Eu(III);wherein said L is a linker consisting of 1 to 4 —CH2— groups;wherein said BHEM is a phosphodiester moiety; andwherein said PPBM is an HSA binding moiety; and(b) subjecting said mammal to MR imaging. | 1. A method of MR imaging of a mammal, said method comprising the steps of: (a) administering to said mammal a contrast agent or a pharmaceutically acceptable salt thereof, said contrast agent having the formula: IEM-[Lm-{BHEMs-PPBMo}p]q,wherein m, s, o, p, and q are one;wherein said IEM is an image enhancing moiety selected from the group consisting of: wherein one of R1–R11 is -[Lm-{BHEMs-PPBMo}p] and the R1–R11 groups that are not -[Lm-{BHEMs-PPBMo}p] are hydrogen;wherein R12, R13, and R14 can be the same or different and are selected from the group consisting of O− and NH2;wherein R15 is H, CH2CH(OH)CH3, hydroxy alkyl, or CH2COR12;wherein said M is a paramagnetic metal ion selected from the group consisting of Gd(III), Fe(III), Mn(II), Mn(III), Cr(III), Cu(II), Dy(III), Tb(III), Ho(III), Er(III), and Eu(III);wherein said L is a linker consisting of 1 to 4 —CH2— groups;wherein said BHEM is a phosphodiester moiety; andwherein said PPBM is an HSA binding moiety; and(b) subjecting said mammal to MR imaging. 2. The method of claim 1, wherein said pharmaceutically acceptable salt is selected from N-methyl-D-glucamine, calcium, and sodium. 3. The method according to claim 1, wherein said mammal is a hiunan. 4. The method according to claim 1, wherein said MR imaging comprises imaging at least a portion of the vasculature of said mammal. 5. The method according to claim 1, wherein a portion of said imaging occurs 10 minutes or later after administration of said contrast agent. 6. The method according to claim 1, wherein said contrast agent is administered to said mammal at a concentration between about 0.001 to about 1 mmol/kg body weight. 7. The method according to claim 6, wherein said contrast agent is administered to said mammal at a concentration between about 0.005 to about 0.05 mmnol/kg body weight. 8. The method according to claim 1, wherein said contrast agent is selected from the group consisting of MS-315, MS-317, MS-321, MS-322, MS-325, MS-326, MS-327, and MS-328. 9. The method according to claim 1, wherein said imaging of said mammal comprises imaging a tumor of said mammal. 10. The method according to claim 1, wherein said imaging of said mammal comprises imaging the brain of said mammal. 11. The method according to claim 10, wherein said imaging of said brain comprises imaging the blood-brain barrier of said brain. 12. The method according to claim 10, wherein said imaging of said brain occurs when said mammal is undergoing cognitive events. | Epix Pharmaceuticals, Inc. | A61B005/055
US07260426 | B2 | 2007-08-21 | Method and apparatus for tracking an internal target region without an implanted fiducial | 2002-11-12 | A method and apparatus for locating an internal target region during treatment without implanted fiducials is presented. The method comprises producing a plurality of first images that show an internal volume including the internal target region, then producing a live image of the internal volume during treatment and matching this live image to one of the plurality of first images. Since the first images show the internal target region, matching the live image to one of the first images identifies the position of the target region regardless of whether the second image itself shows the position of the target region. The first images may be any three-dimensional images such as CT scans, magnetic resonance imaging, and ultrasound. The live image may be, for example, an x-ray image. The invention may be used in conjunction with a real-time sensor to track the position of the target region on a real-time basis. | 1. A method for locating an internal target region during treatment, the method comprising: generating a plurality of first images of a relevant volume of a patient, the relevant volume including the internal target region and each of the plurality of first images showing the internal target region;generating a live image of the relevant volume at a predetermined time during treatment; anddetermining a location of the internal target region at the predetermined time by identifying one of the plurality of first images that best matches the live image. | 1. A method for locating an internal target region during treatment, the method comprising: generating a plurality of first images of a relevant volume of a patient, the relevant volume including the internal target region and each of the plurality of first images showing the internal target region;generating a live image of the relevant volume at a predetermined time during treatment; anddetermining a location of the internal target region at the predetermined time by identifying one of the plurality of first images that best matches the live image. 2. The method of claim 1, wherein the plurality of first images include one of computer tomography scans, ultrasound images, and magnetic resonance images. 3. The method of claim 1, wherein the plurality of first images are three-dimensional images. 4. The method of claim 1, wherein generating the plurality of first images comprises: generating a first computer tomography scan taken at a first point of a respiratory cycle;generating a second computer tomography scan taken at a second point of the respiratory cycle; andgenerating at least one computer-generated intermediate three-dimensional image depicting the relevant volume at a third point between the first point and the second point. 5. The method of claim 4, wherein generating the at least one intermediate three-dimensional image comprises interpolating between the first computer tomography scan and the second computer tomography scan and modifying one of the first and the second computer tomography scans. 6. The method of claim 4, wherein generating the at least one intermediate three-dimensional image comprises continuously deforming a first three-dimensional image into a second three-dimensional image taken at a different point in time from the first three-dimensional image. 7. The method of claim 1, wherein the plurality of first images are generated at preselected points in the patient's respiratory pattern. 8. The method of claim 1, further comprising computing a plurality of digitally reconstructed radiographs for each of the plurality of first images, wherein each of the plurality of digitally reconstructed radiographs represents the relevant volume from a particular viewing angle. 9. The method of claim 8, wherein each of the first images indicate a point in the respiratory cycle and each of the digitally reconstructed radiographs indicate a rotational and translational shift of the relevant volume. 10. The method of claim 8, wherein identifying one of the plurality of first images that best matches the second image comprises comparing the digitally reconstructed radiographs to the second image. 11. The method of claim 10, wherein the comparing comprises comparing the digitally reconstructed radiographs and the second image on a pixel by pixel basis. 12. The method of claim 1, wherein the second image comprises an x-ray image. 13. The method of claim 1, wherein generating the plurality of first images is done independently o... | Accuray Incorporated | A61B005/00
US07013173 | B2 | 2006-03-14 | Optical probe with reference fiber | 2001-11-29 | A system for characterizing tissue includes the steps of generating an emission signal, generating a reference signal, directing the emission signal to and from the tissue, directing the reference signal in a predetermined manner relative to the emission signal, and using the reference signal to compensate the emission signal. In one embodiment compensation is provided for fluctuations in light delivery to the tip of the probe due to cable motion. | 1. A system for producing a signal adapted to characterize tissue and to correct for optical fluctuations, consisting of: a laser module for generating an emission signal and a reference signal,a splitter connected to said laser module for splitting said emission signal and said reference signal,a probe with a needle that is adapted to be inserted into the tissue and that directs said emission signal to the probe and through the needle to the tissue for characterizing the tissue,an emission optical detector,a reference optical detector,a transmission system includingan emission optical fiber connected to said said splitter, to said probe, to said needle, and to said emission optical detector that transmits said emission signal to said probe, to said needle, and from said probe to the tissue for characterizing the tissue, and from the tissue to said needle, to said probe and to said emission optical detector, anda reference optical fiber connected to said said splitter, to said probe, and to said reference optical detector that transmits said reference signal to said probe and from said probe to said reference optical detector, anda compensation system that utilizes said reference signal to correct said emission signal and producing a signal for characterizing the tissue for the optical fluctuations. | 1. A system for producing a signal adapted to characterize tissue and to correct for optical fluctuations, consisting of: a laser module for generating an emission signal and a reference signal,a splitter connected to said laser module for splitting said emission signal and said reference signal,a probe with a needle that is adapted to be inserted into the tissue and that directs said emission signal to the probe and through the needle to the tissue for characterizing the tissue,an emission optical detector,a reference optical detector,a transmission system includingan emission optical fiber connected to said said splitter, to said probe, to said needle, and to said emission optical detector that transmits said emission signal to said probe, to said needle, and from said probe to the tissue for characterizing the tissue, and from the tissue to said needle, to said probe and to said emission optical detector, anda reference optical fiber connected to said said splitter, to said probe, and to said reference optical detector that transmits said reference signal to said probe and from said probe to said reference optical detector, anda compensation system that utilizes said reference signal to correct said emission signal and producing a signal for characterizing the tissue for the optical fluctuations. 2. The system for producing a signal adapted to characterize tissue of claim 1, wherein said laser module comprises multiple lasers that generate said optical emission signal and said optical reference signal. | The Regents of the University of California | A61B005/00
US07018341 | B2 | 2006-03-28 | Methods and apparatus for stroke patient treatment | 2001-12-11 | Methods and apparatus for assessing the condition of and treating patients for stroke during the delivery of continuous positive airway pressure (CPAP) are disclosed. By determining central and obstructive apnea occurrences or the percentile of pressure delivered to the patient from patient airflow, stroke indicators may be calculated and analyzed to provide information on the type of stroke a patient has suffered and the degree of a patient's neuro-recovery. The apparatus may be programmed with a particular protocol intended to identify between alternative forms of CPAP treatment appropriate for the patient based upon an assessment of the patient's condition. The device can select and recommend between the various treatment forms, identify billing reimbursement codes and generate data to track patient response to stroke treatments. | 1. A method for evaluating the condition of a patient after stroke comprising the steps of: delivering breathable gas at a pressure above atmospheric to a patient after stroke;measuring the airflow of said patient; anddetermining a stroke indicator from said airflow, said determining step including analyzing said airflow to determine occurrences of central apnea and occurrences of obstructive apnea,wherein said stroke indicator is calculated as a function of a ratio of said occurrences of central apnea and said occurrences of obstructive apnea. | 1. A method for evaluating the condition of a patient after stroke comprising the steps of: delivering breathable gas at a pressure above atmospheric to a patient after stroke;measuring the airflow of said patient; anddetermining a stroke indicator from said airflow, said determining step including analyzing said airflow to determine occurrences of central apnea and occurrences of obstructive apnea,wherein said stroke indicator is calculated as a function of a ratio of said occurrences of central apnea and said occurrences of obstructive apnea. 2. The method of claim 1 further comprising recording said stroke indicator over time. 3. The method of claim 2 further comprising the step of analyzing said stroke indicator over time to determine a change in said stroke indicator. 4. The method of claim 1 wherein said determining step includes analyzing said airflow to determine a first percentile pressure from pressure delivered to a patient in a current session and comparing said first percentile pressure with a second percentile pressure from pressure delivered to said patient in a prior session. 5. The method of claim 4 wherein said first and second percentile pressures are 95th percentile pressures. 6. The method of claim 3 or 5 wherein said stroke indicator indicates a degree of neuro-recovery of the patient. 7. The method of claim 1 wherein said stroke indicator indicates a type of stroke that the patient has suffered. 8. The method of claim 1 further comprising the step of recording said stroke indicator in a database of stroke indicators for multiple patients. 9. The method of claim 1 further comprising the step of identifying subsequent treatment based upon said stroke indicator. 10. The method of claim 9 wherein said treatment is a form of CPAP therapy. 11. The method of claim 9 wherein said treatment is a change in drug therapy. 12. A method for evaluating the condition of a patient after stroke comprising the steps of: delivering breathable gas at a pressure above atmospheric to a patient after stroke;measuring the airflow of said patient;determining a stroke indicator from said airflow, said stroke indicator representing information about the patient's condition, andevaluating changes in said stroke indicator to assess the efficacy of an administered drug, said evaluating step including storing said stroke indicator in a database of patient information, said stroke indicator being a function of the number of occurrences of central apnea and the number of occurrences of obstructive apnea,wherein said stroke indicator is a function of a ratio of said number of occurrences of central apnea and said number of occurrences of obstructive apnea. | Resmed Limited | A61B005/08
US06858006 | B2 | 2005-02-22 | Cardiopulmonary monitoring | 2001-09-07 | According to the invention, a system for monitoring health information of a plurality of patients is disclosed. Included in the system are a data gathering device, a node, a wide area network, and a monitoring center. The data gathering device includes a respiration sensor, a wireless transmitter, and a self-contained power source. The node is wirelessly coupled to the data gathering device and the wide area network is coupled to the node. Located remote to the node, the monitoring center is coupled to the wide area network. | 1. A system for monitoring health information of a plurality of patients, the system comprising: a data gathering device comprising: a respiration sensor, a wireless transmitter, and a self-contained power source; a node wirelessly coupled to the data gathering device; a wide area network coupled to the node; and a monitoring center located remote to the node and coupled to the wide area network, wherein: the respiration sensor measures movement between two points on skin of a one of the plurality of patients where the two points are proximate to the one's chest, the two points are between the respiration sensor and the skin of the one, and an area of two square inches or less is covered by the data gathering device. | 1. A system for monitoring health information of a plurality of patients, the system comprising: a data gathering device comprising: a respiration sensor, a wireless transmitter, and a self-contained power source; a node wirelessly coupled to the data gathering device; a wide area network coupled to the node; and a monitoring center located remote to the node and coupled to the wide area network, wherein: the respiration sensor measures movement between two points on skin of a one of the plurality of patients where the two points are proximate to the one's chest, the two points are between the respiration sensor and the skin of the one, and an area of two square inches or less is covered by the data gathering device. 2. The system for monitoring health information of the plurality of patients as recited in claim 1, further comprising a plurality of nodes where each is linked to both a different patient and the wide area network. 3. The system for monitoring health information of the plurality of patients as recited in claim 1, wherein the monitoring center reports medical data relating to diagnosing at least one of: lung disease, heart disease, sleep apnea, arrhythmia, ischemia, congestive heart failure, and sudden infant death syndrome. 4. The system for monitoring health information of the plurality of patients as recited in claim 1, wherein the data gathering device further includes a digitizer to digitize an analog signal from the respiration sensor. 5. The system for monitoring health information of the plurality of patients as recited in claim 1, wherein the node comprises: a portable device worn on one of the plurality of patients, and a wireless connection to the wide area network. 6. The system for monitoring health information of the plurality of patients as recited in claim 1, wherein the wide area network includes at least one of: a wireless datalink, a telephone line, an Internet connection, and a private network. 7. The system for monitoring health information of the plurality of patients as recited in claim 1, wherein the wireless transmitter has a range of up to about 300 feet. 8. The system for monitoring health information of the plurality of patients as recited in claim 1, further comprising a second data gathering device comprising at least three of the following: a heart rate sensor, a pulse pressure sensor, a blood pressure sensor, an oxymetry sensor, an exhaled CO2 sensor, an electrocardiogram sensor, a peripheral tissue edema sensor, a second wireless transmitter, and a second self-contained power source. 9. The system for monitoring health information of the plurality of patients as recited in claim 1, further comprising a medical professional interface proximate to the monitoring center, wherein the medical professional interface provides heart rate and respiration data to a medical professional. 10. The system for monitoring health information of the plurality of patients as recited in claim 1, wherein: th... | Wireless Medical, Inc. | A61B005/00
US06904315 | B2 | 2005-06-07 | Atrial aware VVI: a method for atrial synchronous ventricular (VDD/R) pacing using the subcutaneous electrode array and a standard pacing lead | 2000-12-14 | An algorithm is implemented in a circuit for sensing P-waves in a pacemaker to ensure ventricular pacing synchronization with sensed atrial depolarization waves. VDD and VDDR pacing (atrial synchronized, ventricular inhibited pacing) are implemented via a single standard ventricular pacing lead (unipolar or bipolar) and preferably a subcutaneous electrode array (SEA). Specifically, an implanted ventricular lead provides ventricular pacing and ventricular sensing while the SEA enable atrial sensing, thus eliminating the need for an implanted atrial lead or a specialized single pass VDD lead. The algorithm manages the sensed cardiac waves to effect a desired pacing regimen based on the input from the single lead and SEA. | 1. A software system implemented in a circuit for sensing P-waves in a pacemaker, the system in combination with the circuit comprising: means for detecting a plurality of atrial depolarization P-wave signals wherein said means for detecting comprises at least two subcutaneous electrodes in data communication with means for pacing including a P-wave morphology detector; and means for pacing a ventricle synchronously with each one of said plurality of atrial depolarization P-wave signals detected by the at least two subcutaneous electrodes; wherein said means for pacing further comprises at least one pacing lead and wherein a first of said at least one pacing lead is adapted to electrically couple to a ventricular chamber. | 1. A software system implemented in a circuit for sensing P-waves in a pacemaker, the system in combination with the circuit comprising: means for detecting a plurality of atrial depolarization P-wave signals wherein said means for detecting comprises at least two subcutaneous electrodes in data communication with means for pacing including a P-wave morphology detector; and means for pacing a ventricle synchronously with each one of said plurality of atrial depolarization P-wave signals detected by the at least two subcutaneous electrodes; wherein said means for pacing further comprises at least one pacing lead and wherein a first of said at least one pacing lead is adapted to electrically couple to a ventricular chamber. 2. The system of claim 1 wherein said means for pacing is a single chamber ventricular-inhibited pacemaker. 3. The system of claim 2 wherein said pacemaker includes a hermetically sealed case including said at least two subcutaneous electrodes being peripherally distributed about the perimeter of the case. 4. The system of claim 2 wherein the at least one pacing lead comprises a ventricular pacing lead. 5. The system of claim 4 wherein said ventricular pacing lead is one of a unipolar pacing lead and a bipolar pacing lead. 6. A sensing circuit operating in co-operation with a pacemaker, a lead and at least one subcutaneous electrode array (SEA) implemented for pacing the ventricle synchronous with atrial depolarization signals, the circuitry comprising: an analog to digital converter (ADC) for converting a plurality of cardiac depolarization signals; a detector for detecting at least one of said plurality of cardiac depolarization signals coupled to said analog to digital converter (ADC); a digital to analog converter (DAC) coupled to the detector to convert at least some of the signals passing through said detector; a means for R-wave detection adapted to mechanically and electrically couple to a ventricular chamber; and a means for P-wave detection, wherein the means for P-wave detection is disposed on an external portion of an implantable medical device, and wherein signals from both the means for R-wave detection and signals from the means for P-wave detection are electrically coupled to a P-wave morphology detector having a single P-wave detection output signal and the P-wave detection signal in electrically coupled to said digital to analog converter (DAC). 7. The circuit of claim 6 wherein said circuit further comprises: a plurality of signal inputs, wherein said plurality of signal inputs further comprise: a signal input into said analog to digital converter (ADC) for a ventricular electrogram (VEGM) data signal from said lead; a signal input into said analog to digital converter (ADC) for a electrocardiogram (EGG) data signal from said at least one subcutaneous electrode array (SEA); and a signal input into said analog to digital converter (ADC) for an electrocardiogram (EGG) data signal from an external lead. ... | Medtronic, Inc. | A61N001/18 A61B005/04
US06970734 | B2 | 2005-11-29 | Flexible marker bands | 2002-12-02 | A catheter assembly comprises a catheter shaft having a first region and a second region. The first region is less radiopaque than the second region. The first region is constructed of a shaft material and the second region is constructed from the shaft material in combination with at least one radiopaque material. At least a portion of the shaft material of the second region defines at least one grooved region. The at least one radiopaque material is positioned at least partially within the at least one grooved region. | 1. A catheter assembly comprising: a catheter shaft, the catheter shaft comprising a first region and a second region, the first region being less radiopaque than the second region, the first region being constructed of a material, at least a portion of the second region being defined by a grooved pattern of the material in combination with at least one radiopaque material, the at least one radiopaque material comprises at least one radiopaque metal, the at least one radiopaque metal is at least one wire, the grooved pattern comprising a plurality of interconnected grooves, wherein at least some of the grooves have a wave-like pattern extending about the circumference of the second region. | 1. A catheter assembly comprising: a catheter shaft, the catheter shaft comprising a first region and a second region, the first region being less radiopaque than the second region, the first region being constructed of a material, at least a portion of the second region being defined by a grooved pattern of the material in combination with at least one radiopaque material, the at least one radiopaque material comprises at least one radiopaque metal, the at least one radiopaque metal is at least one wire, the grooved pattern comprising a plurality of interconnected grooves, wherein at least some of the grooves have a wave-like pattern extending about the circumference of the second region. 2. The catheter assembly of claim 1 wherein the first region comprises an outer diameter and the second region comprises an outer diameter, the outer diameter of the first region being substantially equal to the outer diameter of the second region. 3. The catheter assembly of claim 1 wherein the material is a polymeric material. 4. The catheter assembly of claim 1 wherein the material is selected from at least one member of the group consisting of: HDPE, PTFE, Nylon, block copolymers of polyamide and polyether, polyether-ester copolymer, and any combinations thereof. 5. The catheter assembly of claim 1 wherein the at least one radiopaque metal is selected from at least one member of the group consisting of: gold, tungsten, iridium, rhodium, platinum, barium, bismuth, CoFeO4, MnFe2O4, MgFe2O4, and any combinations thereof. 6. The catheter assembly of claim 1 wherein the at least one wire comprises a plurality of wires. 7. The catheter assembly of claim 1 wherein the second region defines a wall thickness, the wall thickness being about 0.002 to about 0.008 inches. 8. The catheter assembly of claim 7 wherein the at least one radiopaque material defines about 0.001 to about 0.004 inches of the wall thickness. 9. The catheter of claim 1 wherein the at least one radiopaque material is at least one polymer in combination with at least one member of the group consisting of gold, tungsten, iridium, rhodium, platinum, barium, bismuth, CoFeO4, MnFe2O4, MgFe2O4, and any combinations thereof. 10. The catheter of claim 1 wherein the at least one grooved region has a depth of about 0.001 to about 0.004 inches. 11. The catheter of claim 1 wherein the grooved pattern corresponds to the pattern of interconnected members. | Boston Scientific SciMed, Inc. | A61B005/05
US06850798 | B2 | 2005-02-01 | Method and apparatus for measuring body fat in animals | 2002-10-08 | A method for measuring a body composition measure of an animal includes steps of measuring a first electrical impedance between a front foot pad of the animal and a rear foot pad of the animal; measuring a second electrical impedance between two front foot pads of the animal in parallel with one another and two rear foot pads of the animal in parallel with one another; and utilizing the first electrical impedance, the second electrical impedance, and a regression relationship to determine the body composition measure of the animal. | 1. An apparatus for measuring at least one body composition measure of an animal, said apparatus comprising: a mat having four pairs of electrodes on a surface thereof, each said electrode pair configured to contact a different paw of an animal standing on the mat; electronic circuitry electrically coupled to the four pairs of electrodes and configured to utilize the four pairs of electrodes to measure a first impedance of the animal standing on the mat through one front paw of the animal and one rear paw of the animal and a second impedance of the animal standing on the mat though both front paws of the animal in parallel and both rear paws of the animal in parallel; and the electronic circuitry including a processor configured to determine a body composition measure of the animal utilizing the first measured impedance and the second measured impedance. | 1. An apparatus for measuring at least one body composition measure of an animal, said apparatus comprising: a mat having four pairs of electrodes on a surface thereof, each said electrode pair configured to contact a different paw of an animal standing on the mat; electronic circuitry electrically coupled to the four pairs of electrodes and configured to utilize the four pairs of electrodes to measure a first impedance of the animal standing on the mat through one front paw of the animal and one rear paw of the animal and a second impedance of the animal standing on the mat though both front paws of the animal in parallel and both rear paws of the animal in parallel; and the electronic circuitry including a processor configured to determine a body composition measure of the animal utilizing the first measured impedance and the second measured impedance. 2. An apparatus in accordance with claim 1 wherein the animal for which the pairs of electrodes, electronic circuitry, and processor are configured is a dog. 3. An apparatus in accordance with claim 1 wherein the processor is configured to determine at least one member of the group consisting of body fat and body fat percentage of the animal utilizing an empirically determined regression relationship. 4. An apparatus in accordance with claim 3 wherein the processor is configured to determine body fat percentage. 5. An apparatus in accordance with claim 1 wherein the processor is configured to determine at least one body composition measure selected from the group consisting of total body water, fat free mass, lean body mass, and body fat of the animal. 6. An apparatus in accordance with claim 5 wherein the processor is configured to determine at least one body composition measure selected from the group consisting of total body water, fat free mass, lean body mass, and body fat of the animal, as a percentage of the weight of the animal. 7. An apparatus in accordance with claim 1 wherein the electronic circuitry comprises a current source coupled to a first electrode of each pair of electrodes, a voltmeter coupled to a second electrode of each pair of electrodes, and a controller operatively coupled to the current source to energize electrodes of selected pairs of electrodes to pass a current through one front paw of the animal and one rear paw of the animal for the first impedance measurement and though both front paws of the animal in parallel and both rear paws of the animal in parallel for taking the second impedance measurement; the voltmeter also coupled to the current source to determine the first impedance and the second impedance. 8. An apparatus in accordance with claim 7 wherein the voltmeter is a phase sensitive voltmeter and the current source is an alternating current source. 9. An apparatus in accordance with claim 1 further comprising a display coupled to the processor and a keypad coupled to the processor, and wherein the processor is configured to input zoometric data fr... | Nestec, Ltd. | A61B005/05
US06944491 | B2 | 2005-09-13 | Method of acquiring an image of a non-dermatoglyphic zone of the skin or of a zone of the hair by means of acquisition apparatus including a non-optical sensor | 2001-07-20 | The invention relates to a method of acquiring an image of a non-dermatoglyphic zone of the skin or the hair by means of acquisition apparatus comprising a non-optical sensor. | 1. A method of acquiring an image comprising: providing at least one non-optical sensor for obtaining information concerning microrelief of a zone; and utilizing said at least one non-optical sensor to acquire an image of one of a non- dermatoglyphic zone of skin and a zone of hair. | 1. A method of acquiring an image comprising: providing at least one non-optical sensor for obtaining information concerning microrelief of a zone; and utilizing said at least one non-optical sensor to acquire an image of one of a non- dermatoglyphic zone of skin and a zone of hair. 2. A method according to claim 1, wherein the non-optical sensor includes a sensor having an active surface that is sensitive to variations in temperature. 3. A method according to claim 1, wherein the non-optical sensor includes a sensor having an active surface sensitive to an electrical property measured by measuring one of capacitance and conductance. 4. A method according to claim 1, wherein the non-optical sensor includes a sensor having an active surface that is sensitive to variations in pressure. 5. A method according to claim 2, wherein the active surface is defined by a plurality of individual detection cells disposed in at least one row. 6. A method according to claim 5, wherein the acquisition apparatus is arranged to deliver the image in digital form. 7. A method according to claim 1, wherein the acquisition apparatus is arranged to acquire an image of a zone that is large enough to be statistically representative, including an area lying in a range of from about 0.2 cm2 to about 2 cm2. 8. A method according to claim 1, wherein the image is acquired statically, without moving the non-optical sensor relative to the zone under study during image acquisition. 9. A method according to claim 1, wherein the image is acquired dynamically, with relative movement between the non-optical sensor and the zone under study during image acquisition. 10. A method according to claim 9, wherein the non-optical sensor includes an active surface in the form of a strip of individual detection cells. 11. A method according to claim 1, wherein the image is acquired without the non-optical sensor coming into contact with the zone under study. 12. A method according to claim 1, wherein the image is acquired with the non-optical sensor in contact with the zone under study. 13. A method according to claim 12, including measuring pressure of contact between the non-optical sensor and the zone under study during image acquisition. 14. A method according to claim 12, wherein the image is acquired at a substantially constant contact pressure. 15. A method according to claim 1, wherein the acquired image is a three dimensional image of the zone under study. 16. A method according to claim 1, wherein the acquired image is a two dimensional image of the zone under study. 17. A method according to claim 1, wherein the non-optical sensor has a spatial resolution lying in a range of from 10 μm to 100 μm. 18. A method according to claim 1, further including processing the image in order to determine characteristic parameters of the zone under study. 19. A method according to claim 18, wherein the processing provides information concerning a surface density of skin lines. ... | L'Oreal | A61B005/05
US06976965 | B2 | 2005-12-20 | Ultra miniature pressure sensor | 2002-09-19 | A guide wire having pressure sensing capabilities for measuring the pressure of liquid in a vessel comprising a flexible elongate member and having proximal and distal extremities and having an outside diameter of 0.018″ or less. The distal extremity of said flexible elongate member is adapted to be disposed in the liquid in said vessel. A housing is carried by the flexible elongate member and has a diameter substantially the same as the diameter of the flexible elongate member. The housing has a space therein with a pressure sensor mounted in the space in the housing. The pressure sensor has a diaphragm that is sensitive to changes of pressure in the liquid in the vessel. The diaphragm is rectangular in shape and is bordered by a rim surrounding the well. A backing plate is formed of an insulating material bonded to the crystal and serves to reinforce the rim of the crystal of semiconductor material. The backing plate has a cavity therein underlying the diaphragm and is in substantial registration with the diaphragm. The cavity serves to provide a pressure reference. | 1. A pressure sensor apparatus comprising: a guidewire;a sensor housing having an external wall, a lumen, and a hole in the external wall extending from an external surface of the sensor housing to the lumen and being open to ambient fluid, the sensor housing disposed near a distal extremity of the guidewire;a first coil disposed proximally from the sensor housing;a second coil comprising a radiopaque material and disposed distally from the sensor housing; anda solid state pressure sensor having a pressure sensitive region located at or near one end thereof, the solid state pressure sensor mounted within the sensor housing such that the pressure sensitive region projects into the lumen of the sensor housing without contacting the external wall of the sensor housing and such that a portion of the pressure sensitive region is disposed opposite the hole and is in fluid communication with the ambient fluid via the hole. | 1. A pressure sensor apparatus comprising: a guidewire;a sensor housing having an external wall, a lumen, and a hole in the external wall extending from an external surface of the sensor housing to the lumen and being open to ambient fluid, the sensor housing disposed near a distal extremity of the guidewire;a first coil disposed proximally from the sensor housing;a second coil comprising a radiopaque material and disposed distally from the sensor housing; anda solid state pressure sensor having a pressure sensitive region located at or near one end thereof, the solid state pressure sensor mounted within the sensor housing such that the pressure sensitive region projects into the lumen of the sensor housing without contacting the external wall of the sensor housing and such that a portion of the pressure sensitive region is disposed opposite the hole and is in fluid communication with the ambient fluid via the hole. 2. The pressure sensor apparatus of claim 1, wherein the solid state pressure sensor is mounted within the sensor housing using an elastic medium. 3. The pressure sensor apparatus of claim 2, wherein the elastic medium is an epoxy. 4. The pressure sensor apparatus of claim 1, wherein the pressure sensitive region includes a pressure sensitive diaphragm. 5. The pressure sensor apparatus of claim 1, wherein the guidewire has an outside diameter of 0.018″ or less. 6. The pressure sensor apparatus of claim 1, wherein the solid state pressure sensor comprises a diaphragm structure supported by a base plate. 7. The pressure sensor apparatus of claim 1, further comprising a core wire extending from a proximal extremity of the guidewire, through the sensor housing, and to the distal extremity of the guidewire. 8. The pressure sensor apparatus of claim 7, wherein the core wire comprises a tapered distal extremity. 9. The pressure sensor apparatus of claim 1, wherein the lumen is filled with a viscous fluid. 10. The pressure sensor apparatus of claim 9, wherein the viscous fluid comprises an oil. 11. The pressure sensor apparatus of claim 9, wherein the hole has a diameter configured to hold the viscous fluid within the lumen and prevent bleeding of the viscous fluid into the ambient fluid. 12. The pressure sensor apparatus of claim 1, wherein the hole has a diameter of approximately 5 mils or less. 13. The pressure sensor apparatus of claim 1, comprising a temperature sensitive resistor disposed on a distal extremity of the guidewire. | Volcano Corporation | A61B005/00
US06947783 | B2 | 2005-09-20 | System for the navigation-assisted positioning of elements | 2001-03-15 | A system for applying an element to a body, wherein the position of a device for preparing and/or producing a connection between said body and said element relative to the position of said body is detected, it also relating to a system for applying an element to a body comprising a device for preparing and/or producing a connection between said element and said body and a position detection device for detecting the position of said device, for preparing and/or producing a connection between said element and said body, and for detecting said position of said body. | 1. A system for applying an element to a living body, comprising: a) a device for physically altering a part of the living body to prepare and/or produce a connection between said element and said living body; and b) a position detection device (i) for detecting a position of said device for physically altering a part of the living body to prepare and/or produce a connection between said element and said living body, and (ii) for detecting a position of said living body, wherein plural markers are applied to said device for physically altering for three-dimensional tracking of said device for physically altering. | 1. A system for applying an element to a living body, comprising: a) a device for physically altering a part of the living body to prepare and/or produce a connection between said element and said living body; and b) a position detection device (i) for detecting a position of said device for physically altering a part of the living body to prepare and/or produce a connection between said element and said living body, and (ii) for detecting a position of said living body, wherein plural markers are applied to said device for physically altering for three-dimensional tracking of said device for physically altering. 2. The system as set forth in claim 1, wherein said device for preparing and/or producing a connection between said body and said element is a drill, a saw, a surface-working device, a screwing device or a nailing device. 3. The system as set forth in claim 1, wherein said position detection device detects optical, acoustical and/or radio signals. 4. The system as set forth in claim 1, wherein one or more markers are applied to said body. 5. The system as set forth in claim 1, wherein a processor is provided for preparing or implementing one or more of the steps in a method comprising: detecting the position of a device for preparing and/or producing a connection between said body and said element relative to the position of said body. 6. The system as set forth in claim 1, wherein a display device is provided, connected to said processor. 7. The system as set forth in claim 1, wherein an input device is provided, connected to said processor. 8. The system as set forth in claim 1, wherein a data recording device is provided for recording data during the operation of said system. 9. The system as set forth in claim 1, comprising means for applying the element to the prepared or produced connection, wherein the means for applying the element to the prepared or produced connection aligns the element to the body. | BrainLAB AG | A61B005/05 A61B019/00
US06912415 | B2 | 2005-06-28 | Method for acquiring MRI data from a large field of view using continuous table motion | 2001-11-26 | MRA data is acquired from a large region of interest by translating the patient through the bore of the MRI system as a three-dimensional MRA data set are acquired. Patient table movement is controlled to track a bolus of contrast agent as it passes through the region of interest. Fluoroscopic images may be acquired during the scan to enable accurate bolus tracking. A seamless image of the entire region of interest is reconstructed. | 1. In a magnetic resonance imaging (MRI) system having a defined field of view (FOV), a method for producing an image of a subject over an extended field of view (FOVtot) which is larger than the FOV, the steps comprising: a) moving the subject through the MRI system such that the extended field of view (FOVtot) passes through the defined field of view (FOV); b) continuously acquire NMR data from the subject as it is moved through the FOV by repeatedly performing an imaging pulse sequence which acquires NMR data comprising a view of the subject using a readout gradient directed along the direction of subject movement; c) adjusting each view acquired in step b) using subject position information; d) storing each adjusted view in a data matrix; and e) reconstructing an image using the data matrix. | 1. In a magnetic resonance imaging (MRI) system having a defined field of view (FOV), a method for producing an image of a subject over an extended field of view (FOVtot) which is larger than the FOV, the steps comprising: a) moving the subject through the MRI system such that the extended field of view (FOVtot) passes through the defined field of view (FOV); b) continuously acquire NMR data from the subject as it is moved through the FOV by repeatedly performing an imaging pulse sequence which acquires NMR data comprising a view of the subject using a readout gradient directed along the direction of subject movement; c) adjusting each view acquired in step b) using subject position information; d) storing each adjusted view in a data matrix; and e) reconstructing an image using the data matrix. 2. The method as recited in claim 1 in which the MRI system has a table, and step a) is performed by: i) placing the subject on the table; and ii) moving the table. 3. The method as recited in claim 2 in which the table is moved continuously while performing step a). 4. The method as recited in claim 2 in which the table is moved at different velocities while performing step a). 5. The method as recited in claim 2 which includes: injecting the subject with a contrast agent; and in which the table is moved at a velocity which tracks the contrast agent as it moves through the extended field of view (FOVtot). 6. The method as recited in claim 5 which includes: reconstructing monitoring images during the performance of step a) from data stored in the data matrix. 7. The method as recited in claim 2 in which step c) includes adjusting the location in the data matrix in which the view is stored in step d) along the direction of subject movement as a function of the table location at the time the view is acquired in step b). 8. The method as recited in claim 2 in which step c) includes adjusting the phase of the NMR data in the view as a function of the table location at the time the view is acquired in step b). 9. The method as recited in claim 2 in which step b) further includes: i) acquiring table location information as each view is acquired; and the table location information is used in step c) to adjust each corresponding view. 10. The method as recited in claim 9 in which step c) includes: i) performing a Fourier transformation of the NMR data in the view; and ii) calculating a location in the data matrix for the transformed view as a function of the table location at the time the view was acquired in step b). 11. The method as recited in claim 1 in which step c) includes adjusting the location in the data matrix in which the view is stored in step d) along the direction of subject movement. 12. The method as recited in claim 1 in which step c) includes adjusting the phase of the NMR data in the view. 13. The method as recited in claim 1 in which step c) includes: i) adjusting the phase of the NMR data in the view; ii) Fourier transformi... | Mayo Foundation for Medical Education and Research | A61B005/05
US06929610 | B2 | 2005-08-16 | Non-invasive measurement of blood pressure | 2002-03-25 | By a method to measure the blood pressure of a patient, a series of measurements for the diastolic, the systolic blood pressure and for the pulse rate are made. The results are analysed in order to determine, and to display whether the patient is properly settled to a rest condition, and whether the patients blood pressure variation is found as being normal. Further it is investigated, whether there are high measurement result alterations due to artefacts. Highly altered results are excluded. If the results are sufficiently similar, an average of the result is made. If the results are not sufficiently similar, but still below pre-defined alteration limits, a weighted average is applied to the systolic and the diastolic blood pressure. | 1. A method for non-invasively measuring the blood pressure of a patient, comprising the steps of a) measuring the diastolic and the systolic blood pressure value and the pulse rate of the patient; b) repeating step a) at least once thereby generating a series of consecutive measurement results of the systolic and diastolic blood pressure and of the pulse rate; c) automatically analyzing said series of consecutive results for determining the validity of said results by verifying whether a first condition is fulfilled, whereby said first condition is fulfilled if results of the systolic blood pressure are continuously decreasing and if the results of the diastolic blood pressure are continuously decreasing or increasing and if the results of the pulse rate are continuously decreasing. | 1. A method for non-invasively measuring the blood pressure of a patient, comprising the steps of a) measuring the diastolic and the systolic blood pressure value and the pulse rate of the patient; b) repeating step a) at least once thereby generating a series of consecutive measurement results of the systolic and diastolic blood pressure and of the pulse rate; c) automatically analyzing said series of consecutive results for determining the validity of said results by verifying whether a first condition is fulfilled, whereby said first condition is fulfilled if results of the systolic blood pressure are continuously decreasing and if the results of the diastolic blood pressure are continuously decreasing or increasing and if the results of the pulse rate are continuously decreasing. 2. A method according to claim 1, wherein in step c), a variation index of the diastolic blood pressure, the systolic blood pressure and the pulse rate is calculated according to the following formula: VI=1n=∑i=1n⁢Ri-meanmean=∑i=1n⁢Ri-mean∑i=1⁢Riwhere⁢ ⁢mean=∑i=1n⁢Rin⁢ Ri are the respective measurement results and n is the number of measurements for each series of results of the diastolic blood pressure, the systolic blood pressure and the pulse rate and whereby it is verified whether a second condition is fulfilled, whereby the second condition is fulfilled if the variation index is above a predetermined variation index limit for at least one of the parameters diastolic blood pressure, systolic blood pressure and pulse rate. 3. A method according to claim 2, comprising the further steps of verifying whether a third condition is fulfilled, whereby the third condition is fulfilled when the differences between subsequent measurement results of the systolic and of the diastolic blood pressure measurement are below a predetermined or predeterminable similarity limit and verifying whether a fourth condition is fulfilled, whereby the fourth condition is fulfilled when the differences between subsequent measurement results of the systolic and the diastolic blood pressure are above a predetermined or predeterminable alteration limit. 4. A method according to claim 1, wherein if the first condition is met, a signal is generated advising the user to repeat the measurement or an additional measurement is automatically initiated after a predetermined period of time, and/or a signal is generated, said signal indicating that the user has not been properly settled to a rest condition in the course of the measurement. 5. A method according to claim 2, wherein if the second condition is met, a message is displayed to the user. 6. A method according to claim 3, wherein if the third condition is met, an arithmetic average of the measurement results of the diastolic blood pressure and of the systolic blood pressure and of the pulse frequency is calculated and displayed and wherein if the fourth condition is met, results above the alteration limit are disregarded and a new m... | Microlife Intellectual Property GmbH | A61B005/02
US06944485 | B1 | 2005-09-13 | Hyperbaric resuscitation system and method | 2000-04-20 | A hyperbaric resuscitation system (10) includes a hyperbaric chamber (20) having a volume sufficient to enclose a human patient (1) and at least two operating personnel (60). The system (10) also includes a device for pressurizing the hyperbaric chamber (20) to at least 1.5 atmospheres with air. The concentration of oxygen in high pressure, oxygen-rich gas to be breathed by the patient (1) provided by an independent system (41) at chamber pressure is automatically regulated by a regulating system (33) which receives information about the amount of oxygen in cerebral tissue of the patient (1) from a spectrophotometer (51, 52). Although devices for measuring the exact amount of oxygen in cerebral tissue do not yet exist, the presently available devices can show trends in the amount of oxygen in the tissue. Since the physician working on a patient in a hyperbaric resuscitation system is more concerned about trending than exact values, the present system can still be of great benefit in resuscitating patients. | 1. A system comprising: (a) a light source;(b) a pickup optode unit for detecting light from the light source;(c) a spectrophotometer coupled to the pickup optode unit for sensing and recording a NIR wavelength interval including cytochrome oxidase, water and hemoglobin data;(d) a personal computer with a software algorithm to extract via model-free mathematical operations the cytochrome oxidase, water and hemoglobin data from the NIR wavelength interval for evaluation and display. | 1. A system comprising: (a) a light source;(b) a pickup optode unit for detecting light from the light source;(c) a spectrophotometer coupled to the pickup optode unit for sensing and recording a NIR wavelength interval including cytochrome oxidase, water and hemoglobin data;(d) a personal computer with a software algorithm to extract via model-free mathematical operations the cytochrome oxidase, water and hemoglobin data from the NIR wavelength interval for evaluation and display. 2. The system of claim 1, wherein the light source is a stabilized pulsed light. 3. A method of using the system of claim 1 to monitor the change of any natural or manmade chromophore existing in a person's brain to assist in the diagnosis or treatment of a neurological or psychotic disorder, comprising: using the light source to illuminate a person's cerebral tissue;using the pickup optode unit to detect light from the person's cerebral tissue;using the spectrophotometer to sense and record a NIR wavelength interval including cytochrome oxidase, water and hemoglobin data;using the personal computer and the software algorithm to extract the cytochrome oxidase, water and hemoglobin data from the NIR wavelength interval for evaluation. 4. The method of claim 3, wherein the light source is a stabilized pulsed light. 5. The method of claim 3, wherein the spectrophotometer monitors relative changes in redox levels in real-time. 6. The method of claim 3, wherein Fourier transforms are used in analyses of near infrared data obtained from the spectrophotometer. 7. The method of claim 3, wherein: the spectrophotometer includes: a background pickup device which receives photons that have traversed the patient's scalp and skull but not deep enough to reach the patient's cerebral cortex,a sample pickup device that is positioned to receive photons that have traversed the patient's scalp, skull dura matter, and pia, andthe background signal is subtracted from the sample signal by the algorithm to result in a signal representing the patient's cerebral cortex. 8. The method of claim 3, wherein the light source is a quartz halogen 150 watt light source. 9. The method of claim 3, wherein the NIR wavelength interval is about 700–1050 nm. 10. The method of claim 3, wherein oxygen in cerebral tissue is monitored by monitoring cytochrome oxidase in the cerebral tissue. 11. The method of claim 3, wherein oxygen in cerebral tissue is monitored by monitoring the redox ratio of cytochrome oxidase in the patient's cerebral tissue. 12. The invention of claim 1, wherein the spectrophotometer monitors relative changes in redox levels in real-time. 13. The invention of claim 1, wherein the software algorithm uses Fourier transforms in analyses of near infrared data obtained from the spectrophotometer. 14. The invention of claim 1, wherein: the spectrophotometer includes: a background pickup device which receives photons that have traversed a patient's scalp and skull but not deep enough... | - | A61B005/00
US06902535 | B2 | 2005-06-07 | Guide-wire mounted balloon modulation device and methods of use | 2002-08-26 | An apparatus for modulating the pressure of a fluid such as a gas within the expandable portion of a guide wire catheter. A preferred embodiment apparatus features a mechanism for controllably gripping and releasing the open, proximal end of a tubular guide wire, a mechanism for introducing a fluid to a desired pressure and volume into the expandable portion of the tubular guide wire through the open end, and, while maintaining the pressure and volume of fluid in the tubular guide wire, a mechanism for introducing a sealing member into the open end of the tubular guide wire to seal the fluid in the tubular guide wire. In a particularly preferred embodiment, the apparatus also features a deflation tool for piercing the seal and letting the fluid out. Using this apparatus, the tubular guide wire can be re-sealed and re-opened as necessary. | 1. An apparatus for modulating the pressure of a fluid within an expandable portion of a tubular guide wire or catheter, the apparatus comprising an inflation tool comprising: means for controllably gripping an open, proximal end of a tubular guide wire or catheter, the gripping means comprising a deformable member and a deforming mechanism, the deforming mechanism comprising a lever mechanism arranged to controllably act upon the deformable member to cause gripping of the tubular guide wire or catheter by the deformable member; means for introducing a fluid into said expandable portion of said tubular guide wire or catheter through said open, proximal end, thereby inflating said expandable portion of said tubular guide wire or catheter to a desired extent; and means for introducing a sealing member into said open proximal end of said tubular guide wire or catheter to seal said fluid in said tubular guide wire or catheter. | 1. An apparatus for modulating the pressure of a fluid within an expandable portion of a tubular guide wire or catheter, the apparatus comprising an inflation tool comprising: means for controllably gripping an open, proximal end of a tubular guide wire or catheter, the gripping means comprising a deformable member and a deforming mechanism, the deforming mechanism comprising a lever mechanism arranged to controllably act upon the deformable member to cause gripping of the tubular guide wire or catheter by the deformable member; means for introducing a fluid into said expandable portion of said tubular guide wire or catheter through said open, proximal end, thereby inflating said expandable portion of said tubular guide wire or catheter to a desired extent; and means for introducing a sealing member into said open proximal end of said tubular guide wire or catheter to seal said fluid in said tubular guide wire or catheter. 2. The apparatus of claim 1, further comprising a means for re-opening said proximal end of said tubular guide wire or catheter. 3. The apparatus of claim 2, wherein said means for re-opening said proximal end of said tubular guide wire or catheter comprises a deflation tool comprising a needle for piercing said sealing member. 4. The apparatus of claim 3, wherein said deflation tool further comprises a handle having a bore that is open on at least one end, and wherein said needle is mounted in said bore coaxially with a centerline of said bore. 5. The apparatus of claim 4, further comprising a lead-in taper to said bore, and a spring coiled around said needle to help support said needle. 6. The apparatus of claim 4, wherein said needle arranged in said bore creates a channel between said needle and an inner wall of said handle, said channel having a spring arranged therein such that said spring maintains concentricity of said needle in said bore. 7. The apparatus of claim 1, wherein said sealing member comprises wax. 8. The apparatus of claim 1, wherein said means for introducing said sealing member comprises a rod mounted in said apparatus, said rod being capable of moving axially and radially. 9. The apparatus of claim 8, wherein said sealing member is housed in a bore in a flat surface of said rod that is to contact said open end of said tubular guide wire or catheter, said bore being offset from a longitudinal axis of said rod. 10. The apparatus of claim 9, wherein said open, proximal end of said tubular guide wire or catheter is urged against said flat surface. 11. The apparatus of claim 10, wherein said introducing means for said sealing member is arranged to rotate said rod until said bore containing said sealing member comes into alignment with said proximal end of said tubular guide wire or catheter, and press said rod axially toward said tubular guide wire or catheter, such that said proximal end of said tubular guide wire or catheter is sealed. 12. The apparatus of claim 9, wherein said flat surface is s... | Kansey Nash Corporation | A61B005/00 A61M025/00
US06885883 | B2 | 2005-04-26 | Methods for improving performance and reliability of biosensors | 2001-05-14 | The present invention relates to a predictive-kinetic method for use with data processing of a sensor-generated signal, as well as, microprocessors and monitoring systems employing such a predictive-kinetic method. Data from a transient region of a signal is used with suitable models and curve-fitting methods to predict the signal that would be measured for the system at the completion of the reaction. The values resulting from data processing of sensor response using the methods of the present invention are less sensitive to measurement variables. | 1. One or more microprocessors, comprising programming to control (i) operating a sensing device for obtaining a measured signal over time, comprising a measured signal response curve, from a sample comprising glucose collected from a biological system, wherein said measured signal is specifically related to the amount or concentration of glucose present in the biological system, and said measured signal response curve comprises kinetic and equilibrium regions; and (ii) a computation method using (a) a mathematical model comprising selected parameters, wherein said model describes the measured signal response curve, and said mathematical model is selected from the group consisting of a first order process, combined first order and zero order process, a parallel multiple first order process, a flux process, and an nth order process, and (b) an error minimization method to iteratively estimate values of the parameters using said model to fit a predicted response curve to said measured signal response curve, wherein (a′) the error minimization method provides a calculated error based on differences between said predicted and measured signal response curves, and (b′) said estimating is iteratively performed until the calculated error between the predicted and measured signal response curves falls within an acceptable range or until no further statistically significant change is seen in the calculated error, at which time iterative estimation of the parameters is stopped, said iterative estimation and error minimization results in a predicted response curve corresponding to said measured signal response curve, said predicted response curve yields a predicted end-point value and a measurement correlated to the amount or concentration of the glucose. | 1. One or more microprocessors, comprising programming to control (i) operating a sensing device for obtaining a measured signal over time, comprising a measured signal response curve, from a sample comprising glucose collected from a biological system, wherein said measured signal is specifically related to the amount or concentration of glucose present in the biological system, and said measured signal response curve comprises kinetic and equilibrium regions; and (ii) a computation method using (a) a mathematical model comprising selected parameters, wherein said model describes the measured signal response curve, and said mathematical model is selected from the group consisting of a first order process, combined first order and zero order process, a parallel multiple first order process, a flux process, and an nth order process, and (b) an error minimization method to iteratively estimate values of the parameters using said model to fit a predicted response curve to said measured signal response curve, wherein (a′) the error minimization method provides a calculated error based on differences between said predicted and measured signal response curves, and (b′) said estimating is iteratively performed until the calculated error between the predicted and measured signal response curves falls within an acceptable range or until no further statistically significant change is seen in the calculated error, at which time iterative estimation of the parameters is stopped, said iterative estimation and error minimization results in a predicted response curve corresponding to said measured signal response curve, said predicted response curve yields a predicted end-point value and a measurement correlated to the amount or concentration of the glucose. 2. The one or more microprocessors of claim 1, further programmed (i) to operate said sensing device to obtain a series of measured signals over time resulting in a series of measured signal response curves, and (ii) to provide a predicted response curve corresponding to each measured signal response curve. 3. The one or more microprocessors of claim 2, wherein after estimation of each predicted response curve for each measured signal response curve in the series of measurements an amount or concentration of the glucose is determined based on the predicted response curve. 4. The one or more microprocessors of claim 2, wherein a first mathematical model is selected to estimate at least one predicted end-point value corresponding to a measured signal response curve and at least one different mathematical model is used to estimate at least one predicted end-point value corresponding to a different measured signal response curve. 5. The one or more microprocessors of claim 2, wherein a single mathematical model is selected to estimate the predicted end-point values for all measurements in the series. 6. The one or more microprocessors of claim 1, wherein said measured signal response curve comprises a mea... | Cygnus, Inc. | A61B005/05
US06958043 | B2 | 2005-10-25 | Apparatus and method for displacing the partition between the middle ear and the inner ear using a manually powered device | 2002-05-21 | Apparatus and method for treating a disease or condition of the ear, such as Meniere's disease or endolymphatic hydrops, includes a manual device operable by hand to pressurize air and pressure transmitting arrangement for delivering the pressurized air to the middle ear, preferably in pulses, to displace the partition between the middle ear and the inner ear to influence fluid in the inner ear. | 1. Apparatus for treating a human ear having an outer ear canal separated from a middle ear by a tympanic membrane with a vent therethrough and an inner ear separated from the middle ear by a partition to treat the symptoms of Meniere's disease comprising a manual device operable by hand to pressurize air and including a reservoir for storing air pressurized by said manual device, said reservoir being capable of storing a quantity of air sufficient to produce a series of positive air pressure pulses; and pressure transmitting means communicating with said reservoir and configured to communicate with the outer ear canal to deliver air from said reservoir as a series of positive air pressure pulses to the middle ear via the vent to displace the partition to influence fluid distribution in the inner ear, said pressure transmitting means being capable of delivering the series of positive air pressure pulses from said quantity of air stored in said reservoir without additional air being pressurized by said manual device. | 1. Apparatus for treating a human ear having an outer ear canal separated from a middle ear by a tympanic membrane with a vent therethrough and an inner ear separated from the middle ear by a partition to treat the symptoms of Meniere's disease comprising a manual device operable by hand to pressurize air and including a reservoir for storing air pressurized by said manual device, said reservoir being capable of storing a quantity of air sufficient to produce a series of positive air pressure pulses; and pressure transmitting means communicating with said reservoir and configured to communicate with the outer ear canal to deliver air from said reservoir as a series of positive air pressure pulses to the middle ear via the vent to displace the partition to influence fluid distribution in the inner ear, said pressure transmitting means being capable of delivering the series of positive air pressure pulses from said quantity of air stored in said reservoir without additional air being pressurized by said manual device. 2. The apparatus recited in claim 1 wherein said manual device includes a chamber having resilient walls and squeezable by hand to pressurize air therein and said reservoir receives air from said chamber. 3. The apparatus recited in claim 1 wherein said pressure transmitting means includes a plug for engaging the outer ear canal and an aperture in the plug through which the series of positive air pressure pulses is delivered. 4. Apparatus for treating a human ear having an outer ear canal separated from a middle ear by a tympanic membrane with a vent therethrough and an inner ear separated from the middle ear by a partition to treat the symptoms of Meniere's disease comprising a manual device, a chamber having resilient walls and squeezable by hand to pressurize air therein, and; pressure transmitting means communicating with said manual device and configured to communicate with the outer ear canal to deliver air pressure to the middle ear via the vent to displace the partition to influence fluid distribution in the inner ear, said pressure transmitting means including a reservoir communicating with said chamber for holding air pressure therein greater than ambient air pressure for delivery to the middle ear via the outer ear canal and the vent. 5. The apparatus recited in claim 4 wherein said pressure transmitting means includes means for delivering air pressure pulses to the middle ear via the outer ear canal and the vent. 6. The apparatus recited in claim 4 wherein said pressure transmitting means includes a passage communicating with said reservoir for delivering the air pressure from said reservoir as a static air pressure greater than ambient air pressure to the middle ear via the outer ear canal and the vent. 7. Apparatus for treating a human ear having an outer ear canal separated from a middle ear by a tympanic membrane with a vent therethrough and an inner ear separated from the middle ear by a partition to treat the... | Medtronic Xomed, Inc. | A61B005/00 A61M031/00
US06840910 | B2 | 2005-01-11 | Method of distributing skin care products | 2001-08-01 | The present invention relates to a method of distributing a skin care product to a subject, the method including the steps of: (a) obtaining a sample of interstitial fluid from the skin of the subject; (b) measuring the amount of a skin analyte in the sample; and (c) distributing a skin care product to the subject to alter the amount of the skin analyte in the skin of the subject. | 1. A method of distributing a skin care product to a subject, said method comprising the steps of: (a) obtaining a sample of interstitial fluid from the skin of said subject; and (b) measuring the amount of a skin analyte in said sample; and (c) distributing a skin care product to said subject to alter the amount of said skin analyte in the skin of said subject. | 1. A method of distributing a skin care product to a subject, said method comprising the steps of: (a) obtaining a sample of interstitial fluid from the skin of said subject; and (b) measuring the amount of a skin analyte in said sample; and (c) distributing a skin care product to said subject to alter the amount of said skin analyte in the skin of said subject. 2. A method of claim 1, wherein said method comprises creating a opening in the stratum corneum of the skin of said patient and accessing said sample through said opening. 3. A method of claim 2, wherein said opening is created with a needle, a blade, a laser, or an electric arc. 4. A method of claim 3, wherein said opening is created with a laser. 5. A method of claim 4, wherein said method further comprises comparing the amount of said analyte is compared to a reference standard. 6. A method of claim 5, wherein said skin care product comprises a vitamin. 7. A method of claim 4, wherein said skin care product comprises a vitamin. 8. A method of claim 3, wherein said sample is extracted through said opening under negative pressure. 9. A method of claim 8, wherein said skin analyte is a vitamin, an immunoglobulin, an insulin-like growth factor, an interleulin, an interleukin receptor antagonist, or a metalloproteinase. 10. A method of claim 8, wherein said skin analyte is vitamin C. 11. A method of claim 10, wherein said method further comprises comparing the amount of said analyte is compared to a reference standard. 12. A method of claim 11, wherein said skin care product comprises a vitamin. 13. A method of claim 10, wherein said skin care product comprises a vitamin. 14. A method of claim 8, wherein said method further comprises comparing the amount of said analyte is compared to a reference standard. 15. A method of claim 14, wherein said skin care product comprises a vitamin. 16. A method of claim 8, wherein said skin care product comprises a vitamin. 17. A method of claim 2, wherein said sample is extracted through said opening under negative pressure. 18. A method of claim 2, wherein said method further comprises comparing the amount of said analyte is compared to a reference standard. 19. A method of claim 18, wherein said skin care product comprises a vitamin. 20. A method of claim 2, wherein said skin care product comprises a vitamin. 21. A method of claim 1, wherein said skin analyte is a vitamin, an immunoglobulin, an insulin-like growth factor, an interleulin, an interleukin receptor antagonist, or a metalloproteinase. 22. A method of claim 1, wherein said skin analyte is vitamin C. 23. A method of claim 1, wherein said method further comprises comparing the amount of said analyte is compared to a reference standard. 24. A method of claim 23, wherein said skin care product comprises a vitamin. 25. A method of claim 1, wherein said skin care product comprises a vitamin. | Johnson & Johnson Consumer Companies, Inc. | A61B005/00 A61M035/00
US06945941 | B2 | 2005-09-20 | Volumetric physiological measuring system and method | 2002-02-28 | A system and method for measuring volumes and areas using electromagnetic induction techniques. A current is generated and fed into one of two coil assemblies to induce voltage into another coil assembly to provide accurate values for volume or area. | 1. A device for measuring the changing area within at least one portion of a mammal due to cardiac function, comprising: a. conductive coil means configured tightly about the various circumferences of at least one portion of the mammal; b. fixed coil means remotely located relative to the conductive coil means about the mammal; and, c. current generating means for selectively providing constant alternating current to either one of the conductive coil means or the fixed coil means to create an induced voltage in the other coil means representative of true area within the coil means that is configured tightly about the various circumferences of the mammal portion or portions, within the signals and area changing over time due to the cardiac function of the mammal, said constant current substantially increasing the accuracy of said measurement system. | 1. A device for measuring the changing area within at least one portion of a mammal due to cardiac function, comprising: a. conductive coil means configured tightly about the various circumferences of at least one portion of the mammal; b. fixed coil means remotely located relative to the conductive coil means about the mammal; and, c. current generating means for selectively providing constant alternating current to either one of the conductive coil means or the fixed coil means to create an induced voltage in the other coil means representative of true area within the coil means that is configured tightly about the various circumferences of the mammal portion or portions, within the signals and area changing over time due to the cardiac function of the mammal, said constant current substantially increasing the accuracy of said measurement system. 2. The device of claim 1 in which the current generating means generates current with a frequency range of about 10 kHz to about 200 kHz and from about 1 milliampere to about 1 ampere. 3. The device of claim 1 in which the conductive coil means comprises a plurality of coil loops which are configured tightly about the thoracic region and the abdominal region of the mammal. 4. The device of claim 1 further comprising sensing and control means for controlling operation of the device, the sensing and control means being electrically connected to the conductive coil means, the fixed coil means, and the current generating means. 5. The device of claim 4 in which the sensing and control means comprises timing multiplex switching means for providing simultaneous volume measurements of both a thoracic region and an abdominal region of the mammal. 6. The device of claim 5 in which the sensing and control means comprises multiplexing means for providing simultaneous measurement of a plurality of regions of the mammal utilizing either phase, frequency or time multiplexing. 7. The device of claim 1, further including a sensing and control means comprising an amplifier and a rectifier electrically connected to the coil means that is receiving the induced voltage from the other coil means. 8. The device of claim 1 in which the fixed coil means comprises a plurality of small coil elements that are configured for matching and positioning to permit the fixed coil means to generate a homogeneous magnetic field similar to a single large coil for either sensing an induced voltage from the conductive coil means or for generating a field to create an induced voltage in the conductive coil means. 9. The device of claim 8 in which the plurality of small coil elements comprises three small coil elements. 10. The device of claim 9 in which each of the small coil elements is wound on a ferrite core and arranged linearly with optimized positions and signal intensity weighting to generate a homogeneous magnetic field at the portion of the mammal being measured. 11. The device of claim 1 in which the conductive coil mean... | VoluSense AS | A61B005/08 A61B005/02 G01F001/56
US07024237 | B1 | 2006-04-04 | Mask system and method for stereotactic radiotherapy and image guided procedures | 1999-10-29 | Repeat fixation for medical procedures is accomplished using a non-invasive locator, specifically a face mask. The face mask has at least three fiducial markers on it. By detecting the position of the markers, the position of features within the patient can be determined with great precision. The features can be the location of head or neck cancer or other target to which radiation or other medical procedure is to be applied. Since the face mask has been molded to uniquely fit to the patient's face, it may be removed after an initial imaging of the patient. The face mask may then be re-attached one or more times to the teeth. The face mask is molded to conform to the patient's face by using thermoplastic material. | 1. A system for medical procedures, the system comprising: a locator attachable to a patient, having at least 3 fiducial markers thereon;a medical device for performing a diagnostic imaging or a therapeutic medical procedure on a patient; anda sensing subsystem for sensing the positions of the fiducial markers when the patient is in a position for performing the medical procedure using the medical device; andwherein the locator has a registration portion for registration with a portion of a patient's body, the locator being mechanically free such that a patient is positionable without applying forces to the locator during patient positioning; and wherein the locator is molded to fit external features of a specific patient. | 1. A system for medical procedures, the system comprising: a locator attachable to a patient, having at least 3 fiducial markers thereon;a medical device for performing a diagnostic imaging or a therapeutic medical procedure on a patient; anda sensing subsystem for sensing the positions of the fiducial markers when the patient is in a position for performing the medical procedure using the medical device; andwherein the locator has a registration portion for registration with a portion of a patient's body, the locator being mechanically free such that a patient is positionable without applying forces to the locator during patient positioning; and wherein the locator is molded to fit external features of a specific patient. 2. The system of claim 1 wherein the locator is a face mask molded to fit in registry with the face of a specific patient. 3. The system of claim 2 wherein the registration portion allows removal of the locator from the patient and re-attachment to the patient with an identical orientation relative to the portion of the patient as when the locator was previously attached. 4. The system of claim 3 wherein the face mask includes thermoplastic molded to fit the face of a patient. 5. The system of claim 1 wherein the locator includes thermoplastic molded to fit a portion a portion of a patient. 6. A medical method comprising the steps, not necessarily in order, of positioning a patient for a first medical procedure; molding a locator to external features of a patient, the locator being placed in registry with a portion of the patient;using at least 3 fiducial markers that are fixed relative to the locator a first time to get precise positioning information relative to at least part of the patient, the locator remaining mechanically free during this step;performing a first medical procedure on the patient, the locator remaining mechanically free during this step;after the first medical procedure, removing the locator from the patient;at a later time after the removing of the locator, reattaching the locator to the patient, the locator again being in registry with the portion of the patient and having an identical relative to the portion of the patient as when the locator was previously attached;after the re-attaching step, using the fiducial markers a second time to get precise positioning information relative to the at least part of the patient, the locator remaining mechanically free during this step; andafter the re-attaching step, performing a second medical procedure on the patient, the locator remaining mechanically free during this step. 7. The medical method of claim 6 wherein; before performing the second medical procedure, the patient is positioned using a positioner independent of the locator to adjust and secure at least the portion of the patient in a desired position. 8. The medical method of claim 6 wherein the attaching and re-attaching of the locator is non-invasive. 9. The medical method of claim 8 wherein ... | University of Florida Research Foundation, Inc. | A61B005/00
US07018338 | B2 | 2006-03-28 | Method and device for pulse rate detection | 2002-09-26 | Portable pulse rate detecting device for contact with human body tissue, including a light-emitting source for emitting radiant energy directed at through human body tissue; at least first and second light detectors for detecting intensity of radiant energy after propagation through human body tissue and for providing first and second input signals as a function of such propagation, a detecting device for providing a motion reference signal, and processing means for removing motion-related contributions from the first and second input signals and subtracting a calculated model based on the motion reference signal from each of the first and second input signals, wherein the processing means is also for removing measurement noise and residual non-modeled contributions from the first and second enhanced signals using a noise reduction algorithm. | 1. Portable pulse rate detecting device for contact with human body tissue, comprising: a light-emitting source for emitting radiant energy through human body tissue;at least first and second light detectors for detecting intensity of radiant energy after propagation through human body tissue and for providing first and second input signals (y1(t), y2(t)) as a function of such detected radiant energy;a device for measuring motion of the detecting device with respect to human body tissue and for providing a motion reference signal (ax(t), ay(t), az(t)) as a function of said motion;processing means for removing, from said first and second input signals (y1(t), y2(t)), motion-related contributions due to movement of the detecting device with respect to human body tissue and producing first and second enhanced signals,said processing means being arranged to calculate a model of said motion-related contributions based on said motion reference signal (ax(t), ay(t), az(t)) and to subtract this model from each of said first and second input signals (y1(t), y2(t)),wherein said processing means is also for removing measurement noise and residual non-modelled contributions from said first and second enhanced signals using a noise reduction algorithm; andmeans for detecting non-stationary signal segments in said motion reference signal (ax(t), ay(t), az(t)) and for disabling at least one of said light-emitting source, said light detectors and said processing means when non-stationary signal segments are detected. | 1. Portable pulse rate detecting device for contact with human body tissue, comprising: a light-emitting source for emitting radiant energy through human body tissue;at least first and second light detectors for detecting intensity of radiant energy after propagation through human body tissue and for providing first and second input signals (y1(t), y2(t)) as a function of such detected radiant energy;a device for measuring motion of the detecting device with respect to human body tissue and for providing a motion reference signal (ax(t), ay(t), az(t)) as a function of said motion;processing means for removing, from said first and second input signals (y1(t), y2(t)), motion-related contributions due to movement of the detecting device with respect to human body tissue and producing first and second enhanced signals,said processing means being arranged to calculate a model of said motion-related contributions based on said motion reference signal (ax(t), ay(t), az(t)) and to subtract this model from each of said first and second input signals (y1(t), y2(t)),wherein said processing means is also for removing measurement noise and residual non-modelled contributions from said first and second enhanced signals using a noise reduction algorithm; andmeans for detecting non-stationary signal segments in said motion reference signal (ax(t), ay(t), az(t)) and for disabling at least one of said light-emitting source, said light detectors and said processing means when non-stationary signal segments are detected. 2. Portable pulse rate detecting device according to claim 1, wherein said model is a nonlinear model. 3. Portable pulse rate detecting device according to claim 2, wherein the model is selected using a Minimum Description Length (MDL) criterion. 4. Portable pulse rate detecting device according to claim 2, wherein said nonlinear model is a third order polynomial moving average model. 5. Portable pulse rate detecting device according to claim 1, wherein said noise reduction algorithm is based on Principal Component Analysis (PCA). 6. Portable pulse rate detecting device according to claim 1, wherein said processing means is for detecting peaks in said first and second enhanced signals and for extracting inter-beat intervals from said detected peaks to provide a pulse rate measurement, the detection of said peaks being inhibited during the refractory period of cardiac cells. 7. Portable pulse rate detecting device according to claim 1, wherein said detecting device comprises a strap for being worn around a wrist, said light-emitting source and said light detectors being located on a bottom side of said detecting device in contact with human body tissue. 8. Portable pulse rate detecting device according to claim 1, wherein said processing means is a digital signal processor DSP housed within said detecting device. 9. Portable pulse rate detecting device according to claim 1, wherein said device for providing the motion reference signal (ax(t),... | CSEM Centre Suisse d'Electronique et de Microtechnique SA | A61B005/02 A61B006/00
US07025425 | B2 | 2006-04-11 | Method, system, and computer program product for the evaluation of glycemic control in diabetes from self-monitoring data | 2001-03-29 | A method, system, and computer program product related to the diagnosis of diabetes, and is directed to predicting the long-term risk of hyperglycemia, and the long-term and short-term risks of severe hypoglycemia in diabetics, based on blood glucose readings collected by a self-monitoring blood glucose device. The method, system, and computer program product pertain directly to the enhancement of existing home blood glucose monitoring devices, by introducing an intelligent data interpretation component capable of predicting both HbA1c and periods of increased risk of hypoglycemia, and to the enhancement of emerging continuous monitoring devices by the same features. With these predictions the diabetic can take steps to prevent the adverse consequences associated with hyperglycemia and hypoglycemia. | 1. A computerized method for evaluating the HbA1c of a patient based on blood glucose (BG) data collected over a predetermined duration, said method comprising: computing weighted deviation toward high blood glucose (WR) and estimated rate of change of blood glucose (Dr) based on said collected BG data; andestimating HbA1c using a predetermined mathematical formula based on said computed WR and Dr. | 1. A computerized method for evaluating the HbA1c of a patient based on blood glucose (BG) data collected over a predetermined duration, said method comprising: computing weighted deviation toward high blood glucose (WR) and estimated rate of change of blood glucose (Dr) based on said collected BG data; andestimating HbA1c using a predetermined mathematical formula based on said computed WR and Dr. 2. The method of claim 1, wherein: said computed WR is mathematically defined from a series of BG readings x1, x2, . . . xn taken at time points t1, t2, . . . , tn as: WR=1n⁢∑i=1n⁢w⁢⁢r⁡(xi;1)where: wr(BG;b)=10f(BG)b if f(BG)>0 and 0 otherwise,b=1, representing a weighting parameter, andsaid computed Dr is mathematically defined as: Dr=average of sk+1−sk,where: sk=10S(k+t1)2 for k=0, 1, . . . , tn−t1,S(tj)=f(xj), for j=1, . . . , n. 3. The method of claim 1, wherein said estimate of HbA1c from said BG monitoring data is mathematically defined as: Estimated HbA1c=0.9008(WR)−0.89207(Dr)+6.7489. 4. The method of claim 1, further comprising: defining predetermined categories for the estimate of HbA1c, each of said HbA1c estimate categories representing a range of values for estimated HbA1c; andassigning said estimated HbA1c to at least one of said HbA1c estimate categories. 5. The method of claim 4, wherein said HbA1c estimate categories are defined as follows: classified category 1, wherein said estimated HbA1c is less than about 7.8;classified category 2, wherein said estimated HbA1c is between about 7.8 and about 8.5;classified category 3, wherein said estimated HbA1c is between about 8.5 and about 9.0;classified category 4, wherein said estimated HbA1c is between about 9.0 and about 9.6;classified category 5, wherein said estimated HbA1c is between about 9.6 and about 10.3;classified category 6, wherein said estimated HbA1c is between about 10.3 and about 11.0; andclassified category 7, wherein said estimated HbA1c is above about 11.0. 6. The method of claim 5, further comprising: defining predicted confidence intervals for corresponding said HbA1c estimate categories, wherein said predicted confidence intervals are defined as follows: said classified category 1 corresponds with a predicted HbA1c less than about 8.0;said classified category 2 corresponds with a predicted HbA1c between about 8.0 and about 8.5;said classified category 3 corresponds with a predicted HbA1c between about 8.5 and about 9.0;said classified category 4 corresponds with a predicted HbA1c between about 9.0 and about 9.5;said classified category 5 corresponds with a predicted HbA1c between about 9.5 and about 10.1;said classified category 6 corresponds with a predicted HbA1c between about 10.1 and about 11.0; andsaid classified category 7 corresponds with a predicted HbA1c above about 11.0. 7. The method of claim 4, further comprising: defining predicted confidence intervals for corresponding said HbA1c, each of said predicted confidence intervals representing a range of val... | University of Virginia Patent Foundation | A61B005/05
US06984374 | B2 | 2006-01-10 | Method for the evaluation of implantable materials | 2003-01-30 | A method for the evaluation of a material to determine whether the material is susceptible to bacterial contamination or colonization comprising providing bacteria which are modified to produce a first detectable signal, exposing the material being evaluated to the bacteria and determining whether the first signal is present determining whether the first signal is present on the material or within the material. | 1. A method for an evaluation of an implantable material to determine whether the implantable material is susceptible to bacterial contamination or colonization when implanted into a living animal or human comprising: a) providing bacteria which are modified to produce a first detectable signal;b) exposing the implantable material being evaluated to the modified bacteria by intravenously introducing the modified bacteria into a living animal or human; andc) determining by in vivo detection whether the first signal is present on the implantable material or within the implantable material;where absence of the first signal on the implantable material or within the implantable material indicates that the implantable material is not susceptible to bacterial contamination or colonization and where presence of the first signal on the implantable material or within the implantable material indicates that the implantable material is susceptible to bacterial contamination or colonization; and where the implantable material is non-living. | 1. A method for an evaluation of an implantable material to determine whether the implantable material is susceptible to bacterial contamination or colonization when implanted into a living animal or human comprising: a) providing bacteria which are modified to produce a first detectable signal;b) exposing the implantable material being evaluated to the modified bacteria by intravenously introducing the modified bacteria into a living animal or human; andc) determining by in vivo detection whether the first signal is present on the implantable material or within the implantable material;where absence of the first signal on the implantable material or within the implantable material indicates that the implantable material is not susceptible to bacterial contamination or colonization and where presence of the first signal on the implantable material or within the implantable material indicates that the implantable material is susceptible to bacterial contamination or colonization; and where the implantable material is non-living. 2. The method of claim 1, additionally comprising quantifying the susceptibility of the implantable material to bacterial contamination or colonization by quantifying the amount of the first signal on the implantable material or within the implantable material; where increasing amounts of the first signal on the implantable material or within the implantable material indicates increasing susceptibility of the implantable material to bacterial contamination or colonization. 3. The method of claim 2, where exposing the implantable material being evaluated to the modified bacteria comprises using the implantable material as wound closure material in an animal or human. 4. The method of claim 3, where exposing the implantable material being evaluated to the modified bacteria additionally comprises administering the modified bacteria intravenously to the animal or human. 5. The method of claim 1, where the bacteria are modified to produce a second detectable signal, and where the method additionally comprises determining whether the second signal is present on the implantable material or within the implantable material; where absence of the second signal on the implantable material or within the implantable material indicates that the implantable material is not susceptible to bacterial contamination or colonization and where presence of the second signal on the implantable material or within the implantable material indicates that the implantable material is susceptible to bacterial contamination or colonization. 6. The method of claim 1, where the first signal is light emission in the visible spectrum. 7. The method of claim 5, where the second signal is light emission in the visible spectrum. 8. The method of claim 1, where the bacteria are modified to incorporate a functional green fluorescent protein. 9. The method of claim 1, where the bacteria are modified to incorporate a functional luciferase. 10. The method o... | Loma Linda University | A61K049/00 A61B005/055 A61B010/00 A61B005/00 A61B008/00
US07020508 | B2 | 2006-03-28 | Apparatus for detecting human physiological and contextual information | 2002-08-22 | A detecting apparatus includes a housing support section(s), a housing removably attached thereto, one or more sensors and a processor. An alternate apparatus measures heat flux and includes a known resistivity base member, a processing unit and two temperature measuring devices, one in thermal communication with the body through a thermal energy communicator and the other in thermal communication with the ambient environment. A further alternate apparatus includes a housing or flexible section having an adhesive material on a surface thereof for removably attaching the apparatus to the body. A further alternate apparatus includes a housing having an inner surface having a concave shape in a first direction and convex shape in a second direction substantially perpendicular thereto. Also, an apparatus for detecting heart related parameters includes one or more filtering sensors for generating filtering signals related to the non-heart related motion of the body. | 1. An apparatus for detecting at least one of human physiological and contextual information from the body of a wearer, comprising: a flexible section for engaging at least a portion of said body;a housing removably attached to and supported by said flexible section;a strap attached to said flexible section for securing said flexible section to said body;one or more sensors supported by said housing, said one or more sensors selected from the group consisting of physiological sensors and contextual sensors; anda processing unit supported by said housing in electronic communication with said one or more sensors. | 1. An apparatus for detecting at least one of human physiological and contextual information from the body of a wearer, comprising: a flexible section for engaging at least a portion of said body;a housing removably attached to and supported by said flexible section;a strap attached to said flexible section for securing said flexible section to said body;one or more sensors supported by said housing, said one or more sensors selected from the group consisting of physiological sensors and contextual sensors; anda processing unit supported by said housing in electronic communication with said one or more sensors. 2. An apparatus according to claim 1, further comprising a groove provided in said housing and a tongue provided on said flexible section, said groove being suitable to receive said tongue. 3. An apparatus according to claim 1, further comprising a tongue provided on said housing and a groove provided in said flexible section, said groove being suitable to receive said tongue. 4. An apparatus according to claim 1, further comprising an adhesive material for removably attaching said housing to said flexible section. 5. An apparatus according to claim 1, further comprising a magnet supported by each of said housing and said flexible section. 6. An apparatus according to claim 1, further comprising a stretchable band and a groove provided in said housing, said band suitable to fit around a portion of said flexible section and into said groove when said flexible section is placed over said housing. 7. An apparatus according to claim 1, said housing further comprising a display. 8. An apparatus according to claim 1, said one or more sensors comprising at least one of a sensor for measuring GSR including at least two electrical contacts, a skin temperature sensor, an ambient temperature sensor, an accelerometer, an ambient light sensor, an ambient sound sensor, an EMG sensor, an ECG sensor, a skin impedance sensor, a heat flux sensor, and a sensor for measuring heart related parameters. 9. An apparatus according to claim 1, said one or more sensors comprising at least two sensors. 10. An apparatus for detecting at least one of human physiological and contextual information from the body of a wearer, comprising: a housing support section for engaging at least a portion of said body;a housing removably attached to and supported by said housing support section;one or more sensors supported by said housing, said one or more sensors selected from the group consisting of physiological sensors and contextual sensors;a processing unit supported by said housing in electronic communication with said one or more sensors; andone or more additional housing support sections that may be selectively removably attached to said housing. 11. An apparatus according to claim 10, said housing support section and at least one of said additional housing support sections being suitable to engage different body portions of said wearer. 12. An apparatus accordi... | BodyMedia, Inc. | A61B005/04
US06920351 | B2 | 2005-07-19 | Neural prosthetic using temporal structure in the local field potential | 2002-11-25 | In an embodiment, neural activity of a subject may be measured with an implant in the sensory motor cortex of the subject and used to predict an intended movement. The measured neural activity may be the local field potential (LFP) at an electrode or single unit (SU) activity. The spectral structure of the LFP and the SU activity may be estimated using spectral analysis techniques. The estimated LFP and SU responses may be used to predict and intended movement by the subject. | 1. A method comprising: measuring a local field potential response from neural activity in a subject; estimating a temporal structure of the measured local field potential response; and generating a signal indicative of an intended movement in response to the estimated temporal structure of said measured local field potential response. | 1. A method comprising: measuring a local field potential response from neural activity in a subject; estimating a temporal structure of the measured local field potential response; and generating a signal indicative of an intended movement in response to the estimated temporal structure of said measured local field potential response. 2. The method of claim 1, further comprising: cueing a subject to perform a plurality of movements; measuring a local field potential response in response to each of said movements; and storing the local field potential response corresponding to each of said movements in a memory device. 3. The method of claim 2, wherein generating the signal comprises: comparing the measured local field potential response to a plurality of the stored local field potential responses; and determining the intended movement in response to said comparison. 4. The method of claim 1, wherein the signal comprises an electric signal. 5. The method of claim 1, wherein the measuring is invasive. 6. The method of claim 1, wherein the temporal structure of the measured local field potential response is approximately localized in time and frequency. 7. The method of claim 6, further comprising estimating the temporal structure of using spectral analysis. 8. The method of claim 6, wherein the temporal structure is represented as a spectrogram. 9. The method of claim 1, wherein the intended movement comprises a reach. 10. The method of claim 9, wherein the reach comprises target. 11. The method of claim 9, wherein the reach comprises trajectory. 12. The method of claim 9, wherein the intended movement comprises a grasp. 13. The method of claim 1, wherein the intended movement comprises a saccade. 14. The method of claim 1, wherein the intended movement comprises a head movement. 15. The method of claim 1, wherein the intended movement comprises a smooth pursuit eye movement. 16. The method of claim 1, wherein the intended movement comprises a leg movement. 17. The method of claim 1, wherein the intended movement comprises a go signal. 18. The method of claim 1, wherein the intended movement comprises a scrub signal. 19. The method of claim 1, wherein the neural activity comprises activity in the sensory motor cortex. 20. The method of claim 1, wherein the neural activity comprises activity in the parietal reach region. 21. The method of claim 1, wherein the neural activity comprises activity in the anterior intraparietal cortex. 22. The method of claim 1, wherein the neural activity comprises activity in the lateral intraparietal cortex. 23. The method of claim 1, wherein the neural activity comprises activity in area 5 of the superior parietal lobule. 24. The method of claim 1, wherein the neural activity comprises activity in frontal cortical areas involved in movement, said area selected from the motor cortex, the ventral and dorsal premotor cortex, the supplementary motor area, the supplementary eye field, the fronta... | California Institute of Technology | A61B005/04
US06879851 | B2 | 2005-04-12 | Fiber optic endoscopic gastrointestinal probe | 2001-06-07 | A fiber optic probe and a balloon catheter used in conjunction with optical imaging systems, in particular with systems which deliver and collect a single spatial mode beam, such as a single photon, a multiphoton, confocal imaging and ranging systems, such as fluorescence imaging systems. | 1. A balloon catheter comprising: an inflatable balloon having an inner surface; a catheter having a longitudinal axis and a bore, said bore in communication with said balloon; and an optical fiber, positioned within said bore, and having a lens positioned at a distal end of said optical fiber, said lens positioned at a working distance from said inner surface of said balloon and said distal end oriented substantially at an acute angle as measured from said longitudinal axis of said catheter to said inner surface of said balloon; wherein said fiber is capable of transmitting light having a predetermined distribution of wavelengths, and wherein said balloon is transparent to at least a portion of said distribution of wavelengths of light. | 1. A balloon catheter comprising: an inflatable balloon having an inner surface; a catheter having a longitudinal axis and a bore, said bore in communication with said balloon; and an optical fiber, positioned within said bore, and having a lens positioned at a distal end of said optical fiber, said lens positioned at a working distance from said inner surface of said balloon and said distal end oriented substantially at an acute angle as measured from said longitudinal axis of said catheter to said inner surface of said balloon; wherein said fiber is capable of transmitting light having a predetermined distribution of wavelengths, and wherein said balloon is transparent to at least a portion of said distribution of wavelengths of light. 2. The balloon catheter of claim 1 wherein said optical fiber is rotatable such that said distal end of said optical fiber describes a path of motion about said long axis of said catheter substantially along said inner surface of said balloon. 3. The balloon catheter of claim 1 wherein said optical fiber is translationally moveable such that said distal end of said optical fiber describes a path of motion substantially parallel to said long axis of said catheter. 4. The balloon catheter of claim 1 wherein the balloon has a radius between 2.5 mm and 20 mm when inflated. 5. The balloon catheter of claim 1 wherein said balloon when uninflated and said optical fiber are dimensioned to fit within said catheter. 6. The balloon catheter of claim 1 wherein said working distance is about 1.0 to about 5.0 mm. 7. The balloon catheter of claim 1 wherein said distal end of said optical fiber is flexible. 8. The balloon catheter of claim 7 further comprising a pullwire wherein said pullwire is operably joined to said distal end of said optical fiber. 9. The balloon catheter of claim 1 wherein said lens comprises a GRIN lens. 10. A balloon catheter comprising: an inflatable balloon having an inner surface, a lumen, and a diameter between about 2.5 mm and 20 mm when inflated; a catheter having a bore in communication with said lumen of said balloon; and an optical fiber, positioned within said bore, and having a GRIN lens positioned at a distal end of said optical fiber, said GRIN lens positioned within about 1.0 mm to about 5.0 mm from said inner surface of said balloon when inflated, said optical fiber being rotatably and translationally moveable relative to said catheter; wherein said fiber is capable of transmitting light having a predetermined distribution of wavelengths, wherein said balloon is transparent to at least a portion of said distribution of wavelengths of light, and wherein said optical fiber, when moved translationally and rotationally, focuses said light onto and receives backscatter from structures substantially adjacent to the exterior of said balloon. | Lightlab Imaging, LLC | A61B005/00
US06849046 | B1 | 2005-02-01 | System and method for detecting the state of hydration of a living specimen | 2000-09-21 | The invention provides a system (2) for non-invasive detection of the state of hydration in a living specimen, the system including an RF generator/transmitter (416) for emitting RF radiation signals and for transmitting the radiation signals through a tissue of a living specimen; an RF receiver (8) for receiving RF signals transmitted through the tissue, and for feeding the signals to a processor (10) for comparison of relative attenuation of RF frequencies passing through the tissue with a reference attenuation ratio signal, and an indicator unit (12) for providing an output signal representative of the water content level of the tissue. A method for non-invasive detection of the state of hydration in a living specimen is also described and claimed. | 1. A system for non-invasive detection of the state of hydration in a living specimen, said system comprising: an RF generator/transmitter for emitting RF radiation signals at a plurality of frequencies and for transmitting said radiation signals through a tissue of a living specimen; a receiver adapted to receive RF radiation Signals passing through the living specimen at the plurality of frequencies; a processor; and an indicator unit for providing an output signal representative of a hydration state of said tissue; wherein said processor receives signals from said receiver and compares the intensity of the transmitted RF signals, at the plurality of frequencies, passing through said tissue with that of at least one reference value, indicating the state of hydration of said living tissue. | 1. A system for non-invasive detection of the state of hydration in a living specimen, said system comprising: an RF generator/transmitter for emitting RF radiation signals at a plurality of frequencies and for transmitting said radiation signals through a tissue of a living specimen; a receiver adapted to receive RF radiation Signals passing through the living specimen at the plurality of frequencies; a processor; and an indicator unit for providing an output signal representative of a hydration state of said tissue; wherein said processor receives signals from said receiver and compares the intensity of the transmitted RF signals, at the plurality of frequencies, passing through said tissue with that of at least one reference value, indicating the state of hydration of said living tissue. 2. The system as claimed in claim 1, wherein said generator/transmitter emits RF radiation at frequencies of between 10 kHz and 300 GHz. 3. The system as claimed in claim 1, wherein said indicator unit provides an audible, visual or palpating output signal. 4. The system as claimed in claim 1, further including means for attaching said system to a wrist the user's body. 5. The system as claimed in claim 1, wherein the plurality of frequencies include at least one low frequency at which RF energy causes induction of ionic motion of ions dissolved in the body water. 6. The system as claimed in claim 5, wherein the plurality of frequencies include a frequency around 40 MHz. 7. The system as claimed in claim 5, wherein the plurality of frequencies include at least one high frequency at which RE energy is absorbed by body water by rotational excitation of water molecules. 8. The system as claimed in claim 7, wherein the plurality of frequencies include a frequency around 2.4 GHz. 9. The system as claimed in claim 1, wherein the plurality of frequencies include more than two frequencies. 10. The system as claimed in claim 1, wherein the indicator unit provides a warning signal when the living specimen is dehydrated. 11. The system as claimed in claim 1, wherein the indicator unit provides a warning signal when the living specimen is dehydrated. 12. The system as claimed in claim 1, wherein the processor is adapted to store the intensity of the transmitted RF signals, at the plurality of frequencies, passing through said tissue, under standard conditions, for use as the at least one reference value. 13. The system as claimed in claim 1, wherein the processor is preset with the at least one reference values, in accordance with known average body water content levels. 14. The system as claimed in claim 1, wherein the transmitter and receiver have separate ports leading to the living specimen. 15. The system as claimed in claim 1, wherein said at least one reference value comprises at least one attenuation ratio. 16. The system as claimed in claim 1, wherein said receiver is adapted to receive a signal transmitted through said tissue. 17. A method for n... | - | A61B005/00
US06939309 | B1 | 2005-09-06 | Electrophysiology mapping system | 2000-04-12 | A mapping catheter is positioned in a heart chamber, and active electrode sites are activated to impose an electric field within the chamber. The blood volume and wall motion modulates the electric field, which is detected by passive electrode sites on the preferred catheter. Electrophysiology measurements, as well as geometry measurements, are taken from the passive electrodes and used to display a map of intrinsic heart activity. | 1. A system for displaying electrophysiology data collected from a chamber in the patient's heart, to a user, said heart having intrinsic electrical activity, said system comprising: a) a set of passive electrodes adapted for placement into position within said heart; b) a set of active electrodes adapted for placement into position within said heart; c) a pulse generator coupled to said set of active electrodes for creating and applying an electric field within said heart; d) a signal digitizer coupled to said set of passive electrodes for generating a numerical representation of electric field strength at each of said set of passive electrodes, generating a set of measured potentials; e) a signal conditioner for extracting from said measured potentials a map representation of the interior surface of said heart, and for extracting from said measured potentials a representation of said intrinsic electrical activity of said heart chamber; f) display means for displaying said representation of said electrical activity on said map representation. | 1. A system for displaying electrophysiology data collected from a chamber in the patient's heart, to a user, said heart having intrinsic electrical activity, said system comprising: a) a set of passive electrodes adapted for placement into position within said heart; b) a set of active electrodes adapted for placement into position within said heart; c) a pulse generator coupled to said set of active electrodes for creating and applying an electric field within said heart; d) a signal digitizer coupled to said set of passive electrodes for generating a numerical representation of electric field strength at each of said set of passive electrodes, generating a set of measured potentials; e) a signal conditioner for extracting from said measured potentials a map representation of the interior surface of said heart, and for extracting from said measured potentials a representation of said intrinsic electrical activity of said heart chamber; f) display means for displaying said representation of said electrical activity on said map representation. 2. The system of claim 1, wherein said map representation is a dynamic representation reflecting motion of said heart chamber over time. 3. The system of claim 1, wherein said map representation is a representation reflecting a static position of said heart chamber in time. 4. The system of claim 1, wherein said passive electrodes are located on a basket; said basket located on the distal end of a catheter; said basket having a plurality of conformal fingers adapted for engagement with said heart chamber. 5. The system of claim 4, wherein said active electrodes are located on said catheter proximate said passive electrodes. 6. The system of claim 1, wherein said passive electrodes are located proximate a balloon; said balloon located on the distal end of a catheter. 7. The system of claim 6, wherein said active electrodes are located on said catheter proximate said passive electrodes. 8. The system of claim 1, wherein said signal conditioner extracts said map representation from said intrinsic electrical activity of said heart from said measured potentials in the frequency domain by filtering said measured potentials. 9. A system for measuring and displaying electrophysiologic signals originating in a patient's heart, to an observer, said system comprising: a) an active electrode set adapted for placement into position within said patient's heart; b) a passive electrode set adapted for placement into position within said patient's heart in a known physical relationship with said active electrode set, defining a set of passive electrode sites; c) an electric field generator coupled to said active electrode set for generating an oscillating electric interrogation field around said passive electrode sites; d) a signal conditioner coupled to said passive electrode set for extracting low frequency electrophysiologic measurements from said passive electrode sites, generating an electrophysiologic activi... | Endocardial Solutions, Inc. | A61B005/04
US06866640 | B2 | 2005-03-15 | Adaptor for use with point-of-care testing cartridge | 2002-03-26 | An adaptor is provided to facilitate delivery of a fluid specimen from a syringe to a point-of-care testing cartridge. The adaptor includes a tube with an outlet end, an inlet end and a passage extending between the ends. The outlet end of the tube is dimensioned for mating with the entry port of the testing cartridge. The inlet end of the tube is dimensioned for mating with the syringe. The adaptor further includes a support wall extending transversely from the tube and at least one positioning wall extending from the support wall. The positioning wall is spaced transversely from the outlet of the tube by a distance substantially equal to the distance between a side wall of the testing cartridge and the entry port of the testing cartridge. A bottom wall and a lock arm extend from the positioning wall. The bottom wall of the adaptor engages the bottom wall of the testing cartridge and the lock arm of the adaptor engages the top wall of the testing cartridge. | 1. An adaptor for use with a testing cartridge and a syringe to facilitate delivery of a fluid specimen from said syringe to said testing cartridge, said testing cartridge having a housing with top and bottom walls defining a thickness for said testing cartridge, said testing cartridge further having a pair of opposed side walls and a pair of opposed end walls extending between said too and bottom walls, said entry port of said testing cartridge extending through said top wall at a first distance from one said side wall and at a second distance from one said end wall, said adaptor comprising: a bottom wall, a positioning wall extending angularly from said bottom wall, a support wall cantilevered from said positioning wall and spaced from said bottom wall, a tapered tube connected to said support wall at a location spaced from said positioning wall, said tapered tube having an outlet end, an inlet end and a passage extending between said ends, said outlet end being spaced from said bottom wall, of said adaptor by a distance greater than the thickness of said testing cartridge and being spaced from said positioning wall by a distance substantially equal to said first distance, whereby said outlet end of said tapered tube substantially aligns with said entry port of said testing cartridge when said bottom wall and said positioning wall of said adaptor are engaged with said bottom wall and said side wall of said testing cartridge. | 1. An adaptor for use with a testing cartridge and a syringe to facilitate delivery of a fluid specimen from said syringe to said testing cartridge, said testing cartridge having a housing with top and bottom walls defining a thickness for said testing cartridge, said testing cartridge further having a pair of opposed side walls and a pair of opposed end walls extending between said too and bottom walls, said entry port of said testing cartridge extending through said top wall at a first distance from one said side wall and at a second distance from one said end wall, said adaptor comprising: a bottom wall, a positioning wall extending angularly from said bottom wall, a support wall cantilevered from said positioning wall and spaced from said bottom wall, a tapered tube connected to said support wall at a location spaced from said positioning wall, said tapered tube having an outlet end, an inlet end and a passage extending between said ends, said outlet end being spaced from said bottom wall, of said adaptor by a distance greater than the thickness of said testing cartridge and being spaced from said positioning wall by a distance substantially equal to said first distance, whereby said outlet end of said tapered tube substantially aligns with said entry port of said testing cartridge when said bottom wall and said positioning wall of said adaptor are engaged with said bottom wall and said side wall of said testing cartridge. 2. The adaptor of claim 1, wherein said positioning wall is a first positioning wall, and wherein said adaptor further comprises a second positioning wall connected to said bottom wall and spaced from said outlet end of said tube by a distance substantially equal to said second distance. 3. The adaptor of claim 1, further comprising a lock arm extending from said positioning wall and configured for locked engagement with said top wall of said testing cartridge when said outlet end of said tube is aligned with said entry port of said testing cartridge. 4. The adaptor of claim 1, wherein said adaptor is a formed of plastic, and wherein said bottom wall, said positioning wall, said support wall and said tapered tube are unitary with one another. 5. The adaptor of claim 1, wherein said bottom wall of said adaptor is substantially perpendicular to said positioning walls and to said tube. 6. The adaptor of claim 1, wherein said inlet of said tube of said adaptor is tapered. 7. The adaptor of claim 6, wherein said inlet of the tube of said adaptor comprising a pair of opposite Luer projections for threaded engagement with a Luer collar. 8. An adaptor for a testing cartridge, said adaptor comprising a bottom wall, first and second positioning walls angularly aligned to one another and projecting rigidly and unitarily from the bottom wall, a support wall cantilevered rigidly and unitarily from said first positioning wall and a tapered tube extending rigidly and unitarily from an end of said support wall spaced from said first... | Becton, Dickinson and Company | A61B005/00 B65D081/00
US06976953 | B1 | 2005-12-20 | Maintaining the alignment of electric and magnetic fields in an x-ray tube operated in a magnetic field | 2002-09-26 | A new technique for maintaining the alignment of electric and magnetic fields in an x-ray tube so the tube can be operated in the presence of a static external magnetic field without being negatively affected thereby. Deflection of the electron beam of the x-ray tube by the high magnetic field is reduced or eliminated by modifying or canceling, at a location near the electron beam, the magnetic field components transverse to the beam. In a preferred embodiment, a set of electromagnet coils are positioned on or near the tube and oriented in a way that when current is applied internal magnetic fields are produced in a direction opposite to the transverse magnetic fields, thereby causing cancellation. In one implementation, one or more sensors are used to detect the transverse magnetic fields. The sensor is positioned near the electron beam, either inside or outside the x-ray tube. The sensor produces a signal dependent on a static magnetic field component transverse to the desired direction of the electron beam. This signal is used to control the amount of current applied to the coils. A controller and a feedback circuit may be included to adjust in real time the amount of current being applied. | 1. An x-ray source positioned within an external magnetic field, said external magnetic field having a component transverse to an axis of said x-ray source, said x-ray source comprising: a magnetic field sensor for generating a signal sensitive to said transverse component. | 1. An x-ray source positioned within an external magnetic field, said external magnetic field having a component transverse to an axis of said x-ray source, said x-ray source comprising: a magnetic field sensor for generating a signal sensitive to said transverse component. 2. The x-ray source of claim 1, further comprising: a first electromagnetic coil positioned such that when a selected current is applied to said coil, a first internal magnetic field is created within said x-ray source and a first transverse component of said first internal magnetic field opposes said transverse component of said external magnetic field. 3. The x-ray source of claim 2, wherein said x-ray source includes an x-ray tube; andsaid first internal magnetic field is created in a sensitive portion of said x-ray tube between a cathode and an anode. 4. The x-ray source of claim 2, further comprising: a second electromagnetic coil positioned to produce a second internal magnetic field having a second transverse component that is transverse to said x-ray source axis and substantially perpendicular to said first transverse component of said first internal magnetic field. 5. The x-ray source of claim 4, wherein said sensor is sensitive separately to said first and second transverse components of said first and second internal magnetic fields for controlling respective current applied to said first and second electromagnetic coils. 6. The x-ray source of claim 5, further comprising: a feedback means for adjusting said respective current according to an output of said sensor. 7. The x-ray source of claim 1, further comprising: a feedback means being operatively coupled to said magnetic field sensor and responsive to said signal from said magnetic field sensor. 8. The x-ray source of claim 1, wherein said sensor is a Hall effect probe. 9. A method for maintaining alignment of electric and magnetic fields in an x-ray source operated in a high magnetic field, wherein said high magnetic field has a component transverse to an axis of said x-ray source, said method comprising the steps of: detecting said transverse component with a magnetic field sensor;generating a corresponding signal dependent on said transverse component; andadjusting current applied to a first electromagnetic coil positioned in said x-ray source in accordance with said signal. 10. The method of claim 9, further comprising the step of: repeating said detecting, said generating, and said adjusting steps until said signal is effectively zero. 11. The method of claim 9, wherein said x-ray source includes an x-ray tube; and whereinsaid first electromagnetic coil is positioned such that when a selected current is applied to said coil, a first internal magnetic field is created within said x-ray tube in a sensitive portion between a cathode and an anode and a first transverse component of said first internal magnetic field opposes said transverse component of said external magnetic field. 12. The method of ... | The Board of Trustees of the Leland Stanford Junior University | A61B005/05
US06899684 | B2 | 2005-05-31 | Method of respiratory gas analysis using a metabolic calorimeter | 2002-06-10 | A method of determining a respiratory parameter for a subject using an indirect calorimeter is provided. The indirect calorimeter includes a respiratory connector for passing inhaled and exhaled gases, a flow pathway operable to receive and pass inhaled and exhaled gases having a flow tube within the flow pathway through which the inhaled and exhaled gases pass, a flow meter for determining an instantaneous flow volume of the inhaled and exhaled gases, a component gas concentration sensor for determining an instantaneous fraction of a predetermined component gas and a computation unit having a processor and a memory. The method includes the steps of initializing the indirect calorimeter and the subject breathing into the respiratory connector if the indirect calorimeter is initialized, sensing the flow volume of the inhaled and exhaled gases passing through the flow pathway using the flow meter and transmitting a signal representing the sensed flow volume to the computation unit. The method also includes the steps of sensing a concentration of a predetermined component gas as the inhaled and exhaled gases pass through the flow pathway using the component gas sensor, and transmitting a signal representing the sensed concentration of the predetermined component gas to the computation unit. The method further includes the steps of calculating at least one respiratory parameter for the subject as the subject breathes through the calorimeter using the sensed flow volume and the sensed concentration of the predetermined component gas, and providing the subject with the at least one respiratory parameter. | 1. A method of determining a respiratory parameter for a subject using an indirect calorimeter, comprising: initializing the indirect calorimeter, wherein the indirect calorimeter includes a respiratory connector for passing inhaled and exhaled gases as the subject breathes, a flow pathway operable to receive and pass the inhaled and exhaled gases and having a flow tube within the flow pathway through which the inhaled and exhaled gases pass, a flow meter, a component gas concentration sensor, and a computation unit having a processor and a memory; passing the inhaled and exhaled gases through the respiratory connector, if the indirect calorimeter is initialized; sensing a flow volume of the inhaled and exhaled gases passing through the flow pathway, using the flow meter, and transmitting a signal representing the sensed flow volume to the computation unit; sensing a concentration of a predetermined component gas in at least one of the inhaled and exhaled gases passing through the flow pathway, using the component gas concentration sensor, and transmitting a signal representing the sensed concentration to the computation unit; determining if the subject's breathing has reached steady state; calculating at least one respiratory parameter for the subject at steady state using the sensed flow volume and the sensed concentration; and providing the at least one respiratory parameter to the subject. | 1. A method of determining a respiratory parameter for a subject using an indirect calorimeter, comprising: initializing the indirect calorimeter, wherein the indirect calorimeter includes a respiratory connector for passing inhaled and exhaled gases as the subject breathes, a flow pathway operable to receive and pass the inhaled and exhaled gases and having a flow tube within the flow pathway through which the inhaled and exhaled gases pass, a flow meter, a component gas concentration sensor, and a computation unit having a processor and a memory; passing the inhaled and exhaled gases through the respiratory connector, if the indirect calorimeter is initialized; sensing a flow volume of the inhaled and exhaled gases passing through the flow pathway, using the flow meter, and transmitting a signal representing the sensed flow volume to the computation unit; sensing a concentration of a predetermined component gas in at least one of the inhaled and exhaled gases passing through the flow pathway, using the component gas concentration sensor, and transmitting a signal representing the sensed concentration to the computation unit; determining if the subject's breathing has reached steady state; calculating at least one respiratory parameter for the subject at steady state using the sensed flow volume and the sensed concentration; and providing the at least one respiratory parameter to the subject. 2. The method as set forth in claim 1, wherein the predetermined component gas is oxygen, and the component gas concentration sensor is an oxygen sensor. 3. The method as set forth in claim 2, wherein the oxygen sensor is a fluorescence quench type oxygen sensor. 4. The method as set forth in claim 2, wherein initializing the indirect calorimeter includes heating the oxygen sensor to a predetermined temperature in order to sense a concentration of oxygen in at least one of the inhaled and exhaled gases. 5. The method as set forth in claim 1, wherein initializing the indirect calorimeter includes initializing the flow meter by determining if there is zero flow through the flow meter and zeroing the flow meter if a flow is indicated. 6. The method as set forth in claim 1, wherein the respiratory connector is a mask. 7. The method as set forth in claim 1, wherein determining if the subject's breathing has reached steady state includes: determining an amount of oxygen consumed by the subject using the sensed flow volume and the sensed concentration; and determining if the amount of oxygen consumed has reached steady state. 8. The method as set forth in claim 7, further including estimating a steady state amount of oxygen consumed if the subject's breathing does not reach steady state. 9. The method as set forth in claim 7, further including notifying the subject if the subject's breathing does not reach steady state. 10. The method as set forth in claim 7, further including waiting a predetermined period of time after detecting a flow of the inhaled ... | HealtheTech, Inc. | A61B005/08
US06842635 | B1 | 2005-01-11 | Optical device | 1999-07-30 | There is described a sensor device which comprises light source means for emitting a light beam, photodetector means for receiving the light beam after passing through or being reflected within living tissue and arranged to provide signals corresponding to the intensities of the respective wavelength of light received by the photodetector means characterised in that the sensor device measures blood oxygen saturation. The device can be used in conjunction with a conventional pulse oximeter. There is also described a method of measuring blood oxygen saturation. | 1. A method for monitoring oxygenation of blood comprising the steps of: determining a first reference spectrum over a plurality of wavelengths; determining a second reference spectrum over the plurality of wavelengths; transmitting light containing the plurality of wavelengths into blood; measuring a remitted spectrum over the plurality of wavelengths; calculating a measured blood absorption spectrum as a function of the remitted spectrum, the first reference spectrum and the second reference spectrum; and, removing effects of light scattering from the measured blood absorption spectrum by a removing effects method comprising the steps of: calculating a correction function that is a function of a plurality of isobestic points of the measured blood absorption spectrum; correcting the measured blood absorption spectrum by the correction function; normalizing the measured blood absorption spectrum following the correcting step; calculating an optimal spectrum as a function of a substantially oxygenated reference absorption spectrum and a substantially deoxygenated reference absorption spectrum, so that the optimal spectrum best matches the measured blood absorption spectrum in a determined sense; and, calculating an estimate of SO2 as a function of the optimal spectrum. | 1. A method for monitoring oxygenation of blood comprising the steps of: determining a first reference spectrum over a plurality of wavelengths; determining a second reference spectrum over the plurality of wavelengths; transmitting light containing the plurality of wavelengths into blood; measuring a remitted spectrum over the plurality of wavelengths; calculating a measured blood absorption spectrum as a function of the remitted spectrum, the first reference spectrum and the second reference spectrum; and, removing effects of light scattering from the measured blood absorption spectrum by a removing effects method comprising the steps of: calculating a correction function that is a function of a plurality of isobestic points of the measured blood absorption spectrum; correcting the measured blood absorption spectrum by the correction function; normalizing the measured blood absorption spectrum following the correcting step; calculating an optimal spectrum as a function of a substantially oxygenated reference absorption spectrum and a substantially deoxygenated reference absorption spectrum, so that the optimal spectrum best matches the measured blood absorption spectrum in a determined sense; and, calculating an estimate of SO2 as a function of the optimal spectrum. 2. The method for monitoring oxygenation of blood according to claim 1, wherein the step of normalizing the measured blood absorption spectrum following the correcting step of said removing effects method comprises the steps of: determining an area under the measured blood absorption spectrum; and, dividing the measured blood absorption spectrum by a function of the area. 3. The method for monitoring oxygenation of blood according to claim 1, wherein the correction function is linear. 4. The method for monitoring oxygenation of blood according to claim 1, wherein the correction function is a function of two isobestic points of the measured blood absorption spectrum. 5. The method for monitoring oxygenation of blood according to claim 4, wherein the correction function is a line that passes through the two isobestic points. 6. The method for monitoring oxygenation of blood according to claim 5, wherein the two isobestic points are at 522 nanometers and 586 nanometers, respectively. 7. The method for monitoring oxygenation of blood according to claim 1, wherein the first reference spectrum is a spectrally neutral “white” spectrum and the second reference spectrum represents an ambient “dark” spectrum. | Edwards Lifesciences LLC | A61B005/00
USRE038800 | E1 | 2005-09-20 | NIR clinical opti-scan system | 2002-06-27 | The present invention relates to three-dimensional optical imaging techniques and, more particularly, to the detection and three-dimensional imaging of absorbing and/or scattering structures in complex random media, such as human body tissue, by detecting scattered light emerging from the medium. An apparatus for optical tomographic imaging of tissue structures with non-uniform surface geometries in accordance with the invention comprises: an optical source capable of providing light having a wavelength capable of at least attenuated transmission through the tissue; a fiber array consisting of fiber bundles for transmitting light from the optical source to the tissue to be imaged, and a second fiber array consisting of fiber bundles for receiving light scattered by the tissue; an adjustable assembly comprising an adjustable support member supporting one end of each fiber bundle, the fiber bundles being distributed along the portion of each support member which conforms to a surface of a specimen being imaged so as to transmit light into, and collect light emanating from, the surface of the imaged tissue at a multitude of spaced-apart points; and a detector array receiving light collected by the fibers in the second fiber bundle. | 1. An apparatus for optical tomographic imaging of tissue structures with non-uniform surface geometries, comprising: an optical source capable of providing light having a wavelength capable of at least attenuated transmission through the tissue; a first fiber array including fiber bundles for transmitting light from the optical source to the tissue to be imaged; a second fiber array including fiber bundles for receiving light scattered by the tissue; an adjustable assembly including anat least one adjustable support member supporting one end of each fiber bundle, the fiber bundles being from the first and second fiber array and being distributed along the portion of each said one or more adjustable support member which conforms to a surface of a specimenthe tissue being imaged so as to transmit light into, and collect light emanating from, the surface of the imaged tissue at a multitude of spaced-apart points; and a detector array receiving primarily steady-state scattered light collected by the fibers in the second fiber bundle, an output of the detector array that corresponds to the steady-state scattered light useable as a variable in a radiation transport equationarray. | 1. An apparatus for optical tomographic imaging of tissue structures with non-uniform surface geometries, comprising: an optical source capable of providing light having a wavelength capable of at least attenuated transmission through the tissue; a first fiber array including fiber bundles for transmitting light from the optical source to the tissue to be imaged; a second fiber array including fiber bundles for receiving light scattered by the tissue; an adjustable assembly including anat least one adjustable support member supporting one end of each fiber bundle, the fiber bundles being from the first and second fiber array and being distributed along the portion of each said one or more adjustable support member which conforms to a surface of a specimenthe tissue being imaged so as to transmit light into, and collect light emanating from, the surface of the imaged tissue at a multitude of spaced-apart points; and a detector array receiving primarily steady-state scattered light collected by the fibers in the second fiber bundle, an output of the detector array that corresponds to the steady-state scattered light useable as a variable in a radiation transport equationarray. 2. The apparatus of claim 1 wherein the optical source is selected from the group consisting of: diode lasers, Ti:Sapphire lasers, dye lasers, multi-wavelength lasers, continuous-wave (CW) lasers, and pulsed-power lasers. 3. The apparatus of claim 2 wherein the optical source emits light in the near infrared portion of the electromagnetic spectrum. 4. The apparatus of claim 1 wherein the first fiber array has translatable light-receiving ends adjacent to the light source. 5. The apparatus of claim 1 further comprising a focusing lens disposed between the optical source and the first fiber array. 6. The apparatus of claim 1 wherein the one or more adjustable support members are constructed ofcomprise deformable material. 7. The apparatus of claim 1 wherein the one or more adjustable support members comprise adjustable irises. 8. The apparatus of claim 1 further comprising an attenuator disposed between the fiber bundles which transmit light away from the adjustable assemblyof the second fiber array and the detector array. 9. The apparatus of claim 8 wherein the attenuator comprises: an inner cylinder housing another end the fiber bundles from the adjustable assemblyof the second fiber array; an outer cylinder housing theone end of a plurality of fiber bundles that direct light to the detector array; and a plurality of attenuators disposed between the inner and outer cylinders. 10. The apparatus of claim 9 wherein the plurality of attenuators comprise neutral density filters having a respective plurality of density values. 11. The apparatus of claim 9 wherein the plurality of attenuators comprise polarizers. 12. The apparatus of claim 1 further comprising a reference detector for monitoring the optical source from a portion of the first fiber array receiving light ... | The Research Foundation of State University of New York | A61B005/00 G01N021/00
US06890310 | B2 | 2005-05-10 | Adaptor for use with point-of-care testing cartridge | 2002-03-27 | An adaptor is provided to facilitate delivery of a fluid specimen from a syringe to a point-of-care testing cartridge. The adaptor includes a mounting sleeve for snapped engagement over a portion of the testing cartridge, including the entry port to the testing cartridge. The adaptor further includes a Luer fitting unitarily formed with the mounting sleeve. The Luer fitting has an inlet end configured for mating with the syringe and an outlet end configured for alignment with the entry port of the testing cartridge. The adaptor can be snapped onto the testing cartridge. The Luer tip of the syringe then can be mated with the Luer fitting of the adaptor for safe and convenient delivery of the specimen through the adaptor and into the testing cartridge. | 1. An adaptor for use with a testing cartridge and a syringe to facilitate delivery of a fluid specimen from said syringe to said testing cartridge, said testing cartridge having a bottom wall, a top wall and at least one side wall extending between said top and bottom walls, an entry port extending through said top wall for delivering the specimen to the testing cartridge, said adaptor comprising a mounting sleeve configured for slidable mounting over portions of said top and bottom wall of said testing cartridge including portions of said top wall having said entry port therein, at least one locking detent formed on said mounting sleeve for locked engagement with said testing cartridge, and a Luer filling having an outlet end extending through said mounting sleeve at a location for alignment with said entry port when said mounting sleeve is mounted on said testing cartridge, an inlet end spaced from said outlet end and a passage extending between said ends, said passage having at least one non-linear portion so that said inlet end is aligned at an angle to said outlet end, said inlet end being configured for fluid-tight mounting with said syringe. | 1. An adaptor for use with a testing cartridge and a syringe to facilitate delivery of a fluid specimen from said syringe to said testing cartridge, said testing cartridge having a bottom wall, a top wall and at least one side wall extending between said top and bottom walls, an entry port extending through said top wall for delivering the specimen to the testing cartridge, said adaptor comprising a mounting sleeve configured for slidable mounting over portions of said top and bottom wall of said testing cartridge including portions of said top wall having said entry port therein, at least one locking detent formed on said mounting sleeve for locked engagement with said testing cartridge, and a Luer filling having an outlet end extending through said mounting sleeve at a location for alignment with said entry port when said mounting sleeve is mounted on said testing cartridge, an inlet end spaced from said outlet end and a passage extending between said ends, said passage having at least one non-linear portion so that said inlet end is aligned at an angle to said outlet end, said inlet end being configured for fluid-tight mounting with said syringe. 2. The adaptor of claim 1, wherein the mounting sleeve comprises a bottom wall for slidable engagement with said bottom wall of said testing cartridge and a top wall configured for sliding engagement with said top wall of said testing cartridge, said locking detents being formed respectively on said top and bottom walls of said adaptor. 3. The adaptor of claim 2, wherein said mounting sleeve further comprises a side wall extending between said top and bottom walls and disposed for slidable engagement against said side wall of the testing cartridge. 4. The adaptor of claim 3, further comprising an end wall extending between said top and bottom walls and angularly aligned to said side wall. 5. The adaptor of claim 1, wherein portions of said passage adjacent said inlet end of said Luer fitting define a conical taper configured for mating with a Luer tip of said syringe. 6. The adaptor of claim 5, wherein portions of said Luer fitting adjacent said inlet end include a pair of projections configured for threaded engagement with a Luer collar on said syringe. 7. A adaptor for use with a testing cartridge and a syringe to facilitate delivery of a fluid specimen from said syringe to said testing cartridge, said testing cartridge having a bottom wall, a top wall and at least one side wall extending between said ton and bottom walls, an entry port extending through said top wall for delivering the specimen to the testing cartridge, said adaptor comprising a mounting sleeve configured for slidable mounting over portions of said too and bottom wall of said testing cartridge including portions of said too wall having said entry port therein, and a Luer fitting having an outlet end extending through said mounting sleeve at a location for alignment with said entry port when said mounting sleeve is mounted on ... | Becton, Dickinson and Company | A61B005/00 B65D081/00
US06859661 | B2 | 2005-02-22 | Surgical system for use in the course of a knee replacement operation | 2002-01-23 | A system for use by a surgeon during a knee replacement operation comprises a navigation computer whose memory holds data relating to size and shape of at least one implant as well as data obtained during pre-operative scanning. The system also registers the three dimensional position and orientation of the patient's tibia and/or femur during the course of the knee replacement operation. The system is capable of applying tension to the ligaments of the patient's knee in flexion or in extension after resection of the patient's tibia. Moreover the computer is programmed to display images of the bones on the computer's screen at different stages of the operation and with potential planes of cut. | 1. A system for use by a surgeon during implantation in a patient of a tibial implant and of a femoral implant in the course of a knee replacement operation, the system comprising: a computer having memory for holding data relating to size and shape of at least one tibial implant and at least one femoral implant as well as data obtained during pre-operative scanning of the patient relating to the tibia and femur of the patient and also data relating to the three dimensional position and orientation of at least one bone selected from the patient's tibia and the patient's femur obtained during the course of the knee replacement operation; display means for displaying images derived from data held in the memory of the computer relating to the patient's tibia, to the patient's femur, to the at least one tibial implant, and to the at least one femoral implant; registration means for registering the three dimensional position and orientation of the at least one bone during the course of the knee replacement operation; detector means connected to the computer and operatively coupled to the registration means for detecting the position and orientation of the at least one bone and for supplying to the memory of the computer data relating to the three dimensional position and orientation of at least one bone selected from the patient's tibia and the patient's femur obtained during the course of the knee replacement operation; and tensor means for applying tension to the ligaments of the patient's knee in flexion in extension; the computer being configured to display on the display means, at least one more of the following images derived from data held in the memory of the computer and depicting the position and orientation of: (i) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in flexion; (ii) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in extension; (iii) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in flexion and potential planes of cut for resection of the patient's femur; (iv) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in extension and potential planes of cut for resection of the patient's femur; (v) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in flexion and with at least one component selected from the tibial component and the femoral component as it will appear after implantation thereof; and (vi) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in extension and with at least one component selected from the tibial component and the femoral component as it will appear after implantation thereof. | 1. A system for use by a surgeon during implantation in a patient of a tibial implant and of a femoral implant in the course of a knee replacement operation, the system comprising: a computer having memory for holding data relating to size and shape of at least one tibial implant and at least one femoral implant as well as data obtained during pre-operative scanning of the patient relating to the tibia and femur of the patient and also data relating to the three dimensional position and orientation of at least one bone selected from the patient's tibia and the patient's femur obtained during the course of the knee replacement operation; display means for displaying images derived from data held in the memory of the computer relating to the patient's tibia, to the patient's femur, to the at least one tibial implant, and to the at least one femoral implant; registration means for registering the three dimensional position and orientation of the at least one bone during the course of the knee replacement operation; detector means connected to the computer and operatively coupled to the registration means for detecting the position and orientation of the at least one bone and for supplying to the memory of the computer data relating to the three dimensional position and orientation of at least one bone selected from the patient's tibia and the patient's femur obtained during the course of the knee replacement operation; and tensor means for applying tension to the ligaments of the patient's knee in flexion in extension; the computer being configured to display on the display means, at least one more of the following images derived from data held in the memory of the computer and depicting the position and orientation of: (i) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in flexion; (ii) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in extension; (iii) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in flexion and potential planes of cut for resection of the patient's femur; (iv) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in extension and potential planes of cut for resection of the patient's femur; (v) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in flexion and with at least one component selected from the tibial component and the femoral component as it will appear after implantation thereof; and (vi) the patient's tibia and the patient's femur prior to resection of the femur with the knee tensed by the tensor means in extension and with at least one component selected from the tibial component and the femoral component as it will appear after implantation thereof. 2. A system according to claim ... | Finsbury (Development) Limited | A61B005/05
US06882875 | B1 | 2005-04-19 | Visible display for an interventional device | 2000-10-18 | A display is located on an interventional device insertible into a body. More specifically, the display is located on a distal end of the device and is coupled to a sensing system. The sensing system senses a bodily condition. The display receives the sensed bodily condition and displays a signal indicative of the bodily condition within a pre-existing field of view. | 1. An interventional device comprising: an elongated member having a distal end insertible into a body and an opposite proximal end; a sensing system disposed at the distal end of the elongated member for sensing a bodily condition and generating a visual signal; and an intracorporeal display also located at the distal end of the elongated member for displaying the signal inside the body, the signal being indicative of a bodily condition and visible in a pre-existing field of view. | 1. An interventional device comprising: an elongated member having a distal end insertible into a body and an opposite proximal end; a sensing system disposed at the distal end of the elongated member for sensing a bodily condition and generating a visual signal; and an intracorporeal display also located at the distal end of the elongated member for displaying the signal inside the body, the signal being indicative of a bodily condition and visible in a pre-existing field of view. 2. The interventional device of claim 1 wherein the display further comprises an indicator for providing the signal indicative of the bodily condition. 3. The interventional device of claim 2 wherein the indicator comprises at least one light emitter. 4. The interventional device of claim 3 wherein the at least one light emitter comprises (a) a light emitting diode, (b) a liquid crystal display, or (c) a projection display. 5. The interventional device of claim 2 wherein the indicator comprises at least one optical emitter. 6. The interventional device of claim 2 wherein the indicator comprises at least one chemical indicator. 7. The interventional device of claim 6 wherein the at least one chemical indicator comprises a litmus paper. 8. The interventional device of claim 2 wherein the indicator comprises at least one polymeric emitter. 9. The interventional device of claim 8 wherein the at least one polymeric emitter comprises at least one organic light emitting diode. 10. The interventional device of claim 1 wherein the display receives information regarding the bodily condition from the sensing system. 11. The interventional device of claim 1 wherein the sensing system comprises at least one sensor. 12. The interventional device of claim 11 wherein the at least one sensor comprises an electrode capable of sensing the bodily condition. 13. The interventional device of claim 12 wherein the sensing system further comprises a light source and a light detector. 14. The interventional device of claim 12 wherein the sensing system comprises a light source and a light detector. 15. The interventional device of claim 14 wherein the light detector is adapted for monitoring light emission from the light source. 16. The interventional device of claim 14 wherein the light detector is adapted for monitoring light emission from the light source. 17. The interventional device of claim 16 further comprising a filter disposed adjacent the light detector for selectively detecting light emission with in a range of pre-determined wavelength. 18. The interventional device of claim 16 further comprising a filter disposed adjacent the light detector for selectively detecting light emission within a range of pre-determined wavelength. 19. The interventional device of claim 1, further comprising a power source for providing power to the display. 20. The interventional device of claim 1, further comprising a processor for receiving and processing the sensed bodily conditio... | Boston Scientific Corporation | A61B005/05
US07041063 | B2 | 2006-05-09 | Noninvasive measurement of chemical substances | 2003-02-06 | Utilization of a contact device placed on the eye in order to detect physical and chemical parameters of the body as well as the non-invasive delivery of compounds according to these physical and chemical parameters, with signals being transmitted continuously as electromagnetic waves, radio waves, infrared and the like. One of the parameters to be detected includes non-invasive blood analysis utilizing chemical changes and chemical products that are found in the conjunctiva and in the tear film. A transensor mounted in the contact device laying on the cornea or the surface of the eye is capable of evaluating and measuring physical and chemical parameters in the eye including non-invasive blood analysis. The system utilizes eye lid motion and/or closure of the eye lid to activate a microminiature radio frequency sensitive transensor mounted in the contact device. The signal can be communicated by wires or radio telemetered to an externally placed receiver. The signal can then be processed, analyzed and stored. Several parameters can be detected including a complete non-invasive analysis of blood components, measurement of systemic and ocular blood flow, measurement of heart rate and respiratory rate, tracking operations, detection of ovulation, detection of radiation and drug effects, diagnosis of ocular and systemic disorders and the like. | 1. An apparatus for measuring temperature, said apparatus comprising: a contact device to engage a surface of the eye and for detecting temperature on the surface of said eye, said contact device transmitting a detection signal representative of a detected temperature to activate another device based upon the detected temperature, said another device including a detector for detecting infectious agents, said detector being included as part of the contact device. | 1. An apparatus for measuring temperature, said apparatus comprising: a contact device to engage a surface of the eye and for detecting temperature on the surface of said eye, said contact device transmitting a detection signal representative of a detected temperature to activate another device based upon the detected temperature, said another device including a detector for detecting infectious agents, said detector being included as part of the contact device. 2. The apparatus of claim 1, wherein the device includes a wireless transmission device. 3. Method for measuring temperature on the eye, said method comprising the steps of: placing a contact device on a surface of the eye,detecting a temperature on the surface of said eye with the contact device,transmitting a detection signal representative of a detected temperature level, andactivating another device based upon the detected temperature level, the another device including a detector for detecting infectious agents, the detector being included as part of the contact device. | - | A61B005/00
US07261692 | B2 | 2007-08-28 | Skin permeable gas collector and skin permeable gas measuring apparatus | 2001-12-26 | An apparatus for measuring skin permeable gas, which can be used to obtain clinical information by noninvasive and bloodless operations and is useful for in-home medical care and health management. The apparatus includes a device for collecting skin permeable gas including a container for storing skin permeable gas; a partition, disposed at a lower section inside the container, having an opening at the center thereof; and an operating member for opening and closing the opening of the partition. | 1. A device for collecting skin permeable gas, comprising: a container having an opening section, the opening section configured to be brought into intimate contact with skin to collect emitted skin permeable gas, the collected skin permeable gas being delivered and then measured,wherein the container includes: an upper chamber, for storing skin permeable gas and including a ventilation port and delivery port for removing the stored skin permeable gas;a lower chamber configured to be brought into intimate contact with the skin;a partition, having an opening at the center thereof, for separating the upper chamber from the lower chamber; anda push-pull operating member for controlling the flow of the skin permeable gas flowing between the lower chamber and the upper chamber,wherein the lower chamber is configured to be pressed against the skin and then fixed thereto in such a state that the operating member is pressed down to close the opening of the partition using a valve member disposed at the lower end of the operating member, the skin permeable gas emanating from the skin is collected in the lower chamber, the valve member of the operating member is drawn up after the collection to open the opening of the partition to allow the skin permeable gas collected in the lower chamber to flow into the upper chamber, and the valve member of the operating member is then pressed down after a predetermined time to close the opening of the partition to store the skin permeable gas. | 1. A device for collecting skin permeable gas, comprising: a container having an opening section, the opening section configured to be brought into intimate contact with skin to collect emitted skin permeable gas, the collected skin permeable gas being delivered and then measured,wherein the container includes: an upper chamber, for storing skin permeable gas and including a ventilation port and delivery port for removing the stored skin permeable gas;a lower chamber configured to be brought into intimate contact with the skin;a partition, having an opening at the center thereof, for separating the upper chamber from the lower chamber; anda push-pull operating member for controlling the flow of the skin permeable gas flowing between the lower chamber and the upper chamber,wherein the lower chamber is configured to be pressed against the skin and then fixed thereto in such a state that the operating member is pressed down to close the opening of the partition using a valve member disposed at the lower end of the operating member, the skin permeable gas emanating from the skin is collected in the lower chamber, the valve member of the operating member is drawn up after the collection to open the opening of the partition to allow the skin permeable gas collected in the lower chamber to flow into the upper chamber, and the valve member of the operating member is then pressed down after a predetermined time to close the opening of the partition to store the skin permeable gas. | Japan Science and Technology Corporation | A61B005/00
US07047054 | B2 | 2006-05-16 | Laser diode optical transducer assembly for non-invasive spectrophotometric blood oxygenation monitoring | 2001-09-14 | A non-invasive near infrared spectrophotometric monitoring transducer assembly includes a housing member, which is adhered directly on a patient's skin. The housing member contains a prism coupled to a flexible and lightweight single core optical light guide, which provides a means of transferring narrow spectral bandwidth light from multiple distant laser diodes of different wavelengths by use of a multi-fiber optic light combining assembly. Different wavelengths are needed to monitor the level of blood oxygenation in the patient. The assembly also contains a planar light guide mounted on the prism located in the housing member, which light guide contacts the patient's skin when the housing member is adhered to the patient's skin. The light guide controls the spacing between the prism and the patient's skin, and therefore controls the intensity of the area on the patient's skin which is illuminated by the laser light. The housing member contains a photodiode assembly, which detects the infrared light at a second location on the skin to determine light absorption. The photodiode assembly is preferably shielded from ambient electromagnetic interference (EMI) by an optically transparent EMI attenuating window. This rigid window placed over the photodiode also provides a planar interface between the assembly and the skin, improving optical coupling and stability as well as reducing the capacitive coupling between skin and the photodiode resulting in further EMI attenuation. The housing may be associated with a disposable sterile hydrogel coated adhesive envelope, or pad, which when applied to the patient's skin will adhere the housing to the patient's skin. The transducer assembly will thus be reusable, and skin-contacting part of the device, i.e., the envelope or pad can be discarded after a single use. The assembly also includes a laser safety interlock means, which is operable to turn off the laser light output in the event that the assembly accidentally becomes detached from the patient's skin. | 1. A near infrared spectrophotometric (NIRS) monitoring assembly for noninvasive monitoring of blood oxygenation levels in a subject's body, said assembly comprising: a) at least one flexible housing having a window and which can be attached directly to a subject's body;b) a plurality of laser light sources which are operatively associated with said housing, each of said light sources being operative to emit a near infrared light signal of a different wavelength, which light signals are transmitted through said housing window;c) a photodiode assembly which can be attached directly to the subject's body and which is operably connected to said light sources, said photodiode assembly being operative to measure light intensity values emanating from light emitted by said light sources and passing through the subject's body; andd) a laser safety interlock system designed to disable laser operation in case of assembly dislodgment from the patient's body, said safety interlock comprising:I) means for monitoring ambient light and emitted laser light; andII) means for processing signals received by said photodiode assembly so as to determine secure assembly attachment to the subject's body:i) by monitoring ambient light conditions and disabling laser operation when ambient light levels are greater than a predetermined level;ii) by monitoring laser output and disabling laser operation when laser output is outside of a predetermined range; andiii) by monitoring changes in detected laser output over a predetermined measuring period and disabling laser operation if changes in laser output are above or below predetermined levels during a predetermined measuring period. | 1. A near infrared spectrophotometric (NIRS) monitoring assembly for noninvasive monitoring of blood oxygenation levels in a subject's body, said assembly comprising: a) at least one flexible housing having a window and which can be attached directly to a subject's body;b) a plurality of laser light sources which are operatively associated with said housing, each of said light sources being operative to emit a near infrared light signal of a different wavelength, which light signals are transmitted through said housing window;c) a photodiode assembly which can be attached directly to the subject's body and which is operably connected to said light sources, said photodiode assembly being operative to measure light intensity values emanating from light emitted by said light sources and passing through the subject's body; andd) a laser safety interlock system designed to disable laser operation in case of assembly dislodgment from the patient's body, said safety interlock comprising:I) means for monitoring ambient light and emitted laser light; andII) means for processing signals received by said photodiode assembly so as to determine secure assembly attachment to the subject's body:i) by monitoring ambient light conditions and disabling laser operation when ambient light levels are greater than a predetermined level;ii) by monitoring laser output and disabling laser operation when laser output is outside of a predetermined range; andiii) by monitoring changes in detected laser output over a predetermined measuring period and disabling laser operation if changes in laser output are above or below predetermined levels during a predetermined measuring period. 2. The assembly of claim 1, wherein said laser safety interlock system includes a laser start up sequence upon initial attachment of said assembly to the subject's body to determine secure assembly attachment, wherein the ambient light level is monitored and verified to be below a predetermined level prior to allowing laser operation to begin. 3. A near infrared spectrophotometric (NIRS) monitoring assembly for noninvasive monitoring of blood oxygenation levels in a subject's body, said assembly comprising: a) at least one flexible housing having a window which housing can be attached directly to a subject's body;b) a plurality of light sources which are operably associated with said housing, each of said light sources being operable to emit a near infrared light signal of a different wavelength, which light signals are transmitted through said housing window;c) a photodiode assembly which can be attached directly to the subject's body and which is operatively connected to said light sources, said photodiode assembly being operative to measure light intensity values emanating from light emitted by said light sources and passing through the subject's body; andd) a processor which controls actuation of said light sources and which processes light intensity signals from said photodiode assembly; a... | CAS Medical Systems, Inc. | A61B005/00
US07294112 | B1 | 2007-11-13 | Motion monitoring apparatus | 2000-05-15 | The present invention relates to a motion monitoring apparatus and method for monitoring a patient under medical care. A sensor arrangement is provided in the form of a pad which the patient lies on. The sensor arrangement provides a signal which can be monitored to observe motion of the patient and provide an alarm should the motion meet certain predetermined conditions. The invention is particularly applicable for monitoring patients under sedation, recovering from anaesthesia, or in intensive care. The device is particularly useful for veterinary patients. | 1. A method of monitoring a patient under medical care, the method comprising steps of: providing a sensor arrangement which is arranged to detect motion of a head, limbs and trunk of the patient, and respiratory motion of the patient and produce a motion signal in response to the motion;monitoring the motion signal;analysing the motion signal to determine a rate of the motion signal and determine whether the rate of the motion signal is indicative of patient arousal; andproviding an alarm should the monitored motion be indicative of patient arousal;wherein the patient is a non-human animal which is monitored during recovery from anaesthesia or when under sedation. | 1. A method of monitoring a patient under medical care, the method comprising steps of: providing a sensor arrangement which is arranged to detect motion of a head, limbs and trunk of the patient, and respiratory motion of the patient and produce a motion signal in response to the motion;monitoring the motion signal;analysing the motion signal to determine a rate of the motion signal and determine whether the rate of the motion signal is indicative of patient arousal; andproviding an alarm should the monitored motion be indicative of patient arousal;wherein the patient is a non-human animal which is monitored during recovery from anaesthesia or when under sedation. 2. The method in accordance with claim 1, further comprising the step of monitoring a body temperature of the patient and providing an alarm should the body temperature, one of rise above and below predetermined values. 3. The method in accordance with claim 2, wherein a temperature sensor is one of provided proximate and within the patient to constantly monitor the temperature. 4. The method in accordance with claim 3, wherein a control means is arranged to receive signals from the sensor arrangement and the temperature sensor, and process those signals to provide the alarms. 5. The method in accordance with claim 4, wherein the control means is housed in a single unit. 6. The method in accordance with claim 1, wherein the sensor arrangement includes a pad on which the patient lies, the pad mounting a sensor for monitoring the head, the limbs, the trunk, and the respiratory motion of the patient. 7. The method in accordance with claim 6, comprising the further step of providing a separate respiratory motion arrangement for measuring the respiratory motion of the patient, and comparing a signal from the respiratory motion sensor with the signal from the pad sensor, to obtain an indication of bodily motion of the patient. 8. The method in accordance with claim 1, further comprising the step of providing an alarm should the motion of the patient cease to be detected. 9. The method in accordance with claim 8, further comprising the step of providing an alarm should the motion of the patient fall below a predetermined value. 10. The method in accordance with claim 1, further comprising the step of analyzing the motion of the patient by tracking the rate of motion over a period of time. 11. The method in accordance with claim 10, comprising the further step of applying trend analysis to monitor trends in the motion of the patient. 12. The method in accordance with claim 1, wherein the head, the limb, the trunk motion and the respiratory motion are monitored by a single sensor arrangement. 13. The method in accordance with claim 1, wherein the head, the limb, the trunk motion and the respiratory motion are monitored by separate sensor arrangements. 14. The method in accordance with claim 1, wherein the motion signal is a single motion signal generated from a combination of the h... | - | A61B005/08 A61B005/103 A61B005/117
US07081750 | B1 | 2006-07-25 | Dynamic real-time magnetic resonance imaging sequence designer | 2001-05-10 | A system and method for facilitating RF pulse sequence generation and modification and for real-time sequence input modification for use in conjunction with magnetic resonance imaging equipment. A graphical user interface is provided through a display coupled to a digital computer operating as the primary control system for a magnetic resonance imaging scanner and associated hardware. Through the graphical user interface, an operator may choose or design sequences of radiofrequency pulses, gradient waveforms and other input parameters for the magnetic resonance imaging apparatus. Real-time information is also communicated to the operator through the graphical user interface allowing for real-time manipulation of the magnetic resonance imaging inputs and for displaying the magnetic resonance response thereto. | 1. A user interface operable to create, on a display device, a window displaying a plurality of menu editor items for user selection, said menu editor items comprising: a sequence editor item that creates an RF pulse sequence from at least one value; anda sequence tailor editor item configured for user interaction with a graphical representation of a selected pulse sequence, wherein during said user interaction, the selected pulse sequence is graphically displayed to the user,said user interaction including dynamic, non-standard and on-the-fly manipulation of, and modification to, said graphical representation of said selected pulse sequence that is currently undergoing user interaction with real time visual feedback of the interaction on the manipulated pulse sequence to the user, andwherein said menu editor items further comprise an MRI scan setting menu editor item configured for initiation of a magnetic resonance imaging scan. | 1. A user interface operable to create, on a display device, a window displaying a plurality of menu editor items for user selection, said menu editor items comprising: a sequence editor item that creates an RF pulse sequence from at least one value; anda sequence tailor editor item configured for user interaction with a graphical representation of a selected pulse sequence, wherein during said user interaction, the selected pulse sequence is graphically displayed to the user,said user interaction including dynamic, non-standard and on-the-fly manipulation of, and modification to, said graphical representation of said selected pulse sequence that is currently undergoing user interaction with real time visual feedback of the interaction on the manipulated pulse sequence to the user, andwherein said menu editor items further comprise an MRI scan setting menu editor item configured for initiation of a magnetic resonance imaging scan. 2. The user interface in accordance with claim 1, wherein said sequence tailor editor item is activated in response to user selection. 3. The user interface in accordance with claim 1, wherein user selection of said sequence editor item activates a display of at least one sequence parameter operative in creating said pulse sequence, said at least one sequence parameter being operable to accept a default value. 4. The user interface in accordance with claim 3, wherein said at least one sequence parameter is operable to accept a user entered value. 5. The user interface in accordance with claim 3, wherein said at least one sequence parameter is selected from the group consisting of: a gradient resolution parameter, a radio frequency pulse resolution parameter, an echo gathering time parameter, a sequence name parameter, at least one gradient motion compensation parameter, at least one radio frequency pulse characteristic parameter, and at least one data acquisition parameter. 6. The user interface in accordance with claim 3, wherein acceptance, by the user interface, of the at least one sequence parameter activates said sequence tailor editor item. 7. The user interface in accordance with claim 1, wherein activation of said sequence tailor editor item activates display of said pulse sequence and at least one control feature. 8. The user interface in accordance with claim 7, wherein said at least one control feature comprises at least one of a control section, a shape editor, a block editor, and a time scaler. 9. The user interface in accordance with claim 8, wherein said shape editor, when activated, is operable to modify at least one radio frequency pulse characteristic parameter and the radio frequency pulse shape associated with said selected pulse sequence which is displayed to the user. 10. The user interface in accordance with claim 8, wherein said time scaler, when activated, displays at least one vertical line through the graphically displayed selected pulse sequence in order to assist the user in analysi... | Fonar Corporation | G01V003/00 A61B005/055
US07025729 | B2 | 2006-04-11 | Apparatus for detecting sleep apnea using electrocardiogram signals | 2001-09-14 | There is provided a method of determining a diagnostic measure of sleep apnea including the following steps: acquiring an electrocardiogram signal, calculating a set of RR intervals and electrocardiogram-derived respiratory signal from said electrocardiogram, and hence calculating a set of spectral and time-domain measurements over time periods including power spectral density, mean, and standard deviation. These measurements are processed by a classifier model which has been trained on a pre-existing data base of electrocardiogram signals to provide a probability of a specific time period containing apneic episodes or otherwise. These probabilities can be combined to form an overall diagnostic measure. The system also provides a system and apparatus for providing a diagnostic measure of sleep apnea. | 1. An apparatus useful in diagnosing sleep apnea in a human patient, the apparatus comprising: means for measuring a cardiac electrical potential of the patient and for generating an electrocardiogram signal over a time period;means for analyzing the electrocardiogram signal and for deriving a respiratory signal therefrom;means for partitioning the electrocardiogram signal into a set of time periods shorter than the time period of the electrocardiogram signal;means for classifying each time period in the set of time periods as either apneic or normal; andmeans for combining classification results from a plurality of the time periods and for providing a diagnostic measure of sleep apnea for the human patient based on the combined classification results. | 1. An apparatus useful in diagnosing sleep apnea in a human patient, the apparatus comprising: means for measuring a cardiac electrical potential of the patient and for generating an electrocardiogram signal over a time period;means for analyzing the electrocardiogram signal and for deriving a respiratory signal therefrom;means for partitioning the electrocardiogram signal into a set of time periods shorter than the time period of the electrocardiogram signal;means for classifying each time period in the set of time periods as either apneic or normal; andmeans for combining classification results from a plurality of the time periods and for providing a diagnostic measure of sleep apnea for the human patient based on the combined classification results. 2. The apparatus of claim 1, wherein the means for classifying comprises calculating, for each time period in the set of time periods, an associated probability that said each time period is apneic. 3. The apparatus of claim 1, wherein the means for analyzing the electrocardiogram signal comprises means for determining a sequence of flit intervals. 4. The apparatus of claim 3, further comprising means for correcting physiologically unreasonable RR intervals contained in the sequence of RR intervals. 5. The apparatus of claim 1, wherein the means for analyzing the electrocardiogram signal derives the respiratory signal by calculating an area enclosed by the electrocardiogram signal in a region around an R waveform peak. 6. The apparatus of claim 1, further comprising signal processing means for removing interference from the electrocardiogram signal and for producing a processed electrocardiogram signal which is provided to the means for analyzing the electrocardiogram signal. 7. The apparatus of claim 6, further comprising means for recording the processed electrocardiogram signal. 8. The apparatus of claim 1, wherein the means for combining classification results provides a diagnostic measure of sleep apnea for the human patient as at least one of an Apnea-Hyponea Index and a Respiratory Disturbance Index. 9. The apparatus of claim 1, wherein the means for analyzing the electrocardiogram signal comprises a computer software program executed by a processor. 10. The apparatus of claim 1, wherein the means for combining classification results and for providing a diagnostic measure of sleep apnea for the human patient comprises a display interface which provides a visual indication of a diagnosis. 11. The apparatus of claim 10, wherein the visual indication of the diagnosis comprises a color-coded indication. 12. The apparatus of claim 1, further comprising means for transmitting the electrocardiogram signal to a physical location remote from a location of the patient, wherein the means for analyzing the electrocardiogram signal and for deriving a respiratory signal comprise processing means located at the remote physical location. 13. The apparatus of claim 1, wherein only a single channe... | BiancaMed Limited | A61B005/0402
US06994675 | B2 | 2006-02-07 | Non-invasive measurement of suprasystolic signals | 2001-07-19 | An apparatus for assessing cardiovascular status of a mammal comprises a system for locally applying a pressure to an artery capable of restricting blood flow through said artery, a wideband external pulse transducer having an output and situated to measure suprasystolic signals proximate to said artery, and a computing device receiving said output for calculating vascular compliance values. The method described is particularly useful for determining cardiac output, assessing whether a pregnant female has preeclampsia or a patient has cardiac insufficiency, or assessing cardiac arrhythmias. | 1. An apparatus for assessing cardiovascular status of a mammal comprising: a system for locally applying a pressure to an artery, capable of restricting blood flow through said artery;a wideband external pulse transducer having an output and situated to measure suprasystolic signals proximate to said artery; anda computing device receiving said output for assessing large arterial vascular compliance values from said suprasystolic signals. | 1. An apparatus for assessing cardiovascular status of a mammal comprising: a system for locally applying a pressure to an artery, capable of restricting blood flow through said artery;a wideband external pulse transducer having an output and situated to measure suprasystolic signals proximate to said artery; anda computing device receiving said output for assessing large arterial vascular compliance values from said suprasystolic signals. 2. The apparatus of claim 1, wherein said computing device analyzes a temporal amplitude pattern of said output. 3. The apparatus of claim 2, wherein said system for locally applying a pressure to an artery occludes said artery, said temporal amplitude pattern comprises a series of adjacent amplitude peaks and related dips, and the computing device calculates a first ratio of (1) the amplitude difference of a first peak and a subsequent dip and (2) the amplitude difference of a second peak and the subsequent dip as well as a second ratio of (1) the amplitude difference of the first peak and a subsequent dip and (2) the amplitude difference of a third peak and a subsequent dip. 4. The apparatus of claim 3, wherein the computing device calculates said value as a function of the natural logarithm of said first ratio or a power function of said first ratio. 5. The apparatus of claim 1 which assesses whether a pregnant mammal has a preeclampsia condition. 6. The apparatus of claim 1 which assesses whether a pregnant mammal has cardiac insufficiency. 7. The apparatus of claim 1, wherein stroke volume is determined. 8. The apparatus of claim 1, wherein breathing efficacy is determined. 9. The apparatus of claim 1, wherein said computing device calculates a heart rate and value relating to stroke volume from said heart rate and said cardiac output. 10. The apparatus of claim 1, wherein said wideband external pulse transducer comprises a foil electret transducer. 11. The apparatus of claim 1, wherein said wideband external pulse transducer comprises a piezoelectric transducer. 12. The apparatus of claim 11, wherein said wideband external pulse transducer comprises a metalized polyvinylidene fluoride polymer film. 13. The apparatus of claim 1, wherein the large arterial vascular compliance values are aortic compliance values. 14. The apparatus of claim 1 for assessing a physiological factor. 15. A method for assessing the large arterial vascular compliance of a mammal, comprising the steps of: (a) measuring the suprasystolic waveforms of a peripheral artery with blood flow occluded;(b) determining the difference in amplitude between a first suprasystolic peak and a first suprasystolic trough of the suprasystolic waveform;(c) determining the difference in amplitude between a second suprasystolic peak and a second suprasystolic trough of the suprasystolic waveform;(d) determining the ratio of a difference from step (b) to the difference from step (c); and(e) calculating a large arterial vascular compliance va... | - | A61B005/02
US06998842 | B2 | 2006-02-14 | MRI apparatus with low-frequency cable integrated into the patient carrier | 2002-12-18 | In a medical MRI apparatus, it is often desirable to have devices 62, 64 for communicating with or monitoring the patient 58 in the imaging volume 29 of the apparatus. Such devices need DC power or low-frequency connections with an area 70 outside the imaging volume. A device cable that does not interfere with the MRI magnetic fields is known per se. The invention proposes to fix the device cable in a groove 76 in the patient carrier 60 in such a manner that its strip-shaped conductors 78, 80 extend parallel to the field lines of the stationary field B of the apparatus. In this way, it is also possible to clip device carriers 72, 74 to the groove, thus supplying the devices with DC power or establishing an audio connection through the device cable in the groove. | 1. A medical MRI apparatus with a bore provided with a patient table having an elongated patient carrier in the longitudinal direction that coincides with the direction of the stationary magnetic field B in the imaging volume of the MRI apparatus when the patient carrier is placed within the bore of the MRI apparatus, the medical MRI apparatus including: at least one device cable configured for low-frequency electric conduction, which device cable comprises mutually separated segments which are each shorter than a predetermined value, and wherein the segments are separated from each other by frequency-dependent separation elements forming a conductor of low-frequency currents and an insulator of radio-frequency alternating current, the device cable being rigidly connected throughout its length to the patient carrier in the longitudinal direction of the patient carrier, the device cable being configured in order to enable free selective interconnection entirely along the length of the device cable with an associated device also located within the bore of the MRI apparatus. | 1. A medical MRI apparatus with a bore provided with a patient table having an elongated patient carrier in the longitudinal direction that coincides with the direction of the stationary magnetic field B in the imaging volume of the MRI apparatus when the patient carrier is placed within the bore of the MRI apparatus, the medical MRI apparatus including: at least one device cable configured for low-frequency electric conduction, which device cable comprises mutually separated segments which are each shorter than a predetermined value, and wherein the segments are separated from each other by frequency-dependent separation elements forming a conductor of low-frequency currents and an insulator of radio-frequency alternating current, the device cable being rigidly connected throughout its length to the patient carrier in the longitudinal direction of the patient carrier, the device cable being configured in order to enable free selective interconnection entirely along the length of the device cable with an associated device also located within the bore of the MRI apparatus. 2. The medical MRI apparatus as claimed in claim 1, provided with an associated device which is to be arranged in or near the imaging volume, which associated device is provided with a connection cable which is connectable to the device cable, which connection cable comprises mutually separated segments which are each shorter than a predetermined value, and wherein the separation between the segments is the result of frequency-dependent separation elements forming a conductor of low-frequency currents and an insulator of radio-frequency alternating current. 3. The medical MRI apparatus as claimed in claim 2, wherein the associated device is arranged in order to maintain a signal connection to an area outside the imaging volume via a high-frequency carrier wave. 4. A medical MRI apparatus with a bore provided with a patient table having an elongated patient carrier in the longitudinal direction that coincides with the direction of the stationary magnetic field B in the imaging volume when the patient carrier is placed within the bore of the MRI apparatus, the patient carrier including: a groove along the length of the patient carrier's longitudinal directionat least one device cable configured for low-frequency electric conduction, fixedly connected within the groove along the length of the groove, which device cable includes mutually separated segments which are each shorter than a predetermined value, which segments are separated from each other by frequency-dependent separation elements forming a conductor of low-frequency currents and an insulator of radio-frequency alternating current; andan associated device located within the bore of the MRI apparatus, which is configured to be selectively connectable anywhere along the groove of the patient carrier in order to make at least one electrical connection with the segments of the device cable. 5. The medical MRI apparatus a... | Koninklijke Philips Electronics, N.V. | G01V003/00 A61B005/055
US07179230 | B2 | 2007-02-20 | Pulsimeter and measuring system comprising the same | 2001-04-24 | A measuring system includes a measuring apparatus (2) for obtaining pulse data and identification data related to the pulse data, and a central controlling apparatus (1, 3) for receiving the pulse data and the identification data from the measuring apparatus. The central controlling apparatus includes a data processing unit (1) for performing data processing based on the pulse data and the identification data. | 1. A pulsimeter comprising: a waterproof housing;a first panel attached to the housing;a pulse sensor arranged in the housing in facing relationship to the first panel for detecting pulse data;a display for displaying a pulse count per unit time based on the pulse data;a communication unit for performing data communication with an external device; anda controller for transmitting the pulse data to the external device via the communication unit;wherein, in addition to the pulse data, the pulse sensor detects identification data related to the pulse data. | 1. A pulsimeter comprising: a waterproof housing;a first panel attached to the housing;a pulse sensor arranged in the housing in facing relationship to the first panel for detecting pulse data;a display for displaying a pulse count per unit time based on the pulse data;a communication unit for performing data communication with an external device; anda controller for transmitting the pulse data to the external device via the communication unit;wherein, in addition to the pulse data, the pulse sensor detects identification data related to the pulse data. 2. The pulsimeter according to claim 1, wherein the pulse sensor includes a light emitting clement for irradiating a detection object, and a light receiving element for detecting light reflected by the detection object. 3. The pulsimeter according to claim 1, wherein the data communication is performed wirelessly. 4. A pulsimeter comprising: a waterproof housing;a first panel attached to the housing;a pulse sensor arranged in the housing in facing relationship to the first panel for detecting pulse data;a display for displaying a pulse count per unit time based on the pulse data;a communication unit for performing data communication with an external device;a controller for transmitting the pulse data to the external device via the communication unit;an identification data inputting unit provided separately from the pulse sensor for inputting identification data related to the pulse data, the controller transmitting the identification data together with the pulse data to the external device via the communication unit; anda second panel supported by the housing, the identification data inputting unit being arranged in the housing in facing relationship to the second panel. 5. The pulsimeter according to claim 4, wherein the housing is configured like a dome having a top, the first panel being arranged adjacent the top, the second panel being farther from the top than the first panel is. 6. The pulsimeter according to claim 1, further comprising a rechargeable power source provided in the housing, the power source being charged by an external recharging device in a non-contact manner. | Arkray, Inc. | A61B005/02
US07115268 | B1 | 2006-10-03 | Peptides useful for reducing symptoms of toxic shock syndrome and septic shock | 1999-06-18 | This invention relates to compositions and methods which provide protection against, or reduce the severity of toxic shock and septic shock from bacterial infections. More particularly it relates to peptides derived from homologous sequences of the family of staphylococcal and streptococcal toxins, which may be polymeric, and carrier-conjugates thereof. The invention also relates to serum antibodies induced by the peptides and carrier-conjugates and their use to prevent, treat, or protect against the toxic effects of most, if not all, of the staphylococcal and streptococcal toxins.;The invention also relates to diagnostic assays and kits to detect the presence of staphylococcal and streptococcal toxins, or antibodies thereto. The invention also relates isolated and purified to nucleic acids encoding the peptides of the invention and transformed host cells containing those nucleic acids. | 1. A non-toxic, purified peptide consisting of consecutive amino acids having an amino acid sequence set forth in X25X26YGGX1TX2X3X4X5N wherein X1 is V, X2 is L, X3 is H, X4 is E, X5 is G, X25 is C, and X26 is M (SEQ ID NO:34 CMYGGVTLHEGN). | 1. A non-toxic, purified peptide consisting of consecutive amino acids having an amino acid sequence set forth in X25X26YGGX1TX2X3X4X5N wherein X1 is V, X2 is L, X3 is H, X4 is E, X5 is G, X25 is C, and X26 is M (SEQ ID NO:34 CMYGGVTLHEGN). | The Rockefeller University | A61K039/02 A61K039/395 A61K039/40 A61K039/42 A61B005/055
US06990370 | B1 | 2006-01-24 | Method for mapping heart electrophysiology | 2000-04-12 | A mapping catheter is positioned in a heart chamber, and active electrode sites are activated to impose an electric field within the chamber. The blood volume and wall motion modulates the electric field, which is detected by passive electrode sites on the preferred catheter. Electrophysiology measurements, as well as geometry measurements, are taken from the passive electrodes and used to display a map of intrinsic heart activity. | 1. A process for measuring electrophysiologic data in a heart chamber comprising the steps of: positioning a set of passive electrodes within patient's heart; positioning a set of active electrodes within a patient's heart; supplying oscillating current to said set of active electrodes thereby generating an electric field in said heart chamber; detecting said electric field at said passive electrode sites, generating a set of field measurement data; extracting in the time domain, from said field measurement data, that component of said field measurement data corresponding to the underlying electrophysiologic activity of the heart chamber, and generating electrophysiology data; graphically displaying said chamber geometry data; and graphically displaying said electrophysiologic data. | 1. A process for measuring electrophysiologic data in a heart chamber comprising the steps of: positioning a set of passive electrodes within patient's heart; positioning a set of active electrodes within a patient's heart; supplying oscillating current to said set of active electrodes thereby generating an electric field in said heart chamber; detecting said electric field at said passive electrode sites, generating a set of field measurement data; extracting in the time domain, from said field measurement data, that component of said field measurement data corresponding to the underlying electrophysiologic activity of the heart chamber, and generating electrophysiology data; graphically displaying said chamber geometry data; and graphically displaying said electrophysiologic data. 2. A process for determining and displaying the location of a therapy catheter in a heart chamber comprising the steps of: positioning a set of passive electrodes within said heart chamber; positioning a set of active locator electrodes within said heart chamber, said locator electrodes being positioned on a therapy catheter; supplying oscillating current to said set of active locator electrodes thereby generating an electric field in said heart chamber; detecting said electric field at said passive electrode sites, and generating field measurement data; extracting in the frequency domain, from said field measurement data, that component of said field measurement data corresponding to locator electrode location and generating location data; and graphically displaying said location data. | Endocardial Solutions, Inc. | A61B005/04 A61B005/05
US07201725 | B1 | 2007-04-10 | Device and method for determining a depth of an incision | 2000-09-25 | A device for determining the depth of an incision from the epidermal layer to a puncture in a blood vessel in which the device includes an elongated member which has a lumen extending from a distal end to a proximal end. The distal end of the elongated member is configured to locate the blood vessel. After the device locates the blood vessel, a depth indicator is positioned to mark the depth of the blood vessel. In one embodiment, the elongated member includes a tapering interior surface which can receive a portion of the blood vessel. An extending control member may extend from the device at the distal end wherein the extending control member enters the blood vessel before the device locates the vessel. Fluid from the blood vessel enters the extending member through a vent hole and exits through an opening at the proximal end of the device, thereby providing a user with visual feedback that the device is moving in the desired direction toward the blood vessel. | 1. A device for determining a depth of incision that extends from an epidermal layer to a blood vessel having a puncture, the device comprising: an elongated member including a distal end and a proximal end;a lumen extending through the elongated member, the lumen having a proximal end coincident with the elongated member proximal end and a distal end spaced proximally from the elongated member distal end, the lumen accommodating at least one extending member that enters the blood vessel through the puncture, anda tapered surface extending between the elongated member distal end and the lumen distal end, the tapered surface having a larger diameter at the elongated member distal end and a smaller diameter at the lumen distal end, the tapered surface receiving a portion of the blood vessel surrounding the puncture for impeding the distal end of the elongated member from entering the blood vessel. | 1. A device for determining a depth of incision that extends from an epidermal layer to a blood vessel having a puncture, the device comprising: an elongated member including a distal end and a proximal end;a lumen extending through the elongated member, the lumen having a proximal end coincident with the elongated member proximal end and a distal end spaced proximally from the elongated member distal end, the lumen accommodating at least one extending member that enters the blood vessel through the puncture, anda tapered surface extending between the elongated member distal end and the lumen distal end, the tapered surface having a larger diameter at the elongated member distal end and a smaller diameter at the lumen distal end, the tapered surface receiving a portion of the blood vessel surrounding the puncture for impeding the distal end of the elongated member from entering the blood vessel. 2. The device according to claim 1, wherein the elongated member has a constant outer diameter. 3. The device according to claim 1, wherein the elongated member has an outer diameter which progressively decreases to a smaller outer diameter at the distal end. 4. The device according to claim 1, wherein the lumen is centered within the elongated member. 5. The device according to claim 1, wherein the lumen is off-center along the longitudinal axis of the elongated member. 6. The device according to claim 1, wherein the tapered surface has a substantially conic shape. 7. The device according to claim 1, wherein the tapered surface has a substantially concave spherical shape. 8. The device according to claim 1, wherein the tapered surface has a substantially stepped configuration. 9. The device according to claim 1, wherein a diameter of the tapered surface at the distal end of the elongated member is about 50% to 99% of an outer diameter of an exterior surface of the elongated member. 10. The device according to claim 1, wherein the tapered surface at the elongated member distal end has a diameter of about 050 to 160 inches. 11. The device according to claim 1, wherein the elongated member comprises a material with a hardness of at least 50 D. 12. The device according to claim 1, wherein a portion of the elongated member comprises a friction reducing material. 13. The device according to claim 1, wherein the elongated member comprises a friction reducing material. 14. The device according to claim 1, further comprising a depth indicating member positioned on an exterior of the elongated member and movable in an axial direction with respect to the elongated member. 15. The device according to claim 14, wherein the depth indicating member is an elastic ring. 16. The device according to claim 1, where the extending member extends beyond the distal end of the elongated member. 17. The device according to claim 16, wherein the extending member includes a proximal end, a distal end, and a lumen which extends from the proximal end to the distal end... | Sub-Q, Inc. | A61B005/117 A61B005/103
US06955647 | B2 | 2005-10-18 | System and method for repetitive interval clinical evaluations | 2002-12-06 | A healthcare tool allows a patient to record daily parameters associated with the patient's clinical status, for example, body weight for congestive heart failure patients. A graph may be created showing the parameters on a control chart. The parameters are statistically analyzed against a control range, and when a parameter moves out of the control range, the system automatically creates a pop-up window alerting the patient that the parameter is outside the control range, and that the patient should consider informing a healthcare professional. | 1. A computer implemented method of impeding a progression of a disease within a patient comprising the steps of: defining a set of disease associated parameters; performing repetitive measurements on said set of disease associated parameters; evaluating said repetitive measurements of disease associated parameters; performing statistical analysis on a history of said repetitive measurements of disease associated parameters, wherein said statistical analysis comprises a Deming Statistical Method; alerting the patient to those statistical analyses which indicate a potential future problem that requires intervention; and applying secondary prevention techniques to address said potential future problem. | 1. A computer implemented method of impeding a progression of a disease within a patient comprising the steps of: defining a set of disease associated parameters; performing repetitive measurements on said set of disease associated parameters; evaluating said repetitive measurements of disease associated parameters; performing statistical analysis on a history of said repetitive measurements of disease associated parameters, wherein said statistical analysis comprises a Deming Statistical Method; alerting the patient to those statistical analyses which indicate a potential future problem that requires intervention; and applying secondary prevention techniques to address said potential future problem. 2. The method of claim 1, wherein said step of performing repetitive measurements further comprises entering said repetitive measurements on said set of disease associated parameters through a patient interface to the computer. 3. The method of claim 1, wherein said statistical analysis performing step uses a history of said repetitive measurements occurring at a defined period of time preceding the most recent of said performed repetitive measurements. 4. The method of claim 1, further comprising the step of storing in a computer memory said history of said repetitive measurements of disease-associated parameters. 5. The method of claim 1, wherein said steps of evaluating said repetitive measurements, performing statistical analysis on said history of said repetitive measurements, and alerting the patient to statistical analyses which indicate a potential future problem are performed by executing a set of instructions encoded within a computer memory. 6. The method of claim 1, wherein the disease is selected from the group consisting of congestive heart failure, diabetes, asthma, emphysema, cancer, infarction, ischemia, arteriosclerosis, toxicity, mental disease, depression and arrhythmia. 7. The method of claim 6, wherein said set of disease associated chaotic parameters comprise at least one parameter selected from the group consisting of body weight, peak flow, glucose, number of nitroglycerin tablets taken, number of chest pain episodes, minutes walking without pain, ambulation distance without pain, number of emesis, number of episodes of diarrhea, mood charts, depression charts, and food/liquid intake. 8. The method of claim 1, wherein the computer is coupled to a network and communicates with the patient's health care team. 9. The method of claim 1, wherein the computer uses an audio interface to interact with the patient. 10. The method of claim 9, wherein the audio interface uses voice synthesis and recognition over a telephone to automatically gather data from the patient. 11. The method of claim 10, wherein the audio interface connects the patient automatically to a health care provider when said statistical analysis indicate a potential future problem that requires intervention. 12. The method of claim 1, wherein the computer... | - | A61B005/00
US06838552 | B1 | 2005-01-04 | Diagnosis and management of infection caused by Chlamydia | 2000-11-08 | The present invention provides a unique approach for the diagnosis and management of infections by Chlamydia species, particularly C. pneumoniae. The invention is based, in part, upon the discovery that a combination of agents directed toward the various stages of the chlamydial life cycle is effective in substantially reducing infection. Products comprising combination of antichlamydial agents, novel compositions and pharmaceutical packs are also described. | 1. A method of detecting Chlamydia in a sample, said method comprising: (a) providing an antibody that specifically binds to a peptide having a sequence consisting essentially of SEQ ID NO: 97; and (b) performing on a sample an antigen capture assay using said antibody wherein binding of said antibody indicates the presence of Chlamydia in said sample. | 1. A method of detecting Chlamydia in a sample, said method comprising: (a) providing an antibody that specifically binds to a peptide having a sequence consisting essentially of SEQ ID NO: 97; and (b) performing on a sample an antigen capture assay using said antibody wherein binding of said antibody indicates the presence of Chlamydia in said sample. 2. A method for detecting Chlamydia in a sample, said method comprising the steps of: a) providing a sample; b) contacting said sample with a peptide having a sequence consisting essentially of SEQ ID NO: 97; and c) detecting the binding of an antibody in said sample to said peptide, wherein binding of said antibody to said peptide indicates the presence of Chlamydia in said sample. 3. The method of claim 2, wherein said antibody specifically binds to a peptide having a sequence consisting essentially of SEQ ID NO: 97, and the binding of said antibody to said peptide indicates that said Chlamydia is Chlamydia pneumoniae. 4. A method for detecting the presence of Chlamydia in a sample, said method comprising the steps of: a) immobilizing a sample onto a substrate; b) providing an antibody that specifically binds to a peptide having a sequence consisting essentially of SEQ ID NO: 97; c) contacting said immobilized sample with said antibody wherein said antibody becomes immobilized if said sample contains Chlamydia; and d) detecting the presence of immobilized antibody, wherein the presence of immobilized antibody indicates the presence of Chlamydia in said sample. 5. The method of claim 4, wherein said antibody specifically binds to a peptide having a sequence consisting essentially of SEQ ID NO: 97, and the presence of immobilized antibody indicates that said Chlamydia is Chlamydia pneumoniae. | Vanderbilt University | A61K051/00 A61K039/118 A61B005/055 G01N033/55 G01N033/567
US06970729 | B2 | 2005-11-29 | Method and device for determining local distribution of a measuring parameter | 2000-11-30 | A system and method for determining the local distribution of measurement variables being measured relative to or present in a predefined measurement area of a biological sample wherein in a first optical measuring process the local distribution of a first of the measurement variables is determined using a sensor film applied to the measurement area and including a luminescence indicator reacting to the first measurement variable by a change of luminescence decay time. The luminescence decay time or a quantity derived therefrom is recorded by an imaging technique as a function of the first measurement variable. In a second optical measuring process the local distribution of a second of the measurement variables is determined simultaneously or immediately following, using an imaging technique which is effective through the sensor film of the first optical measuring process. | 1. Method for determining by imaging techniques the local distribution of first and second independent measurement variables measured relative to or present in a predefined measurement area of a biological sample comprising the steps of: applying a single sensor film to the predefined measurement area of the biological sample, said single sensor film containing a luminescence indicator which reacts to said first independent measurement variable by a change in luminescence decay time,directing first excitation radiation into said single sensor film to excite said luminescence indicator and cause said luminescence indicator to emit first measurement radiation,detecting said first measurement radiation and determining therefrom local distribution of said first independent measurement variable in a first optical measuring process,directing second excitation radiation into and through said single sensor film and into said biological sample to cause second measurement radiation to be emitted from said biological sample, anddetecting said second measurement radiation and determining therefrom local distribution of said second independent measurement variable in a second optical measuring process. | 1. Method for determining by imaging techniques the local distribution of first and second independent measurement variables measured relative to or present in a predefined measurement area of a biological sample comprising the steps of: applying a single sensor film to the predefined measurement area of the biological sample, said single sensor film containing a luminescence indicator which reacts to said first independent measurement variable by a change in luminescence decay time,directing first excitation radiation into said single sensor film to excite said luminescence indicator and cause said luminescence indicator to emit first measurement radiation,detecting said first measurement radiation and determining therefrom local distribution of said first independent measurement variable in a first optical measuring process,directing second excitation radiation into and through said single sensor film and into said biological sample to cause second measurement radiation to be emitted from said biological sample, anddetecting said second measurement radiation and determining therefrom local distribution of said second independent measurement variable in a second optical measuring process. 2. Method according to claim 1, wherein said biological sample is the epidermis or the surface of an organ. 3. Method according to claim 1, wherein said first measurement variable to be determined is one of the group consisting of the local distribution of oxygen concentration (tcPO2), the local distribution of transcutaneous oxygen concentration and the local distribution of oxygen flow (O2-flux) through an organ surface. 4. Method according to claim 3, wherein said organ surface is skin. 5. Method according to claim 1, wherein said first measurement variable determined is the local distribution of CO2 concentration or the local distribution of CO2 flow. 6. Method according to claim 1, wherein said first measurement variable determined is the local distribution of one of the group consisting of temperature, glucose concentration and an ionic concentration. 7. Method according to claim 6, wherein said ionic concentration is the pH level. 8. Method according to claim 1, wherein said second measuring process is a Laser Doppler imaging (LDI) process being used to determine a perfusion rate in said measurement area as said second measurement variable. 9. Method according to claim 1, wherein said second measuring process is a photometric or photographic method and said measurement area is recorded as an image in a predefined range of wavelengths. 10. Method according to claim 9, wherein autofluorescence in said measurement area is recorded as said image. 11. Method according to claim 9, wherein infrared radiation in said measurement area is recorded as said image. 12. Method according to claim 1, wherein said second measuring process is a profilometric method. 13. Method according to claim 12, wherein said profilometric method is an interferometric metho... | Perimed AB | A61B005/00
US06865411 | B2 | 2005-03-08 | Patient positioning system for radiotherapy | 2001-09-07 | The invention concerns a patient positioning system for radiotherapy, comprising a device for positioning a patient within the framework of treatment planning, using image detection and/or carrying out radiation treatment, comprising a patient bed (3, 42, 50) on which a patient is positioned, wherein means are provided by which the same bed (3, 42, 50) may be arranged both on a planning couch (8) for image detection within the framework of treatment planning as well as on a radiotherapy table (1, 44), and by providing a reference means (54) on the patient bed by which the position of the immobilized patient or of a marked radiation target is determined with respect to the patient bed (3, 42, 50). Within this framework, a corresponding method is also provided, as well as an adjusting means for the transport and the bed. | 1. A device for positioning a patient for treatment planning and radiation treatment comprising: a patient bed on which a patient can be positioned, a reference device secured to the patient bed that can be tracked by a tracking system, wherein the patient bed is configured to be supported by both a planning couch for image-generation and by a radiotherapy table for radiation exposure, and whereby the position of the patient and/or of marked points on the surface of the patient's body can be determined with respect to the patient bed. | 1. A device for positioning a patient for treatment planning and radiation treatment comprising: a patient bed on which a patient can be positioned, a reference device secured to the patient bed that can be tracked by a tracking system, wherein the patient bed is configured to be supported by both a planning couch for image-generation and by a radiotherapy table for radiation exposure, and whereby the position of the patient and/or of marked points on the surface of the patient's body can be determined with respect to the patient bed. 2. The device as set forth in claim 1, wherein the reference device includes a reference star including an arrangement of markers, and further comprising a tracking system which can detect both the positions of markings arranged on the patient as well as the markers on the reference star. 3. The device as set forth in claim 1, further comprising at least one immobilization device provided to secure the patient to the patient bed. 4. The device as set forth in claim 1, further comprising, an adjusting device configured to be positioned between the patient bed and the planning couch or treatment table such that the patient bed may be rotated about at least two axes. 5. The device as set forth in claim 4, wherein the adjusting device is configured to turn independently about at least one axis laterally and/or one axis longitudinally to the bed. 6. The device as set forth in claim 4, wherein the adjusting device consists of a base plated and a cover plate which are movably linked together via a three-point or four-point bearing. 7. The device as set forth in claim 6, wherein the bearing of the cover plate on the base plate includes a joint with two translatory and two rotatory degrees of freedom, that prevents the cover plate from shifting laterally with respect to the base plate. 8. The device as set forth in claim 4, wherein the adjusting device comprises at least two height adjustable, pneumatic, hydraulic, piezoelectric or electromechanical supports that are adjustable directly or by means of a lever system. 9. The device as set forth in claim 8, wherein the adjusting device is configured to cause the bed to rotate about its longitudinal axis by purposefully setting the adjusting speed of at least two supports, whereby the position of the apparent axis of rotation can be placed in a desired position. 10. The device as set forth in claim 4, wherein the adjusting device comprises a front, head-end support and a rear, foot-end support, wherein at least one of the supports is constructed as a ball joint or as a universal joint. 11. The device as set forth in claim 10, wherein the front or the rear support comprise toggle levers with ball joints on both sides. 12. The device as set forth in claim 10, wherein the adjusting device is configured to cause the bed to rotate about its longitudinal axis by counter-adjusting two head-end or foot-end supports. 13. The device as set forth in claim 10, wherein the adju... | BrainLAB AG | A61B005/05
US06873867 | B2 | 2005-03-29 | Referencing or registering a patient or a patient body part in a medical navigation system by means of irradiation of light points | 2001-04-05 | A method for referencing a body part in a camera-assisted, medical navigation system including the following steps: manually manipulating a light beamer to sequentially produce a plurality of light marks on a surface of the body part; using a plurality of cameras to scan the surface of the body part, wherein the plurality of cameras detect the plurality of light marks; determining three-dimensional spatial positions for respective sequential light marks; and referencing or registering the body part based on the three dimensional spatial positions of light marks. | 1. A method for referencing or registering a body part in a camera-assisted, medical navigation system comprising the following steps: manually manipulating a light beamer so as to scan a light beam produced thereby across a surface of the body part to sequentially produce a plurality of light marks on the surface of the body part; using a plurality of cameras to detect the light marks on the surface of the body part; determining three-dimensional spatial positions for respective light marks; and referencing or registering the body part based on the three dimensional spatial positions of light marks. | 1. A method for referencing or registering a body part in a camera-assisted, medical navigation system comprising the following steps: manually manipulating a light beamer so as to scan a light beam produced thereby across a surface of the body part to sequentially produce a plurality of light marks on the surface of the body part; using a plurality of cameras to detect the light marks on the surface of the body part; determining three-dimensional spatial positions for respective light marks; and referencing or registering the body part based on the three dimensional spatial positions of light marks. 2. The method as set forth in claim 1, wherein the step of referencing or registering the body part includes assigning a set of image data, previously produced by an imaging technique for said body part, to the referenced or registered body part. 3. The method as set forth in claim 1, wherein the plurality of light marks are produced by a beam of invisible light. 4.The method as set forth in claim 3, wherein said light beam is a laser light beam. 5. The method as set forth in claim 3, wherein a visible light reflection is created on said surface by means of a second beam of visible light, aimed substantially at the same target area as that of said invisible light beam. 6. The method as set forth in claim 5, wherein said second light beam is a visible laser beam. 7. The method as set forth in claim 6, wherein said light beams are generated by two juxtaposed or nested light sources. 8. The method as set forth in claim 1, wherein the position of the light marks are continuously detected and determined until a sufficient number of positional data is obtained for registering the body part. 9. The method as set forth in claim 2, wherein the image data includes at least one of CT, MRI, PET, SPECT, X-ray and ultrasound scan data. 10. The method as set forth in claim 3, wherein the invisible light is infrared light. 11. A method for referencing or registering a patient or a patient body part in a camera-assisted, medical navigation system comprising the following steps: the patient body part to be referenced is brought into the detecting range of a navigation system assisted by at least two cameras, this navigation system detecting with computer support the three-dimensional, spatial positions of light marks, light marks are generated on the surface of the part of the body to be referenced by means of a light beam, the three-dimensional position of the light marks being determined by the camera-assisted navigation system, the three-dimensional position of the surface of the part of the body to be referenced is determined by means of the positional data for the light marks, wherein either the camera arrangement or said body part to be referenced is moved during referencing so that camera shades are eliminated, a relative movement of said body part being tracked in said navigation system by means of a marker array fixedly positioned relative to said b... | BrainLAB AG | A61B005/05 A61B019/00
US06974424 | B2 | 2005-12-13 | Palatometer and nasometer apparatus | 2001-09-19 | A palatometer and nasometer apparatus is disclosed for the diagnosis and treatment of speech impairment, new language sound learning, and other uses. The palatometer includes a palatal body containing electrodes which fits in a user's mouth. Contact between the user's tongue or lips and the electrodes is indicated on display equipment. The palatometer and display equipment operate to indicate the position and movement of the tongue and lips. The nasometer includes interchangeable sound separator plates to fit users with different facial curvatures. Microphones are attached to the top and bottom of the sound separator plate to measure sound emitted from the nose and mouth for determining the nasality of speech. The palatometer and nasometer are combined to optimize the ability to define speech normality as well as diagnose and remediate speech. | 1. A fonometer comprising: palatometer means for indicating the position and movement of a tongue and lips;nasometer means for determining the nasality of speech; andmeans for processing output from the palatometer and nasometer for display. | 1. A fonometer comprising: palatometer means for indicating the position and movement of a tongue and lips;nasometer means for determining the nasality of speech; andmeans for processing output from the palatometer and nasometer for display. 2. The fonometer of claim 1 wherein the palatometer means comprises a palatal body configured and dimensioned to conform generally to the palate of a person, said palatal body having a lingual surface, and a plurality of electrodes disposed on the lingual surface. 3. The fonometer of claim 2 wherein the electrodes are disposed on the lingual surface in a plurality of rows and columns, wherein said rows and columns are perpendicular to each other. 4. The fonometer of claim 3 wherein the electrodes are equidistantly spaced apart. 5. The fonometer of claim 4 wherein a spacing between the electrodes is 3.5 millimeters. 6. The fonometer of claim 2 wherein said palatal body comprises a flexible printed circuit. 7. The fonometer of claim 6 wherein the flexible printed circuit comprises a plurality of intercoupled lobes. 8. The fonometer of claim 7 wherein the plurality of intercoupled lobes comprises three lobes of substantially equal surface area. 9. The fonometer of claim 7 wherein the plurality of intercoupled lobes comprises a spatially separate center lobe intercoupled to the other lobes of the flexible printed circuit through a neck portion, a width of the neck portion being less then an average length and width of the center lobe. 10. The fonometer of claim 6 wherein the flexible printed circuit further comprises two pairs of spatially separate lead lines, each pair residing on opposite sides of the flexible printed circuit. 11. The fonometer of claim 2 wherein the palatal body is a one-piece member of unitary construction having electrodes disposed therein. 12. The fonometer of claim 2 wherein the palatal body is made of thermoformable material. 13. The fonometer of claim 6 wherein the flexible printed circuit comprises a butterfly shape. 14. The fonometer of claim 6 wherein the flexible printed circuit comprises a labial sensor. 15. The fonometer of claim 6 wherein the flexible printed circuit defines a flat orientation prior to installation on a base plate and defines an orientation conforming to the base plate when disposed on the base plate. 16. The fonometer of claim 1 wherein said palatometer means comprises means for transmitting electrical current through electrodes responsive to contact of the person's tongue or lips with said electrodes. 17. The fonometer of claim 2 wherein the palatal body comprises a base plate and a flexible printed circuit, said base plate being dimensioned to covers the person's teeth. 18. The fonometer of claim 1 wherein said nasometer means comprises a sound separator plate to separate sounds emitted from a person's nose and mouth, said separator plate being removably connected to the nasometer means such that one of a plurality of sound separator plates may... | Logometrix Corporation | A61B005/103 A61B005/117
US06930481 | B2 | 2005-08-16 | Three dimensional MRI RF coil unit capable of parallel imaging | 2002-09-05 | A magnetic resonance imaging apparatus acquires magnetic resonance signals by the PI method using an RF coil unit having basic coils serving as surface coils which are arrayed with at least two coils along a static magnetic field direction (z direction) and at least two coils along each of two orthogonal x, y directions. The coils are divided into an upper unit and a lower unit. The upper unit and lower unit are fixed by a band or the like to allow them to be mounted on an object to be examined. The signals detected by the respective surface coils are sent to a data processing system through independent receiver units and formed into a magnetic resonance image. | 1. An RF coil unit used in a magnetic resonance imaging apparatus which receives magnetic resonance signals generated from an object to be examined which is placed in a static magnetic field and generates a magnetic resonance image, said RF coil unit comprising: a plurality of surface coils which are arrayed about a common imaged volume in three orthogonal directions, x, y, and z, there being at least two of said coils that are spatially distinct from each other being disposed along the direction z of the static magnetic field; at least two of said coils that are spatially distinct from each other disposed along the orthogonal direction x; and at least two of said coils that are spatially distinct from each other being disposed along the orthogonal direction y, wherein each one of said coils receives magnetic resonance signals in use, with the coils in at least one of said directions being disposed on opposite sides of the common image volume, whereby the RF coil unit is capable of being used for performing parallel imaging (PI) in all three orthogonal directions, x, y, and z. | 1. An RF coil unit used in a magnetic resonance imaging apparatus which receives magnetic resonance signals generated from an object to be examined which is placed in a static magnetic field and generates a magnetic resonance image, said RF coil unit comprising: a plurality of surface coils which are arrayed about a common imaged volume in three orthogonal directions, x, y, and z, there being at least two of said coils that are spatially distinct from each other being disposed along the direction z of the static magnetic field; at least two of said coils that are spatially distinct from each other disposed along the orthogonal direction x; and at least two of said coils that are spatially distinct from each other being disposed along the orthogonal direction y, wherein each one of said coils receives magnetic resonance signals in use, with the coils in at least one of said directions being disposed on opposite sides of the common image volume, whereby the RF coil unit is capable of being used for performing parallel imaging (PI) in all three orthogonal directions, x, y, and z. 2. An RF coil unit according to claim 1, wherein said plurality of surface coils are arranged as first and second coil groups, each group having plural coils arranged in the form of a matrix, and said first and second coil groups are arranged to oppose each other through the object in at least one of the directions. 3. An RF coil unit according to claim 1, further comprising a fixing device that fixes said plurality of surface coils with respect to the object. 4. An RF coil unit according to claim 1, further comprising an alignment portion configured to serve as a reference for alignment of respectively corresponding surface coils in one group with respect to another group. 5. An RF coil unit according to claim 1, further comprising a shape maintaining portion configured to maintain at least some of said plurality of surface coils in a predetermined shape. 6. An RF coil unit according to claim 1, wherein at least one of said plurality of surface coils is a QD coil which detects substantially orthogonal magnetic resonance signals. 7. An RF coil unit according to claim 6, wherein said QD coil includes a looped coil element and an 8-shaped coil element. 8. An RF coil unit according to claim 6, wherein said QD coil has a plurality of 8-shaped coil elements which are overlapped so as to be orthogonal to each other. 9. An RF coil unit according to claim 1, wherein each of said surface coils includes a looped coil element and an 8-shaped coil element which detect substantially orthogonal magnetic resonance signals, said looped coil element being partly overlapped on another looped coil element of an adjacent surface coil in at least one of three directions, and said 8-shaped coil element being partly overlapped on another 8-shaped coil element of an adjacent surface coil in at least one of the three directions. 10. An RF coil unit according to claim 1, wherein each of ... | Kabushiki Kaisha Toshiba | G01V003/00 A61B005/055
US06849607 | B2 | 2005-02-01 | Galectin recognized photosensitizers for photodynamic therapy | 2002-05-07 | Purpurin-carbohydrate conjugates and their method of preparation and use for treatment of cancer cells. The conjugates have the general formula: wherein R6 and R7 taken together are ═NR11 or are independently —OR11, where at least one R11 is preferably a mono or polysaccharide moiety and R1-R8 are various groups formed from carbon and hydrogen and optionally oxygen and nitrogen where R3 and R4 may together from a covalent bond. | 1. A compound of the formula: where R1 is lower alkyl, vinyl, alkyl ether, aryl, lower carboxy, or —CH(OR9)CH3 where R9 is alkyl of 1 to about 20 carbon atoms, a cyclic containing substituent containing 1 to about 20 carbon atoms connected to the a ring through a carbon-carbon or —O— bond; R2 and R5 are independently hydrogen or lower alkyl, R3 and R4 are independently —H, lower alkyl, or —OR10, where R10 is H or lower alkyl, or R3 and R4, together form a covalent bond; R6 and R7 taken together are ═NR11 or are independently —OR11 where R11 is independently —H, alkyl of 1 to about 20 carbon atoms, a cyclic containing substituent containing 1 to about 20 carbon atoms, an amide group or a mono or polysaccharide connected to ═N— or —O— through at least one of a saturated or unsaturated lower alkylene group, a saturated or unsaturated heterocylic or hydrocarbon five or six membered ring, an ether linkage, an amide linkage or an ester group; R8 is —CH2CH2COR12 where R12 is an amino acid residue, —NHR13 or —OR14 where R13 is hydrogen, alkyl of 1 to about 20 carbon atoms, or a mono or polysaccharide connected to —NH— through at least one of a saturated or unsaturated lower alkylene group, a saturated or, unsaturated heterocylic or hydrocarbon five or six membered ring, an ether linkage, an amide linkage or an ester group and R14 is hydrogen, or alkyl of 1 to about 20 carbon atoms, where lower alkyl includes alkyl and alkylene groups of 1 to 5 carbon atoms and alkyl includes linear, branched and cyclic unsubstituted alkyl and linear, branched and cyclic alkyl and alkylene groups substituted with hydroxy, carboxy, alkyl, vinyl, amino, amido, keto, heterocyclic, mono and polysaccharide and amino acid groups; provided that the compound contains at least one mono or polysaccharide group that will combine with galectin-1 and M is a chelated metal or is two hydrogens bound to the unsaturated nitrogens in the a and c rings. | 1. A compound of the formula: where R1 is lower alkyl, vinyl, alkyl ether, aryl, lower carboxy, or —CH(OR9)CH3 where R9 is alkyl of 1 to about 20 carbon atoms, a cyclic containing substituent containing 1 to about 20 carbon atoms connected to the a ring through a carbon-carbon or —O— bond; R2 and R5 are independently hydrogen or lower alkyl, R3 and R4 are independently —H, lower alkyl, or —OR10, where R10 is H or lower alkyl, or R3 and R4, together form a covalent bond; R6 and R7 taken together are ═NR11 or are independently —OR11 where R11 is independently —H, alkyl of 1 to about 20 carbon atoms, a cyclic containing substituent containing 1 to about 20 carbon atoms, an amide group or a mono or polysaccharide connected to ═N— or —O— through at least one of a saturated or unsaturated lower alkylene group, a saturated or unsaturated heterocylic or hydrocarbon five or six membered ring, an ether linkage, an amide linkage or an ester group; R8 is —CH2CH2COR12 where R12 is an amino acid residue, —NHR13 or —OR14 where R13 is hydrogen, alkyl of 1 to about 20 carbon atoms, or a mono or polysaccharide connected to —NH— through at least one of a saturated or unsaturated lower alkylene group, a saturated or, unsaturated heterocylic or hydrocarbon five or six membered ring, an ether linkage, an amide linkage or an ester group and R14 is hydrogen, or alkyl of 1 to about 20 carbon atoms, where lower alkyl includes alkyl and alkylene groups of 1 to 5 carbon atoms and alkyl includes linear, branched and cyclic unsubstituted alkyl and linear, branched and cyclic alkyl and alkylene groups substituted with hydroxy, carboxy, alkyl, vinyl, amino, amido, keto, heterocyclic, mono and polysaccharide and amino acid groups; provided that the compound contains at least one mono or polysaccharide group that will combine with galectin-1 and M is a chelated metal or is two hydrogens bound to the unsaturated nitrogens in the a and c rings. 2. The compound of claim 1 wherein R6 and R7 are ═NR11, where R11 is a mono or polysaccharide connected to ═N— through at least one of a saturated or unsaturated lower alkylene group, a saturated or unsaturated heterocylic or hydrocarbon five or six membered ring, an ether linkage, an amide linkage or an ester group. 3. The compound of claim 2 where the saccharide is selected from the group consisting of lactose and galactose. 4. The compound of claim 2 where R11 is: 5. The compound of claim 2 wherein R11 is: 6. The compound of claim 2 wherein R11 is: 7. The compound of claim 2 wherein R11 is: 8. The compound of claim 2 where R11 is —CH2—R15—O—R16 where R15 is a lower alkylene group and R16 is lower alkyl, alkylether, or a saccharide. 9. The compound of claim 1 where R3 and R4 taken together form a covalent bond. 10. The compound of claim 2 where R3 and R4 taken together for a covalent bond. 11. The compound of claim 1 where R1 and R5 are ethyl. 12. The compound of claim 2 where R1 and R5 are ethyl. 13. The compound of c... | Health Research, Inc. | A01N043/04 C07H015/00 A61B005/055
US06865419 | B2 | 2005-03-08 | Method and apparatus for measurement of mean pulmonary artery pressure from a ventricle in an ambulatory monitor | 2001-11-30 | A system and method for determining mean pulmonary arterial pressure (MPAP) using a pressure sensor located within a ventricle of a heart, and a signal indicative of cardiac electrical activity such as an electrocardiogram (EGM) signal. The pressure may be sensed within the right and/or left ventricle using an implanted pressure sensor. The sensed pressure may be used to determine the Ventricular Systolic Pressure (VSP) and an estimated Pulmonary Arterial Diastolic pressure (ePAD). The VSP, ePAD, and time intervals associated with systole and diastole may then be used to obtain an MPAP that closely approximates mean pulmonary arterial pressure measured using a sensor located in the pulmonary artery. | 1. A system for determining mean pulmonary arterial pressure of a patient, comprising: a first sensor adapted to be located in a ventricle of a heart to measure ventricular pressure; a first circuit to measure electrocardiogram (EGM) signals; and a processing circuit coupled to receive signals indicative of the ventricular pressure and the EGM signals, and to determine there from, mean pulmonary arterial pressure (MPAP). | 1. A system for determining mean pulmonary arterial pressure of a patient, comprising: a first sensor adapted to be located in a ventricle of a heart to measure ventricular pressure; a first circuit to measure electrocardiogram (EGM) signals; and a processing circuit coupled to receive signals indicative of the ventricular pressure and the EGM signals, and to determine there from, mean pulmonary arterial pressure (MPAP). 2. The system of claim 1, wherein the first circuit includes at least one electrode adapted to be located within the cardiac vascular system of the patient. 3. The system of claim 1, wherein the first circuit includes at least two electrodes adapted to be placed on an external surface of the patient. 4. The system of claim 1, wherein the first circuit is located within an implantable device, and wherein the first circuit includes at least one electrode adapted to be positioned adjacent to a housing of the implantable device. 5. The system of claim 1, wherein the first sensor is adapted to be located within a first ventricle of the heart, wherein the system includes a second sensor adapted to be located within the other ventricle of the heart, and wherein the processing circuit includes means to estimate the MPAP from pressure measured by both the first and second sensors. 6. The system of claim 1, wherein the processing circuit is located within an implantable device. 7. The system of claim 1, wherein the processing circuit is adapted to be located in a device external to the patient, and wherein the system further includes a communication circuit to transfer indications of the measured pressure and the EGM signals to the processing circuit. 8. The system of claim 1, wherein the processing circuit includes first and second portions, wherein the first portion is located within an implantable device, wherein the second portion is located within a device adapted to be external to the patient, and wherein the system further includes a communication circuit to transfer data signals between the first and second portions. 9. The system of claim 1, and further including a therapy delivery circuit coupled to the processing circuit to provide therapy to the patient. 10. The system of claim 1, wherein the processing circuit includes means for controlling the therapy delivery circuit based on the estimated MPAP. 11. The system of claim 10, wherein the therapy delivery circuit includes a circuit to provide cardiac resynchronization therapy to the patient. 12. The system of claim 10, wherein the therapy delivery circuit includes a drug delivery device to deliver a biologically-active agent to the patient. 13. A method of determining mean pulmonary arterial pressure (MPAP), comprising: a.) sensing pressure within a ventricle of a heart; b.) sensing an electrocardiogram (EGM) signal of the heart; and c.) using the sensed pressure and the EGM signal to derive the MPAP. 14. The method of claim 13, wherein step c.) includes deriving ... | Medtronic, Inc. | A61B005/0402 A61N001/365
US06862471 | B1 | 2005-03-01 | Method and apparatus for blanking T-waves from combipolar atrial cardiac signals based on expected T-wave locations | 2001-10-25 | The stimulation device blanks T-waves from the atrial channel of an electrical cardiac signal by employing a T-wave blanking interval localized to the expected location and duration of the T-wave. To this end, the stimulation device determines the average interval between an R-wave and a T-wave in the patient in which the device is implanted and also determines the average duration of a T-waves within the patient. A T-wave blanking interval is initiated following the average R-T interval subsequent to detection of an R-wave and lasts for a period of time equal to the average T-wave duration. In this manner, highly localized T-wave blanking is achieved permitting P-waves or other atrial signals to be detected during remaining non-blanked portions of the atrial channel of the cardiac signal at least for the purposes of atrial rate detection. The relatively short T-wave blanking interval of the invention is particularly well suited for use in combipolar sensing systems. Method and apparatus implementations are described. | 1. A method performed by an implantable cardiac stimulation device for analyzing a cardiac signal to generate information representative of the characteristics of R-waves and T-waves found therein, the method comprising: sensing a cardiac signal; identifying pairs of consecutive R-waves and T-waves within the cardiac signal; measuring values representative of characteristics of pairs of R-waves and T-waves; generating statistical information representative of the measured values, the statistical information including an average of each measured value; storing the statistical information generated for the measured values; sensing additional cardiac signals; identifying R-waves in the additional cardiac signals and then applying the stored averaged values to identify expected locations and durations of T-waves within the additional cardiac signals; blanking portions of an atrial channel of the additional cardiac signals to ignore signals occurring within a period of time corresponding to the expected locations and durations of T-waves; identifying an additional pair of consecutive R-waves and T-waves within the additional cardiac signals; measuring values representative of characteristics of the additional pair of R-waves and T-waves; determining an amount of variation between the measured values of the additional pair of R-waves and T-waves and the average of the measured values of previously identified pairs; and determining whether the amount of variation exceeds a predetermined threshold of variation and, if not, updating the statistics to reflect the measured values of the additional pair of R-waves and T-waves. | 1. A method performed by an implantable cardiac stimulation device for analyzing a cardiac signal to generate information representative of the characteristics of R-waves and T-waves found therein, the method comprising: sensing a cardiac signal; identifying pairs of consecutive R-waves and T-waves within the cardiac signal; measuring values representative of characteristics of pairs of R-waves and T-waves; generating statistical information representative of the measured values, the statistical information including an average of each measured value; storing the statistical information generated for the measured values; sensing additional cardiac signals; identifying R-waves in the additional cardiac signals and then applying the stored averaged values to identify expected locations and durations of T-waves within the additional cardiac signals; blanking portions of an atrial channel of the additional cardiac signals to ignore signals occurring within a period of time corresponding to the expected locations and durations of T-waves; identifying an additional pair of consecutive R-waves and T-waves within the additional cardiac signals; measuring values representative of characteristics of the additional pair of R-waves and T-waves; determining an amount of variation between the measured values of the additional pair of R-waves and T-waves and the average of the measured values of previously identified pairs; and determining whether the amount of variation exceeds a predetermined threshold of variation and, if not, updating the statistics to reflect the measured values of the additional pair of R-waves and T-waves. 2. The method of claim 1 wherein the amount of variation includes one or more of variation in an amplitude of the T-waves, variation in an amplitude of the R-waves, variation in a time delay between R-waves and corresponding T-waves, variation in a duration of individual R-waves, and variation in a duration of individual T-waves. | Pacesetter, Inc. | A61B005/0452
US07024370 | B2 | 2006-04-04 | Methods and apparatus for early detection of health-related events in a population | 2002-03-26 | A method is disclosed for providing early detection, classification, and reporting of health-related events in a population. The method includes capturing sets of specific emergency room patient information from a subset of the population as the patient information is first electronically entered into, for example, an electronic medical record (EMR). The patient information is pre-processed, transmitted to and stored in a central database in a central computer facility. The patient information is sorted and analyzed by the central computer facility to detect any health-related events in the population and to generate corresponding alerts. The alerts are electronically reported to designated authorities such as health officials and other government authorities such as the CDC. | 1. A method for providing early detection and reporting of health-related events in a population, said method comprising: automatically capturing patient information from a plurality of members of said population, as said patient information is first electronically entered for electronic recording in a plurality of emergency room environments, to generate patient database information at a central collecting location, wherein said capturing comprises data-mining for said patient information in a plurality of emergency medical record (EMR) systems at a plurality of respective emergency room sites by continuously accessing patient information recorded in said EMR systems through an electronic interface and extracting data corresponding to said patient information;performing analyses on said patient database information of said plurality of members to generate a health-related alert; andreporting said health-related alert to designated authorities. | 1. A method for providing early detection and reporting of health-related events in a population, said method comprising: automatically capturing patient information from a plurality of members of said population, as said patient information is first electronically entered for electronic recording in a plurality of emergency room environments, to generate patient database information at a central collecting location, wherein said capturing comprises data-mining for said patient information in a plurality of emergency medical record (EMR) systems at a plurality of respective emergency room sites by continuously accessing patient information recorded in said EMR systems through an electronic interface and extracting data corresponding to said patient information;performing analyses on said patient database information of said plurality of members to generate a health-related alert; andreporting said health-related alert to designated authorities. 2. The method of claim 1 wherein said patient information includes: emergency room patient triage information;patient contact information; andinitial patient interview information. 3. The method of claim 2 wherein said emergency room patient triage information comprises: instances of basic patient demographic information; andinstances of basic patient clinical information. 4. The method of claim 3 wherein said instances of basic patient demographic information comprises: patient age;patient sex; andpatient zip code. 5. The method of claim 3 wherein said instances of basic patient clinical information comprise at least one of: chief complaint including any of cough, fever, shortness of breath, weakness, malaise, bleeding, rash, and pain;vital signs including any of body temperature, tachycardia, hypotension, blood pressure, and heart rate; andsmart bio-detection chip results. 6. The method of claim 1 wherein said capturing is accomplished electronically over a network from a plurality of emergency rooms. 7. The method of claim 1 wherein said capturing comprises: sending said patient information, from each of said plurality of emergency room sites, over the Internet as at least one packet of information;receiving said at least one packet of information, for each of said plurality of emergency room sites, in a web server on a central collecting computer; andstoring said patient information on said central collecting computer as said patient database information. 8. The method of claim 7 further comprising: verifying said patient information for completeness before said sending;stripping said patient information of certain personal patient data, including patient name, before said sending to make said patient information anonymous;encrypting said patient information before said sending, to make said patient information secure;de-encrypting said patient information after said receiving; andverifying said patient information after said de-encrypting. 9. The method of claim 7 further comprising integratin... | P) CIS, Inc. | G06F159/00 A61B005/00 A61B010/00
US06987998 | B2 | 2006-01-17 | Cardiac rhythm management patient report | 2001-02-28 | A patient oriented report from an implantable medical device such as a cardiac rhythm management device, e.g. a cardiac pacemaker or cardioverter/defibrillator, and a system and method for producing the same. The report provides the patient with various information regarding the treatment, type of device, manufacturer, physician, device settings, or device status. | 1. A cardiac patient report generating system, comprising: a cardiac rhythm management device;a programming unit communicating with said cardiac rhythm management device; andan output device connected to said programming unit, said output device producing a report for a patient, wherein said report directly informs a patient about said cardiac rhythm management device in terminology easy for a non-medically trained patient to interpret. | 1. A cardiac patient report generating system, comprising: a cardiac rhythm management device;a programming unit communicating with said cardiac rhythm management device; andan output device connected to said programming unit, said output device producing a report for a patient, wherein said report directly informs a patient about said cardiac rhythm management device in terminology easy for a non-medically trained patient to interpret. 2. The system of claim 1, wherein said output device is a printer and said report is a printed report for a patient. 3. The system of claim 1, wherein said output device includes a write device and the report for a patient is stored on a removable storage media. 4. The system of claim 3, wherein said write device is an electromagnetic read/write device and the report is written electromagnetically to a removable electromagnetic storage media which is given to a patient. 5. The system of claim 3, wherein said write device is an electro-optical read/write device and the report is written electro-optically to a removable electro-optically storage media which is given to a patient. 6. The system of claim 1, wherein said report includes at least one of patient's name, physician's name, cardiac rhythm management device manufacturer, cardiac rhythm management device model number, and cardiac rhythm management device serial number. 7. The system of claim 6, wherein said cardiac rhythm management device manufacturer, cardiac rhythm management device model number, and cardiac rhythm management device serial number are stored in said cardiac rhythm management device and read by said programming unit. 8. The system of claim 7, wherein said programming unit formats said cardiac rhythm management device manufacturer, cardiac rhythm management device model number, and cardiac rhythm management device serial number into a data stream, and transmits same to said. output device. 9. The system of claim 1, wherein said report includes at least one of a physician's name, physician's contact information and a date of a patient's next visit to the physician. 10. The system of claim 1, wherein said cardiac rhythm management device includes leads and said report includes at least one of lead manufacturer, lead model number, and lead serial number. 11. The system of claim 1, wherein said cardiac rhythm management device is implantable in a patient and is selected from a group consisting of a cardiac pacemaker, a cardioverter/defibrillator, and a combination pacemaker/defibrillator. 12. The system of claim 1, wherein said cardiac rhythm management device is implanted in a patient and said report includes a reason for implantation. 13. The system of claim 12, wherein the reason for implantation is on said report in brief medical terminology and in plain English terminology for a non-medically trained patient's understanding. 14. The system according to claim 1, wherein said cardiac rhythm management device is a pacing device and... | Cardiac Pacemakers, Inc. | A61B005/04
US06987448 | B2 | 2006-01-17 | Medical gas alarm system | 2001-08-20 | A medical gas alarm system for use in a healthcare facility having a medical gas system and having a network of computer devices is provided. The alarm system includes at least one area alarm controller adapted to receive a first signal indicative of a condition of a first portion of the medical gas system. The area alarm controller is adapted to communicate with the network. The alarm system also includes at least one master alarm controller adapted to receive a second signal indicative of a condition of a second portion of the medical gas system. The master alarm controller is adapted to communicate with the network. The area alarm controller is adapted to communicate with the master alarm controller through the network. | 1. A medical gas alarm system for use in a healthcare facility having a medical gas system which delivers a plurality of medical gases to a plurality of locations in the healthcare facility and having a network of computer devices, the medical gas alarm system comprising at least one area alarm controller adapted to receive a first signal indicative of a condition of a first portion of the medical gas system and adapted to communicate with the network, and at least one master alarm controller adapted to receive a second signal indicative of a condition of a second portion of the medical gas system and adapted to communicate with the network, the at least one area alarm controller being adapted to communicate with the at least one master alarm controller through the network. | 1. A medical gas alarm system for use in a healthcare facility having a medical gas system which delivers a plurality of medical gases to a plurality of locations in the healthcare facility and having a network of computer devices, the medical gas alarm system comprising at least one area alarm controller adapted to receive a first signal indicative of a condition of a first portion of the medical gas system and adapted to communicate with the network, and at least one master alarm controller adapted to receive a second signal indicative of a condition of a second portion of the medical gas system and adapted to communicate with the network, the at least one area alarm controller being adapted to communicate with the at least one master alarm controller through the network. 2. The medical gas alarm system of claim 1, wherein the area alarm controller is adapted to couple to a first network hub of the network. 3. The medical gas alarm system of claim 2, wherein the master alarm controller is adapted to couple to a second network hub of the network. 4. The medical gas alarm system of claim 3, wherein the master alarm controller communicates with a server of the network through the second network hub. 5. The medical gas alarm system of claim 3, wherein the alarm controller communicates with a server of the network through the first network hub. 6. The medical gas alarm system of claim 3, wherein the area alarm controller communicates with the master alarm controller through the first and second network hubs. 7. The medical gas alarm system of claim 1, further comprising a plurality of sensor modules configured to sense gas pressures in respective gas pipes of the medical gas system, the area alarm controller including a plurality of display modules, each display module being associated with a respective sensor module, and each display module being configured to display the gas pressure sensed by the respective sensor module. 8. The medical gas alarm system of claim 7, wherein the area alarm controller includes a first electric circuit coupled to the network, the first electric circuit also being coupled to the plurality of display modules and to the plurality of sensor modules, and the first electric circuit operating to communicate the gas pressures sensed by the plurality of sensor modules to the network. 9. The medical gas alarm system of claim 8, wherein the master controller includes a second electric circuit that receives from the network the gas pressures sensed by the plurality of sensor modules. 10. The medical gas alarm system of claim 8, wherein the master controller includes a second electric circuit, the first electric circuit including a first microprocessor, the second electric circuit including a second microprocessor, and each of the plurality of sensor modules includes a respective microprocessor. 11. The medical gas alarm system of claim 7, wherein the area alarm controller activates a first audible alarm when an alarm ... | Hill-Rom Services, Inc. | G08B029/00 G01N009/00 A61B005/00
US07087015 | B1 | 2006-08-08 | Neurological pathology diagnostic apparatus and methods | 2000-01-31 | Apparatus and methods for rapidly diagnosing the presence or absence of the symptoms of neurological pathology caused by physical head trauma (such as occurs in contact sports or automobile collisions), disease (such as occurs in Alzheimer's disease), toxins (substance abuse or environmental toxins) or infection (such as occurs as a side effect of later-stage AIDS infection). These inventions are useful for diagnosing neurological pathology, as well as for monitoring recovery from or maintenance or progression of neurological pathology. | 1. A method of measuring the cognitive performance of an individual comprising: a. the individual completing at least one cognitive test with at least one testing protocol; b. storing results of said at least one cognitive test in a computer readable media; and c. applying a reliable change technique to obtain a statistically meaningful inference of a neurological pathology, wherein the reliable change technique uses at least one baseline of the individual. | 1. A method of measuring the cognitive performance of an individual comprising: a. the individual completing at least one cognitive test with at least one testing protocol; b. storing results of said at least one cognitive test in a computer readable media; and c. applying a reliable change technique to obtain a statistically meaningful inference of a neurological pathology, wherein the reliable change technique uses at least one baseline of the individual. | PanMedix, Inc. | A61B005/00
US07085600 | B2 | 2006-08-01 | Cardiac event detector | 2002-09-18 | A detector for detecting a cardiac event in an IEGM has a bank of digital filters composed of at least two filters with substantially different impulse responses corresponding to different properties of the cardiac event. The filters respectively filter different portions of the IEGM and respectively emit a value per time unit corresponding to the morphological resemblance between the IEGM and the finite impulse response. An analyzing unit combines these output values from the filters in a predetermined manner for comparison with predetermined detection criteria. | 1. A cardiac event detector for detecting cardiac events in an IEGM, comprising: a plurality of digital filters having respectively different filter characteristics, at least two of said filters having substantially orthogonal impulse responses to different properties of a cardiac event to be detected, said at least two filters each emitting a value per time unit corresponding to a morphological resemblance between a portion of said IEGM and the filter impulse response for that filter; andan analyzing unit supplied with respective outputs from said plurality of digital filters, including said at least two filters, for comparing said output signals with a defined detection criterion for said cardiac event to be detected. | 1. A cardiac event detector for detecting cardiac events in an IEGM, comprising: a plurality of digital filters having respectively different filter characteristics, at least two of said filters having substantially orthogonal impulse responses to different properties of a cardiac event to be detected, said at least two filters each emitting a value per time unit corresponding to a morphological resemblance between a portion of said IEGM and the filter impulse response for that filter; andan analyzing unit supplied with respective outputs from said plurality of digital filters, including said at least two filters, for comparing said output signals with a defined detection criterion for said cardiac event to be detected. 2. A cardiac event detector as claimed in claim 1 wherein each of said at least two filters emits an integer as said value per time unit. 3. A cardiac event detector as claimed in claim 1 wherein said analyzing unit employs the respective values per time unit from said at least two filters as weights for modeling a waveform of said portion of said IEGM in a combination of the respective filter impulse responses. 4. A cardiac event detector as claimed in claim 3 wherein said filters generate said weights using previously calculated weights in a recursive procedure. 5. A cardiac event detector as claimed in claim 1 wherein said analyzing unit performs a mathematical operation on said values for obtaining a mathematical computational result, and compares said mathematical computational result to a predetermined threshold value for determining whether said cardiac event exists. 6. A cardiac event detector as claimed in claim 5 wherein said analyzing unit forms a linear combination of said values as said mathematical computation. 7. A cardiac event detector as claimed in claim 5 wherein said analyzing unit calculates T(x)=x(n)TH(HTH)−1HTx(n), wherein x(n)=[xn−m+1, . . . , xn] representing an analog-to-digital converted IEGM, wherein m designates a number of samples, n designates the current time, H designates a weighting matrix determined from respective impulse responses of said filters. 8. A cardiac event detector as claimed in claim 7 wherein said analyzing unit calculates T(x) with a time interval corresponding to k samples, wherein k is an integer. 9. A cardiac event detector as claimed in claim 8 wherein k is less than 10. 10. A cardiac event detector as claimed in claim 5 wherein said analyzing unit calculates a quantity T(x)=a(x)Tsign(a(x)) wherein x(n)=[xn−m+1, . . . , xn] representing an analog-to-digital converted IEGM, wherein m designates a number of samples, n designates the current time, H designates a weighting matrix determined from respective impulse responses of said filters. 11. A cardiac event detector as claimed in claim 10 wherein said analyzing unit calculates T(x) with a time interval corresponding to k samples, wherein k is an integer. 12. A cardiac event detector as claimed in claim 11 wherein k is... | St. Jude Medical AB | A61B005/04
US07043287 | B1 | 2006-05-09 | Method for modulating light penetration depth in tissue and diagnostic applications using same | 1999-10-15 | Devices and methods for non-invasively measuring at least one parameter of a sample, such as the presence of a disease condition, progression of a disease state, presence of an analyte, or concentration of an analyte, in a biological sample, such as, for example, a body part. In these devices and methods, temperature is controlled and is varied between preset boundaries. The methods and devices measure light that is reflected, scattered, absorbed, or emitted by the sample from an average sampling depth, dav, that is confined within a region in the sample wherein temperature is controlled. According to the method of this invention, the sampling depth dav, in human tissue is modified by changing the temperature of the tissue. The sampling depth increases as the temperature is lowered below the body core temperature and decreases when the temperature is raised within or above the body core temperature. Changing the temperature at the measurement site changes the light penetration depth in tissue and hence dav. Change in light penetration in tissue as a function of temperature can be used to estimate the presence of a disease condition, progression of a disease state, presence of an analyte, or concentration of an analyte in a biological sample. According to the method of this invention, an optical measurement is performed on a biological sample at a first temperature. Then, when the optical measurement is repeated at a second temperature, light will penetrate into the biological sample to a depth that is different from the depth to which light penetrates at the first temperature by from about 5% to about 20%. | 1. A method of measuring at least one optical parameter of intact human tissue, wherein said at least one optical parameter is selected from the group consisting of absorption coefficient, scattering coefficient, mean free path, effective attenuation coefficient, and light penetration depth, said method comprising the steps of: (a) setting the temperature of said intact human tissue to a first temperature and allowing said intact human tissue to equilibrate at said first temperature before optical data are collected at said first temperature, said first temperature being within the range of from about 0° C. to about 45° C.;(b) performing an optical measurement on said intact human tissue at said first temperature, wherein said optical measurement is a spatially resolved diffuse reflectance measurement;(c) determining at least one optical parameter of said intact human tissue from said spatially resolved diffuse reflectance measurement at said first temperature, wherein said at least one optical parameter is selected from the group consisting of absorption coefficient, scattering coefficient, mean free path, effective attenuation coefficient, and light penetration depth, said first temperature corresponding to a first depth in said intact human tissue;(d) changing said first temperature of said intact human tissue to at least a second temperature and allowing said intact human tissue to equilibrate at at least said second temperature before optical data are collected at said at least said second temperature, said at least second temperature being within the range of from about 0° C. to about 45° C.;(e) performing said an optical measurement on said intact human tissue at said at least second temperature, wherein said optical measurement is a spatially resolved diffuse reflectance measurement;(f) determining said at least one optical parameter of said intact human tissue from said spatially resolved diffuse reflectance measurement at at least a second temperature, wherein said at least one optical parameter is selected from the group consisting of absorption coefficient, scattering coefficient, mean free path, effective attenuation coefficient, and light penetration depth, said at least second temperature corresponding to a second depth in said intact human tissue; and(g) determining said at least one parameter of said intact human tissue from the functional relationship of said at least one optical parameter on depth in said intact human tissue, wherein said optical measurements in step (b) and step (e) are carried out by transmitting light into a region of said intact human tissue at a light introduction site and collecting light re-emitted from said region of intact human tissue at a light collection site, wherein the distance between any light introduction site and any light collection site is less than three millimeters. | 1. A method of measuring at least one optical parameter of intact human tissue, wherein said at least one optical parameter is selected from the group consisting of absorption coefficient, scattering coefficient, mean free path, effective attenuation coefficient, and light penetration depth, said method comprising the steps of: (a) setting the temperature of said intact human tissue to a first temperature and allowing said intact human tissue to equilibrate at said first temperature before optical data are collected at said first temperature, said first temperature being within the range of from about 0° C. to about 45° C.;(b) performing an optical measurement on said intact human tissue at said first temperature, wherein said optical measurement is a spatially resolved diffuse reflectance measurement;(c) determining at least one optical parameter of said intact human tissue from said spatially resolved diffuse reflectance measurement at said first temperature, wherein said at least one optical parameter is selected from the group consisting of absorption coefficient, scattering coefficient, mean free path, effective attenuation coefficient, and light penetration depth, said first temperature corresponding to a first depth in said intact human tissue;(d) changing said first temperature of said intact human tissue to at least a second temperature and allowing said intact human tissue to equilibrate at at least said second temperature before optical data are collected at said at least said second temperature, said at least second temperature being within the range of from about 0° C. to about 45° C.;(e) performing said an optical measurement on said intact human tissue at said at least second temperature, wherein said optical measurement is a spatially resolved diffuse reflectance measurement;(f) determining said at least one optical parameter of said intact human tissue from said spatially resolved diffuse reflectance measurement at at least a second temperature, wherein said at least one optical parameter is selected from the group consisting of absorption coefficient, scattering coefficient, mean free path, effective attenuation coefficient, and light penetration depth, said at least second temperature corresponding to a second depth in said intact human tissue; and(g) determining said at least one parameter of said intact human tissue from the functional relationship of said at least one optical parameter on depth in said intact human tissue, wherein said optical measurements in step (b) and step (e) are carried out by transmitting light into a region of said intact human tissue at a light introduction site and collecting light re-emitted from said region of intact human tissue at a light collection site, wherein the distance between any light introduction site and any light collection site is less than three millimeters. 2. The method of claim 1, wherein said change in temperature of said intact human tissue results in a change in penetration... | Abbott Laboratories | A61B005/00
US07232416 | B2 | 2007-06-19 | Method and device for determining the anatomical condition of a human being or an animal and a strength training machine and a computer program which is saved to a memory medium | 2001-04-12 | A method for determining the anatomic condition of components of the body of a human or of an animal is disclosed comprising producing a sound signal in a selected component of the body by at least one of passive and active movements of at least one joint. The method includes recording a sound signal at a point adjacent to the selected component of the body and evaluating the sound signal. An auxiliary device selected from the group consisting of a splint, an orthosis and an exercise machine is used which ensures a repeatable movement determined by the design of the auxiliary device. The sound signal caused by the movement is recorded and evaluated by a spectral analysis of frequencies and amplitudes contained in the sound signal. A comparison is made with reference patterns of the spectral analyses for the diagnosis of the anatomic condition. The result of a diagnosis is indicated as at least one of a print out and a screen display and wherein the joint investigated is schematically illustrated on the screen display, optionally with associated soft tissues, tendons and ligaments. The diagnosis is indicated visually, by at least one of coloring, lighting up of and flashing of damaged positions. | 1. A method for determining the anatomic condition of components of the body of a human or of an animal having joints, comprising: utilizing an auxiliary device selected from the group consisting of a brace, an orthosis, and an exercise machine to cause reproducible passive movements or reproducible active movements of at least one joint to produce a sound signal in a selected component of the body associated with said at least one joint;recording said sound signal at a point adjacent to the selected component of the body;evaluating said sound signal, by a spectral analysis of frequencies and amplitudes contained in said sound signal;making a comparison of the outcome of said spectral analysis with a plurality of stored reference patterns for corresponding spectral analyses for the diagnosis of the anatomic condition; andindicating the result of a diagnosis as a screen display in the form of an animation in which the illustrated joint executes movements which correspond to those of said reproducible movements of said at least one joint, with these movements being indicated either synchronously or asynchronously to the actual movement of said at least one joint, said at least one joint being schematically illustrated on the screen display and the diagnosis being indicated visually by highlighting of damaged positions on said animation. | 1. A method for determining the anatomic condition of components of the body of a human or of an animal having joints, comprising: utilizing an auxiliary device selected from the group consisting of a brace, an orthosis, and an exercise machine to cause reproducible passive movements or reproducible active movements of at least one joint to produce a sound signal in a selected component of the body associated with said at least one joint;recording said sound signal at a point adjacent to the selected component of the body;evaluating said sound signal, by a spectral analysis of frequencies and amplitudes contained in said sound signal;making a comparison of the outcome of said spectral analysis with a plurality of stored reference patterns for corresponding spectral analyses for the diagnosis of the anatomic condition; andindicating the result of a diagnosis as a screen display in the form of an animation in which the illustrated joint executes movements which correspond to those of said reproducible movements of said at least one joint, with these movements being indicated either synchronously or asynchronously to the actual movement of said at least one joint, said at least one joint being schematically illustrated on the screen display and the diagnosis being indicated visually by highlighting of damaged positions on said animation. 2. The method in accordance with claim 1, wherein the animation on the screen display is effected to be substantially slower than said passive or active movements of the joint, for the purpose of easier comprehension of the damage which is present. 3. The method in accordance with claim 1, wherein the auxiliary device is used for the therapeutic treatment of the joint and the setting of the auxiliary device is effected with reference to the sound signals that are recorded. 4. The method in accordance with claim 1, wherein the auxiliary device is designed for at least one of the rehabilitation of the joint, the training of the joint and the training of the muscles associated with the joint. 5. The method in accordance with claim 1, wherein said highlighting comprises at least one of colouring of, lighting up of, and flashing of said damaged positions. 6. The method in accordance with claim 1, wherein the result of the diagnosis is also printed out. | - | A61B005/103 A61B005/117 A61B005/00 A61B007/00 A61M025/00 A63B021/005 A63B071/00
US07020578 | B2 | 2006-03-28 | Method for evaluating novel, stroke treatments using a tissue risk map | 2001-02-02 | A method of evaluating novel stroke treatments includes generating a risk map indicative of the probability of tissue infarction on voxel-by-voxel basis and selecting a probability range for evaluating the therapeutic effect of the novel treatment. In one particular embodiment, tissue having a fifty percent probability of tissue infarction is selected. A novel treatment that has a reduced level of overall actual infarction as compared to the predicted value is indicative of therapeutic effect. | 1. A method of evaluating a novel stroke treatment, comprising: generating a risk map indicating voxel-by-voxel statistical probability of tissue infarction; selecting a probability risk threshold for tissue infarction risk in the risk map; evaluating the therapeutic effect of the novel stroke treatment by comparing the selected probability risk threshold and actual tissue infarction state on a voxel-by-voxel basis. | 1. A method of evaluating a novel stroke treatment, comprising: generating a risk map indicating voxel-by-voxel statistical probability of tissue infarction; selecting a probability risk threshold for tissue infarction risk in the risk map; evaluating the therapeutic effect of the novel stroke treatment by comparing the selected probability risk threshold and actual tissue infarction state on a voxel-by-voxel basis. 2. The method according to claim 1, further including combining DWI and PWI data to generate the risk map. 3. The method according to claim 2, wherein the DWI data is selected from one or more types selected from the group consisting of elements of the diffusion tensor such as ADC, low-b EPI T2 image, anisotropy indices, eigenvalues and eigenvectors. 4. The method according to claim 2, wherein the PWI data is selected from one or more types selected from the group consisting of CBF, CBV, MTT and other parameters of the dynamic susceptibility contrast-enhanced images. 5. The method according to claim 2, wherein the risk map is generated using a GLM algorithm combining EPI T2, ADC, DWI, CBF, CBV, and/or MTT data. 6. The method according to claim 1, further including utilizing a GLM algorithm to generate the risk map. 7. The method according to claim 1, further including utilizing a GAM algorithm to generate the risk map. 8. The method according to claim 1, further including determining whether the novel treatment achieves a statically significant difference as compared to a conventional therapy. 9. The method according to claim 8, further including determining whether the novel therapy is effective. 10. The method according to claim 1, further including calculating true positive, true negative, false positive, and false negative values. 11. The method according to claim 1, further including generating a plurality of risk maps, each of which corresponds to a particular therapy. 12. The method according to claim 11, further including selecting one of the plurality of therapy based upon the plurality of risk maps. 13. The method according to claim 12, further including selecting a therapy corresponding to a respective one of the plurality of risk maps having a smallest volume of tissue risk of infarction. 14. The method according to claim 12, further including selecting a therapy corresponding to a respective one of the plurality of risk maps that minimizes a risk threshold of infarction. 15. The method according to claim 1, further including selecting the risk threshold based upon which tissue is likely to infarct if a patient receives conventional therapy. 16. The method according to claim 1, further including selecting the risk threshold based upon an optimal operating point for discriminating tissue likely to infarct front tissue not likely to infarct. 17. The method according to claim 1, further including selecting the risk threshold of about thirty percent. 18. The method according to claim 1, further including sele... | The General Hospital Corporation | A61B005/055
US06990364 | B2 | 2006-01-24 | Noninvasive measurement of glucose through the optical properties of tissue | 2002-01-25 | Methods and apparatus for noninvasive determination of blood analytes, such as glucose, through NIR spectroscopy utilize optical properties of tissue as reflected in key spectroscopic features to improve measurement accuracy and precision. Physiological conditions such as changes in water distribution among tissue compartments lead to complex alterations in the measured absorbance spectrum of skin and reflect a modification in the effective pathlength of light, leading to a biased noninvasive glucose measurement. Changes in the optical properties of tissue are detected by identifying key features responsive to physiological variations. Conditions not conducive to noninvasive measurement of glucose are detected. Noninvasive glucose measurements that are biased by physiological changes in tissue are compensated. In an alternate embodiment, glucose is measured indirectly based on natural physiological response of tissue to glucose concentration. A spectroscopic device capable of such measurements is provided. | 1. A method for noninvasive measurement of a target analyte in a tissue sample, comprising the steps of: measuring a spectrum of a tissue sample; detecting at least one optical property of said tissue sample related to fluid shift between extravascular and intravascular tissue compartments triggered by changes in concentration of said analyte, as manifested by at least one spectral feature reflecting said fluid shift; and either correcting a direct analyte measurement on the basis of said at least one optical property; or measuring said analyte indirectly on the basis of said at least one optical property. | 1. A method for noninvasive measurement of a target analyte in a tissue sample, comprising the steps of: measuring a spectrum of a tissue sample; detecting at least one optical property of said tissue sample related to fluid shift between extravascular and intravascular tissue compartments triggered by changes in concentration of said analyte, as manifested by at least one spectral feature reflecting said fluid shift; and either correcting a direct analyte measurement on the basis of said at least one optical property; or measuring said analyte indirectly on the basis of said at least one optical property. 2. The method of claim 1, wherein said step of measuring a spectrum comprises measuring a spectrum, m, of said tissue, said spectrum denoted by a vector m∈1×N of absorbance values pertaining to a set of N wavelengths, λ∈N spanning near IR region of approximately 700 to 2500 nm. 3. The method of claim 2, wherein m comprises a set of wavelengths in the near IR that have been optimized for extraction of features. 4. The method of claim 1, further comprising the step of preprocessing said spectrum, preprocessing comprising any of: filtering; calculating a first or second derivative of said spectrum; and scatter correction. 5. The method of claim 1, further comprising the step of extracting said at least one features, feature extraction comprising any mathematical transformation that enhances a quality or aspect of a sample measurement for interpretation so that structural properties and physiological state of said tissue sample are concisely represented. 6. The method of claim 5, wherein said feature extraction step comprises the steps of: identifying distinct absorption bands that change in differing manners with respect to changes in pathlength, scattering and absorption properties of the sample; developing a set of features that represents or reflects the optical properties of the tissue; and applying said features to identify; conditions suitable for analyze measurement; and conditions unsuitable for analyte measurement. 7. The method of claim 5, wherein features include any of: at least one simple features, comprising values of a spectral measurement or a processed spectral measurement at critical points, wherein a critical point is a point having a slope of zero; a derived features, comprising at least one features derived from simple features through a mathematical transformation; and an abstract features; wherein simple and derived features generally have a physical interpretation, and wherein abstract features do not necessarily have a specific interpretation related to a physical system. 8. The method of claim 7, wherein an abstract features comprises scores from a principal component analysis, wherein leading principal components represent variation related to the structural properties and physiological state of the tissue sample, so that the optical properties of the tissue are represented. 9. The method of claim 7, wherein... | Sensys Medical, Inc. | A61B005/00
US07069072 | B2 | 2006-06-27 | Method and device for determining the segmental volume and electrical parallel conductance of a cardiac chamber or blood vessel | 2001-08-30 | The segment volume of a cardiac chamber or blood vessel of a patient is determined by injecting a first indicator in the blood stream of a patient, which influences the conductance of the blood. The electrical conductance in the cardiac chamber or blood vessel is measured. An injected quantity of indicator is determined and the development of the concentration of this indicator in the blood is measured wherein the cardiac output is calculated from the injected quantity of indicator and the development of the concentration in the blood. Subsequently, the segment volume and electrical parallel conductance of a cardiac chamber or blood vessel are calculated from the calculated cardiac output, the injected quantity of conductance indicator and the measured conductance. | 1. A method for determining the segmental volume and the electrical parallel conductance of a cardiac chamber or a blood vessel of a patient, the method comprising: injecting a conductance indicator, which influences electrical conductance of blood, into a bloodstream of a patient, wherein an injected quantity of the conductance indicator is determined;measuring the electrical conductance in said cardiac chamber or blood vessel;calculating a sum of segmental volume of said cardiac chamber or blood vessel and electrical parallel conductance from said measured electrical conductance;measuring development of concentration of a concentration indicator in the bloodstream;calculating the cardiac output from the injected quantity of concentration indicator and the development of the concentration thereof; andcalculating the segmental volume and the electrical parallel conductance of the cardiac chamber or the blood vessel from the calculated cardiac output, the injected quantity of conductance indicator and the measured electrical conductance. | 1. A method for determining the segmental volume and the electrical parallel conductance of a cardiac chamber or a blood vessel of a patient, the method comprising: injecting a conductance indicator, which influences electrical conductance of blood, into a bloodstream of a patient, wherein an injected quantity of the conductance indicator is determined;measuring the electrical conductance in said cardiac chamber or blood vessel;calculating a sum of segmental volume of said cardiac chamber or blood vessel and electrical parallel conductance from said measured electrical conductance;measuring development of concentration of a concentration indicator in the bloodstream;calculating the cardiac output from the injected quantity of concentration indicator and the development of the concentration thereof; andcalculating the segmental volume and the electrical parallel conductance of the cardiac chamber or the blood vessel from the calculated cardiac output, the injected quantity of conductance indicator and the measured electrical conductance. 2. The method according to claim 1, wherein the conductance indicator comprises the concentration indicator, and wherein measuring the development of the concentration comprises measuring the temperature or the color of the concentration indicator. 3. The method according to claim 1, and further comprising injecting the concentration indicator. 4. The method according to claim 1, and further comprising: using the calculated parallel conductance a correction value; anddetermining the segmental volume exclusively from the conductance signal using said correction value. 5. The method according to claim 1, and further comprising: dividing a period of a heartbeat into equal time segments; andmonitoring continuously the segmental volume and the electric parallel conductance of the cardiac chamber or the blood vessel within the for multiple heartbeat. 6. An apparatus for determining a segmental volume of a cardiac chamber or a blood vessel of a patient, the apparatus comprising: a processing unit having a control output for initiating injection of a conductance indicator into a bloodstream of the patient, which influences electrical conductance of blood;a first detector configured to be placed in the cardiac chamber or blood vessel for measuring a conductance signal, which depends on the electrical conductance at the location of the first detector, wherein said processing unit is arranged for calculating the segmental volume of the cardiac chamber or the blood vessel from said conductance signal;a second detector configured to be placed in the bloodstream for measuring an indicator signal depending on a concentration of a concentration indicator in the blood, wherein said processing unit is arranged for calculating the cardiac output from an injected quantity of the conductance indicator and the indicator signal, and calculating the segmental volume and electrical parallel conductance of the cardiac chamber or blood... | Jozef R. C. Jansen | A61B005/00
US07076291 | B2 | 2006-07-11 | Heart rate monitor | 2002-01-14 | The invention relates to a heart rate monitor. The essential point of the invention is that a display element (201) of a display (20) of the heart rate monitor for displaying a settable minimum limit for a heart rate level is located at a first end (211) of a display element unit (210) controlled according to the heart rate level, on the same side of the display as the first end of the display element unit. Similarly, a display element (202) for displaying a settable maximum limit for a heart rate level is located at a second end (220) of the display element unit (210), on the same side of the display as the second end of the display element unit. | 1. A carry-on heart rate monitor measuring a person's heart rate non-invasively, the heart rate monitor comprising a display for displaying heart rate information about a heart rate signal measured on the person, the display comprising a display element for displaying a settable minimum limit for a desired heart rate level, a display element for displaying a settable maximum limit for a desired heart rate level, a display element unit controlled by the measured heart rate level and provided with a plurality of display element segments, at least one of the display element segments being activated in response to the measured heart rate level and graphically representing the measured heart rate level by a position of the at least one activated display element segment relative to the display element for displaying the settable minimum limit for the desired heart rate level and the display element for displaying the settable maximum limit for the desired heart rate level, wherein the display element for displaying a settable minimum limit for the heart rate level is located at a first end of the display element unit controlled according to the heart rate level, on the same side of the display as the first end of the display element unit, and that the display element for displaying a settable maximum limit for the heart rate level is located at a second end of the display element unit controlled according to the heart rate level, on the same side of the display as the second end of the display element unit, the heart rate monitor adapted for being a carry-on heart rate monitor. | 1. A carry-on heart rate monitor measuring a person's heart rate non-invasively, the heart rate monitor comprising a display for displaying heart rate information about a heart rate signal measured on the person, the display comprising a display element for displaying a settable minimum limit for a desired heart rate level, a display element for displaying a settable maximum limit for a desired heart rate level, a display element unit controlled by the measured heart rate level and provided with a plurality of display element segments, at least one of the display element segments being activated in response to the measured heart rate level and graphically representing the measured heart rate level by a position of the at least one activated display element segment relative to the display element for displaying the settable minimum limit for the desired heart rate level and the display element for displaying the settable maximum limit for the desired heart rate level, wherein the display element for displaying a settable minimum limit for the heart rate level is located at a first end of the display element unit controlled according to the heart rate level, on the same side of the display as the first end of the display element unit, and that the display element for displaying a settable maximum limit for the heart rate level is located at a second end of the display element unit controlled according to the heart rate level, on the same side of the display as the second end of the display element unit, the heart rate monitor adapted for being a carry-on heart rate monitor. 2. A heart rate monitor as claimed in claim 1, wherein the display element for displaying a settable minimum limit for the heart rate level and the first end of the display element unit controlled according to the heart rate level are located on the same side of the display with respect to both a center line of the display parallel to the reading direction of the display and a center line of the display perpendicular to the reading direction of the display. 3. A heart rate monitor as claimed in claim 2, wherein the display element for displaying a settable maximum limit for the heart rate level and the second end of the display element unit controlled according to the heart rate level are located on the same side of the display with respect to both a center line of the display parallel to the reading direction of the display and a center line of the display perpendicular to the reading direction of the display. 4. A heart rate monitor as claimed in claim 1, wherein the display element for displaying a settable maximum limit for the heart rate level and the second end of the display element unit controlled according to the heart rate level are located on the same side of the display with respect to both a center line of the display parallel to the reading direction of the display and a center line of the display perpendicular to the reading direction of the display. 5. A hear... | Polar Electro Oy | A61B005/04
US07070579 | B1 | 2006-07-04 | Device used to connect an external ventricular drainage catheter | 1998-04-30 | A luer connector is described that is angled away from the patient's scalp in the direction of the drip assembly line. In the preferred embodiment, the luer connector has a female luer connector that mates with a male luer connector on the drip assembly line. The luer connector also has a hollow protrusion that extends into a catheter thereby allowing the luer connector to be fluidly connected to the catheter. A fluid passage is formed through the angled luer connector from the female luer connector, through the body of the angled luer connector and through the hollow protrusion. The luer connector has a pair of “wings” that extend outwardly from the luer connector and allow the luer connector to be sutured to the patient's scalp. | 1. A bodily fluid drainage assembly having a catheter in the body and a luer connector for connecting the catheter to a drip assembly line comprising: a hollow barrel having a barrel lumen, the barrel having a barrel axis that is coaxial with the barrel lumen, the hollow barrel having a single terminal;a hollow catheter connection protrusion attached to and extending away from the barrel, the catheter connection protrusion sized to fit within the catheter, the catheter connection protrusion having a protrusion lumen that extends through the catheter connection protrusion, the protrusion lumen being in fluid communication with the barrel lumen, the catheter connection protrusion having a terminal end opposite the barrel;a pair of anchoring protrusions attached to and extending away from the barrel, the anchoring protrusions being formed essentially in a plane:a female luer connector attached to the barrel opposite the catheter connection protrusion, the female luer connector creating the single terminal, the female luer connector having a female luer axis that extends through and is coaxial with the female luer that is not coaxial with the barrel axis, the female luer axis extending away from the plane containing the anchoring protrusions. | 1. A bodily fluid drainage assembly having a catheter in the body and a luer connector for connecting the catheter to a drip assembly line comprising: a hollow barrel having a barrel lumen, the barrel having a barrel axis that is coaxial with the barrel lumen, the hollow barrel having a single terminal;a hollow catheter connection protrusion attached to and extending away from the barrel, the catheter connection protrusion sized to fit within the catheter, the catheter connection protrusion having a protrusion lumen that extends through the catheter connection protrusion, the protrusion lumen being in fluid communication with the barrel lumen, the catheter connection protrusion having a terminal end opposite the barrel;a pair of anchoring protrusions attached to and extending away from the barrel, the anchoring protrusions being formed essentially in a plane:a female luer connector attached to the barrel opposite the catheter connection protrusion, the female luer connector creating the single terminal, the female luer connector having a female luer axis that extends through and is coaxial with the female luer that is not coaxial with the barrel axis, the female luer axis extending away from the plane containing the anchoring protrusions. 2. The assembly of claim 1 wherein the female luer axis intersects the barrel axis at an angle of between 15° to 90°. 3. The assembly of claim 2 wherein the female luer axis intersects the barrel axis at an angle of about 30°. 4. The assembly of claim 1 wherein the pair of anchoring protrusions produce a substantially planar surface. 5. The assembly of claim 4 wherein the female luer axis intersects the substantially planar surface. 6. The assembly of claim 1 wherein the female luer axis is equidistant from each of the anchoring protrusions. 7. The assembly of claim 1 wherein the female luer axis is closer to one of the anchoring protrusions than the other. 8. The assembly of claim 1 wherein the anchoring protrusions each have a suturing hole to allow the anchoring protrusions to be attached to a patient. 9. The assembly of claim 1 wherein the protrusion lumen is coaxial with the central lumen. 10. The assembly of claim 1 wherein the protrusion has an outside diameter that of slightly larger diameter than the inner lumen of the catheter. 11. The assembly of claim 1 further comprising a bulbous end formed on the terminal end of the catheter connection protrusion. 12. A bodily fluid drainage assembly having a catheter in the body and a luer connector for connecting the catheter to drip assembly line comprising: a hollow barrel having a barrel lumen, the barrel having a barrel axis that is coaxial with the barrel lumen, the hollow barrel having a single terminal;a hollow catheter connection protrusion attached to and extending away from the barrel, the catheter connection protrusion sized to fit within the catheter, the catheter connection protrusion having a protrusion lumen that extends through the cat... | Medtronic, Inc. | A61M005/00 A61M005/178 A61B005/14
US06890312 | B1 | 2005-05-10 | Joint angle indication system | 2001-12-03 | A joint angle indication system provides information related to an angular relationship between a first body part and a second body part that are pivotally coupled at a joint. The system includes a first arm member attached to the first body part, and a second arm member attached to the second body part. One end of the second arm member is pivotally coupled to one end of the first arm member. The system includes one or more joint angle variation sensors that provide one or more electrical characteristics that vary based on variation in the joint angle between the first and second arm members. The joint angle is variable over an angular range that includes a first angle and a second angle. A biofeedback circuit generates a first feedback signal when the joint angle is less than or equal to the first angle, generates a second feedback signal when the joint angle is greater than or equal to the second angle, and generates no feedback signal when the joint angle is less than the second angle and greater than the first angle. Preferably, the first and second feedback signals are audio signals, and the second feedback signal has a different frequency from the first feedback signal. The system also includes an angle display circuit that is electrically coupled to the one or more joint angle variation sensors. The angle display circuit visually displays a joint angle value based on the one or more electrical characteristics. | 1. A joint angle indication system for providing information related to an angular relationship between a first body part and a second body part which are pivotally coupled at a joint, the system comprising: a first arm member operable for attachment to the first body part, the first arm member having a first proximal end and a first distal end; a second arm member operable for attachment to the second body part, the second arm member having a second proximal end and a second distal end, the second proximal end of the second arm member pivotally coupled to the first proximal end of the first arm member; at least one joint angle variation sensor for providing at least one electrical characteristic which varies based on variation in a joint angle of the first arm member relative to the second arm member, where the joint angle is variable over an angular range which includes a first angle and a second angle, the at least one joint angle variation sensor comprising first and second potentiometers coupled between the first and second arm members for providing a first electrical resistance of the first potentiometer and a second electrical resistance of the second potentiometer; a biofeedback circuit operable to generate a first feedback signal when the first electrical resistance of the first potentiometer indicates the joint angle is less than or equal to the first angle, operable to generate the second feedback signal when the first electrical resistance of the first potentiometer indicates the joint angle is greater than or equal to the second angle, and operable to generate no feedback signal when the first electrical resistance of the first potentiometer indicates the joint angle is less than the second angle and greater than the first angle; and an angle display circuit for visually displaying a joint angle value based on at least the second electrical resistance. | 1. A joint angle indication system for providing information related to an angular relationship between a first body part and a second body part which are pivotally coupled at a joint, the system comprising: a first arm member operable for attachment to the first body part, the first arm member having a first proximal end and a first distal end; a second arm member operable for attachment to the second body part, the second arm member having a second proximal end and a second distal end, the second proximal end of the second arm member pivotally coupled to the first proximal end of the first arm member; at least one joint angle variation sensor for providing at least one electrical characteristic which varies based on variation in a joint angle of the first arm member relative to the second arm member, where the joint angle is variable over an angular range which includes a first angle and a second angle, the at least one joint angle variation sensor comprising first and second potentiometers coupled between the first and second arm members for providing a first electrical resistance of the first potentiometer and a second electrical resistance of the second potentiometer; a biofeedback circuit operable to generate a first feedback signal when the first electrical resistance of the first potentiometer indicates the joint angle is less than or equal to the first angle, operable to generate the second feedback signal when the first electrical resistance of the first potentiometer indicates the joint angle is greater than or equal to the second angle, and operable to generate no feedback signal when the first electrical resistance of the first potentiometer indicates the joint angle is less than the second angle and greater than the first angle; and an angle display circuit for visually displaying a joint angle value based on at least the second electrical resistance. 2. The joint angle indication system of claim 1 wherein: the first arm member is forked at the first proximal end to form a first prong portion and a second prong portion; the second proximal end of the second arm member is disposed between the first and second prong portions of the first arm member; the first potentiometer is disposed between the second arm member and the first prong of the first arm member, and the second potentiometer is disposed between the second arm member and the second prong of the first arm member. 3. The joint angle indication system of claim 1 wherein the biofeedback circuit further comprises a first piezoelectric buzzer for generating the first feedback signal and a second piezoelectric buzzer for generating the second feedback signal. 4. The joint angle indication system of claim 1 further comprising an audio output circuit for generating an audio angle indication signal having a signal characteristic which varies in relation to a variation in the at least one electrical characteristic. 5. The joint angle indication system of claim 4 wherein the sign... | William B. Priester | A61B005/103
US06837988 | B2 | 2005-01-04 | Biological fluid sampling and analyte measurement devices and methods | 2001-06-12 | A device for sampling a biological fluid and measuring a target analyte within the biological fluid is provided. The device has at least one micro-piercing member used to penetrate the skin to a selected depth and access biological fluid, a sampling means and a measuring means. The sampling means comprises a fluid transfer medium, such as a hydrophilic porous material, by which sampled biological fluid is transferred from the micro-piercing member to the measuring means. The measuring means includes an electrochemical cell having at least one porous electrode and, typically, a reagent material, where the electrochemical cell is configured so as to make an electrochemical measurement of a target analyte in accessed biological fluid present therein. Methods of sampling biological fluids within the skin and measuring the sampled fluids are also provided, as well as kits comprising one or more of the inventive devices. | 1. A biological fluid sampling and analyte concentration measurement device, said device comprising: (a) at least one skin-piercing member comprising a biological fluid access opening; (b) an electrochemical cell for measuring the concentration of analyte within the biological fluid, wherein the cell comprises at least one porous electrode; and (c) a hydrophilic porous material in fluid communication with the at least one piercing member and with the at least one porous electrode, wherein hydrophilic porous material transfers biological fluid present at the access opening of the at least one piercing member to the electrochemical cell, wherein the porous material comprises a distal portion associated with the at least one piercing member and a proximal portion adjacent to the at least one porous electrode, wherein the proximal portion is more porous than the distal portion. | 1. A biological fluid sampling and analyte concentration measurement device, said device comprising: (a) at least one skin-piercing member comprising a biological fluid access opening; (b) an electrochemical cell for measuring the concentration of analyte within the biological fluid, wherein the cell comprises at least one porous electrode; and (c) a hydrophilic porous material in fluid communication with the at least one piercing member and with the at least one porous electrode, wherein hydrophilic porous material transfers biological fluid present at the access opening of the at least one piercing member to the electrochemical cell, wherein the porous material comprises a distal portion associated with the at least one piercing member and a proximal portion adjacent to the at least one porous electrode, wherein the proximal portion is more porous than the distal portion. 2. The device of claim 1 wherein the proximal portion is about 10 to 100 times more porous than the distal portion. 3. The device of claim 1 wherein the hydrophilic porous material transfers biological fluid by means of a capillary force exerted on the biological fluid in its presence. 4. The device of claim 1 wherein the porous material is selected from the group consisting of polymers, ceramics, glass and silica. 5. The device of claim 1 wherein the electrochemical cell comprises two spaced-apart electrodes defining a reaction chamber and a selected reagent for chemically reacting with an analyte targeted for measurement. 6. The device of claim 5 wherein the distance between the electrodes is from about 10 to 300 microns. 7. The device of claim 6 herein the distance between the electrodes is from about 10 to 150 microns. 8. The device according to claim 5 wherein the reagent is located on a surface of at least one electrode facing the reaction chamber. 9. The device of claim 5 wherein both electrodes are porous. 10. The device of claim 9 further comprising a housing having at least one vent hole for venting air from within the electrochemical cell. 11. The device of claim 1 wherein the biological fluid is interstitial fluid and the analyte is glucose. 12. The device of claim 1 further comprising a hydrophilic gel. 13. The device of claim 1 wherein the hydrophilic porous material comprises pores having diameters in the range from about 0.1 to 50 μm. 14. The device of claim 13 wherein the pore diameters are in the range from about 0.1 to 10 μm. 15. A system for sampling biological fluid from the skin of a patient and measuring a target analyte within the biological fluid, the system comprising: (a) at least one device according to claim 1; and (b) a control means in electrical communication with the at least one device, the control means comprising: (1) means for sending an electrical input signal to the device and for receiving an electrical output signal from the device, and (2) a software algorithm which automatically calculates and determines the concentrat... | LifeScan, Inc. | G01N027/327 A61B005/05
US06969350 | B1 | 2005-11-29 | Body fat measurement system | 1999-11-12 | A body fat measurement system for mammals includes means for measuring first and second body dimensions having a high correlation with body fat and low correlation respectively. A lookup table provides an output of the percentage body fat on input of the first and second dimensions. | 1. A method for determining the percentage body fat of a four legged domestic pet mammal, comprising the steps of: measuring a first body dimension having a high correlation with percentage body fat;measuring a second body dimension having a low correlation with percentage body fat; anddetermining the percentage body fat from a relationship between only the first and second body measurements and the percentage body fat using a body fat look-up table comprising a first storage area for storing therein entries of the first body dimension, a second storage area storing therein entries of the second body dimension and an output storage area storing an indication of the percentage body fat. | 1. A method for determining the percentage body fat of a four legged domestic pet mammal, comprising the steps of: measuring a first body dimension having a high correlation with percentage body fat;measuring a second body dimension having a low correlation with percentage body fat; anddetermining the percentage body fat from a relationship between only the first and second body measurements and the percentage body fat using a body fat look-up table comprising a first storage area for storing therein entries of the first body dimension, a second storage area storing therein entries of the second body dimension and an output storage area storing an indication of the percentage body fat. 2. A method according to claim 1, wherein the first body measurement is the circumference of the ribcage, taken at the 9th rib. 3. A method according to claim 1, wherein the second body measurement is a leg index measurement, which is the length of the hind limb measured between the patella (knee) and the calcaneal tuber (hock). 4. A method according to claim 1, wherein the output storage area provides an indication of whether the mammal is under, normal or overweight. 5. A method according to claim 1, wherein the output storage area provides a numerical percentage body fat. 6. A method according to claim 1, wherein the relationship between the percentage body fat and first and second body dimension is given by the equation: PercentageBodyFat=[(RC1-L)C2]-Lwhere R=ribcage circumferenceL=leg index measurementC1=constantC2=constant. 7. A method according to claim 1, wherein the four legged mammal is a cat. 8. A method according to claim 1, comprising the further step of determining a target body weight from a relationship between the percentage body fat, a body weight and the target body weight using a look-up table comprising a first storage area for storing entries of the percentage body fat, a second storage area for storing entries of body weight, and an output storage area storing an indication of the target body weight. 9. A method according to claim 8, wherein the relationship between the target weight, the body weight and the percentage body fat is given by the equation: Target weight=1.33×Body weight (kg)×((100−% body fat)/100). 10. A method according to claim 8, wherein the target weight is an ideal weight. 11. A method according to claim 8, further comprising the step of determining an energy allowance. 12. A method according to claim 11, wherein the energy allowance is a daily allowance in kcal/kg. | Mars U.K. Limited | A61B005/00
US06888779 | B2 | 2005-05-03 | Method and apparatus for a waking control system | 2001-11-20 | A method and apparatus are disclosed for a waking system that wakes an individual gradually over a period of time in order to promote the wellness of that individual. The user sets the system controller with a desired final wakeup time, which is the time that the user must be awake. When the actual time reaches a stimulus introduction time (i.e. some time prior to the desired final wakeup time), the system controller causes the introduction of stimulus. The system further comprises a detection system that measures a parameter correlated to the individual's sleep level (preferably motion). This measured parameter is then used in a feedback loop to control the level of stimulus experienced by the individual. In this manner, the user is woken gradually over a period of time between the stimulus introduction time and the desired final wakeup time, without introducing any sudden shocks to the user's physiological systems. | 1. An apparatus for waking an individual in a manner that promotes said individual's well being comprising: at least one sensor operative to sense at least one parameter correlated to said individual's sleep level; and a controller capable of being set by said individual with a final wakeup time, said controller connected to receive sensed values of the at least one parameter from the sensor and configured to introduce a stimulus at a stimulus introduction time prior to the final wakeup time and, over a period between the stimulus introduction time and the final wakeup time, to adjust an intensity of the stimulus based on feedback which comprises the sensed values of the at least one parameter and based on a difference between a current time and the final wakeup time, so as to bring said individual gradually out of sleep and to an awake state over a period of time between the stimulus introduction time and the final wakeup time. | 1. An apparatus for waking an individual in a manner that promotes said individual's well being comprising: at least one sensor operative to sense at least one parameter correlated to said individual's sleep level; and a controller capable of being set by said individual with a final wakeup time, said controller connected to receive sensed values of the at least one parameter from the sensor and configured to introduce a stimulus at a stimulus introduction time prior to the final wakeup time and, over a period between the stimulus introduction time and the final wakeup time, to adjust an intensity of the stimulus based on feedback which comprises the sensed values of the at least one parameter and based on a difference between a current time and the final wakeup time, so as to bring said individual gradually out of sleep and to an awake state over a period of time between the stimulus introduction time and the final wakeup time. 2. An apparatus according to claim 1, wherein said at least one sensor comprises at least one of: an ultrasonic motion detector and an infrared motion detector. 3. An apparatus according to claim 1, wherein said controller is configured to independently adjust an intensity of a plurality of stimuli based on feedback which comprises the sensed values of the at least one parameter and based on a difference between a current time and the final wakeup time. 4. An apparatus according to claim 1, wherein the at least one parameter sensed by said sensor comprises at least one of: motion, brain waves, skin potential, skin resistance, muscle tone, eye movement, heart rate and breathing rate of said individual. 5. An apparatus according to claim 1 comprising a user interface having an input mechanism operative to receive input from said individual and an output mechanism operative to communicate information to said individual. 6. An apparatus according to claim 5, wherein said input mechanism comprises at least one of: a voice recognition system, a keypad, a touch screen interface, a remote control device and a plurality of buttons and switches. 7. An apparatus according to claim 5, wherein said output mechanism comprises at least one of: a liquid crystal display, a light emitting diode display, an analog clock display, a plasma screen display, a cathode ray tube display, an audio output device and a plurality of switch and button indicators. 8. An apparatus according to claim 5, wherein said input mechanism is operative to receive input information from said individual, the input information comprising at least one of: a particular set of stimuli desired, a desired relative intensity of each stimulus selected, the stimulus introduction time, a personal sleep sensitivity level of said individual, a recent sleep history of said individual, an ambient stimulus level in said individual's sleeping environment and a seasonal amount of daylight. 9. An apparatus according to claim 5, wherein said output mechanism is operative to ... | Pulsar Informatics, Inc. | G04B047/00 G04B023/02 G04C021/00 A61M021/00 A61B005/24
US06979298 | B2 | 2005-12-27 | Apparatus and method for the measurement of the aerodynamics of olfaction in animals and man | 2002-05-31 | A system and method for observing, collecting and analyzing olefactory characteristics of a human or animal subject, such as sniffing, breathing and respiratory patterns and sounds. Example applications include the evaluation and training of dogs for explosive and drug detection, clinical diagnostics, scientific research, and identification. | 1. A system for monitoring olfactory characteristics of a subject, the system comprising a transmitting microphone for attachment in a position generally overlying the olfactory epithelium of the subject, a receiver for receiving a signal indicative of an internal olfactory characteristic from the microphone, and a recorder for recording data corresponding to sounds from the olfactory epithelium as a function of the presentation of at least one stimulus presented to the subject. | 1. A system for monitoring olfactory characteristics of a subject, the system comprising a transmitting microphone for attachment in a position generally overlying the olfactory epithelium of the subject, a receiver for receiving a signal indicative of an internal olfactory characteristic from the microphone, and a recorder for recording data corresponding to sounds from the olfactory epithelium as a function of the presentation of at least one stimulus presented to the subject. 2. The system of claim 1, further comprising a harness for attaching the microphone over the olfactory epithelium of the subject. 3. The system of claim 1, further comprising a display for displaying data corresponding to olfactory characteristics of the subject. 4. A method for monitoring olfactory characteristics of a subject, said method comprising the steps of attaching a microphone at or adjacent a nasal region of the subject, presenting at least one olfactory stimulus to the subject, receiving signals from the microphone corresponding to the olfactory characteristics of the subject, and recording data corresponding to the olfactory characteristics of the subject as a function of the presentation of the at least one olfactory stimulus. 5. The method of claim 4, further comprising monitoring states of sniffing by a dog subject, including a first state of searching for a target olfactory stimulus and a second state wherein the dog subject is presented with the target olfactory stimuli. 6. The method of claim 4, further comprising monitoring long-range detection of olfactory stimuli by a dog subject wherein sniff patterns of the subject change to alert a handler of the dog subject of the presence of a specific materials. 7. The method of claim 4, further comprising the step of monitoring breathing and the respiratory extent and rate of the subject. 8. The method of claim 4, further comprising the step of monitoring a respiratory dysfunction in the subject. 9. The method of claim 8, further comprising conducting a chemical analysis of exhaled breath of the subject. 10. The method of claim 4, wherein the microphone is attached in a position generally overlying the olfactory epithelium of the subject. 11. A system for monitoring olfactory characteristics of an animal, the system comprising: a microphone mounted to a harness adapted to be worn by the animal; anda computer for receiving signals from the microphone and outputting information related to an olfactory characteristic of the animal, wherein the information related to an olfactory characteristic of the animal includes a sniffing frequency of the animal, and wherein the computer compares the sniffing frequency of the animal with a target sniffing frequency. 12. The system of claim 11, wherein the system outputs an alert to a handler of the animal based on the comparison of the sniffing frequency of the animal with the target sniffing frequency. 13. The system of claim 11, wherein the harness holds the mi... | Auburn University | A61B005/00
US06907281 | B2 | 2005-06-14 | Fast mapping of volumetric density data onto a two-dimensional screen | 2001-09-07 | A method and system for constructing from volumetric CT or MRI scan data of a patient region, a virtual two-dimensional fluoroscopic image of the patient region as seen from a selected point in space are disclosed. In practicing the method, a plurality of rays are constructed between a selected view point and each of a plurality of points in an XY array, where at least some of the rays pass through the patient target region, the points in the XY array correspond to the XY array of pixels in a display screen, and each pixel in the display screen includes multiple N-bit registers for receiving digital-scale values for each of multiple colors. For each pixel element, a sum of all M-bit density values associated with voxels along a ray extending from the selected point to the associated pixel element is calculated. The summing is carried out by distributing the M bits in the voxel density values among the multiple N-bit registers of that pixel, such that such that one or more bit positions of each M-bit value is assigned to a selected register. The image constructed of gray-scale values representing the summed M-bit density values at each pixel in the screen is displayed to the user. | 1. A method for constructing from volumetric CT or MRI scan data of a patient region, a virtual two-dimensional fluoroscopic or shadowgraph image of the patient region as seen from a selected point in space, comprising (a) selecting a point in space representing a virtual fluoroscopic irradiation source, (b) defining a plurality of rays extending between the source point and each of a plurality of points in a two-dimensional XY array of points, where (i) the array points correspond to pixels in an XY pixel array in a display screen, (ii) each pixel in the display screen has multiple N-bit registers for receiving digitized values for each of multiple colors, (iii) the scan data is composed of voxels having defined coordinates and associated M-bit tissue-density values, and (iv) at least some of the rays pass through a plurality of such voxels in the patient target region, (c) for each ray, distributing the M bits of the density values of voxels along that ray among the multiple N-bit registers of the associated pixel, such that one or more bit positions of each M-bit value is assigned to a selected register of that pixel, (d) for each ray, summing the M-bit values of the voxels along that ray by (i) individually summing the one or more bit-position values in each associated pixel register, and (ii) determining the sum of the M-bit values along the associated ray from the values in the individual registers, and (e) displaying the image constructed of gray-scale values representing the summed M-bit density values at each pixel in said display screen. | 1. A method for constructing from volumetric CT or MRI scan data of a patient region, a virtual two-dimensional fluoroscopic or shadowgraph image of the patient region as seen from a selected point in space, comprising (a) selecting a point in space representing a virtual fluoroscopic irradiation source, (b) defining a plurality of rays extending between the source point and each of a plurality of points in a two-dimensional XY array of points, where (i) the array points correspond to pixels in an XY pixel array in a display screen, (ii) each pixel in the display screen has multiple N-bit registers for receiving digitized values for each of multiple colors, (iii) the scan data is composed of voxels having defined coordinates and associated M-bit tissue-density values, and (iv) at least some of the rays pass through a plurality of such voxels in the patient target region, (c) for each ray, distributing the M bits of the density values of voxels along that ray among the multiple N-bit registers of the associated pixel, such that one or more bit positions of each M-bit value is assigned to a selected register of that pixel, (d) for each ray, summing the M-bit values of the voxels along that ray by (i) individually summing the one or more bit-position values in each associated pixel register, and (ii) determining the sum of the M-bit values along the associated ray from the values in the individual registers, and (e) displaying the image constructed of gray-scale values representing the summed M-bit density values at each pixel in said display screen. 2. The method of claim 1, wherein the pixels in the display screen each contain four 8-bit registers, the density values in each voxel are 8-bit numbers, and the 8 bit voxel-density numbers are distributed two digits each to each of the registers. 3. The method of claim 1, wherein the pixels in the display screen each contain four 8-bit registers, the density values in each voxel are 8-bit numbers, and the 8 bit voxel-density numbers are distributed 3 digits, 3 digits, 1 digit and 1 digit to each of the four registers. 4. A system for constructing from volumetric CT or MRI scan data of a patient region, a virtual two-dimensional fluoroscopic image of the patient region as seen from a selected point in space, comprising (a) a scan-data file for storing, in digital form, volumetric scan data of a surgical site of a patient, where the scan data is composed of voxels having defined coordinates and associated M-bit tissue-density values, (b) a display screen containing a two-dimensional XY array of pixel elements, each element containing multiple N-bit registers for receiving digital-scale values for each of multiple colors, and (c) a computational device operatively connected to the data file and display screen for (i) defining a plurality of rays extending between a selected source point and each of a plurality of points in a two-dimensional XY array, where (ia) the array points correspond to array pix... | GE Medical Systems | A61B005/05
US06852084 | B1 | 2005-02-08 | Wireless physiological pressure sensor and transmitter with capability of short range radio frequency transmissions | 2000-04-28 | A device and a method of sensing and transmitting physiological pressures and body temperatures are disclosed. The device includes a transducer and a transmitter. The transmitter is adapted to broadcast a signal which is modulated by the output of a transducer. The transmitter is also adapted to limit the power of the output signal. The method includes transducing a physiological parameter and broadcasting a signal which is modulated by the transduced parameter. The power of the output signal is limited so that the signal will attenuate within a predetermined distance. | 1. A new device for monitoring one of a plurality of physiological pressures associated with a patient, the device comprising: a housing for protecting the device, the housing having a first planar surface adapted for placement on a patient; at least one pressure transducer integral with the housing and adapted for placement on the patient, the at least one pressure transducer providing an output signal; a transmitter in operative communication with the at least one pressure transducer, and operatively attached to the housing, the transmitter adapted to broadcast a signal which is modulated by an output of the at least one pressure transducer; and a display secured to the housing and operatively connected to the at least one pressure transducer for displaying a representation of the output from the at least one pressure transducer, the display adapted for placement on the patient, a memory disposed within the housing and operatively connected to the at least one pressure transducer for storing an audio representation of the output signal; wherein the at least one pressure transducer includes a bell portion and a diaphragm portion disposed along said first surface, and the device includes a switch for selecting either an output of the bell or the diaphragm as the output signal. | 1. A new device for monitoring one of a plurality of physiological pressures associated with a patient, the device comprising: a housing for protecting the device, the housing having a first planar surface adapted for placement on a patient; at least one pressure transducer integral with the housing and adapted for placement on the patient, the at least one pressure transducer providing an output signal; a transmitter in operative communication with the at least one pressure transducer, and operatively attached to the housing, the transmitter adapted to broadcast a signal which is modulated by an output of the at least one pressure transducer; and a display secured to the housing and operatively connected to the at least one pressure transducer for displaying a representation of the output from the at least one pressure transducer, the display adapted for placement on the patient, a memory disposed within the housing and operatively connected to the at least one pressure transducer for storing an audio representation of the output signal; wherein the at least one pressure transducer includes a bell portion and a diaphragm portion disposed along said first surface, and the device includes a switch for selecting either an output of the bell or the diaphragm as the output signal. 2. The device of claim 1 further comprising: a receiver which can be tuned to receive a signal from the transmitter. 3. The device of claim 2, further comprising: a computer in operative communication with the receiver. 4. The device of claim 1, further comprising: a temperature sensor, wherein the switch selects the output of the bell portion, the output of the diaphragm portion, or an output of the temperature sensor as the output signal, and wherein the display is further adapted to display a representation of the output from the temperature sensor. 5. The device of claim 1, further comprising: a pulse sensing transducer, wherein the switch selects the output of the bell portion, the output of the diaphragm portion, or an output of the pulse sensing transducer as the output signal, and wherein the display is further adapted to display a representation of the output from the pulse sensing transducer. 6. The device of claim 1 wherein the transmitter is adapted to limit the power of the broadcast signal so that the signal will attenuate to at most a negligible value within a predetermined distance from the transmitter. | - | A61B005/02
US06892089 | B1 | 2005-05-10 | Cardiac motion tracking using cine harmonic phase (HARP) magnetic resonance imaging | 2000-04-14 | The present invention relates to a method of measuring motion of an object such as a heart by magnetic resonance imaging. A pulse sequence is applied to spatially modulate a region of interest of the object and at least one first spectral peak is acquired from the Fourier domain of the spatially modulated object. The inverse Fourier transform information of the acquired first spectral-peaks is computed and a computed first harmonic phase image is determined from each spectral peak. The process is repeated to create a second harmonic phase image from each second spectral peak and the strain is determined from the first and second harmonic phase images. In a preferred embodiment, the method is employed to determine strain within the myocardium and to determine change in position of a point at two different times which may result in an increased distance or reduced distance. The method may be employed to determine the path of motion of a point through a sequence of tag images depicting movement of the heart. The method may be employed to determine circumferential strain and radial strain. | 1. A method of measuring motion of an object by magnetic resonance imaging comprising applying a pulse sequence to spatially modulate a region of interest of said object; acquiring at least one first spectral peak from the Fourier domain of said spatially modulated object; computing the inverse Fourier transform information of said acquired first spectral peaks; computing a first harmonic phase image from each said spectral peak; repeating said process with respect to a different time to create a second harmonic phase image from each second spectral peak; and determining strain from said first and second harmonic phase images. | 1. A method of measuring motion of an object by magnetic resonance imaging comprising applying a pulse sequence to spatially modulate a region of interest of said object; acquiring at least one first spectral peak from the Fourier domain of said spatially modulated object; computing the inverse Fourier transform information of said acquired first spectral peaks; computing a first harmonic phase image from each said spectral peak; repeating said process with respect to a different time to create a second harmonic phase image from each second spectral peak; and determining strain from said first and second harmonic phase images. 2. The method of claim 1 wherein the object includes a heart. 3. The method of claim 2 wherein the step of determining strain comprises determining circumferential Lagrangian strain. 4. The method of claim 3 wherein determining strain includes determining at least one of endocardial, epicardial and midwall strain. 5. The method of claim 3 further including employing positive strain values as an indication of circumferential stretching and negative strain values as an indication of circumferential contraction. 6. The method of claim 2 wherein the determining strain includes determining radial Lagrangian strain. 7. The method of claim 6 determining at least one of endocardial and epicardial radial strain. 8. The method of claim 6 further including employing positive strain values as an indication of radial thickening and negative strain values as an indication of radial thinning. 9. The method of claim 2 wherein determining strain includes determining strain in the myocardium. 10. The method of claim 2 further including employing a refinement procedure in tracking said motion. 11. The method of claim 10 wherein employing a refinement procedure includes employing an anchor point which has relatively small motion as a tracking reference. 12. The method of claim 11 wherein employing a refinement procedure includes employing a sequence of points each of which is separated from the next adjacent point by less than one pixel. 13. The method of claim 12 wherein employing a refinement procedure includes effecting said refinement procedure in a generally radial path. 14. The method of claim 1 wherein the first harmonic phase image and the second harmonic phase image comprise two-dimensional images. 15. The method of claim 1 wherein the object includes a moving human heart. 16. The method of claim 15 wherein determining strain includes measuring motion within the left ventricular of said heart. 17. The method of claim 15 wherein the step of applying a pulse is initiated generally at end-diastole. 18. The method of claim 1 wherein applying a pulse sequence includes employing SPAMM pulse sequence. 19. The method of claim 1 wherein the first harmonic phase image and the second harmonic phase image include three-dimensional images. 20. The method of claim 1 wherein determining strain includes determining the spacing bet... | Johns Hopkins University | A61B005/055
US06911006 | B2 | 2005-06-28 | Pain inferring device and pain inferring method | 2001-04-24 | A pain inferring device for designing a shape giving a subject as little pain as possible when the subject touches the shape. The pain inferring device includes an input unit, an output unit, a main control unit, a learning storage unit, and a neural network. The neural network learns the relationship between the input value of shape data and the output value of the data on the degree of pain when a subject touches a shape. By the learning, an input/output function (namely, the coefficient of coupling of neurons in layers) representing the relationship between the input and output values is defined and stored in the learning storage unit. When the shape data is inputted through the input unit, the neural network infers the degree of pain by using the function stored in the learning storage unit. | 1. A pain inferring method, comprising the steps of: learning a relationship between shape date as an input and a degree of pain, as an output, felt by a person when he or she touches a shape corresponding to the shape data; producing an input-output function indicative of the relationship between the input and the output; and inferring, based on an input of new shape data, a degree of pain according to the input-output function. | 1. A pain inferring method, comprising the steps of: learning a relationship between shape date as an input and a degree of pain, as an output, felt by a person when he or she touches a shape corresponding to the shape data; producing an input-output function indicative of the relationship between the input and the output; and inferring, based on an input of new shape data, a degree of pain according to the input-output function. 2. The pain inferring method according to claim 1, wherein the learning step and the inferring step are performed using a neural network. 3. The pain inferring method according to claim 2, wherein the learning step includes learning the relationship between the input and the output in accordance with a back-propagation algorithm using the neural network. 4. The pain inferring method according to claim 3, further comprising the steps of: inputting predetermined shape data and a predetermined degree of pain; and terminating learning by the neural network, when an error between a degree of pain formed newly and the predetermined degree of pain reduces to a predetermined level or lower in the back-propagation algorithm. | Kabushiki Kaisha Tokai Rika Denki Seisakusho | A61B005/00
US06895617 | B2 | 2005-05-24 | Patient support system and elements thereof | 2002-08-20 | A patient support system for medical treatments is disclosed having an attached table section, an exchangeable table section and a fastener for the sections, the fastener including a sleeve, a rod and a cam-lock mechanism. The system also includes an easily operated and precisely located clamp for mounting accessory devices to the top of the table sections. The sections include peripheral channels and ridges which are used by the clamp to make a tight engagement. The channels also are used to form a coordinate number scale so that the patient, clamp and accessory devices may be consistently located time after time for multiple periodic treatments. | 1. A radiotherapy treatment and imaging patient support system comprising: a base; a patient support table connected to said base, said table having an edge, right and left side peripheral portions, and upper and lower peripheral portions, said peripheral portions being spaced away from said edge; location indicia formed on said table along one of said right and left peripheral portions, said indicia being a continuous sequence wherein a location on a body placed on said table is located relative to said table in a reproducible manner; and a clamp for mounting about said edge of said table, said clamp having an upper jaw portion and an opening therethrough, a lower jaw portion with an opening therethrough, a pivot shaft being operatively connected to said upper and said lower jaw portions through said openings, a lever connected to said upper jaw portion, a cam connected to said lever for bearing against said lower jaw portion, and an element having a hairline, said element being mounted in said opening in said upper jaw wherein viewing of said location indicia is facilitated. | 1. A radiotherapy treatment and imaging patient support system comprising: a base; a patient support table connected to said base, said table having an edge, right and left side peripheral portions, and upper and lower peripheral portions, said peripheral portions being spaced away from said edge; location indicia formed on said table along one of said right and left peripheral portions, said indicia being a continuous sequence wherein a location on a body placed on said table is located relative to said table in a reproducible manner; and a clamp for mounting about said edge of said table, said clamp having an upper jaw portion and an opening therethrough, a lower jaw portion with an opening therethrough, a pivot shaft being operatively connected to said upper and said lower jaw portions through said openings, a lever connected to said upper jaw portion, a cam connected to said lever for bearing against said lower jaw portion, and an element having a hairline, said element being mounted in said opening in said upper jaw wherein viewing of said location indicia is facilitated. 2. The apparatus of claim 1 including: a second location indicia mounted to said table along the other of said right and left peripheral portions, said first mentioned and said second location indicia being identical; third and fourth location indicia mounted to said table along said upper and lower peripheral portions, said third and fourth location indicia being identical; and four longitudinally extending channels formed in said table, one channel along each of said right, said left, said upper and said lower peripheral portions, and wherein said location indicia are formed in said channels. 3. The apparatus as claimed in claim 2 wherein: each of said channels includes two oblique side walls and a flat base, and each of said location indicia is formed on a respective base of each of said channels. 4. The apparatus as claimed in claim 3 including: two recesses in said upper jaw portion; two springs for mounting in said two recesses; and two locator elements for mounting in said two recesses and being biased by said two springs wherein said locator elements help locate said clamp relative to said patient support table. 5. The apparatus as claimed in claim 2 wherein: said patient support table has first and second sections, said first section being connected to said base and said second section having a middle portion; said middle portion includes an opening therethrough; and said middle portion includes oblique walls defining said opening. 6. The apparatus as claimed in claim 4 wherein: said patient support table has first and second sections, said first section being connected to said base and said second section having a middle portion; said middle portion includes an opening therethrough; and said middle portion includes oblique walls defining said opening. 7. The apparatus as claimed in claim 4 including: two upwardly extending projections connected to said... | Aktina Corp. | A61B005/05
US06845263 | B2 | 2005-01-18 | Heart-sound detecting apparatus and pulse-wave-propagation-velocity-relating-information obtaining system using the heart-sound detecting apparatus | 2001-07-02 | A heart-sound detecting apparatus, including: a heart-sound microphone which detects a plurality of heart sounds produced by a heart of a living subject and outputs a heart-sound signal representative of the detected heart sounds; a smoothing device for smoothing, by differentiation, a waveform of the heart-sound signal output from the heart-sound microphone; a squaring device for squaring an amplitude of the smoothed waveform with respect to a base line of the heart-sound signal; and a start-point determining device for determining a start point of a first heart sound I as one of the detected heart sounds, based on that the squared amplitude is greater than a prescribed threshold value. | 1. A system for obtaining information relating to a propagation velocity at which a pulse wave propagates along an artery of a living subject, the system comprising: a heart-sound detecting apparatus comprising: a heart-sound microphone which detects a plurality of heart sounds produced by a heart of a living subject and outputs a heart-sound signal representative of the detected heart sounds; a smoothing device for smoothing, by differentiation, a waveform of the heart-sound signal output from the heart-sound microphone; a squaring device for squaring an amplitude of the smoothed waveform with respect to a base line of the heart-sound signal; and a start-point determining device for determining a start point of a first heart sound I as one of the plurality of detected heart sounds, based on that the squared amplitude being greater than a prescribed threshold value; a pulse-wave detecting device which is adapted to be worn on the subject to detect the pulse wave which propagates along the artery of the subject; and a pulse-wave-propagation-velocity-relating-information obtaining device for obtaining information based on a time of the start point of the first heart sound I determined by the start-point determining device of the heart-sound detecting apparatus, and a time when a rising point of the pulse wave is detected by the pulse-wave detecting device. | 1. A system for obtaining information relating to a propagation velocity at which a pulse wave propagates along an artery of a living subject, the system comprising: a heart-sound detecting apparatus comprising: a heart-sound microphone which detects a plurality of heart sounds produced by a heart of a living subject and outputs a heart-sound signal representative of the detected heart sounds; a smoothing device for smoothing, by differentiation, a waveform of the heart-sound signal output from the heart-sound microphone; a squaring device for squaring an amplitude of the smoothed waveform with respect to a base line of the heart-sound signal; and a start-point determining device for determining a start point of a first heart sound I as one of the plurality of detected heart sounds, based on that the squared amplitude being greater than a prescribed threshold value; a pulse-wave detecting device which is adapted to be worn on the subject to detect the pulse wave which propagates along the artery of the subject; and a pulse-wave-propagation-velocity-relating-information obtaining device for obtaining information based on a time of the start point of the first heart sound I determined by the start-point determining device of the heart-sound detecting apparatus, and a time when a rising point of the pulse wave is detected by the pulse-wave detecting device. 2. The system according to claim 1, wherein the pulse-wave-propagation-velocity-relating-information obtaining device, comprises: a pulse-wave-propagation-time determining device for determining, based on the time of the start point of the first heart sound I, and the time when the rising point of the pulse wave is detected by the pulse-wave detecting device, a propagation time needed for the pulse wave to propagate from the heart to a position where the pulse-wave detecting device is worn on the subject. 3. The system according to claim 1, wherein the pulse-wave-propagation-velocity-relating-information obtaining device comprises: a pulse-wave-propagation-velocity determining device for determining the propagation velocity at which the pulse wave propagates, by dividing a distance from the heart to a position where the pulse-wave detecting device is worn on the subject by a time difference between the time of the start point of the first heart sound I and the time when the rising point of the pulse wave is detected by the pulse-wave detecting device. 4. The system according to claim 1, further comprising: an output device which outputs the information obtained by the pulse-wave-propagation-velocity-relating-information obtaining device, so that an observer can observe the output information. 5. The system according to claim 1, wherein said heart-sound detecting apparatus further comprises: a high-pass filter which passes a component of the heart-sound signal output from the heart-sound microphone, the component having frequencies which are not lower than a lowest signal-pass frequency of... | Colin Medical Technology Corporation | A61B005/02
US07069068 | B1 | 2006-06-27 | Method for determining haemodynamic indices by use of tomographic data | 2000-03-23 | Haemodynamic indices of an organ or a part of tissue are determined from a time series of tomographic data obtained by means of Magnetic Resonance Imaging. Maps of indices are produced, being significant of the dynamics of the capillary tissue flow acquired during rapid bolus injection of a tracer that stays mainly intravascular. The method may be used for evaluating the efficacy of a drug on an organ, or for obtaining information of the likelihood of recovery of an organ or part of tissue upon or during a period of insufficient vascular supply or during the progression of a chronic disease. The method may be used for discriminating between relevant therapy of an organ. | 1. Method for determining haemodynamic indices of an organ or of a part of tissue of a mammal including a) determining a time series of tomographic data pertaining to the organ or part of tissue during and after a bolus injection of a tracer dose to said mammal, the tracer being substantially intravascular in said tissue,b) determining a time series of concentration data being indicative of the concentration of the tracer in arteries of the organ or tissue from the time series of tomographic data,c) determining a residue function of the organ or of the part of tissue by deconvolution of the time series of tomographic data with the time series of concentration data,d) determining a distribution of transit times from the negative slope of the residue function, ande) determining a probability density function (PDF) of a haemodynamic index from the distribution of transit times. | 1. Method for determining haemodynamic indices of an organ or of a part of tissue of a mammal including a) determining a time series of tomographic data pertaining to the organ or part of tissue during and after a bolus injection of a tracer dose to said mammal, the tracer being substantially intravascular in said tissue,b) determining a time series of concentration data being indicative of the concentration of the tracer in arteries of the organ or tissue from the time series of tomographic data,c) determining a residue function of the organ or of the part of tissue by deconvolution of the time series of tomographic data with the time series of concentration data,d) determining a distribution of transit times from the negative slope of the residue function, ande) determining a probability density function (PDF) of a haemodynamic index from the distribution of transit times. 2. Method according to claim 1, wherein a probability density function (PDF) of a normalised haemodynamic index is determined from the distribution of transit times, the index being normalised by the value of the integral of said index. 3. Method according to claim 1, wherein at least one of the haemodynamic indices is a quantitative haemodynamic parameter obtained from the PDF. 4. Method according to claim 3, wherein at least one of the at least one parameter is obtain from comparison of the determined PDF and a previously determined reference PDF. 5. Method according to claim 4, wherein the parameter is obtained by use of the Kolmogorov Smirnov test. 6. Method according to claim 3 and comprising the steps of determining the impulse response function of the organ or of the part of tissue by deconvolution of the time series of tomographic data with the time series of concentration data,determining the relative tissue flow from the impulse response function of the organ or of the part of tissue,normalising said time series of concentration data with the integral of said time series of concentration data with respect to time,determining the normalised relative tissue flow, respectively the normalised blood volume, of the organ or part of tissue by use of the relative tissue flow and the time series of normalised concentration data, andconverting said normalised relative tissue flow, respectively normalised blood volume, to an absolute value for the tissue flow (Ft), respectively the blood volume, by means of a previously determined conversion factor,the quantitative haemodynamic parameter being of metabolic significance and determined from the PDF and the absolute tissue flow (Ft), respectively the absolute blood volume. 7. Method according to claim 6, wherein a parameter (E) significant for the local extraction of a substance is determined, the method further comprising the following steps: calculating the relative flow heterogeneity (w(f)) as a function of the relative flow (f) from the distribution of transit times,estimating a value (P) for the local capillary permeab... | - | A61B005/00
US07137952 | B2 | 2006-11-21 | Intraocular pressure recording system | 2001-06-29 | A new noninvasive approach for intraocular pressure (IOP) measurement allowing continuous monitoring over prolonged periods, regardless of patient's position and activities. The key element of this measurement method is a soft contact lens (1) including at least one strain gage (2) longitudinally arranged around the center of the contact lens and capable of measuring precisely spherical deformations of the eyeball induced by the changes in IOP. This information is transmitted with wires or (preferably) wirelessly in real time to an external recording system (14). The system is placed in the same way as a normal corrective contact lens, no anesthesia is required and patient vision remains almost completely unimpaired. | 1. Intraocular pressure recording system comprising a soft contact lens having a center and an active strain gage fixed to said contact lens, wherein said active strain gage is placed at a distance from the center of the contact lens, is not in direct contact with an eye, and comprises a portion having a circular arc shape which is situated around the center of said soft contact lens. | 1. Intraocular pressure recording system comprising a soft contact lens having a center and an active strain gage fixed to said contact lens, wherein said active strain gage is placed at a distance from the center of the contact lens, is not in direct contact with an eye, and comprises a portion having a circular arc shape which is situated around the center of said soft contact lens. 2. Intraocular pressure recording system according to claim 1 wherein said active strain gage is made of a resistive material. 3. Intraocular pressure recording system according to claim 2 wherein said active strain gage is a continuous longitudinal element. 4. Intraocular pressure recording system according to claim 3 wherein said continuous longitudinal element comprises portions and is placed in such a way that several of said portions are disposed parallel to each other. 5. Intraocular pressure recording system according to claim 3 wherein said continuous longitudinal element is microfabricated. 6. Intraocular pressure recording system according to claim 3 wherein said continuous longitudinal element is a wire. 7. Intraocular pressure recording system according to claim 1 further comprising a passive strain gage. 8. Intraocular pressure recording system according to claim 7 wherein said passive strain gage has a general circular arc shape made of a continuous meridional element comprising portions and is placed in such a way that several of said portions are radially arranged. 9. Intraocular pressure recording system according to claim 8 wherein said continuous meridional element is microfabricated. 10. Intraocular pressure recording system according to claim 8 wherein said continuous meridional element is a wire. 11. Intraocular pressure recording system according to claim 7 including a wireless telemetry system for data transmission with said passive strain gage. 12. Intraocular pressure recording system according to claim 7 including a wireless telemetry system for data transmission with said active and passive strain gages. 13. Intraocular pressure recording system according to claim 1 wherein said active strain gage is shaped in order to be placed on the comeosclera junction. 14. Intraocular pressure recording system according to claim 1 including a wireless telemetry system for data transmission with said active strain gage. 15. Intraocular pressure recording system according to claim 1 wherein said contact lens further comprises other measuring devices. 16. Intraocular pressure recording system according to claim 1 wherein the system comprises several active gages. 17. Intraocular pressure recording system according to claim 1 wherein the system comprises several passive gages. 18. Intraocular pressure recording system according to claim 1 wherein the system comprises gages in a Wheatstone bridge configuration. 19. Intraocular pressure recording system according to claim 18 wherein the Wheatstone bridge configuration comprises two active... | Ecole Polytechnique Federale de Lausanne-Service des Relations Industrielles | A61B003/16 A61B005/00 A61B005/103
US07094198 | B2 | 2006-08-22 | Delivery system and method for interstitial radiation therapy using seed elements with ends having one of projections and indentations | 2002-06-04 | A delivery system and method for interstitial radiation therapy uses a seed strand composed of a plurality of radioactive seeds or encapsulated seeds having ends with projections or indentations connected by a polymer. The seed strand is axially stiff and radially flexible and is bioabsorbable in living tissue. | 1. A method for manufacturing therapeutic elements, comprising the steps of: selecting a plurality of radioactive seed elements, each having a first end and a second end, wherein the first end has one of an indentation and a projection and the second end has one of an indentation and a projection;arranging said plurality of radioactive seed elements end-to-end with a desired spacing in between adjacent radioactive seed elements, said spacing customized to a patient's treatment prescription;introducing a liquid polymeric material into the spacing between the adjacent radioactive seed elements so that the liquid polymeric material engages said projections and said indentations of said first ends and said second ends without completely encapsulating said radioactive seed elements; andallowing said liquid polymeric material to solidify in engagement with said ends, thereby connecting said ends of said plurality of radioactive seed elements to form a treatment strand, said treatment strand customized to the patient's treatment prescription without the use of pre-fabricated spacers and said treatment strand having said plurality of radioactive seed elements connected by segments of the solidified polymeric material. | 1. A method for manufacturing therapeutic elements, comprising the steps of: selecting a plurality of radioactive seed elements, each having a first end and a second end, wherein the first end has one of an indentation and a projection and the second end has one of an indentation and a projection;arranging said plurality of radioactive seed elements end-to-end with a desired spacing in between adjacent radioactive seed elements, said spacing customized to a patient's treatment prescription;introducing a liquid polymeric material into the spacing between the adjacent radioactive seed elements so that the liquid polymeric material engages said projections and said indentations of said first ends and said second ends without completely encapsulating said radioactive seed elements; andallowing said liquid polymeric material to solidify in engagement with said ends, thereby connecting said ends of said plurality of radioactive seed elements to form a treatment strand, said treatment strand customized to the patient's treatment prescription without the use of pre-fabricated spacers and said treatment strand having said plurality of radioactive seed elements connected by segments of the solidified polymeric material. 2. The method of claim 1 wherein the introducing step further comprises using a co-polymer as the polymeric material. 3. The method of claim 1 wherein the introducing step further comprises using a natural bio-compatible material as the polymeric material. 4. The method of claim 1 wherein the introducing step further comprises using a synthetic bio-compatible material as the polymeric material. 5. The method of claim 1 wherein the introducing step further comprises using a natural bio-absorbable material as the polymeric material. 6. The method of claim 1 wherein the introducing step further comprises using a synthetic bio-absorbable material as the polymeric material. 7. The method of claim 1 wherein the introducing step includes forming the treatment strand that has a substantially constant diameter. 8. The method of claim 1 wherein the introducing step includes forming the treatment strand having a diameter substantially similar to a diameter of the plurality of said radioactive seed elements. 9. The method of claim 1 further comprising, before the introducing step, the step of selecting a polymeric material, which material, after the allowing step, provides said treatment strand with sufficient radial flexibility to maintain locational accuracy of said plurality of radioactive seed elements relative to a tumor target as said tumor target shrinks in size. 10. The method of claim 1 wherein the introducing step further comprises ensuring that the thickness of said treatment strand around said plurality of radioactive seed elements is sufficient to decrease normal tissue necrosis from a high local dose of radiation. 11. The method of claim 1 further comprising, before the introducing step, the step of selecting a polymeric materia... | Worldwide Medical Technologies, LLC Ideamatrix, Inc. | A61B005/00
US07006858 | B2 | 2006-02-28 | Implantable, retrievable sensors and immunosensors | 2002-08-09 | A sensor is disclosed, for implantation within a blood vessel to monitor an analyte in blood. In one embodiment, the sensor is retrievable. In another embodiment, the sensor is an immunosensor. A signal representative of a reaction between an analyte and an antigen of the analyte is transmitted to an external receiver. Methods are also disclosed. | 1. A method of using a sensor in a blood vessel, comprising the steps of providing a sensor on a tubular support having a streamlined configuration with respect to the tubular support and an analyte sensing surface thereon and positioning the sensor at a site in a blood vessel such that the sensing surface is positioned radially inwardly from the vessel wall, and removing the sensor during a catheterization process. | 1. A method of using a sensor in a blood vessel, comprising the steps of providing a sensor on a tubular support having a streamlined configuration with respect to the tubular support and an analyte sensing surface thereon and positioning the sensor at a site in a blood vessel such that the sensing surface is positioned radially inwardly from the vessel wall, and removing the sensor during a catheterization process. 2. A method as in claim 1, wherein the positioning step comprises carrying the sensor on a catheter and transluminally advancing the catheter to the site. 3. A method as in claim 1, wherein the positioning step further comprises inflating a balloon on the catheter at the site. 4. A method as in claim 1, wherein the positioning step further comprises removing a restraint from a self expandable sensor support. 5. A method as in claim 1, wherein the removing step further comprises guiding a retrieval snare to a snareable member on the sensor. 6. A method as in claim 1, wherein the removing step further comprises fluoroscopic guidance. | - | A61B005/00
US07107090 | B2 | 2006-09-12 | Devices and methods for presenting and regulating auxiliary information on an image display of a telesurgical system to assist an operator in performing a surgical procedure | 2002-12-05 | Systems and methods for performing robotically-assisted surgical procedures on a patient enable an image display device to provide an operator with auxiliary information related to the surgical procedure, in addition to providing an image of the surgical site itself. The systems and methods allow an operator to selectively access and reference auxiliary information on the image display device during the performance of a surgical procedure. | 1. A robotically controlled surgical system comprising: a surgical instrument;an image capture device positionable so as to capture an image of a surgical site;at least one image display coupled to the image capture device so as to display the image of the surgical sire;a master control input device;a source of auxiliary information; andat least one processor coupled to the surgical instrument, the image capture device, the at least one image display, the master control input device, and the source of auxiliary information, and coupling the master control input device to the surgical instrument and the source of auxiliary information so as to allow manipulation of the surgical instrument in response to manipulation of the master control input device while the master control input device is operatively associated with the surgical instrument, and allow manipulation of the auxiliary information on the image display in response to manipulation of the master control input device while the operative association between the master control input device and the surgical instrument is interrupted. | 1. A robotically controlled surgical system comprising: a surgical instrument;an image capture device positionable so as to capture an image of a surgical site;at least one image display coupled to the image capture device so as to display the image of the surgical sire;a master control input device;a source of auxiliary information; andat least one processor coupled to the surgical instrument, the image capture device, the at least one image display, the master control input device, and the source of auxiliary information, and coupling the master control input device to the surgical instrument and the source of auxiliary information so as to allow manipulation of the surgical instrument in response to manipulation of the master control input device while the master control input device is operatively associated with the surgical instrument, and allow manipulation of the auxiliary information on the image display in response to manipulation of the master control input device while the operative association between the master control input device and the surgical instrument is interrupted. 2. The system of claim 1, wherein the source of the auxiliary information is a data capturing/gathering device. 3. The system of claim 2, wherein the auxiliary information is captured in real-time. 4. The system of claim 1, wherein the source of the auxiliary information is a data storage device. 5. The system of claim 4, wherein the auxiliary information comprises preoperative information specific to a surgical procedure to be performed. 6. The system of claim 4, wherein the auxiliary information comprises preoperative information specific to the patient on which a surgical procedure is to be performed. 7. The system of claim 1, wherein the processor is configured to adjust an image of the auxiliary information in a discrete window on the image display. 8. The system of claim 7, wherein the adjusting of the auxiliary information in a discrete window on the image display further comprises varying the position of the discrete window across the image display. 9. The system of claim 7, wherein the adjusting of the auxiliary information in a discrete window on the image display further comprises varying the dimensions of the discrete window across the image display. 10. The system of claim 7, wherein the adjusting of the auxiliary information in a discrete window on the image display further comprises varying the orientation of the discrete window across the image display. 11. The system of claim 1, wherein the manipulation of the auxiliary information on the image display comprises displaying the auxiliary information on the image display such that the auxiliary information is overlaid on the image of the surgical site. 12. The system of claim 1, wherein the auxiliary information comprises an image corresponding to the surgical site, and wherein the manipulation of the auxiliary information on the image display involves alternately displaying the auxilia... | Intuitive Surgical | A61B005/00
US07079885 | B2 | 2006-07-18 | Method and device for instrument, bone segment, tissue and organ navigation | 2000-12-13 | The invention relates to a method and device for referencing between a data set, which describes geometrically the spatial model of a body, and the real physical environment in which the body is placed. A three-dimensional position reference body is used on the real body. The position reference body has one or more elementary bodies or markers whose 3-dimensional position can be detected with sensors and which define a fixed geometric reference with respect to the center of gravity of the body or to other body reference volumes. For registration, the position reference body and/or its elementary bodies are correlated in the data model and in the physical world. | 1. A method for instrument and bone segment as well as tissue and organ navigation, comprising the steps of: detecting position data of position reference bodies arranged on bones, tissue or organs and contour data and/or surface data of geometric figures arranged on or spatially correlated with the position reference bodies, and mathematically processing the data in such a way that the position data and the contour and/or surface data can be represented in a common coordinate system. | 1. A method for instrument and bone segment as well as tissue and organ navigation, comprising the steps of: detecting position data of position reference bodies arranged on bones, tissue or organs and contour data and/or surface data of geometric figures arranged on or spatially correlated with the position reference bodies, and mathematically processing the data in such a way that the position data and the contour and/or surface data can be represented in a common coordinate system. 2. The method according to claim 1, wherein the position reference bodies define a fixed geometric reference relative to the center of gravity of the body or other body reference volumes and that for the purpose of registration, the position reference body or its elementary bodies are associated with each other in a data model and in the physical world. 3. A device for instrument and bone segment as well as tissue and organ navigation, comprising a 3-D marker position detection unit for detecting signals from 3-D markers, the 3-D marker position detection unit being coupled to an optical 3-D scanner or radar unit for detecting surfaces; and a geometric figure arrangement being one or both of at least one geometric figures connected with each of said 3-D markers or at least one 3-D markers shaped as a geometric figures, wherein the geometric figure is detectrable by the 3-D scanner, thereby allowing a determination of the coordinates of the 3-D markers on the display and processing unit. 4. The device according to claim 3, wherein the display and processing unit enables referencing between an operating site and a patient data set through a computational surface fit andenables referencing between data of the 3-D position detection unit and data of the 3-D scanner based on the geometric figure arrangement. 5. The device according to claim 3, further comprising second 3-D markers for application on a bone segment of a patient or on an instrument for enabling a bone segment navigation or instrument navigation. 6. The device according to claim 3, wherein the geometric figures are applied to or on a frame connecting a plurality of 3-D markers. 7. The device according to claim 3, wherein a base, on which a 3-D marker is placed, is provided as the geometric figure for determining coordinates of the 3-D markers on the display and processing unit from 3-D scanner measurement data. 8. The device according to claim 6, wherein the frame forms the geometric figure for determining coordinates of the 3-D markers on the display and processing unit from 3-D scanner (5) measurement data. 9. The device according to claim 3, wherein at least one of said geometric figures is formed as one chosen from the group consisting of a recess/sulcus, a raised portion/crista (7b) and a sphere. 10. The device according to claim 3, wherein at least one of said geometric figures is color-coded or formed as a bar code. 11. The device according to claim 3, wherein at least two of the geometric ... | LB Medical GmbH | A61B005/05
US07235054 | B2 | 2007-06-26 | Measuring head for a gas analyser | 2002-10-25 | A measuring head for the analysis of respiratory gases to and from a patient connected to a respirator, includes an aperture (7) which is intended to be placed over an adapter (2) through which respiratory gases flow. The measuring head (1) has a light transmitter (9) which includes an IR-emitter (12) on one side of the aperture (7) and a light receiver (10) which includes an IR-detector on the other side of the aperture (7). The measuring head (1) also includes a signal processing unit (20) necessary for the gas analysis. | 1. A measuring head for the analysis of gasses present in a patient's breath, the measuring head adapted for placement in close proximity to the patient's mouth, comprising: a central aperture (7) connecting to a central portion (4) of a breath sampling adapter (2), the breath sampling adapter adapted for placement in close proximity to the patient's mouth to obtain a sampled breath from the patient's mouth;two mutually opposing surfaces (8) facing inwardly towards the aperture;a light transmitter (9) aligned with a light receiver (10) for transmitting and receiving infrared light respectively mounting on respective ones of the two opposing surfaces (8) such that light emitted by the light transmitter passes through the central portion of the adapter and the sampled breath to reach the light receiver;a connection for a fuel cell measuring an oxygen gas content of the sampled breath, the connection located proximate the central aperture and in communication with the sampled breath;a light transmitter board (13) connected to the light transmitter, the transmitter board including a microprocessor and an A/D-converter;a detector board (14), connected to the light receiver, comprising a motor drive with a rotatable output shaft;a rotatable filter wheel (16), comprising optical band-pass filters, carried on the shaft and situated so as to lie in an optical path between the light transmitter and the light receiver so that the infrared light will pass optical band-pass filters of the filter wheel prior to arriving at the light receiver to provide a detector board output signal based on the light passing through the optical band-pass filters; anda signal processing unit (20) operatively connected to the light transmitter board and the detector board output signal, the signal processing unit comprising a microprocessor, an analogue to digital converter processing the detector board output signal, and a calibration data memory (31), the signal processing unit outputting a digital signal (40) containing processed gas data of the multiple gasses present in the sampled breath for direct presentation on a presentation unit of a values system. | 1. A measuring head for the analysis of gasses present in a patient's breath, the measuring head adapted for placement in close proximity to the patient's mouth, comprising: a central aperture (7) connecting to a central portion (4) of a breath sampling adapter (2), the breath sampling adapter adapted for placement in close proximity to the patient's mouth to obtain a sampled breath from the patient's mouth;two mutually opposing surfaces (8) facing inwardly towards the aperture;a light transmitter (9) aligned with a light receiver (10) for transmitting and receiving infrared light respectively mounting on respective ones of the two opposing surfaces (8) such that light emitted by the light transmitter passes through the central portion of the adapter and the sampled breath to reach the light receiver;a connection for a fuel cell measuring an oxygen gas content of the sampled breath, the connection located proximate the central aperture and in communication with the sampled breath;a light transmitter board (13) connected to the light transmitter, the transmitter board including a microprocessor and an A/D-converter;a detector board (14), connected to the light receiver, comprising a motor drive with a rotatable output shaft;a rotatable filter wheel (16), comprising optical band-pass filters, carried on the shaft and situated so as to lie in an optical path between the light transmitter and the light receiver so that the infrared light will pass optical band-pass filters of the filter wheel prior to arriving at the light receiver to provide a detector board output signal based on the light passing through the optical band-pass filters; anda signal processing unit (20) operatively connected to the light transmitter board and the detector board output signal, the signal processing unit comprising a microprocessor, an analogue to digital converter processing the detector board output signal, and a calibration data memory (31), the signal processing unit outputting a digital signal (40) containing processed gas data of the multiple gasses present in the sampled breath for direct presentation on a presentation unit of a values system. 2. The measuring head of claim 1, the light transmitter and light receiver are aligned so that the light emitted by the light transmitter passes straight through the central portion of the adapter to reach the light receiver. 3. The measuring head of claim 1, further comprising: a digital input (41) to the signal processing unit for controlling calibration,an oxygen gas sensor (19), an ambient pressure sensor (23), and a system temperature sensor (25) connected to the signal processing unit via the analogue to digital converter. 4. The measuring head of claim 3, further comprising: a demultiplexer (28) connected to an output side of the A/D-converter, the demultiplexer processing the output of the detector board to obtain the current light intensity at different wavelengths of the filter wheel, the oxygen gas sensor sig... | Phase-In AB | A61B005/08 G01N001/22 G01N031/00 G01N033/497
US07052458 | B2 | 2006-05-30 | Medical diagnostic system | 2002-04-12 | A medical diagnostic system, for example, an X-ray examination apparatus, comprises a local group of functions which are only locally adjustable and a global group of functions which can be remotely adjusted. Preferably, functions in both groups can be remotely installed, but the functions in the local group can be activated only locally. Preferably, local activation is performed on the basis of an authorization. | 1. A medical diagnostic system comprising: a data processor which is controlled by functions including a local group of adjustable functions and a global group of adjustable functions;means for installing an adjustment to the local group of adjustable functions, andmeans for installing an adjustment to the global group of adjustable functions;local activating means for activating the installation of the adjustment to the local group of adjustable functions, said activating being performed only at the site of the medical diagnostic system; andglobal activating means for activating the installation of the adjustments to the global group of adjustable functions, said global activating being from one of: a remote service center or from the site of the medical diagnostic system. | 1. A medical diagnostic system comprising: a data processor which is controlled by functions including a local group of adjustable functions and a global group of adjustable functions;means for installing an adjustment to the local group of adjustable functions, andmeans for installing an adjustment to the global group of adjustable functions;local activating means for activating the installation of the adjustment to the local group of adjustable functions, said activating being performed only at the site of the medical diagnostic system; andglobal activating means for activating the installation of the adjustments to the global group of adjustable functions, said global activating being from one of: a remote service center or from the site of the medical diagnostic system. 2. A medical diagnostic system as claimed in claim 1, wherein the means for installing an adjustment to the local group of adjustable functions comprises an installation platform for adjusting the adjustable functions, the installation platform enabling the installation of remotely adjustable functions. 3. A medical diagnostic system as claimed in claim 2, further comprising an activation means by which the remotely installed functions are activated at the site of the medical diagnostic system on the basis of an authorization. 4. A medical diagnostic system as claimed in claim 3, in which the installation platform is provided with a read unit for receiving the authorization from a data carrier. 5. A medical diagnostic system as claimed in claim 1, further comprising authorization means by which the adjustments in the local group and/or the global group are activated on the basis of an authorization. | Koninklijke Philips Electronics N.V. | A61B005/00
US07090997 | B2 | 2006-08-15 | Diagnostic agent and test method for colon cancer using tannase as index | 2001-10-26 | A diagnostic agent for colon cancer, which comprises a reagent for detecting a tannase high-producing bacterium or measuring an amount of tannase contained in an intracolonic microflora sample, and a test method for colon cancer, which comprises the step of detecting a tannase high-producing bacterium or measuring an amount of tannase contained in an intracolonic microflora sample. | 1. A test method for colon cancer, which comprises the step of detecting Staphylococcus lugdunensis contained in an intracolonic microflora sample. | 1. A test method for colon cancer, which comprises the step of detecting Staphylococcus lugdunensis contained in an intracolonic microflora sample. | Eisai Co., Ltd. | C12Q001/14 A61B005/00 C12R001/44 C12N001/20 C12N009/18 C12Q001/44
US07003336 | B2 | 2006-02-21 | Analyte sensor method of making the same | 2001-02-08 | A sensor and method of making the same for implantation in a body that includes a substrate with notches cut in the substrate to form a necked down region in the substrate; and at least one sensor electrode formed from one or more conductive layers. Preferably, the thickness of the substrate ranges from approximately 25μ to 350μ, but the thickness of the substrate can range from 5μ to 750μ. The sensor may be incorporated in to a sensor assembly includes a slotted needle having a slot. The notches creating the necked down region allow the substrate to slide into the slotted needle, which that has the slot narrow enough to permit passage of the necked down region. However, a non-necked down region of the substrate is prevented from pulling out of the slotted needle through the slot. The slot of the slotted needle may also permit the necked down region of the substrate to slide down the slot. | 1. An analyte sensor for implantation in a body for use with a slotted needle, the sensor comprising: a substrate with notches cut in the substrate to form a necked down region in the substrate and a non-necked down region in the substrate wherein the non-necked down region is distal to the notch which forms the necked down region and is for placement within the body; andat least one sensor electrode formed from one or more conductive layers,wherein the notches creating the necked down region allow the substrate to slide into the slotted needle that has a slot narrow enough to permit passage of the necked down region, but prevents the non-necked down region of the substrate from slipping out of the slotted needle, andwherein the notches cut in the substrate do not expose any sensor electrode to analytes. | 1. An analyte sensor for implantation in a body for use with a slotted needle, the sensor comprising: a substrate with notches cut in the substrate to form a necked down region in the substrate and a non-necked down region in the substrate wherein the non-necked down region is distal to the notch which forms the necked down region and is for placement within the body; andat least one sensor electrode formed from one or more conductive layers,wherein the notches creating the necked down region allow the substrate to slide into the slotted needle that has a slot narrow enough to permit passage of the necked down region, but prevents the non-necked down region of the substrate from slipping out of the slotted needle, andwherein the notches cut in the substrate do not expose any sensor electrode to analytes. 2. The sensor in accordance with claim 1, wherein the thickness of the substrate ranges from approximately 25μ to 350μ. 3. The sensor in accordance with claim 1, wherein the thickness of the substrate ranges from 5μ to 750μ. 4. The sensor in accordance with claim 1, wherein a width of the substrate in the non-necked down region is sized to fit within a slotted needle having a diameter smaller than 21 gauge. 5. The sensor in accordance with claim 1, wherein a width of the substrate in the non-necked down region is sized to fit within a slotted needle having a diameter smaller than 22 gauge. 6. The sensor in accordance with claim 1, wherein a width of the substrate in the non-necked down region is sized to fit within a slotted needle having a diameter smaller than 23 gauge. 7. The sensor in accordance with claim 1, wherein a width of the substrate in the non-necked down region is sized to fit within a slotted needle having a diameter smaller than 24 gauge. 8. The sensor in accordance wit claim 1, wherein at least one of the at least one sensor electrode is formed on a first surface of the substrate. 9. The sensor in accordance with claim 8, wherein all of the at least one sensor electrode are only formed on the first surface. 10. The sensor in accordance with claim 8, wherein at least another one of the at least one sensor electrodes is formed on a second surface of the substrate. 11. The sensor in accordance with claim 10, wherein a third one of the at least one sensor electrode is a reference electrode configured to contact a skin surface. 12. A sensor set for use with the sensor in accordance with claim 1 comprising: a mounting base adapted for mounting onto a patient's skin; andan insertion needle carried by the mounting base to protrude from the mounting base ad having at least a portion of the sensor nested within the insertion needle, the insertion needle defining a longitudinally extending slot along one side to permit sliding withdrawal of the insertion needle from the mounting base and the nested portion of the sensor and to accept the necked down region of the substrate. 13. The sensor of claim 1, wherein the notches are cut ... | Medtronic MiniMed, Inc. | B65D085/38 A61B005/05 A61B005/04
US06895263 | B2 | 2005-05-17 | Real time self-adjusting calibration algorithm | 2002-05-08 | A method of calibrating glucose monitor data includes collecting the glucose monitor data over a period of time at predetermined intervals. It also includes obtaining at least two reference glucose values from a reference source that temporally correspond with the glucose monitor data obtained at the predetermined intervals. Also included is calculating the calibration characteristics using the reference glucose values and the corresponding glucose monitor data to regress the obtained glucose monitor data. And calibrating the obtained glucose monitor data using the calibration characteristics is included. In preferred embodiments, the reference source is a blood glucose meter, and the at least two reference glucose values are obtained from blood tests. In additional embodiments, the calculation of the calibration characteristics is obtained using linear regression and in particular embodiments, least squares linear regression. Alternatively, the calculation of the calibration characteristics is obtained using non-linear regression. | 1. A method of calibrating sensor data collected from a sensor, wherein a calibration formula based on past calibrations is used to interpret the sensor data, the method comprising the steps of: obtaining a calibration reference value for the sensor; calculating a current calibration factor based on the calibration reference value and a current sensor data point; identifying either a possible error or change in sensitivity in the sensor from the current calibration factor; and confirming a sensor failure or recognizing a change in sensor sensitivity by comparing the current calibration factor with an estimated value for the current calibration factor to determine whether the current calibration factor was unexpected, wherein the estimated value for the current calibration factor is determined from the calibration formula and a past calibration factor; and confirming the sensor failure when at least two unexpected calibration factors are received in succession without supporting each other. | 1. A method of calibrating sensor data collected from a sensor, wherein a calibration formula based on past calibrations is used to interpret the sensor data, the method comprising the steps of: obtaining a calibration reference value for the sensor; calculating a current calibration factor based on the calibration reference value and a current sensor data point; identifying either a possible error or change in sensitivity in the sensor from the current calibration factor; and confirming a sensor failure or recognizing a change in sensor sensitivity by comparing the current calibration factor with an estimated value for the current calibration factor to determine whether the current calibration factor was unexpected, wherein the estimated value for the current calibration factor is determined from the calibration formula and a past calibration factor; and confirming the sensor failure when at least two unexpected calibration factors are received in succession without supporting each other. 2. The method of claim 1, further comprising: updating the calibration formula with the calibration reference value if the sensor failure is not confirmed and no change in sensor sensitivity is recognized; and interpreting the collected sensor data using the updated calibration formula. 3. The method of claim 2, wherein the calibration formula is a modified regression method. 4. The method of claim 3, wherein the modified regression method is Gaussian regression method. 5. The method of claim 3, wherein the calibration formula weights past calibrations based on how recent the calibration was performed. 6. The method of claim 2, wherein the calibration is performed while obtaining the sensor data. 7. The method of claim 1, wherein the step of confirming a sensor failure or recognizing a change in sensor sensitivity further comprises: recognizing a change in sensor sensitivity when an unexpected past calibration factor is supported by a subsequent current calibration factor in succession; restarting the calibration formula calculation when a change in sensor sensitivity is recognized. 8. The method of claim 7, further comprising: interpreting the collected sensor data using the restarted calibration formula. 9. The method of claim 1, wherein the sensor is a glucose sensor. 10. The method of claim 1, wherein the current sensor data point is obtained by the steps of: sampling characteristic sensor data at a predetermined rate from a sensor over time; deriving at least one current sensor data point from the sampled characteristic sensor data at a predetermined memory storage rate. 11. The method of claim 1, wherein the current sensor data point is obtained by the steps of: sampling characteristic sensor data; deriving interval values by applying clipping limits and averaging the post-clipped sampled characteristic sensor data over a predetermined interval rate; and deriving at least one current sensor data point by averaging the derived interval valu... | Medtronic MiniMed, Inc. | A61B005/00
US06850797 | B2 | 2005-02-01 | Visceral fat meter | 2001-05-28 | A visceral fat determining device which is easy to operate at home and capable of providing information associated with visceral fat of a subject is disclosed. The visceral fat determining device is comprising means for inputting personal data representing vital characteristics of the subject, a data processing unit that can process various types of calculations, and a display portion capable of displaying the personal data and the calculation results. Quantitative information on the visceral fat of the subject can be obtained through the calculation based on the abdominal girth of the subject, which is a circumferential length of a torso of the subject. | 1. A visceral fat determining device comprising: input means for inputting personal data including an abdominal girth WL which is a circumferential length of a torso of a subject; data processing unit that stores the personal data and calculates quantitative information on abdominal visceral fat of the subject based on the personal data; and a display portion that displays the personal data and a result of the calculation performed by the data processing unit, wherein the quantitative information on the abdominal visceral fat of the subject is calculated based on the abdominal girth WL of the subject. | 1. A visceral fat determining device comprising: input means for inputting personal data including an abdominal girth WL which is a circumferential length of a torso of a subject; data processing unit that stores the personal data and calculates quantitative information on abdominal visceral fat of the subject based on the personal data; and a display portion that displays the personal data and a result of the calculation performed by the data processing unit, wherein the quantitative information on the abdominal visceral fat of the subject is calculated based on the abdominal girth WL of the subject. 2. The visceral fat determining device according to claim 1, wherein the quantitative information on the abdominal visceral fat is an amount of the abdominal visceral fat. 3. The visceral fat determining device according to claim 1, further comprising: body fat ratio measuring means for measuring a bioelectrical impedance Z of the subject through electrodes in contact with end portions of the subject's body and for calculating a body fat ratio FAT of the subject based on the measured bioelectrical impedance and the inputted personal data or part of the data, wherein the body fat ratio FAT obtained by the body fat ratio measuring means is displayed on the display portion. 4. The visceral fat determining device according to claim 1, wherein the abdominal girth WL is a circumferential length of an abdomen at 4th lumbar vertebrae of the subject. 5. The visceral fat determining device according to any one of claim 1, wherein size measuring means is provided for measuring the abdominal girth WL. 6. The visceral fat determining device according to claim 5, wherein the abdominal girth WL measured through the size measuring means is inputted into the data processing unit. 7. A visceral fat determining device comprising: input means for inputting personal data including an abdominal girth WL that is a circumferential length of a torso of a subject; a data processing unit that stores the personal data and calculates an estimated value of an abdominal visceral fat cross sectional area VA of the subject based on the personal data; and a display portion that displays the personal data and a value of the calculation performed by the data processing unit, wherein the data processing unit stores a first regression coefficient and a first regression constant of the abdominal girth WL which are obtained based on statistical analysis of correlation between actually measured values of the abdominal visceral fat cross sectional area VA obtained in abdominal tomography of human bodies of random samples and values of the abdominal girth WL of the human samples, wherein the estimated value of an abdominal visceral fat cross sectional area VA of the subject is obtained based on a value of the abdominal girth WL, the first regression coefficient of the abdominal girth WL of human, and the first regression constant. 8. The visceral fat determining device according to... | Yamoto Scale Company | A61B005/05
US06939304 | B2 | 2005-09-06 | Method and apparatus for non-invasively evaluating endothelial activity in a patient | 2001-10-22 | A method and apparatus for non-invasively evaluating endothelial activity in a patient, particularly for indicating the presence of an endothelial dysfunction condition, by applying an occluding pressure to a predetermined part of an arm or leg of the patient to occlude arterial blood flow therein; maintaining the occluding pressure for a predetermined time period; removing the occluding pressure after the elapse of the predetermined time period to restore arterial blood flow; monitoring a digit of the arm or leg by a digit-probe for changes in the peripheral arterial tone therein before and after the application of the occluding pressure to the arm or leg of the patient; and utilizing any detected changes in the peripheral arterial tone for evaluating endothelial activity in the patient. Particularly important advantages are provided when the occluding cuff is applied to the digit of the patient receiving the monitoring digit-probe, on the proximal side thereof with respect to the patient's heart, and also when a reference digit-probe is applied to another digit, not influenced by the occluding pressure, to compensate for shifts inherent to vascular beds. | 1. Apparatus for non-invasively evaluating endothelial activity in a patient, comprising: an occluding cuff for applying an occluding pressure to a predetermined part of an arm or leg of the patient to occlude blood flow therein for a predetermined time period; a monitoring digit-probe for monitoring a digit of said arm or leg for changes in the peripheral arterial tone therein before and after the application of said occluding pressure to the arm or leg of the patient; and a processor for utilizing any detected changes in said peripheral arterial tone for evaluating endothelial activity in the patient; wherein said processor is controlled by said monitoring digit-probe to indicate when an occluding pressure has been applied to said predetermined part of the patient's arm or leg, which is indicated when said monitoring digit-probe measures no pulsatile volume flow in said digit. | 1. Apparatus for non-invasively evaluating endothelial activity in a patient, comprising: an occluding cuff for applying an occluding pressure to a predetermined part of an arm or leg of the patient to occlude blood flow therein for a predetermined time period; a monitoring digit-probe for monitoring a digit of said arm or leg for changes in the peripheral arterial tone therein before and after the application of said occluding pressure to the arm or leg of the patient; and a processor for utilizing any detected changes in said peripheral arterial tone for evaluating endothelial activity in the patient; wherein said processor is controlled by said monitoring digit-probe to indicate when an occluding pressure has been applied to said predetermined part of the patient's arm or leg, which is indicated when said monitoring digit-probe measures no pulsatile volume flow in said digit. 2. The apparatus according to claim 1, wherein said processor utilizes the output of said monitoring digit-probe for providing a feedback signal to control the application of said occluding pressure. 3. A method for non-invasively evaluating endothelial activity in a patient, comprising: applying an occluding pressure to a predetermined part of an arm or leg of the patient to occlude arterial blood flow therein; maintaining said occluding pressure for a predetermined time period; removing said occluding pressure after the elapse of said predetermined time period to restore arterial blood flow; monitoring a digit of said arm or leg for changes in the peripheral arterial tone therein before, during, and after the application of said occluding pressure to the arm or leg of the patient; and utilizing any detected changes in said peripheral arterial tone, including changes in the dynamic signal time-course following removal of the occluding pressure, for evaluating endothelial activity in the patient; wherein said digit of the arm or leg of the patient is monitored for changes in the peripheral arterial tone by a monitoring digit-probe which is received on said digit and which measures peripheral arterial pulsatile flow therethrough while applying a non-occluding pressure to the outer extremity of the digit sufficiently high to prevent blood pooling in the veins and to unload the tension in the arterial walls, but not so high as to occlude the arteries; and wherein a reference digit-probe is applied to a digit of an arm or leg of the patient not to receive the occluding pressure for measuring changes in the peripheral arterial pulsatile flow therein, said latter changes being used for correcting the measured changes by the monitoring digit-probe to compensate for spontaneous short-term shifts in the peripheral arterial tone of local or systemic origin inherent to vascular beds and/or related to sympathetic nervous system activity due to painful stress resulting from the occlusion. 4. The method according to claim 3, wherein the measurements derived from said monitoring di... | Itamar Medical Ltd. | A61B005/02
US06896713 | B1 | 2005-05-24 | Liquid separator with holder unit | 2000-01-20 | A liquid separator for separating liquid from gases, comprises a water trap (1) that includes a container (3), a connector (5) for incoming gas flow, a separation chamber (4) that includes a filter and at least one connection passageway for leading separated gas to an analysis instrument. The water trap (1) can be removably fitted in a holder unit (2) connected to the analysis instrument, and the holder unit (2) includes connection devices (15, 16) for receiving the connection passageway. | 1. A liquid separator for separating liquid from gases, comprising: the combination of a water trap (1) and a holder unit (2),the water trap (1) including a container (3), a connection (5) for incoming gas flow, a separation chamber (4) that includes a filter (7), and at least one connection passageway (9, 10) for leading liquid-free gas to an analysis instrument via the holder unit,the holder unit being connectable to the analysis instrument,the water trap (1) removably fitted in an external cavity of the holder unit (2),the holder unit (2) provided with at least one connection device (15, 16) for accommodating the connection passageway (9, 10) within the external cavity, andthe holder unit (2) including a first electric contact element (18) which functions to detect the presence of a water trap (1) in the holder unit and to stop the flow of sample gas to the analysis instrument when no water trap is fitted in the holder unit,characterised in that the holder unit (2) includes a second electric contact element (17) which functions to detect the type of water trap (1) fitted in the holder unit and to adjust a flow rate of the analysis instrument in accordance with the type of water trap used. | 1. A liquid separator for separating liquid from gases, comprising: the combination of a water trap (1) and a holder unit (2),the water trap (1) including a container (3), a connection (5) for incoming gas flow, a separation chamber (4) that includes a filter (7), and at least one connection passageway (9, 10) for leading liquid-free gas to an analysis instrument via the holder unit,the holder unit being connectable to the analysis instrument,the water trap (1) removably fitted in an external cavity of the holder unit (2),the holder unit (2) provided with at least one connection device (15, 16) for accommodating the connection passageway (9, 10) within the external cavity, andthe holder unit (2) including a first electric contact element (18) which functions to detect the presence of a water trap (1) in the holder unit and to stop the flow of sample gas to the analysis instrument when no water trap is fitted in the holder unit,characterised in that the holder unit (2) includes a second electric contact element (17) which functions to detect the type of water trap (1) fitted in the holder unit and to adjust a flow rate of the analysis instrument in accordance with the type of water trap used. 2. A liquid separator according to claim 1, characterised in that the connection device (15, 16) is a quick-fastener device for connection to the connection passageway (9, 10). 3. A liquid separator according to claim 1, characterised in that the water trap (1) includes two connection passageways (9, 10), and in that the holder unit (2) includes two connection devices (15, 16). 4. A liquid separator according to claim 1, characterised in that the water trap (1) is designed in different sizes for infants and adults; and in that one size of the water trap includes a contact region to actuate the second electric contact element (17) of the holder unit. 5. A liquid separator, comprising: a holder unit (2); anda water trap (1) externally held by the holder unit,the holder unit with connections to an instrument used to analyze expiration gas,the holder unit configured to externally hold water traps of two different types,the water trap comprisinga separation chamber (4) provided with a connection (5) to receive a gas flow incoming from a patient,a container (3) located beneath the separation chamber,a liquid passageway (6) within the separation chamber and a filter (7) positioned above the passageway,an upper chamber part (8) located above the separation chamber (4) and connecting to an upper side of the filter,two connection passageways (9, 10), within the upper chamber part, connected via the filter to the separation chamber, the two connection passageways being connectable to the holder, and via the holder to the instrument, locking tabs fitted externally to the separation chamber and engageble with corresponding elements of the holder to externally snap-fit the water trap to the holder unit and allow a user to remove the water trap from the holder,the hol... | Artema Medical AB | A61M016/08 A61B005/08
US06878117 | B1 | 2005-04-12 | Handheld sensor for acoustic data acquisition | 2000-09-25 | An apparatus according to an embodiment of the present invention is provided for sensing acoustic signals. The apparatus includes a housing having an apertured posterior member, a sensing unit extending through the apertured posterior member for interfacing by contact with a patient, a cursor control for positioning a cursor on a display, the display located anteriorially on the housing, a button for fixing a position of the cursor in the display, indicating the position on the housing with respect to the patient, and a circuit for causing the cursor to move about the display in response to the cursor control, and for transmitting the acoustic signal. | 1. An apparatus for sensing an acoustic signal, comprising: a housing having an apertured posterior member; a sensing unit extending through the apertured posterior member for interfacing by contact with a patient; a cursor control for positioning a cursor on a display, the display located anteriorially on the housing; a button for fixing a position of the cursor in the display, indicating the position of the housing with respect to the patient; and a circuit for causing the cursor to move about the display in response to the cursor control, and for transmitting the acoustic signal. | 1. An apparatus for sensing an acoustic signal, comprising: a housing having an apertured posterior member; a sensing unit extending through the apertured posterior member for interfacing by contact with a patient; a cursor control for positioning a cursor on a display, the display located anteriorially on the housing; a button for fixing a position of the cursor in the display, indicating the position of the housing with respect to the patient; and a circuit for causing the cursor to move about the display in response to the cursor control, and for transmitting the acoustic signal. 2. The apparatus of claim 1, wherein the sensing unit is passive. 3. The apparatus of claim 1, wherein the sensing unit is mounted to a pod by a mounting spring for reducing a motion artifact in the acoustic signal. 4. The apparatus of claim 1, wherein the cursor control is a rocker switch having two axes of motion. 5. The apparatus of claim 1, wherein the acoustic signal is transmitted via one of a cable and a wireless transmitter to a data processing system. 6. The apparatus of claim 1, wherein the circuit further comprises a pre-amplifier circuit for amplifying the acoustic signal. 7. An apparatus for sensing an acoustic signal comprising: a housing having an apertured posterior member; a sensing unit for interfacing by contact with a patient and capturing the acoustic signal, wherein the sensing unit is supported by a mounting apparatus for reducing a motion artifact in the acoustic signal, wherein the mounting apparatus is connected to the housing; and a telemetry sensor for indicating a position of the housing with respect to the patient. 8. The apparatus of claim 7, wherein the sensing unit is passive. 9. The apparatus of claim 8, wherein the sensing unit is flush mounted with the housing. 10. The apparatus of claim 7, wherein the mounting apparatus is one of a spring, and a rubber mount. 11. The apparatus of claim 10, wherein the sensing unit extends through the apertured posterior member. 12. The apparatus of claim 7, wherein the acoustic signal is transmitted via one of a cable and a wireless transmitter to a data processing system. 13. An apparatus for generating an electrical signal representative of an acoustic signal of a cardiovascular system comprising: a sensing unit for detecting the acoustic signal and converting the acoustic signal to the electrical signal, wherein the sensing unit is supported by a mounting apparatus for reducing a motion artifact in the acoustic signal, wherein the mounting apparatus is connected to a pod housing; a preamplifier for amplifying the electrical signal; and a telemetry unit for transmitting positional data and the electrical signal to a data processing system. 14. The apparatus of claim 13, further comprising: a liquid crystal display located anteriorly on the pod housing the sensor, the pre-amplifier, and the telemetry unit, a rocker switch for positioning a cursor in the display, indicating a positi... | Zargis Medical Corp. | A61B005/02
US06951639 | B2 | 2005-10-04 | 1,4,7,10-tetraazacyclododecane butyltriols, processes for their production and pharmaceutical agents containing them | 2002-04-09 | 1,4,7,10-Tetraazacyclodedecane butyltriols of general formula IA in which R1 means hydrogen or a metal ion equivalent independent of one another andR2 means a butyltriol radicalas well as their salts with organic or inorganic bases or amino acids are valuable pharmaceutical agents. | 1. A method of performing in vivo NMR spectroscopy in a patient comprising administering a compound as a susceptibility agent which has the formula IA in which each R1 is hydrogen or a metal ion equivalent, independent of one another and R2 is butyltriol or a salt thereof with an organic or inorganic base or an amino acid, and at least two of the substitutens R1 are metal ion equivalents. | 1. A method of performing in vivo NMR spectroscopy in a patient comprising administering a compound as a susceptibility agent which has the formula IA in which each R1 is hydrogen or a metal ion equivalent, independent of one another and R2 is butyltriol or a salt thereof with an organic or inorganic base or an amino acid, and at least two of the substitutens R1 are metal ion equivalents. 2. A method of radiation therapy of a patient comprising administering as a photon and electron emitting agent, a compound which has the formula IA in which each R1 is hydrogen or a metal ion equivalent, independent of one another and R2 is butyltriol or a salt thereof with an organic or inorganic base or an amino acid, and at least two of the substitutens R1 are metal ion equivalents. | Schering AG | A61K051/00 A61B005/055
US06980851 | B2 | 2005-12-27 | Method and apparatus for determining changes in heart failure status | 2001-11-15 | A method for operating a cardiac rhythm management device in which a clinical state vector is computed as a combination of a plurality of parameters related to a patient's heart failure status and compared to a previously computed clinical state vector to determine a clinical trajectory indicative of changes in the patient's heart failure status. Such detected changes in status can be used both as a clinical tool to evaluate treatment and to automatically adjust the operation of the device. | 1. A method for delivering pacing therapy to a heart failure patient, comprising: operating an implantable cardiac rhythm management device that generates sensing signals from sensed cardiac activity and delivering cardiac pacing therapy through one or more pacing channels to one or more heart chambers;computing a clinical state vector as a combination of a plurality of parameters related to a patient's heart failure status, wherein the plurality of parameters define an n-dimensional vector space with n being the number of parameters and with each parameter mapped to an ordinal scale that represents a coordinate axis in the n-dimensional vector space, and further wherein the n parameters include at least one parameter derived from a sense signal; and,computing a difference vector between the computed clinical state vector and a previously computed state vector, wherein the magnitude of the difference vector indicates the extent of change in the patient's heart failure status and the direction of the difference vector indicates whether the patient's heart failure status is improving or worsening. | 1. A method for delivering pacing therapy to a heart failure patient, comprising: operating an implantable cardiac rhythm management device that generates sensing signals from sensed cardiac activity and delivering cardiac pacing therapy through one or more pacing channels to one or more heart chambers;computing a clinical state vector as a combination of a plurality of parameters related to a patient's heart failure status, wherein the plurality of parameters define an n-dimensional vector space with n being the number of parameters and with each parameter mapped to an ordinal scale that represents a coordinate axis in the n-dimensional vector space, and further wherein the n parameters include at least one parameter derived from a sense signal; and,computing a difference vector between the computed clinical state vector and a previously computed state vector, wherein the magnitude of the difference vector indicates the extent of change in the patient's heart failure status and the direction of the difference vector indicates whether the patient's heart failure status is improving or worsening. 2. The method of claim 1 wherein the computation of the clinical state vector is performed by an external programmer. 3. The method claim 1 wherein the computation of the clinical state vector is performed by a controller of the cardiac rhythm management device. 4. The method of claim 1 wherein the parameter derived from a sense signal corresponds to a PR interval in an electrogram. 5. The method of claim 1 wherein the parameter derived from a sense signal corresponds to a QRS duration in an electrogram. 6. The method of claim 1 wherein the parameter derived from a sense signal corresponds to an inter-ventricular delay between senses in the right and left ventricles. 7. The method of claim 1 wherein the plurality of parameters includes a frequency of atrial fibrillation occurrence over a specified period of time. 8. The method of claim 1 wherein the plurality of parameters includes a frequency at which a ventricular tachycardia converts to ventricular fibrillation over a specified period of time. 9. The method of claim 1 wherein the plurality of parameters includes a measure of heart rate variability. 10. The method of claim 1 wherein the plurality of parameters includes a measured body weight of the patient. 11. The method of claim 1 wherein the plurality of parameters includes a measured or derived left ventricular end diastolic pressure. 12. The method of claim 1 wherein the plurality of parameters includes a measured or derived systolic pressure index. 13. The method of claim 1 wherein the plurality of parameters includes a measured or derived pulse pressure index. 14. The method of claim 1 wherein the plurality of parameters includes a measured or derived maximum left ventricular dP/dt index. 15. The method of claim 1 wherein the plurality of parameters includes a frequency of ectopic beats over a specified period of time. 16. The meth... | Cardiac Pacemakers, Inc. | A61B005/0402
US06977068 | B1 | 2005-12-20 | Method for detection of fibrin clots | 1998-07-23 | A method for the detection of fibrin in a source, in particular the in vivo detection of a fibrin in a patient, the method comprising supplying to the source or patient an amount of a detectable reagent comprising a plurality of discrete particles, each of the particles comprising a plurality of layers of carbon and being capable of binding to fibrin; and detecting the presence of the particles in the source. The particles may also comprise a detectable marker encased in said plurality of layers of carbon, the presence of said marker being capable of detection in said source. | 1. A method for the in vivo detection of fibrin, said method comprising the steps of: administering to said patient an effective amount of a detectable reagent comprising discrete diagnostic particles dispersed in a pharmaceutically or veterinarily acceptable carrier, diluent, excipient, adjuvant or any combination thereof, wherein said diagnostic particles comprise a detectable marker encased in at least two layers of carbon, wherein upon administration of said reagent said diagnostic particles are dispersed in the aqueous medium and form a stable colloid and wherein said particles are made by heating a carbon crucible having deposited thereon a detectable marker to a temperature in the range of 2250° C. to 3000° C. in an inert gas and in a sealed container, thereby generating particles suspended in said inert gas, and precipitating said particles suspended in said inert gas to form said diagnostic particles;binding said diagnostic particles to said fibrin, wherein said diagnostic particles exhibit a specific affinity for said fibrin; anddetecting the presence of said detectable marker in said patient. | 1. A method for the in vivo detection of fibrin, said method comprising the steps of: administering to said patient an effective amount of a detectable reagent comprising discrete diagnostic particles dispersed in a pharmaceutically or veterinarily acceptable carrier, diluent, excipient, adjuvant or any combination thereof, wherein said diagnostic particles comprise a detectable marker encased in at least two layers of carbon, wherein upon administration of said reagent said diagnostic particles are dispersed in the aqueous medium and form a stable colloid and wherein said particles are made by heating a carbon crucible having deposited thereon a detectable marker to a temperature in the range of 2250° C. to 3000° C. in an inert gas and in a sealed container, thereby generating particles suspended in said inert gas, and precipitating said particles suspended in said inert gas to form said diagnostic particles;binding said diagnostic particles to said fibrin, wherein said diagnostic particles exhibit a specific affinity for said fibrin; anddetecting the presence of said detectable marker in said patient. 2. The method of claim 1 wherein said particles suspended in said inert gas are precipitated using an electrostatic precipitator. 3. The method according to claim 1, wherein the carrier is an aqueous solution. 4. The method according to claim 1, wherein the aqueous solution is 5% glucose in water. 5. The method according to claim 1, wherein a surface of said diagnostic particles is coated with a surfactant coating, thereby generating surfactant coated diagnostic particles, and wherein said surfactant coating increases the binding efficiency of said surfactant coated diagnostic particles with fibrin relative to diagnostic particles not having said surfactant coating. 6. The method of claim 1 wherein said diagnostic particles form a nanocolloid upon administration of said detectable reagent. 7. The method of claim 5, wherein said surfactant coating comprises C16EO6. 8. A method for the in vivo detection of fibrin present in the bloodstream of a subject, said method comprising the steps of: administering to the bloodstream of said subject an effective amount of a detectable reagent comprising discrete diagnostic particles dispersed in a pharmaceutically or veterinarily acceptable carrier, diluent, excipient, adjuvant or any combination thereof, wherein said diagnostic particles comprise a detectable marker encased in about 2 to 20 layers of carbon, wherein said diagnostic particles are made by heating a carbon crucible having deposited thereon a detectable marker to a temperature in the range of 2250° C. to 3000° C., thereby generating suspended particles, and precipitating said suspended particles to form said diagnostic particles, wherein upon administration of said reagent said diagnostic particles are dispersed in the aqueous medium and form a stable colloid;binding said diagnostic particles to said fibrin, wherein said diagnostic particle... | The Australian National University | A61K051/00 A61K049/00 A61M036/14 A61B005/055
US06872380 | B2 | 2005-03-29 | Process for preparation of MR contrast agents | 2001-11-16 | The present invention provides a process for the preparation of an MR contrast agent, said process comprising: i) obtaining a solution in a solvent of a hydrogenatable, unsaturated substrate compound and a catalyst for the hydrogenation of said substrate compound;ii) introducing said solution in droplet form into a chamber containing hydrogen gas (H2) enriched in para-hydrogen (p-1H2) and/or ortho-deuterium (o-2H2) whereby to hydrogenate said substrate to form a hydrogenated imaging agent;iii) optionally subjecting said hydrogenated imaging agent to a magnetic field having a field strength below earth's ambient field strength;iv) optionally dissolving said imaging agent in an aqueous medium;v) optionally separating said catalyst from the solution of said imaging agent in said aqueous medium;vi) optionally separating said solvent from the solution of said imaging agent in said aqueous medium; andvii) optionally freezing the solution of said imaging agent in said aqueous medium. | 1. A process for the preparation of an MR contrast agent comprising: i) obtaining a solution in a solvent of a hydrogenatable, unsaturated substrate compound and a catalyst for the hydrogenation of said substrate compound; and ii) introducing said solution in droplet form into a chamber containing hydrogen gas (H2) enriched in para-hydrogen (p-1H2) and/or ortho-deuterium (o-2H2) to hydrogenate said substrate to form a hydrogenated imaging agent. | 1. A process for the preparation of an MR contrast agent comprising: i) obtaining a solution in a solvent of a hydrogenatable, unsaturated substrate compound and a catalyst for the hydrogenation of said substrate compound; and ii) introducing said solution in droplet form into a chamber containing hydrogen gas (H2) enriched in para-hydrogen (p-1H2) and/or ortho-deuterium (o-2H2) to hydrogenate said substrate to form a hydrogenated imaging agent. 2. The process of claim 1 further comprising subjecting said hydrogenated imaging agent to a magnetic field having a field strength at or below the ambient magnetic field strength of the earth. 3. The process of claim 1 further comprising dissolving said imaging agent in an aqueous medium. 4. The process of claim 3, further comprising separating said catalyst from said solution of imaging agent in aqueous medium. 5. The process of claim 3, further comprising separating said solvent from said solution of imaging agent in aqueous medium. 6. The process of claim 3, further comprising freezing said solution of imaging agent in aqueous medium. 7. The process of claim 2 wherein said field strength in step (iii) is less than 50 μT. 8. The process of claim 2 wherein said field strength in step (iii) is less than 1 μT. 9. The process of claim 2 wherein said field strength in step (iii) is less than or equal to 0.1 μT. 10. The process of claim 2 wherein said field strength in step (iii) is cycled in a first part from earth's ambient field strength to a field strength less than 0.1 μT, and in a second part back to ambient field strength again. 11. The process of claim 10 wherein the first part of the cycle is approximately ≦1 ms and the second part is approximately 10-10000 ms. 12. The process of claim 1 wherein said process is carried out directly in water and wherein both said substrate and said catalyst are water-soluble. | Amersham Health AS | A61B005/055 C07C069/02
US06961604 | B1 | 2005-11-01 | Wireless controller and application interface for an MRI system | 2001-10-09 | A wireless remote control unit (60) that operates in the radio frequency bandwidth is used for interfacing with a sequence control system (B) and an image processing system (C) from within a magnetic resonance suite (A) in the presence of a magnetic field produced by a main magnet assembly (10). The sequence control system and the image processing system are connected with a first wireless transceiver (30) whose antenna (28) is located within a magnetic resonance suite. The first transceiver (30) communicates with the remote control unit (60) and a second transceiver (32) connected with a transmitter (40) and gradient coil amplifiers (34). Resonance signals received by a radio frequency coil (46) are communicated to the first transceiver (30) by the second transceiver (32) or by a transceiver (46′) mounted on the receiving coil. The receiver coil transceiver also engages in a handshake protocol to identify itself. | 1. A magnetic resonance apparatus comprising: a magnet assembly for generating a main magnetic field through a subject disposed in an examination region, the magnet being disposed in a magnetic resonance suite;a sequence control system for generating magnetic resonance sequences;an RF coil which at least receives resonance signals from the examination region, the RF coil being disposed adjacent the examination region;an image processing system which processes the resonance signals received by the RF coil into images and manipulates the reconstructed images;a wireless interface disposed with an antennae in the magnetic resonance suite for wireless communication between a handheld wireless remote control unit and at least one of the sequence control system and the image processing system;the handheld wireless remote control unit including: an image display for displaying images and information from at least one of the sequence control system and the image processing system to an operator;an RF receiver for receiving RF signals via the wireless interface from at least one of the sequence control system and the image processing system which carry the images and information for display on the image display;an input portion for accepting requests from the operator; an RF transmitter for transmitting the operator requests by the wireless interface to at least one of the sequence control system and the image processing system. | 1. A magnetic resonance apparatus comprising: a magnet assembly for generating a main magnetic field through a subject disposed in an examination region, the magnet being disposed in a magnetic resonance suite;a sequence control system for generating magnetic resonance sequences;an RF coil which at least receives resonance signals from the examination region, the RF coil being disposed adjacent the examination region;an image processing system which processes the resonance signals received by the RF coil into images and manipulates the reconstructed images;a wireless interface disposed with an antennae in the magnetic resonance suite for wireless communication between a handheld wireless remote control unit and at least one of the sequence control system and the image processing system;the handheld wireless remote control unit including: an image display for displaying images and information from at least one of the sequence control system and the image processing system to an operator;an RF receiver for receiving RF signals via the wireless interface from at least one of the sequence control system and the image processing system which carry the images and information for display on the image display;an input portion for accepting requests from the operator; an RF transmitter for transmitting the operator requests by the wireless interface to at least one of the sequence control system and the image processing system. 2. The magnetic resonance apparatus as set forth in claim 1, wherein the wireless interface and the remote control unit communicate with carrier frequencies greater than 500 MHZ. 3. The magnetic resonance apparatus as set forth in claim 2, wherein the carrier frequencies are between 2.3 and 2.6 GHz. 4. A magnetic resonance apparatus comprising: a magnet assembly for generating a main magnetic field through a subject disposed in an examination region, the magnet being disposed in a magnetic resonance suite;a sequence control system for generating magnetic resonance sequences;an RF coil which at least receives resonance signals from the examination region, the RF coil being disposed adjacent the examination region;an image processing system which processes the resonance signals received by the RF coil into images and manipulates the reconstructed images;a wireless interface disposed with an antenna in the magnetic resonance suite for wireless communication between (i) at least one of the sequence control systems and the image processing system and (ii) a wireless remote control unit for communicating information to an operator from at least one of the sequence control system and the image processing system, the wireless remote control unit including: a radio frequency receiver for receiving radio frequency signals from the wireless interface. 5. The magnetic resonance apparatus as set forth in claim 4, further including: a microprocessor for processing operator input to the remote control unit. 6. A magnetic resonance apparatus ... | Koninklijke Philips Electroncis N.V. | A61B005/055
US06916283 | B2 | 2005-07-12 | System and method for assessing urinary function | 2002-06-27 | A system and method for assessing urinary function is provided. The system includes a processor, a tubing assembly forming a first fluid conduit between a first fluid inlet and a first fluid outlet, and an insert member having a channel therethrough and coupled to the first fluid outlet so that fluid flowing through the first fluid conduit may flow through the insert member. The insert member is dimensioned for at least partial insertion into a patient's urethral canal distal of the patient's urethral sphincter, and dimensioned so that, when inserted, it substantially blocks fluid flow into or out of the urethral canal other than through the insert member channel. The system further includes a fluid delivery device electrically coupled to and controlled by the processor, and coupled to the first fluid conduit for pumping fluid therethrough, and a pressure detection system for detecting pressure within the urethral canal distal of the urethral sphincter and providing information correlating to the detected pressure to the processor. The method includes coupling the test module to the control device, inserting the insert member into a patient, infusing fluid through the insert member until the urethral sphincter opens, and measuring pressure within the urethral canal as fluid is infused therein. Also provided is a test module for coupling to a control device to form a system for assessing urinary function. | 1. A system for assessing urinary function comprising: a processor; a tubing assembly forming a first fluid conduit between a first fluid inlet and a first fluid outlet; an insert member having a channel therethrough and coupled to the first fluid outlet so that fluid flowing through the first fluid conduit may flow through the insert member, the insert member being dimensioned for at least partial insertion into a patient's urethral canal distal of the patient's urethral sphincter, and being dimensioned so that, when inserted, it substantially blocks fluid flow into or out of the urethral canal other than through the insert member channel; a fluid delivery device electrically coupled to and controlled by the processor, and coupled to the first fluid conduit for pumping fluid therethrough; and a detection system for detecting resistance of the urethral sphincter as fluid is infused therein. | 1. A system for assessing urinary function comprising: a processor; a tubing assembly forming a first fluid conduit between a first fluid inlet and a first fluid outlet; an insert member having a channel therethrough and coupled to the first fluid outlet so that fluid flowing through the first fluid conduit may flow through the insert member, the insert member being dimensioned for at least partial insertion into a patient's urethral canal distal of the patient's urethral sphincter, and being dimensioned so that, when inserted, it substantially blocks fluid flow into or out of the urethral canal other than through the insert member channel; a fluid delivery device electrically coupled to and controlled by the processor, and coupled to the first fluid conduit for pumping fluid therethrough; and a detection system for detecting resistance of the urethral sphincter as fluid is infused therein. 2. The system according to claim 1, wherein the detection system is a pressure detection system for detecting pressure within the urethral canal distal of the urethral sphincter and providing information correlating to the detected pressure to the processor. 3. The system according to claim 2, wherein the pressure detection system is in fluid communication with the first fluid conduit at a location such that pressure within the first fluid conduit substantially correlates to pressure within the urethral canal distal of the urethral sphincter. 4. The system according to claim 3, wherein the pressure detection system further comprises a pressure interface in fluid communication with the first fluid conduit, and a pressure sensor electrically coupled to the processor, wherein the pressure interface is in contract with the pressure sensor and provides pressure information thereto. 5. The system according to claim 1, wherein the pressure detection system detects Urethral Resistance Pressure. 6. The system according to claim 1, wherein the fluid delivery device is a peristaltic pump. 7. The system according to claim 1, further comprising a fluid source coupled to the fluid inlet, wherein the fluid delivery device delivers fluid from the fluid source, through the first fluid conduit and insert member channel and into the patient's urethral canal. 8. The system according to claim 1, further comprising a display device electrically coupled to the processor and capable of displaying data. 9. The system according to claim 8, wherein pressure detected by the pressure detection system is displayed on the display device. 10. The system according to claim 9, wherein the pressure detection system further comprises a pressure interface in fluid communication with the first fluid conduit, and a pressure sensor electrically coupled to the processor, wherein the pressure interface is in contact with the pressure sensor and provides pressure information thereto. 11. The system according to claim 1, wherein the pressure detection system detects Urethral Resistance Pressu... | Ethicon, Inc. | A61F002/00 A61M005/32 A61B005/00
US06843772 | B2 | 2005-01-18 | Inferior-and-superior-limb blood-pressure-index measuring apparatus | 2002-11-21 | An apparatus for measuring an inferior-and-superior-limb blood-pressure index of a patient, including an inferior-limb blood-pressure measuring device, a superior-limb blood-pressure measuring device, an index determining device for determining the inferior-and-superior-limb blood-pressure index, based on a blood pressure of an inferior limb measured by the inferior-limb blood-pressure measuring device and a blood pressure of a superior limb measured by the superior-limb blood-pressure measuring device, an inferior-limb-pulse-wave-propagation-velocity-related-information, obtaining device which obtains inferior-limb-pulse-wave-propagation-velocity-related information, an upper-half-body-pulse-wave-propagation-velocity-related-information obtaining device which obtains upper-half-body-pulse-wave-propagation-velocity-related information, and an evaluation-information obtaining device for obtaining, based on the inferior-limb-pulse-wave-propagation-velocity-related information and the upper-half-body-pulse-wave-propagation-velocity-related information, evaluation information that is related to evaluation of reliability of the inferior-and superior-limb blood-pressure index. | 1. An apparatus for measuring an inferior-and-superior-limb blood-pressure index of a living subject, comprising: an inferior-limb blood-pressure measuring device which includes an inferior-limb cuff adapted to be wound around an inferior limb of the subject and measures a blood pressure of the inferior limb; a superior-limb blood-pressure measuring device which includes a superior-limb cuff adapted to be wound around a superior limb of the subject and measures a blood pressure of the superior limb; an inferior-and-superior-limb-blood-pressure-index determining means for determining the inferior-and-superior-limb blood-pressure index of the subject, based on the blood pressure of the inferior limb measured by the inferior-limb blood-pressure measuring device and the blood pressure of the superior limb measured by the superior-limb blood-pressure measuring device; an inferior-limb-pulse-wave-propagation-velocity-related-information obtaining device which obtains inferior-limb-pulse-wave-propagation-velocity-related information that is related to a velocity at which an inferior-limb pulse wave propagates in a first portion of the subject that includes the inferior limb; an upper-half-body-pulse-wave-propagation-velocity-related-information obtaining device which obtains upper-half-body-pulse-wave-propagation-velocity-related information that is related to a velocity at which an upper-half-body pulse wave propagates in a second portion of an upper half body of the subject; and an evaluation-information obtaining means for obtaining, based on comparison of the inferior-limb-pulse-wave-propagation-velocity-related information obtained by the inferior-limb-pulse-wave-propagation-velocity-related-information obtaining device and the upper-half-body-pulse-wave-propagation-velocity-related information obtained by the upper-half-body-pulse-wave-propagation-velocity-related-information obtaining device, evaluation information that is related to evaluation of reliability of the inferior-and-superior-limb blood-pressure index determined by the inferior-and-superior-limb-blood-pressure-index determining means. | 1. An apparatus for measuring an inferior-and-superior-limb blood-pressure index of a living subject, comprising: an inferior-limb blood-pressure measuring device which includes an inferior-limb cuff adapted to be wound around an inferior limb of the subject and measures a blood pressure of the inferior limb; a superior-limb blood-pressure measuring device which includes a superior-limb cuff adapted to be wound around a superior limb of the subject and measures a blood pressure of the superior limb; an inferior-and-superior-limb-blood-pressure-index determining means for determining the inferior-and-superior-limb blood-pressure index of the subject, based on the blood pressure of the inferior limb measured by the inferior-limb blood-pressure measuring device and the blood pressure of the superior limb measured by the superior-limb blood-pressure measuring device; an inferior-limb-pulse-wave-propagation-velocity-related-information obtaining device which obtains inferior-limb-pulse-wave-propagation-velocity-related information that is related to a velocity at which an inferior-limb pulse wave propagates in a first portion of the subject that includes the inferior limb; an upper-half-body-pulse-wave-propagation-velocity-related-information obtaining device which obtains upper-half-body-pulse-wave-propagation-velocity-related information that is related to a velocity at which an upper-half-body pulse wave propagates in a second portion of an upper half body of the subject; and an evaluation-information obtaining means for obtaining, based on comparison of the inferior-limb-pulse-wave-propagation-velocity-related information obtained by the inferior-limb-pulse-wave-propagation-velocity-related-information obtaining device and the upper-half-body-pulse-wave-propagation-velocity-related information obtained by the upper-half-body-pulse-wave-propagation-velocity-related-information obtaining device, evaluation information that is related to evaluation of reliability of the inferior-and-superior-limb blood-pressure index determined by the inferior-and-superior-limb-blood-pressure-index determining means. 2. An apparatus according to claim 1, wherein the inferior-limb blood-pressure measuring device includes the inferior-limb cuff adapted to be wound around an ankle of the subject and measures a blood pressure of the ankle, wherein the superior-limb blood-pressure measuring device includes the superior-limb cuff adapted to be wound around an upper arm of the subject and measures a blood pressure of the upper arm, wherein the inferior-limb-pulse-wave-propagation-velocity-related-information obtaining device obtains the inferior-limb-pulse-wave-propagation-velocity-related information related to the velocity at which the inferior-limb pulse wave propagates in the first portion of the subject that includes the ankle and the upper arm, and wherein the upper-half-body-pulse-wave-propagation-velocity-related-information obtaining device obtains the upper-half... | Colin Medical Technology Corporation | A61B005/02
US07097620 | B2 | 2006-08-29 | Guidewire with pressure and temperature sensing capabilities | 2002-09-19 | A guide wire having pressure sensing capabilities for measuring the pressure of liquid in a vessel comprising a flexible elongate member and having proximal and distal extremities and having an outside diameter of 0.018″ or less. The distal extremity of said flexible elongate member is adapted to be disposed in the liquid in said vessel. A housing is carried by the flexible elongate member and has a diameter substantially the same as the diameter of the flexible elongate member. The housing has a space therein with a pressure sensor mounted in the space in the housing. The pressure sensor has a diaphragm that is sensitive to changes of pressure in the liquid in the vessel. The diaphragm is rectangular in shape and is bordered by a rim surrounding the well. A backing plate is formed of an insulating material bonded to the crystal and serves to reinforce the rim of the crystal of semiconductor material. The backing plate has a cavity therein underlying the diaphragm and is in substantial registration with the diaphragm. The cavity serves to provide a pressure reference. | 1. A pressure sensor apparatus having pressure sensing capabilities, the pressure sensor apparatus comprising: a guidewire having a maximum diameter of 0.018″ or less;a sensor housing having an external wall and a lumen, the sensor housing disposed near a distal extremity of the guidewire;a coil disposed proximally from the sensor housing;a solid state pressure sensor comprising a diaphragm and secured to the housing; andan ultrasonic transducer disposed distally from the pressure sensor. | 1. A pressure sensor apparatus having pressure sensing capabilities, the pressure sensor apparatus comprising: a guidewire having a maximum diameter of 0.018″ or less;a sensor housing having an external wall and a lumen, the sensor housing disposed near a distal extremity of the guidewire;a coil disposed proximally from the sensor housing;a solid state pressure sensor comprising a diaphragm and secured to the housing; andan ultrasonic transducer disposed distally from the pressure sensor. 2. The pressure sensor apparatus of claim 1, further comprising a control console electrically coupled to the pressure sensor and the ultrasonic transducer. 3. The pressure sensor apparatus of claim 2, wherein the control console is adapted to determine pressure and flow velocity at a distal region of the pressure sensor apparatus based on measurements obtained from the pressure sensor and measurements obtained from the ultrasonic transducer. 4. The pressure sensor apparatus of claim 1, wherein the pressure sensor further comprises a semiconductor crystal. 5. The pressure sensor apparatus of claim 1, wherein the ultrasonic transducer is adapted to measure flow velocity at a distal region of the pressure sensor apparatus. 6. The pressure sensor apparatus of claim 1, wherein the sensor housing comprises a cutout, and the pressure sensor is mounted in the cutout such that an outer surface of the pressure sensor is substantially flush with the external wall of the sensor housing. 7. The pressure sensor apparatus of claim 1, further comprising a second coil comprising a radiopaque material and disposed distally from the sensor housing. 8. A method for determining pressure at a target region within a patient's vasculature, the method comprising: providing a pressure sensor apparatus having a guidewire with a maximum diameter of 0.018″ or less, a sensor housing disposed near a distal extremity of the guidewire, a solid state pressure sensor secured to the housing, and an ultrasonic transducer disposed near the distal extremity;disposing the distal extremity of the pressure sensor apparatus at the target region within the patient's vasculature;obtaining measurements from the pressure sensor;obtaining measurements from the ultrasonic transducer; andanalyzing the measurements from the pressure sensor and the measurements from the ultrasonic transducer to determine pressure and flow velocity at the target region. 9. The method of claim 8, further comprising: providing a control console,wherein obtaining measurements from the pressure sensor, obtaining measurements from the ultrasonic transducer, and analyzing the measurements are performed with the control console. 10. The method of claim 8, wherein the target region comprises a vascular stenosis. 11. The method of claim 10, further comprising: determining a first pressure on a proximal side of the stenosis;determining a second pressure on a distal side of the stenosis; andcomparing the first and second pressures... | Volcano Corporation | A61B005/00
US07239908 | B1 | 2007-07-03 | Assessing the condition of a joint and devising treatment | 2000-09-14 | Methods are disclosed for assessing the condition of a cartilage in a joint, particularly a human knee. The methods include converting an image such as an MRI to a three dimensional map of the cartilage. The cartilage map can be correlated to a movement pattern of the joint to assess the affect of movement on cartilage wear. Changes in the thickness of cartilage over time can be determined so that therapies can be provided. Information on thickness of cartilage and curvature of cartilage or subchondral bone can be used to plan therapy. Information on movement pattern can be used to plan therapy. | 1. A method of treating a human joint disease involving cartigage comprising: obtaining electronic image data of a joint, wherein said image data includes both normal and diseased cartilage tissue; electronically evaluating in three dimensions said image data to obtain information about geometry of the joint, said electronically evaluating including evaluating the joint along one or more vectors tangent or perpendicular to the cartilage or bone surface; and selecting a therapy based on said three-dimensional evaluation. | 1. A method of treating a human joint disease involving cartigage comprising: obtaining electronic image data of a joint, wherein said image data includes both normal and diseased cartilage tissue; electronically evaluating in three dimensions said image data to obtain information about geometry of the joint, said electronically evaluating including evaluating the joint along one or more vectors tangent or perpendicular to the cartilage or bone surface; and selecting a therapy based on said three-dimensional evaluation. 2. The method of claim 1, wherein said electronically evaluating further comprises: estimating the loss of cartilage in a joint, wherein the joint comprises cartilage and accompanying bones on either side of the joint, which method comprises, obtaining a three-dimensional representation of the cartilage at an initial time and calculating the thickness or regional volume of a region of degenerated cartilage evaluated at the initial time, obtaining a three-dimensional representation of the cartilage at a later time, calculating the thickness or regional volume of a region of degenerated cartilage evaluated at the later time, and determining the loss in thickness or regional volume of the region of degenerated cartilage between the later and initial times. 3. The method of claim 1, wherein said electronically evaluating further comprises: assessing the condition of cartilage in a joint of a human, which method comprises, electronically tranferring an electronically generated image of a cartilage of the joint from a transferring device to a receiving device located distant from the transferring device, receiving the transferred image at the distant location, converting the transferred image to a degeneration pattern of the cartilage, and transmitting the degeneration pattern to a site for analysis. 4. The method of claim 1, wherein said electronically evaluating further comprises: determining the volume of cartilage loss in a region of a cartilage defect of a cartilage in a joint of a mammal which method comprises, determining the thickness, DN, of the normal cartilage near the cartilage defect, obtaining the thickness of the cartilage defect, DD, of the region, subtracting DD form DN to give the thickness of the cartilage loss, DL, and multiplying the DL value times the area of the cartilage defect, AD, to give the volume of cartilage loss. 5. The method of claim 1, wherein said electronically evaluating further comprises: estimating the change of cartilage in a joint of a mammal over time, which method comprises, estimating the thickness or width or area or volume of a region of cartilage at an inital time T1, estimating the thickness or width or area or volume of a region of cartilage at an later time T2, and determining the change in the thickness or width or area or volume of the region of cartilage between the initial and the later times. 6. The method of claim 1, wherein said electronically evaluating further compri... | The Board of Trustees of the Leland Stanford Junior University | A61B005/05
US07018336 | B2 | 2006-03-28 | Implantable sensor flush sleeve | 2001-12-27 | A apparatus and method for cleaning a sensor tip of an implantable sensor includes a flush sleeve directed towards the sensor tip, a fluid conduit in fluid communication with the flush sleeve, where fluid injected into a first end of a flush sleeve surrounding the sensor sprays off the sensor tip through at least one orifice located at the second end of the flush sleeve. The first end of the flush sleeve contains a fluid conduit and a septum, where a needle is used to pierce the septum to inject the fluid into the fluid conduit. | 1. An implantable sensor system for taking readings from a patient in vivo, the sensor system comprising: an implantable sensor having a distal end with a sensor tip for direct contact with patient fluids and a proximal end to anchor the implantable sensor within the patient;a flush sleeve directed towards the sensor tip;a rinsing fluid; anda fluid conduit in fluid communication with the flush sleeve,wherein the rinsing fluid received in the fluid conduit is used to spray the sensor tip,wherein the flush sleeve concentrically surrounds the implantable sensor around a substantially common axis, such that the sensor is within the flush sleeve,wherein the fluid conduit contains a septum and a protector sleeve, wherein the septum is pierced by a needle injected into the patient to deliver the rinsing fluid into the fluid conduit, and wherein the protector sleeve acts as a backstop to prevent the needle from penetrating the sensor. | 1. An implantable sensor system for taking readings from a patient in vivo, the sensor system comprising: an implantable sensor having a distal end with a sensor tip for direct contact with patient fluids and a proximal end to anchor the implantable sensor within the patient;a flush sleeve directed towards the sensor tip;a rinsing fluid; anda fluid conduit in fluid communication with the flush sleeve,wherein the rinsing fluid received in the fluid conduit is used to spray the sensor tip,wherein the flush sleeve concentrically surrounds the implantable sensor around a substantially common axis, such that the sensor is within the flush sleeve,wherein the fluid conduit contains a septum and a protector sleeve, wherein the septum is pierced by a needle injected into the patient to deliver the rinsing fluid into the fluid conduit, and wherein the protector sleeve acts as a backstop to prevent the needle from penetrating the sensor. 2. The sensor system of claim 1, further comprising a connector fitting for supporting the implantable sensor within the patient. 3. The sensor system of claim 2, wherein the sensor is plugged into the connector fitting, and the connector fitting is affixable internally to the patient. 4. The sensor system of claim 2, wherein the connector fitting is connected to a telemetry unit to transmit readings from the implantable sensor. 5. The sensor system of claim 1, wherein the flush sleeve surrounds the implantable sensor in a tight fit connection. 6. The sensor system of claim 5, wherein the flush sleeve contains at least one one-way valve near the sensor tip. 7. The sensor system of claim 1, wherein the fluid conduit is located at a proximal end of the sensor. 8. The sensor system of claim 7, wherein the proximal end of the sensor is covered by a protector sleeve. 9. The sensor system of claim 1, wherein the rinsing fluid is a saline solution. 10. The sensor system of claim 1, wherein the rinsing fluid contains an anti-coagulant. 11. An implantable multi-lumen sensor system for taking readings from a patient in vivo, the sensor system comprising: an inner lumen comprising an implantable sensor having a distal end with a sensor tip for direct contact with patient fluids and a proximal end to anchor the implantable sensor within the patient;an outer lumen comprising a flush sleeve surrounding the inner lumen in a substantially coaxial manner, such that the inner lumen is within the outer lumen; anda rinsing fluid received in the flush sleeve to spray the sensor tip,wherein the flush sleeve farther contains a septum and a protector sleeve, wherein the septum is pierced by a needle injected into the patient to deliver the rinsing fluid into the flush sleeve, and wherein the protector sleeve acts as a backstop to prevent the needle from penetrating the sensor. 12. The sensor system of claim 11, wherein the flush sleeve surrounds the inner lumen in a tight fit connection. 13. The sensor system of claim 12, wherein the f... | Medtronic MiniMed, Inc. | A61B005/00
US07215231 | B1 | 2007-05-08 | MRI system | 2002-08-16 | In one embodiment, a ferromagnetic frame for use in an MRI system comprises first and second opposing ferromagnetic elements magnetically coupled by a flux return path. First and second poles are provided, separated by a distance to define an imaging volume therebetween. The first pole is movably supported with respect to the first ferromagnetic element to vary the distance between the first and second poles. The first pole may have a first, imaging position and a second position wherein the first pole is farther from the second pole than in the first position. When the first pole is in the first, imaging position, a gap is provided between the first pole and the first ferromagnetic element. The second pole is supported by the second ferromagnetic element. When the first pole is in the second position, the imaging volume is made larger, providing room for medical personnel to access a patient. Performance of medical procedures on the patient is thereby facilitated. The flux return path may be two or more opposing ferromagnetic posts or plates, for example. Magnet assemblies, MRI systems, methods of conducting an MRI procedure and methods of conducting a medical procedure are also disclosed. | 1. A magnetic resonance imaging (MRI) magnet assembly, comprising: a ferromagnetic frame comprising:opposing, vertical, ferromagnetic elements;opposing, horizontal, upper and lower ferromagnetic pole supports connected to the vertical ferromagnetic supports;opposing upper and lower poles separated by a distance to define an imaging volume therebetween;at least one member extending from each vertical ferromagnetic element; anda source of magnetic flux associated with the ferromagnetic frame;wherein: the lower pole is supported by the lower ferromagnetic pole support;the upper pole is movably supported with respect to the upper pole support to vary the distance between the upper pole and the lower pole between one position, wherein the upper pole is separated from the lower pole by a first distance, and an other position, wherein the upper pole is separated from the lower pole by a second distance greater than the first distance; andthe at least one member is positioned to support the upper pole when the upper pole is in the one position, and not support the upper pole when the upper pole is in the other position. | 1. A magnetic resonance imaging (MRI) magnet assembly, comprising: a ferromagnetic frame comprising:opposing, vertical, ferromagnetic elements;opposing, horizontal, upper and lower ferromagnetic pole supports connected to the vertical ferromagnetic supports;opposing upper and lower poles separated by a distance to define an imaging volume therebetween;at least one member extending from each vertical ferromagnetic element; anda source of magnetic flux associated with the ferromagnetic frame;wherein: the lower pole is supported by the lower ferromagnetic pole support;the upper pole is movably supported with respect to the upper pole support to vary the distance between the upper pole and the lower pole between one position, wherein the upper pole is separated from the lower pole by a first distance, and an other position, wherein the upper pole is separated from the lower pole by a second distance greater than the first distance; andthe at least one member is positioned to support the upper pole when the upper pole is in the one position, and not support the upper pole when the upper pole is in the other position. 2. The MRI magnet assembly of claim 1, wherein: in the one position, the upper pole and the upper pole support are separated by a third distance defining a gap; andin the other position, the upper pole and the upper pole support are separated by a fourth distance, the fourth distance being in a range from about zero to a distance less than the third distance. 3. The MRI magnetic assembly of claim 2, wherein: a sum of the first distance between the upper pole and the lower pole and the third distance between the upper pole and the upper pole support when the upper pole is in the one position is equal to a sum of a fifth distance between the upper pole and the lower pole and the fourth distance between the upper pole and the upper pole support when the upper pole is in the other position. 4. The MRI magnet assembly of claim 1, further comprising: stationary ferromagnetic material around the upper pole;wherein the upper pole and the ferromagnetic material each have a respective center of mass; andthe ferromagnetic material is positioned such that the magnetic center of mass of the ferromagnetic material is above the magnetic center of mass of the upper pole when the upper pole is in the one position and the magnetic center of mass of the ferromagnetic material is below the magnetic center of mass of the upper pole when the upper pole is in the other position. 5. The MRI magnet assembly of claim 4, wherein the ferromagnetic material is connected to and extends from the vertical ferromagnetic elements. 6. The MRI magnet assembly of claim 4, wherein the ferromagnetic material comprises a plurality of ferromagnetic segments extending from the vertical ferromagnetic elements. 7. The MRI magnet assembly of claim 1, wherein each one of the opposing, vertical ferromagnetic elements is a ferromagnetic plate. 8. The MRI magnetic assembly of cla... | Fonar Corporation | H01F005/00 A61B005/05
US07039447 | B2 | 2006-05-02 | Glucose measurement utilizing non-invasive assessment methods | 2003-01-02 | This involves non-invasive glucose measurement processes for determining blood glucose level in the human body. After achieving a static level of glucose at a surface of the skin over some period of time, the glucose may then be measured by a variety of different processes. A sample of the glucose may also first be extracted from the skin and this sample may then be measured. Clearly, these processes are especially suitable for monitoring glucose levels in the human body, and is especially beneficial to users having diabetes mellitus. These procedures may be used for other analyte materials that are found in appropriate regions of the outer skin. | 1. A method for non-invasively determining the concentration of glucose in blood of a subject comprising the steps of: removing glucose from a skin surface;allowing glucose to accumulate on the skin surface over a period of time, wherein the level is correlateable to the level of glucose within the blood;measuring the level of glucose that has accumulated on the skin surface;correlating the level of glucose to the level of glucose within the blood by using predetermined calibration data obtained by analysis of the subject's blood. | 1. A method for non-invasively determining the concentration of glucose in blood of a subject comprising the steps of: removing glucose from a skin surface;allowing glucose to accumulate on the skin surface over a period of time, wherein the level is correlateable to the level of glucose within the blood;measuring the level of glucose that has accumulated on the skin surface;correlating the level of glucose to the level of glucose within the blood by using predetermined calibration data obtained by analysis of the subject's blood. 2. The method of claim 1 wherein the step of removing glucose comprises applying a glucose solvent to the skin surface. 3. The method of claim 1 wherein the step of allowing glucose to accumulate on the skin surface takes less than 10 minutes. 4. The method of claim 3 wherein the step of allowing glucose to accumulate on the skin surface takes less than 1 minute. 5. The method of claim 1 wherein the step of measuring comprises measuring with electrochemical sensors. 6. The method of claim 1 wherein the step of measuring comprises measuring with optochemical sensors. 7. The method of claim 1 wherein the step of measuring comprises measuring with near-infrared spectroscopy. 8. The method of claim 1 wherein the step of measuring comprises measuring with mid-infrared spectroscopy. 9. The method of either of claim 7 or 8 wherein the step of removing glucose does not interfere with IR spectra of glucose obtained during the step of measuring. 10. The method of claim 1 wherein the step of measuring the level of glucose comprises extracting a sample from the skin surface and measuring the glucose level from the sample. 11. The method of claim 10 wherein the step of extracting is selected from the group consisting of blister suction, wick extraction, microdialysis extraction, iontophoretic extraction, sontophoretic extraction, and chemically enhanced extraction. 12. A method for calibrating a non-invasive device for determining the concentration of glucose in a subject's blood comprising the steps of: removing glucose from a surface of skin;allowing glucose to accumulate on the skin surface over a period of time;measuring the level of glucose that has accumulated on the skin surface;comparing the measured level of glucose to a level obtained by a blood sample glucose determination;adjusting the measured level of glucose so that the level is correlated to the level of glucose obtained by the blood sample glucose determination. 13. The method of claim 12 wherein the step of allowing glucose to accumulate on the skin surface takes less than 10 minutes. 14. The method of claim 13 wherein the step of allowing glucose to accumulate on the skin surface takes less than 1 minute. 15. The method of claim 12 wherein the level of glucose on the skin surface is correlateable to the level of glucose found in the blood to a degree that does not vary by more than 15%. 16. A method for non-invasively determining the concentration o... | VivoMedical, Inc. | A61B005/00
US06897655 | B2 | 2005-05-24 | Moving table MRI with frequency-encoding in the z-direction | 2002-09-04 | A system and method are disclosed using continuous table motion while acquiring data to reconstruct MR images across a large FOV without significant slab-boundary artifacts that reduces acquisition time. At each table position, full z-encoding data are acquired for a subset of the transverse k-space data. The table is moved through a number of positions over the desired FOV and MR data are acquired over the plurality of table positions. Since full z-data are acquired for each slab, the data can be Fourier transformed in z, interpolated, sorted, and aligned to match anatomic z locations. The fully sampled and aligned data is then Fourier transformed in remaining dimension(s) to reconstruct the final image that is free of slab-boundary artifacts. | 1. A method of imaging large volumes without resulting slab-boundary artifacts, the method comprising: defining a desired FOV larger than an optimal imaging volume of an MR scanner in at least a first direction; defining desired phase encode locations in at least one dimension perpendicular to the first direction; choosing a plurality of slab thicknesses in the first direction wherein slab thickness varies with each phase encode location and is smaller than the desired FOV and within the optimal imaging volume of the MR scanner; and continuously moving one of the optimal imaging volume and an imaging object in the first direction while repeatedly exciting and encoding spins with readout in the first direction to acquire data that is restricted to a slab thickness for a set of phase encode locations until a set of MR data for the FOV for the first set of phase encode locations is acquired; reversing the direction of motion of one of the optimal imaging volume and the imaging object in the first direction; continuously moving one of the optimal imaging volume and the imaging object in the a reversed direction while repeatedly exciting and encoding spins with readout in the first direction to acquire data that is restricted to the slab thickness for a second set of phase encode locations until a set of MR data for the FOV for the second set of phase encode locations is acquired; and reconstructing at least one image of the FOV from the set of MR data for the first set of phase encode locations and the set of MR data for the second set of phase encode locations. | 1. A method of imaging large volumes without resulting slab-boundary artifacts, the method comprising: defining a desired FOV larger than an optimal imaging volume of an MR scanner in at least a first direction; defining desired phase encode locations in at least one dimension perpendicular to the first direction; choosing a plurality of slab thicknesses in the first direction wherein slab thickness varies with each phase encode location and is smaller than the desired FOV and within the optimal imaging volume of the MR scanner; and continuously moving one of the optimal imaging volume and an imaging object in the first direction while repeatedly exciting and encoding spins with readout in the first direction to acquire data that is restricted to a slab thickness for a set of phase encode locations until a set of MR data for the FOV for the first set of phase encode locations is acquired; reversing the direction of motion of one of the optimal imaging volume and the imaging object in the first direction; continuously moving one of the optimal imaging volume and the imaging object in the a reversed direction while repeatedly exciting and encoding spins with readout in the first direction to acquire data that is restricted to the slab thickness for a second set of phase encode locations until a set of MR data for the FOV for the second set of phase encode locations is acquired; and reconstructing at least one image of the FOV from the set of MR data for the first set of phase encode locations and the set of MR data for the second set of phase encode locations. 2. The method of claim 1 wherein the phase encode locations are divided into a number of sets and wherein each set has a constant slab thickness. 3. The method of claim 1 further comprising acquiring multiple lines of data following a single excitation to minimize overall scan time. 4. The method of claim 1, further comprising selecting a set of phase encode locations to be proximate to a center of k-space in a transverse direction and timing acquisition of the set of phase encode locations to capture optimal image contrast associated with passage of a bolus of extrinsic contrast agent. 5. The method of claim 1 further comprising interleaving with a pre-defined temporal frequency between the exciting and encoding waveforms and alternate RF and gradient waveforms to manipulate image contrast. 6. The method of claim 5 further comprising alternate RF and gradient waveforms designed to manipulate spins in spatially localized regions relative to the slab. 7. The method of claim 1 further comprising acquiring MR data associated with multiple phase encode locations following a single excitation. 8. The method of claim 1 further comprising acquiring MR data associated with a single phase encode location multiple times following a single excitation. 9. The method of claim 1 further comprising acquiring MR data with different temporal frequencies for reconstructing an image and further compris... | General Electric Company The Board of Trustees of the Leland Stanford, Jr. University | G01V003/00 A61B005/05
US06959213 | B2 | 2005-10-25 | Implantable device with digital waveform telemetry | 2002-11-22 | A technique for acquiring and accessing information from a medical implantable device is provided. Analog waveforms of interest are sensed and processed by signal acquisition circuitry. Analog parameters of interest are applied to selector switches which are controlled by a logic circuit. The logic circuit is also coupled an A/D converter for converting the analog signals to digital values. The digital values are stored in dedicated registers and are available for telemetry to an external device upon receipt of a request or prompt signal. When a digitized value is accessed and telemetered, the control logic circuit changes the conductive state of the selector switches to apply the corresponding analog signal to the A/D converter. The resulting digital value is applied to the corresponding register to refresh the accessed and telemetered value. The technique permits the external device to request and configure the implanted device to send only digitized values of interest. The technique also makes efficient use of the A/D converter, which consumes energy only as needed to refresh the memory when digital values are accessed and telemetered. | 1. An integrated circuit to process at least one signal of signals of interest, the integrated circuit comprising: at least two leads to conduct the signals of interest, wherein a first lead of the at least two leads includes an atrial tip conductor and an atrial ring conductor, and wherein a second lead of the at least two leads includes a ventricular tip conductor and a ventricular ring conductor; an atrial sense amplifier coupled to the first lead of the at least two leads; a ventricular sense amplifier coupled to the second lead of the at least two leads; an atrial ring-to-can intracardiac electrogram amplifier coupled to the first lead of the at least two leads; a ventricular ring-to-can intracardiac electrogram amplifier coupled to the second lead of the at least two leads; an impedance sensor coupled to one of the at least two leads a switch bank coupled to the atrial sense amplifier and the ventricular sense amplifier, wherein the switch bank includes a plurality of switching devices; and first and second programmable gain amplifiers coupled to the switch bank. | 1. An integrated circuit to process at least one signal of signals of interest, the integrated circuit comprising: at least two leads to conduct the signals of interest, wherein a first lead of the at least two leads includes an atrial tip conductor and an atrial ring conductor, and wherein a second lead of the at least two leads includes a ventricular tip conductor and a ventricular ring conductor; an atrial sense amplifier coupled to the first lead of the at least two leads; a ventricular sense amplifier coupled to the second lead of the at least two leads; an atrial ring-to-can intracardiac electrogram amplifier coupled to the first lead of the at least two leads; a ventricular ring-to-can intracardiac electrogram amplifier coupled to the second lead of the at least two leads; an impedance sensor coupled to one of the at least two leads a switch bank coupled to the atrial sense amplifier and the ventricular sense amplifier, wherein the switch bank includes a plurality of switching devices; and first and second programmable gain amplifiers coupled to the switch bank. 2. The integrated circuit of claim 1, further comprising a combiner coupled to the atrial ring-to-can intracardiac electrogram amplifier, and wherein the combiner is further coupled to the ventricular ring-to-can intracardiac electrogram amplifier. 3. The integrated circuit of claim 2, further comprising a pair of node switches coupled to the ventricular ring-to-can intracardiac electrogram amplifier and the combiner. 4. The integrated circuit of claim 1, wherein the atrial ring-to-can intracardiac electrogram amplifier includes: a comparing circuit to produce a compared signal; a filtering circuit receptive to the compared signal to produce a filtered signal; and an amplifying circuit receptive to the filtered signal to produce an amplified signal. 5. The integrated circuit of claim 4, wherein the comparing circuit includes an operational amplifier to compare the signals of interest conducted on the at least two leads, wherein the operational amplifier is receptive to the atrial tip conductor of the first lead and the atrial ring conductor of the first lead. 6. The integrated circuit of claim 4, wherein the filtering circuit includes: a first low pass filter to produce a first filtered signal; a set of low pass filters receptive to the first filtered signal to produce a second filtered signal; and a high pass filter receptive to the second filtered signal to produce a third filtered signal. 7. The integrated circuit of claim 6, further comprising: a first set of switches receptive to the first filtered signal; a second set of switches receptive to the second filtered signal; and a third set of switches receptive to the third filtered signal. 8. The integrated circuit of claim 6, further comprising a programmable threshold-crossing circuit to produce at least one indicating pulse if the third filtered signal is at a predetermined level so as to indicate whether a chamber... | Intermedics, Inc. | A61N001/362 A61B005/0402
US06915150 | B2 | 2005-07-05 | Process and device for the preoperative determination of the positioning data of endoprosthetic parts | 2002-03-08 | In order to be able to ascertain the position of the bone in the body without complicated procedures in a process for the preoperative determination of the positioning data of endoprosthetic parts of a central joint relative to the bones forming the central joint, a respective outer articular point is determined by way of movement of the bones about a respective outer joint which is located at the end of the two bones facing away from the central joint. An articular point is determined for each of the two bones in the area of the central joint, and a direction characteristic for each of these bones is determined by way of a straight-line connection of the two articular points obtained in this manner for the two bones. The orientation of the endoprosthetic parts relative to this characteristic direction is then determined. Apparatus for carrying out the process is also provided. | 1. A method for the preoperative determination of positioning data of an endoprosthetic part of a knee joint relative to a thigh bone and a lower leg bone forming the knee joint, comprising the steps of: moving the thigh bone of the knee joint relative to a corresponding hip bone while determining the positions of the thigh bone and the hip bone, to determine a first articular point for said thigh bone in the area of the hip; the thigh and hip bone positions being determined by securely connecting marking elements to said bones and determining the respective positions of the marking elements using a measuring device that provides signals corresponding to the respective positions to a data processing system; determining a second articular point for said thigh bone in the area of the knee joint; determining a first characteristic direction for said thigh bone by way of a straight-line connection of the first and second articular points; and determining an orientation of the endoprosthetic part relative to the first characteristic direction. | 1. A method for the preoperative determination of positioning data of an endoprosthetic part of a knee joint relative to a thigh bone and a lower leg bone forming the knee joint, comprising the steps of: moving the thigh bone of the knee joint relative to a corresponding hip bone while determining the positions of the thigh bone and the hip bone, to determine a first articular point for said thigh bone in the area of the hip; the thigh and hip bone positions being determined by securely connecting marking elements to said bones and determining the respective positions of the marking elements using a measuring device that provides signals corresponding to the respective positions to a data processing system; determining a second articular point for said thigh bone in the area of the knee joint; determining a first characteristic direction for said thigh bone by way of a straight-line connection of the first and second articular points; and determining an orientation of the endoprosthetic part relative to the first characteristic direction. 2. A method in accordance with claim 1, wherein: said first articular point corresponds to a central point of said hip joint; and said second articular point corresponds to an intersection point of a plurality of curves of maximum invariance generated by movement, relative to one another, of the thigh bone and the lower leg bone at the knee joint. 3. A method in accordance with claim 1, comprising the further steps of: moving the lower leg bone of the knee joint relative to a corresponding ankle bone, to determine a third articular point in the area of an ankle joint for said lower leg bone and determining the positions of the lower leg bone and the ankle bone; determining a fourth articular point in the area of the knee joint for said lower leg bone; determining a second characteristic direction for said lower leg bone by way of a straight-line connection of the third and fourth articular points; and determining the orientation of an endoprosthetic part relative to the second characteristic direction. 4. A method in accordance with claim 3, wherein: said third articular point corresponds to an intersection point of pivot axes of the ankle joint about an axis extending transversely in relation to the lower leg bone and about an axis extending longitudinally in relation to the lower leg bone; and said fourth articular point corresponds to an intersection point of a plurality of curves of maximum invariance generated by movement, relative to one another, of the lower leg bone and the thigh bone at the knee joint. 5. A method in accordance with claim 3, wherein, for the orientation of the endoprosthetic part, sawing planes serving as contact surfaces for the endoprosthetic part are determined, said planes taking up predetermined orientations relative to the characteristic directions. 6. A method in accordance with claim 5, wherein the sawing planes are at right angles to the characteristic directions. 7. ... | Aesculap AG & Co. KG Aesculap-ICP S.A. | A61B005/00
US06901371 | B1 | 2005-05-31 | Imaging system protocol handling method and apparatus | 1999-12-30 | A technique is disclosed for providing programs, such as operational protocols, to medical diagnostic institutions and systems. The protocols are created and stored on machine readable media. A description of the protocols is displayed at the diagnostic institution or system. A user may select a desired protocol or program from a user interface, such as a listing of protocols. The protocol listing may include textual and exemplary image descriptions of the protocols. Selected protocols are transferred from the machine readable media to the diagnostic institution or system. The transfer may take place over a network link, and may be subject to fee arrangements, subscription status verifications, and so forth. Protocols may be loaded for execution on system scanners by selection from the same or a similar protocol listing screen. | 1. A method for providing operational protocols to medical diagnostic systems, the method comprising the steps of: storing a protocol on a machine readable medium, the protocol including at least one operating parameter for a medical diagnostic system; displaying user viewable indicia descriptive of the protocol at a medical diagnostic location, wherein the user viewable indicia include an exemplary image obtainable via the protocol; performing a protocol exchange transaction including selecting the protocol via a user interface and loading the protocol at the medical diagnostic location from the machine readable medium via a network connection to the medical diagnostic location; and storing an accounting record of the transaction. | 1. A method for providing operational protocols to medical diagnostic systems, the method comprising the steps of: storing a protocol on a machine readable medium, the protocol including at least one operating parameter for a medical diagnostic system; displaying user viewable indicia descriptive of the protocol at a medical diagnostic location, wherein the user viewable indicia include an exemplary image obtainable via the protocol; performing a protocol exchange transaction including selecting the protocol via a user interface and loading the protocol at the medical diagnostic location from the machine readable medium via a network connection to the medical diagnostic location; and storing an accounting record of the transaction. 2. The method of claim 1, wherein the accounting record includes data for invoicing fees associated with the protocol. 3. The method of claim 1, comprising the further step of transmitting an authorization prompt to the medical diagnostic location prior to loading the protocol. 4. The method of claim 1, comprising the further step of verifying a service subscription of the medical diagnostic location, the accounting record referencing the subscription. 5. The method of claim 4, wherein the subscription verified includes data representative of a time expiring subscription. 6. The method of claim 1, wherein the user viewable indicia include a textual description of the protocol. 7. The method of claim 1, wherein the user viewable indicia are viewed at a computer workstation coupled to the medical diagnostic system. 8. The method of claim 1, wherein the protocol includes data for filming, viewing, reconstructing or processing images reconstructed from image data. 9. The method of claim 1, wherein the network connection includes the Internet. 10. A method for obtaining an operational protocol for a medical diagnostic system or institution, the method comprising the steps of: ordering a protocol by viewing a protocol list on a user interface at the medical diagnostic system or institution, and selecting a desired protocol from the list, wherein the protocol list includes an exemplary image obtainable via the protocol; performing a transaction by accessing data from a protocol library defining the desired protocol via a network link between the diagnostic system or institution and the library, and transmitting the data from the library to the diagnostic system; and storing a record of the transaction. 11. The method of claim 10, wherein the library includes protocols for a plurality of diagnostic system modalities, and wherein the protocol list includes only protocols for a modality of the medical diagnostic system. 12. The method of claim 10, comprising the further step of transmitting data descriptive of the protocol to the medical diagnostic system for addition to the protocol list. 13. The method of claim 10, comprising the further step of authorizing a fee for the protocol. 14. The method of claim 10, whe... | GE Medical Systems Global Technology Company, LLC | G06F017/60 A61B005/00 A61B010/00 G06T007/00
US06860857 | B2 | 2005-03-01 | Implantable intravascular pressure determining device and method | 2002-02-19 | In an implantable intravascular pressure determining device and method, a pressure sensor generates a raw pressure signal and an acceleration sensor measures acceleration in a patient. Time intervals are identified wherein the raw pressure signal accurately represents the intravascular pressure, these intervals being identified as the time intervals wherein the measured acceleration is below a predetermined threshold. The raw pressure signal is processed, to generate a processed signal which is used as an intravascular pressure signal, only in the aforementioned time intervals. | 1. An implantable intravascular pressure determining device comprising: a pressure sensor adapted for placement in a vascular system of a subject for generating a raw pressure signal; an acceleration sensor, separate from said pressure sensor, adapted for measuring acceleration associated with said subject and for generating an acceleration signal corresponding to the measured acceleration, said acceleration signal being separate from said raw pressure signal; an evaluation unit supplied with said raw pressure signal and said acceleration signal for identifying time intervals, in which said measured acceleration is below a predetermined threshold, and for generating, a control signal in said time intervals said evaluation unit identifying a maximum signal strength of said acceleration signal, and setting said predetermined threshold to a predetermined fraction of said maximum signal strength; and a processor supplied with said raw pressure signal and said control signal for processing said raw pressure signal, to generate a processed signal corresponding to intravascular pressure, only in the time intervals represented by said control signal. | 1. An implantable intravascular pressure determining device comprising: a pressure sensor adapted for placement in a vascular system of a subject for generating a raw pressure signal; an acceleration sensor, separate from said pressure sensor, adapted for measuring acceleration associated with said subject and for generating an acceleration signal corresponding to the measured acceleration, said acceleration signal being separate from said raw pressure signal; an evaluation unit supplied with said raw pressure signal and said acceleration signal for identifying time intervals, in which said measured acceleration is below a predetermined threshold, and for generating, a control signal in said time intervals said evaluation unit identifying a maximum signal strength of said acceleration signal, and setting said predetermined threshold to a predetermined fraction of said maximum signal strength; and a processor supplied with said raw pressure signal and said control signal for processing said raw pressure signal, to generate a processed signal corresponding to intravascular pressure, only in the time intervals represented by said control signal. 2. An implantable intravascular pressure determining device as claimed in claim 1 wherein said evaluation unit uses an amplitude of said acceleration signal, as a representative of said signal strength, for identifying said maximum signal strength. 3. An implantable intravascular pressure determining device as claimed in claim 1 wherein said evaluation unit samples an amplitude of said acceleration signal, to obtain a sampled momentary value of said acceleration signal, and wherein said evaluation unit uses said sampled momentary value, as a representative of said signal strength, for identifying said maximum signal strength. 4. An implantable intravascular pressure determining device as claimed in claim 1 wherein said evaluation unit repeatedly samples an amplitude of said acceleration signal to obtain a plurality of sampled momentary values, over a plurality of diastolic heart cycles intervals, and wherein said evaluation unit forms an average value of said plurality of sampled momentary values and compares said average value, as a representative of said signal strength, to said predetermined threshold to identify said time intervals. 5. An implantable intravascular pressure determining device as claimed in claim 4 wherein said evaluation unit forms said average value over a plurality of said diastolic heart cycle intervals in a range between 30 and 40. 6. An implantable intravascular pressure determining device as claimed in claim 4 wherein said processor processes said raw pressure signal to generate said processed signal corresponding to said intravascular pressure in at least one of said time intervals wherein said average value is substantially zero. 7. An implantable intravascular pressure determining device as claimed in claim 1 wherein said evaluation unit sets said predetermined threshold a... | St. Jude Medical AB | A61B005/00
US06950690 | B1 | 2005-09-27 | Method for delivering a device to a target location | 1999-10-21 | Capsule (60) moves through the gastrointestinal tract (62) in a first pass to generate a map of the gastrointestinal tract, and to identify a location of interest. In its second pass, capsule (60) moves through the gastrointestinal tract, and is controlled to perform a job at the identified location. Repeated localizations generate generate a map of the route taken by the capsule in the gastrointestinal tract (62). images displayed on the image monitor (61) are compared with the generated map displayed on the position monitor (63) to identify the location of a pathology (72). | 1. A method comprising: recording a first set of images from a gastrointestinal tract using a swallowable imaging device;identifying an image showing a location of interest;recording a second set of images from the gastrointestinal tract using a swallowable imaging device; andperforming image analysis by comparing images from the first set with images from the second set to identify in the second set an image of interest corresponding to an image in the fist set, thereby determining the location of the imaging device relative to the location of interest. | 1. A method comprising: recording a first set of images from a gastrointestinal tract using a swallowable imaging device;identifying an image showing a location of interest;recording a second set of images from the gastrointestinal tract using a swallowable imaging device; andperforming image analysis by comparing images from the first set with images from the second set to identify in the second set an image of interest corresponding to an image in the fist set, thereby determining the location of the imaging device relative to the location of interest. 2. The method according to claim 1, wherein the imaging device is to generate data, and wherein the imaging device is a capsule comprising: means for signal analysis of data generated in a first pass and a second pass; andmeans for controlling the imaging device in a second pass according to said signal analysis; andmeans for performing a job in the gastrointestinal tract. 3. The method according to claim 1, comprising the steps of: inserting the imaging device into the gastrointestinal tract in a first pass;locating the imaging device; anddisplaying the location on a position monitor. 4. The method according to claim 3, further comprising a step of displaying the location of the imaging device two or three dimensionally. 5. The method according to claim 4, wherein the location of the imaging device is displayed as an overlay to a schematic presentation of the gastrointestinal tract. 6. The method according to claim 1, wherein the step of recording a first set of images using a swallowable imaging device comprises recording the first set of images in a first pass, and wherein the step of recording a second set of images using a swallowable imaging device comprises: recording the second set of images from the gastrointestinal tract in a second pass;performing image analysis by comparing images from the first pass and of images from in the second pass; anddetermining the location of the imaging device relative to the location of interest. 7. The method according to claim 6, wherein the first pass and the second pass are one or more passes. 8. The method according to claim 1, wherein the location of interest is a location of a pathology. | Given Imaging LTD | A61B005/05
US06928742 | B2 | 2005-08-16 | Method and apparatus for finding the position of a mechanical axis of a limb | 2001-05-17 | A method of finding the position of a mechanical axis running in the longitudinal direction through a limb that is rotatably supported by its first end in a center of rotation that is not fixed and is not accessible for the mechanical determination of coordinates, by means of an optical coordinate-measurement device with an indicator to detect measurement-point coordinates in one rotational position of the limb, such that, for each rotational position, a multi-point indicator rigidly attached near the second end of the limb signals several measurement-point coordinates, and, from the sets of measurement-point coordinates detected in a plurality of rotational positions, at least one group is selected that can be assigned to the same site of the center of rotation, and the measurement-point-coordinate sets assigned to this site of the center of rotation are used to calculate the mechanical axis. | 1. A method of finding the position of a mechanical axis running in the longitudinal direction through a limb, the limb rotatably supported by a first end via a movable and inaccessible center of rotation for the mechanical determination of coordinates, comprising: providing an optical coordinate-measurement device with at least one multi-point indicator to signal measurement-point coordinates in a rotational position of the limb; rigidly attaching said multi point indicator near a second end of the limb; detecting several measurement-point coordinates (P11, P21, P31, Pn1) per rotational position of the limb; recording said measurement-point coordinates; selecting from measurement-point-coordinate sets (P11-Pn1, P12-Pn2, P1m-Pnm) at least one group; assigning said at least one group to a same site of the center of rotation; and calculating the position of the first end of the limb and a mechanical axis running through the first and second ends using the measurement-point-coordinate sets (P11-Pn1, P12-Pn2, P1m-Pnm) assigned to said site of the center of rotation without piercing a skin of the limb above said second end. | 1. A method of finding the position of a mechanical axis running in the longitudinal direction through a limb, the limb rotatably supported by a first end via a movable and inaccessible center of rotation for the mechanical determination of coordinates, comprising: providing an optical coordinate-measurement device with at least one multi-point indicator to signal measurement-point coordinates in a rotational position of the limb; rigidly attaching said multi point indicator near a second end of the limb; detecting several measurement-point coordinates (P11, P21, P31, Pn1) per rotational position of the limb; recording said measurement-point coordinates; selecting from measurement-point-coordinate sets (P11-Pn1, P12-Pn2, P1m-Pnm) at least one group; assigning said at least one group to a same site of the center of rotation; and calculating the position of the first end of the limb and a mechanical axis running through the first and second ends using the measurement-point-coordinate sets (P11-Pn1, P12-Pn2, P1m-Pnm) assigned to said site of the center of rotation without piercing a skin of the limb above said second end. 2. The method according to claim 1, further comprising: calculating several spherical surfaces that represent paths through which the second end of the limb moves from the at least one group of measurement point coordinate sets (P11-Pn1, P12-Pn2, P1m-Pnm) assigned to said same site of the center of rotation; and calculating the coordinates of a point on which said spherical surfaces are centered to represent the center of rotation of the limb as a point on the mechanical axis, said calculation employing the equation position P(n m) relative to P(n+i m)=position P(n m+j) relative to P(n+im+j). 3. The method of claim 1, wherein the selection of the at least one group of measurement-point-coordinate sets (P11-Pn1, P12-Pn2, P1m-Pnm) from a number of sets that have been collected in a plurality of rotational positions of the limb is performed by proceeding through several steps of an iteration procedure, in each step of which one measurement-point-coordinate set is eliminated and an improvement thereby achieved is evaluated. 4. The method according to claim 3, wherein the evaluation is accomplished by means of an iterative calculation by a least-squares method. 5. The method of claim 1, further comprising: calibrating the multi-point indicator before beginning the collection of the measurement-point coordinates; comparing measurement-point coordinate relations resulting from said calibration with the relations of the recorded measurement-point coordinates for every rotational position of the limb; and deleting a measurement-point-coordinate set lying outside a predetermined tolerance limit. 6. The method of claim 1, further comprising specifying a second point required for the calculation of the mechanical axis, in terms of bony reference points by employing the optical coordinate measurement device with an additional indicator... | Plus Orthopedics AG | A61B005/103
US06875421 | B2 | 2005-04-05 | Method for detecting chitin-containing organisms | 2002-07-22 | The present invention provides a method for detecting chitin-containing organisms on an area of a person or animal by contacting the particular area with a dye that is capable of binding or conjugating to chitin and emitting fluorescence upon exposure to light. If a chitin-containing organism is present in the treated area, the chitin of the organism will bind or conjugate the dye and, upon exposure to light, the chitin-containing organisms may be visualized and removed. Also provided by the present invention are solutions and suspensions that contain a dye capable of binding or conjugating to a chitin-containing organism and emitting fluorescence upon exposure to light. The solutions and suspensions provided herein may be in the form of a shampoo, cream, lotion or detergent for the detection of chitin-containing organisms present in body hair, on the skin, clothing or the fur of animals. The solutions and suspensions may also be in a form suitable for spraying on clothes, as well as on the hair, skin and fur. Further provided by the present invention are kits comprising a solution or suspension that contains a dye capable of binding or conjugating to a chitin-containing organism and a source of light that emits a wavelength capable of exciting fluorescence from the chitin-bound or chitin-conjugated dye. | 1. A kit comprising a solution or suspension comprising a dye that binds to a chitin-containing organism and a source of light that emits a wavelength capable of exciting fluorescence from the dye, wherein the dye is present in an amount effective to bind the chitin-containing organism and emit fluorescence upon exposure to light, wherein the solution or suspension is formulated in a shampoo, lotion, cream, detergent or spray and, when applied to skin, hair, fur or clothing, the solution or suspension renders an arthropod detectable, and wherein the dye is calcofluor white or distyryl biphenyl fluorescent whitening agent. | 1. A kit comprising a solution or suspension comprising a dye that binds to a chitin-containing organism and a source of light that emits a wavelength capable of exciting fluorescence from the dye, wherein the dye is present in an amount effective to bind the chitin-containing organism and emit fluorescence upon exposure to light, wherein the solution or suspension is formulated in a shampoo, lotion, cream, detergent or spray and, when applied to skin, hair, fur or clothing, the solution or suspension renders an arthropod detectable, and wherein the dye is calcofluor white or distyryl biphenyl fluorescent whitening agent. 2. The kit of claim 1, wherein the solution or suspension is formulated in a shampoo. 3. The kit of claim 1, wherein the solution or suspension is formulated in a lotion. 4. The kit of claim 1, wherein the solution or suspension is formulated in a cream. 5. The kit of claim 1, wherein the solution or suspension is formulated in a detergent. 6. The kit of claim 1, wherein the solution or suspension is formulated in a spray. 7. The kit of claim 1, wherein the dye comprises 0.02-5% of the solution or suspension. 8. The kit of claim 1, wherein the dye is calcofluor white. 9. The kit of claim 1, wherein the dye is distyryl biphenyl fluorescent whitening agent. | Albert Einstein College of Medicine of Yeshiva University | A61B010/00 A61B005/00 A61B008/00
US06969357 | B1 | 2005-11-29 | Gas analyzer calibration checking device | 2000-06-07 | A calibration checking device, and methods of mandating its use at regular time intervals, to ensure maintenance of the accuracy of gas analyzers, especially for use with breath tests, Each calibration checking device is designed to be used for a predetermined number of tests. After first connection of a new calibration checking device, a volume of known calibration checking gas is released into the instrument, and a calibration checking measurement is initiated. A signal is sent to a counting mechanism which both enables the use of the instrument, and commences a count of the number of tests performed by the breath tester. When the number of tests is exceeded, the instrument is disenabled. The device can include a moisture filter having an interface with the instrument, which prevents its operation if the filter is used beyond the recommended number of times, or if excess moisture renders it saturated. | 1. A calibration checking device for use with a gas analyzer, comprising: a calibration checking unit; anda calibration checking unit releasing a calibration checking gas of known composition into said gas analyzer; andan enabling mechanism actuated by use of said calibration checking unit and enabling operation of said gas analyzer. | 1. A calibration checking device for use with a gas analyzer, comprising: a calibration checking unit; anda calibration checking unit releasing a calibration checking gas of known composition into said gas analyzer; andan enabling mechanism actuated by use of said calibration checking unit and enabling operation of said gas analyzer. 2. A calibration checking device for use with a gas analyzer according to claim 1 and wherein said enabling mechanism is operative to count the number of tests performed by said gas analyzer. 3. A calibration checking device for use with a gas analyzer according to claim 1 and wherein said enabling mechanism is operative to accumulate the time of operation of said gas analyzer. 4. A calibration checking device for use with a gas analyzer according to claim 1 and also comprising a filter for removing fluids from a gas to be analyzed. 5. A calibration checking device for use with a gas analyzer according to claim 4 and wherein said enabling mechanism for enabling operation of said gas analyzer is operated by said filter. 6. A calibration checking device for use with a gas analyzer according to claim 4 and wherein said filter is a section of a sampling tube having built-in fluid filtering properties. 7. A calibration checking device for use with a gas analyzer according to claim 1 and wherein said enabling mechanism is communicative with said gas analyzer by means of a signal selected from a group consisting of electrical, electronic, optical, mechanical, magnetic, pneumatic and gaseous signals. 8. A calibration checking device for use with a gas analyzer according to claim 1 and wherein said enabling mechanism is operative to ensure proper location of said calibration checking unit. 9. A calibration checking device for use with a gas analyzer according to claim 1 and wherein said enabling mechanism comprises an optical transmitter and a receiver, the optical path between which is completed by reflection from said calibration checking unit only when said calibration checking unit is properly located in said gas analyzer. 10. A calibration checking device for use with a gas analyzer according to claim 1 and wherein said enabling mechanism is actuated by release of said calibration checking gas. 11. A calibration checking device for use with a gas analyzer according to claim 1 and wherein said enabling mechanism is actuated by means of an active integrated circuit disposed on said calibration checking device. 12. A calibration checking device for use with a gas analyzer according to claim 11 and also comprising a filter for removing fluids from a gas to be analyzed. 13. A calibration checking device for use with a gas analyzer according to claim 12 and wherein said filter comprises a drying agent disposed in proximity to at least part of an inside wall of said sampling tube. 14. A calibration checking device for use with a gas analyzer according to claim 1 and wherein said enabling mechanism is operative to a... | Oridion Breathid Ltd. | A61B005/08 G01N001/22
US06901262 | B2 | 2005-05-31 | Methods and systems for providing mobile subscriber surveillance | 2002-04-03 | Methods and systems for providing mobile subscriber surveillance identify certain call signaling messages as candidate messages for mobile subscriber surveillance. From the candidate messages, messages associated with mobile subscribers under surveillance are identified. Mobile subscriber location information is obtained for the messages associated with the mobile subscribers under surveillance. The location information is forwarded to a surveillance center, such as a state or federal law enforcement or security agency. The original call signaling messages are forwarded to their intended destinations so that surveillance is performed transparently to the mobile subscriber under surveillance. | 1. A method for providing mobile subscriber location information to a surveillance center, the method comprising: (a) receiving a first signaling message in a wireless communications network; (b) determining whether the first message is associated with a mobile subscriber under surveillance by comparing subscriber identification information extracted from the first message with mobile subscriber surveillance watch list information stored in a database; (c) in response to determining that the first message is associated with a mobile subscriber under surveillance, obtaining location information regarding the mobile subscriber; and (d) forwarding the location information to a surveillance center. | 1. A method for providing mobile subscriber location information to a surveillance center, the method comprising: (a) receiving a first signaling message in a wireless communications network; (b) determining whether the first message is associated with a mobile subscriber under surveillance by comparing subscriber identification information extracted from the first message with mobile subscriber surveillance watch list information stored in a database; (c) in response to determining that the first message is associated with a mobile subscriber under surveillance, obtaining location information regarding the mobile subscriber; and (d) forwarding the location information to a surveillance center. 2. The method of claim 1 wherein receiving a first message includes receiving a MAP update location message. 3. The method of claim 1 wherein receiving a first message includes receiving an RRLP measure position response message. 4. The method of claim 1 wherein receiving a first message includes receiving a GPRS mobility management message. 5. The method of claim 1 wherein determining whether the first message is associated with a mobile subscriber under surveillance includes extracting a mobile subscriber identifier from the first message and performing a lookup in a table containing identifiers associated with mobile subscribers under surveillance using the mobile subscriber identifier. 6. The method of claim 5 wherein performing a lookup using the mobile subscriber identifier comprises performing a lookup using an international mobile subscriber identifier (IMSI) extracted from the first message. 7. The method of claim 5 wherein performing a lookup using the mobile subscriber identifier comprises performing a lookup using a mobile station ISDN number (MSISDN) extracted from the first message. 8. The method of claim 5 wherein performing a lookup using the mobile subscriber identifier includes performing a lookup using an IP address extracted from the first message. 9. The method of claim 1 wherein obtaining location information includes extracting the location information from the first message. 10. The method of claim 1 wherein obtaining location information includes querying a mobile location center for the location information. 11. The method of claim 1 wherein obtaining location information includes querying an HLR for the location information. 12. The method of claim 1 wherein forwarding the location information to a surveillance center includes formulating a second message including the location information and sending the second message to the surveillance center over a network. 13. A method for providing mobile subscriber location information to a surveillance center, the method comprising: (a) receiving a first signaling message in a wireless communications network; (b) determining whether the first message is associated with a mobile subscriber under surveillance; (c) in response to determining that the first message is associated... | Tekelec | H04Q007/20 G01S003/02 A61B005/02
US06932772 | B2 | 2005-08-23 | Noninvasive blood pressure measuring method and apparatus | 2001-10-16 | A method and a device for non-invasive blood pressure measurement wherein the angle between the hand and the wrist, and the turning angle of the wrist relative to the middle part of the forearm, are kept to the most suitable degree for measuring the blood pressure of the radial artery. At least one pressure bladder and one arterial pulse transducer array are placed on the skin over the radial artery of the wrist to apply the external pressure to the artery and to detect the change of the arterial pulse signals. This method and device can correctly measure the intermittent or continuous blood pressure of the radial artery or the ulnar artery based on the principles of oscillation method and volume compensation method, and effectively eliminate the influence on the measurement due to body movement and the influence on blood circulation and neural function of the hand caused by long-term blood pressure measurement. | 1. A method for non-invasively measuring arterial blood pressure at a wrist of a patient, said method comprising the steps of: a) keeping a wrist of a patient at a posture which lowers a position of at least one tendon of the wrist near to a radial artery to be measured, and causes the radial artery to be close to a radius of the wrist; b) applying a changing external pressure to the skin above a point where the radial artery crosses a most protuberant spot on a volar aspect of the radius of the wrist; c) detecting a pulse wave signal of the radial artery along with a change in said external pressure on the skin above said point; d) measuring blood pressure of the radial artery by measuring said external pressure applied to the radial artery when said pulse wave signal changes. | 1. A method for non-invasively measuring arterial blood pressure at a wrist of a patient, said method comprising the steps of: a) keeping a wrist of a patient at a posture which lowers a position of at least one tendon of the wrist near to a radial artery to be measured, and causes the radial artery to be close to a radius of the wrist; b) applying a changing external pressure to the skin above a point where the radial artery crosses a most protuberant spot on a volar aspect of the radius of the wrist; c) detecting a pulse wave signal of the radial artery along with a change in said external pressure on the skin above said point; d) measuring blood pressure of the radial artery by measuring said external pressure applied to the radial artery when said pulse wave signal changes. 2. The method as defined in claim 1, wherein said step of keeping the wrist at said posture comprises forming a flexing angle of between approximately 100 and 170 degrees between a dorsal side of the wrist and a dorsal side of the hand attached to the wrist. 3. The method as defined in claim 1, wherein said step of keeping the wrist at said posture comprises forming a flexing angle of between approximately 100 and 170 degrees between a dorsal side of the wrist and a dorsal side of the hand attached to the wrist, and synchronously forming a turning angle of the wrist relative to the forearm attached to the wrist of between approximately 30 and 100 degrees towards a medial side of the patient's body. 4. The method as defined in claim 1, wherein said step of keeping the wrist at said posture comprises forming a flexing angle of between approximately 100 and 170 degrees between a dorsal side of the wrist and a dorsal side of the hand attached to the wrist, and synchronously forming a deflecting angle from a central line of the palm of the hand relative to a central line of a volar side of the wrist at between approximately 10 and 40 degrees towards the little finger of the hand. 5. The method as defined in claim 1, wherein said step of keeping the wrist at said posture comprises forming a flexing angle of between approximately 100 and 170 degrees between a dorsal side of the wrist and a dorsal side of the hand attached to the wrist, and synchronously forming a turning angle of the wrist relative to the forearm attached to the wrist of between approximately 30 and 100 degrees towards a medial side of the patient's body and a deflecting angle from a central line of the palm of the hand relative to a central line of a volar side of the wrist at between approximately 10 and 40 degrees towards the little finger of the hand. 6. The method as defined in claim 1, wherein said step of detecting a pulse wave signal of the radial artery includes measuring an oscillation in said external pressure caused by a pulsation of the radial artery. 7. The method as defined in claim 1, wherein said step of detecting a pulse wave signal of the radial artery includes measuring an oscillation i... | - | A61B005/02
US06955650 | B2 | 2005-10-18 | Metabolic calorimeter employing respiratory gas analysis | 2002-05-31 | An indirect calorimeter for measuring the metabolic rate of a subject includes a respiratory connector configured to be supported in contact with the subject so as to pass inhaled and exhaled gases as the subject breathes, a flow pathway, and a hygiene barrier positioned to block a predetermined pathogen from the exhaled gases. The indirect calorimeter also includes a flow pathway having a first end in fluid communication with the respiratory connector and a second end in fluid communication with a source and sink for respiratory gases. The flow pathway includes a flow tube through which the inhaled and exhaled gases pass, an outer housing surrounding the flow tube, and a chamber disposed between the flow tube and the first end. The indirect calorimeter also includes a flow meter configured to generate electrical signals as a function of the instantaneous flow volume of inhaled and exhaled gases passing through the flow pathway, and a component gas concentration sensor operable to generate electrical signals as a function of the instantaneous fraction of a predetermined component gas in the exhaled gases as the gases pass through the flow pathway. The indirect calorimeter further includes a computation unit operable to receive the electrical signals from the flow meter and the concentration sensor and operative to calculate at least one respiratory parameter for the subject as the subject breathes through the calorimeter. | 1. An indirect calorimeter, comprising: a respiratory connector configured to be supported in contact with a subject so as to pass inhaled and exhaled gases as said subject breathes; a hygiene barrier module operatively connected to said respiratory connector, such that said inhaled and exhaled gases pass therethrough said hygiene barrier module for blocking a predetermined pathogen from said exhaled gases; a flow pathway operatively connected to said hygiene barrier module for receiving and passing said inhaled and exhaled gases, said flow pathway having a first end in fluid communication with said respiratory connector and a second end in fluid communication with a source and sink for respiratory gases, said flow pathway comprising: (1) a flow tube through which said inhaled and exhaled gases pass, (2) an outer housing surrounding said flow tube, and (3) a chamber disposed between said flow tube and said first end, said chamber being a concentric chamber surrounding one end of said flow tube and being defined between said flow tube and said outer housing; a flow meter configured to generate electrical signals as a function of an instantaneous flow volume of said inhaled and exhaled gases passing through said flow pathway; a component gas concentration sensor operable to generate electrical signals as a function of an instantaneous fraction of a predetermined component gas in said exhaled gases passing through said flow pathway; and a computation unit operable to receive said electrical signals from said flow meter and said concentration sensor and to calculate at least one respiratory parameter for said subject as said subject breathes through said calorimeter. | 1. An indirect calorimeter, comprising: a respiratory connector configured to be supported in contact with a subject so as to pass inhaled and exhaled gases as said subject breathes; a hygiene barrier module operatively connected to said respiratory connector, such that said inhaled and exhaled gases pass therethrough said hygiene barrier module for blocking a predetermined pathogen from said exhaled gases; a flow pathway operatively connected to said hygiene barrier module for receiving and passing said inhaled and exhaled gases, said flow pathway having a first end in fluid communication with said respiratory connector and a second end in fluid communication with a source and sink for respiratory gases, said flow pathway comprising: (1) a flow tube through which said inhaled and exhaled gases pass, (2) an outer housing surrounding said flow tube, and (3) a chamber disposed between said flow tube and said first end, said chamber being a concentric chamber surrounding one end of said flow tube and being defined between said flow tube and said outer housing; a flow meter configured to generate electrical signals as a function of an instantaneous flow volume of said inhaled and exhaled gases passing through said flow pathway; a component gas concentration sensor operable to generate electrical signals as a function of an instantaneous fraction of a predetermined component gas in said exhaled gases passing through said flow pathway; and a computation unit operable to receive said electrical signals from said flow meter and said concentration sensor and to calculate at least one respiratory parameter for said subject as said subject breathes through said calorimeter. 2. The calorimeter according to claim 1, wherein said hygiene barrier module includes a hygiene barrier housing including two generally parallel and spaced apart side walls and a hygiene barrier disposed between said side walls. 3. The calorimeter according to claim 2, wherein said hygiene barrier is a filtering material. 4. The calorimeter according to claim 2, wherein said hygiene barrier housing includes a lower wall and a saliva retention wall extending upwardly from said lower wall and adjacent said hygiene barrier to form a saliva trap. 5. The calorimeter according to claim 2, wherein said hygiene barrier module includes a respiration port for operatively connecting said hygiene barrier module to a corresponding port for said respiratory connector and a calorimeter port for operatively connecting said hygiene barrier module to a corresponding inlet conduit in fluid communication with said flow pathway. 6. The calorimeter according to claim 1, wherein said respiratory connector is a mask. 7. The calorimeter according to claim 1, wherein said hygiene barrier module and said respiratory connector are disposable. 8. The calorimeter according to claim 1, wherein said respiratory connector is a mouthpiece. 9. An indirect calorimeter, comprising: a respiratory connector configure... | HealtheTech, Inc. | A61B005/08
US06843774 | B2 | 2005-01-18 | Technique for diagnosing attention deficit hyperactivity disorder | 2002-11-21 | This invention features a method of determining whether an individual has Attention Deficit Hyperactivity Disorder by sampling the peripheral skin temperature of an individual in an inactive state and manipulating the resulting temperature data to produce a value which is indicative of whether or not the individual has ADHD. | 1. A method of determining whether a subject has Attention Deficit Hyperactivity Disorder, comprising: sampling the peripheral skin temperature of the subject during a pre-determined time interval when the subject is in an inactive state to provide sampled peripheral skin temperature data; dividing the data into a series of windows of equal time intervals across the pre-determined time interval; for each window applying a fast Fourier Transform algorithm (FFT) to produce a data set of FFT magnitudes and determining a range of FFT magnitude variation, which is the difference between the FFT magnitude at a maximum frequency fmax and the FFT magnitude at the minimum frequency fmin; and aggregating the magnitude ranges for one or more of the windows to produce an aggregate magnitude range value which is indicative of whether or not the subject has ADHD; wherein after said sampling, temperature data sampled during a noise event is eliminated through manual or automatic means. | 1. A method of determining whether a subject has Attention Deficit Hyperactivity Disorder, comprising: sampling the peripheral skin temperature of the subject during a pre-determined time interval when the subject is in an inactive state to provide sampled peripheral skin temperature data; dividing the data into a series of windows of equal time intervals across the pre-determined time interval; for each window applying a fast Fourier Transform algorithm (FFT) to produce a data set of FFT magnitudes and determining a range of FFT magnitude variation, which is the difference between the FFT magnitude at a maximum frequency fmax and the FFT magnitude at the minimum frequency fmin; and aggregating the magnitude ranges for one or more of the windows to produce an aggregate magnitude range value which is indicative of whether or not the subject has ADHD; wherein after said sampling, temperature data sampled during a noise event is eliminated through manual or automatic means. 2. The method of claim 1 wherein said aggregate magnitude range value is the mean of said magnitude range values of all of said windows. 3. The method of claim 1 wherein said aggregate magnitude range value is the median of said magnitude range values of all of said windows. 4. The method of claim 1 wherein said aggregate magnitude range value is the variance of said magnitude range values of all of said windows. 5. The method of claim 1 wherein said aggregate magnitude range value is determined by an aggregation statistical algorithm of said magnitude range values of all of said windows. 6. The method of claim 1 wherein said sampling includes sampling the peripheral skin temperature of said subject during two or more pre-determined time intervals, wherein said aggregate magnitude range value is aggregated for one or more of said pre-determined time intervals. 7. The method of claim 1 wherein said sampling includes sampling the peripheral skin temperature of one or both hands of said subject. 8. The method of claim 1 wherein said the magnitude range of one or more windows at the beginning of said pre-determined time interval is not used to produce said aggregate magnitude range value. 9. The method of claim 1 wherein environmental data including subject motion, audible environment and ambient temperature data are recorded. 10. The method in claim 9 wherein said environmental data is synchronized to the temperature data. 11. The method of claim 9 wherein environmental data are manually analyzed by a test evaluator for the occurrence of one or more noise events. 12. The method of claim 11 wherein temperature data sampled during a noise event is eliminated prior to calculating the aggregate magnitude range value. 13. The method of claim 9 wherein said environment data is automatically evaluated using scene change algorithms and peak noise detection to determine whether a noise event occurred during the sampling. 14. The method of claim 13 wherein automatic analysis tec... | The Mclean Hospital Corporation | A61B005/00
US07233818 | B1 | 2007-06-19 | Methods and apparatus for mapping internal and bulk motion of an object with phase labeling in magnetic resonance imaging | 2000-08-04 | Magnetic resonance imaging method and apparatus are provided for mapping the internal or bulk motion of an object by labeling the phase of a specimen magnetization with a selected spatial function and measuring changes in the phase of the magnetization. The spatial function is selectable to provide magnetization phase modulation corresponding to displacements in a selected direction, such as a radial or azimuthal direction. Methods and apparatus for producing images based on magnetization phase modulation acquire image data based on stimulated echos and stimulated anti-echos. In an embodiment, a series of 180 degree pulses produces alternating stimulated and stimulated anti-echos that are measured and assigned to respective images. | 1. A magnetic resonance imaging method, comprising: applying a first excitation radio-frequency (RF) pulse along a first excitation axis to produce a first transverse magnetization;applying a first position encoding phase label to the first transverse magnetization;storing a component of the position encoded first transverse magnetization as a first stored longitudinal magnetization by applying a first storing RF pulse along a first storing axis;allowing a first mixing time interval of duration TM to elapse;after the first mixing time has elapsed, producing a transverse magnetization based on the first stored longitudinal magnetization;decoding the transverse magnetization produced after the first mixing time;generating a first image signal based on the decoded transverse magnetization produced after the first mixing time;applying a second excitation RF pulse along a second excitation axis to produce a second transverse magnetization;applying a second position encoding phase label to the second transverse magnetization;storing a component of the position encoded second transverse magnetization as a second stored longitudinal magnetization by applying a second storing RF pulse along a second storing axis, wherein a difference between an angle from the first excitation axis to the first storing axis and an angle from the second excitation axis to the second storing axis is not an integer multiple of 180 degrees;allowing a second mixing time of duration TM to elapse;after the second mixing time has elapsed, producing a transverse magnetization based on the second stored longitudinal magnetization;decoding the transverse magnetization produced after the second mixing time;generating a second image signal based on the decoded transverse magnetization produced after the second mixing time; andcombining the first image signal and the second image signal to reduce a contribution associated with a complex conjugate of a phase-labeled magnetization to obtain an artifact-reduced image of a specimen. | 1. A magnetic resonance imaging method, comprising: applying a first excitation radio-frequency (RF) pulse along a first excitation axis to produce a first transverse magnetization;applying a first position encoding phase label to the first transverse magnetization;storing a component of the position encoded first transverse magnetization as a first stored longitudinal magnetization by applying a first storing RF pulse along a first storing axis;allowing a first mixing time interval of duration TM to elapse;after the first mixing time has elapsed, producing a transverse magnetization based on the first stored longitudinal magnetization;decoding the transverse magnetization produced after the first mixing time;generating a first image signal based on the decoded transverse magnetization produced after the first mixing time;applying a second excitation RF pulse along a second excitation axis to produce a second transverse magnetization;applying a second position encoding phase label to the second transverse magnetization;storing a component of the position encoded second transverse magnetization as a second stored longitudinal magnetization by applying a second storing RF pulse along a second storing axis, wherein a difference between an angle from the first excitation axis to the first storing axis and an angle from the second excitation axis to the second storing axis is not an integer multiple of 180 degrees;allowing a second mixing time of duration TM to elapse;after the second mixing time has elapsed, producing a transverse magnetization based on the second stored longitudinal magnetization;decoding the transverse magnetization produced after the second mixing time;generating a second image signal based on the decoded transverse magnetization produced after the second mixing time; andcombining the first image signal and the second image signal to reduce a contribution associated with a complex conjugate of a phase-labeled magnetization to obtain an artifact-reduced image of a specimen. 2. The method of claim 1, further comprising estimating specimen motion based on the combined first image signal and second image signal. 3. The method of claim 2, wherein the first excitation axis and the first storing axis are orthogonal. 4. The method of claim 2, further comprising: dividing the mixing time intervals of duration TM into a first interval of duration t1 and a second interval of duration t2, where the first and second intervals are selected based on at least one longitudinal decay time T1; andapplying a 180° radio-frequency (RF) pulse to the specimen after the time interval t1 during acquisition of both the first image signal and the second image signal to reduce contributions associated with free induction decay. 5. The method of claim 4, wherein time points within the mixing time intervals divide the mixing time intervals into a first interval of duration t1 and a second interval of duration t2 such that t2=ln [2/(1+exp(−TM/T1))], wherein ... | United States of America as represented by the Secretary of the Department of Health and Human Services | A61B005/055
US07087047 | B2 | 2006-08-08 | Predonation blood sampling apparatus | 2001-02-13 | This invention discloses a donor blood donation and sampling system including a donor needle adapted for drawing blood from a body, a blood collection bag coupled to a blood collection conduit which is coupled to the donor needle, a blood sampling conduit also coupled to the donor needle, and a sampling tube assembly arranged for selectable fluid engagement with the blood sampling conduit and for selectable clamping engagement with at least said blood collection conduit. A donor blood donation and sampling method is also disclosed. | 1. A donor blood donation and sampling system comprising: a donor needle adapted for drawing blood from a body;a blood collection bag coupled to a blood collection conduit which is coupled to said donor needle;a blood sampling conduit also coupled to said donor needle;a sampling tube assembly including: a housing element; anda clamping assembly disposed within said housing element and arranged for selectable fluid engagement with said blood sampling conduit and for selectable clamping engagement with at least said blood collection conduit, thereby selectably stopping blood flow in at least said blood collection conduit; anda blood supply conduit coupled to said donor needle, said blood collection conduit and said blood sampling conduit each being coupled to said blood supply conduit. | 1. A donor blood donation and sampling system comprising: a donor needle adapted for drawing blood from a body;a blood collection bag coupled to a blood collection conduit which is coupled to said donor needle;a blood sampling conduit also coupled to said donor needle;a sampling tube assembly including: a housing element; anda clamping assembly disposed within said housing element and arranged for selectable fluid engagement with said blood sampling conduit and for selectable clamping engagement with at least said blood collection conduit, thereby selectably stopping blood flow in at least said blood collection conduit; anda blood supply conduit coupled to said donor needle, said blood collection conduit and said blood sampling conduit each being coupled to said blood supply conduit. 2. A donor blood donation and sampling system according to claim 1 and also comprising a breakaway cannula located on said blood collection conduit. 3. A donor blood donation and sampling system according to claim 2 and wherein said sampling tube assembly also comprises: a needle having first and second sharpened points at its opposite ends;a vacuum sampling tube receiving socket for removably receiving a vacuum sampling tube in fluid engagement with said needle at said second sharpened point, thereby providing fluid communication between blood in said blood sampling conduit and the interior of said vacuum sampling tube; andresilient material disposed about said first sharpened point for preventing inadvertent piercing of said blood sampling conduit by said needle prior to clamping engagement therebetween,wherein said clamping assembly is arranged for selectable clamping engagement of said blood sampling conduit in fluid engagement with said needle, said clamping engagement causing piercing of said blood sampling conduit by said first sharpened point of said needle and wherein said resilient material comprises a sealing material operative in association with said blood sampling conduit and said needle for preventing leakage of blood from said blood sampling conduit after piercing thereof by said needle. 4. A donor blood donation and sampling system according to claim 3 and wherein said resilient material comprises a foam material. 5. A donor blood donation and sampling system according to claim 3 and wherein said sampling tube assembly is operative upon selectable clamping engagement with said blood sampling conduit and a blood collection conduit to simultaneously pierce said blood sampling conduit and block said blood collection conduit. 6. A donor blood donation and sampling system according to claim 3 and wherein said clamping assembly is also operative upon selectable clamping engagement with said blood sampling conduit to break said breakaway cannula, thereby permitting blood flow therethrough. 7. A donor blood donation and sampling system according to claim 3 and wherein said clamping assembly is operative to have the following functionalities: a. breakin... | Teva Medical Ltd. | A61B019/00 A61B005/00 A61M037/00 A61M001/00
US07263710 | B1 | 2007-08-28 | Medical diagnostic system with on-line real-time video training | 1999-12-31 | A central service facility is connected via a network and a uniform service interactive platform to a multiplicity of remotely located medical diagnostic systems. The interactive platform includes a video/audio plug-in application which can be used to access a library of videos via a network. The video library is stored at the central service facility or on a local area network at the remote facility. The interactive platform of the medical diagnostic system comprises a web browser which enables the end-user to search and select a training video from the database, and a video/audio player for viewing the selected video directly from the operator's console. Depending on bandwidth availability, the video file can be downloaded to the remotely located medical diagnostic equipment prior to display or streamed in real-time. | 1. A system comprising a central service facility connected to a remotely located medical diagnostic imaging system of a particular imaging modality via a network, wherein said central service facility comprises a memory that stores a video library comprising a multiplicity of training videos, each training video being classified as a member of one of a plurality of training video classes, each class corresponding to a respective imaging modality, said particular imaging modality belonging to a first training video class, and each training video showing one or more of the following: how to operate a medical diagnostic imaging system of said respective imaging modality, how to perform an examination of a particular exam type on a patient using a medical diagnostic imaging system of said respective imaging modality or how to make a diagnosis during an examination of said particular exam type performed using a medical diagnostic imaging system of said respective imaging modality;wherein said remotely located medical diagnostic imaging system comprises a display monitor and a web browser for displaying interactive web pages on said display monitor, said web browser being programmed to display a first interactive web page comprising a plurality of operational mode graphical buttons, only one of which is useable for navigating to a second interactive web page to be displayed on said display monitor; to display said second interactive web page on said display monitor in response to selection of said one operational mode graphical button, said second interactive web page comprising a plurality of imaging modality graphical buttons each of which corresponds to a respective one of said plurality of training video classes, including a first imaging modality graphical button corresponding to said first training video class; and to display an updated version of said second interactive web page on said display monitor in response to clicking on said first imaging modality graphical button, said updated version of said second interactive web page comprising a plurality of training video graphical buttons each of which corresponds to a respective one of a plurality of training videos belonging to said first imaging modality class;wherein said remotely located medical diagnostic imaging system further comprises means for sending a training video request to said central service facility via said network in response to selection of a training video belonging to said first imaging modality class by clicking on one of said plurality of training video graphical buttons displayed on said display monitor, said training video request comprising an identifier identifying said selected training video and a source system identifier identifying said medical diagnostic imaging system;wherein said central service facility further comprises: means for verifying whether the medical diagnostic imaging system identified by said source system identifier has a valid subscription in ... | 1. A system comprising a central service facility connected to a remotely located medical diagnostic imaging system of a particular imaging modality via a network, wherein said central service facility comprises a memory that stores a video library comprising a multiplicity of training videos, each training video being classified as a member of one of a plurality of training video classes, each class corresponding to a respective imaging modality, said particular imaging modality belonging to a first training video class, and each training video showing one or more of the following: how to operate a medical diagnostic imaging system of said respective imaging modality, how to perform an examination of a particular exam type on a patient using a medical diagnostic imaging system of said respective imaging modality or how to make a diagnosis during an examination of said particular exam type performed using a medical diagnostic imaging system of said respective imaging modality;wherein said remotely located medical diagnostic imaging system comprises a display monitor and a web browser for displaying interactive web pages on said display monitor, said web browser being programmed to display a first interactive web page comprising a plurality of operational mode graphical buttons, only one of which is useable for navigating to a second interactive web page to be displayed on said display monitor; to display said second interactive web page on said display monitor in response to selection of said one operational mode graphical button, said second interactive web page comprising a plurality of imaging modality graphical buttons each of which corresponds to a respective one of said plurality of training video classes, including a first imaging modality graphical button corresponding to said first training video class; and to display an updated version of said second interactive web page on said display monitor in response to clicking on said first imaging modality graphical button, said updated version of said second interactive web page comprising a plurality of training video graphical buttons each of which corresponds to a respective one of a plurality of training videos belonging to said first imaging modality class;wherein said remotely located medical diagnostic imaging system further comprises means for sending a training video request to said central service facility via said network in response to selection of a training video belonging to said first imaging modality class by clicking on one of said plurality of training video graphical buttons displayed on said display monitor, said training video request comprising an identifier identifying said selected training video and a source system identifier identifying said medical diagnostic imaging system;wherein said central service facility further comprises: means for verifying whether the medical diagnostic imaging system identified by said source system identifier has a valid subscription in ... | General Electric Company | H04N007/173 H04N007/18 H04N007/16 G09B023/28 G09B007/00 A61B005/00
US07249422 | B2 | 2007-07-31 | Device by gioniometric measurements | 2002-12-30 | A device (1) for measuring the relative orientation according to at least one degree of freedom for two objects, including a constraint generator (10) suitable for causing a goniometric sensor (40) to move in a plane, having the function of measuring the variation of relative orientation of the two objects in this plane. The goniometric sensor (40) is arranged in a housing (41) that crosses longitudinally the constraint generator (10), which has high flexional stiffness in a first longitudinal plane (β) and a low flexional stiffness in a second longitudinal plane (φ) orthogonal to the first (β). The sensor measures rotations in a plane and the constraint generator induces a rotation in that plane. With the device (1) a data suit (50) can be made for measuring the movement of limbs of an individual. For example, arranging three devices (10a, 10b, 10c) in series, but capable of measuring angles in orthogonal planes, the rotation can be measured of the arm (76) with respect to a shoulder (75) of an individual. | 1. Device for measuring relative orientation of at least a first object with respect to a second object free from said first object, or connected to the first object but with independent movement, wherein said orientation is carried out according to at least one degree of freedom, comprising: at least one goniometric sensor for each said degree of freedom, constructed and arranged for measuring variation of orientation in a plane;at least one constraint generator for each said goniometric sensor, constructed and arranged to cause the sensor to move in said plane;wherein each said constraint generator comprises a flexible elongated element having a longitudinal axis with a low flexional stiffness in a first plane passing through said axis, and a high flexional stiffness in a second plane orthogonal to said first plane and passing through the axis, said element being flexible only in said first plane. | 1. Device for measuring relative orientation of at least a first object with respect to a second object free from said first object, or connected to the first object but with independent movement, wherein said orientation is carried out according to at least one degree of freedom, comprising: at least one goniometric sensor for each said degree of freedom, constructed and arranged for measuring variation of orientation in a plane;at least one constraint generator for each said goniometric sensor, constructed and arranged to cause the sensor to move in said plane;wherein each said constraint generator comprises a flexible elongated element having a longitudinal axis with a low flexional stiffness in a first plane passing through said axis, and a high flexional stiffness in a second plane orthogonal to said first plane and passing through the axis, said element being flexible only in said first plane. 2. Device according to claim 1, wherein said constraint generator has high stiffness to torsion and to tensile stress. 3. Device according to claim 1, wherein each said goniometric sensor is disposed in a housing that crosses longitudinally said constraint generator, said goniometric sensor measuring the relative angular movement of said objects in a bending plane of said constraint generator. 4. Device according to claim 1, wherein each said constraint generator comprises a plurality of substantially parallel portions of larger cross section alternating with a plurality of portions of smaller cross section, said portions of smaller cross sections conferring to the structure a larger flexional capacity in a first plane and larger flexional stiffness in a plane orthogonal to the first plane, and a high stiffness to torsion and to tensile stress. 5. Device according to claim 1, wherein each said constraint generator comprises an plate shaped element having bellow-like projections extending therefrom. 6. Device according to claim 5, wherein said projections have structure selected from the group consisting of helical and alternating annular portions having larger and smaller cross sections. 7. Device according to claim 1, wherein at least one said constraint generator comprises flanges orthogonal to said axis for fastening at ends thereof further constraint generators in series or for fastening to objects whose rotation will be measured. 8. Device according to claim 1, wherein said goniometric sensor comprises: a flexible elongated element extending between a first and a second object, said element having a neutral axis which does not change in length when bending, and at least one fiber spaced apart from said neutral axis and extending from the first to the second object;means for measuring the length variation of said fiber as the relative rotation varies between the first and the second object, the relative rotation being proportional to said length variation. 9. Device for measuring rotation of a wrist of an individual with respect to a point... | - | A61B005/103
US06981950 | B2 | 2006-01-03 | Noninvasive intraductal fluid sampling device | 2002-02-08 | Disclosed is an apparatus for noninvasive intraductal fluid aspiration. Intraductal fluid is noninvasively aspirated using compression, heating and suction cycles. The device is configured to provide a compression force, deliver heat, and apply a vacuum pressure to encourage the aspiration of patient fluid. The device includes an adjustable support and a movable wall for applying compression. Withdrawn fluid may be used either for diagnostic or therapeutic purposes. | 1. A nipple aspirate fluid aspiration device, comprising: an adjustable support, defining a concavity, the support comprising a plurality of petals, movable throughout an adjustment range; at least three inflatable bladders within the concavity; a vacuum source in communication with the concavity; a heat source; and a fluid circulation pathway for circulating a fluid through the bladders. | 1. A nipple aspirate fluid aspiration device, comprising: an adjustable support, defining a concavity, the support comprising a plurality of petals, movable throughout an adjustment range; at least three inflatable bladders within the concavity; a vacuum source in communication with the concavity; a heat source; and a fluid circulation pathway for circulating a fluid through the bladders. 2. A nipple aspirate fluid aspiration device as in claim 1, wherein each petal carries an inflatable bladder. 3. A nipple aspirate fluid aspiration device as in claim 1, wherein the heat source is in thermally conductive contact with the bladder. 4. A nipple aspirate fluid aspiration device as in claim 1, wherein the heat source is in thermally conductive contact with the fluid so that the fluid heats the bladder. 5. A nipple aspirate fluid aspiration device as in claim 1, further comprising a control for inflating and deflating the bladder in accordance with a predetermined program. 6. A nipple aspirate fluid aspiration device as in claim 5 wherein the predetermined program comprises alternating inflation and deflation cycles. 7. A nipple aspirate fluid aspiration device as in claim 6 wherein the predetermined program inflates the bladder within the range of from about 2 to about 40 cycles per minute. 8. A nipple aspirate fluid aspiration device as in claim 7 wherein the predetermined program inflates the bladder within the range of from about 3 to about 12 cycles per minute. 9. A nipple aspirate fluid aspiration device as in claim 6 wherein the predetermined program maintains the bladder inflated within the range of from about 4 to about 8 seconds per cycle. 10. A nipple aspirate fluid aspiration device as in claim 1, wherein the bladder is inflatable from a reduced profile along an axis transverse to the support and an inflated profile along the axis. 11. A nipple aspirate fluid aspiration device as in claim 10, wherein the bladder has a maximum thickness in the inflated profile along the axis within the range of from about 0.2 inches to about 2.0 inches. 12. A nipple aspirate fluid aspiration device as in claim 1, wherein the bladder bas an inflated thickness of no more than about 1 inch. 13. A nipple aspirate fluid aspiration device as in claim 12, wherein the bladder has an inflated thickness of no more than about 0.5 inches. 14. A nipple aspirate fluid aspiration device as in claim 1, further comprising a central processing unit for controlling the inflatable bladders. 15. A nipple aspirate fluid aspiration device as in claim 1, further comprising: a control unit; a patient interface unit, carrying the adjustable support; and a control line extending between the control unit and the patient interface unit. 16. A nipple aspirate fluid aspiration device as in claim 15, further comprising a closed fluid circulation loop, having a reservoir removably carried by the control unit in communication with the bladder carried by the patient interfac... | Neomatrix, LLC | A61B005/00 B65D081/00
US07269516 | B2 | 2007-09-11 | Systems and methods for monitoring behavior informatics | 2002-05-15 | A system and method used to assess animal behavior includes a module having sensors that collects a variety of physical and biological data from a test subject. Interpretation of the data is provided to assess the test subject's behavior, neurology, biochemistry and physiology. The module is useful in observing the effects of a drug on the test animal and providing information on the drug's signature. Another advantage is the module's portability that allows it to be used in standard laboratory cages. This portability allows the animal to be tested in its own habitat, that can reduce any erroneous data due to stressing the animal when removed to a test cage. Additionally, the module's design allows for parallel data collection and interpretation from several laboratory animals undergoing different experiments. Multi-dimensional modeling of the test subject based the system's interpretation of the data allows pattern recognition of the drug signature, and predictive drug analysis. | 1. A method of predicting whether a test treatment belongs to a class of test treatments having a pattern of animal behavior associated with said class of test treatments, said method comprising: (a) generating an electronic database of stored class signature information of at least one treatment class signature by (i) collecting behavioral data from animals administered at least one known treatment which is a member of a class of treatments; and (ii) using a computer to define said signatures for said treatment class by correlating features of said behavioral data with said known treatment class member; and (iii) storing said signature information in a computer;(b) administering said test treatment to at least one animal;(c) using a computer to capture behavioral data of said at least one animal administered said test treatment;(d) using a computer to compare information derived from an analysis of the captured data from (c) with stored class signature information from (a) to predict a probability that the test treatment is a member of at least one treatment class represented in the stored class signature information; and(e) outputting a result from (d) that is indicative of whether said test treatment is a member of at least one treatment class. | 1. A method of predicting whether a test treatment belongs to a class of test treatments having a pattern of animal behavior associated with said class of test treatments, said method comprising: (a) generating an electronic database of stored class signature information of at least one treatment class signature by (i) collecting behavioral data from animals administered at least one known treatment which is a member of a class of treatments; and (ii) using a computer to define said signatures for said treatment class by correlating features of said behavioral data with said known treatment class member; and (iii) storing said signature information in a computer;(b) administering said test treatment to at least one animal;(c) using a computer to capture behavioral data of said at least one animal administered said test treatment;(d) using a computer to compare information derived from an analysis of the captured data from (c) with stored class signature information from (a) to predict a probability that the test treatment is a member of at least one treatment class represented in the stored class signature information; and(e) outputting a result from (d) that is indicative of whether said test treatment is a member of at least one treatment class. 2. The method of claim 1 wherein the prediction is generated based on matching one or more elements of the data of step (c) to one or more elements of at least one stored class signature. 3. The method of claim 1 wherein the signatures are continuously updated by adding behavioral data of one or more animals affected by at least one known test treatment which is a member of a class of treatments. 4. The method of claim 1 wherein one or more animals affected by a known test treatment is subject to at least one mechanical challenge. 5. The method of claim 1 wherein one or more animals administered a test treatment is subject to at least one mechanical challenge. 6. The method of claim 1, wherein the test treatment comprises: administering a drug. 7. The method of claim 1, wherein the test treatment comprises: introducing a genetic mutation. 8. The method of claim 1, wherein the test treatment comprises: introducing a lesion. 9. The method of claim 1, wherein the test treatment comprises: performing a surgical procedure. 10. The method of claim 1, wherein the test treatment comprises: exposure to a chemical. 11. The method of claim 1, wherein the test treatment comprises: exposure to an environmental change. 12. The method of claim 1, wherein the test treatment is performed shortly before automatic capture of the information. 13. The method of claim 1, wherein the test treatment is performed chronically before automatic capture of the information. 14. The method of claim 1, wherein the test treatment is performed chronically during automatic capture of the information. 15. The method of claim 1, wherein the test treatment is performed during an early developmental stage of the animals. 16. T... | PsychoGenics, Inc. Carnegie Mellon University | G06F019/00 G06K009/00 A61B005/00
US07112175 | B2 | 2006-09-26 | Tele-diagnostic device | 2001-12-18 | A system (10), probe, and method for collecting a plurality of diagnostic information and transmitting the diagnostic information to a remote location. The system (10) comprises a member (12) contoured to at least a portion of a person's hand, and an interface unit (20) in electrical communication with the member (12). The interface unit (20) is capable of transmitting information to a remote location. The member (12) comprises at least eight sensors. | 1. A system for collecting diagnostic information and transmitting the diagnostic information to a remote location, the system comprising: a member contoured to at least a portion of a person's hand, the member comprising at least an EKG diagnostic device, the diagnostic device comprising at least eight EKG sensors, the member comprising a palm portion, a wrist portion and a plurality of phalange portions with the EKG sensors being located on the member on at least two of the palm portion, the wrist portion, and at least one of the phalange portions; andan interface unit in electrical communication with the member, wherein the interface unit is capable of transmitting information to a remote location; wherein the plurality of phalange portions comprise an index finger phalange portion and a middle finger phalange portion, wherein the index finger phalange portion is at least as long as the middle finger phalange portion of the member. | 1. A system for collecting diagnostic information and transmitting the diagnostic information to a remote location, the system comprising: a member contoured to at least a portion of a person's hand, the member comprising at least an EKG diagnostic device, the diagnostic device comprising at least eight EKG sensors, the member comprising a palm portion, a wrist portion and a plurality of phalange portions with the EKG sensors being located on the member on at least two of the palm portion, the wrist portion, and at least one of the phalange portions; andan interface unit in electrical communication with the member, wherein the interface unit is capable of transmitting information to a remote location; wherein the plurality of phalange portions comprise an index finger phalange portion and a middle finger phalange portion, wherein the index finger phalange portion is at least as long as the middle finger phalange portion of the member. 2. The system of claim 1 wherein the EKG sensors are located on the member on the palm portion, the wrist portion and at least one of the phalange portions. 3. The system of claim 1 wherein the EKG diagnostic device comprises at least 10 sensors. 4. The system of claim 3 wherein the EKG diagnostic device comprises 11 sensors. 5. The system of claim 1 wherein the index finger phalange portion is longer than the middle finger phalange portion of the member. 6. The system of claim 1 wherein the plurality of phalange portions comprise an index finger phalange portion and a middle finger phalange portion, at least four of the EKG sensors are located on the index finger phalange portion of the member. 7. The system of claim 1 wherein at least five of the EKG sensors are located on the index finger phalange portion of the member. 8. The system of claim 1 wherein the plurality of phalange portions comprise a thumb portion, with at least one of the EKG sensors being located on the thumb portion of the member. 9. The system of claim 1 wherein the plurality of phalange portions comprise a pinky finger portion, with at least one of the EKG sensors being located on the pinky finger portion of the member. 10. The system of claim 1 wherein at least one of the EKG sensors is located on a palmer surface of the palm portion of the member. 11. The system of claim 1 wherein the member has a shape that corresponds to at least a substantial portion of a person's hand such that the member is capable of being worn on a person's hand. 12. The system of claim 11 wherein the member has a portion shaped to contour to a person's palm. 13. The system of claim 11 wherein the member has a portion shaped to contour to a person's finger. 14. The system of claim 13 wherein the member has a portion shaped to contour to a person's palm. 15. The system of claim 11 wherein the member comprises a plurality of diagnostic devices. 16. The system of claim 15 wherein the plurality of diagnostic devices includes the EKG diagnostic device, a bloo... | Ineedmd.com | A61B005/02 A61B005/0205 A61B005/024 A61B005/0402 A61B095/01
US07011633 | B2 | 2006-03-14 | Blood flow measuring apparatus | 2002-01-10 | A blood flow measuring apparatus has a measuring electrode and a counter-electrode adapted to be brought into contact with the blood of a patient. A measuring unit measures a voltage or a current between the measuring electrode and the counter-electrode and determines the flow of blood therefrom. The measuring electrode is ring-shaped and sized to fit within a blood vessel in the body of the patient. An inner surface of the ring-shaped electrode forms an active measuring electrode surface. The blood flow measuring apparatus is suitable for use in a cardiac pacemaker. | 1. A blood flow measuring apparatus comprising: a measuring electrode and a counter electrode adapted for in vivo contact with blood of a patient;said measuring electrode being ring-shaped and having a size and dimensions adapted to fit within a blood vessel of the patient, and having an inner surface forming an active measuring electrode surface; anda measuring unit connected to said measuring electrode and to said counter-electrode for measuring an electrical quantity, selected from the group consisting of voltage and current, between said measuring electrode and said counter-electrode, and for determining a flow of blood dependent on said electrical quantity. | 1. A blood flow measuring apparatus comprising: a measuring electrode and a counter electrode adapted for in vivo contact with blood of a patient;said measuring electrode being ring-shaped and having a size and dimensions adapted to fit within a blood vessel of the patient, and having an inner surface forming an active measuring electrode surface; anda measuring unit connected to said measuring electrode and to said counter-electrode for measuring an electrical quantity, selected from the group consisting of voltage and current, between said measuring electrode and said counter-electrode, and for determining a flow of blood dependent on said electrical quantity. 2. A blood flow measuring apparatus as claimed in claim 1 wherein said electrical quantity is current, and wherein said measuring unit maintains a constant voltage between said measuring electrode and said counter-electrode while measuring said current. 3. A blood flow measuring apparatus as claimed in claim 1 wherein said active surface of said measuring electrode is of a same order of magnitude in size as said counter-electrode. 4. A blood flow measuring apparatus as claimed in claim 1 wherein said ring-shaped electrode is comprised of plastic material, and comprising an electrically conducting electrode material at least partially coating said inner surface. 5. A blood flow measuring apparatus as claimed in claim 4 wherein said electrically conducting electrode material comprises a platinum foil attached to said inner surface. 6. A blood flow measuring apparatus as claimed in claim 1 wherein said measuring electrode has an outer surface, and comprising insulation covering said outer surface. 7. A blood flow measuring apparatus as claimed in claim 1 wherein said measuring electrode comprises a metal net. 8. A blood flow measuring apparatus as claimed in claim 1 wherein said measuring electrode is foldable to reduce a diameter of said measuring electrode for implantation. 9. A blood flow measuring apparatus as claimed in claim 8 wherein said measuring electrode is comprised of a plurality of mutually foldable segments. 10. A blood flow measuring apparatus as claimed in claim 8 further comprising a catheter adapted for insertion into the blood vessel, and a stylet insertable into said catheter, and wherein said measuring electrode, when folded for implantation, is insertable into said catheter and is adapted to be forced out of said catheter by said stylet, to re-assume said ring-shape. 11. A blood flow measuring apparatus as claimed in claim 1 wherein said measuring electrode has a slot allowing said measuring electrode to be wound together to reduce a cross-sectional area of said measuring electrode for implantation. 12. A blood flow measuring apparatus as claimed in claim 8 further comprising a catheter adapted for insertion into the blood vessel, and a stylet insertable into said catheter, and wherein said measuring electrode, when wound for implantation, is insertable into... | St. Jude Medical AB | A61B005/02
US07295988 | B1 | 2007-11-13 | Computer system for optical scanning, storage, organization, authentication and electronic transmitting and receiving of medical records and patient information, and other sensitive legal documents | 2000-05-25 | The invention disclosed herein relates to an improved method and system for the optically scanning, storage, management, retrieval and electronic mailing of a persons medical records and identification information on a 24 hour a day basis, primarily for use in a medical emergency of other medical scenario. The invention disclosed also encompasses unique and novel methods of verifying the authenticity of original medical records via a unique physicians digital signature embedded into the documents, a means for standardizing and prioritizing the history and prior medical records of a patient so as to provide an edited or abbreviated medical chart for emergencies which is current and clinically significant, a means of encrypting medical records for security, and a means of providing a unique alpha numerical identified code for each patient and digitally embedding the identifier into said patient records within the system and a means of 24 hour a day electronic access, transmission and updating of said records using a unique telephone exchange system, Internet, wesbite Intranet or other appropriate electronic or wireless means. | 1. A computer system for inputting, storing, organizing, retrieving, and authenticating, medical records, clinical data, and patient data, the system comprising: an optical scanner for converting medical records, clinical data, and patient data to digital records using a digitizing processsaid optical scanner creating a digital data matrix layer of said digital recordssaid optical scanner simultaneously assigning, embedding and matrixing a unique patient identifier watermark into each said digital record matrix layer during said digitizing processsaid optical scanner simultaneously assigning, embedding and matrixing a first digital physician signature watermark into each said digital record matrix layers during said digitizing processmemory for storing said digital records, containing said unique patient identifier watermark and said digital physician signature watermark, within said computer systemmeans for storing said digital records, containing said unique patient identifier watermark and said digital physician signature watermark, in said memory and retrieving said digital records, containing said unique patient identifier watermark and said digital physician signature watermark, from said memory,a physician signature database having a plurality of physician names and corresponding second digital physician signature watermarks stored therein,means for authenticating said stored digital records including means for comparing said first digital physician signature watermark retrieved from a selected one of said stored digital records with a corresponding one of said second digital physician signature watermarks retrieved from said physician signature database, andmeans for organizing and ranking said authenticated digital records based on their chronology and clinical utility in treating said patient,means for outputting said authenticated digital records based on said organization and ranking. | 1. A computer system for inputting, storing, organizing, retrieving, and authenticating, medical records, clinical data, and patient data, the system comprising: an optical scanner for converting medical records, clinical data, and patient data to digital records using a digitizing processsaid optical scanner creating a digital data matrix layer of said digital recordssaid optical scanner simultaneously assigning, embedding and matrixing a unique patient identifier watermark into each said digital record matrix layer during said digitizing processsaid optical scanner simultaneously assigning, embedding and matrixing a first digital physician signature watermark into each said digital record matrix layers during said digitizing processmemory for storing said digital records, containing said unique patient identifier watermark and said digital physician signature watermark, within said computer systemmeans for storing said digital records, containing said unique patient identifier watermark and said digital physician signature watermark, in said memory and retrieving said digital records, containing said unique patient identifier watermark and said digital physician signature watermark, from said memory,a physician signature database having a plurality of physician names and corresponding second digital physician signature watermarks stored therein,means for authenticating said stored digital records including means for comparing said first digital physician signature watermark retrieved from a selected one of said stored digital records with a corresponding one of said second digital physician signature watermarks retrieved from said physician signature database, andmeans for organizing and ranking said authenticated digital records based on their chronology and clinical utility in treating said patient,means for outputting said authenticated digital records based on said organization and ranking. 2. The computer system of claim 1 wherein said stored digital records are accessible using at least one of said patient identifier, said first physician signature, a biometric characteristic of a user, and a system password. 3. The computer system of claim 1 wherein said digital records are accessible via a plurality of means including: a computer network, a telephone, a voice recognition system, a data access system. 4. The computer system of claim 1 wherein said digital records are stored in a format including at least one discrete data field, wherein said data records are retrievable based on a plurality of means including: the age and date of said clinical data, the severity of said patient's medical condition, and the medical relevance of said clinical data in treating said patient. 5. The computer system of claim 1 wherein said digital records can be updated on a 24 hour basis via a plurality of means including: a computer network, a telephone system, a data access system. 6. The computer system of claim 1 wherein said stored digital records a... | William Reeves | A61B005/00 G06F019/00
US07123010 | B2 | 2006-10-17 | Magnetic resonance apparatus and operation method for hyperthermic treatment | 2002-06-18 | In a magnetic resonance system and operating method, antenna elements in an antenna array are disposed around an examination volume, and each antenna element has a separate transmission channel and reception channel associated therewith. The magnetic resonance apparatus is operated to obtain magnetic resonance signals, from which amplitude and phase information are derived for the individual antenna elements, and this information is used to subsequently operate the antenna elements in the array to emit RF energy with a predetermined phase and amplitude so as to generate focused RF energy for hyperthermic treatment. The magnetic resonance apparatus can also be used to obtain magnetic resonance signals in intervals which during which the hyperthermic treatment is interrupted, from which the temperature of the region being treated can be ascertained. | 1. A magnetic resonance apparatus comprising: a basic field magnet that generates a basic magnetic field in an examination volume;at least one gradient coil that generates at least one gradient magnetic field in the examination volume;a plurality of radio-frequency antennas disposed in and forming a single array, around the examination volume, configured for radiating RF energy into, and receiving RF energy from, the examination volume;a plurality of separate transmission channels respectively connected to the plurality of RF antennas;a plurality of separate reception channels respectively connected to the plurality of RF antenna; anda control unit connected to said at least one gradient coil and to said single array of RF antennas for operating said at least one gradient coil and said single array of RF antennas in order to excite and acquire magnetic resonance signals having an amplitude and a phase, from the subject, said control unit determining, for each antenna in said plurality of RF antennas forming the single array, the amplitude and the phase of the magnetic resonance signal received by the respective antenna, and said control unit also activating the same plurality of RF antennas of said single array independently of each other in order to emit RF energy of a prescribed phase and amplitude into the examination volume in order to generate an RF field focused in the examination volume configured for hyperthermic treatment of the subject. | 1. A magnetic resonance apparatus comprising: a basic field magnet that generates a basic magnetic field in an examination volume;at least one gradient coil that generates at least one gradient magnetic field in the examination volume;a plurality of radio-frequency antennas disposed in and forming a single array, around the examination volume, configured for radiating RF energy into, and receiving RF energy from, the examination volume;a plurality of separate transmission channels respectively connected to the plurality of RF antennas;a plurality of separate reception channels respectively connected to the plurality of RF antenna; anda control unit connected to said at least one gradient coil and to said single array of RF antennas for operating said at least one gradient coil and said single array of RF antennas in order to excite and acquire magnetic resonance signals having an amplitude and a phase, from the subject, said control unit determining, for each antenna in said plurality of RF antennas forming the single array, the amplitude and the phase of the magnetic resonance signal received by the respective antenna, and said control unit also activating the same plurality of RF antennas of said single array independently of each other in order to emit RF energy of a prescribed phase and amplitude into the examination volume in order to generate an RF field focused in the examination volume configured for hyperthermic treatment of the subject. 2. A magnetic resonance apparatus as claimed in claim 1 wherein each of said separate transmission channels comprises a power amplifier and a modulator. 3. A magnetic resonance apparatus as claimed in claim 1 wherein each antenna in said plurality of RF antennas is a resonator rod. 4. A magnetic resonance apparatus as claimed in claim 1 wherein said control unit interrupts generation of said focused RF field during hyperthermic treatment of the subject at predetermined times of an interval and, in said interval, activates said at least one gradient coil and at least one antenna in said plurality of RF antennas in order to acquire magnetic resonance signals indicating a temperature of a region of the subject in which said focused RF field is present. 5. A magnetic resonance apparatus as claimed in claim 1 wherein said subject in the examination volume has impedance characteristics associated therewith, and wherein said magnetic resonance apparatus comprises a matching circuit connected to said plurality of transmission channels that matches a line impedance of each transmission channel to said impedance characteristics. 6. A magnetic resonance apparatus as claimed in claim 1 wherein said array comprises a cylindrical arrangement of said plurality of RF antennas around said examination volume. 7. A magnetic resonance apparatus as claimed in claim 1 wherein said basic field magnet generates said basic magnetic field with a field strength of at least 2T. 8. A method for operating a magnetic resonance ... | Siemens Aktiengesellschaft | G01V003/00 A61B005/055 A61N005/02
US07201724 | B2 | 2007-04-10 | Stylet unit | 2002-04-09 | A stylet unit has a flexible tubular stylet sleeve, a flexible inner stylet wire inserted into a channel of the stylet sleeve with at a least a portion of the stylet wire and at least a portion of the channel each having a non-circular cross section for preventing rotation of the wire inside the sleeve. The stylet unit has a handle with which the sleeve and the stylet wire are movable relative to each other in a longitudinal direction. The sleeve and the wire are connected at one end thereof to the handle, with the sleeve being rotationally arranged within the handle. | 1. A stylet unit comprising: a flexible, tubular stylet sleeve defining a channel therein;a flexible inner stylet wire disposed in said channel of said stylet sleeve, at least a portion of said stylet wire and at least a portion of said channel each having a non-circular cross section for preventing rotation of said wire inside said sleeve;a handle connected to one end of said sleeve and one end of said wire for moving said sleeve and said wire relative to each other in a longitudinal direction; andsaid sleeve being rotationally arranged within said handle to allow rotation of said handle relative to said sleeve. | 1. A stylet unit comprising: a flexible, tubular stylet sleeve defining a channel therein;a flexible inner stylet wire disposed in said channel of said stylet sleeve, at least a portion of said stylet wire and at least a portion of said channel each having a non-circular cross section for preventing rotation of said wire inside said sleeve;a handle connected to one end of said sleeve and one end of said wire for moving said sleeve and said wire relative to each other in a longitudinal direction; andsaid sleeve being rotationally arranged within said handle to allow rotation of said handle relative to said sleeve. 2. A stylet unit as claimed in claim 1 wherein said stylet wire is non-rotationally fixed to at least a part of said handle. 3. A stylet unit as claimed in claim 1 wherein said handle comprises a handling element engaging said stylet sleeve and movable in a longitudinal direction relative to said stylet wire, said stylet sleeve being rotatable relative to said handling element. 4. A stylet unit as claimed in claim 3 comprising a rigid tube fastened to said stylet sleeve, said rigid tube engaging said handling element and being rotatable relative to said handling element. 5. A stylet unit as claimed in claim 4 wherein said tube has a recess and wherein said handling element has a projection engaging said recess. 6. A stylet unit as claimed in claim 4 wherein said handling element has a recess and wherein said tube has a projection engaging said recess. 7. A stylet unit as claimed in claim 4 comprising an intermediate element disposed between said handling element and said tube and engaging said tube and said handling element, said intermediate element being rotatable relative to at least one of said handling element and said tube. 8. A stylet unit as claimed in claim 7 wherein said intermediate element has a recess and wherein said handling element has a projection engaging said recess. 9. A stylet unit as claimed in claim 7 wherein said handling element has a recess and wherein said intermediate element has a projection engaging said recess. 10. A stylet unit as claimed in claim 7 wherein said intermediate element is a bracket that is rotationally arranged around said tube and non-rotationally arranged relative to said handling element. 11. A stylet unit as claimed in claim 3 wherein said stylet wire is non-rotationally fixed to at least a part of said handle, and comprising a guide element along which said handling element is movable toward and from said part of said handle. 12. A stylet unit as claimed in claim 11 wherein said handling element is non-rotationally arranged relative to said guide element. 13. A stylet unit as claimed in claim 11 wherein, at least temporarily during use of said stylet unit, said part of said handle is non-rotationally arranged relative to said guide element and relative to said handling element. 14. A stylet unit as claimed in claim 1 wherein each of the non-circular portion of the channel of... | St. Jude Medical AB | A61B005/00
US07136695 | B2 | 2006-11-14 | Patient-specific template development for neurological event detection | 2001-10-12 | An epileptiform activity patient-specific template creation system permits a user to efficiently develop an optimized set of patient-specific parameters for epileptiform activity detection algorithms. The epileptiform activity patient template creation system is primarily directed for use with an implantable neurostimulator system having EEG storage capability, in conjunction with a computer software program operating within a computer workstation having a processor, disk storage and input/output facilities for storing, processing and displaying patient EEG signals. The implantable neurostimulator is operative to store records of EEG data when neurological events are detected, when it receives external commands to record, or at preset or random times. The computer workstation operates on stored and uploaded records of EEG data to derive the patient-specific templates. | 1. A method for developing a patient-specific detection template from at least one record of electrographic signal data stored by an implantable device, the method comprising the steps of: uploading the at least one record of electrographic signal data from the implantable device to an external apparatus; processing the at least one record and an event annotation to generate at least one parameter; and forming the patient-specific detection template from the at least one parameter; and downloading the patient-specific detection template from the external apparatus to the implantable device, wherein the patient-specific detection template comprises a prediction template. | 1. A method for developing a patient-specific detection template from at least one record of electrographic signal data stored by an implantable device, the method comprising the steps of: uploading the at least one record of electrographic signal data from the implantable device to an external apparatus; processing the at least one record and an event annotation to generate at least one parameter; and forming the patient-specific detection template from the at least one parameter; and downloading the patient-specific detection template from the external apparatus to the implantable device, wherein the patient-specific detection template comprises a prediction template. 2. A method for developing a patient-specific detection template from at least one record of electrographic signal data stored by an implantable device, the method comprising the steps of: uploading the at least one record of electrographic signal data from the implantable device to an external apparatus; processing the at least one record and an event annotation to generate at least one parameter; and forming the patient-specific detection template from the at least one parameter; and downloading the patient-specific detection template from the external apparatus to the implantable device. 3. The method for developing a patient-specific detection template of claim 2, wherein forming the patient-specific detection template comprises creating a data structure including at least one data item representative of the at least one parameter. 4. The method for developing a patient-specific detection template of claim 2, further comprising the step of creating the event annotation by annotating at least one event in the at least one record. 5. The method for developing a patient-specific detection template of claim 4, wherein the step of creating the event annotation comprises the steps of: viewing the at least one record of electrographic signal data; and indicating a start time of the event with the event annotation. 6. The method for developing a patient-specific detection template of claim 5, wherein the step of creating the event annotation further comprises the step of indicating an end time of the event with the event annotation. 7. The method for developing a patient-specific detection template of claim 4, wherein the step of creating the event annotation comprises the steps of: obtaining data representative of an event log containing an event entry; and indicating a start time of the event with the event annotation by processing the event entry. 8. The method for developing a patient-specific detection template of claim 7, wherein the step of creating the event annotation further comprises the steps of: viewing the at least one record of electrographic signal data; and modifying the event annotation to refine the start time of the event. 9. The method for developing a patient-specific detection template of claim 2, further comprising the step of augmenting the at lea... | - | A61B005/0476 A61N001/18
US07192400 | B2 | 2007-03-20 | Device and method for vascular monitoring | 2002-10-24 | A vascular coupler adapted to be fixed into position upon a blood vessel and there retain and position a matable sensing device. In a preferred embodiment, the vascular coupler is an anastomotic coupler and the sensing device comprises an ultrasonic Doppler probe. | 1. A combination comprising a vascular coupler adapted to be fixed into position upon a blood vessel and there retain and position one or more sensing devices in a predetermined position and orientation, and one or more sensing devices, wherein the coupler is an anastomotic coupler formed by the connection of a pair of matching fastener portions, either or both of which is provided with a receptacle for receiving one of the one or more sensing devices and wherein at least one sensing device is removably retained within the receptacle. | 1. A combination comprising a vascular coupler adapted to be fixed into position upon a blood vessel and there retain and position one or more sensing devices in a predetermined position and orientation, and one or more sensing devices, wherein the coupler is an anastomotic coupler formed by the connection of a pair of matching fastener portions, either or both of which is provided with a receptacle for receiving one of the one or more sensing devices and wherein at least one sensing device is removably retained within the receptacle. 2. A combination according to claim 1 wherein the combination of the coupler and the one or more sensing devices is adapted to be made once the coupler itself has been positioned within the body. 3. A combination according to claim 1 wherein the combination of the coupler and the one or more sensing devices is adapted to be made outside the body. 4. A combination according to claim 1 wherein at least one of the fastener portions is mated, or adapted to be mated, with at least one of the one or more sensing devices. 5. A combination according to claim 1 wherein at least one of the one or more sensing devices include a device, or component thereof adapted to indicate or measure one or more blood flow parameters, oxygen, temperature, pH, or the presence or concentration of one or more biochemical compounds. 6. A combination according to claim 1 wherein the fastener portions are identical portions, adapted to be connected by the use of axially directed pins and corresponding intermediate holes. 7. A combination according to claim 6 wherein at least one of the one or more sensing devices comprises an ultrasonic Doppler probe adapted to be releasably positioned within and retained by the receptacle of either fastener portion under microscopic visualization. 8. A combination according to claim 7 wherein the Doppler probe comprises a circular piezoelectric crystal. 9. A method of monitoring or evaluating a tissue site following vascular surgery, comprising the steps of providing and using a combination according to claim 1. 10. A method according to claim 9 wherein the combination of the coupler and the one or more sensing devices is made once the coupler itself has been positioned within the body. 11. A method according to claim 9 wherein the combination of the coupler and the one or more sensing devices is made outside the body. 12. A method according to claim 9 wherein at least one of the fastener portions is mated, or adapted to be mated, with at least one of the one or more sensing devices. 13. A method according to claim 9 wherein at least one of the one or more sensing devices include a device, or component thereof, adapted to indicate or measure one or more blood flow parameters, oxygen, temperature, pH, or the presence or concentration of one or more biochemical compounds. 14. A method according to claim 9 wherein the fastener portions are identical portions, adapted to be connected by the use of axial... | Synovis Life Technologies, Inc. | A61B005/00 A61B005/02
US07236817 | B2 | 2007-06-26 | Animation technology | 2001-03-05 | A method for creating an animated image of the bones of a body part is described. The steps include converting Computed Axial Tomography (CAT or CT) scan and Magnetic Resonance Imaging (MRI) 2-dimensional cross-sectional images (slices) into 3-dimensional images of individual bones (FIG. 1). A first ordered series of slices of a body part in a first position is converted into a 3-dimensional representation of the skeleton of the body part and then a second ordered series of slices of the same body part in a different position is converted. The converted images are then used to create a step frame animation of the movement of the skeleton of the body part. Additional ordered series of slices of the body part in other intermediate positions could be used in the step frame animation. The steps of the method may also include the process of separating the bones if they are depicted as co-joined in a slice (FIGS. 6 and 7). The method can also include the identification of isometric points (FIGS. 8 to 11) on the separated bones to aid a clinician in the diagnosis of a problem or abnormality as well as assist in the application of “what if” surgical procedures on the digital representation of the skeleton of the scanned body part. The method can also include steps that create movement of the bones of the body part by applying rules of motion for one or more points on the separated bones of the scanned body part. | 1. A method of producing 3-dimensional visualisations of digital representations of cross-sectional slices of a vertebrate animal or human body part includes the steps of: a) obtaining a first ordered series of slices of a portion of said body part in a first position;b) applying one or more filters to each of said digital representations of said first ordered series of slices to identify the skeletal portions of said body part;c) converting said first filtered series into a 3-dimensional representation of the skeleton of said body part in said first position;d) obtaining a second ordered series of slices of said portion of said body part in a second position different to said first position wherein said first and said second position of said portion of said body part is representative of extremes of the achievable movement of said portion of said body part;e) applying one or more filters to each of said digital representations of said second ordered series of slices to identify the skeletal portions of said body part;f) converting said second filtered series into a 3-dimensional representation of the skeleton of said body part in said second position;g) combining said 3-dimensional representations to form a step frame animation having as many steps as there are ordered series of slices. | 1. A method of producing 3-dimensional visualisations of digital representations of cross-sectional slices of a vertebrate animal or human body part includes the steps of: a) obtaining a first ordered series of slices of a portion of said body part in a first position;b) applying one or more filters to each of said digital representations of said first ordered series of slices to identify the skeletal portions of said body part;c) converting said first filtered series into a 3-dimensional representation of the skeleton of said body part in said first position;d) obtaining a second ordered series of slices of said portion of said body part in a second position different to said first position wherein said first and said second position of said portion of said body part is representative of extremes of the achievable movement of said portion of said body part;e) applying one or more filters to each of said digital representations of said second ordered series of slices to identify the skeletal portions of said body part;f) converting said second filtered series into a 3-dimensional representation of the skeleton of said body part in said second position;g) combining said 3-dimensional representations to form a step frame animation having as many steps as there are ordered series of slices. 2. A method according to claim 1 further includes the further steps of: h) obtaining a one or more further ordered series of slices of said portion of said body part in a one or more further positions between the extremes of the achievable movement;i) converting said further filtered series into a 3-dimensional representation of the skeleton of said body part in said further positions;j) combining said 3-dimensional representations to form a step frame animation having as many steps as there are ordered series of slices. 3. A method according to claim 1 further includes the steps of: k) selecting a plurality of points on two or more skeletal object s in each said 3-dimensional representation; andl) analysing whether the distance between pairs of said points on different skeletal objects are isometric, if the distance between pairs of points remain the same or within a predetermined variance of distance during said step frame animation. 4. A method according to claim 3 further includes the step of: m) comparing the changes, if any, of the distances between predetermined pairs of points with expected changes, wherein the result of said comparison provides assistance to a cliniclan to form a diagnosis regarding the body part. 5. A method according to claim 3 wherein isometric points determined equate substantially to the fixation location of the ends of ligaments associated with said portion of said body part and the further step of: n) comparing the isometric points determined in step 3 to predetermined isometric points of a typical or contra-lateral body part and if said isometric points vary a greater than a predetermined amount a problem or abnormality or a ... | True Life Creations (SA) Pty Ltd. Macropace Products Pty Ltd. | A61B005/05
US06866637 | B2 | 2005-03-15 | Apparatus and method for the estimation of flow -independent parameters which characterize the relevant features of nitric oxide production and exchange in the human lungs | 2001-04-26 | The invention provides an estimation of key flow-independent parameters characteristic of NO exchange in the lungs, namely: 1) the steady state alveolar concentration, Calv,ss; 2) the maximum flux of NO from the airways, JNO,max; and 3) the diffusing capacity of NO in the airways, DNO,air. The parameters were estimated from a single exhalation maneuver comprised of a pre-expiratory breathhold, followed by an exhalation in which the flow rate progressively decreased. The mean values for JNO,max, DNO,air, and Calv,ss do not depend on breathhold time for breathhold times greater than approximately 10 seconds. A priori estimates of the parameter confidence intervals demonstrates that a breathhold no longer than 20 seconds may be adequate, and that JNO,max be can estimated with the smallest uncertainty, and DNO,air the largest. | 1. A method for determining flow independent parameters characteristic of nitrogen monoxide exchange and of lung function in a subject during one or more exhalations during which flow rate of exhalation varies in time comprising: exhaling into a mouthpiece in which the resistance to airflow can vary in time and; measuring exhaled nitrogen monoxide concentration, Cexh, from said subject as a function of time; and simultaneously measuring volumetric exhalation flow rate, , as a function of time with measurement of exhaled nitrogen monoxide concentration; backwards integrating said volumetric exhalation flow rate, , over time to convert exhalation time to residence time of each exhaled bolus of gas in said airway compartment of said subject; and selecting said flow independent parameters to fit said measured nitrogen monoxide concentration as a function of time given residence time obtained by backwards integration. | 1. A method for determining flow independent parameters characteristic of nitrogen monoxide exchange and of lung function in a subject during one or more exhalations during which flow rate of exhalation varies in time comprising: exhaling into a mouthpiece in which the resistance to airflow can vary in time and; measuring exhaled nitrogen monoxide concentration, Cexh, from said subject as a function of time; and simultaneously measuring volumetric exhalation flow rate, , as a function of time with measurement of exhaled nitrogen monoxide concentration; backwards integrating said volumetric exhalation flow rate, , over time to convert exhalation time to residence time of each exhaled bolus of gas in said airway compartment of said subject; and selecting said flow independent parameters to fit said measured nitrogen monoxide concentration as a function of time given residence time obtained by backwards integration. 2. The method of claim 1 further comprising: estimating Vair, which is the volume of said airway compartment of said subject, and; measuring Vds which is the dead space volume or volume expired prior to observing an NO signal, and where backwards integrating said exhalation flow rate, , over time to convert exhalation time to residence time of each exhaled bolus of gas in an airway compartment in said subject is performed in a computer according to a Case I or Case II integration where: Case⁢ ⁢I⁡(Phase⁢ ⁢I⁢ ⁢and⁢ ⁢II⁢ ⁢of⁢ ⁢exhalation):∫t-τres⁡(t)t+tds⁢VE&amp;⁡(t′)⁢ ⁢ⅆt′=0Case⁢ ⁢II⁡(Phase⁢ ⁢III⁢ ⁢of⁢ ⁢exhalation):∫t-τres⁡(t)t+tds⁢VE&amp;⁡(t′)⁢ ⁢ⅆt′=Vair+Vdswhere t is time, tds is convective transport delay time in the dead space volume, Tres: is residence time of each differential gas bolus in said airway compartment. 3. The method of claim 2 where selecting said flow independent parameters to fit said measured nitrogen monoxide concentration as a function of time given residence time obtained by backwards integration comprises estimating three flow independent parameters Calv,ss, DNO,air, JNO,max from Cexh as a function of Tres(t), where Calv,ss is the steady state alveolar concentration of nitrogen monoxide, DNO,air is the diffusing capacity of nitrogen monoxide in said airway compartment, JNO,max is the maximum total molar flux of nitrogen monoxide from the airway wall, Cexh is the exhaled concentration of nitrogen monoxide, and Tres(t) is the residence time of each differential gas bolus in said airway compartment as a function of time, t. 4. The method of claim 3 where estimating said three flow-independent parameters Calv,ss, DNO,air, JNO,max from Cexh as a function of Tres(t), comprising performing said estimation by relating said three flow-independent parameters Calv,ss, DNO,air, JNO,max and Cexh to measured exhaled concentration in a computer in the relationship given by: Cexh⁡(t+tds)=(Cin⁡(t-τres)-JNO, ⁢maxDNO, ⁢air)⁢ⅇ-DNO, ⁢airVair⁢τres⁡(t)+JNO, ⁢maxDNO, ⁢airwhere Cin is the inlet concentration of nitrogen monoxide to s... | The Regents of the University of California | A61B005/00
US06915149 | B2 | 2005-07-05 | Method of pacing a heart using implantable device | 2002-01-11 | A method of determining an optimal location for implanting a pacemaker electrode includes the steps of pacing a heart from a first location and generating a first map of the heart associated with pacing at the first location. The heart is paced from a second location and a second map is generated of the heart associated with pacing at the second location. The first and second maps are compared in order to diagnose the effect of the pacing and an optimal location for implanting the pacemaker electrode based on the comparison of the maps is selected. | 1. A method of determining an optimal location for implanting a pacemaker electrode comprising: pacing a heart from a first location; generating a first map of the heart associated with pacing at the first location; pacing the heart from a second location; generating a second map of the heart associated with pacing at the second location; comparing the first and second maps to diagnose the effect of the pacing; and selecting an optimal location for implanting the pacemaker electrode based on comparing the maps. | 1. A method of determining an optimal location for implanting a pacemaker electrode comprising: pacing a heart from a first location; generating a first map of the heart associated with pacing at the first location; pacing the heart from a second location; generating a second map of the heart associated with pacing at the second location; comparing the first and second maps to diagnose the effect of the pacing; and selecting an optimal location for implanting the pacemaker electrode based on comparing the maps. 2. A method according to claim 1, wherein generating the first map comprises: (a) bringing an invasive probe into contact with a site on the wall of the heart; (b) determining a position of the invasive probe; (c) determining a cardiac characteristic at the site; (d) repeating (a)-(c) for a plurality of sites of the heart; and (e) combining the positions to form a time-dependent map of at least a portion of the heart. 3. A method according to claim 2, wherein determining the position comprises determining the position at least two different phases of a heart cycle of the heart. 4. A method according to claim 2, wherein determining the cardiac characteristic at the site comprises determining an electrical value at the site. 5. A method according to claim 2, wherein determining the cardiac characteristic at the site comprises determining a non-electrical physiological value at the site. 6. A method according to claim 1, wherein generating the first map comprises: (a) bringing an invasive probe into contact with a site on the wall of the heart; (b) determining a position of the invasive probe; (c) determining a cardiac characteristic at the site; (d) repeating (a)-(c) for a plurality of sites of the heart; and (e) combining the positions to form a map of at least a portion of the heart. 7. A method according to claim 6, wherein determining the cardiac characteristic comprises determining the cardiac characteristic at the site at a plurality of different phases of the heart cycle. 8. A method according to claim 6, wherein determining the cardiac characteristic at the site comprises determining a non-electrical physiological value at the site. 9. A method according to claim 6, wherein determining the cardiac characteristic at the site comprises determining an electrical value at the site. 10. A method according to claim 6, comprising determining at least a second position of the invasive probe, which position is different from the position determined in (b). 11. A method according to claim 10, comprising determining at least one local relationship between changes in positions of the invasive probe and a determined local cardiac characteristic. 12. A method according to claim 6, wherein the cardiac characteristic is determined using a sensor external to the probe. 13. A method according to claim 12, wherein the sensor is external to a body which comprises the heart. 14. A method according to claim 6, wherein the cardiac character... | Biosense, Inc. | A61B005/05
US06921379 | B2 | 2005-07-26 | Breast pump | 2001-12-20 | The invention relates to a breast pump, which in addition to a drive motor (7), vacuum diaphragm pump (1) and programme-controlled electronics (4), comprises a vacuum chamber (12) with a separation diaphragm (13), which is connected between a pump intake conduit (9) and an extraction hood (11). The first chamber (14) is equipped with a device (16) for the negative or positive aeration of said chamber. Said device (16) is also programme-controlled by the electronics (4) during the normal continuous operation of the pump (1). | 1. A breast pump, comprising: a diaphragm vacuum pump, drive motor control electronics for setting a motor speed of the vacuum pump as a function of a preprogrammed vacuum profile, a vacuum chamber coupled between the vacuum pump and a suction cap, the vacuum chamber including a separating diaphragm, a first chamber of the vacuum chamber communicating with a suction line of the vacuum pump, and a second chamber which can be connected to the suction cap, the first chamber being equipped with a controlled device for canceling a vacuum which has been built up. | 1. A breast pump, comprising: a diaphragm vacuum pump, drive motor control electronics for setting a motor speed of the vacuum pump as a function of a preprogrammed vacuum profile, a vacuum chamber coupled between the vacuum pump and a suction cap, the vacuum chamber including a separating diaphragm, a first chamber of the vacuum chamber communicating with a suction line of the vacuum pump, and a second chamber which can be connected to the suction cap, the first chamber being equipped with a controlled device for canceling a vacuum which has been built up. 2. The breast pump as claimed in claim 1, wherein the controlled device for canceling the vacuum in the vacuum chamber comprises a valve which is controlled by control electronics and which, when opened, connects the first chamber to atmosphere and passively removes air from the vacuum chamber. 3. The breast pump as claimed in claim 1, characterized in that the device for canceling the vacuum in the vacuum chamber consists of a valve switch which, when actuated, interrupts the suction line between the vacuum pump and the first chamber and connects said first chamber of the vacuum chamber to the ejection line of the vacuum pump building up a positive pressure, i.e. actively cancels the vacuum in the vacuum chamber by positive pressure buildup. | Medela AG | A61M001/06 A01J005/04 A61B005/00 F04B017/03
US06971993 | B2 | 2005-12-06 | Method for utilizing oral movement and related events | 2001-11-14 | A method for utilizing oral movements is used in speech assessment, speech therapy, language development, and controlling external devices. A device is used which includes a sensor plate having sensors to detect contact of the tongue with the sensor plate. One aspect of the invention allows viewing representations of contact of the tongue and palate during speech and comparing the representations with model representations displayed in a split screen fashion. The model representations may be generated by another speaker utilizing a sensor plate or by computer generated representations which have been electronically stored. The representations may be analyzed to assess speech proficiency and the model may be mimicked for speech enhancement. | 1. A method for providing speech therapy comprising: (A) displaying a model representation of a position of contact between a model tongue and mouth during speech;(B) displaying a representation of a position of contact between a learner's tongue and mouth during speech;(C) displaying a model representation of a contact between a model lip and teeth during speech; and(D) representing the position of contact between the learner's tongue and mouth during speech by a grid of dots on a computer screen, wherein said dots expand and change color responsive to contact between the learner's tongue and mouth. | 1. A method for providing speech therapy comprising: (A) displaying a model representation of a position of contact between a model tongue and mouth during speech;(B) displaying a representation of a position of contact between a learner's tongue and mouth during speech;(C) displaying a model representation of a contact between a model lip and teeth during speech; and(D) representing the position of contact between the learner's tongue and mouth during speech by a grid of dots on a computer screen, wherein said dots expand and change color responsive to contact between the learner's tongue and mouth. 2. The method of claim 1 further comprising instructing the learner to mimic the model representation of the position of contact between the model tongue and mouth during speech. 3. The method of claim 2 further comprising the step of comparing the representation of the position of contact between the learner's tongue and mouth during speech with the model representation of position of contact between the model tongue and mouth during speech. 4. The method of claim 2 further comprising the step of generating a numerical score representing the closeness of fit between the representation of position of contact between the learner's tongue and mouth during speech and the model representation of position of contact between the model tongue and mouth during speech. 5. The method of claim 2 further comprising the step of providing positive reinforcement when the learner mimics the model representation of position of contact between the model tongue and mouth during speech. 6. The method of claim 5 wherein the positive reinforcement comprises computer generated congratulations. 7. The method of claim 5 wherein the positive reinforcement comprises dynamic displays on a computer monitor. 8. The method of claim 2 further comprising the step of providing model acoustic representations of the speech. 9. The method of claim 8 wherein auditory mimicry accuracy between the learner and the model acoustic representation are analyzed acoustically. 10. The method of claim 9 wherein a numerical score representing the closeness of acoustic fit is generated. 11. The method of claim 8 wherein sound spectrographic displays are generated from the model acoustic representation and the learner's speech. 12. The method of claim 2 wherein the model representation is designed to exercise the learner's tongue when the learner mimics the model representation. 13. The method of claim 2 wherein the model representation is designed to instruct the learner to compensate for physical deficiencies. 14. The method of claim 1 wherein the representations of position of contact between the learner's tongue and mouth during speech and the model representations of position of contact between the model tongue and mouth during speech are displayed on a split-screen. 15. The method of claim 1 wherein the learner is unable to overcome speech disorders through traditional auditory bas... | Logometrix Corporation | A61B005/103 A61F005/58 G09B019/04
US06960167 | B2 | 2005-11-01 | System and method for patient monitoring using a reference baseline for use in automated patient care | 2001-08-13 | A system and method for patient monitoring using a reference baseline for use in automated patient. One or more reference physiological measures relating to individual patient information recorded during an initial observation period are processed into a reference baseline. One or more updated physiological measures recorded subsequent to the initial observation period are periodically received. The updated physiological measures are compared to the reference physiological measures in the reference baseline to generate a patient status indicator identifying any such substantially non-conforming updated physiological measure. | 1. A system for patient monitoring using a reference baseline for use in automated patient care, comprising: a processing module processing one or more reference physiological measures relating to individual patient information recorded during an initial observation period by an implantable medical device into a reference baseline; an analysis module periodically receiving one or more updated physiological measures regularly recorded by the implantable medical device subsequent to the initial observation period; a database module storing the reference physiological measures and one or more reference quality of life measures recorded by the individual patient during the initial observation period into a patient care record for the individual patient, and storing the updated physiological measures and one or more updated quality of life measures recorded by the individual patient subsequent to the initial observation period into the patient care record; a comparison module comparing the updated physiological measures to the reference physiological measures in the reference baseline to generate a patient status indicator identifying any such substantially non-conforming updated physiological measure; and a quality of life comparison submodule comparing the updated quality of life measures to the reference quality of life measures in the patient care record to generate the patient status indicator identifying any such substantially non-conforming updated quality of life measure. | 1. A system for patient monitoring using a reference baseline for use in automated patient care, comprising: a processing module processing one or more reference physiological measures relating to individual patient information recorded during an initial observation period by an implantable medical device into a reference baseline; an analysis module periodically receiving one or more updated physiological measures regularly recorded by the implantable medical device subsequent to the initial observation period; a database module storing the reference physiological measures and one or more reference quality of life measures recorded by the individual patient during the initial observation period into a patient care record for the individual patient, and storing the updated physiological measures and one or more updated quality of life measures recorded by the individual patient subsequent to the initial observation period into the patient care record; a comparison module comparing the updated physiological measures to the reference physiological measures in the reference baseline to generate a patient status indicator identifying any such substantially non-conforming updated physiological measure; and a quality of life comparison submodule comparing the updated quality of life measures to the reference quality of life measures in the patient care record to generate the patient status indicator identifying any such substantially non-conforming updated quality of life measure. 2. A system according to claim 1, further comprising: a revision submodule analyzing the updated physiological measures relative to one or more further updated physiological measures to determine a revised patient status indicator. 3. A system according to claim 1, further comprising: a feedback module providing automated feedback based on the patient status indicator to the individual patient. 4. A system according to claim 1, further comprising: a feedback module comparing the updated measures to the reference measures in the patient care record to generate the patient status indicator identifying any such substantially non-conforming updated measures. 5. A system according to claim 1, further comprising: a reevaluation submodule processing one or more reference physiological measures relating to individual patient information recorded subsequent to the initial observation period into a revised reference baseline. 6. A system according to claim 1, further comprising: a monitoring submodule monitoring the individual patient while the individual patient is performing a prescribed set of timed physical stressors during the initial observation period. 7. A system according to claim 6, wherein the prescribed set of activities are representative of substantially normal activity, further comprising: an activity submodule determining relative abnormal activity response based on any updated physiological measure identified as being substantially non-conforming to the cor... | Cardiac Intelligence Corporation | A61B005/04
US06939532 | B2 | 2005-09-06 | Versatile hydrophilic dyes | 2001-01-09 | Dye-peptide conjugates useful for diagnostic imaging and therapy are disclosed. The dye-peptide conjugates include several cyanine dyes with a variety of bis- and tetrakis(carboxylic acid) homologues. The small size of the compounds allows more favorable delivery to tumor cells as compared to larger molecular weight imaging agents. The various dyes are useful over the range of 350-1300 nm, the exact range being dependent upon the particular dye. Use of dimethylsulfoxide helps to maintain the fluorescence of the compounds. The molecules of the invention are useful for diagnostic imaging and therapy, in endoscopic applications for the detection of tumors and other abnormalities and for localized therapy, for photoacoustic tumor imaging, detection and therapy, and for sonofluorescence tumor imaging, detection and therapy. | 1. A compound of formula wherein W5 and X5 are —CR1R2; Y5 is selected from the group consisting of —(CH2)a—CONH-Bm, —CH2—(CH2OCH2)b—CH2—CONH-Bm, —(CH2)a—NHCO-Bm, —CH2—(CH2OCH2)b—CH2—NHCO-Bm, —(CH2)a—N(R3)—(CH2)b—CONH-Bm, (CH2)a—N(R3)—(CH2)c—NHCO-Bm, —(CH2)a—N(R3)—CH2—(CH2OCH2)b—CH2—CONH-Bm, —(CH2)a—N(R3)—CH2—(CH2OCH2)b—CH2—NHCO-Bm, —CH2—(CH2OCH2)b—CH2—N(R3)—(CH2)a—CONH-Bm, —CH2—(CH2OCH2)b—CH2—N(R3)—(CH2)a—NHCO-Bm, —CH2—(CH2OCH2)b—CH2—N(R3)—CH2—(CH2OCH2)d—CONH-Bm, —CH2—(CH2OCH2)b—CH2—N(R3)—CH2—(CH2OCH2)d—NHCO-Bm, —(CH2)a—NR3R4, and —CH2(CH2OCH2)b—CH2NR3R4; Z5 is selected from the group consisting of —(CH2)a—CONH-Dm, —CH2—(CH2OCH2)b—CH2—CONH-Dm, —(CH2)a—NHCO-Dm, —CH2—(CH2OCH2)b—CH2—NHCO-Dm, —(CH2)a—N(R3)—(CH2)b—CONH-Dm, (CH2)a—N(R3)—(CH2)c—NHCO-Dm, —(CH2)a—N(R3)—CH2—(CH2OCH2)b—CH2—CONH-Dm, —(CH2)a—N(R3)—CH2—(CH2OCH2)b—CH2—NHCO-Dm, —CH2—(CH2OCH2)b—CH2—N(R3)—(CH2)a—CONH-Dm, —CH2—(CH2OCH2)b—CH2—N(R3)—(CH2)a—NHCO-Dm, —CH2—(CH2OCH2)b—CH2—N(R3)—CH2—(CH2OCH2)d—CONH-Dm, —CH2—(CH2OCH2)b—CH2—N(R3)—CH2—(CH2OCH2)d—NHCO-Dm, —(CH2)a—NR3R4, and —CH2(CH2OCH2)b—CH2NR3R4; A3 is a single or a double bond; B3, C3, and D3 are independently selected from the group consisting of —O—, —S—, —Se—, —P—, —CR1R2, —CR1, alkyl, NR3, and —C═O; A3, B3, C3, and D3 may together form a 6- to 12-membered carbocyclic ring or a 6- to 12-membered heterocyclic ring optionally containing one or more oxygen, nitrogen, or sulfur atom; a5 vary from 0 to 5; R1 to R4, and R58 to R66 are independently selected from the group consisting of hydrogen, C1-C10 alkyl, C5-C20 aryl, C1-C10 alkoxyl, C1-C10 polyalkoxyalkyl, C1-C20 polyhydroxyalkyl, C5-C20 polyhydroxyaryl, C1-C10 aminoalkyl, cyano, nitro, halogen, saccharide, peptide, —CH2(CH2OCH2)b—CH2—OH, —(CH2)a—CONH-Bm, —CH2—(CH2OCH2)b—CH2—CONH-Bm, —(CH2)a—NHCO-Bm, —CH2—(CH2OCH2)b—CH2—NHCO-Bm, and —CH2—(CH2OCH2)b—CO2H; Bm and Dm are independently selected from the group consisting of bioactive peptide, protein, cell, antibody, antibody fragment, saccharide, glycopeptide, peptidomimetic, drug, drug mimic, hormone, metal chelating agent, radioactive or nonradioactive metal complex, and echogenic agent; a and c independently vary from 1 to 20; b and d independently vary from 1 to 100. | 1. A compound of formula wherein W5 and X5 are —CR1R2; Y5 is selected from the group consisting of —(CH2)a—CONH-Bm, —CH2—(CH2OCH2)b—CH2—CONH-Bm, —(CH2)a—NHCO-Bm, —CH2—(CH2OCH2)b—CH2—NHCO-Bm, —(CH2)a—N(R3)—(CH2)b—CONH-Bm, (CH2)a—N(R3)—(CH2)c—NHCO-Bm, —(CH2)a—N(R3)—CH2—(CH2OCH2)b—CH2—CONH-Bm, —(CH2)a—N(R3)—CH2—(CH2OCH2)b—CH2—NHCO-Bm, —CH2—(CH2OCH2)b—CH2—N(R3)—(CH2)a—CONH-Bm, —CH2—(CH2OCH2)b—CH2—N(R3)—(CH2)a—NHCO-Bm, —CH2—(CH2OCH2)b—CH2—N(R3)—CH2—(CH2OCH2)d—CONH-Bm, —CH2—(CH2OCH2)b—CH2—N(R3)—CH2—(CH2OCH2)d—NHCO-Bm, —(CH2)a—NR3R4, and —CH2(CH2OCH2)b—CH2NR3R4; Z5 is selected from the group consisting of —(CH2)a—CONH-Dm, —CH2—(CH2OCH2)b—CH2—CONH-Dm, —(CH2)a—NHCO-Dm, —CH2—(CH2OCH2)b—CH2—NHCO-Dm, —(CH2)a—N(R3)—(CH2)b—CONH-Dm, (CH2)a—N(R3)—(CH2)c—NHCO-Dm, —(CH2)a—N(R3)—CH2—(CH2OCH2)b—CH2—CONH-Dm, —(CH2)a—N(R3)—CH2—(CH2OCH2)b—CH2—NHCO-Dm, —CH2—(CH2OCH2)b—CH2—N(R3)—(CH2)a—CONH-Dm, —CH2—(CH2OCH2)b—CH2—N(R3)—(CH2)a—NHCO-Dm, —CH2—(CH2OCH2)b—CH2—N(R3)—CH2—(CH2OCH2)d—CONH-Dm, —CH2—(CH2OCH2)b—CH2—N(R3)—CH2—(CH2OCH2)d—NHCO-Dm, —(CH2)a—NR3R4, and —CH2(CH2OCH2)b—CH2NR3R4; A3 is a single or a double bond; B3, C3, and D3 are independently selected from the group consisting of —O—, —S—, —Se—, —P—, —CR1R2, —CR1, alkyl, NR3, and —C═O; A3, B3, C3, and D3 may together form a 6- to 12-membered carbocyclic ring or a 6- to 12-membered heterocyclic ring optionally containing one or more oxygen, nitrogen, or sulfur atom; a5 vary from 0 to 5; R1 to R4, and R58 to R66 are independently selected from the group consisting of hydrogen, C1-C10 alkyl, C5-C20 aryl, C1-C10 alkoxyl, C1-C10 polyalkoxyalkyl, C1-C20 polyhydroxyalkyl, C5-C20 polyhydroxyaryl, C1-C10 aminoalkyl, cyano, nitro, halogen, saccharide, peptide, —CH2(CH2OCH2)b—CH2—OH, —(CH2)a—CONH-Bm, —CH2—(CH2OCH2)b—CH2—CONH-Bm, —(CH2)a—NHCO-Bm, —CH2—(CH2OCH2)b—CH2—NHCO-Bm, and —CH2—(CH2OCH2)b—CO2H; Bm and Dm are independently selected from the group consisting of bioactive peptide, protein, cell, antibody, antibody fragment, saccharide, glycopeptide, peptidomimetic, drug, drug mimic, hormone, metal chelating agent, radioactive or nonradioactive metal complex, and echogenic agent; a and c independently vary from 1 to 20; b and d independently vary from 1 to 100. 2. The compound of claim 1 wherein W5 and X5 are independently selected from the group consisting of —C(CH3)2, —C((CH2)aOH)CH3, —C((CH2)aCO2H)CH3, —C((CH2)aNH2)CH3, C((CH2)aNH2)2, C((CH2)aNR3R4)2,; Y5 is selected from the group consisting of —(CH2)a—CONH-Bm, —CH2—(CH2OCH2)b—CH2—CONH-Bm, —(CH2)a—NHCO-Bm, —CH2—(CH2OCH2)b—CH2—NHCO-Bm, —(CH2)a—NR3R4, and —CH2(CH2OCH2)b—CH2NR3R4; Z5 is selected from the group consisting of —(CH2)a—CONH-Dm, —CH2—(CH2OCH2)b—CH2—CONH-Dm, —(CH2)a—NHCO-Dm, —CH2—(CH2OCH2)b—CH2—NHCO-Dm, —(CH2)a—NR3R4, and —CH2(CH2OCH2)b—CH2NR3R4; A3 is a single or a double bond; B3, C3, and D3 are independently selected from the group consisting of —O—, —S—, NR3, (CH2)a-CR1R2, and —CR1; A3, B3, C3, and D3 may together form a 6- to 10-membered carbocyclic ring or a 6-... | Mallinckrodt, Inc. | A61B010/00 A61B005/00 A61B008/00
US06855309 | B2 | 2005-02-15 | Cascade polymer bound complexing compounds, their complexes and conjugates, processes for their production, and pharmaceutical agents containing them | 2002-12-11 | Cascade polymers, containing complex-forming ligands, optionally at least five ions of an element of atomic numbers 21-29, 39, 42, 44 or 57-83, as well as, if desired, cations of inorganic and/or organic bases, amino acids or amino acid amides, are valuable complexing compounds and complexes for diagnostics and therapy. | 1. A method for radiodiagnosis and/or radiotherapy of a patient, comprising conducting radiodiagnosis and/or radiotherapy of a patient who has been administered a cascade polymer complex having complex-forming ligands complexing at least five ions of an element of atomic numbers 27, 29, 31, 32, 37-39, 43, 49, 62, 64, 70, 75 or 77, as well as, optionally, cations of inorganic and/or organic bases, amino acids or amino acid amides. | 1. A method for radiodiagnosis and/or radiotherapy of a patient, comprising conducting radiodiagnosis and/or radiotherapy of a patient who has been administered a cascade polymer complex having complex-forming ligands complexing at least five ions of an element of atomic numbers 27, 29, 31, 32, 37-39, 43, 49, 62, 64, 70, 75 or 77, as well as, optionally, cations of inorganic and/or organic bases, amino acids or amino acid amides. 2. A method of claim 1, wherein the complex is administered in a dose of 0.0001-5 mmol/kg weight of the patient. 3. A method of claim 1, wherein the cascade polymer complex contains more than 15% by weight of the elements of atomic numbers 27, 29, 31, 32, 37-39, 43, 49, 62, 64, 70, 75 or 77. 4. A method of claim 1, wherein at least one complex-forming ligand is a macrocyclic ligand. 5. A method of claim 4, wherein the at least one macrocyclic complex-forming ligand is a tetraazacyclododecane macrocyclic ring ligand. 6. The method of claim 1, wherein the cascade polymer complex is of the formula I wherein A represents a nitrogen-containing cascade nucleus of basis multiplicity b, S represents a reproduction unit, independently for each generation, N represents a nitrogen atom, Z1 and Z2 for the first to penultimate generation, independently each represent, but for the last generation, Z1 represents a hydrogen atom, a C1-C10-alkyl, C2-C10-acyl or C1-C10-alkylsulfonyl radical optionally containing 1 to 3 carboxy groups, from 1 to 3 sulfonic acid groups, from 1 to 5 hydroxy groups and/or 1 to 3 oxygen atoms, or represents the radical of a complex former or complex K and Z2 represents the radical of a complex K, b represents a number 1 to 50, and s represents a number 1 to 3. 7. The method of claim 6, wherein in formula (I), K represents a radical of the formula IA, IB or IC bonded to the terminal nitrogen atoms of the last generation via V: wherein n and m in each case independently are 0, 1, 2, 3 or 4; n and m adding up to no more than 4, k represents the numbers 1, 2, 3, 4 or 5, l represents the numbers 0, 1, 2, 3, 4 or 5, q represents the numbers 0, 1 or 2, U represents CH2X or V, X represents the radicals —COOH, optionally in the form of an ester or an amide thereof, B, D and E which are identical or different represents the group —(CH2)a wherein a represents the numbers 2, 3, 4 or 5, R1 represents V or a hydrogen atom, with the proviso that R1 represents V only when U at the same time represents CH2X, and that U represents V only when R1 at the same time represents a hydrogen atom, V representing a straight-chain, branched, saturated or unsaturated C1-C20-alkenylene group optionally containing imino, phenylene, phenyleneoxy, phenyleneimino, amide, hydrazide, ureido, thioureido, carbonyl ester groups(s) and oxygen, sulphur and/or nitrogen atom(s), and optionally substituted by hydroxy, mercapto, imino, epoxy, oxo, thioxo and/or by amino group(s). 8. A method for detecting tumors in a patient, comprising cond... | Schering, AG | A61B005/055
US06853856 | B2 | 2005-02-08 | Diagnostic imaging interventional apparatus | 2001-11-21 | A medical imaging system for conducting an image-guided medical procedure on a subject and a method for performing the same is provided. The system includes a medical imaging apparatus, such as a CT scanner, magnetic resonance imaging system, or ultrasonic imaging system, etc., for obtaining volumetric images of the subject. Through intervention planning techniques, an interventional procedure on a subject using the volumetric images is determined. A mechanical arm assembly disposed in proximity to the medical imaging apparatus carries out the interventional procedure. The mechanical arm assembly includes a base support, a distal end, a plurality of arm segments, and a plurality of joints between the arm segments for carrying out the interventional procedure. An end-effector is disposed at the distal end of the mechanical arm assembly. The end-effector includes gripping means for selectively gripping and releasing a surgical instrument during the interventional procedure. | 1. A medical imaging system for conducting an image-guided medical procedure on a subject, the system comprising: a medical imaging apparatus for obtaining volumetric images of the subject; means for planning an interventional procedure on a subject using the volumetric images; a mechanical arm assembly disposed in proximity to the medical imaging apparatus, the mechanical arm assembly comprising a base support, a distal end, a plurality of arm segments, and a plurality of joints between the arm segments for carrying out the interventional procedure; and an end-effector disposed at the distal end of the mechanical arm assembly, the end-effector comprising gripping means for selectively gripping a surgical instrument during the interventional procedure with a gripping force ranging from zero to a force which prevents relative movement between the gripping means and the surgical instrument wherein the end-effector further comprises; a first finger portion having a first gripping surface; a second finger portion having a second gripping surface, the first and second gripping surfaces being opposed to one another for applying a gripping force to the surgical instrument; a first surgical instrument guide disposed on the first finger portion and extending towards the second finger portion; and a second surgical instrument guide disposed on the second finger portion end extending towards the first finger portion. | 1. A medical imaging system for conducting an image-guided medical procedure on a subject, the system comprising: a medical imaging apparatus for obtaining volumetric images of the subject; means for planning an interventional procedure on a subject using the volumetric images; a mechanical arm assembly disposed in proximity to the medical imaging apparatus, the mechanical arm assembly comprising a base support, a distal end, a plurality of arm segments, and a plurality of joints between the arm segments for carrying out the interventional procedure; and an end-effector disposed at the distal end of the mechanical arm assembly, the end-effector comprising gripping means for selectively gripping a surgical instrument during the interventional procedure with a gripping force ranging from zero to a force which prevents relative movement between the gripping means and the surgical instrument wherein the end-effector further comprises; a first finger portion having a first gripping surface; a second finger portion having a second gripping surface, the first and second gripping surfaces being opposed to one another for applying a gripping force to the surgical instrument; a first surgical instrument guide disposed on the first finger portion and extending towards the second finger portion; and a second surgical instrument guide disposed on the second finger portion end extending towards the first finger portion. 2. A medical imaging system according to claim 1 further comprising incremental movement means for incrementally inserting the surgical instrument along a trajectory into the subject. 3. A medical imaging system according to claim 2 wherein the trajectory is linear. 4. A medical imaging system according to claim 1 wherein the gripping force has a magnitude that allows the surgical instrument to be manually inserted into the subject. 5. A medical imaging system according to claim 1 wherein the first surgical instrument guide applies a first force to the surgical instrument in a direction perpendicular to the first gripping surface and the second surgical instrument guide applies a second force to the surgical instrument in a direction perpendicular to the second gripping surface. 6. A medical imaging system according to claim 5 wherein the first force and the second force are substantially zero. 7. A method of conducting an image-guided medical procedure on a subject, the method comprising: generating at least one volumetric image of the subject; planning an interventional procedure on the subject using the at least one volumetric image; gripping a surgical instrument with an end-effector, the end-effector being disposed at a distal end of a mechanical arm assembly; moving the surgical instrument into the subject along a trajectory in accordance with the planned interventional procedure using the mechanical arm assembly; and releasing the surgical instrument wherein the step of gripping comprises the stem of: applying a gripping force to ... | Koninklijke Philips Electronics N.V. | A61B005/05 A61B019/00
US06970737 | B1 | 2005-11-29 | Portable ECG device with wireless communication interface to remotely monitor patients and method of use | 2000-09-13 | A portable ECG monitor and an overall system for remotely monitoring cardiac function of a patient is disclosed, together with a method of use. The portable ECG includes a multi-lead, multi-channel ECG monitor and a wireless communication device connected to the ECG monitor to receive patient ECG data and transmit the patient ECG data to a centralized facility, such as a hospital. The wireless communication device can include a mobile phone and/or an interactive Internet appliance. A method of remotely monitoring ECG data is also disclosed. The method and apparatus are particularly useful with patients experiencing symptomatic ischemia. The method includes providing a portable ECG device with wireless communication capabilities to such a patient, acquiring ECG data from the patient at a location remote from a health care facility, then transmitting the ECG data to the centralized facility, and assessing the ECG data at the centralized facility. The patient is then provided with instructions based on the ECG assessment. The centralized facility and the health care facility may be one in the same, or may be two different and distinct facilities. | 1. An apparatus comprising: a lead wire assembly, each lead wire having a transducer capable of receiving an ECG signal from a patient;a portable ECG device including: a portable, ECG monitor adapted to be connected to the lead wire assembly, the ECG monitor having a data link port and a processor to process the ECG signals from the lead wire assembly and produce standard 12-lead ECG data representative of cardiac condition of the patient; anda wireless communication interface integrated with the ECG monitor to receive patient ECG data from the ECG monitor and transmit patient ECO data to a remote health care provider;an information management system connectable to the data link port of the ECG monitor to maintain ECO monitoring during patient transport to a health care facility, the information management system having data storage to maintain an ECG history that is downloadable at the health care facility; andwherein the processor of the portable ECG device is programmed to prompt the patient if assistance is needed to acquire an ECG, and if so, open a data transmission link to the health care provider otherwise, receive and process the ECG signals, then open a data transmission link and transmit the ECG data to the health care provider. | 1. An apparatus comprising: a lead wire assembly, each lead wire having a transducer capable of receiving an ECG signal from a patient;a portable ECG device including: a portable, ECG monitor adapted to be connected to the lead wire assembly, the ECG monitor having a data link port and a processor to process the ECG signals from the lead wire assembly and produce standard 12-lead ECG data representative of cardiac condition of the patient; anda wireless communication interface integrated with the ECG monitor to receive patient ECG data from the ECG monitor and transmit patient ECO data to a remote health care provider;an information management system connectable to the data link port of the ECG monitor to maintain ECO monitoring during patient transport to a health care facility, the information management system having data storage to maintain an ECG history that is downloadable at the health care facility; andwherein the processor of the portable ECG device is programmed to prompt the patient if assistance is needed to acquire an ECG, and if so, open a data transmission link to the health care provider otherwise, receive and process the ECG signals, then open a data transmission link and transmit the ECG data to the health care provider. 2. The apparatus of claim 1 wherein the wireless communication interface is a wireless phone capable of allowing audio and ECO data transmission concurrently. 3. The apparatus of claim 1 wherein the wireless communication interface is an interactive Internet TV appliance capable of allowing voice, video, and ECG data transmission concurrently. 4. The apparatus of claim 1 wherein the processor is further programmed to: allow selection of a desired transmission mode; andallow concurrent transmission of ECG data in addition to at least audio communication data. 5. The apparatus of claim 4 wherein the processor is further programmed to include bi-directional video and audio transmission with the transmission of ECG data. 6. The apparatus of claim 1 further comprising: an interactive Internet appliance that is connectable to a video and audio monitor to receive ECG data from the wireless communication interface and to transmit the ECG data to the health care provider; anda video camera and a microphone connected to the interactive Internet appliance to transmit video and audio data from the patient to the health care provider. 7. The apparatus of claim 6 wherein the apparatus is adapted to transmit the ECG data and the audio and video data to the health care provider through an interconnected global computer system. 8. The apparatus of claim 6 wherein the apparatus is adapted to transmit the ECG data and the audio and video data to the health care provider at least partially through an electromagnetic transmission wave. 9. The apparatus of claim 6 wherein the wireless communication interface includes an infrared transmitter and an infrared receiver to communicate with the interactive Internet appliance, and ... | GE Medical Systems Information Technologies, Inc. | A61B005/0432
US06852085 | B2 | 2005-02-08 | Fever alarm system | 2003-01-20 | A body temperature measurement device includes a unit that continually measures body temperature and transmits the measurement through an RF transmitter to a display unit. The display unit includes an RF receiver, a processor, and a display that shows the temperature. The display unit includes an adjustable threshold alarm circuit that turns on an alarm whenever the temperature rises above a defined threshold. | 1. A method of calculating body temperature, the method comprising: contacting a patient's skin with a temperature sensing unit; measuring skin temperature using the temperature sensing unit; obtaining a plurality of ambient temperature measurements over a time period substantially before measuring the skin temperature, the plurality of ambient temperature measurements obtained while the sensing unit is in contact with the skin; and correlating at least the measured skin temperature and the plurality of ambient temperature measurements to a body temperature using empirical data stored in look-up tables. | 1. A method of calculating body temperature, the method comprising: contacting a patient's skin with a temperature sensing unit; measuring skin temperature using the temperature sensing unit; obtaining a plurality of ambient temperature measurements over a time period substantially before measuring the skin temperature, the plurality of ambient temperature measurements obtained while the sensing unit is in contact with the skin; and correlating at least the measured skin temperature and the plurality of ambient temperature measurements to a body temperature using empirical data stored in look-up tables. 2. The method of claim 1, wherein the plurality of ambient temperature measurements comprises a history of changes in ambient temperature. 3. A temperature measurement system comprising: a sensing unit configured to be attached in contact with the skin of a patient, wherein the sensing unit comprises a first transmitter configured to transmit a first signal carrying data related to a temperature of the patient; an auditory sensing unit comprising a microphone and an amplifier, the auditory sensing unit configured to collect sound and to transmit an audio signal corresponding to the sound; an intermediate receiving and transmitting unit configured to receive the first signal and further configured to transmit a second signal carrying data related to the temperature of the patient wherein the data carried by the second signal is based at least upon the data carried by the first signal, the intermediate receiving and transmitting unit comprising the auditory sensing unit; and a monitoring unit configured to receive the second signal and the audio signal, and further configured to display the temperature of the patient based at least upon the data carried by the second signal, the monitoring unit comprising a speaker configured to reproduce the sound in response to the audio signal. 4. The system of claim 3, wherein the data carried by the second signal comprises the data carried by the first signal. 5. The system of claim 3, wherein the data carried by the first signal comprises the temperature of the patient. 6. The system of claim 3, wherein the data carried by the second signal comprises the temperature of the patient. 7. The system of claim 3, wherein the intermediate receiving and transmitting unit is configured to transmit the second signal over a greater distance than the first transmitter is configured to transmit the first signal. 8. The system of claim 3, further comprising a processing unit configured to determine the temperature of the patient. 9. The system of claim 8, wherein the sensing unit comprises the processing unit. 10. The system of claim 8, wherein the intermediate receiving and transmitting unit comprises the processing unit. 11. The system of claim 8, wherein the monitoring unit comprises the processing unit. 12. A method of calculating a body temperature of a patient, the method comprising: measuring a plurality... | Eliahu Rubinstein | A61B005/00
US06961608 | B2 | 2005-11-01 | Interventional MR imaging with detection and display of device position | 2001-06-04 | A magnetic resonance imaging system for interventional MR imaging involves operations to insert a device such as a catheter into an object. According to the system, a tip position of the catheter is detected, data indicative of moved loci of the catheter are produced from data indicative of the detected tip position, and the produced movement locus data are displayed. In addition, even when an operator changes the progress direction of the device such as a puncture needle, an imaging cross-section automatically tracks movements of the device. Appropriate preoperative planning is also provided. | 1. A magnetic resonance imaging system for interventional MRI involving an operation to insert a device into an object, the device having a tip and being gradually inserted into the object from the tip thereof, comprising: detection means for detecting a position of a tip of the device within the object; and movement state display means for displaying information in relation to a movement state of the tip of the device on the basis of data indicative of the position of the tip detected by the detection means; wherein the movement state display means includes production means for producing, as the information in relation to the movement state, movement locus data indicative of an inserted part of the device from the data indicative of the position of the tip detected by the detection means; and movement locus display means for displaying the movement locus data produced by the production means. | 1. A magnetic resonance imaging system for interventional MRI involving an operation to insert a device into an object, the device having a tip and being gradually inserted into the object from the tip thereof, comprising: detection means for detecting a position of a tip of the device within the object; and movement state display means for displaying information in relation to a movement state of the tip of the device on the basis of data indicative of the position of the tip detected by the detection means; wherein the movement state display means includes production means for producing, as the information in relation to the movement state, movement locus data indicative of an inserted part of the device from the data indicative of the position of the tip detected by the detection means; and movement locus display means for displaying the movement locus data produced by the production means. 2. The magnetic resonance imaging system according to claim 1, wherein the production means includes means for marking the data indicative of the position of the tip at every appropriate time instant, and means for grouping the data existing between the marked time instants at every marking, the grouped data being outputted as the movement locus data, wherein the movement locus display means is configured to display the grouped data in different modes group by group. 3. The magnetic resonance imaging system according to claim 2, wherein the device has an operating handle portion handed by an operator, the operating handle portion being located at an end opposite to the tip, and the production means includes operation means, which is arranged at the operating handle portion of the device, for generating a signal for the marking the data. 4. The magnetic resonance imaging system according to claim 2, wherein the movement locus display means is configured to display the grouped data in either one of different degrees of intensity or different kinds of color. 5. The magnetic resonance imaging system according to claim 4, wherein the movement locus display means is configured to display the grouped data in either one of different degrees of intensity or different kinds of color depending on a period of time elapsing after each grouping. 6. The magnetic resonance imaging system according to claim 1, wherein the movement locus display means is configured to display the movement locus data produced by the production means superimposed on a reference image. 7. A magnetic resonance imaging system for interventional MRI involving an operation to insert a device into an object, the device having a tip and being gradually inserted into the object from the tip thereof, comprising: detection means for detecting a position of a tip of the device within the object; and movement state display means for displaying information in relation to a movement state of the tip of the device on the basis of data indicative of the position of the tip detected by the detection me... | Kabushiki Kaisha Toshiba | A61B005/05
US06996430 | B1 | 2006-02-07 | Method and system for displaying cross-sectional images of a body | 1999-12-12 | This invention is a method of displaying cross-sectional images of a body so as to render the cross-sectional images more interpretable. The method is effected by using a first imaging device (110, 120), obtaining at least one cross-sectional image of a portion of the body, at least one cross-sectional image being defined by at least one first plane; and projecting each of at least one cross-sectional image of the portion of the body onto a predefined second plane, so as to obtain at least one projected cross-sectional image of the portion of the body, each of the at least one projected cross-sectional image becoming more interpretable by a user. | 1. A method of displaying cross-sectional images of a body so as to render the cross-sectional images more interpretable, the method comprising the steps of: (a) using a first intra-body imaging device deployed on a flexible catheter, obtaining at least one cross sectional image of a portion of the body, said at least one cross sectional image being defined by at least one first plane;(b) employing an inductive electromagnetic-field-sensing position-measuring system, including a locating sensor implemented on said first imaging device, to establish a location and orientation of each of said first plane in six degrees of freedom; and(c) projecting each of said at least one cross-sectional image of said portion of the body onto a predefined second plane, so as to obtain at least one projected cross-sectional image of said portion of the body, each of said at least one projected cross-sectional image becoming more interpretable by a user. | 1. A method of displaying cross-sectional images of a body so as to render the cross-sectional images more interpretable, the method comprising the steps of: (a) using a first intra-body imaging device deployed on a flexible catheter, obtaining at least one cross sectional image of a portion of the body, said at least one cross sectional image being defined by at least one first plane;(b) employing an inductive electromagnetic-field-sensing position-measuring system, including a locating sensor implemented on said first imaging device, to establish a location and orientation of each of said first plane in six degrees of freedom; and(c) projecting each of said at least one cross-sectional image of said portion of the body onto a predefined second plane, so as to obtain at least one projected cross-sectional image of said portion of the body, each of said at least one projected cross-sectional image becoming more interpretable by a user. 2. The method of claim 1, wherein said first imaging device is a single plane imaging device. 3. The method of claim 2, wherein said single plane imaging device is a single plane ultrasound device. 4. The method of claim 3, wherein said single plane ultrasound device is an intrabody single plane ultrasound device. 5. The method of claim 1, wherein said first imaging device is a multiple plane imaging device. 6. The method of claim 1, wherein said predefined second plane is at a preselected angle to a line-of-sight of a second imaging device used to image the body. 7. The method of claim 6, wherein said second imaging device is selected from the group consisting of a fluoroscope, a positron emission tomography device, a computerized tomography device, a magnetic resonance imaging device and a three dimensional-ultrasound device. 8. The method of claim 6, wherein said predefined second plane coincides with said line-of-sight. 9. The method of claim 6, wherein said predefined second plane is perpendicular to said line-of-sight. 10. The method of claim 1, wherein projecting each of said at least one cross-sectional image of said portion of the body onto said predefined second plane is effected by: (i) establishing a location and orientation of said predefined second plane in six degrees of freedom; and(ii) projecting each of said first plane onto said second predefined plane, thereby projecting each of said cross sectional image of said portion of the body onto said predefined second plane. 11. The method of claim 10, wherein said predefined second plane is at a pre selected angle to a line-of-sight of a second imaging device used to image the body. 12. The method of claim 11, wherein establishing said location and orientation of said predefined second plane in six degrees of freedom is effected by a locating sensor implemented on said second imaging device. 13. A method of displaying a three dimensional image rendered from a plurality of consecutive cross-sectional images of a body so as to render the rend... | Super Dimension LTD | A61B005/05
US07144376 | B2 | 2006-12-05 | Liable cardiac muscle diagnosing apparatus and liable cardiac muscle analyzing method by magnetic field measurements | 2001-07-17 | A magnetic field distribution measurement device (1) provides a non-contact magnetic measurement on a subject's chest at a plurality of coordinates and forms therefrom time-series magnetic field distribution data. A first arithmetic device (2) in response generates image data representing a three-dimensional, intramyocardial current density distribution. A second arithmetic device (3) receives a plurality of tomographic image data separately obtained by a tomographic diagnosis apparatus and processes the data to generate three-dimensional, anatomical image data. A display device (4) receives these data and displays on an anatomical image an image representing an intramyocardial current density. Thus in a ventricle a lesioned or viable part of myocardium can be readily identified in location, size, geometry and level. Furthermore, the anatomical image may be replaced with an image representing a normal stimulation propagation circuit and serving as a template. | 1. An apparatus employing a magnetic measurement to diagnose viable myocardium, comprising: a magnetic field distribution measurement device performing a non-contact magnetic measurement on a subject's chest at a plurality of coordinates to obtain a plurality of time-series magnetic data corresponding to said plurality of coordinates, and also using said plurality of time-series magnetic data to generate time-series, magnetic field distribution data on said chest;a first arithmetic device using said generated time-series, magnetic field distribution data to generate time-series, intramyocardial current density distribution data of said subject;a second arithmetic device processing separately provided, tomographic, thoracic data of said subject to generate data representative of an anatomical image;a display device displaying an image of said intramyocardial, current density distribution represented by said data generated by said first arithmetic device, as superimposed on said anatomical image represented by said data generated by said second arithmetic device, thereby capable of three-dimensionally identifying a lesioned part of myocardium presenting an abnormal current density distribution or a physiologically viable part of myocardium; anda database containing information used to determine a relationship between an intramyocardial current density represented by data generated by said first arithmetic device and a myocardial lesion current. | 1. An apparatus employing a magnetic measurement to diagnose viable myocardium, comprising: a magnetic field distribution measurement device performing a non-contact magnetic measurement on a subject's chest at a plurality of coordinates to obtain a plurality of time-series magnetic data corresponding to said plurality of coordinates, and also using said plurality of time-series magnetic data to generate time-series, magnetic field distribution data on said chest;a first arithmetic device using said generated time-series, magnetic field distribution data to generate time-series, intramyocardial current density distribution data of said subject;a second arithmetic device processing separately provided, tomographic, thoracic data of said subject to generate data representative of an anatomical image;a display device displaying an image of said intramyocardial, current density distribution represented by said data generated by said first arithmetic device, as superimposed on said anatomical image represented by said data generated by said second arithmetic device, thereby capable of three-dimensionally identifying a lesioned part of myocardium presenting an abnormal current density distribution or a physiologically viable part of myocardium; anda database containing information used to determine a relationship between an intramyocardial current density represented by data generated by said first arithmetic device and a myocardial lesion current. 2. An apparatus employing a magnetic measurement to diagnose viable myocardium, comprising: a magnetism distribution measurement device performing a non-contact magnetic measurement on a subject's chest at a plurality of coordinates to obtain a plurality of time-series magnetic data corresponding to said plurality of coordinates, and also using said plurality of time-series magnetic data to generate time-series magnetic field distribution data on said chest;an arithmetic device using said generated time-series, magnetism distribution data to generate time-series, intramyocardial current density distribution data of said subject; anda display device using the data generated by said arithmetic device to superimpose together an image representing a stimulation propagation path of said subject extending from a sinoatrial node to a bundle of His-Purkinje fiber network and an image representing an intramyocardial current density distribution and display said images, thereby capable of three-dimensionally identifying a lesioned part of myocardium presenting an abnormal current density distribution or a physiologically viable part of myocardium. 3. The apparatus of claim 2, wherein said arithmetic device operates based on an anatomical or functional factor to divide right and left ventricles into any plurality of regions and generate time-series data of an intramyocardial current density distribution for each said region to allow a lesioned part of myocardium in the right and left ventricles to be three-dimensiona... | Japan Science and Technology Corporation Takenaka Corporation | A61B005/02
US07050842 | B2 | 2006-05-23 | Method of tissue modulation for noninvasive measurement of an analyte | 2001-07-11 | Disclosed are a device and methods for obtaining signals related to blood analytes and bone features by collecting spectra, such as fluorescence and Raman spectra, through the nail of a finger or toe. The invention additionally provides methods of modulating blood flow in a tissue of a subject. | 1. A method of noninvasive spectroscopic measurement of an analyte in a subject comprising: (a) irradiating tissue of the subject through a nail with electromagnetic radiation having an excitation wavelength while the tissue is in a blood-replete state;(b) collecting the Raman spectra emitted by the tissue in the blood-replete state;(c) applying a tissue modulation device having an orifice to the nail of the subject so as to achieve a blood-depleted state within the tissue;(d) irradiating the tissue through the nail with electromagnetic radiation having an excitation wavelength while the tissue is in the blood-depleted state;(e) collecting the Raman spectra emitted by the tissue in the blood-depleted state; and(f) analyzing the collected Raman spectra to determine a concentration of analyte present in the tissue, wherein the analyzing comprises determining the difference between the Raman spectra collected in the blood-replete and blood-depleted states. | 1. A method of noninvasive spectroscopic measurement of an analyte in a subject comprising: (a) irradiating tissue of the subject through a nail with electromagnetic radiation having an excitation wavelength while the tissue is in a blood-replete state;(b) collecting the Raman spectra emitted by the tissue in the blood-replete state;(c) applying a tissue modulation device having an orifice to the nail of the subject so as to achieve a blood-depleted state within the tissue;(d) irradiating the tissue through the nail with electromagnetic radiation having an excitation wavelength while the tissue is in the blood-depleted state;(e) collecting the Raman spectra emitted by the tissue in the blood-depleted state; and(f) analyzing the collected Raman spectra to determine a concentration of analyte present in the tissue, wherein the analyzing comprises determining the difference between the Raman spectra collected in the blood-replete and blood-depleted states. 2. The method of claim 1, wherein the nail is a fingernail. 3. The method of claim 1, wherein the nail is a toenail. 4. The method of claim 1, wherein the analyte comprises glucose. 5. The method of claim 1, wherein the analyte comprises hydroxyapatite. 6. The method of claim 1, wherein the blood-replete stare is achieved by hindering blood flow through an artery that serves the tissue. 7. The method of claim 1, wherein the blood-replete state is achieved by altering the temperature of the tissue. 8. The method of claim 7, wherein the temperature is altered by heating. 9. The method of claim 7, wherein the temperature is altered by cooling. 10. A method of noninvasive spectroscopic measurement of an analyte in a subject comprising: (a) irradiating tissue of the subject through a nail with electromagnetic radiation having an excitation wavelength by directing the electromagnetic radiation at the nail side of the finger or toe of the subject;(b) collecting the spectra emitted by the tissue through the nail side of the finger or toe of the subject in blood-replete and blood-depleted states;(c) analyzing collected Raman spectra to determine a signal that is proportional to the amount of analyte present in the tissue by determining the difference between Raman spectra collected in the blood-replete state and the blood-depleted state;(d) analyzing spectra collected in step (b) that are indicative of blood volume; and(e) comparing the spectra analyzed in step (c) with the spectra analyzed in step (d) to determine the amount of analyte per unit blood volume. 11. The method of claim 10, wherein the spectra indicative of blood volume comprise fluorescence, phosphorescence or absorption spectra. 12. The method of claim 10, wherein the nail is a fingernail. 13. The method of claim 10, wherein the nail is a toenail. 14. The method of claim 10, wherein the analyte comprises glucose. 15. The method of claim 10, wherein the analyte comprises hydroxyapatite. 16. The method of claim 10, wherein the b... | LighTouch Medical, Inc. | A61B005/00
US07043296 | B1 | 2006-05-09 | Bioelectric signal detector and massage machine | 2000-11-17 | A biosignal sensing device (2) includes an electrode section (12) including a pair of electrodes (12a,12b) mounted so that a living body can be in close proximity to or in contact with the electrodes. An oscillation section (11) supplies a high frequency signal to the electrodes. A reflected-wave level from the electrodes receiving the high frequency signal changes according to a change in impedance in the electrodes that are in close proximity to or in contact with the living body. This change in reflected-wave level reflects a change in impedance resulting from body movement of the living body such as respiration and pulsation. A sensing section (13) senses a signal indicating the change in reflected-wave level. Based on this signal, a signal amplifying section (14) calculates a desired biosignal such as pulse rate and respiration rate. | 1. A biosignal sensing device for sensing a biosignal from a living body as a subject, comprising: an oscillating section for supplying a high frequency signal;an electrode section including an electrode mounted so that said living body as a subject can be in close proximity to or in contact with said electrode, said electrode being coupled to receive said high frequency signal from said oscillating section;a sensing section for sensing a change in impedance that occurs in said electrode receiving said high frequency signal when said living body as a subject is in close proximity to or in contact with said electrode;a signal amplifying section having a filter function to extract a signal according a desired biosignal from a signal corresponding to said sensed change in impedance; anda signal processing section for producing said desired biosignal from said signal extracted by said signal amplifying section, whereinsaid sensing section produces and supplies a signal indicating a change in reflected-wave level of said high frequency signal corresponding to said change in impedance that occurs in said electrode according to body movement of said living body, anda frequency of said high frequency signal supplied from said oscillating section is preset to a value in a region where said reflected-wave level changes at a high rate according to change in said frequency with respect to a region where said reflected-wave level changes at a minimum rate according to change in said frequency,wherein said desired biosignal is indicative of a respiration rate. | 1. A biosignal sensing device for sensing a biosignal from a living body as a subject, comprising: an oscillating section for supplying a high frequency signal;an electrode section including an electrode mounted so that said living body as a subject can be in close proximity to or in contact with said electrode, said electrode being coupled to receive said high frequency signal from said oscillating section;a sensing section for sensing a change in impedance that occurs in said electrode receiving said high frequency signal when said living body as a subject is in close proximity to or in contact with said electrode;a signal amplifying section having a filter function to extract a signal according a desired biosignal from a signal corresponding to said sensed change in impedance; anda signal processing section for producing said desired biosignal from said signal extracted by said signal amplifying section, whereinsaid sensing section produces and supplies a signal indicating a change in reflected-wave level of said high frequency signal corresponding to said change in impedance that occurs in said electrode according to body movement of said living body, anda frequency of said high frequency signal supplied from said oscillating section is preset to a value in a region where said reflected-wave level changes at a high rate according to change in said frequency with respect to a region where said reflected-wave level changes at a minimum rate according to change in said frequency,wherein said desired biosignal is indicative of a respiration rate. 2. A biosignal sensing device for sensing a biosignal from a living body as a subject, comprising: an oscillating section for supplying a high frequency signal;an electrode section including an electrode mounted so that said living body as a subject can be in close proximity to or in contact with said electrode, said electrode being coupled to receive said high frequency signal from said oscillating section;a sensing section for sensing a change in impedance that occurs in said electrode receiving said high frequency signal when said living body as a subject is in close proximity to or in contact with said electrode;a signal amplifying section having a filter function to extract a signal according a desired biosignal from a signal corresponding to said sensed change in impedance; anda signal processing section for producing said desired biosignal from said signal extracted by said signal amplifying section, whereinsaid sensing section produces and supplies a signal indicating a change in reflected-wave level of said high frequency signal corresponding to said change in impedance that occurs in said electrode according to body movement of said living body, anda frequency of said high frequency signal supplied from said oscillating section is preset to a value in a region where said reflected-wave level changes at a minimum rate according to change in said frequency,wherein said desired biosignal is indic... | Omron Corporation | A61B005/05
US07096061 | B2 | 2006-08-22 | Apparatus for monitoring CHF patients using bio-impedance technique | 2002-07-03 | A method is described for generating an impedance image of the chest, in which electrical data of the chest is acquired, and a finite volume method is used to calculate an impedance image from the electrical data, using an analytical expression for the Jacobian. | 1. A method for generating an impedance image of a subject's chest, comprising: acquiring electrical data of the chest;formulating an initial impedance image;using a finite volume method to calculate an expected set of electrical data if the impedance distribution of the chest matched the initial impedance image;determining a difference between the acquired electrical data and the expected electrical data; andgenerating a new impedance image based on said difference;wherein generating a new impedance image comprises expressing a Jacobian analytically. | 1. A method for generating an impedance image of a subject's chest, comprising: acquiring electrical data of the chest;formulating an initial impedance image;using a finite volume method to calculate an expected set of electrical data if the impedance distribution of the chest matched the initial impedance image;determining a difference between the acquired electrical data and the expected electrical data; andgenerating a new impedance image based on said difference;wherein generating a new impedance image comprises expressing a Jacobian analytically. 2. A method according to claim 1, wherein calculating an expected set of electrical data and generating a new impedance image are iterated at least one time, using the new impedance image generated in at least one previous iteration to calculate the expected set of electrical data in each iteration except the first iteration. 3. A method according to claim 2, wherein calculating an expected set of electrical data and generating a new impedance image are iterated until the difference between the acquired electrical data and the expected set of electrical data is small enough to satisfy a stopping condition. 4. A method according to claim 1, wherein generating the new impedance image comprises generating with a Newton-Raphson method. 5. A method according to claim 1, wherein generating the new impedance image comprises generating with a modified Newton-Raphson method. 6. A method according to claim 5, wherein generating with a modified Newton-Raphson method comprises adding a regularization matrix to a Hessian matrix. 7. A method according to claim 1, wherein formulating the initial impedance image comprises ascribing typical impedances to different parts of the chest according to at least one image of the chest. 8. A method according to claim 7, wherein ascribing impedances according to at least one image of the chest comprises ascribing impedances according to at least one x-ray image. 9. A method according to claim 8, wherein ascribing impedances according to at least one x-ray image comprises ascribing impedances according to at least one x-ray computed tomography image. 10. A method according to claim 7, wherein ascribing impedance according to at least one image of the chest comprises ascribing impedances according to at least one magnetic resonance image. 11. A method according to claim 7, wherein ascribing impedances according to at least one image of the chest comprises ascribing impedances according to at least one ultrasound image. 12. A method according to claim 1, wherein using the finite volume method comprises inverting a matrix with a technique that is adapted for inverting sparse matrixes. 13. A method according to claim 12, wherein inverting a matrix comprises inverting a matrix with a successive over relaxation method. 14. A method according to claim 1, wherein acquiring electrical data of the chest comprises measuring potentials at a plurality of locations on the subje... | Tel-Aviv University Future Technology Development L.P. | A61B005/05
US07058443 | B2 | 2006-06-06 | Diagnostic features in biatrial and biventricular pacing systems | 2001-04-26 | A biatrial and/or biventricular pacing system is used in a diagnostic context. By placing a pacing/sensing lead in three or four chambers of the heart, various conduction sequences can be determined and the originating chamber of various arrhythmias can be identified. This information is stored temporarily in the pacemaker until it is extracted for analysis. | 1. A method of utilizing a biventricular pacing system to determine the distribution of ventricle to ventricle conduction sequences in a patient having a conductive disorder, the method comprising: placing sensing leads in both ventricular chambers;sensing conduction sequences occurring from one ventricular chamber to another ventricular chamber;determining which ventricular chamber the conduction sequence originated in and which ventricular chamber it propagated to;recording the determining information in a memory such that the information can be used to identify the relative distribution of conduction sequences;identifying a conductive disorder in response to the determined relative distribution; andadjusting a therapy delivered by the pacing system in response to the identified conductive disorder, wherein the conductive disorder comprises a conductive disorder amenable to termination via anti tachycardia pacing (ATP) therapy delivery and wherein the therapy comprises ATP therapy and the adjusting further comprises: initiating the ATP therapy in the ventricular that one of: initiated a recent ventricular depolarization and initiated a majority of ventricular depolarizations over a predetermined time period. | 1. A method of utilizing a biventricular pacing system to determine the distribution of ventricle to ventricle conduction sequences in a patient having a conductive disorder, the method comprising: placing sensing leads in both ventricular chambers;sensing conduction sequences occurring from one ventricular chamber to another ventricular chamber;determining which ventricular chamber the conduction sequence originated in and which ventricular chamber it propagated to;recording the determining information in a memory such that the information can be used to identify the relative distribution of conduction sequences;identifying a conductive disorder in response to the determined relative distribution; andadjusting a therapy delivered by the pacing system in response to the identified conductive disorder, wherein the conductive disorder comprises a conductive disorder amenable to termination via anti tachycardia pacing (ATP) therapy delivery and wherein the therapy comprises ATP therapy and the adjusting further comprises: initiating the ATP therapy in the ventricular that one of: initiated a recent ventricular depolarization and initiated a majority of ventricular depolarizations over a predetermined time period. 2. The method of claim 1, further comprising: measuring the temporal dimension from the beginning to the conclusion of a plurality of prior conductive sequences; andstoring the measured information in the memory by a range of said temporal dimensions so that the measured information can be utilized to select and adjust the anti tachycardia pacing therapy based on at least one recently measured temporal dimension. 3. The method of claim 2, wherein each measured conductive sequence increments a unit counter representing one of a plurality of temporal dimensions. 4. The method of claim 2, further comprising: pacing one ventricular chamber in order to generate a conductive sequence. 5. A biventricular pacing system for determining the distribution of conduction sequences from a first ventricle (V1) to a second ventricle (V2) in a patient having a conductive disorder, comprising: sensing means located in both ventricular chambers (V1,V2) for sensing conduction sequences occurring from one ventricular chamber to another ventricular chamber;means for determining which ventricular chamber the conduction sequence originated in and which ventricular chamber it propagated to;means for recording the determined information in a memory such that the information can be used to identify the relative distribution of conduction sequences; andmeans for detecting an arrhythmia susceptible to termination via anti tachycardia pacing (ATP) therapy and adjusting the ATP therapy based at least in part upon the ventricular chamber the arrhythmia originated in. 6. The biventricular pacing system of claim 5, further comprising: means for measuring the timing of each conductive sequence and including the measured timing information in the memory to identify relati... | Medtronic, Inc. | A61B005/0402
US07057389 | B2 | 2006-06-06 | Nuclear magnetic resonance diagnosing device and diagnosing system | 2002-06-17 | A nuclear magnetic resonance diagnosing apparatus for continuously taking tomograms in a space formed between a pair of magnets includes a cylindrical table top support of a gantry side, a pair of magnets vertically arranged, a table top support of the gantry side being interposed between the pair of magnets; and a column for supporting at least the magnet arranged in an upper side between the pair of magnets. The column is formed so that between a gantry and an examining table body, a width between a bridge construction, which is provided between the column and the table top support of the gantry side, and the table top support of the gantry side may be smaller than an examination width L determined by lines R′ passing through a center of the table top support of the gantry side perpendicular to a center line R. | 1. A nuclear magnetic resonance diagnosing device configured for continuously taking tomograms of an examined person on a bed provided in order to insert the examined person into a space formed between a pair of an upper magnet and a lower magnet, which comprise: a cylindrical tabletop support of a gantry side;said pair of magnets vertically arranged, said table top support of the gantry side being interposed between said pair of said upper magnet and said lower magnet;a column configured for supporting at least said upper magnet; anda bridge construction arranged between said column and said table top support of the gantry side, wherein said bed comprises a slide table sliding on said tabletop support of the gantry side;a slide table support means for mounting said slide table in order to be slidable; anda support system having a slide drive unit configured for the slide table, said table top support of the gantry side having a circular side face portion along which said bed is operatively connected so as to be slidably movable along at least a 200° periphery of said circular side face portion,said upper magnet is covered with a top magnet cover and said lower magnet is covered with an examining table body, and first and second connecting parts between said table top support of the gantry side and said bridge construction are placed in quadrants adjacent said support column away from a straight line which passes through a center of said table top support of the gantry side and which crosses at right angles with a straight line passing through the center of said table top support of the gantry side and a center of said support column,said first and second connecting parts forming a portion having a width smaller than a width of said table top support of the gantry side on said straight line crossing at right angles, said table top support of the gantry side further having a contacting surface configured for positioning an extended line in a longitudinal direction of said bed arranged around said table top support of the gantry side so as to pass a center of said table top support of the gantry side, said contacting surface being formed with a circular arc continuously connecting said first and second connecting parts, andat least one control panel of said nuclear magnetic resonance diagnosing device being located on at least one of said upper magnet and said support column. | 1. A nuclear magnetic resonance diagnosing device configured for continuously taking tomograms of an examined person on a bed provided in order to insert the examined person into a space formed between a pair of an upper magnet and a lower magnet, which comprise: a cylindrical tabletop support of a gantry side;said pair of magnets vertically arranged, said table top support of the gantry side being interposed between said pair of said upper magnet and said lower magnet;a column configured for supporting at least said upper magnet; anda bridge construction arranged between said column and said table top support of the gantry side, wherein said bed comprises a slide table sliding on said tabletop support of the gantry side;a slide table support means for mounting said slide table in order to be slidable; anda support system having a slide drive unit configured for the slide table, said table top support of the gantry side having a circular side face portion along which said bed is operatively connected so as to be slidably movable along at least a 200° periphery of said circular side face portion,said upper magnet is covered with a top magnet cover and said lower magnet is covered with an examining table body, and first and second connecting parts between said table top support of the gantry side and said bridge construction are placed in quadrants adjacent said support column away from a straight line which passes through a center of said table top support of the gantry side and which crosses at right angles with a straight line passing through the center of said table top support of the gantry side and a center of said support column,said first and second connecting parts forming a portion having a width smaller than a width of said table top support of the gantry side on said straight line crossing at right angles, said table top support of the gantry side further having a contacting surface configured for positioning an extended line in a longitudinal direction of said bed arranged around said table top support of the gantry side so as to pass a center of said table top support of the gantry side, said contacting surface being formed with a circular arc continuously connecting said first and second connecting parts, andat least one control panel of said nuclear magnetic resonance diagnosing device being located on at least one of said upper magnet and said support column. 2. A nuclear magnetic resonance diagnosing device configured for continuously taking tomograms of an examined person on a bed provided in order to insert the examined person into a space formed between a pair of an upper magnet and a lower magnet, which comprises: a cylindrical tabletop support of a gantry side;said pair of magnets vertically arranged, said table top support of the gantry side being interposed between said pair of said upper magnet and said lower magnet;a column configured for supporting at least said magnet arranged in an upper side between said pair of magn... | Hitachi Medical Corporation | G01V003/00 A61B005/055 G01N023/04
US07190986 | B1 | 2007-03-13 | Non-adhesive oximeter sensor for sensitive skin | 2002-10-18 | The present invention provides non-adhesive oximeter sensors for patients with sensitive skin. Sensors of the present invention include a light emitting diode (LED) and a photodetector. The LED and the photodetector may be covered by a reflective mask and a faraday shield. Sensors of the present invention have a non-adhesive laminated layer. The non-adhesive layer contacts, but does not stick to, the patient's skin. When the sensor is removed from the patient, the non-adhesive layer does not tear or irritate the patient's skin. The non-adhesive layer preferably has a large static coefficient of friction. Sensors of the present invention can also have hook-and-loop layers. The sensor can be attached to the patient's body by wrapping the sensor around the patient and engaging the hook layer to the loop layer. | 1. An oximeter sensor comprising: a light emitting element;a light detecting element;a substrate adapted to support the light emitting element and the light detecting element, the substrate comprising a non-adhesive patient-contacting surface, the patient contacting-surface having a static coefficient of friction greater than 5 with respect to a patient's skin. | 1. An oximeter sensor comprising: a light emitting element;a light detecting element;a substrate adapted to support the light emitting element and the light detecting element, the substrate comprising a non-adhesive patient-contacting surface, the patient contacting-surface having a static coefficient of friction greater than 5 with respect to a patient's skin. 2. The oximeter sensor of claim 1 wherein: the substrate includes first and second receptacles, the light emitting element residing in the first receptacle, and the light detecting element residing in the second receptacle; and comprising:a transparent polyurethane layer that covers the first receptacle and the second-receptacle. 3. The oximeter sensor of claim 1 comprising: a hook portion and a loop portion coupled to the substrate, the hook portion being releasably coupleable to the loop portion to facilitate attaching the sensor to a patient. 4. The oximeter sensor of claim 1 wherein the first substrate predominantly has a thickness of less than 0.1 inch. 5. The oximeter sensor of claim 1 wherein: the substrate comprises polyvinyl chloride foam. 6. The oximeter sensor of claim 1 further comprising: a cable attached to the sensor; anda strap movably attached to the cable, wherein the strap can be attached to the patient. 7. The oximeter sensor of claim 1 further comprising: reflective mask layers over one or both of the light emitting element and light detecting element. 8. The oximeter sensor of claim 1 further comprising: faraday shields over one or both of the light emitting and light detecting elements. 9. The oximeter sensor of claim 1 further comprising: an opaque film that covers the inner surface of the first laminated layer. 10. The oximeter sensor of claim 9 wherein: the opaque film is dark in color. 11. A method for forming an oximeter sensor, the method comprising: providing a substrate having a surface that contacts a patient, the substrate comprising a non-adhesive material that has a static coefficient of friction greater than 5 with respect to the patient's skin, the non-adhesive material retaining the oximeter sensor on the patient without the use of an adhesive that contacts skin of the patient, the substrate having first and second receptacles; andplacing a light emitting element in the first receptacle;placing a light detecting element in the second receptacle. 12. The method of claim 11 further comprising: placing a first transparent polyurethane layer over the light emitting element; andplacing a second transparent polyurethane layer over the light detecting element. 13. The method of claim 11 further comprising: placing a first reflective layer over the light emitting element; andplacing a second reflective layer over the light detecting element. 14. The method of claim 11 further comprising: placing a hook layer and a loop layer on the substrate, wherein the loop layer is adapted to engage with the hook layer. 15. The method of claim 11 wherein the s... | Nellcor Puritan Bennett Inc. | A61B005/00
US06997877 | B2 | 2006-02-14 | Method for the measurement of post arteriolar pressure | 2002-08-30 | A method for the measurement of post arteriolar pressure in a patient which is determined from information regarding the blood pressure (BP), systemic vascular resistance (R) and cardiac output (CO) by plotting the regression line for the equation BP=R×Q+c where c is the post arteriolar pressure, and determining the value of c from the regression line. | 1. A method for determining the post arteriolar pressure in a patient, which method comprises the steps of: (i) recording and storing the arterial blood pressure waveform of a patient from a blood pressure monitoring device over a period of time;(ii) determining the nominal or actual stroke volume associated with each arterial blood pressure waveform;(iii) obtaining the nominal or actual cardiac output from the data in step (ii); and(iv) plotting the regression line for the equation BP=R×Q+c where BP is the blood pressure,R is the systemic vascular resistance,Q is the cardiac output, andc is the post arteriolar pressure, and(v) determining the value of the post arteriolar pressure from the regression line. | 1. A method for determining the post arteriolar pressure in a patient, which method comprises the steps of: (i) recording and storing the arterial blood pressure waveform of a patient from a blood pressure monitoring device over a period of time;(ii) determining the nominal or actual stroke volume associated with each arterial blood pressure waveform;(iii) obtaining the nominal or actual cardiac output from the data in step (ii); and(iv) plotting the regression line for the equation BP=R×Q+c where BP is the blood pressure,R is the systemic vascular resistance,Q is the cardiac output, andc is the post arteriolar pressure, and(v) determining the value of the post arteriolar pressure from the regression line. 2. A method as claimed in claim 1 wherein the arterial blood pressure is plotted in step (i) for a period of up to ten seconds. 3. A method as claimed in claim 1 wherein the arterial blood pressure is analysed on a beat-to-beat basis. 4. A method as claimed in claim 1 wherein the stroke volume in step (ii) is determined via a cardiac output measuring method. 5. A method as claimed in claim 4 wherein the cardiac output measuring method is a pressure waveform analysis. 6. A method as claimed in claim 5 wherein the data obtained in step (i) is subjected to Fourier analysis in order to obtain the modulus of the first harmonic and the nominal stroke volume is obtained from the modulus of the first harmonic and data relating to the arterial blood pressure and the heart rate. 7. A method as claimed in claim 6 wherein the nominal stroke volume is obtained from the following equation nominal⁢⁢stroke⁢⁢volume=modulus⁢⁢of⁢⁢first⁢⁢harmonic⁢⁢ofblood⁢⁢pressure⁢⁢waveforme+0.0092×MAP×HRwhere e is the base of natural logarithmsMAP is the mean arterial blood pressure, andHR is the heart rate. 8. A method as claimed in claim 1 wherein the regression line in step (iv) is plotted by means of a computer. 9. A method as claimed in claim 8 wherein the value of the post arteriolar pressure is displayed on a computer screen and optionally continuously updated. 10. A method as claimed in claim 2 wherein the arterial blood pressure is analysed on a beat-to-beat basis. 11. A method as claimed in 2 wherein the stroke volume in step (ii) is determined via a cardiac output measuring method. 12. A method as claimed in 3 wherein the stroke volume in step (ii) is determined via a cardiac output measuring method. 13. A method as claimed in 2 wherein the regression line in step (iv) is plotted by means of a computer. 14. A method as claimed in 3 wherein the regression line in step (iv) is plotted by means of a computer. 15. A method as claimed in 4 wherein the regression line in step (iv) is plotted by means of a computer. 16. A method as claimed in 5 wherein the regression line in step (iv) is plotted by means of a computer. 17. A method as claimed in 6 wherein the regression line in step (iv) is plotted by means of a computer. 18. A method as claimed in 7 wherei... | Lidco Group PLC | A61B005/00
US07141026 | B2 | 2006-11-28 | Body movement monitoring system and method | 2001-10-01 | The present invention relates to a system and a method for monitoring body movement, in particular but not exclusively for evaluating the falling risk and/or for monitoring sitting, standing, lying, walking and running periods. The system is characterized by the fact that it comprises means for determining the time of postural transition. | 1. A body movement monitoring system comprising: a sensor to be attached to a trunk of a subject,processing means for deriving information from said sensor,display means for displaying said information to an operator,wherein said system includes means for determining the time of postural transition by calculating a trunk tilt θ by integrating an angular rate signal measured by said sensor, calculating sin θ, and obtaining the postural transition corresponding to the minimum peak of sin θ. | 1. A body movement monitoring system comprising: a sensor to be attached to a trunk of a subject,processing means for deriving information from said sensor,display means for displaying said information to an operator,wherein said system includes means for determining the time of postural transition by calculating a trunk tilt θ by integrating an angular rate signal measured by said sensor, calculating sin θ, and obtaining the postural transition corresponding to the minimum peak of sin θ. 2. A monitoring system according to claim 1 wherein the sensor comprises a miniature gyroscope. 3. A monitoring system according to claim 1 wherein said system includes means to determine the duration and the type of postural transition. 4. A monitoring system according to claim 1 comprising an alarm which is activated when the quality of postural transition is behind a threshold. 5. A monitoring system according to claim 1 comprising furthermore a vertical accelerometer. 6. A monitoring system according to claim 5 comprising furthermore a forward accelerometer. 7. A method for measuring the time of postural transition comprising: attaching a sensor to a trunk of a subject,processing information from said sensor,displaying said information to an operator,determining the time when a postural transition occurs, wherein a trunk tilt θ is first calculated by integrating an angular rate signal measured by said sensor, sin θ is then calculated, and the postural transition corresponds to the minimum peak of sin θ. 8. A method according to claim 7 wherein a duration of postural transition is determined based on the duration between the maximum peaks of sin θ which occur before and after the instant of postural transition. 9. A method according to claim 7 wherein the method further determines a duration and a type of postural transition and wherein falling risk is determined by a combination of three parameters comprising an average and standard deviation of transition duration and an occurrence of abnormal successive transition. 10. A method according to claim 7 further comprising monitoring and measuring physical activity via a vertical accelerometer. | Ecole Polytechnique Federale de Lausanne/ Service des Relations Industrielles (SRI) | A61B005/103 A61B005/117
US06916291 | B2 | 2005-07-12 | Systems, methods and products for diagnostic hearing assessments distributed via the use of a computer network | 2002-02-05 | The systems, methods and associated devices performing diagnostic hearing tests which use a computer network to allow interaction between a test administration site and one or a plurality of remote patient sites. The test can be administered by an audiologist or clinician at a site remote from the patient, in a manner, which can allow interaction between the user and the clinician during at least a portion of the administration of the test. The diagnostic hearing tests can be performed such that they meet standardized guidelines such as ANSI requirements or certification standards and can include distortion product emission level measurements or middle ear compliance measurements. | 1. A method for performing a hearing evaluation test over a computer network, comprising the steps of: administering a hearing evaluation test to a patient using a computer network, the hearing evaluation test comprising a plurality of hearing assessment signals at selected frequencies and hearing levels; transmitting commands from a test administration site to a local patient testing site during said administering step; generating the hearing assessment signals at the local patient site in response to said transmitting step; interactively relaying information between the patient located at the local site and a clinician located at the test administration site during said administering step so that the clinician can evaluate the patient's response to the hearing assessment signals, the test administration site being remote from the local site; and controlling the output of the hearing assessment signals during the administering step such that the clinician at the test administration site determines which hearing assessment signals of the generating step are relayed locally to the patient. | 1. A method for performing a hearing evaluation test over a computer network, comprising the steps of: administering a hearing evaluation test to a patient using a computer network, the hearing evaluation test comprising a plurality of hearing assessment signals at selected frequencies and hearing levels; transmitting commands from a test administration site to a local patient testing site during said administering step; generating the hearing assessment signals at the local patient site in response to said transmitting step; interactively relaying information between the patient located at the local site and a clinician located at the test administration site during said administering step so that the clinician can evaluate the patient's response to the hearing assessment signals, the test administration site being remote from the local site; and controlling the output of the hearing assessment signals during the administering step such that the clinician at the test administration site determines which hearing assessment signals of the generating step are relayed locally to the patient. 2. A method according to claim 1, wherein the hearing evaluation test assessment signals are sufficient in number and variation of frequency and sound intensity to allow the clinician to perform a diagnostic hearing evaluation. 3. A method according to claim 2, wherein the auditory tone presentation of the hearing assessment signals during said administering step meets predetermined test standards. 4. A method according to claim 1, wherein said interactively relaying step includes relaying dynamic audiovisual communications of the patient at the local site to the test administration site during said administering step. 5. A method according to claim 4, further comprising the step of determining the ambient noise level at the local site and transmitting data regarding the determined noise level to the remote site prior to and/or during said administering step. 6. A method according to claim 1, further comprising the step of controlling the sound intensity of the hearing assessment signals delivered to the patient from said test administration site. 7. A method according to claim 1, wherein the hearing evaluation test assessment signals are sufficient to allow the clinician to measure (a) middle ear pressure and compliance characteristics and (b) acoustics reflexes. 8. A method according to claim 1, wherein the hearing evaluation test measures at least one of transient and distortion product emission levels in the ear. 9. A method according to claim 6, further comprising the step of accepting user input by the patient during said administering test. 10. A method according to claim 9, wherein said accepting step is carried out by the patient contacting at least one of a switch and a key on a keyboard. 11. A method according to claim 9, wherein said accepting step is carried out by the patient speaking into a microphone. 12. A method according to claim 1... | East Carolina University | A61B005/00
US06899683 | B2 | 2005-05-31 | Metabolic calorimeter employing respiratory gas analysis | 2002-05-31 | An indirect calorimeter for measuring the metabolic rate of a subject includes a disposable portion and a reusable portion. The disposable portion includes a respiratory connector configured to be supported in contact with the subject so as to pass inhaled and exhaled gases as the subject breathes. The disposable portion also includes a flow pathway operable to receive and pass inhaled and exhaled gases, having a first end in fluid communication with the respiratory connector and a second end in fluid communication with a source and sink for respiratory gases. The disposable portion is disposed within the reusable portion, which includes a flow meter, a component gas concentration sensor, and a computation unit. The flow meter generates a signal as a function of the instantaneous flow volume of respiratory gases passing through the flow pathway and the component gas concentration sensor generates a signal as a function of the instantaneous fraction of a predetermined component gas in the exhaled gases. The computation unit receives the electrical signals from the flow meter and the concentration sensor and calculates at least one respiratory parameter for the subject as the subject breathes through the calorimeter. | 1. An indirect calorimeter, comprising: a disposable portion having a respiratory connector configured to be supported in contact with a subject so as to pass inhaled and exhaled gases as said subject breathes; a flow pathway within said disposable portion operable to receive and pass said inhaled and exhaled gases, said flow pathway having a first end in fluid communication with said respiratory connector and a second end in fluid communication with a source and sink for respiratory gases; a reusable portion having a housing with a recess defined therein for receiving said disposable portion; a flow meter within said housing configured to generate electrical signals as a function of an instantaneous flow volume of said inhaled and exhaled gases passing through said flow pathway; a component gas concentration sensor within said housing operable to generate electrical signals as a function of an instantaneous fraction of a predetermined component gas in said exhaled gases passing through said flow pathway; and a computation unit operable within said housing to receive said electrical signals from said flow meter and said concentration sensor and to calculate at least one respiratory parameter for said subject as said subject breathes through said calorimeter. | 1. An indirect calorimeter, comprising: a disposable portion having a respiratory connector configured to be supported in contact with a subject so as to pass inhaled and exhaled gases as said subject breathes; a flow pathway within said disposable portion operable to receive and pass said inhaled and exhaled gases, said flow pathway having a first end in fluid communication with said respiratory connector and a second end in fluid communication with a source and sink for respiratory gases; a reusable portion having a housing with a recess defined therein for receiving said disposable portion; a flow meter within said housing configured to generate electrical signals as a function of an instantaneous flow volume of said inhaled and exhaled gases passing through said flow pathway; a component gas concentration sensor within said housing operable to generate electrical signals as a function of an instantaneous fraction of a predetermined component gas in said exhaled gases passing through said flow pathway; and a computation unit operable within said housing to receive said electrical signals from said flow meter and said concentration sensor and to calculate at least one respiratory parameter for said subject as said subject breathes through said calorimeter. 2. The calorimeter according to claim 1, wherein said flow pathway includes an elongated flow tube through which said inhaled and exhaled gases flow and a chamber disposed between said flow tube and said first end. 3. The calorimeter according to claim 2, wherein said disposable portion is received in said recess in a direction perpendicular to said flow tube. 4. The calorimeter according to claim 2, wherein said chamber directs said inhaled gases into said flow tube from a plurality of radial directions to induce turbulence in said inhaled gases within said flow tube. 5. The calorimeter according to claim 1, wherein said concentration sensor is an oxygen sensor. 6. The calorimeter according to claim 5, wherein said oxygen sensor is a fluorescence quench oxygen sensor. 7. The calorimeter according to claim 1, wherein said flow meter includes an upper ultrasonic transducer and a lower ultrasonic transducer in fluid communication with said inhaled and exhaled gases passing through said flow pathway. 8. The calorimeter according to claim 7, wherein at least one of said upper ultrasonic transducer and said lower ultrasonic transducer includes a plurality of microscopic transducers arranged in an array. 9. The calorimeter according to claim 8, wherein said microscopic transducers measure temperature, pressure and humidity for correcting a gas flow rate due to corresponding effects of temperature, pressure and humidity. 10. The calorimeter according to claim 1, further comprising a communication means for communicating with a remote computing device. 11. The calorimeter according to claim 10, wherein said communications means is a radio frequency transceiver. 12. The calorimeter accordi... | Healthetech, Inc. | A61B005/08
US06924150 | B1 | 2005-08-02 | Magnetic resonance imaging contrast system | 2002-04-17 | A method for improving the magnetic resonance imaging contrast of a selected portion of a sample. A selected ferrimagnetic constituent is associated with the sample portion, by molecular or biological attachment or by some other means, so that the sample portion is distinguished from other adjacent portions of the sample that do not have the selected ferrimagnetic constituent associated therewith. Alternatively, the sample portion is suspected, but not yet confirmed, to have the selected ferrimagnetic constituent associated with it. Ferromagnetic resonance imaging is performed on the sample portion, and at least one resonance frequency ω0 of the sample portion is provided. At least one material parameter of the sample portion is measured at or near the frequency ω0, with enhanced imaging contrast relative to adjacent portions of the sample that do not have the selected ferrimagnetic constituent in association. Alternatively, the measured value is compared with a corresponding reference value for the selected ferrimagnetic constituent. If (and only if) the measured and corresponding reference values substantially agree, the selected ferrimagnetic constituent is present in association with the selected sample portion. Other parameters of the selected ferrimagnetic consituent can also be measured. | 1. A method for performing magnetic resonance imaging on a selected portion of a sample, the method comprising: providing a selected ferrimagnetic material into the sample by chemically bonding the ferrimagnetic material in particle format to the sample, the ferrimagnetic material having known ferromagnetic spin resonance characteristics, at least one resonant frequency ω=ω0, and a diameter no greater than 1 μm, the ferrimagnetic material now a constituent of a selected portion of the sample; applying an excitation electromagnetic field to the sample using an electromagnetic wave generator, the excitation electromagnetic field having at least one of a static component and a time varying component with a frequency ω=ω1≈ω0, a selected magnetic field strength, and a selected magnetic field direction relative to the sample portion; applying an external gradient magnetic field to the sample, the external gradient magnetic field having at least one of a static component and a time varying component, a selected magnetic field direction, and at least first and second selected magnetic field strengths at at least first and second selected locations, respectively, within the sample portion; measuring at or near the resonant frequency ω0 of the ferrimagnetic material at least one of (i) a real component μ′(ω′) of the complex permeability parameter, and (ii) an imaginary component μ″(ω1) of the complex permeability parameter of the sample portion at each of the selected locations to determine at least one value of a material parameter at each of the selected locations; and constructing an image of the sample portion, by comparing the measured parameter values with corresponding reference values for the selected ferrimagnetic constituent for each of the selected locations. | 1. A method for performing magnetic resonance imaging on a selected portion of a sample, the method comprising: providing a selected ferrimagnetic material into the sample by chemically bonding the ferrimagnetic material in particle format to the sample, the ferrimagnetic material having known ferromagnetic spin resonance characteristics, at least one resonant frequency ω=ω0, and a diameter no greater than 1 μm, the ferrimagnetic material now a constituent of a selected portion of the sample; applying an excitation electromagnetic field to the sample using an electromagnetic wave generator, the excitation electromagnetic field having at least one of a static component and a time varying component with a frequency ω=ω1≈ω0, a selected magnetic field strength, and a selected magnetic field direction relative to the sample portion; applying an external gradient magnetic field to the sample, the external gradient magnetic field having at least one of a static component and a time varying component, a selected magnetic field direction, and at least first and second selected magnetic field strengths at at least first and second selected locations, respectively, within the sample portion; measuring at or near the resonant frequency ω0 of the ferrimagnetic material at least one of (i) a real component μ′(ω′) of the complex permeability parameter, and (ii) an imaginary component μ″(ω1) of the complex permeability parameter of the sample portion at each of the selected locations to determine at least one value of a material parameter at each of the selected locations; and constructing an image of the sample portion, by comparing the measured parameter values with corresponding reference values for the selected ferrimagnetic constituent for each of the selected locations. 2. The method of claim 1, further comprising: comparing at least one of said measured parameter values for each of said selected locations with said corresponding reference value for said selected ferrimagnetic constituent; and when said measured value and said corresponding reference value substantially agree, interpreting this condition as indicating that said selected ferrimagnetic constituent is associated with the sample portion at said selected location at which said material parameter measurement is taken. 3. The method of claim 1, further comprising: when said measured value and said corresponding reference value do not substantially agree, interpreting this condition as indicating that said selected ferrimagnetic constituent is not associated with said sample portion at said selected location at which said material parameter measurement is taken. 4. A method for performing magnetic resonance imaging on a selected portion of a sample, the method comprising: providing a selected ferrimagnetic material into the sample by chemically bonding the ferrimagnetic material in particle format to the sample, the ferrimagnetic material having known ferromagnetic spin resonance characteris... | Intematix Corporation | A61B005/055 A61B008/00 G01N025/18 G01R033/00 G01V003/00
US06908435 | B1 | 2005-06-21 | Method and monitor for enhancing angiogenesis in the heart by exercise follow-up | 2000-11-03 | Various methods and devices for treating a patient who has lost, or is at risk of losing cardiac function by cardiac ischemia are disclosed. Treatment includes first imaging a patient's heart, or a portion thereof, to identify underperfused regions of cardiac muscle, and a source of oxygenated blood that is proximate to the underperfused region. Between the underperfused regions and the oxygenated blood source, a target area is selected where thermal or mechanical injury is introduced, and optionally reintroduced, to convert initial capillary blush, resulting from the injury. | 1. A method of inducing transformation of capillary blush into arterioles in a patient after a surgical procedure, comprising the steps of: (a) providing a monitor having a sensor for measuring the patient's heart rate, and a processor; (b) equipping the patient with the monitor; (c) using the sensor to measure the patient's resting heart rate; (d) elevating the patient's heart rate to a selected threshold by having the patient undergo exercise; (e) using the sensor to detect when the patient's heart rate has reached the selected threshold; (f) determining the duration of time that the patient's heart rate was above the selected threshold during exercise; and (g) using the processor to estimate an amount of time of exercise needed to induce transformation of capillary blush into arterioles based on the selected heart rate threshold and the duration of time determined in step (f). | 1. A method of inducing transformation of capillary blush into arterioles in a patient after a surgical procedure, comprising the steps of: (a) providing a monitor having a sensor for measuring the patient's heart rate, and a processor; (b) equipping the patient with the monitor; (c) using the sensor to measure the patient's resting heart rate; (d) elevating the patient's heart rate to a selected threshold by having the patient undergo exercise; (e) using the sensor to detect when the patient's heart rate has reached the selected threshold; (f) determining the duration of time that the patient's heart rate was above the selected threshold during exercise; and (g) using the processor to estimate an amount of time of exercise needed to induce transformation of capillary blush into arterioles based on the selected heart rate threshold and the duration of time determined in step (f). 2. The method of claim 1, wherein the monitor further comprises a display device, and wherein the method further comprises the step of displaying the duration of time determined in step (f). 3. The method of claim 2, wherein the method further comprises displaying the amount of time of exercise needed to reach a desired level of capillary blush into arterioles. 4. The method of claim 1, wherein the monitor further comprises a storage device, and wherein the method further comprises storing the resting heart rate measured in step (c). 5. The method of claim 4, wherein the method further comprises storing the duration of time determined in step (f). 6. The method of claim 1, wherein the monitor is configured to alert the patient that a selected heart rate threshold has been reached. 7. The method of claim 1, wherein the storage device maintains data on the dates upon which heart-rate exercise was measured, the integrated elevated heart rate/duration value of each exercise date, and the total dates remaining in an exercise regimen. 8. The method of claim 1, wherein the monitor is configured to send an signal to the patient to initiate an exercise program. 9. The method of claim 1, wherein the monitor is adapted to be strapped to the patient. 10. The method of claim 1, wherein the sensor comprises at least one electrode. 11. The method of claim 1, wherein the selected heart rate threshold is about 50 to about 100 percent of the patient's resting heart rate. 12. The method of claim 1, wherein exercise is performed on a daily basis, for at least about one-half hour per day. 13. The method of claim 12, wherein the exercise is performed on a daily basis for about 10 weeks. 14. The method of claim 1, wherein a pacemaker is utilize to stimulate the patient's heart. | SciMed Life Systems, Inc. | A61B005/02
US06926885 | B2 | 2005-08-09 | In-vivo diagnostic method by means of near infrared radiation | 2002-06-26 | This invention relates to an in-vivo diagnostic method based on near infrared radiation (NIR radiation) that uses water-soluble dyes and their biomolecule adducts, each having specific photophysical and pharmaco-chemical properties, as a contrast medium for fluorescence and transillumination diagnostics in the NIR range, to new dyes and pharmaceuticals containing such dyes. | 1. A cyanine dye as represented by the general formula V: wherein Q is a fragment selected from the group consisting of where R30 represents a hydrogen atom, a hydroxy group, a carboxy group, an alkoxy residue containing 1 to 4 carbon atoms or a chlorine atom, b is an integer 2 or 3, R31 represents a hydrogen atom or an alkyl residue containing 1 to 4 carbon atoms, X and Y independently represent an —O—, —S—, —CH═CH— or —C(CH2R32)(CH2R32)— fragment each, R20 to R29, R32 and R33 independently represent a hydrogen atom, a hydroxy group, a carboxy-, a sulfonic acid residue or a carboxyalkyl-, alkoxycarbonyl or alkoxyoxoalkyl residue containing up to 10 C atoms or a sulfoalkyl residue containing up to 4 C atoms, or a non-selectively bonding macromolecule or a residue of the general formula VI —(O)v—(CH2)o—CO—NR34—(CH2)s—(NH—CO)q—R35 (VI) on the condition that, where X and Y are O, S, —CH═CH— or —C(CH3)2— at least one of the residues R20 to R29 corresponds to a non-selectively bonding macromolecule or a compound of the general formula VI, where o and s equal 0 or independently represent an integer between 1 and 6, q and v independently represent 0 or 1, R34 represents a hydrogen atom or a methyl residue, R35 represents an alkyl residue containing 3 to 6 C atoms and comprising 2 to n−1 hydroxy groups, with n being the number of C atoms, or an alkyl residue containing 1 to 6 C atoms that carries 1 to 3 additional carboxy groups, an aryl residue containing 6 to 9 C atoms or arylalkyl residue containing 7 to 15 C atoms, or a residue of the general formula IIId or IIIe on the condition that q is 1, or a non-selectively bonding macromolecule, or R20 and R21, R21 and R22, R22 and R23, R24 and R25, R25 and R26, and/or R26 and R27, together with the interspersed carbon atoms, form a 5- or 6-member aromatic or saturated annelled ring, and its physiologically tolerable salts. | 1. A cyanine dye as represented by the general formula V: wherein Q is a fragment selected from the group consisting of where R30 represents a hydrogen atom, a hydroxy group, a carboxy group, an alkoxy residue containing 1 to 4 carbon atoms or a chlorine atom, b is an integer 2 or 3, R31 represents a hydrogen atom or an alkyl residue containing 1 to 4 carbon atoms, X and Y independently represent an —O—, —S—, —CH═CH— or —C(CH2R32)(CH2R32)— fragment each, R20 to R29, R32 and R33 independently represent a hydrogen atom, a hydroxy group, a carboxy-, a sulfonic acid residue or a carboxyalkyl-, alkoxycarbonyl or alkoxyoxoalkyl residue containing up to 10 C atoms or a sulfoalkyl residue containing up to 4 C atoms, or a non-selectively bonding macromolecule or a residue of the general formula VI —(O)v—(CH2)o—CO—NR34—(CH2)s—(NH—CO)q—R35 (VI) on the condition that, where X and Y are O, S, —CH═CH— or —C(CH3)2— at least one of the residues R20 to R29 corresponds to a non-selectively bonding macromolecule or a compound of the general formula VI, where o and s equal 0 or independently represent an integer between 1 and 6, q and v independently represent 0 or 1, R34 represents a hydrogen atom or a methyl residue, R35 represents an alkyl residue containing 3 to 6 C atoms and comprising 2 to n−1 hydroxy groups, with n being the number of C atoms, or an alkyl residue containing 1 to 6 C atoms that carries 1 to 3 additional carboxy groups, an aryl residue containing 6 to 9 C atoms or arylalkyl residue containing 7 to 15 C atoms, or a residue of the general formula IIId or IIIe on the condition that q is 1, or a non-selectively bonding macromolecule, or R20 and R21, R21 and R22, R22 and R23, R24 and R25, R25 and R26, and/or R26 and R27, together with the interspersed carbon atoms, form a 5- or 6-member aromatic or saturated annelled ring, and its physiologically tolerable salts. 2. An agent for in-vivo diagnosis comprising the cyanine dye according to claim 1 together with physiologically acceptable adjuvants, substrates, and/or diluents. 3. A method of diagnosing tumours tissue comprising administering the cyanine dye claimed in claim 1, irradiating the tumours tissue with light in the visible to near infrared range and recording fluorescent radiation produced form the cyanine dye. | Institut fur Diagnostikforschung GmbH Freien Universitat Berlin | A61B010/00 A61B005/00 A61B008/00
US06839579 | B1 | 2005-01-04 | Temperature indicating oximetry sensor | 2001-11-02 | A pulse oximetry sensor having a temperature indicator and a method of operating such a sensor to visually indicate to an operator the temperature of the sensor and hence a measure of the temperature of the tissue location to which the sensor is attached. In one embodiment, the temperature indicator is a color changing liquid crystal temperature monitor. The liquid crystal temperature monitor includes a number of segments, each of which is activateable within a predetermined temperature range to be monitored. In other embodiments of the present invention, the oximetry sensor includes an active regulated heating element to enhance blood perfusion within the tissue location being monitored. In either embodiment, the temperature indicator provides the operator with an effective way of monitoring the temperature of the sensor and the adjacent tissue location. In case of a heated sensor, the temperature indicator informs the operator that the warming function of the sensor is functioning correctly. In case of a nonheated sensor, the temperature indicator informs the operator regarding a measure of the degree of perfusion. | 1. An oximeter sensor comprising: a body configured to be placed near a patient; a light source connected to said body to provide light to said patient; a light collector connected to said body to collect light from said patient; and a passive non-electronic temperature indicator mounted on said body to provide a visual indication of the temperature of said sensor, said temperature indicator not being in electrical communication with said sensor or with a processor of said sensor. | 1. An oximeter sensor comprising: a body configured to be placed near a patient; a light source connected to said body to provide light to said patient; a light collector connected to said body to collect light from said patient; and a passive non-electronic temperature indicator mounted on said body to provide a visual indication of the temperature of said sensor, said temperature indicator not being in electrical communication with said sensor or with a processor of said sensor. 2. The sensor of claim 1 further comprising an active heating element connected to said body. 3. The sensor of claim 1 wherein said body has an external surface, and wherein said temperature indicator is mounted on said external surface such that said temperature indicator is visible to a clinician when said sensor is in use. 4. The sensor of claim 1 wherein said body has an internal surface, and wherein said temperature indicator is mounted on said internal surface such that said temperature indicator is visible to a clinician before said sensor is attached to a patient. 5. The sensor of claim 1 wherein said temperature indicator comprises a color-changing temperature indicator. 6. The sensor of claim 1 wherein said temperature indicator comprises a liquid crystal temperature monitor. 7. The sensor of claim 1 wherein said temperature indicator comprises a number of segments of a color-changing temperature indicators, each activateable within a predetermined temperature range so as to provide a color-changing visual indication of the temperature of said sensor. 8. The sensor of claim 1 wherein said temperature indicator comprises a number of deposits of liquid crystals, each activateable within a predetermined temperature range so as to provide a color-changing visual indication of the temperature of said sensor. 9. The sensor of claim 1 wherein said temperature indicator produces a reversible color change. 10. The sensor of claim 1 wherein said temperature indicator produces an irreversible color change. 11. The sensor of claim 1 wherein said sensor is configured to be applied to one of an ear lobe and a pinna. 12. An oximeter sensor comprising: a body configured to be placed near a patient; a light source connected to said body to provide light to said patient; a light collector connected to said body to collect light from said patient; and a passive non-electronic color-changing temperature indicator mounted on said body to provide a visual indication of the temperature of said sensor, wherein said color-changing temperature indicator comprises a liquid crystal temperature monitor, said temperature indicator not being in electrical communication with said sensor or with a processor of said sensor. 13. The sensor of claim 12 further comprising an active heating element connected to said body. 14. The sensor of claim 12 wherein said color-changing temperature indicator produces a reversible color change. 15. The sensor of claim 12 wherein said sensor is... | Nellcor Puritan Bennett Incorporated | A61B005/00
US06975891 | B2 | 2005-12-13 | Raman spectroscopic system with integrating cavity | 2001-12-21 | The present invention provides an apparatus for measurement of Raman scattered radiation comprising. The apparatus comprises at least one source of electromagnetic radiation for producing an electromagnetic radiation beam characterized by a narrow spectral width, an integrating cavity having an interior and an exterior, wherein a sample is placed in said interior. The integrating cavity further having at least one port for insertion of the sample in the interior and for transmission of the electromagnetic radiation into and out from the interior, the at least one port extending from the exterior to said interior of said integrating cavity. The integrating cavity also comprises a first optical element for transmitting the electromagnetic radiation into the interior of the integrating cavity through the at least one port, and a second optical element for collecting Raman scattered electromagnetic radiation from the sample through the at least one port. The apparatus also comprises a spectrum analyzer for determining spectral composition of the Raman scattered electromagnetic radiation, a detector for measuring the Raman scattered electromagnetic radiation; and a system for determining concentration of at least one chemical compound from the measured Raman scattered electromagnetic radiation. The apparatus may also comprise a radiation expanding element. A method for measuring the concentration of one or more chemical compounds in a sample using Raman scattering is also provided. | 1. An apparatus for measurement of Raman scattered radiation comprising: a) one or more than one source of electromagnetic radiation for producing an electromagnetic radiation beam characterized by a narrow spectral width; b) an integrating cavity comprising: (i) an interior and an exterior, wherein a sample is placed in said interior, said integrating cavity having one or more than one port for insertion of said sample in said interior and for transmission of said electromagnetic radiation into and out from said interior, said one or more than one port extending from said exterior to said interior of said integrating cavity, and (ii) a radiation expanding element for expanding said electromagnetic radiation before said electromagnetic radiation beam comes into contact with said sample; c) a first optical element for transmitting said electromagnetic radiation into said interior of said integrating cavity through said one or more than one port; d) a second optical element for collecting Raman scattered electromagnetic radiation from said sample through said one or more than one port; e) a spectrum analyzer for determining spectral composition of said Raman scattered electromagnetic radiation; f) a detector for measuring said Raman scattered electromagnetic radiation; and g) a system for determining a concentration of one or more than one chemical compound in said sample from Raman scattered electromagnetic radiation measured by said detector. | 1. An apparatus for measurement of Raman scattered radiation comprising: a) one or more than one source of electromagnetic radiation for producing an electromagnetic radiation beam characterized by a narrow spectral width; b) an integrating cavity comprising: (i) an interior and an exterior, wherein a sample is placed in said interior, said integrating cavity having one or more than one port for insertion of said sample in said interior and for transmission of said electromagnetic radiation into and out from said interior, said one or more than one port extending from said exterior to said interior of said integrating cavity, and (ii) a radiation expanding element for expanding said electromagnetic radiation before said electromagnetic radiation beam comes into contact with said sample; c) a first optical element for transmitting said electromagnetic radiation into said interior of said integrating cavity through said one or more than one port; d) a second optical element for collecting Raman scattered electromagnetic radiation from said sample through said one or more than one port; e) a spectrum analyzer for determining spectral composition of said Raman scattered electromagnetic radiation; f) a detector for measuring said Raman scattered electromagnetic radiation; and g) a system for determining a concentration of one or more than one chemical compound in said sample from Raman scattered electromagnetic radiation measured by said detector. 2. The apparatus according to claim 1 wherein said source of electromagnetic radiation is selected from the group consisting of a laser, a light emitting diode (LED) and a superluminescent diode. 3. The apparatus according to claim 1, wherein said integrating cavity comprises electromagnetic radiation scattering material of sufficient thickness to back scatter a sufficient amount of electromagnetic radiation into said interior of said integrating cavity to enhance production, and analysis, of Raman scattered radiation with said spectrum analyzer and detector. 4. The apparatus according to claim 3, wherein said electromagnetic radiation scattering material comprises electromagnetic radiation scattering material selected from the group consisting of: an electromagnetic radiation scattering opalescent glass, an electromagnetic radiation scattering material comprising polytetrafluoroethylene (PTFE), and an electromagnetic radiation scattering ceramic. 5. The apparatus according to claim 1, wherein said interior is delimited by an internal surface which comprises an electromagnetic radiation scattering coating of sufficient thickness to back scatter a sufficient amount of electromagnetic radiation into said interior of said integrating cavity to enhance production and analysis of Raman scattered radiation with said spectrum analyzer and detector. 6. The apparatus according to claim 5, wherein said electromagnetic radiation scattering coating comprises an electromagnetic radiation scattering material select... | NIR Diagnostics Inc. | A61B005/00 G01J003/44 G01N021/65
US06951545 | B2 | 2005-10-04 | Integral sample collection tip | 2002-12-04 | Assemblies are provided for collecting a sample from a mouth using an integrated sample collection tip. A sample collection assembly includes a sample collection body configured for being placed within the mouth, wherein the sample collection body includes a bore and one or more pores. The sample collection assembly further includes a conduit that disposed within the bore of the sample collection body and that is in fluid communication with the one or more pores. The conduit is bonded within the bore. The sample collection assembly can further include a hand piece that has a tip on which the sample collection body is mounted, and through which the conduit can extend. The pores can include a plurality of micropores. The sample collection body can be hydrophobic, and a hydrophilic surfactant can be disposed on the outer surface of the sample collection body. | 1. A sample collection assembly for collecting a sample from a mouth, comprising: a sample collection body configured for being placed within said mouth, said sample collection body having an outer surface and an inner surface; a bore and one or more pores; and a conduit disposed within said bore in fluid communication with said one or more pores; and wherein the outer surface of said sample collection body is hydrophilic and the inner surface of said sample collection body is hydrophobic. | 1. A sample collection assembly for collecting a sample from a mouth, comprising: a sample collection body configured for being placed within said mouth, said sample collection body having an outer surface and an inner surface; a bore and one or more pores; and a conduit disposed within said bore in fluid communication with said one or more pores; and wherein the outer surface of said sample collection body is hydrophilic and the inner surface of said sample collection body is hydrophobic. 2. The sample collection assembly of claim 1, wherein said conduit is bonded within said bore. 3. The sample collection assembly of claim 1, further comprising a hand piece having a tip on which said sample collection body is mounted. 4. The sample collection assembly of claim 3, wherein said sample collection body comprises a rear surface that is bonded to a front surface of said hand piece tip. 5. The sample collection assembly of claim 3, wherein said conduit extends through said hand piece. 6. The sample collection assembly of claim 1, wherein said conduit is flexible. 7. The sample collection assembly of claim 1, wherein said one or more pores comprises a plurality of pores. 8. The sample collection assembly of claim 1, wherein said one or more pores comprises a plurality of micropores. 9. The sample collection assembly of claim 8, wherein said plurality of micropores have a pore size of 135 μm or less. 10. The sample collection assembly of claim 1, wherein said hydrophobic sample collection body comprises high density polyethylene. 11. The sample collection assembly of claim 1, wherein said sample collection body is hemi-dome shaped. 12. A sample collection assembly for collecting a sample from a mouth; comprising: a sample collection body configured for being placed within said mouth, said sample collection body having an outer surface and an inner surface; a bore and one or more pores; and a conduit bonded within said bore in fluid communication with said one or more pores, wherein an adhesive force between said conduit and said bore is greater than a cohesive force of said sample collection body; and wherein the outer surface of said sample collection body is hydrophilic and the inner surface of said sample collection body is hydrophobic. 13. The sample collection assembly of claim 12, further comprising a hand piece having a tip on which said sample collection body is mounted, wherein said sample collection body comprises a rear surface that is bonded to a front surface of said hand piece tip, and an adhesive force between said rear surface of said sample collection body and said front surface of said hand piece tip. 14. The sample collection assembly of claim 13, wherein said conduit extends through said hand piece. 15. The sample collection assembly of claim 12, wherein said conduit is flexible. 16. The sample collection assembly of claim 12, wherein said one or more pores comprises a plurality of pores. 17. The sample collection ass... | Lifepoint, Inc. | A61B005/00 B65D081/00
US06921369 | B1 | 2005-07-26 | Mobile ergospirometry system | 2000-09-21 | A mobile ergospirometry system including a portable measurement unit that can be attached to a subject and that incorporates a mouthpiece or mask to collect respiratory gas, a sensor for the volume or amount of gas as well as sensors to determine the CO2/O2 concentration in the respiratory gas, a signal processing unit, and a telemetry module. The system further includes a computer-assisted base station having a telemetry unit to create a wireless connection to the telemetry module. In the system a bidirectional transmission route to the telemetry unit is set up by the telemetry module to transmit data and commands, for which purpose a transmitter and receiver subassembly is provided in each case, so that from the base station it is possible to communicate to the subject online information or requirements for operation of the measurement unit and/or the progress of the test. The signal processing unit sends coded commands that it receives to a speech module for the conversion and audio output of instructions as to behavior, explanations, or similar information. | 1. Mobile ergospirometry system, comprising: a portable measurement unit configured to be attached to a subject and incorporating a mouthpiece or mask to collect respiratory gas;a sensor configured to sense a volume or amount of gas and at least one respiratory gas sensor configured to determine at least one of a CO2 and an O2 concentration in the respiratory gas;a signal processing unit, a telemetry module, and a computer-assisted base station having a telemetry unit configured to create a wireless connection to the telemetry module;wherein by the telemetry module a bidirectional transmission route to the telemetry unit is set up to transmit data and commands, for which purpose a transmitter and receiver subassembly is provided, so that the base station may communicate online information or requirements for operation of the measurement unit and/or to modify a course of a test, wherein the signal processing unit sends received coded commands to a speech module for conversion and audio output of information,wherein the measurement unit further comprises a heart-rate detector and a multichannel ECG module, such that detection and storage of ergospirometric and cardiological data are performed synchronously by the signal processing unit, andfurther comprising a memory configured to back up data in the measurement unit, such that, in the memory data, sets are stored in a specified format and comprise data specifically related to each consecutive breath and/or complete sequence of data measured for all or selected parameters during selectable intervals with selectable starting times. | 1. Mobile ergospirometry system, comprising: a portable measurement unit configured to be attached to a subject and incorporating a mouthpiece or mask to collect respiratory gas;a sensor configured to sense a volume or amount of gas and at least one respiratory gas sensor configured to determine at least one of a CO2 and an O2 concentration in the respiratory gas;a signal processing unit, a telemetry module, and a computer-assisted base station having a telemetry unit configured to create a wireless connection to the telemetry module;wherein by the telemetry module a bidirectional transmission route to the telemetry unit is set up to transmit data and commands, for which purpose a transmitter and receiver subassembly is provided, so that the base station may communicate online information or requirements for operation of the measurement unit and/or to modify a course of a test, wherein the signal processing unit sends received coded commands to a speech module for conversion and audio output of information,wherein the measurement unit further comprises a heart-rate detector and a multichannel ECG module, such that detection and storage of ergospirometric and cardiological data are performed synchronously by the signal processing unit, andfurther comprising a memory configured to back up data in the measurement unit, such that, in the memory data, sets are stored in a specified format and comprise data specifically related to each consecutive breath and/or complete sequence of data measured for all or selected parameters during selectable intervals with selectable starting times. 2. Mobile ergospirometry system, comprising: a portable measurement unit configured to be attached to a subject and incorporating a mouthpiece or mask to collect respiratory gas;a sensor configured to sense a volume or amount of gas and at least one respiratory gas sensor configured to determine at least one of a CO2 and an O2 concentration in the respiratory gas;a signal processing unit, a telemetry module, and a computer-assisted base station having a telemetry unit configured to create a wireless connection to the telemetry module;wherein by the telemetry module a bidirectional transmission route to the telemetry unit is set up to transmit data and commands, for which purpose a transmitter and receiver subassembly is provided, so that the base station may communicate online information or requirements for operation of the measurement unit and to modify a course of a test, wherein the signal processing unit sends received coded commands to a speech module for conversion and audio output of information,wherein the speech module comprises a digital memory in which are stored words, word groups, or other mechanisms of presenting information acoustically, andwherein by the signal processing unit, from the words or word groups stored in the digital memory command-oriented word combinations appropriate to a current situation are selected, linked, and output in a controlled m... | Cortex Biophysik GmbH | A61B005/08
US06963769 | B1 | 2005-11-08 | Method for enhancing contrast produced by MRI | 2000-04-20 | A method for obtaining an image by MRI, comprising: administering at least one contrast agent to a subject in amounts effective to perform CEDST MRI analysis; and performing CEDST MRI analysis to produce an image of the subject. A number of different contast agents can be used to practice the present method including, without limitation, sugars, animo acids, nitrogen-containing heterocycles, purines and pyrimidines, nucleosides; imidazole and derivatives thereof, imino acids, barbituric acid and analogs thereof, and miscellaneous materials, such as guanidine hydantoin, parabanic acid, and biologically active salts thereof. | 1. A method for obtaining an image by MRI, comprising: administering at least one contrast agent to a subject in amounts effective to perform chemical exchange dependent saturation transfer MRI analysis; and performing chemical exchange dependent saturation transfer MRI analysis to produce an image of the subject. | 1. A method for obtaining an image by MRI, comprising: administering at least one contrast agent to a subject in amounts effective to perform chemical exchange dependent saturation transfer MRI analysis; and performing chemical exchange dependent saturation transfer MRI analysis to produce an image of the subject. 2. The method according to claim 1 where performing chemical exchange dependent saturation transfer MRI analysis comprises: selectively irradiating and saturating an exchangeable proton or protons of the contrast agent with an applied magnetic field; and determining the effect of this saturation on water proton MR signal amplitude. 3. The method according to claim 1 where performing chemical exchange dependent saturation transfer MRI analysis comprises: selectively irradiating and saturating an exchangeable proton or protons of the contrast agent with an applied magnetic field; applying a selective irradiation with an equal but opposite chemical shift difference between the contrast agent proton exchange site and a water proton resonance frequency (ΔωcA), thereby providing an image set with the irradiation ±chemical shift difference between the contrast agent proton exchange site and the water proton resonance frequency (±ΔωcA); producing a second image set using the first image set. 4. The method according to claim 3 where producing a second image set comprises subtracting images of the first image set to produce the second image set. 5. The method according to claim 3 where producing a second image set comprises dividing images of the first image set to produce the second image set. 6. The method according to claim 1 where the contrast agent is administered as an aqueous composition. 7. The method according to claim 1 where the contrast agent is selected from the group consisting of sugars and oligomers and polymers thereof, amino acids and oligomers and polymers thereof, nitrogen-containing heterocycles, nucleosides, imidazole and derivatives thereof, imino acids and analogs thereof, barbituric acid and analogs thereof, guanidine, hydantoin, parabanic acid, biologically active salts of the contrast agents, and mixtures of contrast agents. 8. The method according to claim 7 where the contrast agent is selected from the group consisting of mannitol, sorbitol, fructose, dextrose, galactose, sucrose, maltose, lactose, and mixtures thereof. 9. The method according to claim 7 where the contrast agent is selected from the group consisting of alanine, lysine, glutamine, tryptophan, and mixtures thereof. 10. The method according to claim 1 where the contrast agent is an oligosaccharide. 11. The method according to claim 1 where the contrast agent is a polysaccharide. 12. The method according to claim 11 where the contrast agent is dextran. 13. The method according to claim 1 where the content agent is 5-hydroxy-tryptophan. 14. The method according to claim 1 where the contrast agent is selected from the group consisting of pur... | The United States of America as represented by the Secretary of the Department of Health and Human Services | A61B005/055
US06950696 | B2 | 2005-09-27 | Method and circuit for detecting cardiac rhythm abnormalities by analyzing time differences between unipolar signals from a lead with a multi-electrode tip | 2001-11-27 | A circuit and method for detecting cardiac rhythm abnormalities employ unipolar signals respectively obtained from a cardiac lead having a tip at which a number of separate electrodes are disposed, the electrodes being simultaneously in contact with cardiac tissue. The respective unipolar signals which are obtained from the multiple electrodes exhibit a time relationship relative to each other, and this time relationship is analyzed to determine whether a cardiac rhythm abnormality is present or one or more of the unipolar signals is compared to a template which is known to represent a cardiac abnormality. Analysis of the time relation is undertaken by determining the absolute value of a time offset between any two of the unipolar signals, or by correlating any two of the unipolar signals. | 1. A method for detecting cardiac rhythm abnormality, comprising the steps of: providing a plurality of electrodes at a tip of a cardiac lead, which are electrically separated from each other; placing said tip of said cardiac lead in contact with cardiac tissue so that all of said electrodes are simultaneously in substantially fixed contact with said cardiac tissue; obtaining individual unipolar electrical signals from said cardiac tissue respectively via said electrodes, said unipolar signals exhibiting a time relationship relative to each other; defining a time window; sampling one of said unipolar signals and another of said unipolar signals; correlating samples in said time window of said one of said unipolar signals with samples in said time window of said another of said unipolar signals to obtain a first correlation result; successively shifting said time window by one sample and, after each shift, again correlating the samples of said one of said unipolar signals in said time window with the samples of said another of said unipolar signals in said time window, to obtain a plurality of successive correlation results; identifying a time offset between said one of said unipolar signals and said another of said unipolar signals by a number of samples associated with a shift of said time window which produced a highest correlation result among said first and successive correlation results; and generating a signal indicating a cardiac rhythm abnormality dependent on said time offset. | 1. A method for detecting cardiac rhythm abnormality, comprising the steps of: providing a plurality of electrodes at a tip of a cardiac lead, which are electrically separated from each other; placing said tip of said cardiac lead in contact with cardiac tissue so that all of said electrodes are simultaneously in substantially fixed contact with said cardiac tissue; obtaining individual unipolar electrical signals from said cardiac tissue respectively via said electrodes, said unipolar signals exhibiting a time relationship relative to each other; defining a time window; sampling one of said unipolar signals and another of said unipolar signals; correlating samples in said time window of said one of said unipolar signals with samples in said time window of said another of said unipolar signals to obtain a first correlation result; successively shifting said time window by one sample and, after each shift, again correlating the samples of said one of said unipolar signals in said time window with the samples of said another of said unipolar signals in said time window, to obtain a plurality of successive correlation results; identifying a time offset between said one of said unipolar signals and said another of said unipolar signals by a number of samples associated with a shift of said time window which produced a highest correlation result among said first and successive correlation results; and generating a signal indicating a cardiac rhythm abnormality dependent on said time offset. 2. A method as claimed in claim 1 comprising forming individual electrodes in said plurality of electrodes as respective dot-like electrodes. 3. A method as claimed in claim 1 comprising arranging said plurality of electrodes at said one end of said cardiac lead with respective spacings between the electrodes which are substantially equal. 4. A method as claimed in claim 1 comprising disposing one of said electrodes at a center of said one end of said cardiac lead, and arranging a remainder of said electrodes symmetrically relative to said one of said electrodes. 5. A method as claimed in claim 1 comprising disposing one of said electrodes at a center of said one end of said cardiac lead, and arranging a remainder of said electrodes radially symmetrically relative to said one of said electrodes. 6. A method as claimed in claim 1 wherein the step of obtaining individual unipolar electrical signals comprises obtaining a plurality of individual unipolar electrical signals from said cardiac tissue respectively via said electrodes with one unipolar signal being obtained for each electrode, making the plurality of unipolar signals equal to the plurality of electrodes. 7. A method for detecting cardiac rhythm abnormality, comprising the steps of: providing a plurality of electrodes at a tip of a cardiac lead, which are electrically separated from each other; placing said tip of said cardiac lead in contact with cardiac tissue so that all of said electrodes are simu... | St. Jude Medical AB | A61B005/046
US06939307 | B1 | 2005-09-06 | Method and apparatus for monitoring haemodynamic function | 1998-05-13 | The present invention relates to a method and apparatus for monitoring haemodynamic function in animals and humans during anaesthesia and surgery. During anaesthesia and surgery the subject's haemodynamic, respiratory, neuromuscular and neurological functions are monitored as indicators of the condition of the health of the subject. Commonly, variations in blood pressure are used to imply corresponding variations in cardiac output, i.e. good blood pressure equals good cardiac output. The present invention utilizes a device to monitor changes of blood flow in peripheral blood vessels of the subject as an indicator of cardiac output. This is believed to provide a much more accurate indicator. | 1. A method of monitoring a haemodynamic function of one of a human and animal subject to facilitate application of an anaesthetic dose to the subject, the method comprising the step of: monitoring blood flow in one of a peripheral blood vessel and a tissue bed; generating an indication of changes in the monitored blood flow in one of peripheral blood vessel and as a cardiac output; and varying administration of the anaesthetic dose to the subject based upon the cardiac output. | 1. A method of monitoring a haemodynamic function of one of a human and animal subject to facilitate application of an anaesthetic dose to the subject, the method comprising the step of: monitoring blood flow in one of a peripheral blood vessel and a tissue bed; generating an indication of changes in the monitored blood flow in one of peripheral blood vessel and as a cardiac output; and varying administration of the anaesthetic dose to the subject based upon the cardiac output. 2. The method in accordance with claim 1 further comprising the step of monitoring a relative change in blood flow to provide an indication of a relative change in cardiac output. 3. The method in accordance with claim 1, further comprising the step of carrying out the monitoring blood flow step non-invasively. 4. The method in accordance with claim 1, further comprising the step of carrying out the monitoring blood flow step continuously. 5. The method in accordance with claim 1, further comprising the step of setting a predetermined limit for a blood flow rate, and outputting an alarm which indicates an alarm condition when the blood flow rate reaches the predetermined limit. 6. The method in accordance with claim 1, further comprising the step of pre-setting one of (1) a base reference level for the blood flow rate indicative of a blood flow level of the subject at rest prior to monitoring of the haemodynamic function and (2) an average flow level for a particular type of subject prior to monitoring of haemodynamic function. 7. The method in accordance with claim 6, further comprising the step of processing the signal to produce a display which indicates the trend of the cardiac output. 8. The method in accordance with claim 1, further comprising the step of, during monitoring the blood flow, employing a device which produces a signal which varies with variations in the blood flow rate, and processing the signal to produce an output providing an indication of variations in cardiac output. 9. The method in accordance with claim 8, further comprising the step of modifying the signal by an adjustment factor obtained by a regression analysis of the subject. 10. The method in accordance with claim 8, further comprising the step of modifying the signal by an adjustment factor obtained from a co-variate parameter. 11. The method in accordance with claim 10, further comprising the step of using a heart rate of the subject as the co-variate parameter. 12. The method in accordance with claim 1, further comprising the step of applying a Doppler effect to monitor blood flow. 13. The method in accordance with claim 1, further comprising the step of employing an infrared blood flow sensor to monitor the blood flow. 14. The method in accordance with claim 1, further comprising the step of employing an electromagnetic flow meter to monitor the blood flow. 15. The method in accordance with claim 1, further comprising the step of employing a color chart to monitor the blo... | - | A61B005/00
US06925317 | B1 | 2005-08-02 | Integrated alignment devices, system, and methods for efficient fluid extraction, substance delivery and other applications | 2000-06-12 | An alignment device and related systems and methods for aligning at least one apparatus with respect to a surface of a tissue. The alignment device comprises a tissue interface member suitable for positioning on the surface of the tissue and mating with the apparatus to maintain alignment of the apparatus during an operation of the apparatus. The alignment device is useful to align various apparatus that are part of a continuous analyte monitoring system. | 1. A system comprising an alignment device for aligning at least one apparatus with respect to a surface of a tissue, the alignment device comprising a tissue interface member suitable for positioning on the surface of the tissue and mating with the apparatus to maintain alignment of the apparatus during an operation of the apparatus, the alignment device further comprising a removable energy absorbing layer attached to the tissue interface member, wherein the energy absorbing layer is responsive to energy directed thereon to heat up and to conductively transfer heat to the surface of the tissue to ablate the tissue to cause tissue ablation, and wherein the apparatus is an energy emitter apparatus including at least one energy source for emitting energy, wherein the energy emitter apparatus includes at least one alignment member that mates with the tissue interface member to achieve alignment with the tissue, so that, when the heat is transferred to the tissue, the layer may be removed to expose the ablated tissue. | 1. A system comprising an alignment device for aligning at least one apparatus with respect to a surface of a tissue, the alignment device comprising a tissue interface member suitable for positioning on the surface of the tissue and mating with the apparatus to maintain alignment of the apparatus during an operation of the apparatus, the alignment device further comprising a removable energy absorbing layer attached to the tissue interface member, wherein the energy absorbing layer is responsive to energy directed thereon to heat up and to conductively transfer heat to the surface of the tissue to ablate the tissue to cause tissue ablation, and wherein the apparatus is an energy emitter apparatus including at least one energy source for emitting energy, wherein the energy emitter apparatus includes at least one alignment member that mates with the tissue interface member to achieve alignment with the tissue, so that, when the heat is transferred to the tissue, the layer may be removed to expose the ablated tissue. 2. The system of claim 1, wherein the energy absorbing layer further comprises a first and a second side, wherein the second side comprises adhesive material disposed thereon for adhering to the surface of the tissue. 3. The system of claim 1, wherein the alignment device has an inner periphery, wherein the energy absorbing layer includes a separation line adjacent the periphery, and wherein the energy absorbing layer is removable from the tissue interface member after tissue ablation. 4. The system of claim 1, wherein the tissue interface member mates with a first apparatus that emits energy to the energy absorbing layer to cause tissue ablation and with a second apparatus suitable for detecting a characteristic in a fluid collected from the tissue. 5. The system of claim 1, wherein the tissue interface member comprises of at least one clip that mates with a surface on the apparatus to hold the apparatus with respect to the tisáue interface member. 6. The system of claim 5, wherein the clip is biased to hold the first apparatus and second apparatus under tension. 7. The system of claim 1, wherein the tissue interface member has an exterior and an interior surface. 8. The system of claim 7, wherein interior surfaces of the tissue interface member engage with surfaces of the apparatus to align the apparatus with the tissue interface member. 9. The system of claim 7, wherein the exterior surfaces of the tissue interface member engage with surfaces of the apparatus to align the apparatus with the tissue interface member. 10. The system of claim 7, wherein the tissue interface member comprises a male alignment member or a female alignment member that mates with a complementary female alignment member or a male alignment member, respectively, on the apparatus. 11. The system of claim 1, wherein the tissue interface member comprises at least one magnetic surface portion to mate with at least one complementary magnetic surface porti... | SpectRx, Inc. | A61B005/00
US06952603 | B2 | 2005-10-04 | Subcutaneous analyte sensor | 2001-03-16 | Assembly and method for measuring the concentration of an analyte in a biological matrix. The assembly includes an implantable optical-sensing element that comprises a body, and a membrane mounted on the body in a manner such that the membrane and the body define a cavity. The membrane is permeable to the analyte, but is impermeable to background species in the biological matrix. A refractive element is positioned in the cavity. A light source transmits light of a first intensity onto the refractive element, and a light detector receives light of a second intensity that is reflected from the cavity. A controller device optically coupled to the detector compares the first and second light intensities, and relates the intensities to analyte concentration. | 1. An assembly for measuring the concentration of an analyte in a biological matrix, comprising: an implantable optical-sensing element comprising a body; a first semi-permeable membrane mounted on said body, said first semi-permeable membrane being permeable to said analyte, and impermeable to background species in said biological matrix, said first membrane and said body aligned to define a first cavity; a first refractive element disposed in said first cavity; a second membrane mounted on said body remote from said first membrane, said second membrane and said body aligned to define a second cavity; and a second refractive element disposed in said second cavity; a source for providing light into each of said first and second cavities toward said respective first and second refractive elements; a detector for receiving light from each of said first and second cavities; and a signal-processing and computing element optically coupled to said detector for relating said received light to a concentration of said analyte. | 1. An assembly for measuring the concentration of an analyte in a biological matrix, comprising: an implantable optical-sensing element comprising a body; a first semi-permeable membrane mounted on said body, said first semi-permeable membrane being permeable to said analyte, and impermeable to background species in said biological matrix, said first membrane and said body aligned to define a first cavity; a first refractive element disposed in said first cavity; a second membrane mounted on said body remote from said first membrane, said second membrane and said body aligned to define a second cavity; and a second refractive element disposed in said second cavity; a source for providing light into each of said first and second cavities toward said respective first and second refractive elements; a detector for receiving light from each of said first and second cavities; and a signal-processing and computing element optically coupled to said detector for relating said received light to a concentration of said analyte. 2. The assembly of claim 1, wherein said analyte comprises a first analyte, said first semi-permeable membrane being permeable to said first analyte and impermeable to a second analyte, and wherein said second membrane is permeable to said second analyte. 3. The assembly of claim 2, wherein said second membrane is impermeable to said first analyte. 4. The assembly of claim 3, wherein said second membrane is impermeable to said analyte. 5. The assembly of claim 1, wherein said source comprises a light transmitter for transmitting light into each of said first and second cavities. 6. The assembly of claim 5, wherein said transmitted light has a wavelength between 400 and 1300 nm. 7. The assembly of claim 1, wherein said detector comprises first and second channels, said first channel receiving light reflected from said first refractive element, and said second channel receiving light reflected from said second refractive element. 8. The assembly of claim 1, wherein said received light is convertible by signal-processing and computing element into an electronic signal. 9. The assembly of claim 8, said assembly further comprising a readout device for display of said electronic signal. 10. The assembly of claim 9, wherein said readout device comprises an analog, digital or audio readout. 11. The assembly of claim 1, wherein said body has a “”-shaped cross-section. 12. The assembly of claim 1, wherein said body has a “”-shaped cross-section. 13. An implantable optical-sensing element suitable for measuring the concentration of an analyte in a biological matrix, said optical-sensing element comprising: a body a first semi-permeable membrane mounted on said body, said first semi-permeable membrane being permeable to said analyte, and impermeable to background species in said biological matrix, said first membrane and said body aligned to define a first cavity; a first refractive element disposed in said first cavity; a second m... | Roche Diagnostics Operations, Inc. | A61B005/00 G01B009/02
US06858013 | B1 | 2005-02-22 | Method and apparatus for determining the cardiac output of a patient | 2000-02-17 | In order to determine the cardiac output of a patient, the patient's respiration cycle is determined and an indicator is injected into the patient's bloodstream over a period of at least substantially one respiration cycle. The change in the indicator value in the bloodstream downstream of the injection point is measured over a period of a number (n) of respiration cycles and the injected amount of indicator is established. The cardiac output is determined on the basis of the measured change in the indicator value, the injected amount of indicator blood and the initial value thereof. To this end a first variation in the indicator value is measured over at least substantially the period of one respiration cycle, directly prior to the injection, and the change in the indicator value caused by the injection is determined on the basis of the difference between the measured change in the indicator value measured over a period of n times that of the first variation and n times the measured first variation. | 1. A method for determining a cardiac output of a patient, wherein the patient's respiration cycle is determined and an indicator is injected into the patient's bloodstream over a period of substantially one respiration cycle and an injected amount of indicator is established, wherein a change in an indicator value in the bloodstream downstream of the injection point is integrated over a measurement period (T1) of a number (n) of respiration cycles, first variation in the indicator value is integrated over substantially a period (T2) of one respiration cycle, prior to the injection, and a second variation in the indicator value is integrated over a period (T3) of substantially one respiration cycle after the measurement period (T1), wherein an average of the first and the second integrated variations is determined, and wherein the change in the indicator value caused by the injection is determined on the basis of a difference between the change in the indicator value integrated over said measurement period (T1) and n times the average of the integrated first and second variations, wherein the cardiac output is determined on the basis of the thus determined change in the indicator value, an amount of indicator injected into the blood and an initial value thereof. | 1. A method for determining a cardiac output of a patient, wherein the patient's respiration cycle is determined and an indicator is injected into the patient's bloodstream over a period of substantially one respiration cycle and an injected amount of indicator is established, wherein a change in an indicator value in the bloodstream downstream of the injection point is integrated over a measurement period (T1) of a number (n) of respiration cycles, first variation in the indicator value is integrated over substantially a period (T2) of one respiration cycle, prior to the injection, and a second variation in the indicator value is integrated over a period (T3) of substantially one respiration cycle after the measurement period (T1), wherein an average of the first and the second integrated variations is determined, and wherein the change in the indicator value caused by the injection is determined on the basis of a difference between the change in the indicator value integrated over said measurement period (T1) and n times the average of the integrated first and second variations, wherein the cardiac output is determined on the basis of the thus determined change in the indicator value, an amount of indicator injected into the blood and an initial value thereof. 2. A method according to claim 1, wherein the first variation in the indicator value is measured directly prior to the injection. 3. A method according to claim 1, wherein the second variation in the indicator value is measured directly contiguous to the measurement period (T1). 4. A method according to claim 1, wherein an arterial blood pressure signal is measured, wherein values of the stroke volume and of the cardiac output over a period of one heartbeat are calculated over a period corresponding to the number (n) of respiration cycles, wherein the average of the calculated values is determined, and wherein a proportionality constant is computed from a comparison of the, average output value thus calculated and the cardiac output value determined on the basis of the change in the indicator value, after which the stroke volume and the cardiac output are multiplied by the computed proportionality constant. 5. A method according to claim 4, wherein the determination of the cardiac output from the change in the indicator value is repeated periodically by carrying out a new injection and computing the proportionality constant. 6. Apparatus for deters cardiac output of a patient, which apparatus comprises a processing unit having a control output for controlling injection means, a first sensor for measuring a change in an indicator value in the patient's bloodstream and a second sensor, for determining the patient's respiration cycle, wherein the processing unit is arranged for establishing an injected amount of indicator, and for integrating a change in the indicator value in the bloodstream downstream of an injection point over a measurement period of (T1) of a number (n) of respirati... | - | A61B005/02
US06869589 | B1 | 2005-03-22 | Cryptate compounds and methods for diagnosis and therapy | 2000-01-05 | The present invention relates to cryptate compounds useful as chelating agents. In particular, the present invention relates to functionalized derivatives of certain cryptate compounds. These functionalized derivatives are suitable for use in radiolabelling and similar applications. The present invention also relates to a method for diagnosis or therapy of a disease utilizing functionalized derivatives of cryptase compounds. In one embodiment, the present invention relates to a compound which is capable of being radiolabelled of general formula (I) in which n represents an integer from 2 to 4, where each R4 and R5 is independently selected from —H, CH3, COOH, NO2, CH2OH, H2PO4, HSO3, CN, C═ONH2 and CHO; X and Y are the same or different and are selected from the group of C—R, N, P and C—Z in which R represents a hydrogen or halogen atom or a hydroxyl, nitro, nitroso, amino, optionally substituted alkyl, optionally substituted aryl, optionally substituted aralkyl or cyano group, or a group of the formula —COOR′, COCOOR′, NH—COCH2Br, —NH—CO—CH═CH—COOR′ in which R′ is a hydrogen atom or alkyl group; or, W is selected from the group of NH, S and O; and Z is a functionalized linkage group which is capable of binding said compound of formula (I) to a molecular recognition unit and wherein at least one of X and Y is C—Z; or a pharmaceutically acceptable salt thereof | 1. A compound which is capable of being radiolabelled of general Formula (I): in which n represents an integer from 2 to 4, where each R4 and R5 is independently selected from —H, CH3, COOH, NO2, CH2OH, H2PO4, HSO3, CN, C(═O)NH2 and CHO; X and Y are the same or different and are selected from the group consisting of C—R, N, P and C—Z in which R is selected from hydrogen, halogen, hydroxyl, nitro, nitroso, amino, optionally substituted alkyl, optionally substituted aryl, optionally substituted aralkyl, cyano, —COOR′, COCOOR′, NH—COCH2Br, —NH—CO—CH═CH—COOR′ in which R′ is a hydrogen atom or alkyl group, wherein at least one of X and Y is C—Z; W is selected from the group of NH, S and O; and Z is a functionalised vinyl pyridyl group which is capable of binding said compound of formula (I) to a molecular recognition unit, selected from or a pharmaceutically acceptable salt thereof. | 1. A compound which is capable of being radiolabelled of general Formula (I): in which n represents an integer from 2 to 4, where each R4 and R5 is independently selected from —H, CH3, COOH, NO2, CH2OH, H2PO4, HSO3, CN, C(═O)NH2 and CHO; X and Y are the same or different and are selected from the group consisting of C—R, N, P and C—Z in which R is selected from hydrogen, halogen, hydroxyl, nitro, nitroso, amino, optionally substituted alkyl, optionally substituted aryl, optionally substituted aralkyl, cyano, —COOR′, COCOOR′, NH—COCH2Br, —NH—CO—CH═CH—COOR′ in which R′ is a hydrogen atom or alkyl group, wherein at least one of X and Y is C—Z; W is selected from the group of NH, S and O; and Z is a functionalised vinyl pyridyl group which is capable of binding said compound of formula (I) to a molecular recognition unit, selected from or a pharmaceutically acceptable salt thereof. 2. A compound having the following structure: wherein said compound is capable of binding to a molecular recognition unit. 3. A compound according to claim 2 which is complexed with a metal ion. 4. A compound according to claim 2 which is complexed with a metal ion selected from the group consisting of Cu, Tc, Gd, Ga, In, Co, Re, Fe, Au, Ag, Rh, Pt, Bi, Cr, W, Ni, V, Ir, Pt, Zn, Cd, Mn, Ru, Pd, Hg, and Ti. 5. A compound according to claim 4, wherein the metal ion is a radionuclide selected from the group consisting of 64Cu, 67Cu, Tc, In, Gd, Ga, Fe, Co, Ti and Re. 6. A compound according to claim 5, wherein the radionuclide is selected from 64Cu and 67Cu. 7. A method radioimaging comprising administering to a subject an effective amount of a radiolabelled metal ion complex of a compound, wherein said compound has the structure or a pharmaceutically acceptable salt thereof. 8. The method of claim 7, wherein said metal ion is selected from the group consisting of Cu, Tc, Gd, Ga, In, Y Co, Re, Fe, Au, Ag, Rh, Pt, Bi, Cr, W, Ni, V, Ir, Pt, Zn, Cd, Mn, Ru, Pd, Hg, and Ti. 9. The method of claim 8, wherein the metal ion is a radionuclide selected from the group consisting of 64Cu, 67Cu, Tc, In, Gd, Ga, Fe, Co, Ti, and Re. 10. The method of claim 9, wherein the radionuclide is selected from 64Cu and 67Cu. 11. A compound according to claim 1, wherein the molecular recognition unit is selected from the group consisting of an antibody, protein, peptide, carbohydrate, nucleic acid, oligonucleotide, oligosaccharide and liposome. 12. A compound according to claim 1, wherein W is NH. 13. A compound according to claim 1, wherein said compound is complexed with a metal ion selected from the group consisting of Cu, Tc, Ga, In, Co, Re, Fe, Au, Ag, Rh, Pt, Bi, Cr, W, Ni, V, Pb, Ir, Zn, Cd, Mn, Ru, Pd, Hg, Ti, Y and Sc. 14. A compound according to claim 1, wherein n is 3 or 4 and the compound is complexed with a metal ion is selected from the lanthanide group of elements in the Periodic Table. 15. A compound according to claim 13, wherein the metal ion is selected from th... | Australian Nuclear Science & Technology Organization The Australian National University | A61K051/00 A61B005/005 C07D267/22 C07D225/00
US06918874 | B1 | 2005-07-19 | Attribute compensation for analyte detection and/or continuous monitoring | 1999-09-10 | A system and method for detecting a measuring an analyte in a biological fluid of an animal. A harvesting device (10) is provided suitable for positioning on the surface of tissue of an animal to harvest biological fluid therefrom. The harvesting device (10) comprises an analyte sensor (50) positioned to be contacted by the harvested biological fluid and which generates a measurement signal representative of the analyte. At least one attribute sensor (40) is provided to measure an attribute associated with the biological fluid harvesting operation of the harvesting device (10) or the assay of the biological fluid, and which generates an attribute signal representative of the attribute. Adjustments are made to operational parameters of the harvesting device (10) based on the one or more attributes. | 1. A system for detecting and measuring an analyte in a biological fluid of a animal, comprising: a harvesting device suitable for positioning on the surface of tissue of an animal to harvest biological fluid therefrom, and comprising an analyte sensor positioned to be contacted by the harvested biological fluid and which generates a measurement signal representative of the analyte; at least one attribute sensor to measure an attribute associated with the operation of the harvesting device and which generates an attribute signal representative of the attribute, wherein the attribute sensor detects a condition of the tissue indicative of fluid productivity; and a processor coupled to the attribute sensor and the analyte sensor to receive the attribute signal and the measurement signal, wherein the processor adjusts for an operational parameter of the harvesting device based on attribute signal, and wherein the processor generates a signal to control an amount of suction applied to the harvesting device based on the attribute signal. | 1. A system for detecting and measuring an analyte in a biological fluid of a animal, comprising: a harvesting device suitable for positioning on the surface of tissue of an animal to harvest biological fluid therefrom, and comprising an analyte sensor positioned to be contacted by the harvested biological fluid and which generates a measurement signal representative of the analyte; at least one attribute sensor to measure an attribute associated with the operation of the harvesting device and which generates an attribute signal representative of the attribute, wherein the attribute sensor detects a condition of the tissue indicative of fluid productivity; and a processor coupled to the attribute sensor and the analyte sensor to receive the attribute signal and the measurement signal, wherein the processor adjusts for an operational parameter of the harvesting device based on attribute signal, and wherein the processor generates a signal to control an amount of suction applied to the harvesting device based on the attribute signal. 2. A method for detecting and measuring an analyte in a biological fluid of a subject, comprising steps of: harvesting biological fluid from the surface of tissue of an animal with a harvesting device; contacting an analyte sensor with the biological fluid on the tissue surface; detecting an analyte in the biological fluid with the analyte sensor; sensing an attribute associated with the operation of the harvesting device, wherein the step of sensing an attribute comprises sensing a condition of the issue indicative of fluid productivity; and adjusting an operational parameter of the harvesting device based on attribute, wherein the step of adjusting comprises adjusting a level of suction applied to the harvesting device for drawing fluid from the tissue into contact with analyte sensor. 3. A method for detecting and measuring an analyte in a biological fluid of a subject, comprising steps of: harvesting biological fluid from the surface of tissue of an animal with a harvesting device, wherein the step of harvesting comprises continually harvesting biological fluid from the surface of tissue; contacting an analyte sensor with the biological fluid on the tissue surface; continually detecting an analyte in the biological fluid with the analyte sensor, wherein the step of detecting an analyte in a biological fluid of a subject comprises continually detecting the analyte; sensing an attribute associated with the operation of the harvesting device, wherein the step of sensing an attribute comprises continually sensing an attribute proximate to the analyte sensor; and adjusting an operational parameter of the harvesting device based on attribute, wherein the step of adjusting comprises continually adjusting an operational parameter of the device. | SpectRx, Inc. Altea Technologies, Inc. | A61B005/00
US07133716 | B2 | 2006-11-07 | Determining the hydration status of a patient | 2001-12-12 | A device and method are provided for determining the volume ECVhydr(t) of a body compartment of a patient at a time t by conducting measurements at the time t of the patient to determine at least one anthropometric measure X(t), the extracellular water volume ECV(t), and the intracellular water volume ICV (t) of the patient. The extracellular water volume ECVbasic(t) of a first compartment with weight Wbasic(t) of the patient at the time t is derived by using X (t), the extracellular water volume ECVsec(t) of a second compartment of the patient at the time t is derived by using ICV (t), and ECVhydr(t) as the extracellular water volume of a third compartment of the patient is derived with weight Whydr(t). The extracellular volume ECVhydr(t) is a measure for the hydration status of the patient. | 1. A method for determining the volume of a body compartment ECVhydrhydr(t) of a patient at a time t in order to manage a hydration level of said patient comprising the steps of: determining at least one anthropometric measure X(t) of the patient at the time t;determining the extracellular water volume ECV(t) of the patient at the time;determining the intracellular water volume ICV(t) of the patient at the time t;deriving the extracellular water volume ECVbasic(t) of a first compartment with weight Wbasic(t) of the patient at the time t by using X(t);deriving the extracellular water volume ECVsec(t) of a second compartment with weight Wsec(t) of the patient at the time t by using ICV(t);deriving the volume ECVhydrhydr(t) as the extracellular water volume of a third compartment of the patient with weight Whydr(t) by using the equation ECVhydr(t)=ECV(t)−ECVbasic(t)−ECVsec(t); andmanaging the hydration level of said patient using said derived volume ECVhydr(t). | 1. A method for determining the volume of a body compartment ECVhydrhydr(t) of a patient at a time t in order to manage a hydration level of said patient comprising the steps of: determining at least one anthropometric measure X(t) of the patient at the time t;determining the extracellular water volume ECV(t) of the patient at the time;determining the intracellular water volume ICV(t) of the patient at the time t;deriving the extracellular water volume ECVbasic(t) of a first compartment with weight Wbasic(t) of the patient at the time t by using X(t);deriving the extracellular water volume ECVsec(t) of a second compartment with weight Wsec(t) of the patient at the time t by using ICV(t);deriving the volume ECVhydrhydr(t) as the extracellular water volume of a third compartment of the patient with weight Whydr(t) by using the equation ECVhydr(t)=ECV(t)−ECVbasic(t)−ECVsec(t); andmanaging the hydration level of said patient using said derived volume ECVhydr(t). 2. The method according to claim 1 wherein the second compartment represents the muscle tissue with weight Wmuscle(t) and the extracellular water volume ECVmuscle(t). 3. The method according to claim 1 wherein the intracellular water volume ICVbasic(t) of the first compartment is also derived from X(t) and that ICVbasic(t) is also used to derive ECVsec(t). 4. The method according to claim 1 wherein the weight W(t) of the patient at time t is determined and that the extracellular water volume ECVbasic(t) and/or the intracellular water volume ICVbasic(t) of the first compartment are derived by using X(t) and W(t). 5. The method according to claim 1 wherein at the extracellular and/or intracellular water volumes ECVbasic(t) and ICVbasic(t) of the first compartment are derived by also using ICV(t). 6. The method according to claim 1 wherein the at least one anthropometric measure X(t) is the height H(t) of the patient at the time t. 7. The method according to claim 1 wherein ECV(t) and ICV(t) are derived from a bioimpedance measurement. 8. The method according to claim 1 wherein ECVbasic(t) is derived from various subcompartments by the following equation:ECVbasic⁡(t)=∑i⁢⁢ECVbasic,i⁡(t),where ECVbasic,i(t) is the ECV of the ith subcompartment of the first compartment at the time t. 9. The method according to claim 1 wherein at least one of the i ECVbasic,i(t) values is derived as proportional part of the weight Wbasic,i(t) of the ith subcompartment, i.e. ECVbasic,i(t)=λECV,iWbasic,i(t),where λECV,i is the corresponding proportionality constant. 10. The method according to claim 1 wherein ICVbasic(t) is derived from various subcompartments by the following equation:ICVbasic⁡(t)=∑i⁢⁢ICVbasic,i⁡(t),where ICVbasic,i(t) is the ICV of the ith subcompartment of the first compartment at the time t. 11. The method according to claim 10 wherein at least one of the i ICVbasic,i(t) values is derived as proportional part of the weight Wbaslc,i(t) of the ith subcompartment, i.e. ICVbasic,i(t)=λICV,iW... | Fresenius Medical Care Deutschland GmbH | A61B005/05
US06997883 | B1 | 2006-02-14 | Diagnostic apparatus for determining the remineralization ability of tissue | 1998-04-07 | The invention relates to a diagnostic apparatus with which diseased hard tissue, in particular carious dental tissue, can be examined through a small working channel, to see whether it can be made healthy again by remineralization. This diagnosis apparatus tightly (54) seals the working channel, produced in the tissue, from its environment, and places the cavity in the tissue under pressure, after which the pressurization is ended and a pressure drop is measured by means of a pressure measuring instrument (60). The speed of the drop in pressure is a measure of the extent of the changes to the diseased tissue. The invention also relates to a treatment device for supporting the remineralization or reinfiltration of diseased hard tissue and for preparing this, as well as to diagnosis and treatment agents and treatment substances. | 1. Apparatus for determining the remineralization ability of a hard tissue with a measuring device (46, 60, 86, 88; 104, 114, 116, 86, 88; 130–140, 88; 144–148, 88; 150–154, 88; 156, 86, 88; 158–174) for local measurement of a physical property of the hard tissue, characterized in thatthe measuring device has a measuring head (40) with a rod-like probe section (42), which probe section (42) is introducible into a working channel (32) of the hard tissue and comprises a measuring element (46, 60; 104, 116; 130, 134, 138; 144, 148; 150–154) being connected to the probe section (42) and responding to the physical property of porosity of the hard tissue, the measuring head (40) has a sealing element (54) which cooperates with a section of the tissue surface, and the measuring head (40) is connected to a fluid source (76; 80) which is under a pressure different from normal pressure, and the measuring device measures the fluid leakage through the tissue to be investigated. | 1. Apparatus for determining the remineralization ability of a hard tissue with a measuring device (46, 60, 86, 88; 104, 114, 116, 86, 88; 130–140, 88; 144–148, 88; 150–154, 88; 156, 86, 88; 158–174) for local measurement of a physical property of the hard tissue, characterized in thatthe measuring device has a measuring head (40) with a rod-like probe section (42), which probe section (42) is introducible into a working channel (32) of the hard tissue and comprises a measuring element (46, 60; 104, 116; 130, 134, 138; 144, 148; 150–154) being connected to the probe section (42) and responding to the physical property of porosity of the hard tissue, the measuring head (40) has a sealing element (54) which cooperates with a section of the tissue surface, and the measuring head (40) is connected to a fluid source (76; 80) which is under a pressure different from normal pressure, and the measuring device measures the fluid leakage through the tissue to be investigated. 2. Apparatus according to claim 1, characterized by a position indicator (48–52) for the axial position of the probe section (42) in relation to a contour surface of the tissue. 3. Apparatus according to claim 2, characterized in that the probe position indicator (48–52) has an input part (48) which cooperates with the tissue surface and which is arranged axially displaceably on the probe section (42) and is supported by a spring (50) on a stationary apparatus section, which spring (50) acts on a force sensor or deformation sensor (52;1). 4. Apparatus according to claim 3, characterized by an evaluation unit (88) which receives the output signals of the probe position indicator (48–52) and the output signals of the measuring element, and compares the output signals from the latter, which are obtained when the positions of the probe section (42) differs with one another, for which purpose the signals obtained for each of the different probe positions are stored as a function of the output signals from the probe position indicator (48–52). 5. Apparatus according to claim 1, where the measuring element has a radial measuring axis, relative to the axis of the probe section (42), characterized in that the probe section (42), is arranged rotatably (118–122) on the measuring head (40). 6. Apparatus according to claim 1, characterized in that the measuring head (40) is attached to a grip part (36) having a long axis. 7. Apparatus according to claim 6, characterized in that the long axis of the grip part (36) and the axis of the measuring head (40) are inclined with respect to one another. 8. Apparatus according to claim 7, characterized in that the long axis of the grip part (36) and axis of the measuring head (40) are inclined at an angle between about 30° and about 90°. 9. Apparatus according to claim 7, characterized in that the long axis of the grip part (36) and the axis of the measuring head (40) are inclined at an angle of about 60°. 10. Apparatus according to claim 6, charact... | - | A61B005/00 A61C019/04
US07010435 | B2 | 2006-03-07 | Method for measuring tension state of a material and uses thereof | 2002-10-16 | A method for determining the state of tension of a material at a given time, including the step of calculating the value of at least one parameter extracted from the ultrasonic signal received after propagation of an ultrasonic wave emitted in the material between an ultrasound source and at least an ultrasound emitter located at a fixed distance from the source, the value of the at least one parameter varying with the state of tension of the material, the variation of the at least one parameter being proportional or inversely proportional to the variation of the state of tension of the material. | 1. A method for determining the tension condition of a material at a given time, comprising the steps of: (a) propagating an ultrasonic wave transmitted in a material between an ultrasound source and at least one ultrasound receiver situated at a fixed predetermined distance from the ultrasound source, wherein the material is under a tension condition; and(b) after propagation of the ultrasonic wave, extracting at least one parameter from the ultrasonic wave and calculating its value, wherein the value of the at least one parameter varies with the tension condition of the material, said variation of the at least one parameter being directly proportional or inversely proportional to the variation in the tension condition of said material. | 1. A method for determining the tension condition of a material at a given time, comprising the steps of: (a) propagating an ultrasonic wave transmitted in a material between an ultrasound source and at least one ultrasound receiver situated at a fixed predetermined distance from the ultrasound source, wherein the material is under a tension condition; and(b) after propagation of the ultrasonic wave, extracting at least one parameter from the ultrasonic wave and calculating its value, wherein the value of the at least one parameter varies with the tension condition of the material, said variation of the at least one parameter being directly proportional or inversely proportional to the variation in the tension condition of said material. 2. The method according to claim 1, wherein the at least one parameter calculated on the basis of the ultrasonic signal is selected from among the following parameters: (i) the travel time of the ultrasonic wave over a predetermined distance in said material;(ii) the amplitude of the ultrasonic wave at the reception point of the ultrasonic signal;(iii) the attenuation of the ultrasonic wave at the reception point of the ultrasonic signal;(iv) the average frequency of the ultrasonic wave at the reception point of the ultrasonic signal;(v) the maximum frequency of the ultrasonic wave at the reception point of the sound signal. 3. The method according to claim 1, wherein in the step of extracting and calculating, the values are calculated for at least one of the following parameter combinations: (i) the travel time and the amplitude of the ultrasonic wave;(ii) the travel time and the attenuation of the ultrasonic wave;(iii) the amplitude and the attenuation of the ultrasonic wave;(iv) the travel time, the amplitude and the attenuation of the ultrasonic wave. 4. The method according to claim 1, wherein the material to be tested is a biological material selected from among a tendon, a ligament and a muscle. 5. The method according to claim 1, wherein the frequency of ultrasounds transmitted by the source ranges between 15 KHz and 10 GHz. 6. A method for continuous measurement of the tension condition of a material, comprising implementing the method according to claim 1 for a plurality of time-spread, given instants. 7. A method for selection of a corrective iron fitting intended to modify the tension of a tendon in equines comprising the following steps: a) implementing the method according to claim 6 on said tendon, after fitting the equines with a corrective iron fitting, or successively with a plurality of corrective iron fittings;b) selecting the corrective iron fitting in order to modify the tension condition of the tendon. 8. A method for continuous measurement of the tension condition of a material, comprising implementing the method according to claim 1 for a plurality of time-spread, given instants at a sampling frequency greater than or equal to 2.5 times that of the frequency at which the tension co... | Institut National de la Recherche Agronomique (INRA) | G01L001/00 G01G011/00 A61B005/10
US07201891 | B1 | 2007-04-10 | Pharmaceutical preparation for the diagnosis of helicobacter pylori infection | 2000-08-30 | A coated preparation for the detection of a H. pylori infection according to a urea breath test protocol which comprises a core composition containing at least an isotope C-labeled urea, an excipient and a lubricant in defined proportions and covered with 0.1˜10 weight % of a coating agent based on 100 weight % of the core composition. With this preparation, the influence of the urease-producing bacteria resident in organs other than the stomach, such as the mouth and throat, is excluded to enable a diagnosis of H. pylori infection without the risk for a false-positive test and with reasonable rapidity. | 1. A coated preparation product for use in the detection of Helicobacter pylori infection comprising: (i) a core composition comprising:about 19 to about 89 parts by weight of isotope carbon-labeled urea relative to 100 parts by weight of the core composition,about 10 to about 80 parts by weight of an excipient component relative to 100 parts by weight of the core composition, andabout 0.01 to about 1 parts by weight of a lubricant component relative to 100 parts by weight of the core composition;the excipient component comprising:(a) at least one member selected from the group consisting of lactose, sucrose, and glucose,(b) at least one member selected from the group consisting of crystalline cellulose, low-substitution hydroxypropyl cellulose, carboxymethylcellulose calcium, and croscarmellose sodium, and,(c) at least one member selected from the group consisting of starch, carboxymethylstarch sodium, hydroxypropylstarch, and partially pregelatinized starch; and(ii) a coating agent;the core composition being coated with about 0.1 to about 5 parts by weight of the coating agent relative to 100 parts by weight of the core composition. | 1. A coated preparation product for use in the detection of Helicobacter pylori infection comprising: (i) a core composition comprising:about 19 to about 89 parts by weight of isotope carbon-labeled urea relative to 100 parts by weight of the core composition,about 10 to about 80 parts by weight of an excipient component relative to 100 parts by weight of the core composition, andabout 0.01 to about 1 parts by weight of a lubricant component relative to 100 parts by weight of the core composition;the excipient component comprising:(a) at least one member selected from the group consisting of lactose, sucrose, and glucose,(b) at least one member selected from the group consisting of crystalline cellulose, low-substitution hydroxypropyl cellulose, carboxymethylcellulose calcium, and croscarmellose sodium, and,(c) at least one member selected from the group consisting of starch, carboxymethylstarch sodium, hydroxypropylstarch, and partially pregelatinized starch; and(ii) a coating agent;the core composition being coated with about 0.1 to about 5 parts by weight of the coating agent relative to 100 parts by weight of the core composition. 2. A coated preparation according to claim 1, wherein the amount of the coating agent ranges from about 0.3 to about 5 parts by weight relative to 100 parts by weight of the core composition. 3. A coated preparation according to claim 1, wherein the amount of the coating agent ranges from about 0.3 to about 3 parts by weight relative to 100 parts by weight of the core composition. 4. A coated preparation according to claim 1 containing: from about 25 to about 75 parts by weight of the isotope carbon-labeled urea relative to 100 parts by weight of the core composition,from about 20 to about 70 parts by weight of the excipient component relative to 100 parts by weight of the core composition, andfrom about 0.05 to about 0.8 parts by weight of the lubricant component relative to 100 parts by weight of the core composition. 5. A coated preparation according to claim 1 containing: from about 30 to about 70 parts by weight of the isotope carbon-labeled urea relative to 100 parts by weight of the core composition,from about 35 to about 65 parts by weight of the excipient component relative to 100 parts by weight of the core composition, andfrom about 0.1 to about 0.7 parts by weight of the lubricant component relative to 100 parts by weight of the core composition. 6. A coated preparation according to claim 1, wherein the core composition contains: from about 10 to about 450 parts by weight of the excipient component relative to 100 parts by weight of the isotope carbon-labeled urea, andfrom about 0.01 to about 6 parts by weight of the lubricant component relative to 100 parts by weight of the isotope carbon-labeled urea. 7. A coated preparation according to claim 1, wherein the core composition contains: from about 50 to about 150 parts by weight of the excipient component relative to 100 parts by weight of the isot... | Otsuka Pharmaceutical Co., Ltd. | A61B005/055
US07016721 | B2 | 2006-03-21 | Medical device ECG marker for use in compressed data stream | 2001-10-10 | Triggers and noise should be available as information in recorded electrograms in memories of implantable medical devices. Particularly where the recording of electrogram data is done in the far field, there will be considerable noise and the interpretation of ECG's reproduced from such recorded data will benefit from the storing of information regarding contemporaneous noise. By storing contemporaneous trigger data and noise data directly in the ECG data, recordings of the ECG data become more useful for physician use when played back through an external display system with minimal loss of ECG data, since out of range values are employed for the noise and trigger information and this non-ECG data is limited in size to no longer than individual point values of the ECG signal. | 1. A method comprising: monitoring an electrogram signal to find a trigger event in the electrogram signal, wherein said electrogram signal is obtained via at least a pair of subcutaneous electrodes spaced from a heart; generating a trigger indicator signal upon finding the trigger event; recording a plurality of sampled values of the electrogram signal in a memory structure; and recording the trigger indicator signal in the memory structure contemporaneously with the recording of the sampled values of the electrogram signal, wherein recording the trigger indicator signal in the memory structure comprises replacing one of the sampled values with the trigger indicator signal and wherein the trigger event comprises detection of out-of-range values of the electrogram signal. | 1. A method comprising: monitoring an electrogram signal to find a trigger event in the electrogram signal, wherein said electrogram signal is obtained via at least a pair of subcutaneous electrodes spaced from a heart; generating a trigger indicator signal upon finding the trigger event; recording a plurality of sampled values of the electrogram signal in a memory structure; and recording the trigger indicator signal in the memory structure contemporaneously with the recording of the sampled values of the electrogram signal, wherein recording the trigger indicator signal in the memory structure comprises replacing one of the sampled values with the trigger indicator signal and wherein the trigger event comprises detection of out-of-range values of the electrogram signal. 2. A method according to claim 1, wherein the trigger indicator signal is a first trigger indicator signal corresponding to a first trigger event, the method further comprising: generating a second trigger indicator signal upon finding the second trigger event in the electrogram signal; and recording the second trigger indicator signal in the memory structure contemporaneously with the recording of the sampled values of the electrogram signal. 3. A method according to claim 1, further comprising compressing the plurality of sampled values prior to recording the plurality of sampled values in the memory structure. 4. A method according to claim 1, further comprising: transmitting the recorded sampled values of the electrogram signal and the recorded trigger indicator signal to an external device; parsing the trigger indicator signal from the recorded sampled values of the electrogram signal; and generating a display as a function of the recorded sampled values. 5. A method according to claim 4, wherein generating the display as a function of the recorded sampled values comprises displaying an electrogram waveform. 6. A method according to claim 4, further comprising generating the display as a function of the trigger indicator signal. 7. A method according to claim 6, wherein generating the display as a function of the trigger indicator signal comprises generating a display icon as a function of the trigger indicator signal. 8. A method according to claim 6, further comprising: monitoring the electrogram signal to find noise in the electrogram signal; generating a noise indicator signal upon finding the noise, the noise indicator signal being a function of the noise found in the electrogram signal; and recording the noise indicator signal in the memory structure contemporaneously with the recording of the sampled values of the electrogram signal, wherein recording the noise indicator signal in the memory structure comprises replacing another of the sampled values with the noise indicator signal. 9. A method according to claim 8, further comprising: transmitting the recorded noise indicator signal to an external device; parsing said noise indicator signal from the reco... | Medtronic, Inc. | A61B005/0432
US07184826 | B2 | 2007-02-27 | Device and method for enhancing transdermal flux of agents being delivered or sampled | 1997-06-17 | A percutaneous agent delivery or sampling device comprising a sheet having a plurality of microblades for piercing and anchoring to the skin for increasing transdermal flux of an agent and for improving the attachment of the device to the skin. | 1. A device for piercing the stratum corneum of a body surface to form pathways through which an agent can be introduced or withdrawn, comprising a sheet having a least one opening therethrough and a plurality of blades extending downward therefrom, an adhesive anchor applied to at least one surface of said sheet, wherein said adhesive anchor helps prevent said sheet from being dislodged from said body surface, and an agent delivery or sampling device connected to said sheet and positioned to deliver or sample an agent through said opening, said agent delivery or sampling device being selected from the group consisting of an electrotransport device, a passive diffusion device, an osmotic device, and a pressure driven device, and wherein a plurality of blades has a substantially identical and uniform configuration. | 1. A device for piercing the stratum corneum of a body surface to form pathways through which an agent can be introduced or withdrawn, comprising a sheet having a least one opening therethrough and a plurality of blades extending downward therefrom, an adhesive anchor applied to at least one surface of said sheet, wherein said adhesive anchor helps prevent said sheet from being dislodged from said body surface, and an agent delivery or sampling device connected to said sheet and positioned to deliver or sample an agent through said opening, said agent delivery or sampling device being selected from the group consisting of an electrotransport device, a passive diffusion device, an osmotic device, and a pressure driven device, and wherein a plurality of blades has a substantially identical and uniform configuration. 2. The device of claim 1, wherein said agent comprises a polypeptide or protein. 3. The device of claim 1, further comprising at least one additional anchoring means selected from the group consisting of a projection extending out from at least one blade of said plurality of blades, a barb, at least one opening extending through said plurality of blades, each one of the plurality of blades defines essentially a plane and wherein said additional anchoring means comprises a portion of said plurality of blades being oriented at an angle of about 90° with respect to a remaining portion of said plurality of blades, and wherein each one of said plurality of blades defines essentially a plane and wherein said additional anchoring means comprises a portion of said plurality of blades being oriented at an angle within a range of about 10° to about 89° with respect to a remaining portion of said plurality of blades. 4. The device of claim 3, wherein said projection extends out from a plane defined by at least one blade. 5. The device of claim 1, wherein said anchor is integral with an edge of said plurality of blades and in a plane defined by said plurality of blades. 6. The device of claim 1, wherein a portion of said blades are located along a periphery of an opening through said sheet. 7. The device of claim 1, wherein a portion of said blades are located along peripheries of a plurality of openings through said sheet. 8. The device of claim 1, further comprising a plurality of second openings through said sheet. 9. The device of claim 1, wherein said device has a blade density of about 600 to about 1000 blades/cm2. 10. The device of claim 1, wherein said device has a blade density of at least about 800 blades/cm2. 11. The device of claim 1, wherein at least a portion of said blades has a length sufficient to pierce the stratum corneum of said body surface to a depth of at least about 25 μm. 12. The device of claim 1, wherein said blades are oriented approximately perpendicular to said sheet. 13. The device of claim 1, wherein said blades are oriented at an angle in a range of about 10° to about 89° to said sheet. 14. The device o... | Alza Corporation | A61M001/30 A61M031/00 A61B005/00
US06994841 | B1 | 2006-02-07 | Relating to magnetic resonance imaging | 1998-11-09 | The invention shows a method for obtaining temporally spaced images of tissues, including blood vessels, to reveal blood flow abnormalities within those tissues and vessels. | 1. A method of detecting blood flow abnormality or variation in a vessel or tissue comprising: administering a contrast enhancing amount of a paramagnetic metal containing magnetic resonance magnetic resonance contrast agent into a vessel of a body;imaging at least a portion of the body through which the MR contrast agent passes, with a MR imaging technique, thereby collecting temporally spaced sets of 3-D and 2-D data, each data set collected successively through an acquisition time;forming a time sequence of image data including early image data and later image data;comparing 3-D and 2-D data from the temporally spaced sets of data by evaluating 2-D and 3-D temporally acquired images by comparing ones of said early image data within said acquisition time with ones of said later image data within said acquisition time and their intensity to assess blood flow or angiographic abnormality or variation. | 1. A method of detecting blood flow abnormality or variation in a vessel or tissue comprising: administering a contrast enhancing amount of a paramagnetic metal containing magnetic resonance magnetic resonance contrast agent into a vessel of a body;imaging at least a portion of the body through which the MR contrast agent passes, with a MR imaging technique, thereby collecting temporally spaced sets of 3-D and 2-D data, each data set collected successively through an acquisition time;forming a time sequence of image data including early image data and later image data;comparing 3-D and 2-D data from the temporally spaced sets of data by evaluating 2-D and 3-D temporally acquired images by comparing ones of said early image data within said acquisition time with ones of said later image data within said acquisition time and their intensity to assess blood flow or angiographic abnormality or variation. 2. A method of detecting blood flow abnormality or variation, in a human body, said method comprising the steps of: administering into vasculature of said human body a contrast enhancing amount of a paramagnetic metal containing magnetic resonance contrast agent;subjecting said human body to a magnetic resonance image procedure capable of generating from magnetic resonance signals from said human body successive images of temporally spaced images taken over an acquisition time period of at least part of said human body into which said contrast agent passes, said procedure being a magnetic resonance imaging procedure;detecting temporal variations in said signals or images; andfrom said temporal variations identifying regions of abnormal or modified blood flow in said human body and providing a quantitative indication of blood flow abnormality or variation. 3. A method of detecting and quantitatively evaluating the severity of blood flow abnormality in a human body, said method comprising the steps of: administering into vasculature of said human body a contrast enhancing amount of a paramagnetic metal containing magnetic resonance contrast agent;subjecting said human body to a magnetic resonance image procedure capable of generating from magnetic resonance signals from said human body successive images of temporally spaced images taken over an acquisition time period of at least part of said human body into which said contrast agent passes, said procedure being a magnetic resonance imaging procedure, to detect temporal variations in said magnetic resonance signals or images;detecting blood flow abnormality or flow variation in obstructed blood vessels in said body; and identifying from said temporal variations in said images the blood flow abnormality. 4. A method of detecting blood flow abnormality or variation in a blood vessel comprising: administering a contrast enhancing amount of a paramagnetic metal containing magnetic resonance contrast agent into a blood vessel of a body;imaging at least a portion of the body through which the MR contrast ... | Kucharczyk and Moseley Partners | A61B005/055
US07211050 | B1 | 2007-05-01 | System for enhancement of neurophysiological processes | 2000-06-15 | Treatment apparatus and method for the enhancement of neurophysiological processes of a patient's visual system utilizes a computer generated display that provides a patient simultaneously with passive visual stimulation images and visual cognitive exercises. The stimulation images include moving elements having edges or points of contrast. The exercises include visual and phonological elements. The invention is used to treat reading, writing and speech disorders, and attention deficit hyperactivity disorder, as well as having broader applications for example in sports sciences. | 1. Apparatus for the enhancement of neurophysiological processes of a patient by the stimulation of receptive cell fields in the visual pathways of the patient between the retina and the visual cortex, the apparatus including first visual display means for viewing by said patient and computer processing means producing an output to said first visual display means to cause a display on said visual display means, said display including at least one visual cognitive exertion exercise and at least one visual stimulation image including one or more therapeutic display elements targeted to stimulate selected ones of said receptive cell fields, said therapeutic display elements including one or more moving contrast edges, wherein said therapeutic display elements are displayed on said visual display means so as to provide therapeutic stimulation to said receptive cell fields of a patient whilst said patient is performing said cognitive exertion exercise; further including second visual display means to be viewed by a therapist and therapist input means allowing said therapist to provide inputs to said processing means to vary said display;further including feedback means providing an indication of said patient's performance to said computer processing means, and means for indicating said performance on said second visual display means; andfurther including means for varying the display on said first visual display means to optimise said patient's brain response. | 1. Apparatus for the enhancement of neurophysiological processes of a patient by the stimulation of receptive cell fields in the visual pathways of the patient between the retina and the visual cortex, the apparatus including first visual display means for viewing by said patient and computer processing means producing an output to said first visual display means to cause a display on said visual display means, said display including at least one visual cognitive exertion exercise and at least one visual stimulation image including one or more therapeutic display elements targeted to stimulate selected ones of said receptive cell fields, said therapeutic display elements including one or more moving contrast edges, wherein said therapeutic display elements are displayed on said visual display means so as to provide therapeutic stimulation to said receptive cell fields of a patient whilst said patient is performing said cognitive exertion exercise; further including second visual display means to be viewed by a therapist and therapist input means allowing said therapist to provide inputs to said processing means to vary said display;further including feedback means providing an indication of said patient's performance to said computer processing means, and means for indicating said performance on said second visual display means; andfurther including means for varying the display on said first visual display means to optimise said patient's brain response. 2. Apparatus according to claim 1 wherein the variations to the display on said first visual display means are made by said computer processing means without input from said therapist. 3. Apparatus for the enhancement of neurophysiological processes of a patient by the stimulation of receptive cell fields in the visual pathways of the patient between the retina and the visual cortex, the apparatus including first visual display means for viewing by said patient and computer processing means producing an output to said first visual display means to cause a display on said visual display means, said display including at least one visual cognitive exertion exercise and at least one visual stimulation image including one or more therapeutic display elements targeted to stimulate selected ones of said receptive cell fields, said therapeutic display elements including one or more moving contrast edges, wherein said therapeutic display elements are displayed on said visual display means so as to provide therapeutic stimulation to said receptive cell fields of a patient whilst said patient is performing said cognitive exertion exercise; and further including means for controlling said patient's position relative to said first visual display means,wherein said means for controlling said patient's position includes an electronically controllable seat capable of moving up, down and laterally, an electronically controllable backrest and an electronically controllable headrest. 4. Apparatus for the enhanc... | - | A61B005/04 A61B013/00 A61B003/02
US07164268 | B2 | 2007-01-16 | Method and apparatus for spin-echo-train MR imaging using prescribed signal evolutions | 2001-12-21 | A magnetic resonance imaging “MRI” method and apparatus for lengthening the usable echo-train duration and reducing the power deposition for imaging is provided. The method explicitly considers the t1 and t2 relaxation times for the tissues of interest, and permits the desired image contrast to be incorporated into the tissue signal evolutions corresponding to the long echo train. The method provides a means to shorten image acquisition times and/or increase spatial resolution for widely-used spin-echo train magnetic resonance techniques, and enables high-field imaging within the safety guidelines established by the Food and Drug Administration for power deposition in human MRI. | 1. A method for generating a spin echo pulse sequence for operating a magnetic resonance imaging apparatus for imaging an object that permits at least one of lengthening usable echo-train duration, reducing power deposition and incorporating desired image contrast into the tissue signal evolutions, said method comprising: a) providing contrast-preparation, said contrast-preparation comprising generating at least one of at least one radio-frequency pulse, at least one magnetic-field gradient pulse, and at least one time delay, whereby said contrast preparation encodes the magnetization with at least one desired image contrast;b) calculating flip angles and phases of refocusing radio-frequency pulses that are applied in a data-acquisition step, wherein said calculation provides desired prescribed signal evolution and desired overall signal level, said calculation comprises: i) selecting values of T1 and T2 relaxation times and selecting proton density;ii) selecting a prescribed time course of the amplitudes and phases of the radio-frequency magnetic resonance signals that are generated by said refocusing radio-frequency pulses; andiii) selecting characteristics of said contrast-preparation step, said data-acquisition step and a magnetization-recovery step, with the exception of the flip angles and phases of the refocusing radio-frequency pulses that are to be calculated; andc) providing said-data acquisition step based on a spin echo train acquisition, said data-acquisition step comprises: i) an excitation radio-frequency pulse having a flip angle and phase;ii) at least two refocusing radio-frequency pulses, each having a flip angle and phase as determined by said calculation step; andiii) magnetic-field gradient pulses that encode spatial information into at least one of said radio-frequency magnetic resonance signals that follow at least one of said refocusing radio-frequency pulses;d) providing magnetization-recovery, said magnetization-recovery comprises a time delay to allow magnetization to relax; ande) repeating steps (a) through (d) until a predetermined extent of spatial frequency space has been sampled. | 1. A method for generating a spin echo pulse sequence for operating a magnetic resonance imaging apparatus for imaging an object that permits at least one of lengthening usable echo-train duration, reducing power deposition and incorporating desired image contrast into the tissue signal evolutions, said method comprising: a) providing contrast-preparation, said contrast-preparation comprising generating at least one of at least one radio-frequency pulse, at least one magnetic-field gradient pulse, and at least one time delay, whereby said contrast preparation encodes the magnetization with at least one desired image contrast;b) calculating flip angles and phases of refocusing radio-frequency pulses that are applied in a data-acquisition step, wherein said calculation provides desired prescribed signal evolution and desired overall signal level, said calculation comprises: i) selecting values of T1 and T2 relaxation times and selecting proton density;ii) selecting a prescribed time course of the amplitudes and phases of the radio-frequency magnetic resonance signals that are generated by said refocusing radio-frequency pulses; andiii) selecting characteristics of said contrast-preparation step, said data-acquisition step and a magnetization-recovery step, with the exception of the flip angles and phases of the refocusing radio-frequency pulses that are to be calculated; andc) providing said-data acquisition step based on a spin echo train acquisition, said data-acquisition step comprises: i) an excitation radio-frequency pulse having a flip angle and phase;ii) at least two refocusing radio-frequency pulses, each having a flip angle and phase as determined by said calculation step; andiii) magnetic-field gradient pulses that encode spatial information into at least one of said radio-frequency magnetic resonance signals that follow at least one of said refocusing radio-frequency pulses;d) providing magnetization-recovery, said magnetization-recovery comprises a time delay to allow magnetization to relax; ande) repeating steps (a) through (d) until a predetermined extent of spatial frequency space has been sampled. 2. The method of claim 1, wherein said calculation of the flip angles and phases is generated using an appropriate analytical or computer-based algorithm. 3. The method of claim 1, wherein said calculation of the flip angles and phases is generated to account for, the effects of multiple applications of: said contrast-preparation, said data-acquisition and said magnetization-recovery steps, which are required to sample the desired extent of spatial-frequency space. 4. The method of claim 1, wherein a two-dimensional plane of spatial-frequency space is sampled. 5. The method of claim 1, wherein a three-dimensional volume of spatial-frequency space is sampled. 6. The method of claim 1, wherein at least one of said contrast-preparation and magnetization-recovery steps is omitted. 7. The method of claim 1, wherein said calculation step i... | University of Virginia Patent Foundation | G01V003/00 A61B005/055
USRE038695 | E1 | 2005-02-08 | Extravasation detection electrode patch | 2001-09-07 | A medical extravasation device has an electrode patch that can attach to the skin for sensing electrical information. Tissue impedance is calculated from the electrode patch signals. The patch has elongate pick-up electrodes inboard of elongate energizing electrodes. The measuring zone determined by the elongate space between the pick-up electrodes enhances sensitivity and specificity. The presence of an extravasation is determined by interpreting the tissue impedance measurement. The method for determining the extravasation includes a first step of determining a pre-injection baseline measurement of the tissue impedance. Then, the tissue impedance is monitored during the procedure itself. A predetermined amount of change in tissue impedance is determined to indicate an extravasation. | 1. An electrode patch for use in a non-invasive device for detecting extravasation that may occur when a needle with a tip is inserted into a patient in order to introduce fluid into the vascular system of a patient, comprising: a body of the patch adapted to be affixed to the skin of a patient, an outer pair of elongated electrodes and an inner pair of elongate electrodes, the length of each of said electrodes being deployed along the body of said patch, said inner pair of electrodes being spaced from one another on either side of a center line, said inner pair defining a measuring zone, said measuring zone being shaped and dimensioned to encompass a needle tip within said zone, said zone being small enough to optimize sensitivity, said inner pair being long enough to facilitate placement of the patch over a needle tip inserted into a patient, each of said outer pair of electrodes being outward, relative to said center line, of a respective one of said inner electrodes, energization of said outer electrodes, when said patch is affixed to the skin of a patient, providing a field which induces a signal in said inner electrodes that is a function of the impedance of the tissue in said measuring zone. | 1. An electrode patch for use in a non-invasive device for detecting extravasation that may occur when a needle with a tip is inserted into a patient in order to introduce fluid into the vascular system of a patient, comprising: a body of the patch adapted to be affixed to the skin of a patient, an outer pair of elongated electrodes and an inner pair of elongate electrodes, the length of each of said electrodes being deployed along the body of said patch, said inner pair of electrodes being spaced from one another on either side of a center line, said inner pair defining a measuring zone, said measuring zone being shaped and dimensioned to encompass a needle tip within said zone, said zone being small enough to optimize sensitivity, said inner pair being long enough to facilitate placement of the patch over a needle tip inserted into a patient, each of said outer pair of electrodes being outward, relative to said center line, of a respective one of said inner electrodes, energization of said outer electrodes, when said patch is affixed to the skin of a patient, providing a field which induces a signal in said inner electrodes that is a function of the impedance of the tissue in said measuring zone. 2. The patch of claim 1 wherein each of said electrodes further includes a first portion which is relatively short as compared to said elongate portion of each electrode, said relatively short first portions of said electrodes define a coupling region which can be connected to a single clip containing electrical contacts. 3. The patch of claim 1 wherein each of said elongate electrodes are substantially the same length and wherein each of said outer pair of electrodes is adjacent to and spaced from one of said inner pair of electrodes. 4. The patch of claim 3 wherein each of said elongate electrodes has a total length of about 3 inches, and a width of about {fraction (3/16)}th of an inch, and wherein said inner pair of electrodes is spaced apart by about 0.75 inches and said outer pair of electrodes is spaced apart by about 1.5 inches, and where said inner pair of electrodes is centralized relative to said outer pair of electrodes. 5. The patch of claim 4 wherein said body is about two inches in width and about three inches in length. 6. The patch of claim 1 wherein said electrodes are silver/silver chloride strips. 7. A device for use with a non-invasive detection system to detect extravasation when fluid is delivered into a patient's vascular system, comprising: a) an outer pair of elongated electrodes and an inner pair of elongated electrodes, b) said inner pair of electrodes being spaced from one another on either side of a center line, said inner pair defining a measuring zone, said measuring zone being shaped and dimensioned to encompass within said zone an end of a channel inserted in the patient's vascular system for delivering fluid into the vascular system, said zone being small enough to optimize sensitivity, said inner pair being lon... | E-Z-EM, Inc. | A61B005/04
US06952605 | B1 | 2005-10-04 | Pneumatic release mechanism for a patient contacting article | 2001-08-08 | A patient contacting assembly comprising a patient contacting member that secures to the surface of a patient by means of an adhesive disposed between the patient contacting member and the patient. A channel is defined through the patient contacting member to enable a release fluid to be injected between the patient contacting member that the patient. Providing the release fluid between the patient contacting member and the patient at least partially reduces the bonding strength of the adhesive, making it easier to remove the patient contacting member from the patient and provides a pneumatic lift, urging the patient contacting member off of the surface of the patient. | 1. A patient contacting assembly, comprising: a patient contacting member including a first surface adapted to overlie a portion of a patient and a second surface;an adhesive disposed on the first surface and adapted to secure the first surface to a surface of such a patient;a stem having a first end portion operatively coupled to the second surface of the patient contacting member and an exposed second end portion; anda channel defined through the stem from the first end portion to the second end portion and defined through the patient contacting member from the second surface to the first surface, wherein the second end portion of the stem is adapted to receive a release fluid for introduction into the channel, wherein the channel communicates the release fluid proximate to the first surface so as to dispense such a release fluid from the channel between the first surface and a surface of a patient to which the patient contacting member is adhered, wherein the channel is sized and configured such that introducing the release fluid into the channel urges the first surface and the surface of the patient apart, thereby dislodging the patient contacting assembly from the patient in a pneumatic fashion, and wherein the channel includes multiple branches emanating from the channel to disperse the release fluid over an area of the first surface. | 1. A patient contacting assembly, comprising: a patient contacting member including a first surface adapted to overlie a portion of a patient and a second surface;an adhesive disposed on the first surface and adapted to secure the first surface to a surface of such a patient;a stem having a first end portion operatively coupled to the second surface of the patient contacting member and an exposed second end portion; anda channel defined through the stem from the first end portion to the second end portion and defined through the patient contacting member from the second surface to the first surface, wherein the second end portion of the stem is adapted to receive a release fluid for introduction into the channel, wherein the channel communicates the release fluid proximate to the first surface so as to dispense such a release fluid from the channel between the first surface and a surface of a patient to which the patient contacting member is adhered, wherein the channel is sized and configured such that introducing the release fluid into the channel urges the first surface and the surface of the patient apart, thereby dislodging the patient contacting assembly from the patient in a pneumatic fashion, and wherein the channel includes multiple branches emanating from the channel to disperse the release fluid over an area of the first surface. 2. The patient contacting assembly according to claim 1, wherein the adhesive is a hydrogel adhesive disposed on the first surface or a surface of a patient. 3. The patient contacting assembly according to claim 1, wherein the stem is integral with the patient contacting member. 4. The patient contacting assembly according to claim 1, further comprising an electrode provided on the first surface of the patient contacting member. 5. The patient contacting assembly according to claim 1, further comprising means, associated with the patient contacting member, for applying a distending force on a surface of a patient to which the patient contacting member is adhered. 6. The patient contacting assembly according to claim 1, wherein the patient contacting member is defined, at least in party, by a substantially rigid material. 7. A patient contacting assembly comprising: a patient contacting member including a first surface adapted to overlie a portion of a patient and a second surface;adhering means for securing the first surface to a surface of such a patient;a stem having (a) a first end portion operatively coupled to the second surface of the patient contacting member and (b) an exposed second end portion; andreleasing means associated with the patient contacting member and the stem for delivering a release fluid, which is introduced into second end portion of the stem, to a location between the first surface and a surface of a patient to which the patient contacting member is adhered, wherein the release means is sized and configured such that introducing the release fluid into the stem urges the first su... | Respironics, Inc. | A61B005/04
US06858015 | B2 | 2005-02-22 | Blood withdrawal system | 2002-05-03 | Blood withdrawal system for withdrawing blood for diagnostic purposes, comprising a housing with an outlet opening for the tip of a lancet, a lancet guide and a lancet drive. The lancet drive comprises a plural lever coupling mechanism, forming a connection between the drive element and the lancet during the puncturing movement and comprising two levers connected to each other via a first swivel joint, wherein one of the levers is coupled to the lancet with its end facing towards the lancet by means of a second swivel joint, and the second lever comprises a third swivel joint at the end facing away from the lancet. During the cocking movement, the freedom of movement of the lancet is limited in such a manner that the lancet tip does not protrude from the outlet opening. | 1. A blood withdrawal system for diagnostic purposes, comprising a housing with an outlet opening for the tip of a lancet, which is movable in the housing along a predetermined puncture path, a lancet guide for guiding the lancet on the predetermined puncture path and a lancet drive by which a movement of a drive element is converted into a puncturing movement after the actuation of a trigger, whereby the lancet is moved with high speed along the predetermined puncture path in puncturing direction until its tip protrudes from the outlet opening, wherein the lancet drive comprises a plural lever coupling mechanism forming a connection between the drive element and the lancet during the puncturing movement and comprising two levers connected to each other by means of a first swivel joint, one of the levers is coupled to the lancet by means of a second swivel joint, and the second lever comprises a third swivel joint at its end facing away from the lancet, and the plural lever coupling mechanism is moved, during the puncturing movement, from a first bent position via a straight position corresponding to the maximum puncturing depth to a second bent position, and is moved, during a cocking movement, from the second bent position via the straight position to the first bent position, the freedom of movement of the lancet is limited during the cocking movement in such a manner that the lancet tip does not protrude from the outlet opening, and the third swivel joint is connected to a moveable bearing part which during the cocking movement moves inside the housing in a direction opposite to the puncturing direction, and the freedom of movement of which is limited during the puncturing movement in the direction opposite to the puncturing direction. | 1. A blood withdrawal system for diagnostic purposes, comprising a housing with an outlet opening for the tip of a lancet, which is movable in the housing along a predetermined puncture path, a lancet guide for guiding the lancet on the predetermined puncture path and a lancet drive by which a movement of a drive element is converted into a puncturing movement after the actuation of a trigger, whereby the lancet is moved with high speed along the predetermined puncture path in puncturing direction until its tip protrudes from the outlet opening, wherein the lancet drive comprises a plural lever coupling mechanism forming a connection between the drive element and the lancet during the puncturing movement and comprising two levers connected to each other by means of a first swivel joint, one of the levers is coupled to the lancet by means of a second swivel joint, and the second lever comprises a third swivel joint at its end facing away from the lancet, and the plural lever coupling mechanism is moved, during the puncturing movement, from a first bent position via a straight position corresponding to the maximum puncturing depth to a second bent position, and is moved, during a cocking movement, from the second bent position via the straight position to the first bent position, the freedom of movement of the lancet is limited during the cocking movement in such a manner that the lancet tip does not protrude from the outlet opening, and the third swivel joint is connected to a moveable bearing part which during the cocking movement moves inside the housing in a direction opposite to the puncturing direction, and the freedom of movement of which is limited during the puncturing movement in the direction opposite to the puncturing direction. 2. A blood withdrawal system according to claim 1, wherein the freedom of movement of the lancet is limited during the cocking movement by means of a trigger for triggering the puncturing movement. 3. A blood withdrawal system according to claim 1, wherein the freedom of movement of the lancet is limited by a stop element. 4. A blood withdrawal system according to claim 1, wherein the freedom of movement of the bearing part is limited during the puncturing movement by a stop element. 5. A blood withdrawal system according to claim 1, wherein the freedom of movement of the bearing part is limited during the puncturing movement elastically by a spring. 6. A blood withdrawal system according to claim 5, wherein the spring is pretensioned. 7. A blood withdrawal system according to claim 1, wherein the plural lever coupling mechanism comprises three levers. 8. A blood withdrawal system according to claim 1, wherein the angular separation of the swivel movement of the second lever between the first bent position and the straight position corresponding to the maximum puncturing depth exceeds 90°. 9. A blood withdrawal system according to claim 8, wherein the drive element is a torsion spring acting on the sec... | Roche Diagnostics Operations, Inc. | A61B005/00
US06853855 | B2 | 2005-02-08 | Magnetic resonance tomography apparatus with improved spatial and time stabilization of the homogeneity of the magnetic basic field | 2002-07-09 | A magnetic resonance apparatus has a basic field magnet for providing a uniform magnetic basic field, a gradient coil system for coding within the basic field as well as a passive shim system for homogenizing the magnetic basic field. This passive shim system is composed of shim iron plates that are placed in the magnetic basic field. A device is provided with which the temperature of these shim iron plates is controlled. | 1. A magnetic resonance apparatus comprising: a basic field magnet which generates a uniform basic magnetic field; a gradient coil system for encoding signals within said basic magnetic field; a passive shim system having a plurality of shim iron plates disposed to interact with said basic field magnet to homogenize said basic magnetic field; and a thermal control device comprising a bridge circuit containing transistors, operatonal amplifiers and a Peltier element in thermal communication with at least one of said shim iron plates, and a temperature regulating circuit, connected to said bridge circuit, for generating an output signal for regulating a temperature of said at least one of said shim iron plates, said output signal switching said transistors for selectively placing said Peltier element in one of an operating mode for elevating the temperature said at least one of the shim iron plates or a mode for decreasing the temperature said at least one of the shim iron plates. | 1. A magnetic resonance apparatus comprising: a basic field magnet which generates a uniform basic magnetic field; a gradient coil system for encoding signals within said basic magnetic field; a passive shim system having a plurality of shim iron plates disposed to interact with said basic field magnet to homogenize said basic magnetic field; and a thermal control device comprising a bridge circuit containing transistors, operatonal amplifiers and a Peltier element in thermal communication with at least one of said shim iron plates, and a temperature regulating circuit, connected to said bridge circuit, for generating an output signal for regulating a temperature of said at least one of said shim iron plates, said output signal switching said transistors for selectively placing said Peltier element in one of an operating mode for elevating the temperature said at least one of the shim iron plates or a mode for decreasing the temperature said at least one of the shim iron plates. 2. A magnetic resonance apparatus as claimed in claim 1 wherein said temperature control device stabilizes the temperature of said at least one of said shim iron plates. 3. A magnetic resonance apparatus as claimed in claim 1 wherein said temperature control device homogenizes said basic magnetic field by causing said at least one of said shim iron plates to exhibit a targeted temperature. 4. A magnetic resonance apparatus as claimed in claim 1 wherein said at least one of said shim iron plates is mounted on a non-magnetic, thermally conductive plate at a first side thereof, and wherein said Peltier element is mounted at a second side of said thermally conductive plate, opposite to said first side. 5. A magnetic resonance apparatus as claimed in claim 1 wherein said temperature control device comprises a databus on which a digital drive signal is supplied for said temperature regulating circuit, and a digital-to-analog converter connected between said databus and said temperature regulating circuit. 6. A magnetic resonance apparatus as claimed in claim 1 wherein a first pair of said transistors in said bridge circuit are charged with current for elevating a temperature of said at least one of said shim iron plates and are mounted in thermal contact with said at least one of said shim iron plates. | Siemens Aktiengesellschaft | A61B005/055
US06931274 | B2 | 2005-08-16 | Processing EEG signals to predict brain damage | 2002-05-06 | Rapid and accurate in-vivo assessment of cerebral white matter injury particularly for pre-term infants, for timely treatment and/or prediction of outcomes has been very limited. This invention exploits the discovery that reduced high-frequency EEG intensity, particularly as shown by the upper spectral edge frequency, is a good indicator of cerebral white matter neural injury and is well correlated with MRI results. With more experience of clinical cases, a set of simple rules such as “if the spectral edge value is below 8 Hz there is a high likelihood of injury” may be validated, yet the EEG technology involved is largely invisible to the user. In the invention, EEG signals are processed by software to obtain, store, and graphically display bilaterally collected EEG spectral edge and intensity values over from hours to weeks. Rejection of corrupted signals by filtering and gating means is responsive to incoming signal characteristics, to additional inputs such as motion sensors or impedance tests, and to patient data (gestational age in particular). The invention includes the software and methods of use. | 1. An apparatus for processing an electroencephalogram (EEG) signal from a subject to assist in management of possible cerebral white-matter neural injury in the subject, comprising: (a) apparatus to acquire at least one channel of an EEG signal from the subject over a period of time, and (b) computing apparatus programmed to analyze the frequency distribution of the EEG signal so acquired, from within a range of from 1 Hz to 50 Hz, and to produce condensed output information descriptive and/or predictive of cerebral white-matter neural injury in the subject. | 1. An apparatus for processing an electroencephalogram (EEG) signal from a subject to assist in management of possible cerebral white-matter neural injury in the subject, comprising: (a) apparatus to acquire at least one channel of an EEG signal from the subject over a period of time, and (b) computing apparatus programmed to analyze the frequency distribution of the EEG signal so acquired, from within a range of from 1 Hz to 50 Hz, and to produce condensed output information descriptive and/or predictive of cerebral white-matter neural injury in the subject. 2. The apparatus of claim 1 further comprising apparatus programmed to analyze the frequency distribution by determining a numerical value for the upper spectral edge of the frequency distribution of the intensity of the EEG signals. 3. The apparatus of claim 2 further comprising apparatus to determine the spectral edge for each of a series of time intervals, to store a corresponding series of spectral edge values, and to include the series of spectral edge values on a display, thereby to enable forecasting of an outcome of cerebral white-matter neural injury in the subject. 4. The apparatus of claim 3 further comprising apparatus to determine, within the range of 1 Hz to 35 Hz, the upper spectral edge below which about 95% of the EEG intensity occurs. 5. The apparatus of claim 4 further comprising apparatus to provide a comparison of the determined upper spectral edge with stored EEG spectral edge and associated neurological outcome information from previous cases, thereby to enable forecasting of an outcome of cerebral white-matter neural injury in the subject. 6. The apparatus of claim 5 further comprising apparatus to acquire a plurality of channels of EEG signals, so that spectral edge recordings are made from a plurality of positions on the head of the subject, thereby to improve confidence in forecasting outcome of cerebral white-matter neural injury in the subject. 7. A program embodied on a computer readable medium in combination with the apparatus of claim 5, the program comprising a routine to display comparable portions of both a current record and a previously collected record thereby permitting a comparison of the current record with stored information derived from the EEG and neurological outcome information from previous cases, thereby enabling prediction of an outcome of cerebral white-matter neural injury in the subject. 8. The apparatus of claim 4 further comprising apparatus programmed to eliminate an artifact in output information by action of at least one event-responsive filtering procedure to delete from the output information data likely to be in error. 9. The apparatus of claim 8 where the event-responsive filtering procedure is a signal interruption procedure that rejects the determined upper spectral edge during detection of a condition within the EEG signal selected from abnormally low signal amplitude, abnormally high signal amplitude, a signal indicati... | Tru-Test Corporation Limited | A61B005/00
US06895265 | B2 | 2005-05-17 | Implantable sensor | 2002-06-25 | A sensor is disclosed, for implantation within a blood vessel to monitor an analyte in blood. In one embodiment, the analyte is glucose. A signal indicative of glucose level is transmitted to an external receiver. The signal may also be used to drive an internal or externally worn insulin pump. Methods are also disclosed. | 1. A blood glucose detector for implantation within a blood vessel, comprising: a tubular support, having a first side for contacting the wall of the vessel and a second side for facing radially inwardly toward the center of the vessel; a sensor carried by the support, the sensor having a streamlined configuration with respect to the support and having a sensing surface thereon; wherein the sensing surface is spaced radially inwardly from the first side by a distance of at least about 0.2 mm, such that the velocity of blood in the vessel inhibits obstruction of the sensing surface; and a transmitter on the support, for transmitting information from the sensor to an external receiver. | 1. A blood glucose detector for implantation within a blood vessel, comprising: a tubular support, having a first side for contacting the wall of the vessel and a second side for facing radially inwardly toward the center of the vessel; a sensor carried by the support, the sensor having a streamlined configuration with respect to the support and having a sensing surface thereon; wherein the sensing surface is spaced radially inwardly from the first side by a distance of at least about 0.2 mm, such that the velocity of blood in the vessel inhibits obstruction of the sensing surface; and a transmitter on the support, for transmitting information from the sensor to an external receiver. 2. A blood glucose detector as in claim 1, wherein the distance is at least about 0.5 mm. 3. A blood glucose detector as in claim 1, wherein the distance is within the range of from about 0.3 mm to about 2.5 mm. 4. A blood glucose detector as in claim 1, wherein the distance is sufficient to increase the blood flow velocity at the sensing surface to from about 125% to about 200% of the blood flow velocity immediately proximal to the support. 5. A blood glucose detector as in claim 1, wherein an inductive link supplies power to the transmitter. 6. A blood glucose detector as in claim 1, further comprising a thin film rechargeable battery carried by the support. 7. A blood glucose detector as in claim 1, wherein the support comprises an enlargeable frame. 8. A blood glucose detector as in claim 7, wherein the support comprises an expandable tubular body. 9. A blood glucose detector as in claim 1, wherein the support comprises a balloon expandable stent. 10. A blood glucose detector as in claim 1, wherein the support comprises a self expandable stent. 11. A blood glucose detector as in claim 8, further comprising a tubular sheath on the tubular body. 12. A blood glucose detector as in claim 11, wherein the sheath is on the radially outwardly facing surface of the tubular body. 13. A blood glucose sensor as in claim 1, wherein the sensor comprises an analyte permeable membrane and an enzyme gel layer. 14. A blood glucose detector as in claim 13, wherein the analyte permeable membrane comprises a glucose permeable membrane. 15. An implantable sensor for sensing the presence of an analyte in a vessel, comprising: a tubular support structure for anchoring the sensor in the vessel, the support having a side wall with a luminal side facing toward the center of the vessel and an abluminal side facing toward the wall of the vessel; a sensor housing carried by the support structure, the housing having a streamlined configuration with respect to the support to minimize blood flow turbulence; a power supply and electrical circuitry in the housing; a sensing surface exposed to the exterior of the housing; wherein the sensing surface is positioned on the radially inwardly most portion of the luminal side of the housing; and a transmitter on the support structure for t... | - | A61B005/00
US06915162 | B2 | 2005-07-05 | Implantable medical device for measuring ventricular pressure | 2001-11-15 | An implantable medical device has a pressure sensing arrangement to measure right ventricular pressure of a heart including a pressure sensor adapted to be positioned in the right ventricle of the heart, to measure the pressure and to generate a pressure signal in response to the measured pressure. The pressure sensing arrangement also has a pressure signal processor and a timing unit. The processor determines from the pressure signal, using diastolic timing signals from the timing unit based on the pressure signal identifying the diastolic phase, a diastolic pressure signal representing the ventricular pressure only during the diastolic phase of the heart cycle. | 1. An implantable medical device comprising: a pressure sensor adapted to be positioned in the right ventricle of a heart, for measuring right ventricular pressure and for generating a pressure signal corresponding to the measured right ventricular pressure; a timing unit supplied with said pressure signal which determines, from said pressure signal, diastolic timing signals identifying a beginning and an end of a diastolic phase of the heart; and a signal processor connected to the timing unit and also supplied with said pressure signal, said signal processor, using said diastolic timing signals, determining from said pressure signal a diastolic pressure signal representing right ventricular pressure only during the entire diastolic phase of the heart cycle defined by said beginning and said end. | 1. An implantable medical device comprising: a pressure sensor adapted to be positioned in the right ventricle of a heart, for measuring right ventricular pressure and for generating a pressure signal corresponding to the measured right ventricular pressure; a timing unit supplied with said pressure signal which determines, from said pressure signal, diastolic timing signals identifying a beginning and an end of a diastolic phase of the heart; and a signal processor connected to the timing unit and also supplied with said pressure signal, said signal processor, using said diastolic timing signals, determining from said pressure signal a diastolic pressure signal representing right ventricular pressure only during the entire diastolic phase of the heart cycle defined by said beginning and said end. 2. An implantable medical device as claimed in claim 1 wherein said timing unit comprises a differentiator which differentiates said pressure signal to obtain a differentiated pressure signal, and a comparator supplied with said pressure signal and said differentiated pressure signal, said timing unit also being supplied with respective threshold values for said pressure signal and for said differentiated pressure signal and comparing said pressure signal and said differentiated pressure signal with the respective threshold values to generate said diastolic timing signals. 3. An implantable medical device as claimed in claim 2 wherein said timing unit comprises a further differentiator which differentiates said differentiated pressure signal to obtain a second differentiated pressure signal, and wherein said timing unit uses said second differentiated pressure signal in combination with said pressure signal as the respective threshold values. 4. An implantable medical device as claimed in claim 2 wherein said timing unit comprises a further differentiator which differentiates said differentiated pressure signal to obtain a second differentiated pressure signal, and wherein said timing unit uses said second differentiated pressure signal in combination with the differentiated pressure signal as the respective threshold values. 5. An implantable medical device as claimed in claim 2 wherein said timing unit comprises a further differentiator which differentiates said differentiated pressure signal to obtain a second differentiated pressure signal, and wherein said timing unit uses said second differentiated pressure signal in combination with the pressure signal and the differentiated pressure signal as the respective threshold values. 6. An implantable medical device as claimed in claim 1 wherein said signal processor generates said diastolic pressure signal substantially continuously during an entirety of said diastolic phase. 7. An implantable medical device as claimed in claim 1 wherein said signal processor comprises a median filter which generates a smoothed combination of respective diastolic pressure signals from a plurality of successive he... | St. Jude Medical AB | A61N001/365 A61B005/04
US06901286 | B1 | 2005-05-31 | Method and system for producing a higher quality electromyographic signal from an electrode array | 2000-07-07 | A method and system for producing higher quality electromyogram (EMG) signals utilizes an array of electrodes for sensing a plurality of EMG signals in an electrically active region of a subject's muscle. A weighting function is applied to the EMG signals to produce weighted signals. This weighting function contains correction features for the relative locations of the center of the electrically active region and the electrodes. The quality of the weighted EMG signals is evaluated, and the weighted signals or sum or mean of the weighted signals whose evaluated quality is insufficient are replaced. A sum or mean of a feature of the weighted signals is calculated to produce a higher quality electromyocardiographic signal. The method and system can also be used to determine signal strength or frequency contents of a signal falling outside the array of electrodes. | 1. A method of producing a higher quality electromyographic signal describing myoelectrical activity of an electrically active region of a subject's muscle, comprising: sensing through an array of electrodes a plurality of EMG signals representative of the myoelectrical activity of the electrically active region of the subject's muscle; applying a weighting function to the detected EMG signals and thereby producing weighted signals, the electrically active region of the subject's muscle comprising a center and the weighting function containing correction features for the relative locations of the center of the electrically active region and the electrodes; and combining the weighted signals and thereby producing the higher quality electromyographic signal. | 1. A method of producing a higher quality electromyographic signal describing myoelectrical activity of an electrically active region of a subject's muscle, comprising: sensing through an array of electrodes a plurality of EMG signals representative of the myoelectrical activity of the electrically active region of the subject's muscle; applying a weighting function to the detected EMG signals and thereby producing weighted signals, the electrically active region of the subject's muscle comprising a center and the weighting function containing correction features for the relative locations of the center of the electrically active region and the electrodes; and combining the weighted signals and thereby producing the higher quality electromyographic signal. 2. A method of producing a higher quality electromyographic signal as defined in claim 1, wherein the weighting function comprises correction features for both cancellation and distance damping effects. 3. A method of producing a higher quality electromyographic signal as defined in claim 1, wherein combining the weighted signals comprises: adding a feature of the weighted signals together. 4. A method of producing a higher quality electromyographic signal describing myoelectrical activity of an electrically active region of a subject's muscle, comprising: sensing through an array of electrodes a plurality of EMG signals representative of the myoelectrical activity of the electrically active region of the subject's muscle; applying a weighting function to the detected EMG signals and thereby producing weighted signals; and combining the weighted signals and thereby producing the higher quality electromyographic signal; wherein: the electrically active region of the subject's muscle comprises a center; the electrodes are separated from the center of the electrically active region by respective distances; the electrodes are separated from each other by an inter-electrode distance; and the weighting function comprises correction features for: the relative location of the center of the electrically active region and the electrodes; the distance separating the center of the electrically active region and the electrodes; the size of the electrically active region; and the inter-electrode distance. 5. A method of producing a higher quality electromyographic signal describing myoelectrical activity of an electrically active region of a subject's muscle, comprising: sensing through an array of electrodes a plurality of EMG signals representative of the myoelectrical activity of the electrically active region of the subject's muscle; applying a weighting function to the detected EMG signals and thereby producing weighted signals, the weighting function containing correction features for the relative locations of the electrically active region and the electrodes; and combining the weighted signals and thereby producing the higher quality electromyographic signal; wherein the electrically active re... | Universite de Montreal | A61B005/0488
US06901281 | B2 | 2005-05-31 | Method of magnetic resonance imaging | 2001-03-22 | The invention provides a method of magnetic resonance imaging of regional blood oxygenation which comprises administering into the vasculature of a vascularized human or non-human animal subject a T2 blood pool contrast agent, detecting a magnetic resonance signal from at least part of the vasculature of said subject into which said contrast agent distributes, and manipulating said signal to generate an indication of the partial pressure of oxygen (pO2) in at least part of said vasculature. | 1. A method of magnetic resonance imaging of regional blood oxygenation which comprises administering into the vasculature of a vascularised human or non-human animal subject a T2 blood pool contrast agent, detecting a magnetic resonance signal from at least part of the vasculature of said subject into which said contrast agent distributes, and manipulating said signal to generate an indication of the partial pressure of oxygen (pO2) in at least part of said vasculature wherein R1 of blood in an artery is determined and hematocrit is determined therefrom. | 1. A method of magnetic resonance imaging of regional blood oxygenation which comprises administering into the vasculature of a vascularised human or non-human animal subject a T2 blood pool contrast agent, detecting a magnetic resonance signal from at least part of the vasculature of said subject into which said contrast agent distributes, and manipulating said signal to generate an indication of the partial pressure of oxygen (pO2) in at least part of said vasculature wherein R1 of blood in an artery is determined and hematocrit is determined therefrom. 2. The method of claim 1 wherein hematocrit and R2* of blood in a vein are determined and pO2 in said vein is determined therefrom. 3. The method of claim 1 wherein said indication is an image indicative of pO2. 4. The method of claim 3 wherein said contrast agent is administered in a quantity such that R2 of venous blood is greater than R2 of arterial blood. 5. The method of claim 4 wherein a non-spoiled gradient echo sequence is used to generate a magnetic resonance image. 6. The method of claim 5 wherein a spoiled gradient echo sequence is used to generate a further magnetic resonance image and the images from the spoiled and non-spoiled gradient echo sequences are subtracted to produce an image in which the signal difference between arteries and veins is enhanced. 7. The method of claim 1 wherein said contrast agent is administered in a quantity such that R2* of venous blood is less than R2* of arterial blood. | Amersham Health AS | A61B005/05
US06901283 | B2 | 2005-05-31 | Adjusting a condition of a fluid medium to produce an image of a patient | 2002-03-04 | A system for producing an image of a patient preferably includes a source of contrast medium, a pressurizing unit operably associated with the source of contrast medium for injecting the contrast medium into the patient, an imaging unit providing an internal image of the patient based upon a condition of the contrast medium in the patient, a control unit adapted to adjust the condition of the contrast medium injected into the patient, an electronic interface in communication with at least one of the pressurizing unit and the imaging unit, and a fluid path disposed between the pressurizing unit and the patient. | 1. A system for producing an image of a patient, comprising: a source of fluid medium; a pressurizing unit configured to increase pressure of the fluid medium being injected into the patient; an imaging unit providing an internal image of the patient based upon a condition of the fluid medium in the patient; an electronic interface that enables information input and output; a control unit adapted to adjust the condition of the fluid medium injected into the patient, the control unit being communicatively coupled to the electronic interface and the pressurizing unit; and a hospital information system that provides infomation about the patient, wherein the electronic interface is configured to communicate with the pressurizing unit, the imaging unit, and the hospital information system. | 1. A system for producing an image of a patient, comprising: a source of fluid medium; a pressurizing unit configured to increase pressure of the fluid medium being injected into the patient; an imaging unit providing an internal image of the patient based upon a condition of the fluid medium in the patient; an electronic interface that enables information input and output; a control unit adapted to adjust the condition of the fluid medium injected into the patient, the control unit being communicatively coupled to the electronic interface and the pressurizing unit; and a hospital information system that provides infomation about the patient, wherein the electronic interface is configured to communicate with the pressurizing unit, the imaging unit, and the hospital information system. 2. The system of claim 1 wherein the fluid medium is a contrast medium is a contrast medium. 3. The system of claim 2 wherein the contrast medium is ultrasound contrast. 4. The system of claim 1 wherein the condition of the fluid medium in the patient corresponds to at least one parameter selected from the group of concentration, flow rate and volume. 5. The system of claim 1 wherein the condition of the fluid medium is concentration. 6. The system of claim 1 wherein the pressurizing unit is adapted to inject at least one bolus of the fluid medium into the patient. 7. The system of claim 1, further comprising a concentration measurement device operable to measure the concentration of the fluid medium. 8. The system of claim 1, further comprising a source of diluent. 9. The system of claim 8, further comprising a mixer for mixing the fluid medium and the diluent. 10. The system of claim 8, further comprising a pump to configured to draw the diluent from the diluent source. 11. The system of claim 1 wherein the pressurizing unit comprises a pump. 12. The system of claim 1, further comprising a fluid path connected to the pressurizing unit for delivering the fluid medium to the patient. 13. The system of claim 12, further comprising a filter disposed in the fluid path. 14. The system of claim 12, further comprising a back-flow valve disposed in the fluid path. 15. The system of claim 12, further comprising a fluid assurance detector in the fluid path. 16. The system of claim 1, further comprising a heating element to adjust the temperature of the fluid medium injected into the patient. 17. The system of claim 1, further comprising a pump to configured to draw the fluid medium from the fluid medium source. 18. A system for producing an image of a patient, comprising: a source of fluid medium; a pressurizing unit configured to increase pressure of the fluid medium being injected into the patient; an imaging unit providing an internal image of the patient based upon a condition of the fluid medium in the patient; an electronic interface that enables information input and output; a control unit adapted to adjust the condition of the fluid medium injected ... | Medrad, Inc. | A61B005/00
US06908443 | B2 | 2005-06-21 | Intravascular guide wire and method for manufacture thereof | 2001-11-20 | A guide wire, and a method for the manufacture thereof, having a core and a plastic jacket enclosing at least a portion of the core. In some embodiments, the guidewire also includes a radiopaque coil disposed about a first portion of the core wire. The polymer jacket includes a first portion disposed about the radiopaque coil, and a second portion adjacent the first portion and overlaying and in contact with a second portion of the core wire. In some embodiments, the plastic jacket comprises a proximal portion formed of a first plastic material and a distal jacket portion formed of a second plastic material. | 1. A guide wire comprising: an elongated core having a proximal region and a distal region, the distal region of the core having a proximal portion and a distal portion; a radiopaque coil disposed about the distal portion of the distal region of the core; a plastic jacket disposed about the distal region of the core and the radiopapue coil, the plastic jacket being in intimate contact with the proximal portion of the distal region of the core alone a length thereof; and wherein a distal end of the radiopaque coil is connected to a distal end of the core to form a connection, and wherein the plastic jacket encapsulates the connection. | 1. A guide wire comprising: an elongated core having a proximal region and a distal region, the distal region of the core having a proximal portion and a distal portion; a radiopaque coil disposed about the distal portion of the distal region of the core; a plastic jacket disposed about the distal region of the core and the radiopapue coil, the plastic jacket being in intimate contact with the proximal portion of the distal region of the core alone a length thereof; and wherein a distal end of the radiopaque coil is connected to a distal end of the core to form a connection, and wherein the plastic jacket encapsulates the connection. 2. A guidewire as in claim 1, wherein the distal portion of the core has a relatively flat cross-sectional shape. 3. A guidewire as in claim 2, wherein the proximal of the core has a relatively round cross-sectional shape. 4. A guide wire comprising: an elongated core having a proximal region and a distal region, the distal region of the core having a proximal portion and a distal portion, the distal portion of the distal region having a relatively flat cross-sectional shape and the proximal portion of the distal region of the core having a relatively round cross-sectional shape; a radiopaque coil disposed about the distal portion of the distal region of the core, with a distal end of the radiopaque coil connected to a distal end of the core to form a junction; and a plastic jacket having a smooth outer surface disposed about the proximal portion of the distal region of the core, the radiopaque coil, and the junction, the plastic jacket being in intimate contact with the radiopaque coil and the proximal portion of the distal region of the core. 5. A guide wire comprising: an elongated core wire including a first portion and a second portion disposed adjacent the first portion; a radiopaque coil disposed about the first portion of the core wire; a polymer jacket including a first section disposed about the radiopaque coil, and a second section adjacent the first section and overlaying and in contact with the second portion of the core wire. 6. A guide wire comprising: an elongated body assembly including a first portion and a second portion disposed adjacent the first portion, the body assembly including an elongated core wire extending through both the first and second portions; the first portion including a radiopaque coil disposed about the core wire, and a first section of a polymer jacket disposed about the radiopaque coil; the second portion including a second section of the polymer jacket overlaying and in contact with the core wire. | SciMed Life Systems, Inc. | A61B005/00 A61M025/00
US06928311 | B1 | 2005-08-09 | Compact device for measuring, tissue analytes | 2000-08-31 | A compact device (20) for non-invasively monitoring concentration levels of blood constituents, including glucose, cholesterol, alcohol, blood gases and various ions. The device includes a finger receptor (140) having a channel for receiving a finger of a user. The channel has a light entrance and a light exit so that light can be passed from a light source (91) through a finger located in the channel in a direction generally normal to the finger. Certain heat generating components, including a stable power supply for the device, are external to the device housing so as to reduce heat generation and thereby increase stability of the device. The device includes a communications interface for interacting with a computer. The device can be used for clinical use or for home use and the memory of the computer can be used to assist with record keeping and with dosage calculations. | 1. A measuring device for non-invasively measuring levels of constituents in blood and tissue in a subject, the measuring device comprising: (a) a housing having an opening for receiving a hand of the subject;(b) a polychromatic light source that emits a broad spectrum of light in the near infrared range and adjacent visible light;(c) a part receptor shaped for receiving a finger of the subject, the part receptor being located relative to the light source so that when the finger of the subject is placed in the part receptor, the light source can be activated and light from the light source can be directed onto the finger, the part receptor being shaped to receive the finger in close alignment, so as to reduce extraneous light;(d) a light receptor for collecting a continuum of wavelengths over the broad spectrum of light after the broad spectrum of light has been directed onto the finger;(e) a dispersion element coupled to the light receptor for dispersing light collected by the light receptor into a dispersed spectrum of component wavelengths;(f) a photodetector coupled to the dispersion element for taking absorbance measurements from the dispersed spectrum and producing a measurement signal;(g) a communications interface connectable to an external computer for communicating the measurement signal to the external computer, and(h) a power interface connectable to an external stabilized power source. | 1. A measuring device for non-invasively measuring levels of constituents in blood and tissue in a subject, the measuring device comprising: (a) a housing having an opening for receiving a hand of the subject;(b) a polychromatic light source that emits a broad spectrum of light in the near infrared range and adjacent visible light;(c) a part receptor shaped for receiving a finger of the subject, the part receptor being located relative to the light source so that when the finger of the subject is placed in the part receptor, the light source can be activated and light from the light source can be directed onto the finger, the part receptor being shaped to receive the finger in close alignment, so as to reduce extraneous light;(d) a light receptor for collecting a continuum of wavelengths over the broad spectrum of light after the broad spectrum of light has been directed onto the finger;(e) a dispersion element coupled to the light receptor for dispersing light collected by the light receptor into a dispersed spectrum of component wavelengths;(f) a photodetector coupled to the dispersion element for taking absorbance measurements from the dispersed spectrum and producing a measurement signal;(g) a communications interface connectable to an external computer for communicating the measurement signal to the external computer, and(h) a power interface connectable to an external stabilized power source. 2. The device of claim 1, further comprising a hand support at the housing opening, the hand support being adjustable to vary the size of the opening. 3. The device of claim 2, wherein the hand support receives the palm of a human hand and the top of the opening is curved to generally fit the profile of the human hand across the top of the hand. 4. The device of claim 1, wherein the polychromatic light source is connected to the power interface. 5. The device of claim 1, further comprising the external computer, wherein the external computer controls at least one function of the measuring device, the computer including means for receiving the measurement signal. 6. The device of claim 5, further comprising an analog to digital converter for converting the measurement signal into a digital measurement signal for communication to the computer. 7. The device of claim 5, wherein, the external computer includes a memory element, a storage element, and a software element for storing a plurality of measurement signals for a plurality of measurements. 8. The device of claim 5, wherein, sad the external computer includes a memory element, a storage element, and a software element for storing, retrieving and displaying dosage information corresponding to measurement signals received by the computer from the device. 9. The device of claim 5, further comprising the external stabilized power source. 10. The device of claim 9, wherein the external stabilized power source is provided by the external computer. 11. The device of claim 1, further comprising th... | NIR Diagnostics Inc. | A61B005/00
US07041057 | B1 | 2006-05-09 | Tissue interface device | 2000-11-17 | A tissue interface device (10) suitable for positioning on or about one or more artificial openings in a biological membrane of an organism and for coupling to a monitor and control unit and a vacuum source. The tissue interface device (10) comprises a housing (100), a sensor channel (130), and a sensor (150). The housing (100) defines an orifice (120), the orifice (120) having an open inlet port (122) on the bottom end (102) of the housing (100) and a distal end (124) that is in fluid communication with the sensor channel (130). The orifice (120) is in fluid communication with fluid that flows from the artificial opening formed in the biological membranes. The sensor channel (130) is for coupling to, and fluid communication therewith, the vacuum source. The sensor (150) is positioned in the sensor channel (130) in a flow path of the fluid for sensing a characteristic of the fluid as it flows out from the artificial opening. The sensor generates a sensor signal representative thereof. | 1. A tissue interface device for positioning about an artificial opening in a biological membrane to measure tissue characteristics, comprising: a. a housing having a bottom end,b. at least one sensor, andc. a sensor channel, wherein, within the housing, the sensor channel couples the bottom end with the sensord. wherein the housing defines an orifice extending into the housing, the orifice having an open inlet port on the bottom end of the housing to receive liquid and a distal end that is in fluid communication with the sensor channel and wherein a portion of the bottom end of the orifice is radiused. | 1. A tissue interface device for positioning about an artificial opening in a biological membrane to measure tissue characteristics, comprising: a. a housing having a bottom end,b. at least one sensor, andc. a sensor channel, wherein, within the housing, the sensor channel couples the bottom end with the sensord. wherein the housing defines an orifice extending into the housing, the orifice having an open inlet port on the bottom end of the housing to receive liquid and a distal end that is in fluid communication with the sensor channel and wherein a portion of the bottom end of the orifice is radiused. 2. The tissue interface device of claim 1, wherein the sensor is positioned in the sensor channel that is in fluid communication with the sensor channel for sensing a characteristic of fluid flowing from the artificial opening in the biological membrane, and whereby, in use, the orifice is in fluid communication with the artificial opening in the biological membrane. 3. The tissue interface device of claim 1, wherein the inlet port has a first diameter and the distal end has a second diameter. 4. The tissue interface device of claim 3, wherein the orifice is tapered and wherein the second diameter is less than the first diameter. 5. The tissue interface device of claim 4, wherein the orifice has a bell-shaped curve in cross-section. 6. The tissue interface device of claim 4, wherein the orifice has a shaped orifice surface that extends from the inlet port to the distal end of the orifice, and wherein at least a portion of the orifice surface is curved in cross-section. 7. The tissue interface device of claim 1, wherein the sensor channel has an outlet port to discharge fluid. 8. The tissue interface device of claim 7, wherein the outlet port is suitable for connection to a vacuum source. 9. The tissue interface device of claim 1, wherein the sensor generates a signal representing a characteristic of the biological fluid. 10. The tissue interface device of claim 1, wherein the sensor generates an electrical sensor signal representing a characteristic of the biological fluid. 11. The tissue interface device of claim 10, further comprising an electrical connector and an electrical lead line in operative communication with the sensor, the electrical lead line coupling the electrical sensor signal to the electrical connector. 12. The tissue interface device of claim 11, wherein the electrical connector is coupled to a monitor and control unit. 13. The tissue interface device of claim 1, wherein the housing is disposable. 14. The tissue interface device of claim 1, wherein, in use, at least a portion of the bottom end of the housing is positioned in sealed contact with the biological membrane around the artificial opening. 15. The tissue interface device of claim 1 wherein the housing has a notch extending therein, wherein the housing further comprises a flexible insert having a base surface, the flexible insert shaped to be complementaril... | SpectRx, Inc. | A61B005/00 A61B005/05 A61B005/04
US07244232 | B2 | 2007-07-17 | Process for identifying cancerous and/or metastatic cells of a living organism | 2002-04-24 | A process for identifying and treating cells in a living organism. The cells are labeled, circulated within the organism, detected with an implanted detector, and then either isolated or ablated. | 1. A process for identifying cancerous and/or metastatic cells of a living organism, comprising the steps of: feeding a fluid stream comprising said cancerous and/or metastatic cells, to a first assembly comprising a reservoir, wherein said reservoir contains markers capable of labeling the cancerous and/or metastatic cells;labeling at least a portion of said cells with said markers so as to produce labeled cells;feeding said fluid stream comprising said labeled cells to a second assembly comprising a particle analyzer, wherein said particle analyzer comprises a stent, wherein a plurality of optical assemblies capable of emitting and detecting radiation are disposed on the inner surface of said stent; thendetecting said labeled cancerous and/or metastatic cells with said particle analyzer. | 1. A process for identifying cancerous and/or metastatic cells of a living organism, comprising the steps of: feeding a fluid stream comprising said cancerous and/or metastatic cells, to a first assembly comprising a reservoir, wherein said reservoir contains markers capable of labeling the cancerous and/or metastatic cells;labeling at least a portion of said cells with said markers so as to produce labeled cells;feeding said fluid stream comprising said labeled cells to a second assembly comprising a particle analyzer, wherein said particle analyzer comprises a stent, wherein a plurality of optical assemblies capable of emitting and detecting radiation are disposed on the inner surface of said stent; thendetecting said labeled cancerous and/or metastatic cells with said particle analyzer. 2. The process as recited in claim 1 further comprising the step of estimating the positions of said labeled cells in or downstream from the stent by means of a three-dimensional, multivariate particle velocity and location estimation algorithm. 3. The process as recited in claim 1, wherein said particle analyzer further comprises a flow cytometer. 4. The process as recited in claim 1, further comprising the step of stripping said markers from the labeled cells within the fluid stream after exiting the particle analyzer, so as to produce a purified fluid stream. 5. The process as recited in claim 1, wherein said marker is a non-toxic fluorescent marker. 6. The process as recited in claim 5, wherein said marker is green fluorescent protein. 7. The process as recited in claim 1, wherein said first and second assemblies are implanted within said living organism. 8. A process for identifying and treating cancerous and/or metastatic cells of a living organism, comprising conducting the process of claim 1, and further comprising ablating said labeled cells, which have been detected by the particle analyzer, with an ablation device. 9. The process as recited in claim 8, wherein the ablation device is positioned downstream from the particle analyzer. 10. The process as recited in claim 9, wherein the distance between the particle analyzer and the ablation device is from about 100 microns to 2 millimeters. 11. The process of claim 8, wherein the ablation device is a part of the stent of the particle analyzer. 12. The process as recited in claim 8, wherein the ablation device comprises a plurality of lasers. | Biomed Solutions, LLC | A61B005/1455
US07039451 | B1 | 2006-05-02 | Phase ordering with automatic window selection (PAWS) for motion resistant MRI | 2000-08-18 | A magnetic imaging technique which is resistant to changes in breathing while allowing the use of phase ordering to provide effective motion artifact reduction in an optimal time. This is provided by apparatus for magnetic resonance imaging a target object subject to periodic motion, comprising a magnetic resonance imaging scanner for exciting said target object and recovering imaging data in k-space; a sensor for detecting a signal indicative of a position of said target object; classifying logic for classifying said at least one line of imaging data into one of a plurality of groups of lines of imaging date in dependence upon said position detected by said sensor as said target object was excited, each group of lines corresponding to one of a plurality of contiguous ranges of position and scan terminating logic for detecting when two or more groups of lines corresponding to contiguous ranges of position together containing a set of lines of imaging date spanning k-space from which an image can be derived. | 1. Apparatus for magnetic resonance imaging a target object, said apparatus comprising: a magnetic resonance imaging scanner for exciting said target object and recovering at least one line of imaging data in k-space;a first sensor for detecting a signal indicative of a position of said target object relative to said magnetic resonance imaging scanner;classifying logic for classifying said at least one line of imaging data into one of a plurality of groups of lines of imaging data in dependence upon said position detected by said first sensor as said target object was excited, each group of lines corresponding to one of a plurality of contiguous ranges of position relative to said magnetic resonance imaging scanner; andscan terminating logic for detecting when two or more groups of lines corresponding to contiguous ranges of position together contain a set of lines of imaging data spanning k-space from which an image can be derived and terminating data acquisition. | 1. Apparatus for magnetic resonance imaging a target object, said apparatus comprising: a magnetic resonance imaging scanner for exciting said target object and recovering at least one line of imaging data in k-space;a first sensor for detecting a signal indicative of a position of said target object relative to said magnetic resonance imaging scanner;classifying logic for classifying said at least one line of imaging data into one of a plurality of groups of lines of imaging data in dependence upon said position detected by said first sensor as said target object was excited, each group of lines corresponding to one of a plurality of contiguous ranges of position relative to said magnetic resonance imaging scanner; andscan terminating logic for detecting when two or more groups of lines corresponding to contiguous ranges of position together contain a set of lines of imaging data spanning k-space from which an image can be derived and terminating data acquisition. 2. Apparatus as claimed in claim 1, wherein said target object is subject to periodic motion, further comprising a second sensor for detecting when said target object is at a predetermined state within said periodic motion, and wherein said magnetic resonance imaging scanner is responsive to said second sensor detecting said target object is at said predetermined state to trigger said excitation. 3. Apparatus as claimed in claim 1, wherein said target object is an in vivo organ. 4. Apparatus as claimed in claim 3, wherein said target object is a heart. 5. Apparatus as claimed in claim 4, wherein said second sensor is an ECG sensor. 6. Apparatus as claimed in claim 1, wherein said first sensor detects changes in position of said target object due to respiration. 7. Apparatus as claimed in claim 6, wherein said first sensor is a magnetic resonance sensor for sensing diaphragm position. 8. Apparatus as claimed in claim 1, wherein when a line of imaging data can be selected to be acquired on either side of k-space corresponding to two different groups of lines, said line is acquired in that group closest to forming one of two or more groups of lines corresponding to contiguous ranges of position together containing a set of lines of imaging data spanning k-space from which an image can be derived. 9. Apparatus as claimed in claim 1, wherein at least one group of lines close to a central position detected by said first sensor corresponds to a smaller range of positions than a group of lines further from said central position. 10. A method of magnetic resonance imaging a target object, said method comprising the steps of: exciting said target object and recovering at least one line of imaging data in k-space using a magnetic resonance imaging scanner;detecting a signal indicative of a position of said target object relative to said magnetic resonance imaging scanner;classifying said at least one line of imaging data into one of a plurality of groups of lines of imaging data in depen... | Imperial College Innovations Limited | A61B005/05 G01V003/00
US07115098 | B2 | 2006-10-03 | Adjustable thermal scanning system and method | 2002-06-14 | A thermal scanning system and method. The thermal scanning system comprises a handheld thermal scanning device with wheels and three infrared sensors that are spaced to scan skin temperature on opposite sides of patient spinal columns of different sizes. Two out of the three sensors are selected for a particular patient, and the wheels are adjustable to accommodate the two selected sensors. The handheld device further calculates the distance traveled by the wheels by measuring the rotation of the wheels. The device has a trigger that allows a user to annotate the location of various anatomical landmarks, such as particular regions of a patient's spinal column. The device is adapted to take discrete temperature readings or rolling temperature readings. The thermal scanning system provides graphical displays and data analysis of the temperature readings from the sensors. | 1. A thermal scanning system adapted to scan skin temperature on opposite sides of a patient's spinal column for patients of different sizes, the thermal scanning system comprising: a handheld body having a portion which is adapted to be gripped by a user at least when scanning;a plurality of temperature sensors oriented in a fixed position relative to said body, including a first sensor, a second sensor and a third sensor, wherein the sensors are mounted on the body and arranged in a linear fashion such that the first and second sensors are spaced to scan skin temperature on opposite sides of a patient's spinal column of a first size range, the second and third sensors are spaced to scan skin temperature on opposite sides of a patient's spinal column of a second size range, and the first and third sensors are spaced to scan skin temperature on opposite sides of a patient's spinal column of a third size range, wherein the size range defined between any two of the three temperature sensors is different from the size range between any other combination of two sensors, each sensor producing a data signal indicative of sensed temperatures;a computer receiving and processing the data signals received from two or more of the sensors; anda selector defining at least first, second and third operational modes for the thermal scanning system, wherein the computer in the first operational mode uses the data signals received from the first and second sensors to determine the skin temperature on opposite sides of a patient's spinal column, wherein the computer in the second operational mode uses the data signals received from the first and third sensors to determine the skin temperature on opposite sides of a patient's spinal column, and wherein the computer in the third operational mode uses the data signals received from the second and third sensors to determine the skin temperature on opposite sides of a patient's spinal column. | 1. A thermal scanning system adapted to scan skin temperature on opposite sides of a patient's spinal column for patients of different sizes, the thermal scanning system comprising: a handheld body having a portion which is adapted to be gripped by a user at least when scanning;a plurality of temperature sensors oriented in a fixed position relative to said body, including a first sensor, a second sensor and a third sensor, wherein the sensors are mounted on the body and arranged in a linear fashion such that the first and second sensors are spaced to scan skin temperature on opposite sides of a patient's spinal column of a first size range, the second and third sensors are spaced to scan skin temperature on opposite sides of a patient's spinal column of a second size range, and the first and third sensors are spaced to scan skin temperature on opposite sides of a patient's spinal column of a third size range, wherein the size range defined between any two of the three temperature sensors is different from the size range between any other combination of two sensors, each sensor producing a data signal indicative of sensed temperatures;a computer receiving and processing the data signals received from two or more of the sensors; anda selector defining at least first, second and third operational modes for the thermal scanning system, wherein the computer in the first operational mode uses the data signals received from the first and second sensors to determine the skin temperature on opposite sides of a patient's spinal column, wherein the computer in the second operational mode uses the data signals received from the first and third sensors to determine the skin temperature on opposite sides of a patient's spinal column, and wherein the computer in the third operational mode uses the data signals received from the second and third sensors to determine the skin temperature on opposite sides of a patient's spinal column. 2. The system of claim 1, wherein the first size range relates to an infant, the second size range relates to a child and the third size range relates to an adult. 3. The system of claim 1, wherein the sensors scan at least a first point and a second point, the points being located on opposite sides of a patient's spinal column or vertebra selected by the user. 4. The system of claim 1, wherein the sensors scan a plurality of points on opposite sides of the patient's spinal column. 5. The system of claim 1, further comprising a display device, the display device displaying a first temperature reading from a point located on a left side of a patient's vertebra and a second temperature reading from a point located on a right side of the patient's vertebra. 6. The system of claim 1, further comprising a display device, the display device displaying a first comparison value and a second comparison value, the first comparison value comprising a first temperature reading from a point located on a left side of a patient's vertebra ad... | - | A61B005/00
US07113814 | B2 | 2006-09-26 | Tissue interrogation spectroscopy | 2001-07-13 | In an emergency medicine patient, accurate measurement of change or lack thereof from non-shock, non-ischemic, non-inflammation, non-tissue injury, non-immune dysfunction conditions is important and is provided, as practical, real-time approaches for accurately characterizing a patient's condition, using Raman (3) and/or fluorescence (30) spectroscopy with a high degree of accuracy. Measurement times are on the order of seconds. High-accuracy measurement is achieved with Raman spectroscopy interrogation of tissue. Simultaneous interrogation by NADH fluorescence spectroscopy may he used. Measurements may be non-invasive to minimally invasive. Preclinical (ultra-early) states of shock can be detected (5), severity can be determined, effectiveness of various treatments can be determined. | 1. A tissue analysis method, comprising: interrogating a biological tissue with Raman spectroscopy and fluorescence spectroscopy to obtain spectroscopy results; analyzing the obtained spectroscopy results for at least one mediator or marker associated with a shock state. | 1. A tissue analysis method, comprising: interrogating a biological tissue with Raman spectroscopy and fluorescence spectroscopy to obtain spectroscopy results; analyzing the obtained spectroscopy results for at least one mediator or marker associated with a shock state. 2. The method of claim 1, wherein the tissue interrogating is noninvasive. 3. The method of claim 1, wherein the tissue is in vivo and in situ. 4. The method of claim 1, wherein the tissue is removed from a patient before the tissue interrogation. 5. The method of claim 1, including measuring NADH presence and/or accumulation by fluorescence spectroscopy. 6. The method of claim 1, including measuring tissue hemoglobin oxygen saturation by Raman spectroscopy and measuring NADH presence and/or accumulation by fluorescence spectroscopy. 7. The method of claim 1, including determining whether the tissue has insufficient oxygen delivery to meet metabolic demands of the tissue while simultaneously determining whether mitochondrial dysfunction or injury exists. 8. The method of claim 1, including measuring myoglobin oxygenation saturation. 9. The method of claim 1, including determining cytochrome oxidase redox status. 10. The method of claim 1, including determining absolute concentration of hemoglobin in the tissue. 11. The method of claim 1, including determining pH of the tissue. 12. The method of claim 1, including intermittently or continuously interrogating the tissue of a patient. 13. The method of claim 1, including determining tissue viability. 14. The method of claim 1, including diagnosing shock. 15. A tissue analysis method, comprising: interrogating a biological tissue with Raman spectroscopy and fluorescence spectroscopy to obtain spectroscopy results, including:measuring tissue hemoglobin oxygen saturation including amount of oxyhemoglobin and deoxyhemoglobin by Raman spectroscopy; and/ormeasuring tissue hemoglobin oxygen saturation by Raman spectroscopy and measuring NADH presence and/or accumulation by fluorescence spectroscopy; and/ordetermining whether the tissue has insufficient oxygen delivery to meet metabolic demands of the tissue while simultaneously determining whether mitochondrial dysfunction or injury exists. 16. The tissue analysis method of claim 15, wherein the obtained spectroscopy results are for at least one mediator or marker associated with a shock state and/or tissue injury. 17. The method of claim 16, wherein the obtained spectroscopy results are for presence and/or proportions for the at least one shock state and/or tissue injury mediator or marker. 18. The method of claim 16 including determining the concentration of at least one mediator or marker. 19. The method of claim 18, including determining absolute concentration. 20. The method of claim 18, including determining relative concentration. 21. The method of claim 16 including determining the presence of at least one mediator or marker. 22. The method of claim 16, wherein... | Virginia Commonwealth University | A61B005/00
US06983181 | B2 | 2006-01-03 | Spatial encoding MR data of a moving subject using a higher-order gradient field | 2002-05-01 | The present invention includes a technique for use with magnetic resonance imaging that includes the application of a non-linear, higher-order gradient field in the presence of a moving object to be scanned. MR data is acquired as the object moves through the non-linear gradient field. Resulting images are contiguous and do not require the patching together of data in either k-space or image space and result in an image with expanded FOV in a longitudinal direction of the moving object. | 1. A method of acquiring spatially encoded magnetic resonance (MR) data comprising the steps of: applying a gradient field having an exponential relation higher than a single order along an object direction;moving an object to be scanned in the object direction through the gradient field; andacquiring MR data as the object moves with contiguous continuity in the object direction. | 1. A method of acquiring spatially encoded magnetic resonance (MR) data comprising the steps of: applying a gradient field having an exponential relation higher than a single order along an object direction;moving an object to be scanned in the object direction through the gradient field; andacquiring MR data as the object moves with contiguous continuity in the object direction. 2. The method of claim 1 wherein the gradient field is applied in a Z-direction and is a pulsed phase encoding gradient having a constant value such that at any given Z-location, a value of the gradient field is a linear function of the given Z-location. 3. The method of claim 2 wherein the object is moved in the Z-direction on successive TRs to acquire a full range of kz data. 4. The method of claim 1 wherein the gradient field is a pulsed Z2 gradient field. 5. The method of claim 1 further comprising: phase correcting acquired MR data along the object direction;Fourier transforming the phase corrected MR data; andreconstructing an image with spatial encoding in the object direction. 6. An magnetic resonance (MR) apparatus to acquire continuous imaging of a moving object over an extended field-of-view (FOV) comprising: a magnetic resonance imaging (MRI) system having a plurality of gradient coils positioned about a bore of a magnet to impress a polarizing magnetic field and an RF transceiver system and an RF switch controlled by a pulse module to transmit RF signals to an RF coil assembly to acquire MR images, wherein at least one of the gradient coils is constructed to imparta higher-order gradient field;a table moveable along a longitudinal axis into and out of the bore of the magnet; anda computer programmed to: transmit a series of RF pulses;apply a higher-order gradient field between each RF pulse in the series of RF pulses; andmove the table and acquire data. 7. The MR apparatus of claim 6 wherein the computer is further programmed to reconstruct a continuous image from the data acquired after each periodic table movement. 8. The MR apparatus of claim 6 wherein the computer is further programmed to step-wise move the table a given increment based on desired Kz encoding resolution. 9. The MR apparatus of claim 6 wherein the computer is further programmed to apply a higher-order gradient of consistent value in each TR. 10. The MR apparatus of claim 6 wherein the computer is further programmed to reconstruct a seamless image of length greater than that of a default FOV for the MRI system without piecing together raw data or multiple out-images. 11. A method of acquiring magnetic resonance (MR) images with expanded FOV in a longitudinal direction of an object comprising the steps of: defining a direction of motion of an object to be scanned;arranging the object such that a longitudinal axis is positioned along the direction of motion;applying a non-linear gradient field in the direction of motion;moving the object through the applied non-linear gradient fiel... | General Electric Company | A61B005/55
US07254946 | B1 | 2007-08-14 | Thermodilution catheter having a safe, flexible heating element | 1995-04-12 | A thermodilution catheter having a heating filament which is fabricated so as to be thin and flexible enough to avoid contact with the patient's blood. The heating filament is either inserted in a preformed catheter lumen, incorporated into a wall of the catheter body itself, or wrapped around the catheter body wall and surrounded by an external sheath. Generally, the covering of the heating filament is minimally thin so as to allow the heat from the heating filament to be transferred to the surrounding blood and to minimally increase the overall cross-sectional area. Since the heating filament does not directly touch the patient's blood, the outer surface may be made smooth so as to prevent inducement of blood clots. In addition, the heating filament may be maintained at a safe temperature by forming the heating element of a flexible material having a high temperature coefficient of resistance, low thermal capacitance and high thermal conductivity. Preferably, the temperature coefficient of resistance of the material forming the heater filament is greater than 0.001 Ω/Ω-° C. and the resistance of the heating filament is proportional or inversely proportional to its temperature. By monitoring this resistance, the application of power to the heating filament may be regulated to prevent overheating of the heating filament. Calibration techniques are also disclosed whereby a memory containing encoded calibration information is disposed within or connected to a connector of the catheter. This memory may also include a program segment used by a cardiac output computer having a cardiac output calculation program, whereby the portion of code stored in the memory must be input into the cardiac output computer before the determination of the cardiac output can be performed. As a result of these improvements, the invention enables the thermodilution cardiac output calculation to be performed quite safely and accurately in a clinical setting. | 1. A multi-lumen, multi-purpose cardiac catheter comprising: (a) a multi-lumen main body portion;(b) a plurality of extension tubes, each one of said plurality of extension tubes being connected to a respective lumen of said multi-lumen main body portion; and(c) an interface connecting said multi-lumen main body portion and said plurality of extension tubes,wherein:(d) said multi-lumen main body portion comprises: (i) at least one lumen for holding and supporting fiber optic filaments;(ii) at least one lumen for receiving thermal element connectors;(iii) at least one lumen for receiving a device for temperature measurement;(iv) at least one lumen associated with a balloon mounted at the distal end of said multi-lumen main body portion for assisting in placement of said multi-lumen main body portion in a patient;(v) a port defined by surfaces of one of the lumens of said multi-lumen main body portion, said port for injecting a fluid into a blood stream of a patient;(vi) a fiber optic apparatus including said fiber optic filaments, said fiber optic filaments disposed in said at least one lumen for holding and supporting said fiber optic filaments, and said fiber optic filaments extending from inside said multi-lumen main body portion to a fiber optic coupler associated with the catheter; and(vii) a necked-down portion near the distal end of said multi-lumen main body portion;(e) a temperature measurement apparatus is mounted at the distal end of said multi-lumen main body portion;(f) wiring extends from the temperature measurement apparatus along said multi-lumen main body portion to a housing associated with said multi-lumen main body portion;(g) an external thermal element is mounted on said multi-lumen main body portion near the distal end of said multi-lumen main body portion; and(h) connectors extend from said external thermal element along said multi-lumen main body portion for connection to a thermal element housing at the proximal end of said multi-lumen main body portion,wherein:(i) said external thermal element and said temperature measurement apparatus are operative with an external apparatus for providing a measurement of continuous cardiac output of a patient;(j) said fiber optic apparatus is cooperative with said external apparatus for providing a measurement of mixed venous oxygen saturation of blood of the patient; and(k) said external thermal element is mounted on said necked-down portion. | 1. A multi-lumen, multi-purpose cardiac catheter comprising: (a) a multi-lumen main body portion;(b) a plurality of extension tubes, each one of said plurality of extension tubes being connected to a respective lumen of said multi-lumen main body portion; and(c) an interface connecting said multi-lumen main body portion and said plurality of extension tubes,wherein:(d) said multi-lumen main body portion comprises: (i) at least one lumen for holding and supporting fiber optic filaments;(ii) at least one lumen for receiving thermal element connectors;(iii) at least one lumen for receiving a device for temperature measurement;(iv) at least one lumen associated with a balloon mounted at the distal end of said multi-lumen main body portion for assisting in placement of said multi-lumen main body portion in a patient;(v) a port defined by surfaces of one of the lumens of said multi-lumen main body portion, said port for injecting a fluid into a blood stream of a patient;(vi) a fiber optic apparatus including said fiber optic filaments, said fiber optic filaments disposed in said at least one lumen for holding and supporting said fiber optic filaments, and said fiber optic filaments extending from inside said multi-lumen main body portion to a fiber optic coupler associated with the catheter; and(vii) a necked-down portion near the distal end of said multi-lumen main body portion;(e) a temperature measurement apparatus is mounted at the distal end of said multi-lumen main body portion;(f) wiring extends from the temperature measurement apparatus along said multi-lumen main body portion to a housing associated with said multi-lumen main body portion;(g) an external thermal element is mounted on said multi-lumen main body portion near the distal end of said multi-lumen main body portion; and(h) connectors extend from said external thermal element along said multi-lumen main body portion for connection to a thermal element housing at the proximal end of said multi-lumen main body portion,wherein:(i) said external thermal element and said temperature measurement apparatus are operative with an external apparatus for providing a measurement of continuous cardiac output of a patient;(j) said fiber optic apparatus is cooperative with said external apparatus for providing a measurement of mixed venous oxygen saturation of blood of the patient; and(k) said external thermal element is mounted on said necked-down portion. 2. The multi-lumen, multi-purpose cardiac catheter of claim 1, wherein: (a) said port is distal said interface and(b) said lumen having the surfaces defining said port is an injectate lumen and said injectate lumen and said port enable injection of an injectate fluid into the blood stream of the patient. 3. The multi-lumen, multi-purpose cardiac catheter of claim 1, wherein said external thermal element comprises a thin film member spirally wound about said multi-lumen main body portion at approximately fourteen centimeters from the distal end ... | Edwards Lifesciences Corporation | A61B005/02
US06984228 | B2 | 2006-01-10 | Phototherapy methods and systems | 2002-08-20 | The invention features methods and systems for treating inflammatory, proliferative skin disorders, such as psoriasis, with ultraviolet phototherapy. The methods and systems use optical techniques to scan a patient's skin, designate areas of affected skin, and selectively deliver high doses of phototherapeutic ultraviolet radiation to the designated areas. To insure that only affected areas of skin affected are designated for the high doses of UV radiation, the methods and systems use one or more optical diagnostics that relate to independent physiological features of affected skin. | 1. A system for treating a proliferative skin disorder in a patient, the system comprising: an illumination system configured to irradiate the patient's skin;a detector system configured to receive radiation emitted from selected areas of the patient's skin in response to the irradiation from the illumination source and generate at least two diagnostic signals based on the received radiation for each of the selected areas, wherein each diagnostic signal provides a different physiological indication of a presence of the skin disorder;an analyzer system connected to the detector system, wherein for each selected area, the analyzer system is configured to (i) separately evaluate whether each diagnostic signal indicates the presence of the skin disorder at the selected area, and (ii) determine whether the selected area should be designated as an affected area based on the separate evaluations of the diagnostic signals; anda therapeutic source configured to deliver doses of therapeutic radiation to the affected areas of skin designated by the analyzer system. | 1. A system for treating a proliferative skin disorder in a patient, the system comprising: an illumination system configured to irradiate the patient's skin;a detector system configured to receive radiation emitted from selected areas of the patient's skin in response to the irradiation from the illumination source and generate at least two diagnostic signals based on the received radiation for each of the selected areas, wherein each diagnostic signal provides a different physiological indication of a presence of the skin disorder;an analyzer system connected to the detector system, wherein for each selected area, the analyzer system is configured to (i) separately evaluate whether each diagnostic signal indicates the presence of the skin disorder at the selected area, and (ii) determine whether the selected area should be designated as an affected area based on the separate evaluations of the diagnostic signals; anda therapeutic source configured to deliver doses of therapeutic radiation to the affected areas of skin designated by the analyzer system. 2. The system of claim 1, further comprising an automated positioning system connected to the analyzer system which controls the delivery of the doses from the therapeutic source to the affected areas of skin designated by the analyzer system. 3. The system of claim 1, wherein during operation the analyzer system causes the therapeutic source to deliver the doses of therapeutic radiation to the affected areas of skin. 4. The system of claim 1, wherein the therapeutic source comprises a xenon chloride excimer laser. 5. The system of claim 1, wherein the first and second diagnostics are diffuse reflectance and fluorescence, and wherein the detector system measures the diffuse reflectance and fluorescence and the analyzer system compares each of the signals relating to the diffuse reflectance and fluorescence to at least one threshold parameter to separately evaluate whether each optical diagnostic signal indicates the presence of the skin disorder at the selected area. 6. The system of claim 5, wherein the analyzer system normalizes the signals before comparing them to the at least one threshold parameter. 7. The system of claim 1, further comprising a manual instrument connected to the therapeutic source and the detector system such that the instrument delivers the radiation emitted from the selected areas of the patient's skin to the detector system and delivers the doses of therapeutic radiation from the therapeutic source to the affected areas. 8. The system of claim 7, wherein the manual instrument is further connected to the illumination system to deliver illumination radiation to the patient's skin from the illumination system. 9. The system of claim 7, wherein the manual instrument comprises one or more optical fibers through which the detector system receives optical radiation emitted from the selected areas of the patient's skin, and through which the therapeutic source delivers t... | The General Hospital Corporation | A61B018/18 A61B005/04
US07200612 | B2 | 2007-04-03 | processing data for interpretation | 2001-03-23 | A system for improving sensor-based decision making provides for the automatic submission of data obtained locally from instrumentation (such as image data) together with the interpretation of that data, which can be the output of some software which has been checked and possibly corrected by a user according to his/her expertise, to a remote database via an internetwork. The submission to the remote database is preferably automatic so that the remote database grows over time. The local site can access the remote database to retrieve information to assist in interpretation of the locally produced data (for example similar images and their corresponding interpretations), or can retrieve updated or improved software or parameters improving the software used for processing the data. The information on the remote database can also be reprocessed by software agents to provide statistical information based on information from a variety of such local sites. The system is particularly useful in improving the interpretation of data which is difficult to interpret such as medical image data (e.g. mammographic or cardiac ultrasound data). | 1. A method of processing data for a user, comprising: accepting input of data acquired from instrumentation as a dataset of an information domain which dataset comprises image data which comprises medical image data; andprocessing the data set under the control of processing software, said processing comprising the input of “user truth” based on interpretation of the dataset by the user, to provide a processed dataset to assist in further interpretation of the dataset;storing via a network said dataset and said processed dataset into a database of information relating to said domain acquired from a plurality of users;estimating automatically and quantitatively the effect of said “user truth” on the data set-in order to combine in said database such data from multiple users while minimizing the effects of user truth;and wherein said processing software comprises a software agent operative, on the basis of at least one of attributes of said dataset, attributes of the processed dataset and user-input, to access the database to retrieve data and/or software to mediate the processing and/or interpretation of said dataset by said processing software wherein the software agent is operative to transfer to the database data concerning manual adjustments made to said processed dataset by the user, and constituting said “user truth”, wherein the processing of said dataset is mediated by updating said processing software using at least one of data and software from the database. | 1. A method of processing data for a user, comprising: accepting input of data acquired from instrumentation as a dataset of an information domain which dataset comprises image data which comprises medical image data; andprocessing the data set under the control of processing software, said processing comprising the input of “user truth” based on interpretation of the dataset by the user, to provide a processed dataset to assist in further interpretation of the dataset;storing via a network said dataset and said processed dataset into a database of information relating to said domain acquired from a plurality of users;estimating automatically and quantitatively the effect of said “user truth” on the data set-in order to combine in said database such data from multiple users while minimizing the effects of user truth;and wherein said processing software comprises a software agent operative, on the basis of at least one of attributes of said dataset, attributes of the processed dataset and user-input, to access the database to retrieve data and/or software to mediate the processing and/or interpretation of said dataset by said processing software wherein the software agent is operative to transfer to the database data concerning manual adjustments made to said processed dataset by the user, and constituting said “user truth”, wherein the processing of said dataset is mediated by updating said processing software using at least one of data and software from the database. 2. A method according to claim 1 wherein the storing via the network of said dataset and said processed dataset into the database is automatic. 3. A method according to claim 1 wherein the software agent comprises means for assisting the user to prepare the dataset and, processed dataset for storing in the database. 4. A method according to claim 1, wherein said database comprises information produced by execution of said method by said plurality of users at a plurality of geographically separate sites. 5. A method according to claim 4, wherein said information produced by execution of said method comprises datasets and processed datasets from said plurality of geographically separate sites. 6. A method according to claim 5, further comprising comparing together the datasets from said plurality of geographically separate sites stored on the database. 7. A method according to claim 6, further comprising storing the results of comparing together the datasets from said plurality of geographically separate sites stored on the database, and outputting the results to the plurality of geographically separate sites as said data to mediate the processing or interpretation of said dataset. 8. A method according to claim 6 wherein said datasets are processed under the control of different processing software at different ones of said geographically separate sites. 9. A method according to claim 5 further comprising processing said information in the database to generate data representi... | Mirada Solutions Limited | G06F007/00 G06F017/30 G06F017/00 G06Q010/00 G06Q050/00 A61B005/00 G06F019/00 G06K009/00
US06985869 | B1 | 2006-01-10 | Digital prescription carrier and monitor system | 2000-01-21 | A digital prescription carrier and monitor system includes a pager sized carrier apparatus with an internal processor, a real-time clock/calendar, non-volatile memory, a communication port, a character display, and alert devices. Prescription data for one or more medications is stored within the carrier by a physician and downloaded by a pharmacist to fill the prescriptions called for. The carrier apparatus requires the entry of an decryption key to access prescription data within the carrier to prevent unauthorized access or tampering with the prescription data. The carrier apparatus also functions as a prescription reminder to the patient and records compliance with the prescription, since a switch must be operated to quiet an alert device activated when a scheduled medication dose is due. | 1. A method for conveying a prescribed medication to a patient, the method comprising the steps of: providing a digital prescription carrier including a read/write memory and an infrared communication interface;encrypting prescription data defining a prescription so that the data would be indecipherable without appropriate computer decryption software;uploading, by a prescriber, the prescription data into said carrier through said interface, said prescription calling for the use of a selected medication of a selected dosage on a selected schedule;transferring said carrier by a patient to a pharmacy;downloading said prescription data from said carrier through said interface at said pharmacy;decrypting said prescription data from indecipherable form into a form that would be decipherable;andfilling said prescription at said pharmacy; wherein,the uploading and downloading steps are each accomplished by a data transfer that occurs without physical contact. | 1. A method for conveying a prescribed medication to a patient, the method comprising the steps of: providing a digital prescription carrier including a read/write memory and an infrared communication interface;encrypting prescription data defining a prescription so that the data would be indecipherable without appropriate computer decryption software;uploading, by a prescriber, the prescription data into said carrier through said interface, said prescription calling for the use of a selected medication of a selected dosage on a selected schedule;transferring said carrier by a patient to a pharmacy;downloading said prescription data from said carrier through said interface at said pharmacy;decrypting said prescription data from indecipherable form into a form that would be decipherable;andfilling said prescription at said pharmacy; wherein,the uploading and downloading steps are each accomplished by a data transfer that occurs without physical contact. 2. A method as set forth in claim 1, further including the step of entering a first access code into said carrier to enable access to said prescription data prior to said uploading step. 3. A method as set forth in claim 2, further including the step of entering a second access code into said carrier to enable access to said prescription data prior to said downloading step. 4. A method as set forth in claim 2, further including the steps of: endowing a prescriber with the first access code;updating, by a prescriber, of prescription information including at least one of deleting a piece of stored prescription data;adding a new piece of stored prescription data;changing a piece of stored prescription data;endowing the pharmacist with the second access code; and,updating, by the pharmacist, of prescription information including at least one of noting the filling of a prescription;reducing the number of refills remaining for a piece of stored prescription data; or,updating patient information. 5. A method as set forth in claim 1, further including the steps of: (a) operating a digital clock/calendar within said carrier to generate internal values of time and date;(b) providing said carrier with a prescription compliance switch interfaced to said clock/calendar;(c) operating said compliance switch by a patient upon taking a medication specified by said prescription; and(d) storing in a compliance memory within said carrier respective values of time and date occurring upon operation of said compliance switch. 6. A method as set fort in claim 5, further including the steps of: (a) providing said carrier with an annunciator or element,(b) entering into said carrier by said pharmacist schedule data defining a prescription schedule comprising a plurality of sets of schedule times and dates at which a patient is to take a medication specified by prescription;(c) periodically comparing within said carrier said internal values of fine and date with said schedule times and dates; and(d) activating said annu... | NextMed, LLC | G06F017/60 G04G007/00 A61B005/00
US06963772 | B2 | 2005-11-08 | User-retainable temperature and impedance monitoring methods and devices | 2002-04-17 | A user-retainable monitoring system is disclosed. At least a pair of sensors is provided in association with a support member. The support member is preferably of a type that may be worn by or at least temporarily implanted in a patient. Possible sensor types include temperature sensors and impedance sensors. Temperature sensors may be used to detect a temperature differential between areas of tissue indicative of pathology. Impedance sensors are used to detect subcutaneous fluid detection. The support member may take the form of a bandage, drain or other structure. Monitor structures as described may have stand-alone utility or be connected to a processor or data recorder to enable various functions. | 1. A method of monitoring a wound area of a subject, said method comprising: contacting a subject with a device comprising: a support member, andone or more sensing portions supported by said support member, wherein said one or more sensing portions comprise at least one of: an impedance sensor and two temperature sensors;obtaining measurement data from said one or more sensing portions, wherein said data is indicative of healing or infection of said wound area; andautomatically initiating a secondary medical procedure in view of said obtained measurement data. | 1. A method of monitoring a wound area of a subject, said method comprising: contacting a subject with a device comprising: a support member, andone or more sensing portions supported by said support member, wherein said one or more sensing portions comprise at least one of: an impedance sensor and two temperature sensors;obtaining measurement data from said one or more sensing portions, wherein said data is indicative of healing or infection of said wound area; andautomatically initiating a secondary medical procedure in view of said obtained measurement data. 2. The method of claim 1, further comprising outputting a message indicative of said healing or said infection. 3. The method of claim 2, wherein said message is an audio message or visual message. 4. The method of claim 1, wherein said wound is a surgical wound. 5. The method of claim 1, wherein contacting comprises implanting at least a portion of said device within the body of said subject. 6. The method of claim 1, wherein at least a portion of said device is contacted with a skin surface of said subject. 7. The method of claim 1, wherein said data is indicative of healing. 8. The method of claim 1, wherein said data is indicative of infection. 9. The method of claim 1, wherein said device comprises an impedance sensor. 10. The method of claim 9, wherein said contacting comprises positioning a part of said impedance sensor adjacent a first side of a wound of said subject and positioning a part of said impedance sensor adjacent a second side of said wound. 11. The method of claim 1, further comprising electronic hardware that includes a memory unit and said method further comprises storing said obtained measurement data in said memory unit. 12. The method of claim 1, wherein said obtained measurement data comprises temperature measurement data. 13. The method of claim 12, wherein a decrease in temperature over time is indicative of wound healing. 14. The method of claim 12, wherein an increase in temperature over time is indicative of a wound infection. 15. A method of monitoring a wound area of a subject, said method comprising: contacting a subject with a device comprising: a support member, andfirst and second temperature sensors supported by said support member,wherein said contacting comprises positioning said first temperature sensor adjacent a non-wound area of said subject and positioning said second temperature sensor adjacent a wound area of said subject; andobtaining measurement data from said first and second temperature sensors, wherein said data is indicative of healing or infection of said wound area. 16. The method of claim 15, further comprising outputting a message indicative of said healing or said infection. 17. The method of claim 16, wherein said message is an audio message or visual message. 18. The method of claim 15, wherein said wound is a surgical wound. 19. The method of claim 15, wherein contacting comprises implanting at least a portion of... | The Board of Trustees of the Leland Stanford Junior University | A61B005/053 A61B005/00
US06908442 | B2 | 2005-06-21 | Bending resistant male connector for a guide wire | 2001-11-07 | During use there is a risk that a male connector for a guide wire is bent. With the present male connector (1) more material with high modulus of elasticity can be provided inside the male connector (1), which makes the male connector (1) more bending resistant. The male connector (1) comprises a core wire (3), a plurality of conductive members (4) spaced apart longitudinally along said core wire (3), a plurality of conductors (5) disposed along the core wire (3), each of the conductors (5) being connected to a respective conductive member (4). The core wire (3) has such a shape that least one longitudinal cavity is provided, which longitudinal cavity remains substantially intact when the male connector (1) is bent, thereby protecting the conductors (5), which are disposed in said cavity, from being damaged. | 1. Male connector (1) for a guide wire (2), wherein the guide wire (2) comprises a guide wire core including a section that substantially extends through the guide wire (2) between a sensor at a distal end of the guide wire (2) and the male connector (1) at a proximal end of the guide wire (2), wherein the male connector (1) comprises a core wire (3), a plurality of conductive members (4) spaced apart longitudinally along said core wire (3), and a plurality of conductors (5) disposed along the core wire (3), each of the conductors (5) being connected to a respective conductive member (4), wherein the core wire (3) has such a shape that at least one longitudinal cavity is provided inside the male connector (1), the longitudinal cavity being adapted to remain substantially intact when the male connector (1) is bent, thereby protecting the conductors (5), which are disposed in said cavity, from being damaged if the male connector (1) is bent, wherein the core wire does not have a solid circular cross section, and wherein the core wire (3) is stiffer than the section of the guide wire core that substantially extends through the guide wire (2) substantially from the distal end of the guide wire (2) substantially to the proximal end of the guide wire (2). | 1. Male connector (1) for a guide wire (2), wherein the guide wire (2) comprises a guide wire core including a section that substantially extends through the guide wire (2) between a sensor at a distal end of the guide wire (2) and the male connector (1) at a proximal end of the guide wire (2), wherein the male connector (1) comprises a core wire (3), a plurality of conductive members (4) spaced apart longitudinally along said core wire (3), and a plurality of conductors (5) disposed along the core wire (3), each of the conductors (5) being connected to a respective conductive member (4), wherein the core wire (3) has such a shape that at least one longitudinal cavity is provided inside the male connector (1), the longitudinal cavity being adapted to remain substantially intact when the male connector (1) is bent, thereby protecting the conductors (5), which are disposed in said cavity, from being damaged if the male connector (1) is bent, wherein the core wire does not have a solid circular cross section, and wherein the core wire (3) is stiffer than the section of the guide wire core that substantially extends through the guide wire (2) substantially from the distal end of the guide wire (2) substantially to the proximal end of the guide wire (2). 2. The male connector (1) according to claim 1, wherein the core wire (3) has a D-shaped cross section, and wherein said cavity is provided between an inner surface of the cylindrical conductive members (4) and the flat part of D-shaped core wire (3). 3. The male connector according to claim 2, wherein a continuous insulating material (6) is disposed between the conductive members (4) and the core wire (3), with a minimum of insulating material (6) being provided between the curved part of the D-shaped core wire (3) and the inner surface of the conductive members (4). 4. The male connector (1) according to claim 2, wherein the core wire (3) is made of titanium, the surface of which is oxidized to titanium dioxide. 5. The male connector (1) according to claim 1, wherein said cavity is provided by a longitudinal recess in the mantle surface of the core wire (3). 6. The male connector (1) according to claim 5, wherein a continuous insulating material (6) is disposed between the conductive members (4) and the core wire (3), with a minimum of insulating material (6) being provided between the cylindrical part of the core wire (3) and an inner surface of the cylindrical conductive members (4). 7. The male connector (1) according to claim 1, wherein several cavities are provided by longitudinal recesses in the mantle surface of the core wire (3), and wherein each conductor (5) is disposed in a separate recess. 8. The male connector (1) according to claim 1, wherein several cavities are provided by longitudinal recesses in the mantle surface of the core wire (3), and wherein each recess accommodates at least one conductor (5). 9. The male connector (1) according to claim 1, wherein said cavity is pro... | Radi Medical Systems AB | A61B005/00 A61M025/00
US06856824 | B1 | 2005-02-15 | Apparatus and method for detecting abnormalities in bodily matter | 1998-10-01 | There is disclosed electrical impedance tomography apparatus adapted to detect abnormalities in bodily matter comprising: electrical signal generating means for generating electrical signals at a plurality of frequencies; an electrode arrangement for applying the electrical signals to the bodily matter and detecting electrical impendance properties of the bodily matter; and data processing means for correlating the detected electrical impedance properties with the presence or absence of abnormalities in the bodily matter; in which electrical signals of a frequency greater than 1 MHz, preferably greater than 2 MHz, more preferably greater than 3 MHz and most preferably greater than 4 MHz are applied to the bodily matter. | 1. An electrical tomographic method for detecting abnormalities in bodily matter comprising the steps of: generating electrical signals having a frequency greater than 4 MHZ; applying the electrical signals to the bodily matter using an electrode arrangement; detecting electrical impedance properties of the bodily matter; and, using data processing means correlating the detected electrical impedance properties with the presence or absence of abnormalities in the bodily matter using a fractal model of tissue impedance. | 1. An electrical tomographic method for detecting abnormalities in bodily matter comprising the steps of: generating electrical signals having a frequency greater than 4 MHZ; applying the electrical signals to the bodily matter using an electrode arrangement; detecting electrical impedance properties of the bodily matter; and, using data processing means correlating the detected electrical impedance properties with the presence or absence of abnormalities in the bodily matter using a fractal model of tissue impedance. 2. The method according to claim 1, wherein the detected electrical impedance properties are selected from the group consisting of Zi, Zx, Zxcom, Zm, Zmcom, Cm and Cmcom. 3. The method according to claim 1, wherein the data processing means correlates detected electrical impedance properties selected from the group consisting of Zicom, Zxcom to Zicom ratio, Zxcom, Zmcom, Cmcom, Zxcom/ZMcom, Zxcom/Cmcom, and Zmcom, Zicom. 4. The method according to claim 1, wherein the data processing means correlates changes and/or large variations of intracellular impedance Zicom with the presence of dyskaryosis. 5. The method according to claim 1, wherein the data processing means correlates abnormal Zxcom, Zxcom to Zicom ratio with abnormal nuclear-to-cytoplasmic ratio (NCR). 6. The method according to claim 1, wherein the data processing means correlates abnormal Zxcom with abnormal inter-cellular cohesion. 7. The method according to claim 1, wherein the data processing means correlates abnormal Zmcom, Cmcom, Zxcom/Zmcom, Zmcom/Cmcom, Zmcom Zicom with abnormal membrane morphology. 8. The method according to claim 1, wherein the data processing means correlates detected impedance properties selected from the group consisting of: i) Zicom; ii) Zxcom; iii) Zmcom, Cmcom; and, iv) ratio Zxcom/Zicom; Zxcom/Zmcom; Zxcom/Cmcom; with the presence of non-infiltrate, early-infiltrate or infiltrate stage cancer. 9. The method of claim 1, wherein the data processing means references the detected electrical impedance properties of the bodily matter to the detected electrical impedance of other bodily matter. 10. The method according to claim 9, being adapted to detect a breast carcinoma, wherein the detected impedance properties are of breast tissue and are referenced to detected electrical impedance properties of fatty tissue in the breast. 11. The method according to claim 1, wherein the data processing means is adapted to compare the detected electrical impedance properties with a database of impedance properties corresponding to bodily matter of known composition. 12. The method of claim 1, adapted to detect a carcinoma. 13. The method according to claim 1, adapted to detect a breast carcinoma. 14. The method according to claim 1, adapted to detect at least one of the group consisting of Stage 3, Stage 2 and Stage 1 carcinomas. 15. The method according to claim 1, wherein the electrode arrangement is disposed in a woman's brassiere. 16. T... | De Montfort University | A61B005/05 A61B006/00
US06913580 | B2 | 2005-07-05 | Method of body fluid specimen collection | 2002-01-23 | In one aspect of the invention, a method of fluid specimen collection comprises the steps of: preparing a site for puncture using an antiseptic; piercing the prepared puncture site using a fluid collection needle; collecting a first body fluid specimen using a device for collecting the first body fluid specimen; collecting a second body fluid specimen using a device for collecting the second body fluid specimen and; selecting the second body fluid specimen for a diagnostic test of a kind that detects the presence of organisms in the second body fluid specimen. In another aspect of the invention, a vessel for collecting, transporting, and transferring a body fluid specimen comprises: a hollow body having a first and second end; a first seal at said first end; a plunger disposed within said hollow body between said first end and said second end; said plunger providing a second seal; a plunger lock coupled to said plunger; said plunger lock being configured to selectively maintain said plunger at said second end when at least a portion of said hollow body between said first seal and said second seal is at least partially evacuated; said plunger lock can further be configured to release said plunger, thereby allowing said plunger to move toward said first seal within said body. | 1. A collection vessel for collecting and transferring a body fluid specimen comprising: a hollow body having a first and a second end; a first seal at said first end; a plunger disposed within said hollow body between said first end and said second end; said plunger providing a second seal; a plunger lock coupled to said plunger; said plunger lock being configured to selectively maintain said plunger at said second end when at least a portion of said hollow body between said first seal and said second seal is at least partially evacuated; said plunger lock further configured to release said plunger, thereby allowing said plunger to move toward said first seal within said hollow body. | 1. A collection vessel for collecting and transferring a body fluid specimen comprising: a hollow body having a first and a second end; a first seal at said first end; a plunger disposed within said hollow body between said first end and said second end; said plunger providing a second seal; a plunger lock coupled to said plunger; said plunger lock being configured to selectively maintain said plunger at said second end when at least a portion of said hollow body between said first seal and said second seal is at least partially evacuated; said plunger lock further configured to release said plunger, thereby allowing said plunger to move toward said first seal within said hollow body. 2. A collection vessel for collecting and transferring a body fluid specimen according to claim 1, further comprising: an airtight junction that interrupts said hollow body forming a first section and a second section; said first section having said first seal; said second section having said plunger and said plunger lock; said airtight junction configured to allow for separation of said first and said second section and coupling of a transfer needle to said second section. 3. A collection vessel for collecting and transferring a body fluid specimen according to claim 1, wherein said hollow body has the shape of a hollow cylinder. 4. A collection vessel for collecting and transferring a body fluid specimen according to claim 1, wherein said plunger lock comprises a threaded shaft and said plunger comprises a threaded receiver. 5. A collection vessel for collecting and transferring a body fluid specimen according to claim 1, wherein said plunger lock breaks away from said plunger, thereby allowing said plunger to move towards said first seal within said hollow body. 6. A collection vessel for collecting and transferring a body fluid specimen according to claim 1, wherein said plunger lock remains at least in part with said plunger as it moves toward said first seal within said hollow body. 7. A collection vessel for collecting and transferring a body fluid specimen according to claim 1, further comprising an additive within said collection vessel. 8. A collection vessel for collecting and transferring a body fluid specimen according to claim 1, wherein said collection vessel is sterilized and packaged to maintain sterility. 9. A method for collecting a first body fluid specimen and a second body fluid specimen, said second body fluid specimen having a lower concentration of living contaminants than said first body fluid specimen comprising the steps of: providing a fluid collection needle having a first end and a second end; providing a sterile evacuated specimen tube comprising a sterile hollow body having an open end, a sterile seal at said open end of said sterile evacuated specimen tube, said sterile seal of said sterile evacuated specimen tube configured wherein said sterile seal of said sterile evacuated specimen tube is maintained at said open end of s... | Benjamin Curtis Stone | A61B005/00
US06947789 | B2 | 2005-09-20 | Method for detecting, sizing and locating old myocardial infarct | 2001-11-13 | Algorithms for detecting, sizing and locating old myocardial infarcts by evaluating particular ECG lead data derived from selected ECG leads, where that particular data is presented hierarchically for review and confirmation, and includes selected voltage-amplitude ratios of R/Q and R/S. For a given human subject, the specific hierarchical pattern of lead data to be examined is selected on the basis of predetermined personal and demographic data, and completely in light of various, well-known, so-called confounders and excluders. | 1. A method for detecting and characterizing, in the presence of confounders, a subject's old myocardial infarct (MI) comprising collecting that subject's ECG data from several preselected, standard ECG leads, establishing, in the presence of a history of confounding conditions and in relation to selected characteristics of that subject's personal data, such as, inter alia, sex, age, and/or race, a set of ECG-data criteria to examine, including R/Q and R/S voltage-amplitude ratio criteria, examining such established criteria set in the context of the mentioned history of confounding conditions, and from said examining, generating an output indicative of the desired detecting and characterizing of an old MI. | 1. A method for detecting and characterizing, in the presence of confounders, a subject's old myocardial infarct (MI) comprising collecting that subject's ECG data from several preselected, standard ECG leads, establishing, in the presence of a history of confounding conditions and in relation to selected characteristics of that subject's personal data, such as, inter alia, sex, age, and/or race, a set of ECG-data criteria to examine, including R/Q and R/S voltage-amplitude ratio criteria, examining such established criteria set in the context of the mentioned history of confounding conditions, and from said examining, generating an output indicative of the desired detecting and characterizing of an old MI. 2. The method of claim 1, wherein the established R/Q and R/S ratio criteria are associated variously with one or more of ECG leads I, A VL, A VF, VI, V2, V 4, V5, V6, V4R and V8. | Innovise Medical, Inc. | A61B005/0452
US06878118 | B2 | 2005-04-12 | Esophageal stethoscope with separate cover member | 2003-02-06 | An esophageal stethoscope includes a tubular flexible body having an acoustic input region. A carrier member, which is separate from the tubular flexible body, is selectively engageable with the tubular flexible body either proximally or distally of the acoustic input region. The carrier member carries one or more devices for providing esophageal recording, monitoring or stimulation functions. For transesophageal cardiac pacing, electrodes are mounted to the carrier member such that engagement of the carrier member with the tubular flexible body functions to convert the esophageal stethoscope to a pacing stethoscope. To carry out a sensing, recording or monitoring function, such as esophageal oximetry, one or more oximetry probes are mounted to the carrier member such that engagement of the carrier member with the esophageal stethoscope converts the stethoscope for use in an oximetry sensing application. A “universal” stethoscope includes engagement structures at predetermined locations along the length of the tubular flexible body, for use in engaging a stimulation carrier and/or a recording/monitoring carrier with the esophageal stethoscope. | 1. An esophageal stethoscope assembly, comprising: a tubular flexible body having a proximal end and a distal end, and including an acoustic input region defined on the tubular flexible body toward the distal end; a carrier member separate from the tubular flexible body and engageable with the tubular flexible body, wherein the carrier member includes one or more devices for sensing and/or controlling a condition from or within the esophagus of a patient, wherein the tubular flexible body extends along a longitudinal axis, and wherein the carrier member defines an axial passage and includes wall structure defining a pair of spaced ends, wherein the space between the spaced ends of the wall structure is in communication with the axial passage, and wherein the carrier member is engageable with the tubular flexible body by application of a push-on force transverse to the longitudinal axis of the tubular flexible body such that the tubular flexible body passes through the space between the wall structure ends and into the axial passage of the carrier member, wherein the carrier member wall structure is formed of a resilient material such that the wall structure is separable to enlarge the space between the wall structure ends to allow the tubular flexible body to pass therethrough into the axial passage of the carrier member, and wherein the resiliency of the wall structure is operable to thereafter engage the wall structure with the tubular flexible body; and a retainer structure interposed between the carrier member and the tubular flexible body for maintaining the carrier member in a predetermined axial position relative to the tubular flexible body. | 1. An esophageal stethoscope assembly, comprising: a tubular flexible body having a proximal end and a distal end, and including an acoustic input region defined on the tubular flexible body toward the distal end; a carrier member separate from the tubular flexible body and engageable with the tubular flexible body, wherein the carrier member includes one or more devices for sensing and/or controlling a condition from or within the esophagus of a patient, wherein the tubular flexible body extends along a longitudinal axis, and wherein the carrier member defines an axial passage and includes wall structure defining a pair of spaced ends, wherein the space between the spaced ends of the wall structure is in communication with the axial passage, and wherein the carrier member is engageable with the tubular flexible body by application of a push-on force transverse to the longitudinal axis of the tubular flexible body such that the tubular flexible body passes through the space between the wall structure ends and into the axial passage of the carrier member, wherein the carrier member wall structure is formed of a resilient material such that the wall structure is separable to enlarge the space between the wall structure ends to allow the tubular flexible body to pass therethrough into the axial passage of the carrier member, and wherein the resiliency of the wall structure is operable to thereafter engage the wall structure with the tubular flexible body; and a retainer structure interposed between the carrier member and the tubular flexible body for maintaining the carrier member in a predetermined axial position relative to the tubular flexible body. 2. The esophageal stethoscope assembly of claim 1, wherein the carrier member is engageable with the tubular flexible body distally of the acoustic input region and wherein the one or more devices comprise a series of pacing electrodes carried by the carrier member. 3. The esophageal stethoscope assembly of claim 1, wherein the carrier member is engageable with the tubular flexible body proximally of the acoustic input region and wherein the one or more devices comprise one or more oximetry sensors carried by the carrier member. 4. The esophageal stethoscope assembly of claim 1, wherein the carrier member is engageable with the tubular flexible body distally of the acoustic input region and wherein the one or more devices comprise a series of pacing electrodes, and further comprising a second carrier member separate from the tubular flexible body and engageable with the tubular flexible body proximally of the acoustic input region, wherein the proximal second carrier member carries one or more oximetry sensors. 5. An esophageal stethoscope assembly, comprising: a tubular flexible body having a proximal end and a distal end, wherein the tubular flexible body defines a length and includes an acoustic input region located toward the distal end; a carrier member separate from the tubular flexible bod... | - | A61B005/042
US06926838 | B2 | 2005-08-09 | Method to measure blood flow and recirculation in hemodialysis shunts | 2003-01-31 | The measurement of blood flow in a dialysis shunt is obtained by injection of an indicator material into a venous line leading from dialysis equipment to the shunt. The blood flow in an arterial line leading from the shunt at a location downstream of the venous line to the dialysis equipment is monitored by an arterial line sensor for the presence of the indicator material. A detector connected to the sensor provides a dilution curve in response to the presence of the indicator material and the blood flow in the shunt is calculated from the area under the dilution curve. The locations of the arterial and venous lines in the shunt can be reversed to obtain a measurement of blood recirculation from the venous line into the arterial line. | 1. A process for determining blood flow in an arterio-venous shunt in a cardiovascular circuit, comprising, (a) changing a selected blood parameter in blood flowing in a circulating system outside the cardiovascular circuit to produce a distinguishable blood characteristic in blood which is delivered to an arterio-venous shunt connected in the cardiovascular circuit and carrying a shunt blood flow; (b) delivering blood having the distinguishable blood characteristic from the circulating system into an upstream location in the arterio-venous shunt so that the delivered blood mixes with the shunt blood flow; (c) removing a portion of the mixed blood from the arterio-venous shunt at a location in the shunt which is downstream from the upstream location and delivering the removed portion of mixed blood to the circulating system; (d) measuring the amount of distinguishable blood characteristic in the removed portion of mixed blood; and (e) calculating the rate of flow of the shunt blood flow in the arterio-venous shunt from the measured amount of distinguishable blood characteristic, wherein the rate of flow of the shunt blood flow in the arterio-venous shunt is calculated using at least one of the following indicator dilution equations or an indicator dilution equation that can be derived from at least one of the following equations: Q=V/S; Qshunt=Vven/Sart−Qdial; Qshunt=Qdial(Vven/Vcal×/Sart−1); Qshunt=(Vven/Sart−Vcal/Scal); Qshunt=Qdial(Sven/Sart−1): where Q is a rate of blood flow; V is an amount of indicator; S is an amount of change in a parameter observed in response to the introduction of an indicator; Qshunt is a flow rate of blood flowing through the arterio-venous shunt; Vven is an amount of indicator introduced into the arterio-venous shunt; Sart is an amount of change in a blood parameter observed in an arterial line in response to the introduction of an indicator; Qdial is a flow rate of blood flowing through a dialyzer in the circulating system; Vcal is a known amount of indicator; Scal is an amount of change in a blood parameter observed in response to the introduction of Vcal; and Sven is an amount of change in a blood parameter observed in a venous line in response to the introduction of an indicator. | 1. A process for determining blood flow in an arterio-venous shunt in a cardiovascular circuit, comprising, (a) changing a selected blood parameter in blood flowing in a circulating system outside the cardiovascular circuit to produce a distinguishable blood characteristic in blood which is delivered to an arterio-venous shunt connected in the cardiovascular circuit and carrying a shunt blood flow; (b) delivering blood having the distinguishable blood characteristic from the circulating system into an upstream location in the arterio-venous shunt so that the delivered blood mixes with the shunt blood flow; (c) removing a portion of the mixed blood from the arterio-venous shunt at a location in the shunt which is downstream from the upstream location and delivering the removed portion of mixed blood to the circulating system; (d) measuring the amount of distinguishable blood characteristic in the removed portion of mixed blood; and (e) calculating the rate of flow of the shunt blood flow in the arterio-venous shunt from the measured amount of distinguishable blood characteristic, wherein the rate of flow of the shunt blood flow in the arterio-venous shunt is calculated using at least one of the following indicator dilution equations or an indicator dilution equation that can be derived from at least one of the following equations: Q=V/S; Qshunt=Vven/Sart−Qdial; Qshunt=Qdial(Vven/Vcal×/Sart−1); Qshunt=(Vven/Sart−Vcal/Scal); Qshunt=Qdial(Sven/Sart−1): where Q is a rate of blood flow; V is an amount of indicator; S is an amount of change in a parameter observed in response to the introduction of an indicator; Qshunt is a flow rate of blood flowing through the arterio-venous shunt; Vven is an amount of indicator introduced into the arterio-venous shunt; Sart is an amount of change in a blood parameter observed in an arterial line in response to the introduction of an indicator; Qdial is a flow rate of blood flowing through a dialyzer in the circulating system; Vcal is a known amount of indicator; Scal is an amount of change in a blood parameter observed in response to the introduction of Vcal; and Sven is an amount of change in a blood parameter observed in a venous line in response to the introduction of an indicator. 2. The process of claim 1, wherein the amount of distinguishable blood characteristic in the removed portion of mixed blood is measured by a sensor located on a line of the circulating system through which the removed portion of mixed blood flows. 3. The process of claim 1, wherein the step of measuring the amount of distinguishable blood characteristic includes measuring the characteristic at a location external to the circulating system. 4. The process of claim 1, wherein the step of changing a selected blood parameter includes changing the selected blood parameter at a location within the circulating system. 5. The process of claim 1, wherein the step of changing a selected blood parameter includes changing a concentration of t... | Transonic Systems, Inc. | A61B005/026 A61B005/0275 A61B008/06
US06880968 | B1 | 2005-04-19 | Test element analysis system | 2000-10-26 | Test element analysis system (1) for the analytical investigation of a sample (8), in particular of a body liquid, of human beings or of animals, comprising test elements (3) with a test zone (7), to be brought in contact with the sample to be investigated for the purpose of performing an analysis, in order to measure a measurement quantity characteristic for the analysis, and an evaluation instrument (2) with a test element holder (5) for positioning a test element (3) in a measuring position in order to perform a measurement, and a measurement and evaluation electronics (15) for measuring the characteristic change and for determining a result of the analysis, based on the result of the measurement.;In order to provide increased measuring accuracy by improved temperature compensation, it is proposed, in the scope of the invention, that the evaluation instrument (2) for the determination of the temperature prevailing in the test zone (7) of the test element (3) comprises an infrared detector (20). | 1. Test element analysis system for the analytical investigation of an analyte present in a body liquid of humans or animals, comprising at least one test element with a test zone, which is contacted with a sample to be analyzed in order to measure a measurement quantity which is characteristic for the analysis, and an evaluation instrument with a test element holder for positioning at least one test element in a measuring position in order to perform a measurement, an analyte measurement and evaluation electronics for measuring the characteristic change and for determining, based on this measurement, a result of the analysis, and an infrared detector adapted for the determination of the temperature in the test zone, characterized in that the evaluation instrument comprises an infrared radiation transport device for providing a location-selective connection of the test zone with the infrared detector. | 1. Test element analysis system for the analytical investigation of an analyte present in a body liquid of humans or animals, comprising at least one test element with a test zone, which is contacted with a sample to be analyzed in order to measure a measurement quantity which is characteristic for the analysis, and an evaluation instrument with a test element holder for positioning at least one test element in a measuring position in order to perform a measurement, an analyte measurement and evaluation electronics for measuring the characteristic change and for determining, based on this measurement, a result of the analysis, and an infrared detector adapted for the determination of the temperature in the test zone, characterized in that the evaluation instrument comprises an infrared radiation transport device for providing a location-selective connection of the test zone with the infrared detector. 2. Analysis system according to claim 1, characterized in that the infrared detector is integrated into the measuring and evaluation electronics. 3. The system of claim 2 wherein a rigid mechanic connection is provided the infrared detector and the measuring and evaluation electronics. 4. Analysis system according to claim 1, characterized in that the infrared radiation transport device comprises a hollow conductor with an interior wall reflective for infrared radiation. 5. Analysis system according to claim 4, characterized in that the hollow conductor is made from metal-coated plastic. 6. Analysis system according to claim 4, characterized in that an input opening of the hollow conductor facing the test zone has a larger opening cross-section than an output opening of the hollow conductor facing the infrared detector. 7. Analysis system according to claim 1, characterized in that the location-selective infrared radiation transport device includes an imaging mirror located inside a housing of the evaluation instrument. 8. Analysis system according to claim 1, characterized in that the test element, in the measuring position, sticks out of a housing of the evaluation instrument in such manner that the sample can be contacted to the test zone, while the test element is in the measuring position, the detector is located in a housing of the evaluation instrument, the housing of the evaluation instrument comprises a window which is transparent for infrared radiation, and a transport path of the infrared radiation between the test zone and the infrared detector passes through the optical window. 9. Analysis system according to claim 8, characterized in that the optical window is dust-proof closed by means of a pane transparent to infrared radiation. 10. Analysis system according to claim 8, characterized in that the evaluation instrument comprises an infrared radiation transport device for providing a location-selective connection of the test zone with the infrared detector and the infrared transparent radiation pane of the optical window is c... | Roche Diagnostics GmbH | G01K001/16 G01J005/00 A61B005/00
US06875020 | B2 | 2005-04-05 | Method of providing an individualized online behavior modification program using medical aids | 2002-10-04 | A method of providing an individualized online behavior modification program using medical aids. The method comprises the steps of: registering an individual that desires to modify a specific behavior; consulting with the individual; providing a suggested treatment plan including recommended medical aids based on information about the individual obtained during the registering step and consulting step; delivering the suggested treatment plan including the medical aids to a health care provider selected by the individual; providing a finalized treatment plan by incorporating the suggested treatment plan and suggestions from the health care provider; monitoring the individual's treatment; modifying the finalized treatment plan by adjusting accordingly the final treatment plan based on information received during the monitoring of the individual's treatment; and providing a final assessment. | 1. A method of providing an individualized online behavior modification program using medical aids, comprising the steps of: registering an individual that desires to modify a specific behavior; providing the individual with a personal online account, enabling the individual to access the online program; consulting with the individual; providing a suggested treatment plan including recommended medical aids based on information about the individual obtained during the registering step and consulting step; delivering the suggested treatment plan to a health care provider, to prescribe or recommend the medical aids to be used by the individual based on the recommended medical aids included in the suggested treatment plan; providing a finalized treatment plan by incorporating suggestions received from the health care provider into the suggested treatment plan to form the finalized treatment plan; monitoring the individual's treatment; and providing a final assessment by receiving a personal final assessment from the individual including an effectiveness determination, and receiving a final assessment from the health care provider of the individual. | 1. A method of providing an individualized online behavior modification program using medical aids, comprising the steps of: registering an individual that desires to modify a specific behavior; providing the individual with a personal online account, enabling the individual to access the online program; consulting with the individual; providing a suggested treatment plan including recommended medical aids based on information about the individual obtained during the registering step and consulting step; delivering the suggested treatment plan to a health care provider, to prescribe or recommend the medical aids to be used by the individual based on the recommended medical aids included in the suggested treatment plan; providing a finalized treatment plan by incorporating suggestions received from the health care provider into the suggested treatment plan to form the finalized treatment plan; monitoring the individual's treatment; and providing a final assessment by receiving a personal final assessment from the individual including an effectiveness determination, and receiving a final assessment from the health care provider of the individual. 2. The method as described in claim 1 wherein the medical aids include stabilization medication and behavior alternatives. 3. The method as described in claim 1 wherein the step of registering the individual comprises acquiring personal information from the individual. 4. The method as described in claim 1 wherein the step of consulting with the individual comprises acquiring from the individual habit information and personal preferences as to various options for modifying the specific behavior. 5. The method as described in claim 1 wherein the suggested treatment plan comprises: a suggested time frame for completion of the behavior modification program including a suggested daily schedule; wherein the suggested daily schedule provides a suggested schedule for the behavior modification and a schedule for use of the medical aids. 6. The method as described in claim 1 wherein the step of monitoring treatment comprises: receiving daily updates from the individual via the individual's personal online account; analyzing the daily updates from the individual; and providing the individual with feedback based on the analysis of the individual's daily updates. 7. The method as described in claim 1, further comprising: analyzing daily updates received from the individual, the personal final assessment from the individual, and the final assessment from the health care provider; and providing the individual with an effectiveness assessment and a long term self maintenance procedure. 8. The method as described in claim 1 further comprising modifying the finalized treatment plan by adjusting accordingly the final treatment plan based on information received during the monitoring of the individual's treatment. 9. The method as described in claim 8 wherein the step of adjusting the finalized treatment plan com... | Rx Maxwell, Inc. | G09B011/00 A61B005/00
US06843773 | B2 | 2005-01-18 | Blood pressure monitoring device and method of manufacturing a parts mounting module of a blood pressure monitoring device | 2002-06-12 | The invention is directed to a blood pressure monitoring device comprising a central parts mounting module and to a method of manufacturing this parts mounting module. According to the invention fluid channels are produced by gas injection into the interior of the parts mounting module, the walls bounding the fluid channels being completely formed by the material of the parts mounting module. The injection of gas continues for as long as a core of the parts mounting module is still in the liquid phase in the area of the fluid channel, whilst an outer area surrounding the core has already solidified, so that the pressure fluid expels the liquid core area, producing a tubular fluid channel. | 1. A blood pressure monitoring device comprising: a cuff; a fluid pressure system for filling the cuff with fluid under pressure; and a parts mounting module for components of the fluid pressure system, said parts mounting module made of a material and including a fluid channel integrated therein and serving to connect the components of the fluid pressure system, said fluid channel having a bounding wall that is completely formed by the material of the parts mounting module. | 1. A blood pressure monitoring device comprising: a cuff; a fluid pressure system for filling the cuff with fluid under pressure; and a parts mounting module for components of the fluid pressure system, said parts mounting module made of a material and including a fluid channel integrated therein and serving to connect the components of the fluid pressure system, said fluid channel having a bounding wall that is completely formed by the material of the parts mounting module. 2. The blood pressure monitoring device according to claim 1 wherein the fluid channel is routed in a three-dimensional curvature extending in several planes. 3. The blood pressure monitoring device according to claim 2 wherein the path of the fluid channel extends in a manner free from sharp-edged kinks. 4. The blood pressure monitoring device according to claim 1 wherein the bounding wall is of an essentially uniform thickness. 5. The blood pressure monitoring device according to claim 1 wherein the parts mounting module is injection molded from plastics material and the fluid channel is formed by injecting gas into the plastics material of the parts mounting module prior to its complete cooling. 6. The blood pressure monitoring device according to claim 1 wherein the fluid channel branches into several fluid channels in the interior of the material of the parts mounting module. 7. The blood pressure monitoring device according to claim 1 wherein the fluid channel has ports into which corresponding components of the fluid pressure system are snap-fitted with an essentially complementary connector element. 8. A blood pressure monitoring device comprising: a cuff; a fluid pressure system for filling the cuff with fluid under pressure; a parts mounting module for components of the fluid pressure system, said parts mounting module made of a material and including a fluid channel integrated therein and serving to connect the components of the fluid pressure system, said fluid channel having a bounding wall that is completely formed by the material of the parts mounting module, wherein all of the components of the fluid pressure system are mounted on the parts mounting module in bays specifically provided for each component; and a housing enclosing the parts mounting module and the components of the fluid pressure system. 9. The blood pressure monitoring device according to claim 8 wherein the cuff is arranged on an underside, and wherein the components of the fluid pressure system include a pump, a pressure valve, and a pressure sensor, all of which are arranged on an upper side of the parts mounting module and are coupled to fluid channel ports. 10. The blood pressure monitoring device according to claim 9 wherein the pump, the pressure valve, and the pressure sensor each have an electrical connector, and wherein the electrical connectors of the pump, the pressure valve, and the pressure sensor are arranged to lie essentially in one plane. 11. The blood pressure mon... | Braun GmbH | A61B005/02
US06862472 | B2 | 2005-03-01 | Medical testing system with an illuminating component | 2001-07-26 | A medical testing system includes an instrument for monitoring the electrical activity of a patient's heart. The instrument includes a work surface and an illuminating component for illuminating the work surface. | 1. A medical testing system comprising: a. an instrument for monitoring the electrical activity of a patient's heart, the instrument including a work surface and a printing device configured to print a graphical waveform representing the electrical activity of the heart on a medium moving across the work surface; b. a monitor coupled to the instrument and configured to display images representing the heart; and c. a light source, coupled to the instrument and positioned above the work surface, that directs light toward the work surface to illuminate the work surface; wherein the light source includes at lest one LED. | 1. A medical testing system comprising: a. an instrument for monitoring the electrical activity of a patient's heart, the instrument including a work surface and a printing device configured to print a graphical waveform representing the electrical activity of the heart on a medium moving across the work surface; b. a monitor coupled to the instrument and configured to display images representing the heart; and c. a light source, coupled to the instrument and positioned above the work surface, that directs light toward the work surface to illuminate the work surface; wherein the light source includes at lest one LED. 2. A medical testing system comprising: a. an instrument for monitoring the electrical activity of a patient's heart, the instrument including a work surface and a printing device adjacent the work surface for printing on a medium a graphical waveform representing the electrical activity of the heart; b. a monitor coupled to the instrument and configured to display images representing the heart; and c. a light source, coupled to the instrument and positioned above the work surface, that directs light toward the work surface to illuminate the work surface and is adapted to illuminate the medium. 3. The system of claim 2, further comprising a supporting component that includes a plate adapted to support the monitor at a position above the instrument, the plate including the light source. 4. The system of claim 2, wherein the instrument includes a power source, the light source being coupled to the power source. 5. The system of claim 2, wherein the instrument is an electrocardiograph. 6. A medical testing system comprising: a. an instrument that monitors the electrical activity of a patient's heart, said instrument including (1) a work surface, (2) a printing component configured to print on a medium, moving across the work surface, a graphical waveform representing the electrical activity of the heart, (3) a power source coupled to the printing component and (4) a monitor coupled to the instrument and configured to display images representing the heart; and b. an illuminating component, coupled to the power source and positioned above the work surface, which illuminates the work surface. 7. The system of claim 6, further comprising a supporting component engaging the instrument for supporting both the monitor and the illuminating component above the work surface. 8. The system of claim 7, wherein the instrument includes a keypad adjacent the work surface. 9. The system of claim 8, wherein the illuminating component illuminates the keypad. 10. The system of claim 6, wherein the illuminating component includes at least one light emitting diode. 11. A medical testing system comprising: a. means for monitoring the electrical activity of a patient's heart, the means for monitoring including (1) a work surface, (2) a means for printing on a medium, moving across the work surface, a graphical waveform representing the electrical... | GE Medical Systems Information Technologies, Inc. | A61B005/04
US06890311 | B2 | 2005-05-10 | Methods for performing medical procedures within a breast duct | 2002-03-15 | Methods for performing medical procedures within the duct of a breast are described. An exemplary embodiment of the invention includes the steps of introducing a guiding member into a breast duct; introducing and positioning a member having an internal lumen in the breast duct; introducing a wash fluid into the breast duct through the internal lumen; removing washings from the breast duct through the internal lumen; and collecting the washings for cytological analysis. | 1. A method for performing a medical procedure within a breast duct, said method includes the steps of: introducing a guiding member into said breast duct; traducing and positioning a member having an internal lumen in said breast duct; introducing a wash fluid into said breast duct through said internal lumen; removing washings from said breast duct through said internal lumen; and collecting said washings for cytological analysis. | 1. A method for performing a medical procedure within a breast duct, said method includes the steps of: introducing a guiding member into said breast duct; traducing and positioning a member having an internal lumen in said breast duct; introducing a wash fluid into said breast duct through said internal lumen; removing washings from said breast duct through said internal lumen; and collecting said washings for cytological analysis. 2. The method of claim 1 wherein the collected washings include cellular material from within said breast duct. 3. The method of claim 1 wherein said wash fluid includes saline. 4. The method of claim 1 wherein said member having an internal lumen includes a catheter; and wherein said step of introducing and positioning said member having said internal lumen includes a step of passing said catheter over said guiding member. 5. The method of claim 4 wherein said catheter is passed over said guiding member after said guiding member is positioned within said breast duet. 6. The method of claim 1 wherein said guiding member includes a guidewire; and the method further includes the step of dilating said breast duct before said guidewire is inserted in said breast duct. 7. The method of claim 6 wherein said dilating step includes the step of cannulating said breast duct. 8. The method of claim 7 wherein said cannulating step includes introducing a rigid duct-probe into said breast duct. 9. The method of claim 1 wherein said medical procedure includes the step of performing a ductoscopy. 10. The method of claim 9 wherein said ductoscopy includes the step of positioning an endoscope within said duct for viewing the interior of said duct from which said washings are removed. 11. A method for obtaining cellular material from within a breast duct during a medical procedure, said method includes the steps of: introducing a member having an internal lumen within said breast duct; introducing a fluid into said breast duct; introducing an endoscope within said breast duct; viewing the interior of said breast duct using said endoscope; removing cellular washings from said breast duct, wherein said step of removing said cellular washings includes removing said cellular washings through said internal lumen. 12. The method of claim 11 wherein said method further includes the step of positioning a catheter within said breast duct, and wherein said fluid is introduced into said breast duct through said catheter. 13. The method of claim 12 wherein said cellular washings are removed from within said duct through said catheter. 14. The method of claim 11 wherein said member having said internal lumen includes a device that cannulates an opening to said breast duct. 15. The method of claim 14 wherein said device includes a ductal probe. 16. The method of claim 11 wherein said medical procedure includes a ductoscopy. 17. A method for obtaining cellular material from within a breast duct during a medical procedure, said method... | The Regents of the University of California | A61B005/00
US06962569 | B2 | 2005-11-08 | Isometric system, method and apparatus | 2002-10-10 | System, method and apparatus for carrying out isometric exercises for either diagnostic purposes or therapeutic purposes. When employed in a diagnostic mode, the instrument is programmed to carry out standardized diagnostic regimens and during such regimens provides both visual and aural cues, carries out mathematical computations of force values and provides recordation of diagnostic data in archival memory. When employed in a therapeutic mode the apparatus may only be programmed within mandated therapeutic parameter limitations. During therapeutic trials, the user is visually and aurally cued throughout the test sequence and the therapeutic data evolved during the regimen is recorded and recoverable from archival memory. | 1. A system for carrying out an isometric diagnostic regimen by a user, comprising: a hand grip assembly including a load cell component responsive to compressive squeezing force applied by a hand of the user to derive a load value output corresponding with the value of said force; a display, responsive to a visual input signal to provide a visually perceptible display output; first and second control members hand actuateable to provide respective first and second control conditions; a controller including a processor and a memory operatively associated therewith and controllable to enable said processor and memory; and said processor, when enabled, entering a grip test mode wherein the processor is responsive to said first control condition to enter a maximum test mode to derive a visual input signal providing a first cue to said user to squeeze said hand grip with a maximum force capability, is responsive in the presence of said first cue and a first load value output occurring for a finite interval to derive said visual input signal providing a second cue displaying a first value of force corresponding with said first load value output, responsive to said first control condition in the presence of said second cue to enter said first value of force into said memory defining a maximum grip test, responsive to a subsequently occurring first control condition to derive a visual input signal providing a next occurring said first cue, responsive in the presence of said next occurring first cue and a second load value output occurring for a finite interval to derive a visual input signal providing a next said second cue displaying a second value of force corresponding with said second load value output, responsive to said first control condition in the presence of said next second cue to enter said second value of force into said memory defining another said maximum grip test, responsive to carrying out a mathematical analysis of said first and second values of force, and to derive a visual input signal providing a third visual cue displaying the results of said mathematical analysis, said processor being responsive in said grip test mode, in the presence of said visual input signal providing said second cue displaying said first value of force, and in the absence of said first control condition, to a next occurring said load value output occurring for a finite interval to derive a visual input signal providing a second cue displaying a substitute said first value of force corresponding with said next occurring load value output, and is responsive to said first control condition in the presence of said second cue displayed substitute first value of force to enter said substitute first value of force into said memory. | 1. A system for carrying out an isometric diagnostic regimen by a user, comprising: a hand grip assembly including a load cell component responsive to compressive squeezing force applied by a hand of the user to derive a load value output corresponding with the value of said force; a display, responsive to a visual input signal to provide a visually perceptible display output; first and second control members hand actuateable to provide respective first and second control conditions; a controller including a processor and a memory operatively associated therewith and controllable to enable said processor and memory; and said processor, when enabled, entering a grip test mode wherein the processor is responsive to said first control condition to enter a maximum test mode to derive a visual input signal providing a first cue to said user to squeeze said hand grip with a maximum force capability, is responsive in the presence of said first cue and a first load value output occurring for a finite interval to derive said visual input signal providing a second cue displaying a first value of force corresponding with said first load value output, responsive to said first control condition in the presence of said second cue to enter said first value of force into said memory defining a maximum grip test, responsive to a subsequently occurring first control condition to derive a visual input signal providing a next occurring said first cue, responsive in the presence of said next occurring first cue and a second load value output occurring for a finite interval to derive a visual input signal providing a next said second cue displaying a second value of force corresponding with said second load value output, responsive to said first control condition in the presence of said next second cue to enter said second value of force into said memory defining another said maximum grip test, responsive to carrying out a mathematical analysis of said first and second values of force, and to derive a visual input signal providing a third visual cue displaying the results of said mathematical analysis, said processor being responsive in said grip test mode, in the presence of said visual input signal providing said second cue displaying said first value of force, and in the absence of said first control condition, to a next occurring said load value output occurring for a finite interval to derive a visual input signal providing a second cue displaying a substitute said first value of force corresponding with said next occurring load value output, and is responsive to said first control condition in the presence of said second cue displayed substitute first value of force to enter said substitute first value of force into said memory. 2. The system of claim 1 in which: said processor is responsive in said test grip mode to said first control condition in the said presence of said second cue displayed substitute first value of force to enter the greater of said first... | MD Systems, Inc. | A61B005/103 A61B005/117
US06957094 | B2 | 2005-10-18 | Examination of scattering properties of biological tissue | 2002-11-19 | This invention is a scheme for monitoring a solute in a biological system comprising the steps of delivering light into a biological system (12) containing a solute, the light having a wavelength selected to be in a range wherein the solute is substantially non-absorbing; detecting at least first and second portions of the delivered light, the first portion having traveled through the biological system along one or more paths characterized by a first average path length, and the second portion having traveled through the biological system along one or more paths characterized by a second average path length that is greater than the first average path length; and comparing the first and second portions of the delivered light to monitor concentration of the solute in the biological system. Also described are schemes for monitoring low molecular weight polyhydroxy solutes, generally sugars (mannitol, fructose, sucrose, glucose, sorbitol), alcohols (methanol, ethanol, propanediol), and electrolytes (sodium, potassium, magnesium, calcium, and chloride ions). | 1. A system for in vivo characterization of biological tissue by measuring scattering properties of the tissue, comprising: a source-detector arrangement including at least one light source constructed to deliver into biological tissue light of at least one wavelength, at a known location, and at least one light detector constructed to provide at least one detected signal, wherein said detected signal depends on photon migration paths in the examined biological tissue; an oscillator constructed to generate a first carrier waveform at a frequency on the order of 108 Hz, wherein said light source is coupled to said oscillator and constructed to generate light modulated by said carrier waveform; said detector being constructed to detect said modulated light that has migrated in the tissue; a phase detector constructed to measure a phase shift of said detected light that has migrated in the tissue compared to said introduced light and provide said phase shift signal to said processor; and a processor arranged to determine, based on said detected light and said phase shift signal, scattering properties of the biological tissue. | 1. A system for in vivo characterization of biological tissue by measuring scattering properties of the tissue, comprising: a source-detector arrangement including at least one light source constructed to deliver into biological tissue light of at least one wavelength, at a known location, and at least one light detector constructed to provide at least one detected signal, wherein said detected signal depends on photon migration paths in the examined biological tissue; an oscillator constructed to generate a first carrier waveform at a frequency on the order of 108 Hz, wherein said light source is coupled to said oscillator and constructed to generate light modulated by said carrier waveform; said detector being constructed to detect said modulated light that has migrated in the tissue; a phase detector constructed to measure a phase shift of said detected light that has migrated in the tissue compared to said introduced light and provide said phase shift signal to said processor; and a processor arranged to determine, based on said detected light and said phase shift signal, scattering properties of the biological tissue. 2. The system of claim 1 wherein said processor is further arranged to calculate a scattering coefficient. 3. The system of claim 1 wherein said light source is arranged to deliver said light of said wavelength directly to the skin for in vivo examination of the biological tissue. 4. The system of claim 1 wherein said light detector is arranged to detect said light emitted from the biological tissue through the skin for in vivo examination of the biological tissue. 5. The system of claim 1 wherein said source-detector arrangement is constructed to change said photon migration paths in the examined tissue. 6. The system of claim 1 wherein said source-detector arrangement is constructed to deliver said light of said wavelength being selected with respect to a solute. 7. The system of claim 6 wherein said processor is further arranged to calculate concentration of said solute. 8. The system of claim 1 wherein said source-detector arrangement is constructed to deliver said light of said wavelength being selected with respect to an exogenous contrast agent bonded to a solute. 9. The system of claim 8 wherein said processor is further arranged to calculate concentration of said solute. 10. The system of claim 1 wherein said source-detector arrangement is constructed to deliver said light of said wavelength being selected with respect to one of the following: a low molecular weight carbohydrate, an alcohol, or an electrolyte. 11. The system of claim 1 wherein said source-detector arrangement is constructed to deliver said light of said wavelength being selected with respect to one of the following: mannitol, fructose, sucrose, glucose, propanediol, methanol, ethanol, sodium ion, potassium ion, and chloride ion. 12. The system of claim 1 wherein said source-detector arrangement is constructed to deliver said light to one of... | Non-Invasive Technology, Inc. | A61B005/00
US07054674 | B2 | 2006-05-30 | Method of and apparatus for investigating tissue histology | 2001-01-12 | Apparatus for monitoring the presence of one or more chromophores in a tissue sample, comprises a light source for projecting light to illuminate an area of such tissue sample, a photo-receptor for receiving light remitted by the illuminated area of tissue, and spectroscopic analyser means for monitoring the remitted light, a comparator having means for comparing variations in the intensity and spectral characteristics of the remitted light with respect to the intensity and spectral characteristics of the projected light at different wavelengths and with a record of the intensity and spectral characteristics of light remitted by a reference sample of such tissue and means for emitting a control signal in response to any such variations. Methods of analyzing tissue histology, especially skin histology, are discussed, and a mathematical model is proposed for the analysis and comparison of the remitted light with a reference sample. | 1. A method of monitoring the presence of one or more chromophores in a sample of biological tissue, which method comprises: illuminating an area of such tissue sample by projecting light from a light source; receiving light remitted by the illuminated area of tissue at a photo-receptor; spectroscopically analyzing the light projected from the light source and the remitted light received by the photo-receptor and utilizing said analysis to generate data indicative of differences between light projected from the light source and the remitted light; using the generated data to define a parameter of the tissue; processing the generated data using a predictive mathematical model of the optical properties of the biological tissue to normalize the defined parameter defined by the generated data to a standard value of that parameter; and measuring at least one further parameter of the tissue using said processed data processed to normalize the defined parameter defined by the generated data to a standard value. | 1. A method of monitoring the presence of one or more chromophores in a sample of biological tissue, which method comprises: illuminating an area of such tissue sample by projecting light from a light source; receiving light remitted by the illuminated area of tissue at a photo-receptor; spectroscopically analyzing the light projected from the light source and the remitted light received by the photo-receptor and utilizing said analysis to generate data indicative of differences between light projected from the light source and the remitted light; using the generated data to define a parameter of the tissue; processing the generated data using a predictive mathematical model of the optical properties of the biological tissue to normalize the defined parameter defined by the generated data to a standard value of that parameter; and measuring at least one further parameter of the tissue using said processed data processed to normalize the defined parameter defined by the generated data to a standard value. 2. A method according to claim 1 applied for non-invasive monitoring of the presence of one or more said chromophores in the tissue sample. 3. A method according to claim 2, applied for controlling a treatment which involves the irradiation of a region of tissue with treatment light of predetermined spectral characteristics, wherein the absorption characteristics of tissue supervening the region to be treated for the treatment light are measured and used in calculating a required exposure of the tissue to the treatment light. 4. A method according to claim 2, applied for predicting the outcome of a treatment which involves the irradiation of a region of tissue with treatment light of predetermined spectral characteristics, wherein the absorption characteristics for the treatment light of the tissue region to be treated and of tissue supervening the region to be treated are measured and used in calculating a required therapeutically effective exposure of the tissue to the treatment light, and the required exposure and the absorption characteristics of the supervening tissue are used to predict potential destruction or scarring of the supervening tissue by such exposure. 5. A method according to claim 1 applied for endoscopic monitoring of the presence of one or more said chromophores in the tissue sample. 6. A method according to claim 1, further comprising: processing data processed to normalize the defined parameter defined by the generated data to a standard value generated data using a predictive mathematical model of the optical properties of the biological tissue to normalize a further parameter; and measuring at least one parameter of the tissue using said processed data processed to normalize the further parameter to a standard value. 7. A method in according to claim 1 wherein said sample of biological tissue comprises a sample of epithelial tissue having a thickness of papillary dermis, wherein using the generated data to define ... | Astron Clinica Limited | A61B005/00
US07181297 | B1 | 2007-02-20 | System and method for delivering customized audio data | 1999-09-28 | A method for providing customized audio data products that includes the steps of (1) storing a machine readable hearing profile for a customer, (2) accepting via a data input device or data network for example, machine readable orders from the customer for a particular audio data product, (3) associating the hearing profile for the customer with the particular audio data product, (4) modifying the particular audio data product according to the hearing profile to produce a customized audio data product, (5) delivering the customized audio data product to the customer, for example via a network or by a machine readable storage medium, (6) accepting feedback from the customer concerning the subjective performance of the customized audio data product, and (7) modifying the hearing profile in response to the feedback is provided. A system that includes a source of an audio data product, such as an Internet accessible registry, and processing resources that are configured to receive a request for the audio data product, and to associate the audio data product with a customer hearing profile is provided. | 1. A method for producing customized audio data, comprising: executing an interface at a server or network of servers in a data network for selection of audio data products from, and for access to, a catalog of audio data products;providing a graphical user interface to a customer at a customer terminal, the graphical user interface presenting the catalog of audio data products from which a customer is able to select an audio data product;receiving, at the server or network of servers, a request for a particular audio data product from the catalog of audio data products wherein the request is made by the customer via the graphical user interface;receiving, at the server or network of servers, a customer hearing profile associated with the customer from a machine readable registry of customer hearing profiles, the customer hearing profile comprising data adapted to be applied to transform the particular audio data product according to hearing a characteristic of the customer to produce a customized audio data product;obtaining, at the server or network of servers, the particular audio data product; andassociating, at the server or network of servers, the particular audio data product with the customer hearing profile producing said customized audio data product in response to the particular audio data product and the associated customer hearing profile, and delivering the customized audio data product to the customer terminal for use by the customer. | 1. A method for producing customized audio data, comprising: executing an interface at a server or network of servers in a data network for selection of audio data products from, and for access to, a catalog of audio data products;providing a graphical user interface to a customer at a customer terminal, the graphical user interface presenting the catalog of audio data products from which a customer is able to select an audio data product;receiving, at the server or network of servers, a request for a particular audio data product from the catalog of audio data products wherein the request is made by the customer via the graphical user interface;receiving, at the server or network of servers, a customer hearing profile associated with the customer from a machine readable registry of customer hearing profiles, the customer hearing profile comprising data adapted to be applied to transform the particular audio data product according to hearing a characteristic of the customer to produce a customized audio data product;obtaining, at the server or network of servers, the particular audio data product; andassociating, at the server or network of servers, the particular audio data product with the customer hearing profile producing said customized audio data product in response to the particular audio data product and the associated customer hearing profile, and delivering the customized audio data product to the customer terminal for use by the customer. 2. The method of claim 1, wherein the customer hearing profile comprises coefficients of a transfer function to transform the audio data product according a psycho-acoustic characteristic of a customer. 3. The method of claim 1, wherein the customer hearing profile comprises an identifier of transfer function to transform the audio data product according a hearing characteristic of a customer. 4. The method of claim 1, wherein the customer hearing profile comprises an identifier of a transfer function and coefficients of the identified transfer function to transform the audio data product according a hearing characteristic of a customer. 5. The method of claim 1, wherein the machine readable registry comprises customer hearing profiles corresponding to registered customers, and further including reading, modifying and writing particular customer hearing profiles in the plurality to maintain the registry. 6. The method of claim 1, wherein the catalog of audio data products comprise pre-recorded audio tracks. 7. The method of claim 1, wherein the catalog of audio data products comprise pre-recorded audio music tracks. 8. The method of claim 1, wherein the catalog of audio data products comprise live audio music streams. 9. The method of claim 1, wherein the catalog of audio data products comprise pre-recorded audio voice tracks. 10. The method of claim 1, wherein the catalog of audio data products comprise live audio voice streams. 11. The method of claim 1, wherein the catalog of audio data p... | Sound ID | G06F017/00 H04R029/00 A61B005/00
US06985764 | B2 | 2006-01-10 | Flex circuit shielded optical sensor | 2002-05-02 | A flex circuit optical sensor has an integrated Faraday shield. A conductive trace layer disposed on a substrate is used to form a conductive grid which shields the face of a photodetector. A conductive ink layer is formed on a substrate side opposite the trace layer. The back and sides of the detector are shielded by flex circuit flaps that have the conductive ink layer but substantially exclude the trace layer so as to fold over and closely adhere to the detector body. The shielded substrate flaps advantageously eliminate a separate detector shield, which is typically fabricated with an etched copper part that must be attached to a flex circuit before mounting the detector. | 1. An optical sensor comprising: a flex circuit substrate having a first side and an opposite second side;a trace layer disposed on said first side having a pattern of conductors so as to electrically connect with an emitter and a detector,a conductive grid portion of said trace layer having at least one aperture so as to pass optical radiation from said emitter to said detector, said trace layer configured to mount said detector proximate said grid portion;a shield layer disposed on said second side, said substrate having a plurality of foldable portions including said shield layer and substantially excluding said trace layer, said plurality of foldable portions being configured to substantially enclose said detector so as to form a Faraday shield in conjunction with said conductive grid portion. | 1. An optical sensor comprising: a flex circuit substrate having a first side and an opposite second side;a trace layer disposed on said first side having a pattern of conductors so as to electrically connect with an emitter and a detector,a conductive grid portion of said trace layer having at least one aperture so as to pass optical radiation from said emitter to said detector, said trace layer configured to mount said detector proximate said grid portion;a shield layer disposed on said second side, said substrate having a plurality of foldable portions including said shield layer and substantially excluding said trace layer, said plurality of foldable portions being configured to substantially enclose said detector so as to form a Faraday shield in conjunction with said conductive grid portion. 2. The optical sensor according to claim 1 wherein said plurality of foldable portions comprise a plurality of flaps adapted to adhere to said detector. 3. The optical sensor according to claim 1 further comprising: a stock material retaining said substrate and having a first wrap with a first end and a second wrap with a second end,each of said wraps adapted for finger attachment of said substrate,said wraps configured so that, in a wrapped position, said first end covers said second end and said second end terminates distal to a person's palm. 4. An optical sensor assembly method comprising the steps of: fabricating an elongated flex circuit having a component side, an opposite side and a plurality of foldable portions;creating a plurality of traces on said component side substantially excluding said plurality of foldable portions;creating a conductive layer on said opposite side substantially including said plurality of foldable portions;mounting a detector and an emitter to said component side, said traces providing electrical connections for said detector and said emitter; anddisposing said plurality of foldable portions around said detector so as to shield said detector from electromagnetic interference. 5. The optical sensor assembly method according to claim 4 wherein said step of fabricating the plurality of foldable portions step comprises the substep of forming a back flap and a plurality of side flaps. 6. The optical sensor assembly method according to claim 5 wherein said disposing step comprises the substeps of adhering said plurality of side flaps to first portions of said detector and adhering said back flap to second portions of said detector. 7. The optical sensor assembly method according to claim 4 further comprising the step of creating a conductive grid portion of at least one of said traces. 8. The optical sensor assembly method according to claim 7 wherein said mounting step comprises the substep of positioning said detector proximate said grid portion so as to shield said detector from electromagnetic interference. 9. A noninvasive disposable optical sensor which outputs a signal usable to determine one or more physiolog... | Masimo Corporation | A61B005/00 H05K003/02
US07058441 | B2 | 2006-06-06 | Optical measuring system for diagnosing ear conditions | 2001-11-15 | A system for detecting and diagnosing ear related conditions. The system includes a device capable of obtaining a spectrum of reflected light from an ear of a subject and a processing unit in connection with the device, which is capable of translating the obtained spectrum of reflected light to one or more output values related to the condition of the ear. The invention further provides a method for detecting and diagnosing ear related conditions including the steps of illuminating inside the ear; inserting a device to the ear canal capable of conveying at least one spectrum of reflected light from the ear to a processing unit; and activating the processing unit thereby translating at least one spectrum of reflected light provided at the time of activating to one or more output values related to the condition of the ear. | 1. A system for use in detecting and diagnosing ear related conditions, the system comprising: (a) a device capable of obtaining a spectrum of reflected light from an ear of a subject; and(b) a processing unit in connection with said device which is capable of translating the obtained spectrum of reflected light to one or more output values related to the condition of the ear, the processing unit comprising a spectral analytical instrument in the form of a spectrometer, which receives said spectrum of reflected light and produces an analog signal indicative thereof, the processing unit being preprogrammed with certain models capable of analyzing the analog signal and converting it to at least one digital value presenting a spectrum vector indicative of a determined state of the ear, processing the spectrum vector by comparing to at least one predetermined reference value, and applying a classification model to the processed spectrum vector to thereby classify the determined state of the ear and generate output data indicative of one of the following ear conditions: normal condition, serous otitis media condition, and otitis media condition. | 1. A system for use in detecting and diagnosing ear related conditions, the system comprising: (a) a device capable of obtaining a spectrum of reflected light from an ear of a subject; and(b) a processing unit in connection with said device which is capable of translating the obtained spectrum of reflected light to one or more output values related to the condition of the ear, the processing unit comprising a spectral analytical instrument in the form of a spectrometer, which receives said spectrum of reflected light and produces an analog signal indicative thereof, the processing unit being preprogrammed with certain models capable of analyzing the analog signal and converting it to at least one digital value presenting a spectrum vector indicative of a determined state of the ear, processing the spectrum vector by comparing to at least one predetermined reference value, and applying a classification model to the processed spectrum vector to thereby classify the determined state of the ear and generate output data indicative of one of the following ear conditions: normal condition, serous otitis media condition, and otitis media condition. 2. The system of claim 1, wherein said spectrometer is sensitive to wavelengths of approximately between 400 nm to 1200 nm. 3. The system of claim 1, wherein said spectral analytical, instrument includes at least one filter. 4. The system of claim 1, wherein said processing unit further comprises a microprocessor interfaced with said converter. 5. The system of claim 4, wherein said microprocessor comprises an accessible memory. 6. The system of claim 1, wherein said processing unit comprises a display for displaying said one or more output values. 7. The system of claim 6, wherein said display includes an LCD display. 8. The system of claim 1, wherein said one or more output values indicate the health of the ear. 9. The system of claim 8, wherein said one or more output values indicate whether there is otitis media or serous otitis media. 10. The system of claim 9, wherein the statistical confidence of a decision relating to the existence of otitis media or serous otitis media is additionally indicated. 11. The system of claim 1, wherein said at least one reference value is at least one digital value resulting from diagnosing the other ear of said subject with said system. 12. The system of claim 1, wherein said at least one reference value is at least one digital value resulting from a diagnosis of a healthy ear with said system. 13. The system of claim 1, wherein said at least one reference value is a statistical range of values resulting from healthy ears. 14. The system of claim 1, wherein said at least one reference value is a statistical range of values resulting from ears with otitis media. 15. The system of claim 1, wherein said at least one reference value is a statistical range of values resulting from ears with serous otitis media. 16. The system of claim 1, wherein said processing u... | Chameleon Medical Innovation, Ltd. | A61B005/00
US07198600 | B2 | 2007-04-03 | Body temperature managing method and device, storage medium, body temperature managing system, and program | 2001-05-31 | Data measured by a thermometer is temporarily stored in a connection terminal, then transmitted to a server from the connection terminal via a network, and stored. The server makes analyses based on the stored temperature data, and transmits the analysis results. Thus, a body temperature managing and analyzing method, providing information unobtainable with a thermometer alone, can be provided by managing and analyzing body temperature over a network. | 1. A body temperature managing method, comprising: a body temperature data obtaining step for obtaining body temperature data;a body temperature data storing step for storing said body temperature data obtained in said obtaining step;a body temperature data analyzing step for analyzing body temperature data based on said body temperature data stored in said storing step;an analyzed data transmitting step for transmitting the body temperature data analyzed in said analyzing step to an outside display device,wherein a display format of said analyzed body temperature data differs depending on said outside display device;a presenting step for presenting a list of facilities according to the analysis results analyzed in said body temperature data analyzing step;a counting step for counting a number of reservations, in the event that reservations have been made at an arbitrary facility from the facility list presented in said presenting step; anda cash-back step for giving cash back based on the counted number of reservations. | 1. A body temperature managing method, comprising: a body temperature data obtaining step for obtaining body temperature data;a body temperature data storing step for storing said body temperature data obtained in said obtaining step;a body temperature data analyzing step for analyzing body temperature data based on said body temperature data stored in said storing step;an analyzed data transmitting step for transmitting the body temperature data analyzed in said analyzing step to an outside display device,wherein a display format of said analyzed body temperature data differs depending on said outside display device;a presenting step for presenting a list of facilities according to the analysis results analyzed in said body temperature data analyzing step;a counting step for counting a number of reservations, in the event that reservations have been made at an arbitrary facility from the facility list presented in said presenting step; anda cash-back step for giving cash back based on the counted number of reservations. 2. A body temperature managing method according to claim 1, wherein said body temperature data is basal body temperature data. 3. A body temperature managing method according to claim 1, wherein the cash-back step gives back cash to the individual making reservations, depending on said counted number. 4. A body temperature managing method according to claim 1, wherein said analyzed data transmitting step comprises a second transmitting step for transmitting analyzed data to a facility selected from said facility list presented in said presenting step, and an obtaining step for obtaining results diagnosed based on said body temperature data at said facility. 5. A body temperature managing method according to claim 1, wherein the destination of transmission of said analyzed data transmitted in said analyzed data transmitting step is one of at least a personal computer, cellular phone, or a portable terminal. 6. A body temperature managing method according to claim 1, wherein data obtained in said body temperature data obtaining step is enciphered data. 7. A body temperature managing method according to claim 1, wherein said body temperature data analyzing step further comprises a deciphering step for deciphering said body temperature data. 8. A body temperature managing method according to claim 1, wherein the transmitted body temperature data includes text data or graph and text data, depending on a size of a display area of the outside display device. 9. A body temperature managing method, comprising: a body temperature data obtaining step for obtaining body temperature data;a body temperature data storing step for storing said body temperature data obtained in said obtaining step;a body temperature data analyzing step for analyzing body temperature data based on said body temperature data stored in said storing step;an analyzed data transmitting step for transmitting the body temperature data analyzed in said analyzing s... | Canon Kabushiki Kaisha | A61B005/00
US07222079 | B1 | 2007-05-22 | Method and system for generating statistically-based medical provider utilization profiles | 1999-11-10 | A method and system for analyzing historical medical provider billings to statistically establish a normative utilization profile. Comparison of a medical provider's utilization profile with a normative profile is enabled. Based on historical treatment patterns and a fee schedule, an accurate model of the cost of a specific medical episode can be created. Various treatment patterns for a particular diagnosis can be compared by treatment cost and patient outcome to determine the most cost-effective treatment approach. It is also possible to identify those medical providers who provide treatment that does not fall within the statistically established treatment patterns or profiles. | 1. A computer-implemented process for processing medical claims including the steps of: (a) reading a medical claim data, input as at least one of a plurality of data records, into a computer memory;(b) validating each of the at least one of a plurality of data records for at least one of a diagnosis code and a treatment code;(c) reading at least one pre-defined relationship between the at least one of a diagnosis code and a treatment code in the validated at least one of a plurality of data records and pre-defined episode treatment categories; and(d) grouping the validated at least one of a plurality of data records to an episode treatment category based upon the pre-defined relationship, each episode treatment category having a dynamic time window defining a time period which validated at least one of plurality of data records may be grouped to an episode treatment category. | 1. A computer-implemented process for processing medical claims including the steps of: (a) reading a medical claim data, input as at least one of a plurality of data records, into a computer memory;(b) validating each of the at least one of a plurality of data records for at least one of a diagnosis code and a treatment code;(c) reading at least one pre-defined relationship between the at least one of a diagnosis code and a treatment code in the validated at least one of a plurality of data records and pre-defined episode treatment categories; and(d) grouping the validated at least one of a plurality of data records to an episode treatment category based upon the pre-defined relationship, each episode treatment category having a dynamic time window defining a time period which validated at least one of plurality of data records may be grouped to an episode treatment category. | Ingenix, Inc. | A61B005/00 G06F019/00
US07177677 | B2 | 2007-02-13 | Nerve proximity and status detection system and method | 2002-10-16 | A method and system is disclosed for detecting proximity of a nerve to a proximity electrode and status of nerve. The method of detecting the proximity of the nerve to the proximity electrode applies a first electrical signal with multiple current levels to a calibration electrode. The method applies a second electrical signal with multiple current levels to the proximity electrode. Based on the response of the nerve to the first and second electrical signals, the method determines the proximity of the proximity electrode to the nerve. The method determines the status of the nerve by applying the first electrical signal with multiple signal levels to the calibration electrode. | 1. A method of intraoperatively determining nerve health to assess success of a surgical procedure comprising the steps of: (a) intraoperatively applying an electrical signal to a conductive element located near the nerve;(b) intraoperatively measuring the responses of the nerve to the applied electrical signal;(c) intraoperatively comparing successive responses of the nerve to the applied electrical signal; and(d) intraoperatively determining the nerve health based on the comparisons of the successive responses of the nerve to the applied electrical signal. | 1. A method of intraoperatively determining nerve health to assess success of a surgical procedure comprising the steps of: (a) intraoperatively applying an electrical signal to a conductive element located near the nerve;(b) intraoperatively measuring the responses of the nerve to the applied electrical signal;(c) intraoperatively comparing successive responses of the nerve to the applied electrical signal; and(d) intraoperatively determining the nerve health based on the comparisons of the successive responses of the nerve to the applied electrical signal. 2. The method of claim 1, wherein step c) includes mapping the responses of the nerve to the applied electrical signal for a fixed latency, and wherein step d) includes determining the nerve health based on mapping the responses of the nerve to the applied electrical signal for the fixed latency. 3. The method of claim 2, wherein step a) includes placing a conductive element near the nerve, and applying an electrical signal having a plurality of predetermined current levels to the conductive element. 4. The method of claim 1, wherein each response is an EMG. 5. The method of claim 1, wherein each response is an EMG measured at a muscle physiologically coupled to the nerve. 6. The method of claim 3, wherein step a) includes placing a conductive element in the epidural space above the dura of the spinal column in which the nerve is located. 7. The method of claim 1, wherein step d) of determining nerve health is based on an interpretation of a slope of mapped responses of the nerve to the applied electrical signal. 8. The method of claim 1, wherein step d) of determining nerve health is based on an interpretation of a comparison between mapped responses of the nerve to the applied electrical signal and known mappings for at least one of healthy and unhealthy nerves. 9. The method of claim 1, wherein the conductive element is coupled to a medical device. 10. The method of claim 9, wherein the medical device comprises at least one of a cannula, a catheter, and a needle. 11. A method of determining nerve status comprising the steps of: (a) applying an electrical signal to a conductive element located near the nerve;(b) measuring the responses of the nerve to the applied electrical signal; and(c) determining the nerve status based on the responses of the nerve to the applied electrical signal, including the sub-steps of (i) mapping the responses of the nerve to the applied electrical signal for a fixed latency; and (ii) determining the nerve status based on the mapping of the responses of the nerve to the applied electrical signal for the fixed latency, wherein the fixed latency is a function of peaks in each response of the nerve to the applied electrical signal. 12. The method of claim 11, wherein each response is an EMG. 13. The method of claim 11, wherein each response is an EMG measured at a muscle physiologically coupled to the nerve. 14. An article of manufacture for use in intr... | NuVasive, Inc. | A61B005/04
US07077812 | B2 | 2006-07-18 | Apparatus and method for palpographic characterization of vulnerable plaque and other biological tissue | 2002-11-22 | The present invention discloses a device and methods for characterizing vulnerable plaque and cancer tissue by measuring changes in tissue elasticity compared to that of normal tissue. The system includes a catheter with an expandable element at a proximal end. The expandable element is equipped with pressure sensors to detect changes in tissue elasticity and can be additionally equipped with sensors that detect tissue temperature and pH. For arterial tissue or tissue lining a body cavity, the device can also be equipped with width gauges that measure the diameter of the artery lumen or the width of any section of the body cavity. The distal end of the catheter may be attached to a motorized pullback device connected to a computer. Data collected by the device sensors are sent to the computer for processing and analysis. | 1. A medical device for measuring the stiffness of biological tissue in a patient, comprising: (a) a catheter having a proximal end and a distal end;(b) a shaft extending through the catheter from the proximal end to the distal end, the shaft comprising a guide wire adapted for moving the catheter in longitudinal and circumferential directions;(c) an expandable element connected to the proximal end of the catheter; and(d) at least one pressure sensor attached to the expandable element, the pressure sensor being adapted to measure elasticity of organic tissue. | 1. A medical device for measuring the stiffness of biological tissue in a patient, comprising: (a) a catheter having a proximal end and a distal end;(b) a shaft extending through the catheter from the proximal end to the distal end, the shaft comprising a guide wire adapted for moving the catheter in longitudinal and circumferential directions;(c) an expandable element connected to the proximal end of the catheter; and(d) at least one pressure sensor attached to the expandable element, the pressure sensor being adapted to measure elasticity of organic tissue. 2. The device of claim 1, wherein the pressure sensor is a sensor selected from the group consisting of a silicone membrane sensor, an ultrasound sensor, and a force gauge sensor. 3. The device of claim 2, wherein the ultrasound sensor comprises an ultrasound system that emits and receives ultrasound signals. 4. The device of claim 1, wherein the expandable element is expandable at a deployment step and retractable during an insertion step of the catheter. 5. The device of claim 1, wherein the pressure sensor is adapted to measure pressure when the expandable element contacts a tissue wall. 6. The device of claim 1, wherein the expandable element comprises a plurality of wires in a basket arrangement. 7. The device of claim 6, wherein the pressure sensors attached to the basket arrangement of wires are cantilever sensors. 8. The device of claim 1, wherein the expandable element comprises a plurality of wires in a radial arrangement. 9. The device of claim 8, wherein the pressure sensors attached to the radial arrangement of wires are cantilever sensors. 10. The device of claim 1, wherein the distal end of the catheter is connected to a motorized system for deployment and pullback of the catheter. 11. The device of claim 10, wherein the motorized system is connected to and controlled by a computer and wherein the computer is adapted to analyze data from the sensors on the device with a software reconstruction program. 12. The device of claim 11, wherein the device has at least one temperature sensor and at least one pH sensor attached to the expandable element of the catheter. 13. The device of claim 12, wherein the temperature sensor is selected from the group consisting of a thermocouple sensor, a thermister sensor and an infrared sensor. 14. The device of claim 13, wherein the infrared sensor comprises an optical fiber. 15. The device of claim 12, wherein the device has width gauges attached to the catheter shaft for determining diameter of an artery lumen. 16. The device of claim 15, wherein the software reconstruction program is adapted to create three-dimensional image maps of the lumen using data sent from the width gauges, temperature sensors and pressure sensors. 17. The device of claim 1, wherein the device has at least one pressure sensor, at least one temperature sensor and at least one pH sensor attached to the shaft of the catheter. 18. The device of claim 17, w... | The Board Regents of the University System | A61B005/117 A61B005/103
US07217245 | B1 | 2007-05-15 | Noninvasive methods for detecting abnormalities in a subject such as disease or dysfunction | 2000-04-14 | The systems, methods and apparatus of the present invention comprise noninvasive methods of measuring temperature changes and rates of temperature change in selected body tissues and fluids as a result of absorption and/or dissipation of externally applied heat for the purpose of detecting and monitoring disease or dysfunction, and for preparing diagnostic images of the tissues and related areas from such measurements. Some embodiments of the present invention monitor temperatures of heated tissues directly while other embodiments may measure temperatures of tissues which are heated through flow of heated fluid therethrough. | 1. A noninvasive method for detecting the flow of body fluids in a subject comprising: noninvasively warming urine in the subject's bladder with an external energy source of ultrasound energy; andnoninvasively detecting the flow of said warmed urine from said bladder to the subject's kidney by measuring a temperature change in said kidney. | 1. A noninvasive method for detecting the flow of body fluids in a subject comprising: noninvasively warming urine in the subject's bladder with an external energy source of ultrasound energy; andnoninvasively detecting the flow of said warmed urine from said bladder to the subject's kidney by measuring a temperature change in said kidney. 2. The method of claim 1 wherein said detecting comprises determining an increase in temperature in said kidney indicative of the presence of said warmed urine. 3. The method of claim 1 further comprising determining a rate of temperature change in said kidney. 4. A noninvasive method for detecting the flow of body fluids in a subject comprising: noninvasively warming gastrointestinal fluid in the subject's stomach with an external energy source of ultrasound energy; andnoninvasively detecting the flow of said warmed gastointestinal fluid from said stomach to the subject's esophagus or duodenum by measuring a temperature change in said esophagus or duodenum. 5. The method of claim 4 wherein said detecting comprises determining an increase in temperature in said esophagus or duodenum indicative of the presence of said warmed gastointestinal fluid. 6. The method of claim 4 further comprising determining a rate of temperature change in said esophagus or duodenum. 7. A noninvasive method for detecting the flow of body fluids in a subject comprising: noninvasively warming blood in the subject's heart with an external energy source of ultrasound energy; andnoninvasively detecting the flow of said warmed blood from said heart to the subject's vascular tree by measuring a temperature change in said vascular tree. 8. The method of claim 7 wherein said detecting comprises determining an increase in temperature in said vascular tree indicative of the presence of said warmed blood. 9. The method of claim 7 further comprising determining a rate of temperature change in said vascular tree. 10. A noninvasive method for detecting the flow of body fluids in a subject comprising: noninvasively warming cerebrospinal fluid in the subject's spinal column or cranial vault with an external energy source of ultrasound energy; andnoninvasively detecting the flow of said warmed cerebrospinal fluid in said spinal column or cranial vault by measuring a temperature change in said spinal column or cranial vault. 11. The method of claim 10 wherein said detecting comprises determining an increase in temperature in said spinal column or cranial vault indicative of the presence of said warmed cerebrospinal fluid. 12. The method of claim 10 further comprising determining a rate of temperature change in said spinal column or cranial vault. | University of Utah Research Foundation | A61B005/00
US07186216 | B2 | 2007-03-06 | Quality control system in disease management services for checking adherence to therapy | 2002-08-07 | Quality control system in Disease Management Services involving education, development and motivation of patients suffering from widespread chronic diseases, such as diabetes, asthma or hypertension, possibly involving telemonitoring of critical body values and resultant early identification and avoidance of risk situations, where an automatic delivery monitoring facility is provided which monitors the receipt of all messages and information by the patients and, if appropriate, informs a control station in order to eliminate access problems. | 1. An automated quality control system for a disease management service (DMS) whose patients suffer from diseases requiring long-term attention by the patient, the system comprising: a delivery unit that sends training material to a patient suffering from a disease requiring long-term attention by the patient, the training material including educational material relating to and aiding patient self-treatment and related to the disease and a comprehension test that is to be taken by the patient after having received the educational material and that determines whether the patient has understood the educational material;a receipt acknowledgment unit that automatically confirms receipt of the training material by the patient;a test evaluation unit that automatically evaluates answers to the comprehension test that has been taken by the patient and returned to the DMS; anda telemonitoring system that monitors body parameter values of the patient. | 1. An automated quality control system for a disease management service (DMS) whose patients suffer from diseases requiring long-term attention by the patient, the system comprising: a delivery unit that sends training material to a patient suffering from a disease requiring long-term attention by the patient, the training material including educational material relating to and aiding patient self-treatment and related to the disease and a comprehension test that is to be taken by the patient after having received the educational material and that determines whether the patient has understood the educational material;a receipt acknowledgment unit that automatically confirms receipt of the training material by the patient;a test evaluation unit that automatically evaluates answers to the comprehension test that has been taken by the patient and returned to the DMS; anda telemonitoring system that monitors body parameter values of the patient. 2. The quality control system of claim 1, wherein said test evaluation unit automatically generates a warning based on a negative automatic evaluation of the answers to the comprehension test by said test evaluation unit, the warning being sent from the DMS to at least one of the patient, a care giver for the patient, and a health insurer for the patient. 3. The quality control system of claim 2, wherein said delivery unit sends further training material to the patient based on a positive automatic evaluation of the answers to the comprehension test by said test evaluation unit. 4. The quality control system of claim 1, wherein said delivery unit sends further training material to the patient based on a positive automatic evaluation of the answers to the comprehension test by said test evaluation unit. 5. The quality control system of claim 1, further comprising an expert system operating in accordance with predetermined rules, said expert system generating rules for generation of questions for the comprehension test based on the disease suffered by the patient and on a background of the patient, so that the questions in the comprehension test vary from patient to patient. 6. The quality control system of claim 1, wherein said test evaluation unit operates dynamically and adds or amends a later question in the comprehension test if the patient provides an incorrect response to a question in the comprehension test. 7. The quality control system of claim 1, wherein said test evaluation unit provides a quantified score from the evaluation of the answers to the comprehension test, and wherein the quantified score is used by the quality control system to produce a dynamically structured treatment plan for the patient. 8. The quality control system of claim 7, wherein said test evaluation unit generates a warning based on the quantified score, the warning being sent from the DMS to at least one of the patient, a care giver for the patient, and a health insurer for the patient. 9. The quality control system ... | Siemens Aktiengesellschaft | A61B005/00
US06999904 | B2 | 2006-02-14 | Variable indication estimator | 2002-08-05 | A variable indication estimator which determines an output value representative of a set of input data. For example, the estimator can reduce input data to estimates of a desired signal, select a time, and determine an output value from the estimates and the time. In one embodiment, the time is selected using one or more adjustable signal confidence parameters determine where along the estimates the output value will be computed. By varying the parameters, the characteristics of the output value are variable. For example, when input signal confidence is low, the parameters are adjusted so that the output value is a smoothed representation of the input signal. When input signal confidence is high, the parameters are adjusted so that the output value has a faster and more accurate response to the input signal. | 1. A signal processor capable of reducing unwanted fluctuations in oxygen saturation values while substantially tracking the occurrence and severity of desaturation events, wherein the oxygen saturation values were derived by processing absorption signals representative of light attenuated by body tissue carrying pulsing blood, the signal processor comprising: a buffer capable of storing a plurality of input oxygen saturation values corresponding to a plurality of time values;a signal confidence indicator;a selected time value based upon said signal confidence indicator; andan estimator capable of determining an oxygen saturation output value at the selected time value, said oxygen saturation output value comprising an estimate for the plurality of input oxygen saturation values;wherein the estimator is an oxygen saturation versus time curve fit to the input oxygen saturation values within the buffer;wherein the oxygen saturation output value is a point along said curve fit at the selected time value; andwherein the selected time value is between an average of the time values and a maximum of the time values. | 1. A signal processor capable of reducing unwanted fluctuations in oxygen saturation values while substantially tracking the occurrence and severity of desaturation events, wherein the oxygen saturation values were derived by processing absorption signals representative of light attenuated by body tissue carrying pulsing blood, the signal processor comprising: a buffer capable of storing a plurality of input oxygen saturation values corresponding to a plurality of time values;a signal confidence indicator;a selected time value based upon said signal confidence indicator; andan estimator capable of determining an oxygen saturation output value at the selected time value, said oxygen saturation output value comprising an estimate for the plurality of input oxygen saturation values;wherein the estimator is an oxygen saturation versus time curve fit to the input oxygen saturation values within the buffer;wherein the oxygen saturation output value is a point along said curve fit at the selected time value; andwherein the selected time value is between an average of the time values and a maximum of the time values. 2. The signal processor of claim 1, wherein the curve fit is a least-mean-square linear fit;wherein a slope of the least-mean-square linear fit corresponds to a ratio of a crossvariance of the plurality of input oxygen saturation values with the plurality of time values and an autovariance of the plurality of time values;wherein an intercept of the least-mean-square linear fit corresponds to an average of the plurality of input oxygen saturation values. 3. The signal processor of claim 1, wherein the estimator comprises one of a smoother, a filter, a forward predictor, and a backward predictor. 4. The signal processor of claim 1, wherein the estimator comprises a variable mode averager. 5. The signal processor of claim 4, wherein the signal confidence indicator corresponds to a level of confidence that one or more of the oxygen saturation values represent a value of a desired signal. 6. The signal processor of claim 5, wherein the estimator approximates a smoother when the signal confidence indicator is low. 7. The signal processor of claim 5, wherein the estimator approximates a filter when the signal confidence indicator is high. 8. The signal processor of claim 5, wherein the oxygen saturation output value includes a faster response to the oxygen saturation values when the signal confidence indicator is high. 9. The signal processor of claim 5, wherein the oxygen saturation output value includes a more accurate response to the oxygen saturation values when the signal confidence indicator is high. 10. The signal processor of claim 1, wherein the estimator is capable of causing an alarm to be activated when the oxygen saturation output value is below a threshold value, and wherein the estimator lowers an amount of false alarms through adjustment of the signal confidence indicator. 11. The signal processor of claim 1, wherein the o... | Masimo Corporation | A61B005/00
US07186400 | B2 | 2007-03-06 | Multidentate aza ligands able to complex metal ions and the use thereof in diagnostics and therapy | 2002-07-10 | Compounds of general formula (I): in which: R1 is hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl, arylalkyl, or two R1, taken together, form a straight or cyclic C2–C10 alkylene group or an ortho disubstituted arylene; R2 is hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl or aryl alkyl optionally substituted with functional groups which allow conjugation with a suitable molecule able to interact with physiological systems; R3, R4 and R5 are hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl, arylalkyl; and their chelates with bi-trivalent ions of the metal elements having atomic number ranging between 20 and 31, 39, 42, 43, 44, 49, and between 57 and 83, and radioisotopes chosen among 203Pb 67Ga 68Ga 72As 111In 113In 90Y 97Ru 62Cu 64Cu 52Fe 52mMn 140La 175Yb 153Sm 166Ho 149Pm 177Lu 142Pr 159Gd 212Bi 47Sc 149Pm 67Cu, 111Ag, 199Au, 161Tb and 51Cr as well as the salts thereof with physiologically compatible bases or acids; | 1. A compound of general formula (I):in which: R1 is hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl, arylalkyl; or the two R1 groups, taken together, form a straight or cyclic C2–C10 alkylene group or an ortho-disubstituted arylene;R2 is hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl or arylalkyl optionally substituted with functional groups which allow conjugation with a suitable molecule able to interact with physiological systems;R3, R4 and R5, which are the same or different, are hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl, arylalkyl;FG, which are the same or different, are carboxy, —PO3H2 or —RP(O)OH groups, wherein R is hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl, arylalkyl; or their chelates with bi-trivalent ions of the metal elements having atomic numbers ranging between 20 and 31, 39, 42, 43, 44, 49, or between 57 and 83, and radioisotopes 203Pb, 67Ga, 68Ga, 72As, 111In, 113In, 90Y, 97Ru, 62Cu, 64Cu, 52Fe, 52Mn, 140La, 175Yb, 153Sm, 166Ho, 149Pm, 177Lu, 142Pr, 159Gd, 212Bi, 47Sc, 149Pm, 67Cu, 111Ag, 199Au, 161Tb, and 51Cr as well as salts thereof with physiologically compatible bases or acids. | 1. A compound of general formula (I):in which: R1 is hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl, arylalkyl; or the two R1 groups, taken together, form a straight or cyclic C2–C10 alkylene group or an ortho-disubstituted arylene;R2 is hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl or arylalkyl optionally substituted with functional groups which allow conjugation with a suitable molecule able to interact with physiological systems;R3, R4 and R5, which are the same or different, are hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl, arylalkyl;FG, which are the same or different, are carboxy, —PO3H2 or —RP(O)OH groups, wherein R is hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl, arylalkyl; or their chelates with bi-trivalent ions of the metal elements having atomic numbers ranging between 20 and 31, 39, 42, 43, 44, 49, or between 57 and 83, and radioisotopes 203Pb, 67Ga, 68Ga, 72As, 111In, 113In, 90Y, 97Ru, 62Cu, 64Cu, 52Fe, 52Mn, 140La, 175Yb, 153Sm, 166Ho, 149Pm, 177Lu, 142Pr, 159Gd, 212Bi, 47Sc, 149Pm, 67Cu, 111Ag, 199Au, 161Tb, and 51Cr as well as salts thereof with physiologically compatible bases or acids. 2. A compound of claim 1, wherein one or more of FG is a carboxy group. 3. A compound of claim 1 wherein R2 is a C1–C20 alkyl or aryl group. 4. A compound of claim 1, wherein R3 is hydrogen, R4 is hydrogen or methyl, and R5 is hydrogen. 5. A compound of claim 1, wherein the two R1 groups together form an alkylene. 6. A compound of claim 5, wherein the two R1 groups together form an ethylene group. 7. An MRI contrast agent comprising a chelate of the compound of claim 1 complexed with a metal ion. 8. The MRI contrast agent of claim 7, wherein the metal ion is selected from the group consisting of Gd(3+), Eu(3+), Dy(3+), La(3+), Yb(3+) and Mn(3+). 9. A compound of claim 8 wherein the metal ion is Gadolinium. 10. A radiotherapeutic or radiodiagnostic agent comprising a chelate of the compound of claim 1 complexed with a radioisotope. 11. A pharmaceutical or diagnostic composition comprising a compound of claim 1 in admixture with a suitable carrier. 12. A compound of formula (IX):wherein the two R1 groups form a straight or cyclic C2–C10 alkylene group or an ortho-disubstituted arylene and R2 is hydrogen, C1–C20 alkyl, C3–C10 cycloalkyl, C4–C20 cycloalkylalkyl, aryl or arylalkyl optionally substituted with functional groups which allow conjugation with a suitable molecule able to interact with physiological systems. 13. A compound of claim 12, wherein the two R1 groups together form an ethylene or propylene group. 14. A compound of claim 3, wherein said alkyl or aryl group is substituted with one or more carboxy or amino groups. 15. A compound of claim 14, wherein said carboxy or amino groups are protected. 16. A compound of claim 14, wherein said alkyl group is a methyl group. | Bracco Imaging S.p.A. | A61B005/055
US07006863 | B2 | 2006-02-28 | Method and apparatus for assessing neural function by sparse stimuli | 2001-03-27 | The present invention relates to a method and an apparatus for simultaneously assessing the functional status of component parts of the nervous system by presenting sparse stimuli to one or more parts of the sensory nervous system. Sparse stimuli consist of temporal sequences of stimulus conditions presented against a baseline null stimulus condition, where the non-null stimulus condition, or conditions, are presented relatively infrequently. The low probability of encountering a stimulus differing from a baseline or null stimulus condition in sparse stimulus sequences insures that gain control mechanisms within the nervous system will increase the neural response magnitude and also bias the measured responses to those neurone populations having such gain controls. The consequently increased response amplitudes ensure more reliably recorded responses than are obtained with non-sparse stimuli. | 1. A method for simultaneously assessing the functional status of component parts of the nervous system of a subject, said method comprising: (a) presenting to one or more parts of the sensory nervous system of the subject stimulus sequences having different temporal modulation sequences of the appropriate stimulus modality for each stimulated part of the sensory nervous system, the stimuli having different sequences for each stimulated part;(b) temporally modulating the stimuli between a null stimulus condition and at least one non-null stimulus condition selected from the group consisting of an increment stimulus condition and a decrement stimulus condition, relative to the null stimulus condition, wherein the probability of encountering the null stimulus condition in the stimulus sequences is higher compared to the probability of encountering at least one non-null stimulus condition, and wherein the temporally modulated stimuli permit estimation of linear and non-linear weighting functions characterising measured responses to each stimulus presented to each part of the nervous system, wherein the at least one non-null condition occurs with an average frequency of between about 0.25 Hz and about 6 Hz;(c) estimating some or all of the coefficients of the linear and non-linear weighting functions for each stimulus sequence from the measured responses to said stimuli, to isolate separate responses from the separately and simultaneously stimulated component parts of the nervous system. | 1. A method for simultaneously assessing the functional status of component parts of the nervous system of a subject, said method comprising: (a) presenting to one or more parts of the sensory nervous system of the subject stimulus sequences having different temporal modulation sequences of the appropriate stimulus modality for each stimulated part of the sensory nervous system, the stimuli having different sequences for each stimulated part;(b) temporally modulating the stimuli between a null stimulus condition and at least one non-null stimulus condition selected from the group consisting of an increment stimulus condition and a decrement stimulus condition, relative to the null stimulus condition, wherein the probability of encountering the null stimulus condition in the stimulus sequences is higher compared to the probability of encountering at least one non-null stimulus condition, and wherein the temporally modulated stimuli permit estimation of linear and non-linear weighting functions characterising measured responses to each stimulus presented to each part of the nervous system, wherein the at least one non-null condition occurs with an average frequency of between about 0.25 Hz and about 6 Hz;(c) estimating some or all of the coefficients of the linear and non-linear weighting functions for each stimulus sequence from the measured responses to said stimuli, to isolate separate responses from the separately and simultaneously stimulated component parts of the nervous system. 2. The method of claim 1, wherein a stimulus sequence comprises a null stimulus condition and two non-null stimulus conditions comprising an increment stimulus condition characterised by an increment of a parameter relative to the null stimulus condition, and a decrement stimulus condition characterised by a decrement of a parameter relative to the null stimulus condition, wherein said non-null stimulus conditions are presented at a lower frequency relative to said null stimulus condition. 3. The method of claim 1, wherein a stimulus sequence comprises a null stimulus condition and a single non-null stimulus condition, wherein said non-null stimulus condition is presented at a lower frequency relative to said null stimulus condition. 4. The method of claim 1, wherein a stimulus sequence comprises a null stimulus condition and a number of non-null stimulus conditions wherein said non-null stimulus conditions are presented at a lower frequency relative to said null stimulus condition. 5. The method of claim 1, wherein the at least one non-null condition occurs with an average frequency of between about 1 and about 6 per second. 6. The method of claim 5, wherein the stimulation is selected from tactile stimuli, auditory stimuli, visual stimuli or a combination thereof. 7. The method of claim 1, wherein the at least one non-null condition comprises stimulation by visual stimuli. 8. The method of claim 7, wherein the visual stimuli are selected from different lumin... | The Australian National University | A61B005/00
US06885885 | B1 | 2005-04-26 | Magnetic resonance imaging method and device | 2000-05-26 | A magnetic resonance imaging apparatus has means for detecting the movement of an object so as to correct the variation of the non-uniformity of static magnetic field because of the movement of the object and thereby to improve the quality of the MR image. When actually measured data is corrected by using a signal for correcting the static magnetic field obtained in a preliminary measurement, the influence of the non-uniformity of static magnetic field is eliminated from the image based on the information from the body movement detecting means by using a correcting signal measured at the same time phase of the periodic movement as that of the actually measured signal. A navigation echo measured during the same repetition time as that of the actually measured signal is used for detection of time phase by detecting means. As a result, the quality of image is not deteriorated by the variation of the non-uniformity of static magnetic field due to the movement of the body of the object, realizing imaging of high quality image. | 1. Magnetic resonance imaging method comprising, (a) step for applying static magnetic field, slice selective gradient magnetic field, gradient magnetic field for reading out a signal to an object placed within a measurement space with a defined pulse sequence, and for acquiring signal showing non-uniformity distribution of said static magnetic field in selected slice of said object, (b) step for applying said static magnetic field, slice selective gradient magnetic field, phase encode gradient magnetic field, gradient magnetic field for reading out said signal to said object placed in said measurement space with said defined pulse sequence, and for acquiring signal showing image information in said selected slice of said object, (c) step for making signal showing real image information of said object by calculating between signal showing image information acquired in said step (b) and signal showing static magnetic field non-uniformity acquired in said step (a), (d) step for reconstructing image of said object from signal showing real image information of said object. | 1. Magnetic resonance imaging method comprising, (a) step for applying static magnetic field, slice selective gradient magnetic field, gradient magnetic field for reading out a signal to an object placed within a measurement space with a defined pulse sequence, and for acquiring signal showing non-uniformity distribution of said static magnetic field in selected slice of said object, (b) step for applying said static magnetic field, slice selective gradient magnetic field, phase encode gradient magnetic field, gradient magnetic field for reading out said signal to said object placed in said measurement space with said defined pulse sequence, and for acquiring signal showing image information in said selected slice of said object, (c) step for making signal showing real image information of said object by calculating between signal showing image information acquired in said step (b) and signal showing static magnetic field non-uniformity acquired in said step (a), (d) step for reconstructing image of said object from signal showing real image information of said object. 2. Magnetic resonance imaging method comprising, (e) step for applying static magnetic field, slice selective gradient magnetic field, gradient magnetic field for reading out said signal to an object placed in measurement space with defined pulse, and for acquiring signal showing static magnetic field non-uniformity in said selected slice of said object, (f) step for detecting motion of said object and for setting plural number of time phase in one period of a motion of said object, and for acquiring MR signal that shows said motion of said object in these every time phase, (g) step for applying static magnetic field, slice selective gradient magnetic field, phase encode gradient magnetic field, gradient magnetic field for reading out said signal with the defined pulse sequence to said object placed in said measurement space in each time phase of said one period of said motion of said object, and for acquiring signal that shows image information in said selected slice of said object, (h) step for removing static magnetic field non-uniformity signal including static magnetic field non-uniformity from said signal showing image information of said object being acquired in said step (g) in each time phase by using said signal showing static magnetic field acquired in said step (e) and said signal showing said motion of object acquired in said step (f), (i) step for reconstructing image by using said signal being removed static magnetic field non-uniformity in said step (h). 3. Magnetic resonance imaging method comprising, (j) step for detecting motion of an object and for setting several time phase in one period of said motion of said object, (k) step for applying static magnetic field, slice selective gradient magnetic field, gradient magnetic field for reading out said signal to said object placed in said measurement space with said defined pulse sequence per each time phase, and... | Hitachi Medical Corporation | A61B005/05
US06892086 | B2 | 2005-05-10 | Medical electrode for preventing the passage of harmful current to a patient | 2001-10-17 | A medical electrode designed to prevent the passage of harmful electric current to a patient, thereby preventing tissue damage and electrocution. In the preferred embodiment, the inventive medical electrode is comprised of a proximal end, a distal end, a conductive lead connecting the proximal and distal ends, and a fuse located upon the medical electrode for preventing the passage of harmful electric current to a patient. For additional protection from induction current, the preferred location of the fuse is aft of the conductive lead element of the electrode. While a fuse is used in the preferred embodiment, the fuse could be replaced with a diode or a circuit breaker. | 1. A medical electrode for preventing the passage of harmful current to a patient, the electrode comprising: a distal end for connecting to a medical device; a proximal end including a portion formed in the shaped shape of a contact lens for connecting to a patient; a conductive lead connected between the distal and proximal ends; and a current stoppage means connected to the electrode at a location for preventing the passage of harmful current to the proximal end coupled to the proximal end of the electrode aft of the conductive lead. | 1. A medical electrode for preventing the passage of harmful current to a patient, the electrode comprising: a distal end for connecting to a medical device; a proximal end including a portion formed in the shaped shape of a contact lens for connecting to a patient; a conductive lead connected between the distal and proximal ends; and a current stoppage means connected to the electrode at a location for preventing the passage of harmful current to the proximal end coupled to the proximal end of the electrode aft of the conductive lead. 2. A medical electrode for preventing the passage of harmful current to a patient, the electrode comprising: a distal end for connecting to a medical device; a proximal end for connecting to a patient; a conductive lead connected between the distal and proximal ends; and a current stoppage means connected aft of the conductive lead and immediately adjacent to the proximal end of the electrode, wherein the current stoppage means is a circuit breaker. 3. A medical side-snap electrode for preventing the passage of harmful current to a patient, the electrode comprising: a distal end for connecting to a medical device; a proximal end for connecting to a patient; a conductive lead connected between the distal and proximal ends; and a current stoppage means selected from the group consisting of a fuse or a circuit breaker, the current stoppage means coupled to the proximal end of the side-snap electrode aft of the conductive lead. 4. A medical electrode having separable members for preventing the passage of harmful current to a patient, the electrode comprising: a distal end for connecting to a medical device; a separable structure including top and bottom separable members, the separable structure being located at a proximal end of said electrode, the separable structure for connecting to a patient, the top member being positioned atop the bottom member, the bottom member comprising a pad for contacting a patient's skin, the pad located adjacent a connector portion, and a current stoppage means located between the pad and the connector portion. 5. The bottom member as recited in claim 4, wherein the current stoppage means is selected from the group consisting of a fuse, a diode or a circuit breaker. 6. The bottom member as recited in claim 4, wherein the current stoppage means is an indicator fuse. 7. The bottom member as recited in claim 4, wherein the connector portion further comprises a male connector. | - | A61B005/04
US06925318 | B2 | 2005-08-02 | Medical probe with variable tip length and shape | 2002-06-14 | A tip assembly that can be varied in length and/or shape is coupled to an elongated sheath to form a medical probe, such as a catheter or surgical probe. For example, the medical probe can be an electrophysiology (EP) catheter that that can be used to map the electrical activity of target tissue and ablate the target tissue when medically indicated. In one embodiment, the tip assembly has a tip and a sleeve configured such that the tip can be longitudinally extended and retracted in response to longitudinal forces applied by an actuator. In another embodiment, the tip extends and retracts, on threads, relative to the sleeve in response to rotational forces applied by an actuator. In still another embodiment, the tip assembly has a distal tip, a proximal section and an expandable-compressible assembly between the tip and proximal section. The expandable-compressible assembly is configured to be longitudinally compressed and expanded, expanding radially outward when compressed, and compressing radially inward when longitudinally expanded, thereby altering the tip assembly shape. | 1. A medical probe, comprising: an elongated sheath; a tip assembly mounted to the distal end of the sheath, the tip assembly having a sleeve and a tip, the tip configured to be alternately displaced distally and proximally relative to the sleeve, wherein the tip assembly and the tip have smooth exterior surfaces; an actuator operatively coupled to displace the tip relative to the sleeve, the actuator being configured to displace the tip out of the sleeve and retract the tip into the sleeve; and a mechanical stop to limit distal movement of the tip relative to the sleeve, wherein the mechanical stop comprises a flange disposed on the outside of the tip, and a shoulder disposed on the inside of the sleeve. | 1. A medical probe, comprising: an elongated sheath; a tip assembly mounted to the distal end of the sheath, the tip assembly having a sleeve and a tip, the tip configured to be alternately displaced distally and proximally relative to the sleeve, wherein the tip assembly and the tip have smooth exterior surfaces; an actuator operatively coupled to displace the tip relative to the sleeve, the actuator being configured to displace the tip out of the sleeve and retract the tip into the sleeve; and a mechanical stop to limit distal movement of the tip relative to the sleeve, wherein the mechanical stop comprises a flange disposed on the outside of the tip, and a shoulder disposed on the inside of the sleeve. 2. The probe of claim 1, wherein the tip is displaced relative to the sleeve in the presence of a longitudinal mechanical force. 3. The probe of claim 1, wherein the actuator is a control wire operatively coupled to the tip to displace the tip relative to the sleeve. 4. The probe of claim 1, further comprising a radio frequency/mapping wire connected to the probe tip assembly. 5. The probe of claim 1, wherein the tip is a generally round tip and the sleeve is a generally cylindrical shank. 6. The probe of claim 1, wherein the elongated sheath is a catheter sheath. 7. The probe of claim 1, wherein the tip assembly is composed of an electrically conductive material. | SciMed Life Systems, Inc. | A61B005/042 A61B018/14
US06856829 | B2 | 2005-02-15 | Method for detecting physiological condition of sleeping patient based on analysis of pulse waves | 2001-08-24 | A physiological condition detecting method detects pulse waves from the body of a patient, and creates the envelope of the pulse waves by connecting every peak of the pulse waves. It determines that the patient is in non-REM sleep, if the envelope fluctuates regularly. It determines that the patient is in REM sleep, if the envelope fluctuates irregularly. Further the method calculates and normalizes the amplitude and period of the envelope. It determines whether the patient has obstructive sleep apnea syndrome based on the normalized amplitude. It determines whether the patient has central sleep apnea syndrome based on the normalized period. It determines that the patient has mixed sleep apnea syndrome based on the normalized amplitude and the normalized period. | 1. A method for detecting a physiological condition of a patient comprising the steps of: emitting light into the body of the patient; detecting pulse waves from the body of the patient from light reflected by the body of the patient; creating an envelope of the pulse waves; analyzing the envelope; and determining whether the patient is in a specific physiological condition based on a result of the analysis. | 1. A method for detecting a physiological condition of a patient comprising the steps of: emitting light into the body of the patient; detecting pulse waves from the body of the patient from light reflected by the body of the patient; creating an envelope of the pulse waves; analyzing the envelope; and determining whether the patient is in a specific physiological condition based on a result of the analysis. 2. A method for detecting a physiological condition of a patient as in claim 1, wherein the step of analyzing includes the step of: calculating for every regular time interval a statistical index which represents characteristics of the envelope over the time interval, and wherein the step of determining includes using the statistical index. 3. A method for detecting a physiological condition of a patient comprising: detecting pulse waves from the body of the patient; calculating a pulse height of each of the pulse waves; analyzing the pulse height; and determining whether the patient is in a specific physiological condition based on a result of the analysis, wherein the step of analyzing includes the step of: calculating for every regular time interval a statistical index which represents characteristics of the pulse height over the time interval, and wherein the step of determining includes using the statistical index. 4. A method for detecting a physiological condition of a patient comprising: detecting pulse waves from the body of the patient; calculating a pulse height of each of the pulse waves; analyzing the pulse height; and determining whether the patient is in a specific physiological condition based on a result of the analysis, wherein the step of analyzing includes the steps of: creating a baseline of the pulse waves by connecting middle points of the pulse height; and calculating for every regular time interval a statistical index which represents characteristics of the baseline over the time interval, and wherein the step of determining includes using the statistical index. 5. A method for detecting a sleep condition of a patient comprising the steps of: emitting light into the body of the patient; detecting pulse waves from the sleeping body of the patient from light reflected by the body of the patient; creating an envelope of the pulse waves by connecting one of tops and bottoms of the pulse waves; determining that the patient is in non-REM sleep if the envelope fluctuates regularly; and determining that the patient is in REM sleep if the envelope fluctuates irregularly. 6. A method for detecting a sleep condition of a patient as in claim 5, further comprising the step of: calculating for every regular time interval a statistical index which represents characteristics of the envelope over the time interval, wherein the steps of determining include using the statistical index. 7. A method for detecting a sleep condition of a patient as in claim 6, wherein the statistical index includes at least one of an average, a ... | Denso Corporation Junichiro Hayano Toshiaki Shiomi | A61B005/02 A61B006/00 A61B005/00
US06904801 | B1 | 2005-06-14 | Functional capacity evaluation apparatus and method | 2002-04-29 | An apparatus and method is provided for testing the ability of an individual to lift objects under various conditions. During standard lifting tests, the average acceleration and velocity of each lift, as well as the distribution of force between the hands and feet of the patient, are electronically measured and recorded. These objective factors can then be used to determine whether the patient is exerting maximal effort, and to assess a patient's condition and progress during a rehabilitation program. | 1. An apparatus for testing the lifting ability of an individual comprising: a detached object to be lifted in free space a fixed predetermined distance A first sensor to determine when the lift has begun; A second sensor to determine when the lift has been completed; and force sensors for determining the force applied by each of the individual's feet during the lift; Means for calculating the velocity and acceleration of the lift, said calculating means connected to said first sensor and said second sensor; and, means for recording the average force applied by each of the individual's feet during the lift, said recording means connected to said force sensors. | 1. An apparatus for testing the lifting ability of an individual comprising: a detached object to be lifted in free space a fixed predetermined distance A first sensor to determine when the lift has begun; A second sensor to determine when the lift has been completed; and force sensors for determining the force applied by each of the individual's feet during the lift; Means for calculating the velocity and acceleration of the lift, said calculating means connected to said first sensor and said second sensor; and, means for recording the average force applied by each of the individual's feet during the lift, said recording means connected to said force sensors. 2. The apparatus as described in claim 1 wherein said first sensor is a transmitted beam photoelectric sensor. 3. The apparatus as described in claim 1 wherein said second sensor is a transmitted beam photoelectric sensor. 4. The apparatus of claim 1 further comprising a rack and at least one shelf selectively placed at predetermined locations on said rack wherein said fixed distance is determined by the selective placement of said shelf at height on said rack, and said first sensor is attached to the base of said rack and said second sensor is attached to said shelf, said object being freely movable along any path between said base and shelf. 5. The apparatus of claim 1 further comprising a rack and at least one shelf selectively placed at predetermined locations on said rack wherein said fixed distance is determined by the selective placement of a lower and a higher shelf on said rack, and said first sensor is attached to said lower shelf, and said second sensor is attached to said higher shelf, said object being freely movable along any path between said lower and higher shelf. 6. The apparatus as described in claim 1 wherein said calculating means comprise a microprocessor control system and operator interface. 7. The apparatus described in claim 1, further comprising force sensors for determining the force applied by each of the individual's feet during the lift and means for recording the average force applied by each of the individual's feet during the duration of the lift, said recording means connected to said sensors. 8. The apparatus as described in claim 1 wherein said calculating means comprise a microprocessor control system and operator interface. 9. The apparatus as described in claim 7 further comprising a data retention means for recording the velocity, acceleration, and average force of multiple lifts. 10. The apparatus as described in claim 1, further comprising force sensors for determining the force applied by each of the individual's hands during the lift, and means for recording the average force applied by each of the individual's hands during the duration of the lift, said recording means connected to said force sensors. 11. The apparatus as described in claim 10 wherein said recording means comprise a microprocessor control system and operator interface... | - | A61B005/22